## All Reports
report_number | date_received | date_of_event | event_type | report_source_code | brand_name | generic_name | manufacturer_d_name | model_number | catalog_number | device_class | event_description | manufacturer_narrative | adverse_event_flag | product_problem_flag | device_operator | manufacturer_g1_name | year_month
1038671-2021-00691 | 2021-12-14 | 2021-08-01 | I, n, j, u, r, y | Manufacturer report | LOGIC FEMORAL PS CEM LEFT SZ 4 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC FEMORAL PS CEM LEFT SZ 4 | 02-010-01-0240 |  | AS REPORTED BY THE EXACTECH KNEE REPLACEMENT STUDY, APPROXIMATELY FIVE-YEAR POST LTKA, THE (B)(6) MALE PATIENT EXPERIENCED POLYETHYLENE WEAR WITH EVIDENCE OF OSTEOLYSIS. SWELLING IN LEFT KNEE, SURGEON REPORTS THIS IS DUE TO THE POLYETHYLENE WEAR WHICH IS VISIBLE ON THE X-RAY WITH EVIDENCE OF OSTEOLYSIS. THERE IS NO REPORT OF FOLLOW-UP OR REVISION AT THIS TIME. THE EVENT IS POSSIBLY RELATED TO THE DEVICE BUT NOT RELATED TO THE PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | (H3) THE PROSTHESIS WEAR AND OSTEOLYSIS REPORTED MAY HAVE BEEN THE RESULT OF MALALIGNMENT BETWEEN THE FEMORAL AND TIBIAL COMPONENTS, THIRD BODY WEAR, PATIENT-RELATED CONDITIONS, OR ANY COMBINATION OF THESE POSSIBILITIES. HOWEVER, THIS CANNOT BE CONFORMED AS THE DEVICES WERE NOT AVAILABLE FOR EVALUATION AND X-RAY IMAGES ARE NOT PROVIDED. THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE REPORTED EVENT OF "PROSTHESIS WEAR¿ IS ASSOCIATED WITH MATERIAL DAMAGE TO A SURFACE, USUALLY INVOLVING PROGRESSIVE LOSS OR DISPLACEMENT OF MATERIAL, DUE TO RELATIVE MOTION BETWEEN THAT SURFACE AND A CONTACTING SUBSTANCE OR SUBSTANCES. | Y | Y | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00692 | 2021-12-14 | 2021-11-18 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | HUMERAL STEM PRIMARY, PRESS FIT 15MM | EXACTECH, INC. | 320-02-42 | 300-01-15 |  | AS REPORTED, THIS (B)(6) Y/O FEMALE PATIENT WAS PREVIOUSLY REVISED ON (B)(6) 2020 FOR WITH A LOOSE HUMERAL COMPONENT FROM A SUSPECTED INFECTION. CURRENTLY, PATIENT'S STEM IS LOOSE. REMOVED HUMERAL COMPONENTS AS WELL AS GLENOSPHERE, LOCKING SCREW AND COMPRESSION SCREWS. REPLACED ALL SHOULDER COMPONENTS. THE PATIENT HAS A BMI OF 50+. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT RETURN DUE TO FACILITY POLICY. | DEVICE EVALUATED BY MANUFACTURER? PENDING EVALUATION. CONCOMITANT MEDICAL DEVICE(S): 300-01-15, (B)(4) - EQUINOXE, HUMERAL STEM PRIMARY, PRESS FIT 15MM; 320-10-00, (B)(4) - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0; 320-15-05, (B)(4) - EQ REV LOCKING SCREW; 320-20-00, (B)(4) - EQ REVERSE TORQUE DEFINING SCREW KIT; 320-42-10, (B)(4) - EQUINOXE REVERSE 42MM HUMERAL CONST LINER +0; 320-20-38, (B)(4) - EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 38MM; 320-20-38, (B)(4) - EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 38MM. | Y | N | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00693 | 2021-12-14 | 2021-11-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRULIANT PIN PULLER SYRINGE STYLE | ORTHOPEDIC MANUAL SURGICAL INSTRUMENT | EXACTECH, INC. | TRULIANT PIN PULLER SYRINGE STYLE | 02-029-90-4100 |  | AS REPORTED, DURING SURGICAL USE, THIS INSTRUMENT BROKE WHEN THE SURGEON WAS REMOVING THE HEADED PINS FROM THE TIBIA. THERE WAS NO PARTS OR PIECES OF THE DEVICE FELL INTO THE PATIENT WOUND SITE. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICE WILL BE RETURNED. | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION RECEIVED THE FOLLOWING SECTIONS G3, G6, H2, H3 AND H6 HAVE BEEN UPDATED ACCORDINGLY. (H3) THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF THE INSTRUMENT BEING MODIFIED WITHOUT EXACTECH¿S PERMISSION. THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE REPORTED EVENT OF "BROKEN / NON-FUNCTIONAL¿ IS ASSOCIATED WITH THE DEVICE WHICH HAS BEEN USED FOR AN UNAPPROVED INDICATION OR FOR AN UNAPPROVED INTENDED USE. SECTION D8: DEVICE AVAILABLE FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00695 | 2021-12-14 | 2021-11-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EQUINOXE | MDLR PRM RMR MD PLT TIP | EXACTECH, INC. |  | 315-25-13 |  | AS REPORTED, DURING USE, THE INSTRUMENT BROKE; THE PILOT NOSE SHEARED OFF. SURGEON CONFIDENT OF CLEAN BREAK AND BOTH PIECES RETAINED, NO PATIENT IMPACT. THE SURGEON CHECKED FOR PARTS AND PIECES. NO PATIENT INFORMATION REPORTED. | SECTION H10: (D4) LOT NUMBER: 53636003 (H3) BASED ON HISTORICAL COMPLAINT INVESTIGATIONS AND THE RETURNED DEVICE, THE BROKEN REAMER REPORTED IN WAS LIKELY THE RESULT OF THE REAMER EXPERIENCING HIGH LOADS POSSIBLY DURING OFF-AXIS REAMING AND/OR COMING IN CONTACT WITH HARD BONE, WHICH LED TO BRITTLE FRACTURE. | N | Y | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00694 | 2021-12-14 | 2021-11-22 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | EQ REV GLENOID PLATE | EXACTECH, INC. | 320-15-01 | 320-15-01 |  | APPROXIMATELY 16 MOS POSTOP THE INITIAL RIGHT TSA, THE (B)(6) FEMALE PATIENT HAD GLENOID NOTCHING. WHEN INVESTIGATING, THE SURGEON FOUND THE GLENOID BASEPLATE WAS LOOSE. PATIENT WAS REVISED TO A HEMIARTHROPLASTY AND IS AWAITING BONE GRAFT TAKE-UP OF THE GLENOID. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. | DEVICE EVALUATED BY MFR: PENDING EVALUATION. CONCOMITANT DEVICE(S): 300-01-15, 6332321 - EQUINOXE, HUMERAL STEM PRIMARY, PRESS FIT 15MM. 320-20-00, 6519700 - EQ REVERSE TORQUE DEFINING SCREW KIT. 320-15-05, 6490907 - EQ REV LOCKING SCREW. 320-10-00, 6474248 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0. 320-01-36, 6310869 - 36MM GLENOSPHERE. | Y | N | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00696 | 2021-12-14 | 2021-11-23 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | CAGE GLENOID SMALL, ALPHA | EXACTECH, INC. | 314-13-02 | 314-13-02 |  | APPROXIMATELY 3.5 YR POSTOP THE INITIAL RIGHT TSA, THIS (B)(6) FEMALE PATIENT EXPERIENCED INSTABILITY AND OPTED TO CONVERT ANATOMIC TO REVERSE SHOULDER. NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. SUPERIOR CUFF WAS TORN AND SUBSCAP VERY THIN DURING REVISION. CONVERSION SURGERY WENT WELL, AND PATIENT IS EXPECTED TO DO WELL. PATIENT LEFT OR STABLE. EXPLANTS WILL NOT BE RETURNED PER HOSPITAL POLICY. | PENDING EVALUATION, CONCOMITANT DEVICE(S): 300-51-45, (B)(4) - 4.5 SHORT FA REP PLT KIT 310-01-44, (B)(4) - EQUINOXE, HUMERAL HEAD SHORT, 44MM (ALPHA). | Y | N | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00697 | 2021-12-14 | 2018-04-11 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 46MM GLENOSPHERE | EXACTECH, INC. | 320-01-46 | 320-01-46 |  | AS REPORTED, APPROXIMATELY 5 MONTHS POSTOP THE INITIAL RIGHT TSA, THIS(B)(6) MALE PRESENTED WITH BROKEN SCREWS IN THE RIGHT SHOULDER AND WAS REVISED APPROXIMATELY 3 MONTHS LATER. AT THAT TIME, THE OUTCOME WAS REPORTED AS RESOLVED. THE CASE REPORT FORM INDICATES THIS EVENT IS DEFINITELY RELATED TO DEVICES AND POSSIBLY RELATED TO PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. DEVICES WERE NOT RETURNED PER STUDY PROTOCOL. | PENDING EVALUATION. CONCOMITANT DEVICE(S): 300-01-13 - EQUINOXE, HUMERAL STEM PRIMARY, PRESS FIT 13MM, 320-46-00 - EQUINOXE REVERSE 46MM HUMERAL LINER +0, 320-10-00 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0, AND 320-15-04 - RS GLENOID PLATE R POST AUG, 8 DEG, RIGHT. | Y | N | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00699 | 2021-12-14 | 2021-11-19 | I, n, j, u, r, y | Manufacturer report | CR TIBIAL INSERT SZ 3, 9MM, SLOPE + | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED, | EXACTECH, INC. | CR TIBIAL INSERT SZ 3, 9MM, SLOPE + | 200-63-09 |  | AS REPORTED, SALES REP INFORMED THAT THERE IS A PLAN THAT THE RECALL INSERT WILL RETRIEVE AT REVISION SURGERY AROUND WEEK AFTER NEXT DUE TO POSTERIOR WEAR OF THE INSERT. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00698 | 2021-12-14 | 2021-11-19 | I, n, j, u, r, y | Manufacturer report | PS TIBIAL INSERTS SZ 4, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | PS TIBIAL INSERTS SZ 4, 9MM | 204-24-09 |  | AS REPORTED, SALES REP IS CURRENTLY IN A REVISION KNEE CASE FROM (B)(6) 2014 AND CALLED AND REQUESTED THE EXPIRATION DATE FOR ITEM# 204-24-09 SN# (B)(4) SINCE THE POLY LOOKS BAD. THE SURGEON DID A POLY EXCHANGE OF A REVISION KNEE THAT WAS DONE IN 2014. HE SAID THE POLY WAS SEVERELY DAMAGED. A NEW POLY WAS IMPLANTED, AND X-RAY TAKEN. THE RANGE OF MOTION (ROM), ALIGNMENT, AND STABILITY WAS GOOD PER THE SURGEON. PATIENT DID WELL AND WAS LAST KNOWN TO BE IS STABLE CONDITION. NO OTHER INFORMATION AVAILABLE. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00700 | 2021-12-14 | 2021-11-19 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 2, 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 2, 11MM | 02-012-44-2011 |  | AS REPORTED, THIS (B)(6) PATIENT PRESENTED AFTER APPROXIMATELY 10 YEARS OF WEAR WITH A MISALIGNED LEFT TKR REPLACEMENT. THE PATIENT PRESENTS IN A MORBIDLY (B)(6) STATE. IT WAS REALIZED THAT THE PATIENT HAD SUFFERED A FRACTURE OF THE PROXIMAL FEMUR AND BONE REMODELING HAD OCCURRED AND THE KNEE FALLING INTO VALGUS. SIZE 2 FEMORAL COMPONENT AND POLY REPLACED DUE TO FRACTURE. REPLACED WITH A CC LOGIC FEMUR AND PSC. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICE WILL BE RETURNED. | CONCOMITANT MEDICAL PRODUCTS: LOGIC FEMORAL PS CEM LEFT SZ 2 (CAT# 02-010-01-0220 / SERIAL# (B)(4)). LOGIC FIT TIBIAL TRAY CEMENTED (CAT# 02-012-45-2020 / SERIAL# (B)(4)). OPTETRAK PATELLA CEMENTED (CAT# 200-02-29 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00703 | 2021-12-14 | 2021-11-22 | I, n, j, u, r, y | Manufacturer report | LOGIC CR TIB INSERT SLOPE+, SZ 4, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC CR TIB INSERT SLOPE+, SZ 4, 9MM | 02-012-48-4009 |  | THE PATIENT EXPERIENCED PAIN AND IMPAIRED MOVEMENT, AND ALSO UNDERWENT A BONE GRAFT DURING THE REVISION SURGERY. | THIS FOLLOW-UP REPORT IS BEING SUBMITTED TO RELAY ADDITIONAL AND CORRECTED INFORMATION. THE FOLLOWING SECTIONS WERE UPDATED: H6 COMPONENT CODE. THE FOLLOWING SECTIONS WERE CORRECTED: H6 FINDINGS (4243 AND 3211 REMOVED), CONCLUSIONS. | Y | Y | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00701 | 2021-12-14 | 2021-11-22 | I, n, j, u, r, y | Manufacturer report | LOGIC CR TIB INSERT STD, SZ 4, 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC CR TIB INSERT STD, SZ 4, 11MM | 02-012-47-4011 |  | AS REPORTED, APPROXIMATELY 5 YRS POSTOP THE INITIAL L TKA THIS (B)(6) MALE PATIENT PRESENTING WITH COMPLAINING OF SWELLING OF HIS LEFT KNEE FOR THE PAST 2 YEARS FOLLOWED BY INSTABILITY AND WAS SCHEDULED FOR A REVISION DUE TO INCREASE WEAR ON LEFT TIBIAL SURFACE. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES ARE EXPECTED TO BE RETURNED FOR EVALUATION. | CONCOMITANT MEDICAL PRODUCTS: LOGIC CR FEMORAL CEM, LEFT, SZ 4 (CAT# 02-010-03-0240 / SERIAL# (B)(4)). LGC TIBIAL FIT TRAY CEM SZ 4F / 5T (CAT# 02-012-45-4050 / SERIAL# (B)(4)). THREE PEG PATELLA 38MM (CAT# 200-02-38 / CAT# 4243255). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00702 | 2021-12-14 | 2021-11-22 | I, n, j, u, r, y | Manufacturer report | NV GXL LNR, LIPPED, 36MM ID, GROUP 3 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LNR, LIPPED, 36MM ID, GROUP 3 CUPS | 132-36-53 |  | AS REPORTED, PATIENT WAS REVISED FOR SEVERE OSTEOLYSIS, POLY WEAR, AND NON-DISPLACED FRACTURE OF THE POSTERIOR COLUMN. DR. (B)(6) COMBINED ON THE SURGERY TO DO PELVIC PLATING. DR. (B)(6) DID BRING ME GRAFTING AND CEMENTING OF THE DEFECT AREAS ON BONE LOSS AND DID A HEAD AND LINER EXCHANGE TO A 40 LIPPED LINER AND 40 PLUS 7 HEAD. THE HIP WAS STABLE WITH GOOD RANGE OF MOTION (ROM). PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION. NO FURTHER INFORMATION AVAILABLE BUT WILL SEND PICTURES SHORTLY. | CONCOMITANT MEDICAL PRODUCTS: NV CROWN CUP CLSTR HOLE 56MM GROUP 3 (CAT# 180-01-56 / SERIAL# (B)(4)). BONE SCREW 6.5MM DIA X 15MM LONG (CAT# 120-65-15 / SERIAL# (B)(4)). P-SERIES PF PLASMA H/O COLLARLESS SZ 5 12/14 (CAT# 118-41-05 / SERIAL# (B)(4)). BIOLOX DELTA FEMORAL HEAD 36MM OD, +0MM (CAT# 170-36-00 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00704 | 2021-12-14 | 2021-11-23 | I, n, j, u, r, y | Manufacturer report | NV GXL LNR, +5LAT, 32MM G1-48/50MM CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LNR, +5LAT, 32MM G1-48/50MM CUPS | 136-32-51 |  | AS REPORTED, APPROXIMATELY 2 YRS POSTOP THE INITIAL LTHA, THIS (B)(6) FEMALE PATIENT PRESENTED WITH LEFT HIP PAIN AND UPON REEXAMINATION PATIENT WAS FOUND WITH SEVERE SUPERIOR POLY WEAR. EXACTECH CUP AND LINER WAS REMOVED AND DEPUY CUP AND LINER WAS INSERTED. A NEW EXACTECH HEAD BALL WAS PLACED ON EXISTING ALTEON STEM. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICES WILL NOT BE RETURNED FOR EVALUATION DUE TO HOSPITAL KEPT DEVICES. | CONCOMITANT MEDICAL PRODUCTS: INTEGRIP CC, CLUSTER 50MM, G1 (CAT# 186-01-50 / (B)(4)). BIOLOX DELTA FEMORAL HEAD 32MM OD, -3.5MM (CAT# 170-32-93 / (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00710 | 2021-12-15 | 2018-12-11 | I, n, j, u, r, y | Manufacturer report | EXACTECH | LINER DRIVER HEAD 32MM | EXACTECH, INC. |  | 141-01-32 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGICAL CASE THE SURGEON EXPERIENCED AN ISSUE WITH THE LINER DRIVER HEAD. THE 32MM LINER IMPACTOR BROKE. THIS DID NOT CAUSE ANY DELAY TO SURGERY AND THE PATIENT WAS NOT ADVERSELY AFFECTED. ALL PIECES WERE ACCOUNTED FOR AND THE WOUND SITE WAS WASHED OUT. THE PATIENT WAS STABLE AS THEY LEFT THE OPERATING ROOM. INACTIVE AGENCY WITH NO HOSPITAL/OTHER CONTACT INFORMATION FOR THE EVENT. NO ADDITIONAL INFORMATION. | (B)(4). THE DEVICE WAS RETURNED IN TWO PIECES. ANALYSIS COMPLETED ON (B)(6) 2019 BY (B)(4). VISUAL EVALUATION CONDITIONS/FINDINGS: THE FRACTURED SURFACES APPEAR CONSISTENT WITH CRACK INITIATION THAT LIKELY INITIATED OVER TIME FROM THE CENTER OF THE DEVICE, PROPAGATION, AND ULTIMATE BRITTLE FRACTURE OF THE DRIVER HEAD. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE AND CONSISTENT WITH YEARS OF USE. DESIGN RELATED ISSUES: THE DESIGN OF THE LINER DRIVER HEAD HAS BEEN IN THE FIELD SINCE 2005. EXACTECH IS AWARE OF 2 OTHER SIMILAR COMPLAINT REPORTS INVOLVING BROKEN LINER DRIVER HEADS SINCE 2008. BECAUSE THE LINER DRIVER HEAD IS USED IN TOTAL HIP SURGERIES, SALES DATA FOR ACETABULAR LINERS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 30 UNITS, WHICH HAS BEEN IN THE FIELD SINCE 2005. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS LINER DRIVER HEAD, (B)(4)-RMR-INSTRUMENTS REV F, WAS REVIEWED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS " VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE BROKEN INSTRUMENT REPORTED IN EXPERIENCE (B)(4) WAS LIKELY THE RESULT OF A CRACK INITIATION, THAT INITIATED OVER TIME FROM THE CENTER OF THE DEVICE WHERE THE TIP OF THE LINER DRIVER HANDLE MEETS THE LINER DRIVER HEAD, PROPAGATION, AND ULTIMATE BRITTLE FRACTURE OF THE DRIVER HEAD. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. ¿ CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. ALL PIECES WERE ACCOUNTED FOR FROM THE BROKEN LINER DRIVER HEAD AND THE WOUND SITE WAS WASHED OUT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF A CRACK INITIATION, THAT INITIATED OVER TIME FROM THE CENTER OF THE DEVICE WHERE THE TIP OF THE LINER DRIVER HANDLE MEETS THE LINER DRIVER HEAD, PROPAGATION, AND ULTIMATE BRITTLE FRACTURE OF THE DRIVER HEAD. | Y | N | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00709 | 2021-12-15 | 2021-11-24 | I, n, j, u, r, y | Manufacturer report | THREE PEG PATELLA 32MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | THREE PEG PATELLA 32MM | 200-02-32 |  | AS REPORTED, THIS (B)(6) Y/O FEMALE PATIENT PRESENTED WITH ANTERIOR LEFT KNEE PAIN, SO PATELLA WAS REMOVED, ANOTHER RESECTION MADE AND A LARGER 35 PATELLA WAS CEMENTED. A POPLITEAL RELEASE WAS CONDUCTED AND A 15 CRC POLY REPLACED EXISTING 13 CRC. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO DEVICES DISPOSED BY THE HOSPITAL. | CONCOMITANT MEDICAL PRODUCTS: LOGIC CR FEMORAL CEM, LEFT SZ 1.5 (CAT#: 02-010-03-0215/SERIAL#: (B)(4)); LGC TIBIAL FIT TRAY CEM SZ 1.5F/1.5T (CAT#: 02-012-45-1515/SERIAL#: (B)(4)); LOGIC TIB INSERT IMPL CRC, SZ 1.5, 13MM (CAT#: 02-012-51-1513/SERIAL#: (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00708 | 2021-12-15 | 2021-11-23 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 42MM HUMERAL CONST LINER +2.5 | EXACTECH, INC. | 320-01-42 | 320-42-13 |  | AS REPORTED, CASE ORIGINALLY DONE BY SURGEON IN (B)(6). APPROXIMATELY 3.5 YRS POSTOP THE RIGHT TSA THIS (B)(6) Y/O MALE PATIENT PRESENTED WITH PAIN AND LIMITED MOTION. X-RAYS SHOW MASSIVE INFERIOR NOTCHING, LOOSE STEM, METALLOSIS AND MALPOSITIONING OF GLENOID. HUMERAL LINER WAS MISSING. A REVISION OF THE RIGHT SHOULDER WAS COMPLETED. THE GLENOSPHERE AND LOCKING SCREW WERE CHANGED WITH ALL HUMERAL COMPONENTS REMOVED AND REPLACED WITH LIMA REVISION SYSTEM. THERE WAS NO BREAKAGE TO DEVICE FOUND. NO OTHER PATIENT INFORMATION OR MEDICAL HISTORY REPORTED. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. | (H3) THE REVISION REPORTED WAS LIKELY THE RESULT OF HUMERAL STEM LOOSENING AND HUMERAL LINER DISASSOCIATION, WHICH LED TO THE REPORTED PAIN AND LIMITED RANGE OF MOTION. THE HUMERAL STEM LOOSENING WAS LIKELY A RESULT OF AN INSUFFICIENT BOND BETWEEN THE HUMERAL STEM AND THE BONE, WHICH LED TO ASEPTIC (NON-INFECTED) HUMERAL LOOSENING. THE HUMERAL LINER DISASSOCIATION MAY HAVE BEEN THE RESULT OF EITHER INCOMPLETE SEATING OF THE LINER DURING IMPLANTATION, SCAPULAR NOTCHING, PATIENT-RELATED CONDITIONS, OR ANY COMBINATION OF THESE POSSIBILITIES. THE METAL ON METAL CONTACT BETWEEN THE HUMERAL TRAY AND GLENOSPHERE LIKELY LED TO THE REPORTED METALOSIS. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICES WERE NOT AVAILABLE FOR EVALUATION. SECTION H11: *THE FOLLOWING SECTIONS HAVE CORRECTED INFORMATION: (D2B) COMMON DEVICE NAME: REVERSE 42MM HUMERAL CONST LINER +2.5 (D4) CATALOG NUMBER: 320-42-13, SERIAL NUMBER: (B)(6), EXPIRATION DATE: 26-JAN-2022, UNIQUE IDENTIFIER (UDI) #: (B)(4). (D11) CONCOMITANT DEVICE(S): 320-10-00, 5460030 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0, 300-01-13, 5339173 - EQUINOXE, HUMERAL STEM PRIMARY, PRESS FIT 13MM, 320-42-00, 4506997 - EQUINOXE REVERSE 42MM HUMERAL LINER +0, 320-15-04, 5442223 - RS GLENOID PLATE R POST AUG, 8 DEG, RIGHT, 320-15-05, 5389263 - EQ REV LOCKING SCREW, 320-20-00, 5179533 - EQ REVERSE TORQUE DEFINING SCREW KIT, 320-01-42, 5309395 - EQUINOXE REVERSE 42MM GLENOSPHERE, (H4) DEVICE MANUFACTURE DATE: 30-JAN-2017. | Y | N | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00706 | 2021-12-15 | 2021-10-14 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | 2MM K-WIRE INSERT | EXACTECH, INC. |  | SSI-300-005880 |  | IT WAS REPORTED THAT DURING A SHOULDER SURGERY, 2 K-WIRE GUIDE HOLD SLEEVES FRACTURED. ASKED, BUT UNKNOWN IF ANY PIECES FELL INTO THE WOUND SITE. THE CASE WAS COMPLETED WITHOUT THE GUIDE. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT INFORMATION AND MEDICAL HISTORY REQUESTED, BUT WAS NOT PROVIDED. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. | PENDING EVALUATION. COULD REQUIRE INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE IF IT WERE TO RECUR. | Y | Y | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00705 | 2021-12-15 | 2021-11-22 | I, n, j, u, r, y | Manufacturer report | LOGIC CR TIB INSERT STD, SZ 4, 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC CR TIB INSERT STD, SZ 4, 11MM | 02-012-47-4011 |  | AS REPORTED, APPROXIMATELY 5.5 YRS POSTOP THE INITIAL IMPLANT, THIS (B)(6) PATIENT HAD A REVISION DUE INCREASE WEAR ON THE TIBIAL SURFACE. REPLACED THE POLY INSERT WITH A SIZE 4, 11MM CRC INSERT. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICE IS AVAILABLE FOR RETURN/EVALUATION. | CONCOMITANT MEDICAL PRODUCTS: LOGIC CR FEMORAL CEM, LEFT, SZ 4 (CAT# 02-010-03-0240 / SERIAL# (B)(4)). LGC TIBIAL FIT TRAY CEM SZ 4F / 5T (CAT# 02-012-45-4050 / SERIAL# (B)(4)). THREE PEG PATELLA 38MM (CAT# 200-02-38 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00707 | 2021-12-15 | 2021-10-14 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | 2MM K-WIRE INSERT | EXACTECH, INC. |  | SSI-300-005880 |  | IT WAS REPORTED THAT DURING A SHOULDER SURGERY, 2 K-WIRE GUIDE HOLD SLEEVES FRACTURED. ASKED, BUT UNKNOWN IF ANY PIECES FELL INTO THE WOUND SITE. THE CASE WAS COMPLETED WITHOUT THE GUIDE. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT INFORMATION AND MEDICAL HISTORY REQUESTED, BUT WAS NOT PROVIDED. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. | DEVICE EVALUATED BY MANUFACTURER? PENDING EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00711 | 2021-12-20 | 2021-11-30 | I, n, j, u, r, y | Manufacturer report | PS TIBIAL INSERTS SZ 3, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | PS TIBIAL INSERTS SZ 3, 9MM | 204-23-09 |  | AS REPORTED, APPROXIMATELY 7 YRS POSTOP THE INITIAL L TKA, THIS (B)(6) MALE PATIENT PRESENTED WITH A ¿SORE KNEE¿ ¿ THE INSITU PATELLA (PATIENT'S NATURAL PATELLA) WAS WORN AND THE PE WAS ALSO WORN OUT. THE PROSTHESIS WAS REVISED TO A COMPETITOR'S KNEE. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THERE WAS NO BREAKAGE OF THE DEVICE. THERE WAS NO SURGICAL DELAY/PROLONGATION. DEVICE IS TO BE RETURNED. | SECTION H11: THE FOLLOWING SECTIONS HAVE CORRECTED INFORMATION: (H3) THE REVISION REPORTED WAS LIKELY THE RESULT OF AXIAL ROTATIONAL MISMATCH BETWEEN THE FEMORAL AND TIBIAL COMPONENTS, THIRD BODY WEAR, PATIENT-RELATED CONDITIONS, OR ANY COMBINATION OF THESE POSSIBILITIES, WHICH LED TO PROSTHESIS WEAR AND PROSTHESIS FRACTURE. ADDITIONALLY, A CONTRIBUTING FACTOR TO THE WEAR AND DELAMINATION MAY HAVE BEEN THE RESULT OF BEING PACKAGED IN A NONCONFORMING VACUUM BAG FOR MORE THAN FIVE YEARS. | Y | Y | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00716 | 2021-12-21 | 2021-10-23 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | GLENOSPHERE, 36MM | EXACTECH, INC. | 320-31-36 | 320-31-36 |  | AS REPORTED, APPROXIMATELY 10 MONTHS POSTOP THE PRIMARY L TSA, THIS (B)(6) FEMALE PRESENTED WITH A SUDDEN ONSET OF PAIN WITHOUT AN INCITING EVENT. CLINICAL SCAPULAR SPINE FRACTURE WAS NOT VISUALIZED ON X-RAY. THE PATIENT HAS A PRE-OP HISTORY OF AVASCULAR NECROSIS OF HUMERAL HEAD. THE REPORTED ACTION TAKEN WAS TO WEAR A SLING AND RELATIVE REST. OUTCOME REPORTED AS RESOLVED ON (B)(6) 2021. THE CASE REPORT FORM INDICATES THIS EVENT IS DEFINITELY RELATED TO DEVICES AND PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. DEVICES WILL NOT RETURN DUE TO STUDY POLICY. | (H3) BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN OR MANUFACTURING ISSUES. THE CAUSE OF THE PAIN AND SUBSEQUENT REVISION CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION. | Y | N | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00712 | 2021-12-21 | 2021-11-09 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE TRAY ADAPTER PLATE TRAY +5 | EXACTECH, INC. | 320-10-05 | 320-10-05 |  | AS REPORTED, APPROXIMATELY THIS PATIENT HAD ALL REVERSE SHOULDER IMPLANTS EXPLANTED BECAUSE OF INFECTION AND IMPLANTED AN ANTIBIOTIC SPACER. INITIAL DATE OF IMPLANT IS UNKNOWN. DEVICES WILL NOT RETURN DUE TO FACILITY POLICY. NO PATIENT INFORMATION OR MEDICAL HISTORY. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. | (H3) BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN OR MANUFACTURING ISSUES. THE CAUSE OF THE INFECTION AND SUBSEQUENT REVISION CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION. | Y | N | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00713 | 2021-12-21 | 2021-11-29 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 38MM GLENOSPHERE | EXACTECH, INC. | 320-01-38 | 320-01-38 |  | AS REPORTED, THIS FEMALE PATIENT WAS REVISED DUE TO RIGHT SHOULDER SUBLUXATION AND COMPONENTS WERE FOR EXPANDED GLENOSPHERE AND LARGER CONSTRAINED 2.5MM. X-RAYS WERE REQUESTED AND WILL BE SENT SEPARATELY. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT RETURN DUE TO HOSPITAL POLICY. | CONCOMITANT DEVICE(S): 320-38-10 - EQUINOXE REVERSE 38MM HUMERAL CONST LINER +0. | Y | N | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00714 | 2021-12-21 | 2021-11-24 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 42MM HUMERAL CONST LINER +2.5 | EXACTECH, INC. | 320-42-13 | 320-42-13 |  | AS REPORTED, APPROXIMATELY 6 MOS POSTOP THE INITIAL L TSA, THIS PATIENT PRESENTED AND DESCRIBES TWO PROLONGED SUBLUXATION AND INSTABILITY EVENTS THAT BOTH SELF-REDUCED. THE PROCEDURE WAS COMPLETED IN OUTPATIENT SURGERY. THE PATIENT WAS REVISED, AND THE EVENT WAS REPORTED AS RESOLVED ON (B)(6) 2021. THE CASE REPORT FORM INDICATES THIS EVENT IS UNLIKELY RELATED TO DEVICES AND POSSIBLY RELATED TO PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | (H3) BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN OR MANUFACTURING ISSUES. THE CAUSE OF THE DISLOCATION AND SUBSEQUENT CLOSED REDUCTION CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION AND THE RESULT OF LOOSENED SUPPORTING LIGAMENTS AND MUSCLES, WHICH ALLOWED FOR DISLOCATION OCCURRING DUE TO THE SIGNIFICANT LIGAMENT RECONSTRUCTION. | Y | N | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00715 | 2021-12-21 | 2021-11-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EQUINOXE | GLENOID IMPACTOR TIP | EXACTECH, INC. |  | 315-07-06 |  | AS REPORTED, DURING THE SURGEON¿S ROUTINE IMPACTION OF AN ANATOMIC GLENOID IMPLANT, THE IMPACTOR TIP BROKE INTO TWO PIECES WHICH WERE BOTH EASILY RETRIEVED. NOTHING FELL INTO THE WOUND SITE. THE (B)(6) MALE WAS NOT AFFECTED BY THE BREAK. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICE IS TO BE RETURNED. | SECTION H10: (D4) LOT NUMBER: 92919002 (H3) BASED ON PREVIOUS COMPLAINT INVESTIGATIONS, THE REPORTED EVENT, AND RETURNED DEVICE, THE FRACTURED GLENOID IMPACTOR TIP REPORTED IN WAS LIKELY THE RESULT OF CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE WHICH WAS LIKELY CAUSED BY ABNORMAL OR AGGRESSIVE USE INCONSISTENT WITH THE INTENDED USE OF THIS DEVICE. | N | Y | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00717 | 2021-12-21 | 2021-12-01 | I, n, j, u, r, y | Manufacturer report | ACETABULAR COMPONENTS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | ACETABULAR COMPONENTS | UNKNOWN |  | AS REPORTED BY THE (B)(6) MALE PATIENT IS 6 YRS POSTOP A LEFT THA AND STARTED HAVING PAIN, HE VISITED THE SURGEON TWICE AND WAS TOLD IT WAS OK. ON THE LAST VISIT, THE SURGEON DECIDED TO SCHEDULE THE PATIENT FOR A FUTURE REVISION DUE TO INCREASED PAIN. APPROXIMATELY 2 DAYS BEFORE THE SCHEDULED PROCEDURE, THE PATIENT WAS STANDING OUTSIDE AND SWATTED AT A ¿HORSE FLY¿. WHEN HE SWATTED, HE TURNED SLIGHTLY A FELT A POP, AT WHICH TIME HE FELL AND HAD TO SLOWLY CRAWL TOWARD HIS FRONT DOOR TO CALL FOR HELP. HE WAS ADMITTED AND A REVISION WAS COMPLETED DURING WHICH THE PATIENT PELVIS WAS FOUND TO BE BROKEN. THE PATIENT STATED DURING THE REVISION, OSTEOARTHRITIS, SEVERE OSTEOLYSIS AND METALLOSIS WAS FOUND ALONG WITH ¿BLACKENED SCAR TISSUE¿. THE INCISION SITE WAS IRRIGATED AND DEBRIDED. HE ALSO SAID THE SURGEON STATED HIS LINER ¿WAS GONE¿. THE PATIENT TOOK THAT TO MEAN THE LINER HAD WORN AWAY. THE PATIENT STATED HIS BLOOD WAS TESTED FOR CHROMIUM. HE STATED HIS LEVEL WAS ¿1 POINT ABOVE NORMAL¿ HE STATED THAT AT 2.5 MONTHS POSTOP, HE STILL NEEDS TO USE A WALKER AND OR CANE TO WALK. HE IS VERY CONCERNED ABOUT HIS SLOW REHAB AND IS UNABLE TO RETURN TO WORK. DEVICES WILL NOT RETURN. | (H3) AS REPORTED BY THE 43 Y/O MALE PATIENT IS 6 YRS POSTOP A LEFT THA AND STARTED HAVING PAIN, HE VISITED THE SURGEON TWICE AND WAS TOLD IT WAS OK. ON THE LAST VISIT, THE SURGEON DECIDED TO SCHEDULE THE PATIENT FOR A FUTURE REVISION DUE TO INCREASED PAIN. APPROXIMATELY 2 DAYS BEFORE THE SCHEDULED PROCEDURE, THE PATIENT WAS STANDING OUTSIDE AND SWATTED AT A ¿HORSE FLY¿. WHEN HE SWATTED, HE TURNED SLIGHTLY A FELT A POP, AT WHICH TIME HE FELL AND HAD TO SLOWLY CRAWL TOWARD HIS FRONT DOOR TO CALL FOR HELP. HE WAS ADMITTED AND A REVISION WAS COMPLETED DURING WHICH THE PATIENT PELVIS WAS FOUND TO BE BROKEN. THE PATIENT STATED DURING THE REVISION, OSTEOARTHRITIS, SEVERE OSTEOLYSIS AND METALLOSIS WAS FOUND ALONG WITH ¿BLACKENED SCAR TISSUE¿. THE INCISION SITE WAS IRRIGATED AND DEBRIDED. HE ALSO SAID THE SURGEON STATED HIS LINER ¿WAS GONE¿. THE PATIENT TOOK THAT TO MEAN THE LINER HAD WORN AWAY. THE PATIENT STATED HIS BLOOD WAS TESTED FOR CHROMIUM. HE STATED HIS LEVEL WAS ¿1 POINT ABOVE NORMAL¿ HE STATED THAT AT 2.5 MONTHS POSTOP, HE STILL NEEDS TO USE A WALKER AND OR CANE TO WALK. HE IS VERY CONCERNED ABOUT HIS SLOW REHAB AND IS UNABLE TO RETURN TO WORK. DEVICES WILL NOT RETURN. UPON REVIEW OF THE AVAILABLE INFORMATION, THERE IS NO EVIDENCE THAT THIS IS A DEVICE RELATED PROBLEM AND THERE IS NO ALLEGATION AGAINST THE DEVICE. IMPLANTATION OF A TOTAL JOINT COULD RESULT PAIN, INFECTION, OR LOOSENING OF TOTAL JOINT HARDWARE. THE MOST LIKELY CAUSE OF THE REPORTED EVENT WAS A MALUNION OF A FRACTURE WHICH WAS LIKELY CAUSED BY PATIENT CONDITIONS. | Y | N | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00718 | 2021-12-21 | 2021-07-13 | I, n, j, u, r, y | Manufacturer report | LOGIC CC FEMORAL SIZE 3, LEFT | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC CC FEMORAL SIZE 3, LEFT | 02-010-06-0230 |  | AS REPORTED BY THE EXACTECH KNEE REPLACEMENT STUDY, APPROXIMATELY THIRTEEN MONTHS POST TKA, THE (B)(6) Y/O MALE PATIENT COMPLAINT OF STIFFNESS AND PAIN TO LEFT KNEE WITH DIFFICULTY WALKING, ASCENDING AND DESCENDING STAIRS. THE EVENT IS POSSIBLY RELATED TO THE DEVICE BUT UNLIKELY RELATED TO THE PROCEDURE. | CONCOMITANT MEDICAL PRODUCTS: TIBIAL INSERT (CAT#: 02-012-65-3011), TIBIAL TRAY (CAT#: 02-022-45-3030), TIBIAL STEM EXTENSION (CAT#: 02-012-60-1440), PATELLA (CAT#: 200-03-35). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00719 | 2021-12-21 | 2021-11-30 | I, n, j, u, r, y | Manufacturer report | CC TIBIAL INSERT SZ 4, 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | CC TIBIAL INSERT SZ 4, 13MM | 208-24-13 |  | DURING OFFICE VISIT, SURGEON DID A CONDYLAR CONSTRAINED LINER EXCHANGE ON A REVISION HE DONE IN 2003, ON (B)(6) 2021. THIS WAS A LINER EXCHANGE OF A SIZE 4/13 MM CONDYLAR CONSTRAINED INSERT. THE LINER HAD MEDIAL WEAR. ALSO THE SPINE STIFFENING SCREW HAD SHEARED OFF AT THE THREADED JUNCTION OF THE SCREW. SURGEON WAS UNABLE TO RETRIEVE THE REST OF THE SCREW THAT WAS LOCKED INTO THE STEM AND TIBIAL TRAY. THIS DID NOT AFFECT THE INSERTION OF A NEW CONDYLAR CONSTRAINED LINER, BUT WAS UNABLE TO PLACE A NEW SPINE STIFFENING SCREW IN DUE TO THE BROKEN PIECE STILL IN THERE. THERE WAS NO BREAKAGE TO DEVICE. THERE WAS NO SURGICAL DELAY OR PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICE WILL BE RETURN. | (H3) THE REVISION REPORTED WAS LIKELY THE RESULT OF PROSTHESIS WEAR, WHICH MAY HAVE BEEN DUE TO THE ALIGNMENT BETWEEN THE FEMORAL AND TIBIAL COMPONENTS, THIRD BODY WEAR, AND/OR PATIENT-RELATED CONDITIONS, AND FRACTURING OF THE SPINE STIFFENING SCREW AFTER BEING IMPLANTED FOR APPROXIMATELY 18 YEARS. THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE REPORTED EVENT OF ¿PROSTHESIS WEAR¿ IS ASSOCIATED WITH MATERIAL DAMAGE TO A SURFACE, USUALLY INVOLVING PROGRESSIVE LOSS OR DISPLACEMENT OF MATERIAL, DUE TO RELATIVE MOTION BETWEEN THAT SURFACE AND A CONTACTING SUBSTANCE OR SUBSTANCES. | Y | Y | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00728 | 2021-12-22 | 2021-12-01 | I, n, j, u, r, y | Manufacturer report | LOGIC FEMORAL PS CEM RIGHT SZ 3.5 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC FEMORAL PS CEM RIGHT SZ 3.5 | 02-010-01-0335 |  | AS REPORTED BY THE EXACTECH KNEE REPLACEMENT STUDY, APPROXIMATELY SIX YEARS POST TKA FEMALE PATIENT (B)(6), PATIENT EXPERIENCED ASEPTIC FEMORAL LOOSENING. PATIENT EXPERIENCED INCREASING PAIN AND LAXITY, AT REVISION WAS PERFORMED FOR FEMORAL COMPONENT LOOSE AT IMPLANT CEMENT INTERFACE. WEAR OF POLY ON POSTERIOR POST. THE EVENT IS POSSIBLY RELATED TO THE DEVICE AND TO THE PROCEDURE. | SECTION H10: (H3) THE REVISION REPORTED IN WAS LIKELY THE RESULT OF BOTH PATIENT CONDITIONS AND AN INSUFFICIENT BOND BETWEEN THE FEMORAL COMPONENT AND THE BONE, WHICH LED TO ASEPTIC(NON-INFECTED) FEMORAL LOOSENING. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICES WERE NOT AVAILABLE FOR EVALUATION AND PRE-REVISION X-RAYS WERE NOT ABLE TO BE OBTAINED. | Y | Y | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00729 | 2021-12-22 | 2021-11-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LOCKING FEMORAL IMPACTOR 0-5 | ORTHOPEDIC MANUAL SURGICAL INSTRUMENT | EXACTECH, INC. | LOCKING FEMORAL IMPACTOR 0-5 | 213-64-01 |  | AS REPORTED, DURING IMPACTION, THE HEAD OF THE FEMORAL IMPACTOR IS FRACTURED. THERE WERE NO PARTS OR PIECES OF THE DEVICE FELL INTO THE PATIENT WOUND SITE. THERE WAS NO SURGICAL DELAY/PROLONGATION. THE DEVICE WILL NOT BE RETURNED AS IT WAS DISPOSED BY THE HOSPITAL. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | N | Y | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00727 | 2021-12-22 | 2021-10-19 | I, n, j, u, r, y | Manufacturer report | NV GXL LNR, LIPPED, 28MM ID, GROUP 1 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LNR, LIPPED, 28MM ID, GROUP 1 CUPS | 132-28-51 |  | AS REPORTED, THE PATIENT UNDERWENT TOTAL LEFT HIP ARTHROPLASTY ON (B)(6) 2014 WITH NOVATION GUP 1 METALLIC HEAD PROSTHESIS AND IN THE REVIEW IN (B)(6) 2021 SHE PRESENTS GREAT WEAR OF POLYETHYLENE WITH A LARGE SUPRA-ACETABULAR CYST. | SECTION H10: (H3) THE PROSTHESIS WEAR AND CYST REPORTED MAY HAVE BEEN DUE TO A COMBINATION OF THE RISK FACTORS SPECIFIED, SUCH AS USE ERROR, IMPLANT POSITIONING, IMPLANT SIZE SELECTION, PATIENT FACTORS (FITNESS FOR SURGERY, BIOMECHANICS, ACTIVITY LEVEL AND LOCAL TISSUE OXIDATION POTENTIAL), THE USE OF NYLON VACUUM BAGS WITHOUT EVOH, AND/OR SURGICAL APPROACH. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICES WERE NOT AVAILABLE FOR EVALUATION AND RADIOGRAPHS WERE NOT PROVIDED. H9: Z-2112-2021; Z-2113-2021; Z-2114-2021; Z-2115-2021; Z-2116-2021; Z-2117-2021; Z-2118-2021; Z-2119-2021; Z-2120-2021; Z-2121-2021; Z-2122-2021; Z-2123-2021; Z-2124-2021; Z-2125-2021; Z-2126-2021; Z-2127-2021; Z-2128-2021; Z-2129-2021; Z-2130 THRU 33-2021. | Y | Y | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00731 | 2021-12-22 | 2021-12-01 | I, n, j, u, r, y | Manufacturer report | CC TIBIAL INSERT SZ 2, 18MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | CC TIBIAL INSERT SZ 2, 18MM | 208-22-18 |  | AS REPORTED, (B)(6) YEARS OLD MALE PATIENT RETURNED WITH KNEE PAIN AND INSTABILITY. SALES REP. WAS ALERTED BY SURGEON¿S OFFICE THAT HE WANTED TO PERFORM A FEMORAL REVISION ON AN OPTETRAK CCK. DUE TO THE FACT THAT THERE WAS AN OFFSET TIBIA IN THE PATIENT AND IT WAS DECIDED TO USED THE OLD CCK. THE FEMORAL COMPONENT WAS REMOVED. THE FEMUR WAS RECUT AND 20MM DISTAL AUGMENTS WERE USED (UP FROM 15MM). STEM LENGTH WAS INCREASED FROM 120MM TO 160MM AND AN APPROPRIATE CCK INSERT, AND OFFSET SCREW WERE USED. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | CONCOMITANT MEDICAL PRODUCTS: CC TIBIAL INSERT SZ 2, 18MM (CAT# 208-22-18 / SERIAL# (B)(4)); CC FEMORAL SZ 2 (CAT# 208-01-02 / SERIAL# (B)(4)); LOGIC STEM EXT 14MM X 160MM (CAT# 02-012-60-1416 / SERIAL# (B)(4)); THREE PEG PATELLA 35MM (CAT# 200-02-35 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00730 | 2021-12-22 | 2021-12-03 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 42MM HUMERAL CONST LINER +2.5 | EXACTECH, INC. | 320-42-13 | 320-42-13 |  | AS REPORTED, THIS MALE PATIENT REQUIRED A LEFT 2ND STAGE REVISION, THE POLY HAD DISLOCATED FROM THE TRAY. UPON INSPECTION, THE PATIENT REQUIRED A REVISION OF THE POLY AND THE TRAY, -ZERO TRAY TO A +10 TRAY. THE LOCKING SCREW WAS COLD WELDED AND NEEDED TO BE DIAMOND SAWED FOR IT TO BE REMOVED. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. | PENDING EVALUATION. CONCOMITANT MEDICAL DEVICE(S): 320-20-00, 4590223 - EQ REVERSE TORQUE DEFINING SCREW KIT; 320-15-05, 4573656 - EQ REV LOCKING SCREW; 320-10-00, 4573805 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0; 320-02-42, 3951501 - RS EXPANDED GLENOSPHERE 42MM, +4MM OFFSET. | Y | N | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00726 | 2021-12-22 | 2014-10-06 | I, n, j, u, r, y | Manufacturer report | NV GXL LNR, LIPPED, 28MM ID, GROUP 1 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LNR, LIPPED, 28MM ID, GROUP 1 CUPS | 132-28-51 |  | THE PATIENT UNDERWENT TOTAL LEFT HIP ARTHROPLASTY ON (B)(6) 2014 USING NOVATION PROSTHESIS AND IN THE CURRENT REVIEW SHOWS GREAT POLYETHYLENE WEAR IN A VERY FAST TIME THAT SHOULD NOT HAVE HAPPENED, WITH A LARGE SUPRA ACETABULAR CYST. | SECTION H10: (H3) THE PROSTHESIS WEAR AND CYST REPORTED MAY HAVE BEEN DUE TO A COMBINATION OF THE RISK FACTORS SPECIFIED, SUCH AS USE ERROR, IMPLANT POSITIONING, IMPLANT SIZE SELECTION, PATIENT FACTORS (FITNESS FOR SURGERY, BIOMECHANICS, ACTIVITY LEVEL AND LOCAL TISSUE OXIDATION POTENTIAL), THE USE OF NYLON VACUUM BAGS WITHOUT EVOH, AND/OR SURGICAL APPROACH. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICES WERE NOT AVAILABLE FOR EVALUATION AND RADIOGRAPHS WERE NOT PROVIDED. H9: Z-2112-2021; Z-2113-2021; Z-2114-2021; Z-2115-2021; Z-2116-2021; Z-2117-2021; Z-2118-2021; Z-2119-2021; Z-2120-2021; Z-2121-2021; Z-2122-2021; Z-2123-2021; Z-2124-2021; Z-2125-2021; Z-2126-2021; Z-2127-2021; Z-2128-2021; Z-2129-2021; Z-2130 THRU 33-2021. | Y | Y | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00720 | 2021-12-22 | 2021-11-05 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | EXPANDED GLENOSPHERE, 40MM, FOR SMALL REVERSE | EXACTECH, INC. | 320-32-40 | 320-32-40 |  | AS REPORTED, APPROXIMATELY 2 MONTHS POST OP THE INITIAL RTSA, THIS (B)(6) MALE PATIENT WAS REVISED DUE TO THE GLENOSPHERE LOCKING SCREW BACKING OUT CAUSING THE GLENOSPHERE TO BE LOOSE. THE ORIGINAL GLENOSPHERE AND GLENOSPHERE LOCKING SCREW WERE EXPLANTED AND NEW HARDWARE REPLACED THE EFFECTED IMPLANTS. THERE WAS A SMALL SURGICAL DELAY. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETUNING DUE TO HOSPITAL POLICY. | PENDING EVALUATION CONCOMITANT DEVICE(S): 320-35-01, 7048184 - SMALL GLENOID PLATE; 320-15-05, 7096565 - EQ REV LOCKING SCREW. NO INFORMATION PROVIDED IN THE FOLLOWING SECTION(S). | Y | N | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00724 | 2021-12-22 | 2021-08-27 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 38MM GLENOSPHERE | EXACTECH, INC. | 320-01-38 | 320-01-38 |  | AS REPORTED, APPROXIMATELY 2 YEARS POST OP INITIAL LEFT TSA, THIS (B)(6) FEMALE PATIENT STARTED EXPERIENCING ASEPTIC GLENOID LOOSENING. REVISION FOR STUDY HAS NOT YET OCCURRED, NO ACTION WAS TAKEN AT THIS TIME. THE CASE REPORT FORM INDICATES THIS EVENT IS POSSIBLY RELATED TO DEVICES AND PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | SECTION H10: (H3) THE GLENOID BASEPLATE LOOSENING REPORTED WAS LIKELY THE RESULT OF AN INSUFFICIENT BOND BETWEEN THE GLENOID BASEPLATE AND THE BONE AND/OR PATIENT-RELATED CONDITIONS. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICES AND X-RAYS WERE NOT AVAILABLE FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00723 | 2021-12-22 | 2021-12-03 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 42MM HUMERAL LINER +0 | EXACTECH, INC. | 320-42-00 | 320-42-00 |  | AS REPORTED, APPROXIMATELY 1.5 YEARS POST OP INITIAL RTSA, THIS (B)(6) Y/O PATIENT WAS REVISED DUE TO INSTABILITY. PATIENT HAD BEEN EXPERIENCING EPISODES OF INSTABILITY AND INCREASED RESTING PAIN. THE CASE REPORT FORM INDICATES THIS EVENT IS POSSIBLY RELATED TO DEVICES AND DEFINITELY RELATED TO THE PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | SECTION H10: (H3) BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN OR MANUFACTURING ISSUES. THE CAUSE OF THE INSTABILITY AND SUBSEQUENT REVISION CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION. | Y | N | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00722 | 2021-12-22 | 2021-12-08 | I, n, j, u, r, y | Manufacturer report | VANTAGE | SHARP PERIPHERAL PUNCH | EXACTECH, INC. |  | SK-015442 |  | AS REPORTED, WHILE IMPACTING THE ANKLE PUNCH INTO BONE OF THIS (B)(6) MALE PATIENT, THE TIP OF THAT PUNCH BROKE OFF IN THE BONE. THE FULL TIP REMAINED INTACT AND WAS PULLED OUT IMMEDIATELY. THERE WAS A SMALL SURGICAL DELAY, IT TOOK A FEW SECONDS TO TAKE THE PIECE OUT AND RE-PUNCH WITH A NEW INSTRUMENT. THERE WAS NO ADVERSE OUTCOME. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00721 | 2021-12-22 | 2021-12-07 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | HUMERAL STEM PRIMARY, PRESS FIT 11MM | EXACTECH, INC. | 300-01-11 | 300-01-11 |  | AS REPORTED, APPROXIMATELY 5 YEARS POST OP THE INITIAL RTSA, THIS (B)(6) FEMALE PATIENT WAS REVISED DUE TO THE HUMERUS AND BASEPLATE LOOSENING. THE CABLED HUMERUS FROM THE INITIAL SURGERY HAD FAILED TO SUPPORT THE FRACTURED HUMERUS. THE SURGEON DECIDED TO USE A DEPUY STEM WHO HAD A LARGER HEAD SIZE TO OFFER. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETUNING DUE TO HOSPITAL POLICY. | SECTION H10: (H3) THE REVISION REPORTED WAS LIKELY THE RESULT OF BOTH PATIENT CONDITIONS AND AN INSUFFICIENT BOND BETWEEN THE IMPLANTS AND THE BONES, WHICH LED TO ASEPTIC (NON-INFECTED) HUMERAL AND BASEPLATE LOOSENING. THE FAILED CABLED HUMERUS LIKELY CONTRIBUTED TO THE REPORTED HUMERAL LOOSENING. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICES WERE NOT RETURNED FOR EVALUATION AND X-RAY IMAGES WERE NOT PROVIDED. | Y | N | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00725 | 2021-12-22 | 2021-12-09 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | CAGE GLENOID M, POST AUG, LEFT | EXACTECH, INC. | 314-13-23 | 314-13-23 |  | AS REPORTED, ALMOST 2 YEARS POST OP LEFT TSA, THIS MALE PATIENT EXPERIENCED DISCOMFORT FOR LESS THAN A 1 YEAR AFTER SURGERY. SURGEON REMOVED THE CAGED GLENOID, POLY HAD SHEERED OFF. CENTER CAGE WAS WELL FIXED IN GLENOID BUT POLY + 3 PERIPHERAL PEGS WERE LOOSE. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. | SECTION H10: H3: THE REVISION REPORTED WAS LIKELY THE RESULT OF FAILURE OF THE CAGE GLENOID COMPONENT. POSSIBLE CONTRIBUTING FACTORS INCLUDE BUT ARE NOT LIMITED TO INCOMPLETE OR OFF-AXIS SEATING OF THE GLENOID COMPONENT AT THE TIME OF IMPLANTATION, PATIENT CONDITIONS, AND/OR ABNORMAL FORCES BEING APPLIED TO THE GLENOID COMPONENT DURING PHYSICAL THERAPY, OTHER PHYSICAL ACTIVITY, OR A POST TRAUMATIC EVENT. THE ROOT CAUSE OF THE FAILURE CANNOT BE DETERMINED BECAUSE THE DEVICE WAS NOT RETURNED FOR EVALUATION AND NOT ENOUGH INFORMATION HAS BEEN PROVIDED AT THIS TIME. | Y | N | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00732 | 2021-12-23 | 2018-01-01 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 4 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 4 11MM | 02-012-35-4011 |  | APPROXIMATELY 9 YRS POSTOP THE INITIAL L TKA, THIS (B)(6) Y/O PATIENT PRESENTED WITH A SWOLLEN KNEE (LEFT) DUE TO POLYETHYLENE WEAR WITH EVIDENCE OF OSTEOLYSIS SEEN ON X-RAY. NO ACTION TAKEN, WILL REVIEW KNEE IN 6 MONTHS. THE EVENT IS POSSIBLY RELATED TO THE DEVICE BUT NOT RELATED TO THE PROCEDURE. | THE PROSTHESIS WEAR AND OSTEOLYSIS REPORTED MAY HAVE BEEN DUE TO THE ALIGNMENT BETWEEN THE FEMORAL AND TIBIAL COMPONENTS, THIRD BODY WEAR, AND/OR PATIENT-RELATED FACTORS. HOWEVER, THIS CANNOT BE CONFIRMED AS THE PATIENT HAS NOT YET BEEN REVISED AT THE TIME OF THIS INVESTIGATION. THEREFORE, THE DEVICES WERE NOT AVAILABLE FOR EVALUATION. THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE REPORTED EVENT OF "PROSTHESIS WEAR" IS ASSOCIATED WITH MATERIAL DAMAGE TO A SURFACE, USUALLY INVOLVING PROGRESSIVE LOSS OR DISPLACEMENT OF MATERIAL, DUE TO RELATIVE MOTION BETWEEN THAT SURFACE AND A CONTACTING SUBSTANCE OR SUBSTANCES. | Y | Y | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00733 | 2021-12-23 | 2021-12-02 | I, n, j, u, r, y | Manufacturer report | LOGIC FEMORAL PS CEM RIGHT SZ 1.5 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC FEMORAL PS CEM RIGHT SZ 1.5 | 02-010-01-0315 |  | AS REPORTED VIA LEGAL DOCUMENTATION, THIS PATIENT HAD RIGHT KNEE REPLACEMENT ON (B)(6)2017 DUE TO OSTEOARTHRITIS. THE PATIENT UNDERWENT LEFT KNEE REVISION SURGERY ON (B)(6)2021, APPROXIMATELY 4 YEARS 1 MONTH AFTER THEIR PRIMARY SURGERY. (B)(6) 2021 OP REPORT FOUND CLEAR, STRAW-COLORED JOINT FLUID THAT YIELDED TRACE LEUKOCYTE ESTERASE AND A LOW NUMBER OF NUCLEATED CELLS ON CELL COUNT, EXTENSIVE POLYETHYLENE AND/OR CEMENT SYNOVITIS, MILD LATERAL PATELLAR FACET WEAR, NO PATELLAR LOOSENING, GROSS FEMORAL LOOSENING DUE TO DEBONDING, DISTAL FEMORAL OSTEOLYSIS, A WELL-FIXED TIBIAL IMPLANT WITH MINIMAL POSTERIOR MEDIAL TIBIAL OSTEOLYSIS, AND PITTING AND DELAMINATION OF THE MEDIAL PATELLAR INSERT ON THE TOPSIDE. THE PATIENT WAS TRANSFERRED TO THE RECOVERY ROOM IN STABLE CONDITION. THERE IS NO OTHER PATIENT DEMOGRAPHIC OR MEDICAL HISTORY AVAILABLE. THERE IS NO DEVICE RETURN. THERE ARE NO PHOTOS OR OTHER IMAGES OF THE DEVICE PROVIDED. NO ADDITIONAL INFORMATION IS AVAILABLE. | BASED ON THE RECEIPT OF ADDITIONAL INFORMATION, THE FOLLOWING FIELDS HAVE BEEN UPDATED: B5, D6A, D8, H7, AND H9. FURTHER UPDATES AND/OR RE-OPENED INVESTIGATION RESULTS WILL BE REPORTED WITHIN 30 DAYS OF THEIR RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00734 | 2021-12-23 | 2021-12-02 | I, n, j, u, r, y | Manufacturer report | PS TIBIAL INSERTS SZ 4, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | PS TIBIAL INSERTS SZ 4, 9MM | 204-24-09 |  | AS REPORTED, THIS IS A (B)(6) PATIENT AT THE TIME THAT HAD A LEFT TOTAL KNEE DONE IN 2009. THE PATIENT IS AN ASPHALT WORKER AND OWNER. THE PATIENT HAD WORN THROUGH HIS 9 MM POLY INSERT. CAUSING THE FEMORAL COMPONENT TO BE RUBBING ON THE TIBIAL COMPONENT WEARING DOWN THE POSTERIOR FLANGE OF THE TIBIA. PATIENT HAD A SIGNIFICANT AMOUNT OF OSTEOLYSIS. PATIENT WAS REVISED TO EXACTECH CC COMPONENTS. THE FEMUR WAS DOWNSIZE TO SIZE 3 CONDYLAR CONSTRAINED FEMUR WITH 18X120 STEM. THE TIBIA SIZE 3F/2T WITH A 5MM TIBIAL AUG ON MEDICAL SIDE WITH 18X80 STEM. THEN A 38 THREE PEG PATELLA, AN ENDED WITH A SIZE 3/19MM CONDYLAR CONSTRAINED INSERT. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICE WILL BE RETURN. | CONCOMITANT MEDICAL PRODUCTS: LEFT SIZE 4 LOGIC FEMUR (CAT# 234-02-04), 4F/3T FIN TIBIAL TRAY (CAT# 200-04-43), 35 3-PEG PATELLA (CAT# 200-02-35), 14X25 EXTENSION STEM (CAT# 02-012-60-1425). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | Y | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00735 | 2021-12-23 | 2021-12-02 | I, n, j, u, r, y | Manufacturer report | TRULIANT TIB IMP PS INSERT SZ 3 10MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT TIB IMP PS INSERT SZ 3 10MM | 02-022-35-3010 |  | AS REPORTED, SURGEON SHOWED X-RAYS OF A (B)(6) FEMALE PATIENT, WHERE IT SHOWED THAT THE PATIENT HAD AN INSTABILITY ISSUE AND NEED TO BE REVISED TO A CONSTRAINED INSERT. SURGEON REMOVED A SIZE 3 FEMUR PS AND A SIZE 3/10 MM PS INSERT AND REVISING TO A SIZE 3 CONDYLAR CONSTRAINED FEMUR WITH A 18X120 STEM AND SIZE 3/11MM CONDYLAR CONSTRAINED INSERT. THERE WAS NO BREAKAGE TO THE DEVICE. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICE WILL BE RETURN. | CONCOMITANT MEDICAL PRODUCTS: TRULIANT PS CEM FEM PS CEM RIGHT SZ 3 (CAT# 02-020-11-0330 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00736 | 2021-12-23 | 2018-06-19 | I, n, j, u, r, y | Manufacturer report | OPTETRAK ASY,PS CEMENTED FEMORAL, SZ 3, | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | OPTETRAK ASY,PS CEMENTED FEMORAL, SZ 3, | 234-02-03 |  | AS REPORTED BY THE EXACTECH KNEE REPLACEMENT STUDY, APPROXIMATELY THREE YEARS POST TKA, THE (B)(6) FEMALE PATIENT EXPERIENCED PAIN AND FUNCTIONAL LIMITATION OF THE LEFT KNEE WITH INFLAMMATORY CHANGES IN SURGICAL WOUND, RADIOGRAPHY EVIDENCE LOOSENING OF TIBIAL AND FEMORAL COMPONENT, ORDERS ARE GIVEN IN FIRST TIME WITH SPACER AND CULTURE TAKING. THERE IS NO FURTHER INFORMATION FOUND REGARDING TREATMENT OR A REVISION. THE EVENT IS POSSIBLY RELATED TO THE DEVICE AND TO THE PROCEDURE. | CONCOMITANT MEDICAL PRODUCTS: TIBIAL INSERT (CAT# 204-23-11), TIBIAL TRAY (CAT# 200-04-32), PATELLA (CAT# 200-02-29) . ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00737 | 2021-12-27 | 2021-12-02 | I, n, j, u, r, y | Manufacturer report | LOGIC FEMORAL PS CEM LEFT SZ 3 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC FEMORAL PS CEM LEFT SZ 3 | 02-010-01-0230 |  | AS REPORTED, THE (B)(6), FEMALE PATIENT COMPLAINED OF HAVING PAIN AND FEMUR WAS LOOSE. THERE WAS NO BREAKAGE OF DEVICE. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICE WILL BE RETURN. | (H3) THE REVISION REPORTED WAS LIKELY THE RESULT OF A DEGRADATION OF THE BOND BETWEEN THE FEMORAL COMPONENT AND THE BONE, PATIENT¿S UNDERLYING CONDITIONS, AND LONGEVITY OF THE IMPLANTED DEVICE, WHICH LED TO ASEPTIC (NON-INFECTED) FEMORAL LOOSENING. THE LOOSENING MAY HAVE LED TO INSTABILITY AND AN INCREASE IN CEMENT AND BONE PARTICLES IN THE JOINT SPACE AND RESULTED IN PROSTHESIS WEAR. THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE REPORTED EVENT OF "LOOSENING - FEMORAL¿ IS ASSOCIATED WITH WEAKENED INTEGRATION OF THE FEMORAL COMPONENT AT THE BONE-IMPLANT INTERFACE DUE TO LOSS OF FIBROUS AND/OR BONY TISSUE AND LEADING TO COMPROMISED ANCHORAGE OF THE DEVICE. FURTHERMORE, THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE REPORTED EVENT OF "PROSTHESIS WEAR¿ IS ASSOCIATED WITH MATERIAL DAMAGE TO A SURFACE, USUALLY INVOLVING PROGRESSIVE LOSS OR DISPLACEMENT OF MATERIAL, DUE TO RELATIVE MOTION BETWEEN THAT SURFACE AND A CONTACTING SUBSTANCE OR SUBSTANCES. | Y | Y | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00738 | 2021-12-29 | 2021-12-14 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | CAGE GLENOID L, POST AUG, RIGHT | EXACTECH, INC. |  | 314-13-34 |  | AS REPORTED, APPROXIMATELY 2 YEARS POST OP THE INITIAL RTSA, A (B)(6) MALE PATIENT'S POLY DISASSOCIATED FROM THE PEGS. PATIENT BELIEVES HE FELT A POP WHILE LIFTING SOMETHING. PATIENT WAS CONVERTED FROM A STEMLESS TSA TO A RTSA. THERE WAS NO SURGICAL DELAY OR PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. | THE REVISION REPORTED WAS LIKELY THE RESULT OF FAILURE OF THE CAGE GLENOID COMPONENT. POSSIBLE CONTRIBUTING FACTORS INCLUDE BUT ARE NOT LIMITED TO INCOMPLETE OR OFF-AXIS SEATING OF THE GLENOID COMPONENT AT THE TIME OF IMPLANTATION, PATIENT CONDITIONS, AND/OR ABNORMAL FORCES BEING APPLIED TO THE GLENOID COMPONENT DURING PHYSICAL THERAPY, OTHER PHYSICAL ACTIVITY, OR A POST TRAUMATIC EVENT. THE ROOT CAUSE OF THE FAILURE CANNOT BE DETERMINED BECAUSE THE DEVICE WAS NOT RETURNED FOR EVALUATION AND NOT ENOUGH INFORMATION HAS BEEN PROVIDED AT THIS TIME. NO INFORMATION PROVIDED. | Y | N | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00739 | 2021-12-30 | 2021-12-07 | I, n, j, u, r, y | Manufacturer report | PS TIBIAL INSERTS SZ 2, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | OPTETRAK PS TIBIAL INSERT | 204-22-09 |  | AS REPORTED, APPROXIMATELY SIX MONTHS POST RIGHT TKA IMPLANTATION, A (B)(6) FEMALE PATIENT HAD ANTERIOR KNEE PAIN AND A POLY AND PATELLA EXCHANGE WAS MADE. THERE WERE NO PARTS OR PIECES FELL INTO THE PATIENT WOUND AREA. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICE WILL BE RETURN. | SECTION H11: THE FOLLOWING SECTIONS HAVE CORRECTED INFORMATION: (D1) BRAND NAME: PS TIBIAL INSERTS SZ 2, 9MM, (D4) CATALOG NUMBER: 204-22-09, UNIQUE IDENTIFIER (UDI) #: (B)(4). (D10) CONCOMITANT DEVICES: 200-02-29, (B)(6) - PATELLA. 02-012-60-1280, (B)(6) - LOGIC STEM EXT 12MM X 80MM. 208-06-02, (B)(6) - CC DISTAL FEM AUGMENT SZ 2, 10MM. 208-06-02, (B)(6) - CC DISTAL FEM AUGMENT SZ 2, 10MM. 02-012-44-2013, (B)(6) - LOGIC TIBIA IMPLANT PSC INSERT, SZ 2, 13MM. 02-012-61-6000, (B)(6) - LOGIC OFFSET STEM EXT COUPLER 6MM. 02-010-06-0522, (B)(6) - LOGIC POST. AUG. BLOCK SIZE 2, 10MM. 02-010-06-0220, (B)(6) - LOGIC CC FEMORAL SIZE 2, LEFT. (G4) PMA/510(K)NUMBER: K932690. (H3) THE REVISION REPORTED WAS LIKELY DUE TO AXIAL ROTATIONAL MISMATCH BETWEEN THE FEMORAL AND TIBIAL COMPONENTS, THIRD-BODY WEAR, PATIENT-RELATED CONDITIONS, OR ANY COMBINATION OF THESE POSSIBILITIES, WHICH LED TO PROSTHESIS WEAR ON THE TIBIAL INSERT AND THE PATELLA CAUSING THE REPORTED ANTERIOR KNEE PAIN. | Y | Y | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00740 | 2021-12-30 | 2021-12-13 | I, n, j, u, r, y | Manufacturer report | ACETABULAR COMPONENTS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | ACETABULAR COMPONENTS | UNKNOWN |  | AS REPORTED, PATIENT WAS REVISED. PATIENT'S RIGHT HIP WAS DONE IN 1998. THERE IS NO IMPLANT RECORD OR OP REPORT FOUND. SURGEON REVISED THE RIGHT HIP, BONE GRAFTED, AND USED A SMITH AND NEPHEW CEMENTED LINER. THE MCS CUP WAS RETAINED AND NEW 11/13 HEAD IMPLANTED. X-RAY AND IMAGES WERE RECEIVED, NO OTHER INFORMATION AVAILABLE. PATIENT LAST KNOWN TO BE IN STABLE CONDITION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00741 | 2021-12-30 | 2021-12-18 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED, | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 9MM | 02-012-44-4009 |  | AS REPORTED, A (B)(6) FEMALE PATIENT UNDERWENT INDEX ARTHROPLASTY IN (B)(6) 2019. HER POLY WAS MANUFACTURED IN JUNE 2014. PATIENT DID WELL FOR 6 MONTHS, THEN STARTED TO DEVELOP PERSISTENT, LARGE EFFUSIONS. INFECTION WORK UP NEGATIVE. X-RAYS SHOW CLASSIC PATTERN OF MEDIAL TIBIAL PLATEAU RADIOLUCENCY AND RADIOLUCENCY UNDER FEMORAL FLANGE. SHE UNDERWENT REVISION BY ANOTHER PRACTITIONER RECENTLY. PATIENT WAS TOLD SHE HAD EXTENSIVE DELAMINATION OF THE POLY WITH EXTENSIVE BONE LOSS. NO ADDITIONAL INFORMATION AVAILABLE. | THIS FOLLOW-UP REPORT IS BEING SUBMITTED TO RELAY ADDITIONAL AND/OR CORRECTED INFORMATION. THE REASON FOR THE REVISION REPORTED CANNOT BE CONFIRMED FROM THE INFORMATION PROVIDED BUT MAY BE THE RESULT OF PROSTHESIS WEAR, INSTABILITY, OR DUE TO INCLUSION OF THE POLYETHYLENE IN THE PACKAGING RECALL. POTENTIAL CONTRIBUTIONS OF USER AND PATIENT-RELATED CONSIDERATIONS TO THE EVENT COULD NOT BE ASSESSED AS THE DEVICES WERE NOT AVAILABLE FOR EVALUATION AND IMAGES WERE NOT PROVIDED. IF ANY FURTHER INFORMATION IS OBTAINED THAT WOULD CHANGE OR ALTER ANY INFORMATION PROVIDED, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY. | Y | Y | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00742 | 2021-12-30 | 2021-12-01 | I, n, j, u, r, y | Manufacturer report | FEMORAL COMPONENTS | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | FEMORAL COMPONENTS | UNKNOWN |  | AS REPORTED, A (B)(6) FEMALE PATIENT UNDERWENT INDEX ARTHROPLASTY ON (B)(6) 2015. HER POLY WAS MANUFACTURED IN NOV 2012. PATIENT PRESENTED WITH PAIN AND EFFUSIONS SEVERAL MONTHS AGO AND WAS RECENTLY REVISED BY ANOTHER PRACTITIONER. THERE ARE X-RAYS AVAILABLE FOR EVALUATION. NO ADDITIONAL INFORMATION AVAILABLE. | SECTION H10: (H3) BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN, MANUFACTURING, OR PATIENT RELATED ISSUES. THE CAUSE OF THE PAIN AND SUBSEQUENT REVISION CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY IS RELATED TO THE PATIENT¿S UNDERLYING CONDITION. | Y | N | HEALTH PROFESSIONAL |  | 2021-12
1038671-2021-00743 | 2021-12-30 | 2021-12-01 | I, n, j, u, r, y | Manufacturer report | FEMORAL COMPONENTS | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | FEMORAL COMPONENTS | UNKNOWN |  | AS REPORTED, A (B)(6) MALE PATIENT UNDERWENT INDEX ARTHROPLASTY (B)(6) 2018. POLY MANUFACTURED IN OCT 2010. PATIENT PRESENTED WITH PERSISTENT EFFUSIONS AND OSTEOLYSIS UNDER FEMORAL FLANGE. PATIENT UNDERWENT REVISION SEVERAL MONTHS AGO. X-RAYS AVAILABLE FOR EVALUATION. NO ADDITIONAL INFORMATION AVAILABLE. | THE REVISION REPORTED MAY HAVE BEEN THE RESULT OF WEAR OF THE POLYETHYLENE TIBIAL INSERT, LEADING TO OSTEOLYSIS. THE EFFUSIONS REPORTED MAY HAVE BEEN THE RESULT OF INFECTION. ADDITIONALLY, A CONTRIBUTING FACTOR TO THE OSTEOLYSIS MAY HAVE BEEN THE RESULT OF THE INSERT BEING PACKAGED IN A NON-CONFORMING VACUUM BAG FOR MORE THAN FIVE YEARS. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE DEVICES WERE NOT RETURNED TO EXACTECH FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2021-12
1038671-2022-00001 | 2022-01-04 | 2021-12-15 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | CAGED GLENOID | EXACTECH, INC. | UNK | UNK |  | AS REPORTED, APPROXIMATELY 4 YEARS POST OP THE INITIAL LEFT TSA, THIS PATIENT WAS REVISED DUE TO PAIN AND POLY FAILURE WITH SIGNIFICANT RADIOLUCENT LINES AND RESORPTION. THE GLENOID COMPONENT WAS EASILY REMOVED AS IT HAD LOOSENED FROM ORIGINAL POSITION AND ALL 4 PEGS APPEARED INTACT WITH MINIMAL BONE IN GROWTH. SIGNIFICANT BONE GRAFTING WITH FRESH FEMORAL HEAD WAS DONE TO THE GLENOID AND THE SHOULDER WAS REVISED TO BIOMET REVERSE. THE PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION. NO OTHER INFORMATION AVAILABLE. NO OTHER PATIENT INFORMATION OR MEDICAL HISTORY REPORTED. DEVICES WILL NOT BE RETURNING DUE TO HOSPITAL POLICY. | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00002 | 2022-01-04 | 2021-11-03 | I, n, j, u, r, y | Manufacturer report | TRULIANT TIB IMP PS INSERT SZ 4.5 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT TIB IMP PS INSERT SZ 4.5 11MM | 02-022-35-4511 |  | AS REPORTED BY THE EXACTECH TRULIANT KNEE CLINICAL STUDY, APPROXIMATELY ONE-YEAR POST TKA, THE (B)(6), MALE PATIENT EXPERIENCED A GLOBAL LAXITY. PATIENT WAS SEEN IN CLINIC (B)(6) 2021 FOR COMPLAINTS OF KNEE LOCK, INCREASE IN PAIN, WEAKNESS, SWELLING AND GIVING WAY IN LEFT KNEE WITH HINGED KNEE CAGE BRACE PLACED FOR EXTRA SUPPORT. PATIENT WAS SEEN (B)(6) 2021 IN A TRUPULL BRACE WITH SOME IMPROVEMENT. PATIENT TO BE SCHEDULED FOR A POLYETHYLENE EXCHANGE FOR GLOBAL LAXITY. THE EVENT IS NOT RELATED TO THE DEVICE BUT DEFINITELY RELATED TO THE PROCEDURE. | CONCOMITANT MEDICAL PRODUCTS: TRULIANT FEMORAL COMPONENT (CAT# 02-020-11-0245). TIBIAL TRAY (CAT# 02-022-45-4545). PATELLA (CAT# 200-02-38). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00003 | 2022-01-04 | 2021-12-16 | I, n, j, u, r, y | Manufacturer report | LOGIC CR TIB INSERT SLOPE+, SZ 3, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC CR TIB INSERT SLOPE+, SZ 3, 9MM | 02-012-48-3009 |  | AS REPORTED, A (B)(6) FEMALE PATIENT HAD A REVISION DUE TO EFFECTED POLY CAUSING POLY WEAR WITHIN 15 MONTHS. PATIENT'S LINER WAS REPLACED. THERE WAS NO BREAKAGE OF THE DEVICE. THERE WAS NO SURGICAL DELAY OR PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. | CONCOMITANT MEDICAL PRODUCTS: LOGIC CR FEMORAL CEM, LEFT, SZ 3 (CAT# 02-010-03-0230 / SERIAL# (B)(4)). LGC TIBIAL FIT TRAY CEM SZ 3F / 3T (CAT# 02-012-45-3030 / SERIAL# (B)(4)). THREE PEG PATELLA 35MM (CAT# 200-02-35 / SERIAL#(B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00004 | 2022-01-04 | 2021-12-16 | I, n, j, u, r, y | Manufacturer report | LOGIC FEMORAL PS CEM LEFT SZ 3.5 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC FEMORAL PS CEM LEFT SZ 3.5 | 02-010-01-0235 |  | AS REPORTED, A (B)(6) FEMALE PATIENT COMPLAINED OF HAVING PAIN. FEMUR WAS LOOSE. THERE WAS NO BREAKAGE OF DEVICE. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICE WILL BE RETURN. | CONCOMITANT MEDICAL PRODUCTS: LOGIC TIBIA PS MOD INSRT SZ 3.5 11MM (CAT# 02-012-35-3511 / SERIAL# (B)(4), LGC TIBIAL FIT TRAY CEM SZ 3.5F / 2.5T (CAT# 02-012-45-3525 / SERIAL# (B)(4), THREE PEG PATELLA 32MM (CAT# 200-02-32 / SERIAL# (B)(4), LOGIC FEMORAL PS CEM RIGHT SZ 3.5 (CAT# 02-010-01-0335 / SERIAL# (B)(4), LOGIC TIBIA PS MOD INSRT SZ 3.5 9MM (CAT# 02-012-35-3509 / SERIAL# (B)(4), LGC TIBIAL FIT TRAY CEM SZ 3.5F / 2.5T (CAT# 02-012-45-3525 / SERIAL# (B)(4), THREE PEG PATELLA 35MM (CAT# 200-02-35 / SERIAL# (B)(4), HOLDING PIN MINI SHARP POINT 4 PK (CAT# 201-78-15 / SERIAL# (B)(4). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00010 | 2022-01-05 | 2019-03-19 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 3 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 3 11MM | 02-012-35-3011 |  | AS REPORTED, APPROXIMATELY THREE YEARS POST TKA, PATIENT HAD A REVISION FOR AN UNKNOWN REASON. CC INSERT WITH > 5-YEARS AT IMPLANTATION BUT OUT OF SCOPE FOR THE PE LINER ISSUE. NO ADDITIONAL INFORMATION AVAILABLE. | CONCOMITANT MEDICAL PRODUCTS: TIBIAL AUGMENT BLOCK 1/2-SZ 3 11MM LLRM (CAT# 204-63-89 / SERIAL# (B)(4)); TIBIAL AUGMENT BLOCK 1/2-SZ 3 11MM RLLM (CAT# 204-63-88 / SERIAL# (B)(4)); THREE PEG PATELLA 32MM (CAT# 200-02-32 / SERIAL# (B)(4)); STEM EXTENSION 80L X12 MM (CAT# 204-32-08 / SERIAL# (B)(4)); CC POSTERIOR FEM AUGMENT SZ 2, 5MM (CAT# 208-07-02 / SERIAL# (B)(4)); CC STEM EXT ADAPTOR 5 DEGREE (CAT# 208-09-05 / SERIAL# (B)(4)); TRAPEZOID TIBIAL TRAY SZ 1F/3T, 2F/3T (CAT# 204-04-23 / SERIAL# (B)(4)); CC FEMORAL SZ 2 (CAT# 208-01-02 / SERIAL# (B)(4)); FLUTED STEM EXTENSION 80L X 14 MM (CAT# 204-34-08 / SERIAL# (B)(4)); CC DISTAL FEM AUGMENT SZ 2, 5MM (CAT# 208-05-02 / SERIAL# (B)(4)); CC DISTAL FEM AUGMENT SZ 2, 5MM (CAT# 208-05-02 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00009 | 2022-01-05 | 2020-02-04 | I, n, j, u, r, y | Manufacturer report | PS TIBIAL INSERTS SZ 3, 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | PS TIBIAL INSERTS SZ 3, 11MM | 204-23-11 |  | AS REPORTED, APPROXIMATELY FOUR YEARS POST TKA, PATIENT HAD A REVISION FOR AN UNKNOWN REASON. IT IS UNKNOWN IF THIS IS A SECOND TIME REVISION FOR THE SAME SIDE. IT IS ALSO UNKNOWN WHY CC INSERT OUT OF SCOPE. NO ADDITIONAL INFORMATION AVAILABLE. | CONCOMITANT MEDICAL PRODUCTS: CC POSTERIOR FEM AUGMENT SZ 2, 5MM (CAT# 208-07-02 / SERIAL # (B)(4)), CC DISTAL FEM AUGMENT SZ 2, 10MM (CAT # 208-06-02 / SERIAL # (B)(4)), THREE PEG PATELLA 35MM (CAT # 200-02-35 / SERIAL # (B)(4)), CC POSTERIOR FEM AUGMENT SZ 2, 5MM (CAT # 208-07-02 / SERIAL # (B)(4)), FLUTED STEM EXTENSION 80L X 14 MM (CAT# 204-34-08 / SERIAL # (B)(4)), CC FEMORAL SZ 2 (CAT# 208-01-02 / SERIAL # (B)(4)), CC DISTAL FEM AUGMENT SZ 2, 10MM (CAT # 208-06-02 / SERIAL # (B)(4)), CC STEM EXT ADAPTOR 5 DEGREE (CAT # 208-09-05 / SERIAL # (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00008 | 2022-01-05 | 2016-05-02 | I, n, j, u, r, y | Manufacturer report | PS TIBIAL INSERTS SZ 3, 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | PS TIBIAL INSERTS SZ 3, 13MM | 204-23-13 |  | AS REPORTED, APPROXIMATELY SEVEN YEARS POST TKA, PATIENT HAD A PARTIAL TIBIAL REVISION FOR AN UNKNOWN REASON. THIS IS NOT A PE LINER ISSUE DUE TO NON-CONFORMING EVOH BAGS, THERE IS NO INFORMATION RECEIVED ANY INFO OF INCIDENT AT THAT TIME, THERE IS NO COMPLAINT REGISTERED ON OUR SIDE. NO ADDITIONAL INFORMATION AVAILABLE. | CONCOMITANT MEDICAL PRODUCTS: TRAPEZOID TIBIAL TRAY SZ 3F/2T (CAT# 204-04-32 / SERIAL# (B)(4)); TIBIAL AUGMENT BLOCK 1/2-SZ 2 5MM (CAT# 204-62-05 / SERIAL# (B)(4)); STEM EXTENSION 80L X12 MM (CAT# 204-32-08 / SERIAL# (B)(4)); TIBIAL AUGMENT BLOCK 1/2-SZ 2 5MM (CAT# 204-62-05 / SERIAL# (B)(4)); THREE PEG PATELLA 35MM (CAT# 200-02-35 / SERIAL# (B)(4)); TIBIAL STEM EXT. SCREW (CAT# 204-70-00 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00006 | 2022-01-05 | 2021-12-06 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 42MM GLENOSPHERE | EXACTECH, INC. | 320-01-42 | 320-01-42 |  | AS REPORTED, APPROXIMATELY 3 YEARS POST OP THE INITIAL LEFT TSA, THIS (B)(6) MALE PATIENT WAS REVISED DUE TO DISLOCATION. PATIENT WOKE UP IN SEVERE PAIN. THE CASE REPORT FORM INDICATES THIS EVENT IS DEFINITELY RELATED TO DEVICES AND DEFINITELY RELATED TO THE PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | CONCOMITANT MEDICAL PRODUCTS: 300-01-09 - EQUINOXE, HUMERAL STEM PRIMARY, PRESS FIT 9MM; 320-42-13 - EQUINOXE REVERSE 42MM HUMERAL CONST LINER +2.5; 320-10-00 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0; 320-15-01 - EQ REV GLENOID PLATE. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00005 | 2022-01-05 | 2021-12-15 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 42MM HUMERAL CONST LINER +2.5 | EXACTECH, INC. | 320-42-13 | 320-42-13 |  | AS REPORTED, APPROXIMATELY 3 YEARS POST OP REVERSE LEFT TSA, PATIENT'S POLY DISENGAGED AND THE SURGEON HAD TO REVISE. WHEN THE PATIENT WAS OPENED UP, THE POLY WAS FULLY DISENGAGED AND THERE WAS SOME WEAR ON THE POLY. THE IMPLANT WAS A 42 +2.5MM CONSTRAINED LINER. THERE WAS NO SURGICAL DELAY OR PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. | CONCOMITANT MEDICAL PRODUCTS: 320-01-42, 78517035 - EQUINOXE REVERSE 42MM GLENOSPHERE; 320-10-00 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00007 | 2022-01-05 | 2018-04-16 | I, n, j, u, r, y | Manufacturer report | PS TIBIAL INSERTS SZ 4, 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | PS TIBIAL INSERTS SZ 4, 11MM | 204-24-11 |  | AS REPORTED, PATIENT HAD A REVISION FOR AN UNKNOWN REASON. THIS IS NO PE LINER ISSUE DUE TO NON-CONFORMING EVOH BAGS. NO INCIDENT REPORTED TO EXACTECH (B)(4), NO MATERIOVIGILANCE KNOWN BY EXACTECH CORP OR EXACTECH (B)(4). NO ADDITIONAL INFORMATION AVAILABLE. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00011 | 2022-01-06 | 2021-12-20 | I, n, j, u, r, y | Manufacturer report | CR TIBIAL INSERT SZ4, 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | CR TIBIAL INSERT SZ4, 13MM | 200-24-13 |  | AS REPORTED, APPROXIMATELY NINE YEARS POST RIGHT TKA, DURING REMOVAL/WITHDRAWAL, AN (B)(6) MALE PATIENT HAD SIGNS OF POLY WEAR AND WAS TAKEN TO SURGERY FOR POLY EXCHANGE. THE POLY WAS REMOVED, FEMUR AND TIBIA WERE BOTH CHECKED AND FOUND TO BE LOOSE. THE PATIENT WAS REVISED WITH A TRULIANT CC KNEE SYSTEM. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THERE WAS NO BREAKAGE OF THE DEVICE. THERE WAS NO SURGICAL DELAY/PROLONGATION. THE DEVICES WILL BE RETURN. | (H3) THE REVISION REPORTED WAS LIKELY THE RESULT OF BOTH PROSTHESIS WEAR AND LOOSENING OF THE FEMORAL AND TIBIAL COMPONENTS. THE PROSTHESIS WEAR WAS LIKELY THE COMBINATION OF MALALIGNMENT BETWEEN THE FEMORAL AND TIBIAL COMPONENTS, THIRD BODY WEAR, AND PATIENT-RELATED CONDITIONS. THE ASEPTIC (NON-INFECTED) FEMORAL AND TIBIAL LOOSENING WAS LIKELY THE RESULT OF AN INSUFFICIENT BOND BETWEEN THE IMPLANTS AND THE BONES. THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE REPORTED EVENT OF "PROSTHESIS WEAR¿ IS ASSOCIATED WITH MATERIAL DAMAGE TO A SURFACE, USUALLY INVOLVING PROGRESSIVE LOSS OR DISPLACEMENT OF MATERIAL, DUE TO RELATIVE MOTION BETWEEN THAT SURFACE AND A CONTACTING SUBSTANCE OR SUBSTANCES. FURTHERMORE, THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE REPORTED EVENT OF "LOOSEINING ¿ FEMORAL¿ IS ASSOCIATED WITH WEAKENED INTEGRATION OF THE FEMORAL COMPONENT AT THE BONE-IMPLANT INTERFACE DUE TO LOSS OF FIBROUS AND/OR BONY TISSUE AND LEADING TO COMPROMISED ANCHORAGE OF THE DEVICE. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00012 | 2022-01-06 | 2021-12-13 | I, n, j, u, r, y | Manufacturer report | NV GXL LINER NEUTRAL, 32MM ID GROUP 1 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LINER NEUTRAL, 32MM ID GROUP 1 CUPS | 130-32-51 |  | AS REPORTED, APPROXIMATELY SIX YEARS POST RIGHT THA, A FEMALE PATIENT WAS REVISED FOR POLY WEAR AND TAKEN BACK FOR A POLY LINER EXCHANGE. THE FEMORAL HEAD WAS REMOVED, STEM WAS WELL FIXED, THE LINER WAS REMOVED, AND CUP FOUND TO BE WELL FIXED. A NEW G1 XLE LIPPED LINER WAS IMPLANTED AS WELL AS A NEW BIOLOX HEAD. THERE WAS NO BREAKAGE TO THE DEVICE. THERE WAS NO SURGICAL DELAY OR PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICES ARE NOT AVAILABLE FOR EVALUATION DUE TO THE HOSPITAL DISCARDED THE IMPLANTS. | CONCOMITANT MEDICAL PRODUCTS: COCR FEM HEAD 32MM -3.5 OFFSET 12/14 (CAT# 142-32-93 / SERIAL# (B)(4)) ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00013 | 2022-01-06 | 2021-10-04 | I, n, j, u, r, y | Manufacturer report | ACETABULAR COMPONENTS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | ACETABULAR COMPONENTS | UNKNOWN |  | AS REPORTED, APPROXIMATELY TWENTY YEARS POST THA, A MALE PATIENT POLLY WORE OUT OVER TIME AND NEEDED TO BE REPLACED. CUP WAS WELL FIXED, AND THE POLLY WAS EXCHANGED. THE PATIENT WAS REVISED TO A NEUTRAL POLLY. THERE WAS NO BREAKAGE TO THE DEVICE. THERE WAS NO SURGICAL DELAY OR PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICES ARE NOT AVAILABLE FOR EVALUATION DUE TO HOSPITAL WOULD NOT ALLOW TAKE THE IMPLANTS. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00014 | 2022-01-06 | 2021-12-01 | I, n, j, u, r, y | Manufacturer report | LOGIC FEMORAL PS CEM RIGHT SZ 4 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | N/A | 02-010-01-0340 |  | AS REPORTED, PATIENT WAS IMPLANTED A TOTAL PROSTHESIS OF THE RIGHT KNEE IN (B)(6) 2018 BY SURGEON FROM (B)(6). BARELY 3 YEARS LATER, PATIENT WAS FORCED TO PROCEED WITH A COMPLETE REPLACEMENT OF THE MATERIAL, WHICH EITHER METAL OR POLYETHYLENE CAUSES BONE DEGENERATION ALL AROUND THE TIBIAL IMPLANT. FOR THIS DELICATE OPERATION, PATIENT CONSULTED ANOTHER SURGEON FROM UCL; SURGEON INFORMED PATIENT THAT THIS IS NOT THE FIRST CASE THAT HE MUST RE-OPERATE AND THAT IN THE UNITED STATES OF AMERICA (USA) MANY CASES HAVE BEEN REPORTED. PATIENT IS PARTICULARLY WORRIED ESPECIALLY SINCE THE SAME TYPE OF PROSTHESIS WAS PLACED ON THE LEFT LEG LAST YEAR; THE DAMAGE IS CONSEQUENT SINCE PATIENT HAVE TO UNDERGO A REVISION OPERATION OF PROSTHESIS WITH MONTHS OF REVALIDATION AFTERWARDS AND OF INCAPACITY WORKING AT THE SAME TIME, WITHOUT KNOWING IF THE SAME MISHAP AWAITS PATIENT FOR THE LEFT KNEE. NO ADDITIONAL INFORMATION IS AVAILABLE. | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION RECEIVED THE FOLLOWING SECTIONS A3, B5, B7, D1, D4, D6A, D6B, D10. G3, G4, G6, H2, H4 AND H6 HAVE BEEN UPDATED ACCORDINGLY. SECTION B5: ADDITIONAL INFORMATION RECEIVED INDICATES THAT PATIENT REPORTED THAT HE HAD A REVISION SURGERY ON (B)(6), 2022, OF AN OPTETRAK LOGIC PROSTHESIS IMPLANTED ON (B)(6), 2018. PATIENT EXPERIENCED PATELLA FRACTURE RESULTING IN ASEPTIC NECROSIS ASSOCIATED WITH PERI-PROSTHETIC OSTEOLYSIS PREDOMINANT ON THE TIBIAL SIDE, TWO-PIECE FRACTURE OF THE TIBIAL "PLATEAU" IN POLYETHYLENE WITH DISSEMINATION OF MANY PIECES IN THE KNEE, LOOSENING OF THE MEDALLION PATELLAR ALL CAUSING MECHANICAL LOOSENING OF THE PROSTHESIS. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO SURGEON AND PATIENT NOT WILLING TO SHARE. SECTION D10: CONCOMITANT MEDICAL PRODUCTS LOGIC TIBIA PS RBK INSRT SZ 4 9MM (CAT# 02-012-38-4009 / SERIAL# (B)(6)) LOGIC TIBIA RBKTRAY CEM SZ 4F/ 4T (CAT# 02-012-43-4040 / SERIAL# (B)(6)) THREE PEG PATELLA 38MM (CAT# 200-02-38 / SERIAL# (B)(6)) ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT SECTIONH1; TYPE OF REPORTABLE EVENT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00015 | 2022-01-06 | 2021-12-15 | I, n, j, u, r, y | Manufacturer report | TRULIANT PS CEM FEM PS CEM LEFT SZ 4.5 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT PS CEM FEM PS CEM LEFT SZ 4.5 | 02-020-11-0245 |  | AS REPORTED, A MALE PATIENT IS GOING TO HAVE A REVISION FOR FAILED IMPLANTS AND THE ORIGINAL SURGERY WAS BACK IN 2019. PATIENT REVISION CONSIST OF FEMUR, CEMENTED FEMORAL STEM AND A POLY LINER. THE INDICATION FOR INITIAL TKA IS TO END-STAGE OSTEOARTHRITIS. THE REVISION IS SCHEDULED TENTATIVE IN (B)(6) 2022 DUE TO LEFT KNEE FEMORAL EFFUSION BETWEEN BONE AND IMPLANT. PATIENT¿S PERFORMS ADL¿S INDEPENDENTLY BUT WITH PAIN AND SWELLING. NO ADDITIONAL INFORMATION AVAILABLE. | CONCOMITANT MEDICAL PRODUCTS TRULIANT PS CEM FEM PS CEM LEFT SZ 4.5 (CAT# 02-020-11-0245 / SERIAL# (B)(4)), TRULIANT TIB FIT TRAY CEM SZ 4.5F/3.5T (CAT# 02-022-45-4535 / SERIAL# (B)(4)), ADVANCED PATELLA 35MM 3 PEG IMPLANT (CAT# 200-07-35 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00021 | 2022-01-07 | 2019-07-23 | I, n, j, u, r, y | Manufacturer report | EXACTECH | A+ STERN INSERTER | EXACTECH, INC. |  | 167-00-00 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE INSTRUMENTATION. THE TIP BROKE OFF THE STEM INSERTER WHILE IMPACTING THE STEM. THE PIECE WAS RETRIEVED BUT DISCARDED. NO DELAY TO SURGERY, NO ADVERSE EFFECTS TO THE PATIENT. PATIENT LEFT THE ROOM STABLE. THE REPRESENTATIVE WAS PRESENT AT THE TIME OF SURGERY. THE DEVICE WAS RECEIVED FOR ANALYSIS. ATTEMPTS WERE MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REPRESENTATIVE, BUT NO RESPONSES WERE RECEIVED. THERE HAS BEEN NO OTHER INFORMATION PROVIDED BY THE CONTACTS FOR THIS EVENT. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. THE TIP BROKE OFF THE STEM INSERTER WHILE IMPACTING THE STEM; THE PIECE WAS RETRIEVED BUT DISCARDED. AN INVESTIGATION WAS CONDUCTED: THE FRACTURED STEM INSERTER REPORTED WAS LIKELY THE RESULT OF LONG-TERM, RIGOROUS USE, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE STEM INSERTER TIP. | (B)(4). THE DEVICE WAS AVAILABLE FOR ANALYSIS. ENGINEERING ANALYSIS COMPLETED BY (B)(4) ON 9/6/2019. DESIGN-RELATED ISSUES: THE DESIGN OF THE NOVATION A+ STEM INSERTER HAS BEEN IN THE FIELD SINCE 2009. EXACTECH IS AWARE OF 2 OTHER SIMILAR COMPLAINT REPORTS INVOLVING NOVATION A+ STEM INSERTERS DUE TO FRACTURED TIPS SINCE 2009. BECAUSE THE NOVATION A+ STEM INSERTER IS USED IN TOTAL HIP SURGERIES, SALES DATA FOR FEMORAL STEMS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF (B)(4) UNITS, WHICH HAVE BEEN IN THE FIELD SINCE 2011. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS NOVATION A+ STEM INSERTER, (B)(4) REV C, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 9. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR. THE FRACTURED STEM INSERTER REPORTED IN (B)(4) WAS LIKELY THE RESULT OF LONG-TERM, RIGOROUS USE, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE STEM INSERTER TIP. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. THE TIP BROKE OFF THE STEM INSERTER WHILE IMPACTING THE STEM; THE PIECE WAS RETRIEVED BUT DISCARDED. AN INVESTIGATION WAS CONDUCTED: THE FRACTURED STEM INSERTER REPORTED WAS LIKELY THE RESULT OF LONG-TERM, RIGOROUS USE, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE STEM INSERTER TIP. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00016 | 2022-01-07 | 2021-12-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EQUINOXE | 3.2MM K-WIRE, THREADED | EXACTECH, INC. | 321-52-08 | 321-52-08 |  | IT WAS REPORTED THAT DURING SURGICAL USE THE K-WIRE WAS IN THE GLENOID AND WHEN IT WAS TIME TO USE THE MINI BASEPLATE DRILL TO DRILL THE CENTER CAGE, THE K-WIRE TIP STARTED TO BREAK APART. THE TIP WAS THEN STUCK IN THE DRILL. BOTH DEVICES WERE REMOVED AND NEW ONES WERE USED TO COMPLETE THE CASE. NO PATIENT INFORMATION OR MEDICAL HISTORY REPORTED. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. | PENDING EVALUATION. CONCOMITANT DEVICE(S): 315-52-66, 292002003 - MOD CANNULATED DRILL CAGE, SMALL REVERSE. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00017 | 2022-01-07 | 2021-12-16 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 42MM HUMERAL LINER +0 | EXACTECH, INC. | 320-42-00 | 320-42-00 |  | AS REPORTED, APPROXIMATELY 10 MONTHS THE INITIAL RTSA, THIS (B)(6) MALE PATIENT WAS REVISED DUE TO DISLOCATION. THE PATIENT WAS MOVING A HEAVY OBJECT WITH HIS RIGHT ARM AND FELT WEIGHT OF OBJECT PULL SHOULDER OUT OF JOINT. PATIENT HAS HAD MULTIPLE PREVIOUS INSTANCES OF INSTABILITY. THE CASE REPORT FORM INDICATES THIS EVENT IS DEFINITELY RELATED TO DEVICES AND PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | SECTION H10: (H3) BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN OR MANUFACTURING ISSUES. THE CAUSE OF THE DISLOCATION AND SUBSEQUENT CLOSED REDUCTION CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION AND THE RESULT OF LOOSENED SUPPORTING LIGAMENTS AND MUSCLES, WHICH ALLOWED FOR DISLOCATION OCCURRING DUE TO THE SIGNIFICANT LIGAMENT RECONSTRUCTION. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00018 | 2022-01-07 | 2021-11-06 | I, n, j, u, r, y | Manufacturer report | EQ REV LOCKING SCREW | REVERSE SHOULDER COMPONENTS | EXACTECH, INC. |  | 320-15-05 |  | IT WAS REPORTED THAT THIS FEMALE PATIENT RECEIVED A REVISION FOR UNKNOWN REASONS. NO OTHER PATIENT INFORMATION OR MEDICAL HISTORY IS AVAILABLE. | THIS FOLLOW-UP REPORT IS BEING SUBMITTED TO RELAY ADDITIONAL AND/OR CORRECTED INFORMATION. THE FOLLOWING SECTIONS WERE UPDATED: D6A, G3/G4, H6. THE FOLLOWING SECTIONS WERE CORRECTED: B3, H6. ADDED 510K NUMBER. MDR SECTION CODES UPDATED/CORRECTED: A, B, C, D, E, G. THE REVISION REPORTED WAS LIKELY THE RESULT OF NOT FULLY SEATING/ENGAGING THE GLENOSPHERE LOCKING SCREW AT THE TIME OF IMPLANTATION, POSSIBLY DUE TO MISASSEMBLY OF THE GLENOSPHERE ON THE GLENOID PLATE, WHICH ALLOWED FOR THE SCREW TO FRACTURE AND THE COMPONENTS TO DISASSEMBLE. HOWEVER, THE EXACT CAUSE AND ORDER OF EVENTS CANNOT BE CONFIRMED BECAUSE THE REVISED COMPONENTS WERE NOT RETURNED FOR EVALUATION AND NO IMAGES OR RADIOGRAPHS WERE PROVIDED. IF ANY FURTHER INFORMATION IS OBTAINED THAT WOULD CHANGE OR ALTER ANY INFORMATION PROVIDED, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00019 | 2022-01-07 | 2019-06-13 | I, n, j, u, r, y | Manufacturer report | EXACTECH | HUMERAL TRAY TRIAL +0 | EXACTECH, INC. |  | 321-13-00 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE HUMERAL TRAY TRIAL. AFTER TRAILING THE ERGO +0 TRIAL HUMERAL TRAY, THE SURGEON USED THE SCREWDRIVER TO REMOVE IT FROM THE STEM. THE TRIAL TRAYS SCREW CAME OUT OF THE TRAY AND THE CLIP THAT WAS SUPPOSED TO HOLD THE SCREW IN THE TRAY CAME OUT OF THE TRAY AS WELL. ALL PIECES WERE SEEN AND REMOVED FROM THE WOUND. THERE WAS NO DELAY TO SURGERY AND NO PATIENT ADVERSE EVENT. THE PATIENT LEFT THE OR STABLE. INACTIVE AGENCY WITH NO HOSPITAL/OTHER CONTACT INFORMATION FOR THE EVENT. NO ADDITIONAL INFORMATION. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO MANUFACTURING ISSUES. THE TRIAL TRAYS SCREW CAME OUT OF THE TRAY AND THE CLIP THAT WAS SUPPOSED TO HOLD THE SCREW IN THE TRAY CAME OUT OF THE TRAY AS WELL; ALL PIECES WERE SEEN AND REMOVED FROM THE WOUND. THERE WAS NO PATIENT ADVERSE EVENT. AS DETERMINED BY (B)(4), THE DISASSEMBLED SUBCOMPONENTS REPORTED WERE LIKELY THE RESULT OF NOT ADEQUATELY DEFINING THE DESIGN INPUT AND ACCEPTANCE CRITERIA FOR THE SCREW RETAINING FEATURE RESULTING IN NOT ENOUGH INTERFERENCE BETWEEN THE CLIP AND THE GROOVE. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE AVAILABLE FOR ANALYSIS. ENGINEERING EVAL COMPLETED BY NF ON (B)(6) 2020. (B)(4), CRC2020-01-31-03 A REVIEW OF DESIGN PRINTS, DESIGN VERIFICATION, AND DESIGN VALIDATION, PER (B)(4), IDENTIFIED THE FOLLOWING: THE SCREW / RETAINING CLIP INTERACTION WAS NOT INCLUDED AS PART OF TR-2019-0131- EQUINOXE ERGO TOLERANCE ANALYSIS REPORT. DURABILITY TESTING FOR THE SCREW WAS NOT PERFORMED PER TR-2019-0319-EVALUATION OF THE ERGO HUMERAL TRAY TRIAL LOCKING MECHANISM. THE ERGO PILOT RMR REV ¿ DID NOT EXPLICITLY DEFINE THE FAILURE MODE OF THE SCREW RETAINING CLIP COMING OUT. THE DEVICE HISTORY RECORD (DHR) WAS REVIEWED AS PART OF (B)(4), AND ALL PARTS WERE ACCEPTED WITH CONFORMANCE TO THE DEVICE SPECIFICATIONS. A CORRECTION AND REMOVAL COMMITTEE (CRC) MEETING WAS HELD TO ADDRESS THE NEED FOR FIELD ACTION. THE COMMITTEE DETERMINED THAT DUE TO THE LOW RISK TO HEALTH AND LOW RATE OF OCCURRENCE, NO FIELD ACTION IS NECESSARY AT THIS TIME. SEE (B)(4) FOR DETAILS. THE SYSTEMATIC ISSUES REGARDING DESIGN CONTROLS NOT BEING FOLLOWED ARE BEING FURTHER ADDRESSED AS PART OF (B)(4). THE DESIGN INPUT DOCUMENT WAS REVISED TO INCLUDE ACCEPTANCE CRITERIA FOR THE FEATURE OR COMPONENT TO RETAIN THE SCREW AND DURABILITY ACCEPTANCE CRITERIA. (750-2016-002 DIR REV A) THE RISK DOCUMENTATION WAS REVISED TO SPECIFY THE FAILURE MODE OF THE RETAINING RING COMING OUT. (750-2016-002-RMR REV A) NEW PART NUMBERS WERE CREATED, AND INSTRUMENT DESIGN/DRAWING WERE REVISED. (321-114-00 REV A, 321-114-05 REV A, 321-114-10 REV A) NEW PARTS WERE ORDERED AND RECEIVED, AND A CADAVER VALIDATION LAB WAS COMPLETED BY SURGEON USERS. VERIFICATION AND VALIDATION ON THE NEW DESIGN CONSIDERING FORESEEABLE MISUSE WILL BE PERFORMED, AND DESIGN TRANSFER AND FINAL DESIGN REVIEW WILL BE CONDUCTED. (B)(4) IS STILL OPEN AT THE TIME OF THIS INVESTIGATION, AND CORRECTIVE ACTIONS ARE SUBJECT TO CHANGE. REFER TO (B)(4) FOR THE MOST UP TO DATE INFORMATION. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: AS DETERMINED BY (B)(4), THE DISASSEMBLED SUBCOMPONENTS REPORTED IN CASE (B)(4) WERE LIKELY THE RESULT OF NOT ADEQUATELY DEFINING THE DESIGN INPUT AND ACCEPTANCE CRITERIA FOR THE SCREW RETAINING FEATURE RESULTING IN NOT ENOUGH INTERFERENCE BETWEEN THE CLIP AND THE GROOVE. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. ¿ CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO MANUFACTURING ISSUES. THE TRIAL TRAYS SCREW CAME OUT OF THE TRAY AND THE CLIP THAT WAS SUPPOSED TO HOLD THE SCREW IN THE TRAY CAME OUT OF THE TRAY AS WELL; ALL PIECES WERE SEEN AND REMOVED FROM THE WOUND. THERE WAS NO PATIENT ADVERSE EVENT. AS DETERMINED BY (B)(4), THE DISASSEMBLED SUBCOMPONENTS REPORTED WERE LIKELY THE RESULT OF NOT ADEQUATELY DEFINING THE DESIGN INPUT AND ACCEPTANCE CRITERIA FOR THE SCREW RETAINING FEATURE RESULTING IN NOT ENOUGH INTERFERENCE BETWEEN THE CLIP AND THE GROOVE. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00020 | 2022-01-07 | 2019-06-05 | I, n, j, u, r, y | Manufacturer report | EXACTECH | TRULIANT PIN PULLER SYRINGE STYLE | EXACTECH, INC. |  | 02-029-90-4100 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE INSTRUMENTATION. THE SURGEON WAS PUTTING IN THE PROXIMAL TIBIA AND THE PIN PULLER ¿SNAPPED.¿ ONE OF THE PRONGS CAME OFF AND WAS FOUND ON THE FLOOR. THIS DID CAUSE A 20-MINUTE DELAY TO CASE WITH NO ADVERSE EFFECT TO PATIENT. THE REP WAS PRESENT AT TIME OF SURGERY. THE PATIENT WAS STABLE LEAVING THE OPERATING ROOM. THE DEVICE WAS RECEIVED FOR ANALYSIS. NEW INFORMATION FROM THE REPRESENTATIVE ON (B)(6) 2021: NO INFORMATION ON PATIENT DEMOGRAPHICS. A PART/PIECE FELL INT THE SURGICAL SITE AND WAS RECOVERED BY THE SURGEON. THE STAFF TOOK 10 MINUTES TO GET X-RAY TO ENSURE ALL PIECES WERE RECOVERED. THERE WAS NO INJURY/CLINICAL CONSEQUENCE/ADVERSE EVENT TO THE PATIENT. THERE IS NO OTHER INFORMATION ABOUT THE DEVICE/EVENT/PATIENT FROM THE CONTACTS. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT INJURY CAUSED BY THE REMOVAL OF THE DEVICE FRAGMENTS FROM THE SURGICAL SITE. THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF THE PRONGS OF THE PIN PULLER EXPERIENCING EXCESSIVE FORCES DURING THE EXTRACTION OF BONE PINS. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS AVAILABLE FOR ANALYSIS. ENGINEERING ANALYSIS COMPLETED BY NF ON (B)(6) 2019. C2018-483, CRC2018-09-19-01, (B)(4) FINDINGS: ONE OR MORE OF THE METAL PRONGS USED TO CLAMP A PIN HEAD ARE MISSING FROM THE RECEIVED TRULIANT SYRINGE PIN PULLERS. WHEN ANALYZING THE FRACTURE SURFACE AS PART OF C2018-483, THE FAILURE APPEARS CONSISTENT WITH BRITTLE FRACTURE. THIS TYPE OF FAILURE IS MOST LIKELY THE RESULT OF THE PRONGS EXPERIENCING EXCESSIVE FORCES. FURTHERMORE, IF THE PIN PULLER WAS USED OFF-AXIS RELATIVE TO THE BONE PIN, AN UNEVEN FORCE DISTRIBUTION MAY HAVE BEEN APPLIED TO THE METAL PRONGS AND COULD EXPLAIN WHY THEY FRACTURED FROM ONLY ONE SIDE. OVER TIME, THESE EXCESSIVE OFF-AXIS FORCES COULD POTENTIALLY LEAD TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE MATERIAL. FURTHER INVESTIGATION WILL OCCUR AS PART OF (B)(4). MOST LIKELY UNDERLYING CAUSE: THE BROKEN INSTRUMENT REPORTED IN CASE (B)(4) WAS LIKELY THE RESULT OF THE PRONGS OF THE PIN PULLER EXPERIENCING EXCESSIVE FORCES DURING THE EXTRACTION OF BONE PINS. RMR REVIEW: A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THESE TRULIANT SYRINGE PIN PULLERS, (B)(4), WAS REVIEWED. THE RISK IS CAPTURED IN LINE 11. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT INJURY CAUSED BY THE REMOVAL OF THE DEVICE FRAGMENTS FROM THE SURGICAL SITE. THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF THE PRONGS OF THE PIN PULLER EXPERIENCING EXCESSIVE FORCES DURING THE EXTRACTION OF BONE PINS. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00022 | 2022-01-07 | 2019-12-13 | I, n, j, u, r, y | Manufacturer report | EXACTECH | TRULIANT PS CEM FEM PS CEM RIGHT SZ 3 | EXACTECH, INC. |  | 02-020-11-0330 |  | IT WAS REPORTED FROM THE US THAT THE PATIENT¿S INITIAL RIGHT TKA PROCEDURE WAS DONE ON (B)(6)2018. ON (B)(6)2019 A POLY SWAP WAS PERFORMED CAPTURED IN MFR 1038671-2019-00628. ON (B)(6)2019 THE SURGEON PLANNED TO REVISE THE FEMUR TO ADDRESS A FLEXION CONTRACTURE. HE BEGAN BY USING ¼ INCH OSTEOTOME UNDER FEMORAL FLANGE AND THE FEMUR CAME ENTIRELY LOOSE. NO CEMENT ADHESION TO THE FEMORAL COMPONENT. HE REMOVED AN ADDITIONAL 4MM OF DISTAL FEMUR TO ADDRESS FLEXION CONTRACTURE AND RE-IMPLANTED A SIZE 3 PS FEMUR WITH A 9MM POLY. THE DEVICE WAS RECEIVED ON 12/17/2019. ATTEMPT WAS MADE FOR ADDITIONAL INFORMATION FROM THE REPRESENTATIVE AND A RESPONSE WAS RECEIVED. PER THE REPRESENTATIVE, NO ADDITIONAL INFORMATION IS AVAILABLE ABOUT THIS EVENT. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THE PATIENT HAD A SURGICAL REVISION. THE REVISION REPORTED WAS LIKELY THE RESULT OF THE PATIENT¿S FLEXION CONTRACTURE. THE LOOSENING OF THE FEMORAL COMPONENT DISCOVERED DURING THE REVISION WAS LIKELY THE RESULT OF AN INSUFFICIENT BOND BETWEEN THE BONE CEMENT AND THE FEMORAL COMPONENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS AVAILABLE FOR ANALYSIS. ENGINEERING ANALYSIS COMPLETED BY AD ON 9/9/2020. ACCORDING TO THE INFORMATION PROVIDED, APPROXIMATELY 12 MONTHS AFTER THE INITIAL TOTAL KNEE REPLACEMENT, A SWAP OF THE POLYETHYLENE INSERT WAS COMPLETED. THIS EVENT WAS NOT REPORTED TO EXACTECH WHEN IT OCCURRED BUT HAS SINCE BEEN CAPTURED IN (B)(4). EIGHT MONTHS AFTER THE POLYETHYLENE SWAP, THE SURGEON PLANNED TO REVISE THE FEMUR TO ADDRESS A FLEXION CONTRACTURE. THE SURGEON BEGAN BY USING ¼ INCH OSTEOTOME UNDER FEMORAL FLANGE AND THE FEMUR REPORTEDLY CAME ENTIRELY LOOSE. THERE WAS NO CEMENT ADHESION TO THE FEMORAL COMPONENT. THE SURGEON REMOVED AN ADDITIONAL 4MM OF DISTAL FEMUR TO ADDRESS FLEXION CONTRACTURE AND IMPLANTED A SIZE 3 PS FEMORAL COMPONENT WITH A 9MM INSERT. DESIGN-RELATED ISSUES: AS OF 08/04/2020, EXACTECH HAS RECEIVED 7 COMPLAINT REPORTS INVOLVING LOOSENING OF A TRULIANT FEMORAL COMPONENT SINCE ITS INTRODUCTION IN 2017. THE SALES DATA FOR ALL TRULIANT FEMORAL COMPONENTS WAS USED TO CALCULATE A COMPLAINT OCCURRENCE RATE OF (B)(4). THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF RECEIVING ANY OTHER COMPLAINT REPORTS INVOLVING A PART FROM THIS PACKAGING LOT OF (B)(4) PIECES, WHICH CONSISTED OF TWO (B)(4) MANUFACTURING LOTS. 5 OF THE (B)(4) PIECES REMAIN IN INVENTORY AND, THEREFORE, HAVE NOT YET BEEN IMPLANTED. A REVIEW OF THE DEVICE HISTORY RECORD SHOWED THAT THE NAMED DEVICE WAS ACCEPTED WITH CONFORMANCE TO THE DEVICE REQUIREMENTS. THEREFORE, THIS DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) AND RISK ASSESSMENT AND CONTROLS REPORT (RACR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RACR FOR THIS TRULIANT FEMORAL COMPONENT, (B)(4) WAS REVIEWED. THE RISK IS CAPTURED IN RISK ID 7.8. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE IS ¿VERY LOW¿, LOOSENING AND ¿FLEXION CONTRACTURE/VARUS VALGUS DEFORMITY¿ ARE BOTH LISTED IN THE PRODUCT LABELING UNDER DEVICE SPECIFIC RISKS (IFU 700-096-160 REV C), AND THE RISK IS CAPTURED IN THE RACR. THE REVISION REPORTED IN (B)(4)WAS LIKELY THE RESULT OF THE PATIENT¿S FLEXION CONTRACTURE. THE LOOSENING OF THE FEMORAL COMPONENT DISCOVERED DURING THE REVISION WAS LIKELY THE RESULT OF AN INSUFFICIENT BOND BETWEEN THE BONE CEMENT AND THE FEMORAL COMPONENT. PROPERLY FUNCTIONING IMPLANTS DEPEND ON THEIR APPROPRIATE FIXATION TO THE BONE; FIXATION IS USUALLY ACHIEVED BY CEMENTING THE IMPLANT ONTO THE BONE. ALTHOUGH IMPLANTS ARE FIRMLY FIXED AT THE INITIAL KNEE REPLACEMENT SURGERY, THEY MAY BECOME LOOSE OVER TIME. FRICTION CAUSED BY THE JOINT SURFACES RUBBING AGAINST EACH OTHER WEARS AWAY THE SURFACES OF THE IMPLANT, CREATING TINY PARTICLES THAT ACCUMULATE AROUND THE JOINT. IN A PROCESS CALLED ASEPTIC (NON-INFECTED) LOOSENING, THE BOND OF THE IMPLANT TO THE BONE IS DESTROYED BY THE BODY'S ATTEMPT TO DIGEST THE WEAR PARTICLES. DURING THIS PROCESS, NORMAL BONE IS ALSO DIGESTED (A CONDITION CALLED OSTEOLYSIS), WHICH CAN WEAKEN OR EVEN FRACTURE THE BONE. WHEN THE PROSTHESIS BECOMES LOOSE, THE PATIENT MAY EXPERIENCE PAIN, CHANGE IN ALIGNMENT, OR INSTABILITY. ASEPTIC LOOSENING IS THE MOST COMMON MODE OF FAILURE OF KNEE IMPLANTS. REFERENCE: SCUDERI GR, KOCHHAR T: MANAGEMENT OF FLEXION CONTRACTURE IN TOTAL KNEE ARTHROPLASTY. J ARTHROPLASTY. 2007, 22 (4 SUPPL 1): 20-4. 10.1016/J.ARTH.2006.12.110. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THE PATIENT HAD A SURGICAL REVISION. THE REVISION REPORTED WAS LIKELY THE RESULT OF THE PATIENT¿S FLEXION CONTRACTURE. THE LOOSENING OF THE FEMORAL COMPONENT DISCOVERED DURING THE REVISION WAS LIKELY THE RESULT OF AN INSUFFICIENT BOND BETWEEN THE BONE CEMENT AND THE FEMORAL COMPONENT. | Y | N |  |  | 2022-01
1038671-2022-00025 | 2022-01-10 | 2020-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRUL POROUS BONE SCREW DRILL GUIDE | EXACTECH, INC. |  | 02-029-29-4000 |  | IT WAS REPORTED FROM THE US THAT ON (B)(6) 2020, THE SURGEON WAS PARTICIPATING IN A WEBEX MEETING WITH AN EXACTECH PROJECT ENGINEER. DURING THE WEBEX, THE SURGEON MENTIONED THAT IN AT LEAST ONE CASE HE HAS HAD ISSUES WITH THE TRULIANT POROUS TIBIA PRE-DRILL GUIDE (02-029-29-4000) NOT MATING CORRECTLY WITH THE TRULIANT POROUS TIBIAL TRAY (02-022-55-XXXX). HE DID NOT MENTION A SPECIFIC INSTANCE, ONLY THAT IT HAD OCCURRED, SPECIFICALLY WITH THE ¿UPSIZE¿ TIBIAL TRAYS, WHERE THE FEMORAL SIZE IS SMALLER THAN THE TIBIAL SIZE. FOR EXAMPLE, A 2F/3T TRAY. HE MENTIONED IF THIS OCCURS, THE DRILLED HOLE IN THE BONE IS NOT CONCENTRIC WITH THE SCREW HOLE IN THE TIBIAL TRAY. HE WOULD THEN NEED TO RE-DRILL TO CENTER THE HOLE. THIS CREATED A SMALL SURGICAL DELAY (ABOUT 10 SECONDS) TO RE-DRILL. IT DID NOT LEAD TO ANY INJURY. THE DRILL GUIDE SITS IN THE SCREW HOLE, MEANING THE DRILL GUIDE SITS CLOSER TO THE PERIPHERAL RAIL. IF THE DRILL GUIDE IS USED ON CERTAIN TRAYS WHERE THE FEMORAL SIZE IS SMALLER THAN THE TIBIAL SIZE, IT WILL INTERFERE. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE SURGEON DOES NOT REMEMBER ANYTHING ELSE ABOUT THE EVENT OR PATIENT. HE HAS NOT SEEN THIS TYPE OF ISSUE SINCE. NO ADDITIONAL INFORMATION RELATED TO THE EVENT HAS BEEN PROVIDED. DEVICES USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. THE DRILL GUIDE SITS IN THE SCREW HOLE, MEANING THE DRILL GUIDE SITS CLOSER TO THE PERIPHERAL RAIL. THE DRILLED HOLE IN THE BONE IS NOT CONCENTRIC WITH THE SCREW HOLE IN THE TIBIAL TRAY. THE SURGEON HAD TO RE-DRILL TO CENTER THE HOLE. IF THE DRILL GUIDE IS USED ON CERTAIN TRAYS WHERE THE FEMORAL SIZE IS SMALLER THAN THE TIBIAL SIZE, IT WILL INTERFERE. AN INVESTIGATION WAS CONDUCTED; PER CAPA2019-57, THE MATING ISSUE REPORTED WAS LIKELY THE RESULT OF NOT SPECIFYING DESIGN INPUT REQUIREMENTS ENOUGH TO DRIVE THE TOLERANCE ANALYSIS VERIFICATION ACTIVITIES NEEDED AND NOT INCLUDING AN ADEQUATE SCOPE OF THE TOLERANCE ANALYSIS. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS NOT AVAILABLE FOR ANALYSIS. ENGINEERING EVAL COMPLETED BY NF ON 2020.04.02. {CAPA2019-57 IS STILL OPENED, AND CORRECTIVE ACTIONS ARE SUBJECT TO CHANGE. REFER TO THE CAPA FOR THE MOST UP TO DATE INFORMATION.} CRC2019-09-20-02 AND CAPA2019-57 FINDINGS: ACCORDING TO THE EXPERIENCE REPORT, THE SURGEON EXPERIENCED AN ISSUE WHERE THE TRULIANT POROUS BONE SCREW DRILL GUIDE (02-029-29-4000) WOULD NOT MATE CORRECTLY WITH THE TRULIANT POROUS TIBIAL TRAY (02-022-55-XXXX), SPECIFICALLY WITH THE ¿UPSIZE¿ TIBIAL TRAYS. THE SURGEON NOTED THAT THE DRILLED HOLE IN THE BONE IS NOT CONCENTRIC WITH THE SCREW HOLE IN THE TIBIAL TRAY. CAPA2019-57 WAS INITIATED TO INVESTIGATE THIS ISSUE, AND THE INVESTIGATION IDENTIFIED THE FOLLOWING: RELEVANT DOCUMENTS SINCE A DESIGN CHANGE ON 8/9/2017 INCLUDE TR-2018-1231, WHICH IS A TOLERANCE STACKUP, AND (B)(4), WHICH IS A SAWBONES EVALUATION LOOKING AT FIXATION OF THE TIBIAL TRAY. THE TOLERANCE EVALUATION CONSIDERED MATING CONDITIONS BETWEEN THE SPHERICAL BORE OF THE TRAY AND THE SPHERICAL RADIUS OF THE DRILL GUIDE BUT DID NOT CONSIDER THE POTENTIAL INTERFERENCE BETWEEN THE RAIL AND THE CIRCUMFERENCE OF THE DRILL GUIDE HEAD. THE ACTUAL MATING CONDITIONS TO BE TESTED ARE LAID OUT IN (B)(4). (B)(4) JUST GIVES THE RESULTS OF CONDITIONS EVALUATED IN (B)(4). THE SAWBONES EVALUATION DID NOT INCLUDE THE SPECIFIC TRAY SKUS WHERE INTERFERENCE HAPPENS (SEE CAPA2019-57 INVESTIGATION RESULTS FOR SPECIFIC TRAY SKUS). (B)(4) DOES SPECIFY REVISION LEVEL USED OF 02-029-229-4000, BUT THE LATEST DRILL GUIDE DESIGN WAS USED IN THE REPORT. INTERFERENCE OCCURS BETWEEN THE PERIPHERAL RAIL OF THE TRULIANT POROUS TIBIA (02-022-55-XXXX) AND THE DRILL GUIDE (02-029-29-4000). INTERFERENCE OCCURS BETWEEN TRAYS AND DRILL GUIDES WHEN THE FEMORAL SIZE IS SMALLER THAN THE TIBIAL SIZE, SUCH AS A 3F/4T TRAY, ALSO CALLED AN ¿UPSIZE¿ TRAY AS THE TIBIA IS SIZED ONE SIZE UP FROM THE FEMUR. OTHER STYLES OF TRAYS ARE CALLED ¿SIZE-TO-SIZE¿ OR ¿DOWN-SIZE¿ TRAY. ON THE UPSIZE TRAYS, THE SMALLER FEMUR SIZE NECESSITATES A SMALLER TIBIAL INSERT. THE PERIPHERAL RAIL NEEDS TO BE SMALLER FOR THIS SMALLER TIBIAL INSERT, WHICH IN TURN MEANS THE RAIL IS CLOSER TO THE SCREW HOLES FOR THESE IMPLANTS. THE DRILL GUIDE SITS IN THE SCREW HOLE, MEANING THE DRILL GUIDE SITS CLOSER TO THE PERIPHERAL RAIL. IF THE DRILL GUIDE IS USED ON CERTAIN TRAYS WHERE THE FEMORAL SIZE IS SMALLER THAN THE TIBIAL SIZE, IT WILL INTERFERE. CORRECTIVE ACTION TAKEN: THE DESIGN WILL BE UPDATED TO ELIMINATE THE INTERFERENCE WITH THE PERIPHERAL RAIL OF THE TIBIAL TRAYS. VERIFICATION AND VALIDATION ACTIVITIES WILL BE CONDUCTED TO JUSTIFY THE DESIGN OF THE NEW REVISION. THE QUALITY PLAN/MANUFACTURING QUALITY PLAN WILL BE UPDATED. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: PER CAPA2019-57, THE MATING ISSUE REPORTED IN (B)(4) WAS LIKELY THE RESULT OF NOT SPECIFYING DESIGN INPUT REQUIREMENTS ENOUGH TO DRIVE THE TOLERANCE ANALYSIS VERIFICATION ACTIVITIES NEEDED AND NOT INCLUDING AN ADEQUATE SCOPE OF THE TOLERANCE ANALYSIS. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICES USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. THE DRILL GUIDE SITS IN THE SCREW HOLE, MEANING THE DRILL GUIDE SITS CLOSER TO THE PERIPHERAL RAIL. THE DRILLED HOLE IN THE BONE IS NOT CONCENTRIC WITH THE SCREW HOLE IN THE TIBIAL TRAY. THE SURGEON HAD TO RE-DRILL TO CENTER THE HOLE. IF THE DRILL GUIDE IS USED ON CERTAIN TRAYS WHERE THE FEMORAL SIZE IS SMALLER THAN THE TIBIAL SIZE, IT WILL INTERFERE. AN INVESTIGATION WAS CONDUCTED; PER CAPA2019-57, THE MATING ISSUE REPORTED WAS LIKELY THE RESULT OF NOT SPECIFYING DESIGN INPUT REQUIREMENTS ENOUGH TO DRIVE THE TOLERANCE ANALYSIS VERIFICATION ACTIVITIES NEEDED AND NOT INCLUDING AN ADEQUATE SCOPE OF THE TOLERANCE ANALYSIS. | N | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00023 | 2022-01-10 | 2019-05-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | MDLR CNLTD TRIDRV DRVR | EXACTECH, INC. |  | 315-25-00 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED DIFFICULTY WITH THE TRI-DRIVE MODULAR HANDLE. THE DEVICE WOULD NOT HOLD ANY OF THE FITTINGS. NEW INFORMATION WAS RECEIVED 1 AUG 2021 STATING THAT DURING A ROUTINE WAREHOUSE CHECK THE DEVICE CONDITION WAS DISCOVERED. THE MODULAR TRI DRIVER FROM THE GLENOID TRAY IS NOT ABLE TO SECURE ANY FITTINGS INCLUDING REAMERS AND DRILL BITS. DID NOT CAUSE ANY ISSUES DURING A CASE. PER THE ENGINEERING EVALUATION, THERE WERE NO PIECES LEFT IN THE PATIENT. NO ADDITIONAL INFORMATION. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. NO DELAY IN SURGERY OR PATIENT ADVERSE EVENT. PER THE ENGINEERING EVALUATION, THERE WERE NO PIECES LEFT IN THE PATIENT. AN INVESTIGATION WAS CONDUCTED; THE ASSEMBLY DIFFICULTY REPORTED WAS LIKELY THE RESULT OF A MISSING BALL BEARING, WHICH PREVENTED THE ATTACHMENT FEATURE FROM PROPERLY ENGAGING A MATING DEVICE. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS AVAILABLE FOR ANALYSIS. ENGINEERING EVALUATION COMPLETED BY AD ON 31 MAR 2020. VISUAL EVALUATION: THE RETURNED TRI-DRIVE HANDLE WAS FUNCTIONALLY CHECKED WITH A 38MM MODULAR REVERSE CANNULATED REAMER (C/N: 321-35-38 , L/N: 51569011) AND PRESENTED MATING DIFFICULTIES. THE SLEEVE APPEARS TO RETRACT PROPERLY, HOWEVER THE DEVICE WOULD NOT ENGAGE THE REAMER. THIS IS CONSISTENT WITH THE REPORTED EVENT. THE MATING DIFFICULTY IS LIKELY THE RESULT OF THE MISSING BALL BEARING. DESIGN-RELATED ISSUES: THE DESIGN OF THE TRI-DRIVE CANNULATED HANDLE HAS BEEN IN THE FIELD SINCE 2009. EXACTECH IS AWARE OF 9 SIMILAR COMPLAINT REPORTS INVOLVING 16 TRIDRIVE CANNULATED HANDLE DUE TO ASSEMBLY DIFFICULTIES SINCE 2009. THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS AWARE OF 2 OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 60 UNITS, WHICH HAVE BEEN IN THE FIELD SINCE 2017. THE MANUFACTURING RECORDS FOR THIS DEVICE ONLY EXIST IN PAPER FORM AND ARE STORED IN CARDBOARD BOXES AT IRON MOUNTAIN. DUE TO RISKS ASSOCIATED WITH COVID-19 TRANSMISSION, THE PHYSICAL MANUFACTURING RECORDS FOR THIS COMPLAINT WILL NOT BE REVIEWED. EXACTECH WILL CONTINUE TO MONITOR THIS COMPLAINT TYPE AND, IF WARRANTED, WILL ESCALATE ANY POSITIVE TRENDS TO THE CORRECTIONS AND REMOVAL COMMITTEE (CRC) FOR ADDITIONAL ACTIONS. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RACR. THE ASSEMBLY DIFFICULTY REPORTED IN (B)(4) WAS LIKELY THE RESULT OF A MISSING BALL BEARING, WHICH PREVENTED THE ATTACHMENT FEATURE FROM PROPERLY ENGAGING A MATING DEVICE. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. NO DELAY IN SURGERY OR PATIENT ADVERSE EVENT. PER THE ENGINEERING EVALUATION, THERE WERE NO PIECES LEFT IN THE PATIENT. AN INVESTIGATION WAS CONDUCTED; THE ASSEMBLY DIFFICULTY REPORTED WAS LIKELY THE RESULT OF A MISSING BALL BEARING, WHICH PREVENTED THE ATTACHMENT FEATURE FROM PROPERLY ENGAGING A MATING DEVICE. | N | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00024 | 2022-01-10 | 2021-12-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | NOV 6.5MM HEX DRIVER ARTICULATING | ORTHOPEDIC MANUAL SURGICAL INSTRUMENT | EXACTECH, INC. | NOV 6.5MM HEX DRIVER ARTICULATING | 181-65-02 |  | AS REPORTED, THE TIP OF 6.5 HEX DRIVE BROKE OFF AT THE END OF THE SCREWDRIVER HEAD. THE SURGICAL SITE WAS CHECKED, AND NO INSTRUMENTATION WAS LEFT IN THE SURGICAL FIELD, IT WAS CONFIRMED CLEAN BY SURGEONS. THIS EVENT CAUSED NO DELAY AND THE PROCEDURE PROCEEDED WITH NO ISSUES. PATIENT IS CURRENTLY IN STABLE CONDITION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | N | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00026 | 2022-01-14 | 2021-12-23 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 3 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 3 9MM | 02-012-35-3009 |  | AS REPORTED, A (B)(6) , FEMALE PATIENT WAS SCHEDULED FOR REVISION OF A FEMORAL COMPONENT AND INSERT DUE TO INSTABILITY. THE SURGEON WENT FROM A PS TO A CONSTRAINED COMPONENTS. THE PATIENT¿S POSTOP X-RAYS SHOWED COMPONENT ALIGNMENT. IN TIME THE PATIENT STARTED MOVING INTO A VALGUS KNEE AND BECAME UNSTABLE. THE EVENT IS NOT RELATED TO THE BREAKAGE OF A DEVICE. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICES WILL BE RETURN. | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION RECEIVED THE FOLLOWING SECTIONS D9, G3, G6, H2, H3 AND H6 HAVE BEEN UPDATED ACCORDINGLY. AS REPORTED, A 54 YEARS OLD, FEMALE PATIENT WAS SCHEDULED FOR REVISION OF A FEMORAL COMPONENT AND INSERT DUE TO INSTABILITY. THE SURGEON WENT FROM A PS TO A CONSTRAINED COMPONENTS. THE PATIENT¿S POSTOP X-RAYS SHOWED COMPONENT ALIGNMENT. IN TIME THE PATIENT STARTED MOVING INTO A VALGUS KNEE AND BECAME UNSTABLE. THE EVENT IS NOT RELATED TO THE BREAKAGE OF A DEVICE. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICES WILL BE RETURN. UPON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT A DESIGN OR MANUFACTURING ISSUE CAUSED OR CONTRIBUTED TO THIS EVENT. THE REPORTED REVISION WAS LIKELY THE RESULT OF UNDERLYING PATIENT FACTORS, WHICH LED TO INSTABILITY OF THE JOINT. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00027 | 2022-01-14 | 2021-12-23 | I, n, j, u, r, y | Manufacturer report | TRULIANT PS POR FEM PS POR LEFT SZ 2.5 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT PS POR FEM PS POR LEFT SZ 2.5 | 02-020-12-0225 |  | AS REPORTED, A (B)(6), FEMALE PATIENT HAD A REVISION DUE TO INSTABILITY OF THE KNEE. THIS IS A WELL FIXATED POROUS FEMUR AND TIBIA LESS THAN A YEAR OLD. SURGEON REMOVED THE FEMORAL AND TIBIAL POROUS COMPONENTS. PATIENT REVISED TO A SIZE 2 CONDYLAR CONSTRAINED FEMUR AND HE SIZE 2F/1T TRULIANT TIBIAL FIT TRAY WITH A 15 MM CONDYLAR CONSTRAINED INSERT. THERE WAS NO BREAKAGE TO THE DEVICE. THERE WAS NO SURGICAL DELAY/ PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICES WILL BE RETURN. | CONCOMITANT MEDICAL PRODUCTS: TRULIANT TIB IMP PS INSERT SZ 2.5 10MM (CAT # 02-022-35-2510 / SERIAL # (B)(4)), TRULIANT POR TIB TRAY SIZE 2.5F/1.5T (CAT # 02-022-55-2515 / SERIAL # (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00028 | 2022-01-14 | 2021-11-01 | I, n, j, u, r, y | Manufacturer report | LOGIC CR TIB INSERT SLOPE++, SZ 2, 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC CR TIB INSERT SLOPE++, SZ 2, 13MM | 02-012-49-2013 |  | AS REPORTED, A (B)(6) WOMAN WHO HAD A LEFT TOTAL KNEE REPLACEMENT DONE FOR MEDIAL COMPARTMENT OA BACK IN 2018. PATIENT DID EXCEEDINGLY WELL AND ACHIEVED A RANGE OF MOTION (ROM) OF 0-138 DEGREES. PATIENT RESUMED FULL ACTIVITIES INCLUDING GOING TO THE GYM AND DOING LOW IMPACT EXERCISES AND RETURNED TO HER JOB AS A DENTAL ASSISTANT. PATIENT VISIT SURGEON LAST MONTH COMPLAINING OF WORSENING PROXIMAL TIBIAL PAIN AND SWELLING. NO HISTORY OF TRAUMA. NO CONSTITUTIONAL SYMPTOMS TO SUGGEST INFECTION. PATIENT EXAM SHOWED THAT SHE STILL MAINTAINED HER ROM BUT CLEARLY HAD A 1-2+ EFFUSION WITHOUT ERYTHEMA NOR UNUSUAL WARMTH. HER X-RAYS FROM (B)(6) 2021 WERE COMPARED TO HER X-RAYS FROM (B)(6) 2018 AND THIS CLEARLY SHOWED RADIOLUCENCIES AROUND THE MEDIAL FEMORAL CONDYLE, THE MEDIAL TIBIAL PLATEAU UNDER THE BASEPLATE AND AROUND THE KEEL OF THE TIBIAL BASEPLATE AS WELL. FURTHERMORE, OF GREATER CONCERN WAS THE FACT THAT HER POLY THICKNESS HAD DECREASED BY APPROXIMATELY 50%. PATIENT HAS SUBSEQUENTLY HAD A 3 PHASE BONE SCAN WHICH SHOWS INCREASED ACTIVITY IN ALL 3 OF THOSE SAME AREAS CONSISTENT WITH LOOSENING AND A MAVRIC MRI WHICH SHOWS A STRESS INJURY UNDER HER MEDIAL TIBIAL PLATEAU. HER SEROLOGY HAS BEEN NEGATIVE FOR INFECTION INCLUDING A NORMAL CRP AND SED RATE AND WBC. SHE HAS A JOINT ASPIRATION PENDING BUT AGAIN CLINICALLY HAS NO SIGN OF INFECTION. PATIENT IS ON MY LIST OF PATIENTS WHO HAVE THE GREATER THAN 5 YEAR SHELF LIFE OF OPTETRAK POLY AND SURGEON THINK PATIENT NEEDS A REVISION OF HER TIBIAL COMPONENT AND POSSIBLY HER FEMORAL COMPONENT AS WELL. | THIS FOLLOW-UP REPORT IS BEING SUBMITTED TO RELAY ADDITIONAL AND/OR CORRECTED INFORMATION. THE FOLLOWING SECTIONS WERE UPDATED/CORRECTED: H6 MDR SECTION CODES UPDATED/CORRECTED: A, B, C, D, E, F, G THE REVISION REPORTED WAS LIKELY THE RESULT OF BOTH PATIENT-RELATED CONDITIONS AND AN INSUFFICIENT BOND BETWEEN THE FEMORAL COMPONENT AND THE BONE, WHICH LED TO ASEPTIC (NON-INFECTED) FEMORAL LOOSENING. THE LOOSENING MAY HAVE LED TO INSTABILITY AND AN INCREASE IN CEMENT AND BONE PARTICLES IN THE JOINT SPACE, WHICH RESULTED IN PROSTHESIS WEAR AND OSTEOLYSIS. ADDITIONALLY, A CONTRIBUTING FACTOR TO THE WEAR AND DELAMINATION MAY HAVE BEEN THE RESULT OF BEING PACKAGED IN A NON-CONFORMING VACUUM BAG FOR MORE THAN FIVE YEARS. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WERE NOT AVAILABLE FOR EVALUATION IF ANY FURTHER INFORMATION IS OBTAINED THAT WOULD CHANGE OR ALTER ANY INFORMATION PROVIDED, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY. IF ANY FURTHER INFORMATION IS OBTAINED THAT WOULD CHANGE OR ALTER ANY INFORMATION PROVIDED, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00029 | 2022-01-19 | 2021-12-29 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | CAGE GLENOID MEDIUM, BETA | EXACTECH, INC. | 314-13-13 | 314-13-13 |  | IT WAS REPORTED THAT DURING IMPACTION, THE SURGEON BEGAN TO IMPACT THE GLENOID AND THE CAGE POPPED OFF OF THE GLENOID WHILE THE IT STAYED INTACT TO THE BONE. THE GLENOID WAS REMOVED FROM THE BONE AND A NEW ONE WAS USED. THERE WAS A DELAY BUT IT DID NOT CAUSE ANY ADVERSE EVENTS DUE TO THIS INCIDENT. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETUNING, THEY WERE THROWN OUT AFTER THE CASE. | SECTION H10: (H3) THE DISASSEMBLY/PROSTHESIS FRACTURE OF THE CAGE GLENOID REPORTED MAY HAVE BEEN THE RESULT OF IMPACTING THE GLENOID OFF-AXIS, CREATING A TORQUE/BENDING LOAD ON THE CENTER CAGE. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00030 | 2022-01-19 | 2021-12-30 | I, n, j, u, r, y | Manufacturer report | VANTAGE | TIBIAL PLATE FB SZ 1 LT | EXACTECH, INC. | 350-11-01 | 350-11-01 |  | IT WAS REPORTED THAT DURING IMPACTION OF THE LEFT ANKLE, THE SURGEON WAS TRYING TO FURTHER SEAT THE TIBIAL JUST A BIT MORE AFTER ALREADY IMPACTING IT SEVERAL TIMES AND A SMALL PART OF THE RIM OF THE IMPLANT WHERE THE POLY INSERTS INTO BROKE OFF. THIS PREVENTED A TRIAL POLY FROM BEING PLACED IN THE COMPONENT AND ULTIMATELY THE SURGEON TOOK THAT TIBIAL TRAY OUT AND PUT A NEW ONE IN ITS PLACE. THERE WAS NO SURGICAL DELAY. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. NO OTHER PATIENT INFORMATION OR MEDICAL HISTORY WAS REPORTED. DEVICES WILL BE RETURNING. | PENDING EVALUATION. NO INFORMATION PROVIDED IN THE FOLLOWING SECTION(S). | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00031 | 2022-01-19 | 2021-12-28 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 2 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED, POLYMER/METAL | EXACTECH, INC. | PS TIBIAL INSERT 2X11MM | 02-012-35-2011 |  | AS REPORTED, APPROXIMATELY TEN YEARS POST LEFT TKA, A (B)(6) FEMALE PATIENT HAD A REVISION DUE TO POLY WEAR. UPON REMOVAL OF THE INSERT, IT WOULD APPEAR THAT THE TIBIAL INSERT IS FRACTURED POSTERIOR/MEDIAL AROUND THE AREA OF THE POSTERIOR LOCKING MECHANISM. UPON PUTTING THE KNEE IN FLEXION, THE PS INSERT LIFTED OUT OF THE TIBIAL TRAY ANTERIORLY ABOUT 3-4MM. THERE WERE NO PARTS OR PIECES OF THE DEVICE FELL INTO THE PATIENT WOUND SITE. THERE WAS NO SURGICAL DELAY/PROLONGATION REPORTED. PATIENT REVISED FROM A SIZE 2 11MM PS POLY TO A 13MM PSC POLY. PATIENT DISPOSITION LEFT OR WITH NO ISSUES. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION DUE TO HOSPITAL KEEPS ALL IMPLANT RETRIEVALS. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00036 | 2022-01-23 | 2021-12-30 | I, n, j, u, r, y | Manufacturer report | OPTETRACK CR | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | OPTETRACK CR | UNKNOWN |  | AS REPORTED, APPROXIMATELY 6.5 YEARS POST TKA, PATIENT HAD A REVISION OF AN OPTETRACK CR DUE TO MASSIVE MEDIAL PROXIMAL TIBIAL OSTEOLYSIS. NO ADDITIONAL INFORMATION AVAILABLE. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00034 | 2022-01-23 | 2021-12-29 | I, n, j, u, r, y | Manufacturer report | FEMORAL COMPONENT | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | FEMORAL COMPONENT | UNKNOWN |  | AS REPORTED, APPROXIMATELY THREE YEARS POST TKA, PATIENT HAD A REVISION DUE TO OSTEOLYSIS. NO ADDITIONAL INFORMATION AVAILABLE. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00035 | 2022-01-23 | 2021-12-30 | I, n, j, u, r, y | Manufacturer report | FEMORAL COMPONENTS | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | FEMORAL COMPONENTS | UNKNOWN |  | AS REPORTED, APPROXIMATELY FIVE YEARS POST TKA, PATIENT HAD A REVISION DUE TO OSTEOLYSIS. NONE LEFT ON THE TIBIA EXCEPT FOR THE TT AND A SMALL CENTRAL POST. NO ADDITIONAL INFORMATION AVAILABLE. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00032 | 2022-01-23 | 2022-01-03 | I, n, j, u, r, y | Manufacturer report | TRULIANT CR CEM FEM CR CEM LEFT SZ 4 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT CR CEM FEM CR CEM LEFT SZ 4 | 02-020-13-0240 |  | AS REPORTED, A (B)(6) PATIENT COMPLAINT OF PAIN AND HAD A REVISION DUE TO LOSE FEMUR. THERE WAS NO BREAKAGE OF THE DEVICE. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICE WILL BE RETURN. | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION RECEIVED THE FOLLOWING SECTIONS G3, G6, H2, H3 AND H6 HAVE BEEN UPDATED ACCORDINGLY. (H3) THE REVISION REPORTED WAS LIKELY THE RESULT OF AN INSUFFICIENT BOND BETWEEN THE FEMORAL COMPONENT AND THE BONE, WHICH LED TO ASEPTIC (NON-INFECTED) FEMORAL LOOSENING AND PAIN. THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE REPORTED EVENT OF "LOOSENING - FEMORAL¿ IS ASSOCIATED WITH WEAKENED INTEGRATION OF THE FEMORAL COMPONENT AT THE BONE-IMPLANT INTERFACE DUE TO LOSS OF FIBROUS AND/OR BONY TISSUE AND LEADING TO COMPROMISED ANCHORAGE OF THE DEVICE. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00033 | 2022-01-23 | 2021-12-29 | I, n, j, u, r, y | Manufacturer report | FEMORAL COMPONENTS | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | FEMORAL COMPONENTS | UNKNOWN |  | AS REPORTED, APPROXIMATELY SIX YEARS POST TKA, PATIENT HAD A REVISION DUE TO OSTEOLYSIS. NO ADDITIONAL INFORMATION AVAILABLE. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00040 | 2022-01-24 | 2022-01-05 | I, n, j, u, r, y | Manufacturer report | CC TIBIAL INSERT SZ 3, 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | CC TIBIAL INSERT SZ 3, 11MM | 208-23-11 |  | AS REPORTED, APPROXIMATELY 8 YRS POSTOP THE INITIAL LTKA, THIS (B)(6) FEMALE PATIENT EXPERIENCED PAIN AND POLY WEAR. THE PATIENT EXPERIENCE MASSIVE SWELLING AND NEEDS A NEW POLY. PATIENT HAS A HISTORY OF LEFT KNEE OSTEOARTHRITIS. PATIENT LEFT THE OR WITH NO ISSUES. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO HSS KEEPS ALL RETRIEVALS. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00039 | 2022-01-24 | 2022-01-04 | I, n, j, u, r, y | Manufacturer report | TRULIANT PIN PULLER SYRINGE STYLE | ORTHOPEDIC MANUAL SURGICAL INSTRUMENT | EXACTECH, INC. | TRULIANT PIN PULLER SYRINGE STYLE | 02-029-90-4100 |  | AS REPORTED, DURING SURGICAL USE THE TRULIANT PIN PULLER BROKE DURING A TRULIANT TKA. THE PIN PULLER WAS ENGAGED WITH THE PIN AND UPON PULLING THE PIN VERTICALLY UP OUT OF THE TIBIA TRAY THE PIN PULLER INNER PRONG BROKE. THE BROKEN PRONG WAS ISOLATED AND REMOVED FROM THE STERILE FIELD. THERE WERE NO PARTS OR PIECES OF THE DEVICE FELL INTO THE PATIENT WOUND SITE. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE PIN PULLER IS NOT USABLE AND IS BEING RETURNED TO EXACTECH. | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION RECEIVED THE FOLLOWING SECTIONS G3, G4, G6, H2, H3 AND H6 HAVE BEEN UPDATED ACCORDINGLY. (H3) THE BROKEN TRULIANT PIN PULLERS REPORTED IN WERE LIKELY THE RESULT OF THE BREAKAGE AREA OF THE COMPONENT BEING RELATIVELY THIN WHICH DOES NOT OFFER GREAT SAFETY MARGIN WHEN THE INSTRUMENT IS SUBJECT TO OFF-AXIS LOADING AND/OR OFF-AXIS LOADING WHEN THE USER APPLIES A BENDING MOMENT TO THE DEVICE DURING PIN DRIVING/PULLING. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00037 | 2022-01-24 | 2022-01-04 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | GLENOID, PEGGED BETA, LARGE | EXACTECH, INC. | 314-02-14 | 314-02-14 |  | AS REPORTED, APPROXIMATELY 13 YEARS POST OP THE INITIAL TSA, THIS PATIENT WAS REVISED DUE TO INFECTION. THE STEM WAS WELL FIXED BUT WAS NOT DIFFICULT TO REMOVE WITH SLAPHAMMER. GLENOID WAS LOOSE AND EASILY REMOVED WITH CEMENT REMAINING ON THE BACK OF THE COMPONENT PEGS. PATIENT WAS TREATED WITH ANABIOTIC SHOULDER SPACER AN ANTIBIOTIC BEADS. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION. NO PATIENT INFORMATION OR MEDICAL HISTORY IS AVAILABLE. | DEVICE EVALUATED BY MANUFACTURER? PENDING EVALUATION. CONCOMITANT MEDICAL DEVICE(S): 300-01-13, (B)(4) - EQUINOXE, HUMERAL STEM PRIMARY, PRESS FIT 13MM; 300-10-45, (B)(4) - EQUINOXE REPLICATOR PLATE 4.5MM O/S; 300-20-02, (B)(4) - EQUINOX SQUARE TORQUE DEFINE SCREW DRIVE KIT; 310-01-47, (B)(4) - EQUINOXE, HUMERAL HEAD SHORT, 47MM (BETA). | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00038 | 2022-01-24 | 2017-02-17 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | RS EXPANDED GLENOSPHERE 42MM, +4MM OFFSET | EXACTECH, INC. | 320-42-13 | 320-02-42 |  | AS REPORTED, THIS (B)(6) MALE PATIENT REQUIRED A REVISION ABOUT 5 MONTHS POST OP DUE TO DISLOCATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETURNING. | DEVICE EVALUATED BY MANUFACTURER? PENDING EVALUATION CONCOMITANT MEDICAL DEVICE(S): 320-20-00, (B)(4) - EQ REVERSE TORQUE DEFINING SCREW KIT; 320-10-00, (B)(4) - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0; 320-15-05, (B)(4) - EQ REV LOCKING SCREW. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00041 | 2022-01-25 | 2021-12-22 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | CAGE GLENOID S, POST AUG, LEFT | EXACTECH, INC. | 314-13-22 | 314-13-22 |  | AS REPORTED, APPROXIMATELY 3 YEARS POST OP THE INITIAL LEFT TSA, THIS (B)(6) FEMALE PATIENT'S WAS REVISED DUE TO ASEPTIC GLENOID LOOSENING. PATIENT HAD A SUBSCAP FAILURE AND A LOOSE TSA GLENOID IN STEMLESS IMPLANT. THE CASE REPORT FORM INDICATES THIS EVENT IS POSSIBLY RELATED TO DEVICES AND/OR PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | DEVICE EVALUATED BY MANUFACTURER? PENDING EVALUATION. CONCOMITANT MEDICAL DEVICE(S): 300-62-02 - STEMLESS HUMERAL COMP INTEGRIP, CAGE, SIZE 2; 310-61-44 - STEMLESS HUMERAL HEAD 44MM X 17MM X ALPHA. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00042 | 2022-01-26 | 2020-01-15 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 13MM | 02-012-44-4013 |  | AS REPORTED, APPROXIMATELY TWO YEARS POST INITIAL LEFT TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00043 | 2022-01-26 | 2020-02-18 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 3, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 3, 9MM | 02-012-44-3009 |  | AS REPORTED, APPROXIMATELY FIVE YEARS POST INITIAL RIGHT TKA, THIS FEMALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00044 | 2022-01-26 | 2022-01-06 | I, n, j, u, r, y | Manufacturer report | LOGIC CR TIB INSERT SLOPE ++, SZ 3.5, 11 MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC CR TIB INSERT SLOPE ++, SZ 3.5, 11 MM | 02-012-49-3511 |  | AS REPORTED, PATIENT PRESENTED WITH PAIN AND INSTABILITY AND WAS SCHEDULED FOR A RIGHT REVISION TKA POSSIBLE POLY SWAP. UPON REVISION, THE POLY INSERT WAS REMOVED AND APPEARED TO BE BROKEN INTO PIECES POSTERIORLY. THE POLY WAS NOT INTACT AND SHOWED ABNORMAL WEAR. THE METAL FEMUR AND TIBIA IMPLANTS WERE WELL FIXED, AND THE PATIENT OBTAINED A POLY SWAP AFTER TRAILING WITH A CRC SIZE 13MM INSERT, A NEW CRC SIZE 3.5, 13MM POLY WAS IMPLANTED. THE PATIENT WAS REVISED FROM AN 11MM, 3.5 ++ POLY TO A 13MM 3.5 CRC POLY. THE SURGEON REMOVED AND IRRIGATED OUT THE PARTS/PIECES FROM THE KNEE. THE PATIENT LEFT THE OR ROOM STABLE. THE POLY AND ITS PIECES WERE CONTAINED, SAVES, AND WILL BE RETURNED TO MANUFACTURER. | THE REVISION REPORTED WAS LIKELY THE RESULT OF A COMBINATION OF AXIAL ROTATION MISMATCH OF THE FEMORAL COMPONENT RELATIVE TO THE TIBIAL INSERT, OVERALL POSTERIOR SLOPE OF THE TIBIAL INSERT, AND PATIENT-RELATED CONDITIONS, WHICH ALLOWED FOR EXCESSIVE POSTERIOR CONTACT, ESPECIALLY MEDIALLY, AND LED TO SEVERE WEAR OF THE POLYETHYLENE. ADDITIONALLY, A CONTRIBUTING FACTOR TO THE SEVERE WEAR AND DELAMINATION MAY HAVE BEEN THE RESULT OF BEING PACKAGED IN A NON-CONFORMING VACUUM BAG FOR MORE THAN FIVE YEARS. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00045 | 2022-01-26 | 2020-06-02 | I, n, j, u, r, y | Manufacturer report | FEMORAL COMPONENTS | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | FEMORAL COMPONENTS | UNKNOWN |  | AS REPORTED, APPROXIMATELY THREE YEARS POST INITIAL RTKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO ACUTE INFECTION AND LINER EXCHANGE. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION RECEIVED THE FOLLOWING SECTIONS G3, G6, H2, H3 AND H6 HAVE BEEN UPDATED ACCORDINGLY. (H3) AS REPORTED, APPROXIMATELY THREE YEARS POST INITIAL RTKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO ACUTE INFECTION AND LINER EXCHANGE. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN OR MANUFACTURING ISSUES. THE CAUSE OF SURGICAL REVISION FOR INFECTION CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION. THESE DEVICES ARE USED FOR TREATMENT NOT DIAGNOSIS. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00046 | 2022-01-26 | 2020-08-26 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 4 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 4 9MM | 02-012-35-4009 |  | AS REPORTED, APPROXIMATELY NINE YEARS POST INITIAL RIGHT TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | SECTION H10: (H3) THE REVISION REPORTED WAS LIKELY THE RESULT OF AN INSUFFICIENT BOND BETWEEN THE IMPLANT AND THE BONE AND PATIENT-RELATED CONDITIONS, WHICH LED TO ASEPTIC (NONINFECTED) LOOSENING. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICES WERE NOT AVAILABLE FOR EVALUATION AND PRE-REVISION X-RAYS WERE NOT PROVIDED. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00050 | 2022-01-27 | 2020-01-24 | I, n, j, u, r, y | Manufacturer report | FEMORAL COMPONENTS | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | FEMORAL COMPONENTS | UNKNOWN |  | AS REPORTED, APPROXIMATELY SIX YEARS POST INITIAL RIGHT LTK IMPLANT, THIS FEMALE PATIENT UNDERWENT REVISION DUE TO INSTABILITY. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00057 | 2022-01-27 | 2021-02-10 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 3, 15MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 3, 15MM | 02-012-44-3015 |  | AS REPORTED, APPROXIMATELY THREE YEARS POST INITIAL RIGHT TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | SECTION H10: (H3) THE REVISION REPORTED WAS LIKELY THE RESULT OF AN INSUFFICIENT BOND BETWEEN THE IMPLANT AND THE BONE, WHICH LED TO ASEPTIC (NON-INFECTED) LOOSENING. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICES WERE NOT AVAILABLE FOR EVALUATION AND PRE-REVISION X-RAYS WERE NOT PROVIDED. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00056 | 2022-01-27 | 2020-12-03 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 9MM | 02-012-44-4009 |  | AS REPORTED, APPROXIMATELY SIX YEARS POST INITIAL LEFT TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | SECTION H10: (H3) THE REVISION REPORTED WAS LIKELY THE RESULT OF AN INSUFFICIENT BOND BETWEEN THE IMPLANT AND THE BONE AND PATIENT-RELATED CONDITIONS, WHICH LED TO ASEPTIC (NONINFECTED) LOOSENING. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICES WERE NOT AVAILABLE FOR EVALUATION AND PRE-REVISION X-RAYS WERE NOT PROVIDED. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00055 | 2022-01-27 | 2020-12-07 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 3 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 3 9MM | 02-012-35-3009 |  | AS REPORTED, APPROXIMATELY ELEVEN YEARS POST INITIAL LEFT TKA IMPLANT, THIS FEMALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING AND INSTABILITY. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00054 | 2022-01-27 | 2021-02-10 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 2, 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 2, 13MM | 02-012-44-2013 |  | AS REPORTED, APPROXIMATELY EIGHT YEARS POST INITIAL RIGHT TKA, THIS FEMALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00053 | 2022-01-27 | 2020-08-07 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 3 15MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 3 15MM | 02-012-35-3015 |  | AS REPORTED, APPROXIMATELY TEN YEARS SINCE INITIAL LEFT TKA IMPLANT, THIS FEMALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00049 | 2022-01-27 | 2021-01-06 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 2, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 2, 9MM | 02-012-44-2009 |  | AS REPORTED, APPROXIMATELY THREE YEARS POST INITIAL RIGHT TKA IMPLANT, THIS FEMALE PATIENT UNDERWENT REVISION DUE TO INFECTION. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00051 | 2022-01-27 | 2020-12-03 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 9MM | 02-012-44-4009 |  | AS REPORTED, APPROXIMATELY EIGHT YEARS POST INITIAL LEFT TKA IMPLANT, THIS MALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00058 | 2022-01-27 | 2020-12-01 | I, n, j, u, r, y | Manufacturer report | FEMORAL COMPONENTS | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | FEMORAL COMPONENTS | UNKNOWN |  | AS REPORTED, APPROXIMATELY FIVE YEAR POST INITIAL RIGHT TKA, THIS FEMALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00048 | 2022-01-27 | 2020-03-12 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 5, 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 5, 11MM | 02-012-44-5011 |  | AS REPORTED, APPROXIMATELY ONE YEAR POST INITIAL RIGHT TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO INSTABILITY. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00047 | 2022-01-27 | 2020-12-31 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 2.5, 15MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 2.5, 15MM | 02-012-44-2515 |  | AS REPORTED, APPROXIMATELY 9 YEARS POST INITIAL RIGHT TKA IMPLANT, THIS FEMALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00052 | 2022-01-27 | 2020-12-23 | I, n, j, u, r, y | Manufacturer report | SPECIFIC DEVICE NOT REPORTED | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 3 13MM | UNK |  | AS REPORTED, APPROXIMATELY NINE YEARS POST INITIAL RIGHT TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00059 | 2022-01-27 | 2020-03-10 | I, n, j, u, r, y | Manufacturer report | CC TIBIAL INSERT SZ 2, 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | CC TIBIAL INSERT SZ 2, 13MM | 208-22-13 |  | AS REPORTED, APPROXIMATELY NINE YEARS POST INITIAL RIGHT TKA, THIS FEMALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | SECTION H10: (H3) THE REVISION REPORTED WAS LIKELY THE RESULT OF AN INSUFFICIENT BOND BETWEEN THE IMPLANT AND THE BONE AND PATIENT-RELATED CONDITIONS, WHICH LED TO ASEPTIC (NONINFECTED) LOOSENING. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICES WERE NOT AVAILABLE FOR EVALUATION AND PRE-REVISION X-RAYS WERE NOT PROVIDED. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00068 | 2022-01-28 | 2020-08-10 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 2, 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 2, 11MM | 02-012-44-2011 |  | AS REPORTED, APPROXIMATELY SEVEN YEARS POST INITIAL LEFT TKA, THIS FEMALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | H3: THE REVISION REPORTED WAS LIKELY THE RESULT OF AN INSUFFICIENT BOND BETWEEN THE IMPLANT AND THE BONE AND PATIENT-RELATED CONDITIONS, WHICH LED TO ASEPTIC (NON-INFECTED) LOOSENING. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICES WERE NOT AVAILABLE FOR EVALUATION AND PRE-REVISION X-RAYS WERE NOT PROVIDED. THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE REPORTED EVENT OF "LOOSENING¿ IS ASSOCIATED WITH WEAKENED INTEGRATION OF THE DEVICE AT THE BONE-IMPLANT INTERFACE DUE TO LOSS OF FIBROUS AND/OR BONY TISSUE AND LEADING TO COMPROMISED ANCHORAGE OF THE DEVICE. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00069 | 2022-01-28 | 2021-12-17 | I, n, j, u, r, y | Manufacturer report | NV GXL LINER NEUTRAL, 36MM ID, GROUP 2 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LINER NEUTRAL, 36MM ID, GROUP 2 CUPS | 130-36-52 |  | AS REPORTED, A (B)(6) FEMALE PATIENT REQUITED REVISION HIP SURGERY DUE TO WEAR DEBRIS/OSTEOLYSIS. THE CUP AND LINER WERE REVISED TO ANOTHER SYSTEM. THE STEM REMAINED IN PLACE. THERE WAS NO BREAKAGE OF THE DEVICE. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. | CONCOMITANT MEDICAL PRODUCTS: NV CROWN CUP CLSTR HOLE 52MM GROUP 2 (CAT# 180-01-52 / SERIAL# (B)(4)). BIOLOX DELTA FEMORAL HEAD 36MM OD, +0MM (CAT# 170-36-00 / SERIAL# (B)(4)). ELEMENT-STEM, COLLARED W/HA, STD OFFSET, SZ 13 (CAT# 164-03-13 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00067 | 2022-01-28 | 2020-09-09 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 2 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 2 9MM | 02-012-35-2009 |  | AS REPORTED, APPROXIMATELY 9 YEARS POST RIGHT TKA, THIS FEMALE PATIENT UNDERWENT REVISION DUE TO PERIPROSTHETIC FRACTURE. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00071 | 2022-01-28 | 2021-10-10 | I, n, j, u, r, y | Manufacturer report | ALT XLE LNR EXTCOV G7 40 | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | ALT XLE LNR EXTCOV G7 40 | 01-030-42-0740 |  | AS REPORTED, A (B)(6) PATIENT HAD RIGHT THA REVISION OF THE STEM, HEAD, AND LINER FOR AN UNKNOWN REASON. ALTEON CUP STAYED AND A NEW EXACTECH EXTENDED COVERAGE LINER +40MM HEAD WAS IMPLANTED. THE STEM AND HEAD WERE REVISED TO A DIFFERENT COMPANY. NO OTHER IMPLANT INFORMATION IS KNOWN. THERE WAS NO BREAKAGE OF DEVICES. THERE WAS 30-45 MINUTE SURGICAL DELAY/ PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. X-RAYS WERE RECEIVE; HOWEVER, THE DEVICES ARE NOT AVAILABLE FOR EVALUATION AS IMPLANTS THAT WERE REMOVED WERE THROWN AWAY BY THE STAFF AT THE FACILITY. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00070 | 2022-01-28 | 2021-10-15 | I, n, j, u, r, y | Manufacturer report | NV GXL LINER LIPPED 36MM ID GROUP 2 | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LINER LIPPED 36MM ID GROUP 2 | 132-36-52 |  | AS REPORTED, APPROXIMATELY FOUR YEARS POST THA, A MALE PATIENT UNDERWENT A THA REVISION FOR AN UNKNOWN REASON. THE STEM HEAD AND LINER WERE REMOVED FROM THE PATIENT. A NEW EXACTECH LINER WAS OPENED AND IMPLANTED DURING THE CASE. OTHER COMPANIES MONOBLOCK STEM AND HEAD WERE USED DURING THE REVISION SURGERY. THERE WAS NO BREAKAGE OF DEVICE. THERE WAS 30¿45 MINUTE SURGICAL DELAY/ PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE IMPLANTS THAT WERE REMOVED WERE NOT SAVED AND ARE NOT RETURNING. ANY FURTHER DETAILS OF THIS REVISION PROCEDURE ARE UNKNOWN. | CONCOMITANT MEDICAL PRODUCTS: NECK PRESERVING STEM, STD OFFSET PLASMA SZ 4 (CAT # 162-00-04 / SERIAL # (B)(4)), BIOLOX DELTA FEMORAL HEAD 36MM OD, +0MM (CAT # 170-36-00 / SERIAL # (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00066 | 2022-01-28 | 2020-09-14 | I, n, j, u, r, y | Manufacturer report | PS TIBIAL INSERTS SZ 4, 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | PS TIBIAL INSERTS SZ 4, 11MM | 204-24-11 |  | .AS REPORTED, APPROXIMATELY 20 YEARS POST INITIAL LEFT THA, THIS MALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00061 | 2022-01-28 | 2020-02-20 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 5 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 5 9MM | 02-012-35-5009 |  | AS REPORTED, THIS MALE PATIENT UNDERWENT REVISION DUE TO OSTEOLYSIS. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. THIS IS ONE OF TWO PRODUCTS INVOLVED WITH THE REPORTED EVENT AND THE ASSOCIATED MANUFACTURER REPORT NUMBERS ARE 1038671-2022-00062. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00064 | 2022-01-28 | 2020-01-02 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 4 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 4 9MM | 02-012-35-4009 |  | AS REPORTED, APPROXIMATELY SEVEN YEARS POST INITIAL RIGHT TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO OSTEOLYSIS. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION RECEIVED THE FOLLOWING SECTIONS G3, G6, H2, H3 AND H6 HAVE BEEN UPDATED ACCORDINGLY. (H3) THE REASON FOR THE REVISION REPORTED CANNOT BE DETERMINED FROM THE INFORMATION AVAILABLE. HOWEVER, THE SECONDARY FINDING OF OSTEOLYSIS MAY HAVE BEEN RELATED TO PROSTHESIS WEAR AND/OR LOOSENING. THIS CANNOT BE CONFIRMED BECAUSE THE DEVICES WERE NOT RETURNED FOR EVALUATION AND NO IMAGES OR X-RAYS WERE PROVIDED. THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE REPORTED EVENT OF "OSTEOLYSIS¿ IS ASSOCIATED WITH DESTRUCTION OR DISAPPEARANCE OF BONE TISSUE LIKELY RELATED TO WEAKENED INTEGRATION OF AN IMPLANT AT THE BONE-IMPLANT INTERFACE. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00063 | 2022-01-28 | 2020-02-18 | I, n, j, u, r, y | Manufacturer report | PS TIBIAL INSERTS SZ 2, 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | PS TIBIAL INSERTS SZ 2, 13MM | 204-22-13 |  | AS REPORTED, APPROXIMATELY 11 YEARS POST INITIAL RIGHT TKA, THIS FEMALE PATIENT UNDERWENT REVISION DUE TO DEBRIS SYNOVITITIS. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00062 | 2022-01-28 | 2020-02-20 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 5 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 5 9MM | 02-012-35-5009 |  | AS REPORTED, THIS MALE PATIENT UNDERWENT REVISION DUE TO OSTEOLYSIS. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. THIS IS ONE OF TWO PRODUCTS INVOLVED WITH THE REPORTED EVENT AND THE ASSOCIATED MANUFACTURER REPORT NUMBER ARE 1038671-2022-00061. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00065 | 2022-01-28 | 2020-09-15 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 5 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 5 9MM | 02-012-35-5009 |  | AS REPORTED, APPROXIMATELY 3 YEARS POST INITIAL LEFT TKA. THIS MALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00060 | 2022-01-28 | 2020-05-15 | I, n, j, u, r, y | Manufacturer report | LOGIC CC TIB INSERT SIZE 2, 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC CC TIB INSERT SIZE 2, 13MM | 02-012-65-2013 |  | AS REPORTED, APPROXIMATELY EIGHT MONTHS POST INITIAL RIGHT TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00085 | 2022-01-29 | 2020-10-26 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 4 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 4 11MM | 02-012-35-4011 |  | AS REPORTED, APPROXIMATELY SEVEN YEARS POST RIGHT TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00084 | 2022-01-29 | 2020-06-09 | I, n, j, u, r, y | Manufacturer report | FEMORAL COMPONENTS | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | FEMORAL COMPONENTS | UNKNOWN |  | AS REPORTED, THIS PATIENT UNDERWENT A REVISION FOR UNKNOWN REASON. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00083 | 2022-01-29 | 2020-01-10 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 3, 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 3, 13MM | 02-012-44-3013 |  | AS REPORTED, APPROXIMATELY TWO YEARS POST LEFT TKA, THIS FEMALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00082 | 2022-01-29 | 2020-05-28 | I, n, j, u, r, y | Manufacturer report | CEMENTED CR FEMORAL | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | FEMORAL COMPONENTS | 208-01-03 |  | AS REPORTED, APPROXIMATELY FOUR YEARS POST LEFT TKA, THIS FEMALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00081 | 2022-01-29 | 2021-01-04 | I, n, j, u, r, y | Manufacturer report | FEMORAL COMPONENTS | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | FEMORAL COMPONENTS | UNKNOWN |  | AS REPORTED, APPROXIMATELY EIGHTEEN YEARS POST TKA, THIS FEMALE PATIENT UNDERWENT REVISION FOR AN UNKNOWN REASON. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00080 | 2022-01-29 | 2020-12-09 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 13MM | 02-012-44-4013 |  | AS REPORTED, APPROXIMATELY THREE YEARS POST INITIAL RIGHT TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO RECURRENT HEMARTHROSIS. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00079 | 2022-01-29 | 2020-11-23 | I, n, j, u, r, y | Manufacturer report | PS TIBIAL INSERTS SZ 6, 15MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | PS TIBIAL INSERTS SZ 6, 15MM | 204-26-15 |  | AS REPORTED, APPROXIMATELY EIGHTEEN YEARS POST INITIAL LEFT TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO POLYETHYENE WEAR. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00078 | 2022-01-29 | 2021-02-25 | I, n, j, u, r, y | Manufacturer report | FEMORAL COMPONENTS | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | FEMORAL COMPONENTS | UNKNOWN |  | AS REPORTED, THIS PATIENT UNDERWENT REVISION FOR AN UNKNOWN REASON. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00077 | 2022-01-29 | 2020-11-10 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 3, 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 3, 11MM | 02-012-44-3011 |  | AS REPORTED, APPROXIMATELY 4 YEARS POST LEFT TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO POLYETHLENE WEAR. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00076 | 2022-01-29 | 2021-04-15 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 5 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 5 9MM | 02-012-35-5009 |  | AS REPORTED, APPROXIMATELY 10 YEARS POST INITIAL RIGHT TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO OSTEOLYSIS. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00075 | 2022-01-29 | 2022-01-05 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | L POSTERIOR AUG | EXACTECH, INC. |  | UNK |  | AS REPORTED, APPROXIMATELY 4 YEARS POST OP THE INITIAL LEFT TSA, THIS (B)(6) MALE PATIENT WAS REVISED DUE TO THE GLENOID SHEARING OFF THE CENTRAL PEG. THE PEGS REMAINED IN BONE AND THE POLY DISASSOCIATED FROM THEM. THE SURGEON REVISED THIS PATIENT TO A BIOMET MINI BASEPLATE AND GLENOSPHERE AND EQUINOXE STEM. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETUNING DUE TO HOSPITAL POLICY. | PENDING EVALUATION. CONCOMITANT MEDICAL PRODUCTS: 300-62-03, 03140822 - STEMLESS HUMERAL COMP INTEGRIP, CAGE, SIZE 3; 310-60-47, 104875004 - STEMLESS HUMERAL HEAD EXTRA SHORT, 47MM (BETA). | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00074 | 2022-01-29 | 2022-01-06 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | SHOULDER COMPONENTS | EXACTECH, INC. |  | UNK |  | AS REPORTED, APPROXIMATELY 5 YEARS POST OP THE INITIAL RIGHT RTSA, THIS (B)(6) FEMALE PATIENT WAS REVISED DUE TO AN INFECTION. INFECTION STARTED AFTER PATIENT RECEIVED AN OPEN ACJ EXCISION IN (B)(6) 2021 IN THE (B)(6). PATIENT HAD A SOFT TISSUE ANCHOR IN THE GT, SURGEON WAS UNSURE WHAT IT WAS ORIGINALLY USED TO ATTACHED TO. DEVICES WILL NOT BE RETUNING DUE TO HOSPITAL POLICY. PATIENT REQUIRED PROLONGED HOSPITALIZATION AS A DIRECT RESULT OF THIS INFECTION ISSUE. SURGEON AWARE THE REVISION WAS NOT AS A DIRECT RESULT OF EXACTECH IMPLANTS. | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00073 | 2022-01-29 | 2022-01-03 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REV LOCKING SCREW | EXACTECH, INC. | 320-15-05 | 320-15-05 |  | AS REPORTED, THIS (B)(6) PATIENT'S RIGHT SHOULDER WAS REVISED DUE TO THE GLENOSPHERE LOCKING SCREW AND GLENOSPHERE BACKING OFF OF THE BASEPLATE. THERE WAS NO SURGICAL DELAY/ PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETUNING DUE TO HOSPITAL POLICY. | PENDING EVALUATION. CONCOMITANT DEVICE(S): 320-02-42 - RS EXPANDED GLENOSPHERE 42MM, +4MM OFFSET. NO INFORMATION PROVIDED. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00072 | 2022-01-29 | 2021-12-27 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 42MM HUMERAL LINER +2.5 | EXACTECH, INC. | 320-42-03 | 320-42-03 |  | AS REPORTED, THIS (B)(6) PATIENT'S LEFT SHOULDER WAS REVISED DUE TO THE LINER DISASSOCIATING FROM THE HUMERAL TRAY. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL BE RETUNING. | PENDING EVALUATION . CONCOMITANT DEVICE(S): 320-01-05, 05311223 - HUMERAL TRAY, 300-30-06, 05112876 - EQUINOXE PRESERVE STEM 6MM, REVERSE SHOULDER TORQUE DEFINING SCREW KIT. * | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00093 | 2022-01-30 | 2020-06-30 | I, n, j, u, r, y | Manufacturer report | FEMORAL COMPONENTS | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | FEMORAL COMPONENTS | UNKNOWN |  | AS REPORTED, APPROXIMATELY FIVE YEARS POST INITIAL RIGHT TKA, THIS FEMALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00092 | 2022-01-30 | 2020-09-23 | I, n, j, u, r, y | Manufacturer report | CC TIBIAL INSERT SZ 2, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | CC TIBIAL INSERT SZ 2, 9MM | 208-22-09 |  | AS REPORTED, APPROXIMATELY EIGHT YEARS POST INITIAL RIGHT TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO INFECTION. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00091 | 2022-01-30 | 2021-01-25 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 5 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 5 9MM | 02-012-35-5009 |  | AS REPORTED, APPROXIMATELY TEN YEARS POST RIGHT TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO OSTEOLYSIS. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00090 | 2022-01-30 | 2020-03-13 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 5 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 5 9MM | 02-012-35-5009 |  | AS REPORTED, APPROXIMATELY EIGHT YEARS POST INITIAL LEFT TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00088 | 2022-01-30 | 2021-03-24 | I, n, j, u, r, y | Manufacturer report | FEMORAL COMPONENT | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | FEMORAL COMPONENT | UNKNOWN |  | AS REPORTED, THIS FEMALE PATIENT UNDERWENT LEFT SIDE TKA. THE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00087 | 2022-01-30 | 2020-09-28 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 11MM | 02-012-44-4011 |  | AS REPORTED, APPROXIMATELY FIVE YEARS POST INITIAL LEFT TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00086 | 2022-01-30 | 2020-07-08 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 2, 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 2, 13MM | 02-012-44-2013 |  | AS REPORTED, APPROXIMATELY TWO YEARS POST INITIAL RIGHT TKA, THIS FEMALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00089 | 2022-01-30 | 2020-09-21 | I, n, j, u, r, y | Manufacturer report | FEMORAL COMPONENTS | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | FEMORAL COMPONENTS | UNKNOWN |  | AS REPORTED, APPROXIMATELY FIVE YEARS POST INITIAL RIGHT TKA, THIS FEMALE PATIENT UNDERWENT REVISION OF THE DEVICE DUE TO INSTABILITY. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00094 | 2022-01-31 | 2019-07-16 | I, n, j, u, r, y | Manufacturer report | LOGIC | TIBIA PS MOD INSRT SZ 3.5 11MM | EXACTECH, INC. | 02-012-35-3511 | 02-012-35-3511 |  | AS REPORTED, APPROXIMATELY 5.38 YEARS SINCE INITIAL IMPLANT, THIS MALE PATIENT WITH BMI 21.7 UNDERWENT A RIGHT SIDE TKA AND WAS IMPLANTED WITH AN EXACTECH UHMWPE INSERT, SERIAL # (B)(4). ON (B)(6) 2019, THE PATIENT UNDERWENT REVISION DUE TO SWELLING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00095 | 2022-01-31 | 2022-01-11 | I, n, j, u, r, y | Manufacturer report | THREE PEG PATELLA 35MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | THREE PEG PATELLA 35MM | 200-02-35 |  | AS REPORTED, PATIENT HAD A REVISION DUE TO HOT SPOT UNDER PATELLA. PATIENT¿S PATELLA AND INSERT WERE REVISED DUE TO LOOSENING. THERE WAS NO BREAKAGE OF THE DEVICE. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO DEVICES WERE DISPOSED BY HOSPITAL. | CONCOMITANT MEDICAL PRODUCTS: LGC FEMORAL PS CEM LEFT SZ 2.5 (CAT# 02-010-01-0225 / SERIAL# (B)(4)); LOGIC TIBIA IMPLANT PSC INSERT, SZ 2.5, 9MM (CAT# 02-012-44-2509 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-01
1038671-2022-00112 | 2022-02-01 | 2021-04-14 | I, n, j, u, r, y | Manufacturer report | OPTETRAK | LOGIC TIBIA IMPLANT PSC INSERT, SZ 5, 11MM | EXACTECH, INC. | 02-012-44-5011 | 02-012-44-5011 |  | AS REPORTED, APPROXIMATELY 8.55 YEARS SINCE INITIAL IMPLANT, THIS MALE PATIENT WITH BMI 33.3 UNDERWENT A LEFT SIDE TKA AND WAS IMPLANTED WITH AN EXACTECH UHMWPE INSERT, SERIAL # (B)(4). ON (B)(6) 2021, THE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00113 | 2022-02-01 | 2021-01-05 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 5 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 5 11MM | 02-012-35-5011 |  | IT WAS REPORTED VIA LEGAL DOCUMENTATION THAT APPROXIMATELY 106 MONTHS AFTER A RIGHT TOTAL KNEE REPLACEMENT THE PATIENT UNDERWENT A REVISION SURGERY TO ADDRESS PROSTHESIS WEAR, PAIN, STIFFNESS, DISCOMFORT, WEAKNESS AND LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE. | THIS FOLLOW-UP REPORT IS BEING SUBMITTED TO RELAY ADDITIONAL AND/OR CORRECTED INFORMATION. THE FOLLOWING SECTIONS WERE UPDATED/CORRECTED: H6. MDR SECTION CODES UPDATED/CORRECTED: A, B, D. THE REASON FOR THE REVISION REPORTED CANNOT BE CONFIRMED FROM THE INFORMATION PROVIDED BUT MAY BE THE RESULT OF PROSTHESIS WEAR AND LOOSENING OR DUE TO INCLUSION OF THE POLYETHYLENE IN THE PACKAGING RECALL. POTENTIAL CONTRIBUTIONS OF USER AND PATIENT-RELATED CONSIDERATIONS TO THE EVENT COULD NOT BE ASSESSED AS THE DEVICES WERE NOT AVAILABLE FOR EVALUATION AND IMAGES, RADIOGRAPHS, AND RELEVANT CLINICAL INFORMATION WERE NOT PROVIDED. IF ANY FURTHER INFORMATION IS OBTAINED THAT WOULD CHANGE OR ALTER ANY INFORMATION PROVIDED, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00114 | 2022-02-01 | 2020-08-24 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 3 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 3 9MM | 02-012-35-3009 |  | AS REPORTED, APPROXIMATELY NINE YEARS POST INITIAL LEFT SIDE TKA, THIS FEMALE PATIENT UNDERWENT REVISION OF THE INSERT DUE TO PERSISTENT STIFFNESS. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00115 | 2022-02-01 | 2020-07-17 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 3 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 3 11MM | 02-012-35-3011 |  | AS REPORTED, APPROXIMATELY TEN YEARS POST INITIAL LEFT SIDE TKA, THIS FEMALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING . NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00116 | 2022-02-01 | 2020-11-02 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 3.5, 15MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 3.5, 15MM | 02-012-44-3515 |  | AS REPORTED, APPROXIMATELY 5 YEARS AND 7 MONTHS POST RIGHT SIDE TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | THE REVISION REPORTED WAS LIKELY THE RESULT OF AN INSUFFICIENT BOND BETWEEN THE IMPLANT AND THE BONE AND PATIENT-RELATED CONDITIONS, WHICH LED TO ASEPTIC (NON-INFECTED) LOOSENING. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICES WERE NOT AVAILABLE FOR EVALUATION AND PRE-REVISION X-RAYS WERE NOT PROVIDED. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00117 | 2022-02-01 | 2020-09-02 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 5, 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 5, 11MM | 02-012-44-5011 |  | AS REPORTED, APPROXIMATELY TWO YEARS POST INITIAL LEFT TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00120 | 2022-02-01 | 2020-07-15 | I, n, j, u, r, y | Manufacturer report | TRULIANT TIB IMP PS INSERT SZ 5 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT TIB IMP PS INSERT SZ 5 11MM | 02-022-35-5011 |  | AS REPORTED, APPROXIMATELY EIGHT MONTHS POST INITIAL RIGHT SIDE TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO INFECTION. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | AS REPORTED, APPROXIMATELY EIGHT MONTHS POST INITIAL RIGHT SIDE TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO INFECTION. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN OR MANUFACTURING ISSUES. THE CAUSE OF SURGICAL REVISION FOR INFECTION CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION. THESE DEVICES ARE USED FOR TREATMENT NOT DIAGNOSIS. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00119 | 2022-02-01 | 2021-01-19 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 5 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 5 13MM | 02-012-35-5013 |  | AS REPORTED, APPROXIMATELY 9 YEARS AND 10 MONTHS POST LEFT SIDE TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00125 | 2022-02-01 | 2020-01-07 | I, n, j, u, r, y | Manufacturer report | PS TIBIAL INSERTS SZ 3, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | PS TIBIAL INSERTS SZ 3, 9MM | 204-23-09 |  | AS REPORTED, APPROXIMATELY TWELVE YEARS POST INITIAL LEFT SIDE TKA, THIS FEMALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00122 | 2022-02-01 | 2020-11-11 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 3.5 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED, | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 3.5 13MM | 02-012-35-3513 |  | AS REPORTED, APPROXIMATELY 8 YEARS AND 3 MONTHS POST RIGHT SIDE TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING . NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00123 | 2022-02-01 | 2021-04-07 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 2, 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 2, 11MM | 02-012-44-2011 |  | ADDITIONAL INFORMATION RECEIVED FROM A LEGAL NOTIFICATION. PREOP DIAGNOSIS ¿ LEFT KNEE FEMORAL LOOSENING, OSTEOLYSIS. FINDINGS: GROSSLY LOOSE FEMUR COMPONENT, FEMORAL OSTEOLYSIS, WELL FIXED PATELLA AND TIBIA. INTRAOPERATIVE FROZEN SECTION NOT CONSISTENT WITH INFECTION. PATIENT PRESENTED TO CLINIC WITH LEFT KNEE PAIN. WORK UP REVEALED FEMORAL LOOSENING; C ACNES GREW OUT ON ONE CULTURE IN BROTH ONLY. BELIEVED THIS WAS MOST LIKELY A CONTAMINANT. DECISION TO PROCEED WITH REVISION TKA GIVEN THE PATIENT¿S SIGNIFICANT PAIN. THE FEMORAL COMPONENT WAS FOUND TO BE GROSSLY LOOSE AND WAS EASILY REMOVED WITH A TAMP. THERE WAS EXTENSIVE OSTEOLYSIS OF THE POSTERIOR AND DISTAL FEMORAL CONDYLES LATERALLY AND MEDIALLY. THE TIBIAL COMPONENT WAS FOUND TO BE WELL FIXED. THE PATELLA WAS WELL FIXED AND TRACKED WELL AND WAS NOT REVISED. THE PATIENT¿S OTHER COMPONENTS WERE REVISED TO A COMPETITOR¿S DEVICES. NO FURTHER INFORMATION. | THE REVISION REPORTED WAS LIKELY THE RESULT OF AN INSUFFICIENT BOND BETWEEN THE IMPLANT AND THE BONE, WHICH LED TO ASEPTIC (NON-INFECTED) LOOSENING. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICES WERE NOT AVAILABLE FOR EVALUATION AND PRE-REVISION X-RAYS WERE NOT PROVIDED. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00124 | 2022-02-01 | 2021-08-05 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 13MM | 02-012-44-4013 |  | AS REPORTED, APPROXIMATELY 4.32 YEARS POST RIGHT SIDE TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00118 | 2022-02-01 | 2020-08-31 | I, n, j, u, r, y | Manufacturer report | FEMORAL COMPONENTS | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | FEMORAL COMPONENTS | UNKNOWN |  | AS REPORTED, APPROXIMATELY EIGHT YEARS POST INITIAL RIGHT SIDE TKA, THIS MALE PATIENT UNDERWENT REVISION OF THE DEVICE DUE TO ASEPTIC LOOSENING AND OSTEOLYSIS. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION RECEIVED THE FOLLOWING SECTIONS D9, G3, G6, H2, H3 AND H6 HAVE BEEN UPDATED ACCORDINGLY. (H3) THE REVISION REPORTED WAS LIKELY THE RESULT OF AN INSUFFICIENT BOND BETWEEN THE IMPLANT AND THE BONE, WHICH LED TO ASEPTIC (NON-INFECTED) LOOSENING AND OSTEOLYSIS. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICES WERE NOT AVAILABLE FOR EVALUATION AND PRE-REVISION X-RAYS WERE NOT PROVIDED. THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE REPORTED EVENT OF "LOOSENING" IS ASSOCIATED WITH WEAKENED INTEGRATION OF THE DEVICE AT THE BONE-IMPLANT INTERFACE DUE TO LOSS OF FIBROUS AND/OR BONY TISSUE AND LEADING TO COMPROMISED ANCHORAGE OF THE DEVICE. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00111 | 2022-02-01 | 2019-10-02 | I, n, j, u, r, y | Manufacturer report | OPTETRAK | CC TIBIAL INSERT SZ 2, 13MM | EXACTECH, INC. | 208-22-13 | 208-22-13 |  | AS REPORTED, APPROXIMATELY 8.01 YEARS SINCE INITIAL IMPLANT, THIS FEMALE PATIENT WITH BMI 20.3 UNDERWENT A LEFT SIDE TKA AND WAS IMPLANTED WITH AN EXACTECH UHMWPE INSERT, SERIAL # (B)(4) ON (B)(6) 2019, THE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00121 | 2022-02-01 | 2020-08-05 | I, n, j, u, r, y | Manufacturer report | SPECIFIC DEVICE NOT REPORTED | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | SPECIFIC DEVICE NOT REPORTED | UNK |  | AS REPORTED, APPROXIMATELY 5 YEARS AND 7 MONTHS POST LEFT SIDE TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00109 | 2022-02-01 | 2021-06-02 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 15MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 15MM | 02-012-44-4015 |  | AS REPORTED, APPROXIMATELY 11 YEARS AND 4 MONTHS POST LEFT SIDE TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00110 | 2022-02-01 | 2021-04-28 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 2, 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 2, 11MM | 02-012-44-2011 |  | AS REPORTED, APPROXIMATELY 7 YEARS AND 6 MONTHS POST RIGHT SIDE TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00097 | 2022-02-01 | 2019-07-24 | I, n, j, u, r, y | Manufacturer report | LOGIC | TIBIA PS MOD INSRT SZ 1 9MM | EXACTECH, INC. | 02-012-35-1009 | 02-012-35-1009 |  | AS REPORTED, APPROXIMATELY 7.59 YEARS SINCE INITIAL IMPLANT, THIS FEMALE PATIENT WITH BMI 20.7 UNDERWENT A RIGHT SIDE TKA AND WAS IMPLANTED WITH AN EXACTECH UHMWPE INSERT, SERIAL #:(B)(4). ON (B)(6) 2019, THE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00098 | 2022-02-01 | 2019-07-19 | I, n, j, u, r, y | Manufacturer report | LOGIC | TIBIA IMPLANT PSC INSERT, SZ 2, 9MM | EXACTECH, INC. | 02-012-44-2009 | 02-012-44-2009 |  | AS REPORTED, APPROXIMATELY 5.08 YEARS SINCE INITIAL IMPLANT, THIS FEMALE PATIENT WITH BMI 31 UNDERWENT A LEFT SIDE TKA AND WAS IMPLANTED WITH AN EXACTECH UHMWPE INSERT, SERIAL # (B)(4). ON (B)(6) 2019 , THE PATIENT UNDERWENT REVISION DUE TO INSTABILITY. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00099 | 2022-02-01 | 2019-07-03 | I, n, j, u, r, y | Manufacturer report | OPTETRAK | CC TIBIAL INSERT SZ 2, 13M | EXACTECH, INC. | 208-22-13 | 208-22-13 |  | AS REPORTED, APPROXIMATELY 18.20 YEARS SINCE INITIAL IMPLANT, THIS FEMALE PATIENT WITH BMI 27.6 UNDERWENT A RIGHT SIDE TKA AND WAS IMPLANTED WITH AN EXACTECH UHMWPE INSERT, SERIAL # :(B)(4). ON (B)(6) 2019, THE PATIENT UNDERWENT REVISION DUE TO PERSISTENT STIFFNESS AND INSTABILITY. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00100 | 2022-02-01 | 2019-07-16 | I, n, j, u, r, y | Manufacturer report | OPTETRAK LOGIC | TIBIA PS MOD INSRT SZ 3 13MM | EXACTECH, INC. | 02-012-35-3013 | 02-012-35-3013 |  | AS REPORTED, APPROXIMATELY 5.61 YEARS SINCE INITIAL IMPLANT, THIS MALE PATIENT WITH BMI 21.7 UNDERWENT A LEFT SIDE TKA AND WAS IMPLANTED WITH AN EXACTECH UHMWPE INSERT, SERIAL # (B)(4). ON (B)(6) 2019, THE PATIENT UNDERWENT REVISION DUE TO SWELLING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. | PENDING EVALUATION. NO INFORMATION PROVIDED. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00101 | 2022-02-01 | 2022-01-06 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 38MM GLENOSPHERE | EXACTECH, INC. | 320-01-38 | 320-01-38 |  | AS REPORTED, APPROXIMATELY 1.5 MONTHS POST OF THE INITIAL LEFT TSA, THIS FEMALE PATIENT WAS REVISED DUE TO THE GLENOSPHERE BEING DISCONNECTED FROM THE BASEPLATE. THIS CAUSED THE PATIENT DISCOMFORT IN HER FUNCTION. THE SURGEON REVIEWED HER AND MANAGED TO REDUCE IT, ENSURING IT SAT AGAIN IN ITS NORMAL PLACE BY FLUOROSCOPY. THE REVISION WAS PLANNED TO DETERMINE THE REASON FOR THE DISCONNECTION, BUT WHEN HE OPENED, THE GLENOSPHERE WAS IN PLACE AND WELL SCREWED. SINCE THE SURGEON HAD TO UNSCREW THE GLENOSPHERE AND LOCKING SCREW TO ENSURE THERE WAS NOT DAMAGE, HE CHANGED THEM. PATIENT REQUIRED A REVISION (OUTPATIENT) AS A DIRECT RESULT OF THIS ISSUE. THERE WAS NOT SURGICAL DELAY/PROLONGATION. DEVICES WILL NOT BE RETUNING DUE TO THEM BEING THROWN AWAY DURING THE TIME OF THE SURGERY. | PENDING EVALUATION. CONCOMITANT DEVICE(S): 320-15-05, (B)(4) - EQ REV LOCKING SCREW. 320-15-01, (B)(4) - EQ REV GLENOID PLATE. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00096 | 2022-02-01 | 2019-10-24 | I, n, j, u, r, y | Manufacturer report | OPTETRAK | CC TIBIAL INSERT SZ 4, 9MM | EXACTECH, INC. | 208-24-09 | 208-24-09 |  | AS REPORTED, APPROXIMATELY 6.13 YEARS SINCE INITIAL IMPLANT, THIS MALE PATIENT WITH BMI 30.8 UNDERWENT LEFT SIDE TKA AND WAS IMPLANTED WITH AN EXACTECH UHMWPE INSERT, SERIAL #: (B)(4). ON (B)(6) 2019, THE PATIENT UNDERWENT REVISION DUE TO INSTABILITY AND POLYETHYLENE WEAR. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00103 | 2022-02-01 | 2022-01-12 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | CAGED GLENOID | EXACTECH, INC. |  | UNK |  | AS REPORTED, THIS (B)(6) PATIENT'S RIGHT SHOULDER WAS REVISED DUE TO IMPLANT FAILURE. THE POLY SHEERED OFF CAGE GLENOID. PATIENT WAS NOT REVISED TO EXACTECH DEVICES. THERE WAS NOT DELAY. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETUNING DUE TO HOSPITAL POLICY. | PENDING EVALUATION. CONCOMITANT MEDICAL PRODUCTS: 310-01-50, 5579902 - EQUINOXE, HUMERAL HEAD SHORT, 50MM (BETA); 300-10-15, 51707001 - EQUINOXE REPLICATOR PLATE 1.5MM O/S. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00104 | 2022-02-01 | 2021-04-02 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 2.5, 15MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 2.5, 15MM | 02-012-44-2515 |  | AS REPORTED, APPROXIMATELY 6 YEARS AND 10 MONTHS POST RIGHT SIDE TKA, THIS FEMALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00105 | 2022-02-01 | 2021-01-06 | I, n, j, u, r, y | Manufacturer report | SPECIFIC DEVICE NOT REPORTED | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | SPECIFIC DEVICE NOT REPORTED | UNK |  | AS REPORTED, APPROXIMATELY 6 YEARS POST RIGHT SIDE TKA, THIS FEMALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00106 | 2022-02-01 | 2020-02-05 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 2, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 2, 9MM | 02-012-44-2009 |  | AS REPORTED, APPROXIMATELY 6 YEARS POST LEFT SIDE TKA, THIS FEMALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00107 | 2022-02-01 | 2020-07-02 | I, n, j, u, r, y | Manufacturer report | CC TIBIAL INSERT SZ 2, 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | CC TIBIAL INSERT SZ 2, 13MM | 208-22-13 |  | IT WAS REPORTED VIA LEGAL DOCUMENTATION APPROXIMATELY 32 MONTHS AFTER A RIGHT TOTAL KNEE REPLACEMENT THE PATIENT UNDERWENT A REVISION PROCEDURE TO ADDRESS PROSTHESIS WEAR, PAIN AND RIGHT TOTAL KNEE TIBIAL COMPONENT LOOSENING. A REVISION OPERATIVE REPORT WAS PROVIDED. INTRAOPERATIVELY THE SURGEON OBSERVED CLEAR AND GOLDEN YELLOW JOINT FLUID WITH NO SIGNS OF PURULENCE. THERE WERE SYNOVIAL FIBRILLATIONS CONSISTENT WITH WEAR DEBRIS IRRITATION TO THE MEDIAL AND LATERAL GUTTERS. PATHOLOGY RESULTS CAME BACK AS NON-CONCERNING FOR INFECTION. THE POLYETHYLENE INSERT WAS REMOVED WITHOUT INCIDENT. THE TIBIAL COMPONENT WAS NOTED TO BE SIGNIFICANTLY LOOSE. FEMORAL AND PATELLA COMPONENTS WERE NOTED TO BE WELL FIXED. THERE WERE CYSTS MEDIALLY ON THE PROXIMAL TIBIA WITH NO SIGNIFICANT BONE LOSS. NO SURGICAL OR MEDICAL INTERVENTIONS WERE REPORTED. NO ADDITIONAL INFORMATION IS AVAILABLE. | THIS FOLLOW-UP REPORT IS BEING SUBMITTED TO RELAY ADDITIONAL AND/OR CORRECTED INFORMATION. THE REASON FOR THE REVISION REPORTED CANNOT BE CONFIRMED FROM THE INFORMATION PROVIDED BUT MAY BE THE RESULT OF PROSTHESIS WEAR AND TIBIAL LOOSENING. POTENTIAL CONTRIBUTIONS OF USER AND PATIENT-RELATED CONSIDERATIONS TO THE EVENT COULD NOT BE ASSESSED AS THE DEVICES WERE NOT AVAILABLE FOR EVALUATION AND IMAGES AND RADIOGRAPHS WERE NOT PROVIDED IF ANY FURTHER INFORMATION IS OBTAINED THAT WOULD CHANGE OR ALTER ANY INFORMATION PROVIDED, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00108 | 2022-02-01 | 2021-05-04 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 4 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 4 9MM | 02-012-35-4009 |  | AS REPORTED, APPROXIMATELY 2 YEARS AND 6 MONTHS POST RIGHT SIDE TKA, THIS MALE PATIENT UNDERWENT REVISION OF THE INSERT DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00102 | 2022-02-01 | 2021-12-01 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | 7X80MM DISTAL STEM MODULAR CEMENTED | EXACTECH, INC. | 308-01-07 | 308-01-07 |  | AS REPORTED, APPROXIMATELY 5 MONTHS POST-OP THE INITIAL TSA, THIS (B)(6) FEMALE PATIENT PRESENTED ASEPTIC HUMERAL LOOSENING WITH PAIN AND SWELLING IN THE ARM. NO TRAUMATIC EVENT WAS REPORTED. THE CASE REPORT FORM INDICATES THIS EVENT IS POSSIBLY RELATED TO DEVICES AND/OR PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. NO ACTION WAS TAKEN, PATIENT IS WEARING A BRACE. OUTCOME IS REPORTED AS "ONGOING". | PENDING EVALUATION. CONCOMITANT DEVICE(S): 320-36-10 - 36MM HUMERAL LINER +0MM CONSTRAINED, 320-10-00 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0, 320-35-02 - SMALL SUPERIOR AUGMENT GLENOID PLATE, 308-05-19 - DISTAL FIXATION RING HA 19.5, 308-09-00 - SMALL PROX BODY +0, 320-32-36 - EXPANDED GLENOSPHERE, 36MM, FOR SMALL REVERSE. NO INFORMATION PROVIDED. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00140 | 2022-02-02 | 2020-09-21 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 2 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 2 11MM | 02-012-35-2011 |  | AS REPORTED, APPROXIMATELY 9 YEARS AND 7 MONTHS POST LEFT SIDE TKA, THIS FEMALE PATIENT UNDERWENT REVISION DUE TO INFECTION. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00141 | 2022-02-02 | 2020-01-23 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 5, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 5, 9MM | 02-012-44-5009 |  | AS REPORTED, APPROXIMATELY FIVE YEARS POST INITIAL LEFT TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00142 | 2022-02-02 | 2022-01-12 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 15MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 15MM | 02-012-44-4015 |  | AS REPORTED, PATIENT PRESENTED WITH INSTABILITY AND A POLY EXCHANGE WAS PERFORMED WHERE THE SURGEON WENT TO A THICKER INSERT. THE TIBIA AND FEMUR WERE WELL FIXED. IMAGES WERE RECEIVED. NO ADDITIONAL INFORMATION AVAILABLE. | CONCOMITANT MEDICAL PRODUCTS: LOGIC FEMORAL PS CEM LEFT SZ 4 (CAT# 02-010-01-0240 / SERIAL# (B)(4)); LGC TIBIAL FIT TRAY CEM SZ 4F / 4T (CAT# 02-012-45-4040 / SERIAL# (B)(4)); THREE PEG PATELLA 35MM (CAT# 200-02-35 / SERIAL# (B)(4)); COLLAR FIXATION PIN 2PK 40MM, QUICK RELEASE (CAT# 201-78-82 / SERIAL# (B)(4)); 3" DRILL BIT, MOD. HEX 2 PACK (CAT# 201-78-89 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00144 | 2022-02-02 | 2021-11-01 | I, n, j, u, r, y | Manufacturer report | TRULIANT CR CEM FEM CR CEM RIGHT SZ 2.5 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT CR CEM FEM CR CEM RIGHT SZ 2.5 | 02-020-13-0325 |  | AS REPORTED BY (B)(6) STUDY, APPROXIMATELY THREE MONTHS POST RIGHT TKA, PATIENT HAS INCREASED PAIN AND SWELLING WHEN BENDING KNEE. A JOINT ASPIRATION DONE, SYNOVIAL FLUID ANALYSIS WAS PERFORMED AND LABS SHOW POSITIVE FOR PERSISTENT INFECTION (CONTINUED CHRONIC PROSTHETIC JOINT INFECTION). SURGEON PERFORMED A REVISION SURGERY ON (B)(6) 2021 WITH REMOVAL OF IMPLANTS, PLACEMENT OF SPACER, AND IMPLANTED REMEDY MANUFACTURER COMPONENTS. PATIENT HAS BEEN STARTED ON 6-WEEKS IV ANTIBIOTICS. THE EVENT IS UNLIKELY RELATED TO THE DEVICE BUT RELATED TO THE PROCEDURE. | CONCOMITANT MEDICAL PRODUCTS: TRULIANT TIB IMP CRC INSERT SZ 2.5, 15MM (CAT # 02-022-51-2515 / SERIAL# (B)(4)), TRULIANT TIB FIT TRAY CEM SZ 2.5F / 2.5T (CAT # 02-022-45-2525 / CAT # 6758300). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00147 | 2022-02-02 | 2020-09-21 | I, n, j, u, r, y | Manufacturer report | PS TIBIAL INSERTS SZ 4, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | PS TIBIAL INSERTS SZ 4, 9MM | 204-24-09 |  | AS REPORTED, APPROXIMATELY 1 YEAR POST RIGHT SIDE TKA, THIS FEMALE PATIENT UNDERWENT REVISION DUE TO PERSISTENT STIFFNESS. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00146 | 2022-02-02 | 2020-08-07 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 11MM | 02-012-44-4011 |  | AS REPORTED, APPROXIMATELY 4 YEARS AND 7 MONTHS POST RIGHT SIDE TKA, THIS FEMALE PATIENT UNDERWENT REVISION DUE TO INFECTION. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | H2: THESE DEVICES ARE USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN OR MANUFACTURING ISSUES. THE CAUSE OF THE SUBSEQUENT REVISION CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00148 | 2022-02-02 | 2020-05-11 | I, n, j, u, r, y | Manufacturer report | CC TIBIAL INSERT SZ 4, 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | CC TIBIAL INSERT SZ 4, 13MM | 208-24-13 |  | AS REPORTED, APPROXIMATELY 2.52 YEARS POST LEFT SIDE TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO PERSISTENT STIFFNESS. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. 510K: K954208. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00149 | 2022-02-02 | 2022-01-13 | I, n, j, u, r, y | Manufacturer report | FEMORAL COMPONENTS | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | FEMORAL COMPONENTS | UNKNOWN |  | AS REPORTED, THE PATIENT UNDERWENT REVISION OF THE INSERT DUE TO UNKNOWN REASON. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00150 | 2022-02-02 | 2021-03-19 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 3 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 3 11MM | 02-012-35-3011 |  | AS REPORTED, APPROXIMATELY 10 YEARS AND 3 MONTHS POST LEFT SIDE TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO PERIPROSTHETIC FRACTURE. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00139 | 2022-02-02 | 2021-05-21 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 4 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 4 9MM | 02-012-35-4009 |  | AS REPORTED, APPROXIMATELY 9.51 YEARS POST RIGHT SIDE TKA, THIS PATIENT UNDERWENT REVISION DUE POLYETHYLENE WEAR. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00145 | 2022-02-02 | 2021-04-13 | I, n, j, u, r, y | Manufacturer report | LGC TIBIAL FIT TRAY CEM SZ 4F / 4T | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LGC TIBIAL FIT TRAY CEM SZ 4F / 4T | 02-012-45-4040 |  | AS REPORTED BY THE EXACTECH KNEE REPLACEMENT STUDY, APPROXIMATELY FOUR YEARS POST INITIAL TKA, THIS (B)(6) FEMALE PATIENT EXPERIENCED ASEPTIC TIBIAL LOOSENING. PATIENT CAME IN WITH INCREASED PAIN. X-RAYS SHOW LOOSENING OF TIBIAL COMPONENT WITH MEDIAL & POSTERIOR SUBSIDENCE AND LUCENCY AROUND COMPONENT. ON (B)(6) 2021, PATIENT SAW SURGEON AND KNEE WAS ASPIRATED AND SENT FOR CULTURES. ON (B)(6) 2021, IT'S CONCLUDED NEGATIVE FOR INFECTION. A REVISION WAS DONE ON (B)(6) 2021 USING SMITH-NEPHEW PRODUCTS. THE EVENT IS UNLIKELY RELATED TO THE DEVICE BUY DEFINITELY RELATED TO THE PROCEDURE. | (H3) THE REVISION REPORTED WAS LIKELY THE RESULT OF AN INSUFFICIENT BOND BETWEEN THE TIBIAL TRAY AND THE BONE, POOR BONE QUALITY, PATIENT-RELATED CONDITIONS, OR ANY COMBINATION OF THESE POSSIBILITIES, WHICH LED TO ASEPTIC (NON-INFECTED) TIBIAL LOOSENING AND TIBIAL SUBSIDENCE. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICES AND PRE-REVISION X-RAYS WERE NOT AVAILABLE FOR EVALUATION. THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE REPORTED EVENT OF "LOOSENING - TIBIAL¿ IS ASSOCIATED WITH WEAKENED INTEGRATION OF THE TIBIAL COMPONENT AT THE BONE-IMPLANT INTERFACE DUE TO LOSS OF FIBROUS AND/OR BONY TISSUE AND LEADING TO COMPROMISED ANCHORAGE OF THE DEVICE. FURTHERMORE, THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE REPORTED EVENT OF "TIBIAL SUBSIDENCE¿ IS ASSOCIATED WITH A TIBIAL TRAY THAT SUBSIDES INTO THE PREPARED TIBIA FURTHER THAN EXPECTED. THIS TYPICALLY OCCURS POST-OPERATIVELY. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00138 | 2022-02-02 | 2021-01-14 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 3 15MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | 9999 | 9999 |  | AS REPORTED, APPROXIMATELY 4 YEARS AND 1 MONTH POST RIGHT SIDE TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO POLYETHYLENE WEAR. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | CONCOMITANT MEDICAL PRODUCTS: 02-012-35-3015/2024607/LOGIC TIBIA PS MOD INSRT SZ 3 15MM/510K K033883. ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00143 | 2022-02-02 | 2021-05-27 | I, n, j, u, r, y | Manufacturer report | OPTETRAK LOGIC | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | PS TIBIAL INSERTS SZ 4, 9MM | 204-24-09 |  | AS REPORTED, APPROXIMATELY 12 YEARS AND 8 MONTHS POST RIGHT SIDE TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO INFECTION. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00136 | 2022-02-02 | 2022-01-14 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 38MM HUMERAL LINER +2.5 | EXACTECH, INC. | 320-38-03 | 320-38-03 |  | AS REPORTED, APPROXIMATLY 8 MONTHS POST OP RTSA, THIS 80 Y/O MALE PATIENT WAS REVISED DUE TO AN INFECTION. THE SURGEON WASHED OUT AND EXCHANGED THE LINER, ADAPTER TRAY AND GLENOSPHERE. THERE WAS NO SURGICAL DELAY. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETUNING DUE TO THE HOSPITAL DISPOSING OF THEM. | DEVICE EVALUATED BY MFR: PENDING EVALUATION. CONCOMITANT DEVICE(S): 300-01-13, 6850563 - EQUINOXE, HUMERAL STEM PRIMARY, PRESS FIT 13MM, 320-10-00, 6862676 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0, 320-15-01, 6887213 - EQ REV GLENOID PLATE, 320-15-05, 6896642 - EQ REV LOCKING SCREW, 320-20-00, 6895957 - EQ REVERSE TORQUE DEFINING SCREW KIT, 320-20-26, S214748 - EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 26MM, 320-20-30, S222366 - EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 30MM, 320-02-38, 6574296 - RS EXPANDED GLENOSPHERE 38MM, +4MM OFFSET. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00126 | 2022-02-02 | 2021-08-11 | I, n, j, u, r, y | Manufacturer report | OPTETRAK | LOGIC TIBIA PS MOD INSRT SZ 3 11MM | EXACTECH, INC. | 02-012-35-3011 | 02-012-35-3011 |  | AS REPORTED, APPROXIMATELY 8.83 YEARS SINCE INITIAL IMPLANT, THIS FEMALE PATIENT WITH BMI 19.8 UNDERWENT A RIGHT SIDE TKA AND WAS IMPLANTED WITH AN EXACTECH UHMWPE INSERT, SERIAL # (B)(4). ON (B)(6) 2021, THE PATIENT UNDERWENT REVISION DUE TO A COMPONENT FAILURE. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. | DEVICE EVALUATED BY MANUFACTURER? PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00127 | 2022-02-02 | 2021-07-08 | I, n, j, u, r, y | Manufacturer report | OPTETRAK | LOGIC TIBIA PS MOD INSRT SZ 2 13MM | EXACTECH, INC. | 02-012-35-2013 | 02-012-35-2013 |  | AS REPORTED, APPROXIMATELY 10.33 YEARS SINCE INITIAL IMPLANT, THIS FEMALE PATIENT WITH BMI 20.6 UNDERWENT A RIGHT SIDE TKA AND WAS IMPLANTED WITH AN EXACTECH UHMWPE INSERT, SERIAL # (B)(4). ON (B)(6) 2021, THE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. | DEVICE EVALUATED BY MANUFACTURER? PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00137 | 2022-02-02 | 2020-01-29 | I, n, j, u, r, y | Manufacturer report | CC TIBIAL INSERT SZ 4, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | CC TIBIAL INSERT SZ 4, 9MM | 208-24-09 |  | AS REPORTED, APPROXIMATELY TWENTY YEARS POST INITIAL LEFT SIDE TKA, THIS FEMALE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING AND OSTEOLYSIS. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00129 | 2022-02-02 | 2021-09-01 | I, n, j, u, r, y | Manufacturer report | OPTETRAK | LOGIC TIBIA IMPLANT PSC INSERT, SZ 3, 9MM | EXACTECH, INC. | 02-012-44-3009 | 02-012-44-3009 |  | AS REPORTED, APPROXIMATELY 7.36 YEARS SINCE INITIAL IMPLANT, THIS FEMALE PATIENT WITH BMI 25.7 UNDERWENT A RIGHT SIDE TKA AND WAS IMPLANTED WITH AN EXACTECH UHMWPE INSERT, SERIAL # (B)(4). ON (B)(6) 2021, THE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00130 | 2022-02-02 | 2021-07-29 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED, POLYMER/METAL | EXACTECH, INC. | 02-012-44-4009 | 02-012-44-4009 |  | AS REPORTED, APPROXIMATELY 6.72 YEARS SINCE INITIAL IMPLANT, THIS MALE PATIENT WITH BMI 36.3 UNDERWENT A LEFT SIDE TKA AND WAS IMPLANTED WITH AN EXACTECH UHMWPE INSERT, SERIAL # (B)(4). ON (B)(6) 2021, THE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00128 | 2022-02-02 | 2021-07-23 | I, n, j, u, r, y | Manufacturer report | OPTETRAK | CC TIBIAL INSERT SZ 2, 11MM | EXACTECH, INC. | 208-22-11 | 208-22-11 |  | AS REPORTED, APPROXIMATELY 5.29 YEARS SINCE INITIAL IMPLANT, THIS FEMALE PATIENT WITH BMI 18.7 UNDERWENT A (LEFT) SIDE TKA AND WAS IMPLANTED WITH AN EXACTECH UHMWPE INSERT, SERIAL # (B)(4). ON (B)(6) 2021, THE PATIENT UNDERWENT REVISION DUE TO INSTABILITY. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. | DEVICE EVALUATED BY MFR: PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00132 | 2022-02-02 | 2021-12-15 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 34MM | EXACTECH, INC. | 320-20-34 | 320-20-34 |  | AS REPORTED, THIS (B)(6) Y/O MALE PATIENT'S LEFT SHOULDER WAS REVISED DUE TO THE BASEPLATE SCREWS BREAKING. THE BASEPLATE, GLENOSPHERE, TRAY, AND LINER WERE REVISED. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETUNING DUE TO HOSPITAL POLICY. | DEVICE EVALUATED BY MANUFACTURER? PENDING EVALUATION. CONCOMITANT MEDICAL DEVICE(S): 320-01-42 - EQUINOXE REVERSE 42MM GLENOSPHERE, (B)(4); 320-10-00 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0, (B)(4); 320-15-03 - RS GLENOID PLATE L POST AUG, 8 DEG, LEFT, (B)(4); 320-15-05 - EQ REV LOCKING SCREW, (B)(4); 320-20-00 - EQ REVERSE TORQUE DEFINING SCREW KIT, (B)(4); 320-20-26 - EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 26MM, (B)(4); 320-42-00 - EQUINOXE REVERSE 42MM HUMERAL LINER +0, (B)(4). | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00133 | 2022-02-02 | 2021-12-15 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 34MM | EXACTECH, INC. | 320-20-34 | 320-20-34 |  | AS REPORTED, THIS (B)(6) MALE PATIENT'S LEFT SHOULDER WAS REVISED DUE TO THE BASEPLATE SCREWS BREAKING. THE BASEPLATE, GLENOSPHERE, TRAY, AND LINER WERE REVISED. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETUNING DUE TO HOSPITAL POLICY. | SECTION H10: (H3) THE REVISION REPORTED WAS LIKELY THE RESULT OF THE GLENOID BASEPLATE NOT ACHIEVING LONG-TERM FIXATION, CAUSING THE COMPRESSION SCREWS TO BEAR THE JOINT LOAD AND EVENTUALLY FRACTURE. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICES WERE NOT RETURNED FOR EVALUATION AND ADEQUATE INFORMATION WAS NOT PROVIDED TO DETERMINE THE ROOT CAUSE OF THE INSUFFICIENT GLENOID BASEPLATE FIXATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00134 | 2022-02-02 | 2020-07-08 | I, n, j, u, r, y | Manufacturer report | CC TIBIAL INSERT SZ 2, 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | CC TIBIAL INSERT SZ 2, 13MM | 208-22-13 |  | AS REPORTED, APPROXIMATELY FIVE MONTHS POST INITIAL LEFT SIDE TKA, THIS FEMALE PATIENT UNDERWENT REVISION DUE TO INFECTION AND PERIPROSTHETIC JOINT INFECTION. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00135 | 2022-02-02 | 2022-01-13 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | MDLR RV RMR 46MM PLT TIP | EXACTECH, INC. |  | 321-25-46 |  | IT WAS REPORTED THAT DURING SURGICAL USE, THE TIP OF THE REAMER BROKE OFF WHILE REAMING THE GLENOID. BROKEN PIECE WAS EASILY REMOVED. NO TIME WAS ADDED TO THE SURGERY AND PATIENT LEFT THE OR IN GOOD HEALTH. | SECTION H10:(H3) THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE REAMER DURING USE, WHICH LET TO BRITTLE FRACTURE OF THE PILOT TIP. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00131 | 2022-02-02 | 2020-12-28 | I, n, j, u, r, y | Manufacturer report | OPTETRAK | LOGIC TIBIA PS MOD INSRT SZ 3 11MM | EXACTECH, INC. | 02-012-35-3011 | 02-012-35-3011 |  | AS REPORTED, APPROXIMATELY 10.76 YEARS SINCE INITIAL IMPLANT, THIS MALE PATIENT WITH BMI 26.4 UNDERWENT A LEFT SIDE TKA AND WAS IMPLANTED WITH AN EXACTECH UHMWPE INSERT, SERIAL #:(B)(4). ON (B)(6) 2020, THE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. 510K: K033883. | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00161 | 2022-02-03 | 2022-01-14 | I, n, j, u, r, y | Manufacturer report | PS TIBIAL INSERTS SZ 4, 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | PS TIBIAL INSERTS SZ 4, 11MM | 204-24-11 |  | AS REPORTED PER SALES REP, APPROXIMATELY FIVE YEARS POST INITIAL TKA, THIS PATIENT HAS A SCHEDULE KNEE REVISION ON (B)(6) 2022. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00160 | 2022-02-03 | 2022-01-07 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | GLENOSPHERE EXP 38MM +4MM OFFSET | EXACTECH, INC. | 320-08-38 | 320-08-38 |  | AS REPORTED, APPROXIMATELY 2 MONTHS POST OP INITIAL TSA, THIS (B)(6) Y/O MALE PATIENT WAS REVISED DUE TO DISLOCATION. ON (B)(6) 2022, IN SUBJECT'S SLEEP, JERKED ARM, AND WOKE UP W/ SEVERE PAIN AND WAS UNABLE TO MOVE ARM. AT CLINIC VISIT ON (B)(6) 2022, X-RAYS SHOWED ANTERIOR DISLOCATION. REVISION SCHEDULED FOR THAT NIGHT. THE CASE REPORT FORM INDICATES THIS EVENT IS DEFINITELY NOT RELATED TO DEVICES OR PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | PENDING EVALUATION. CONCOMITANT DEVICE(S): 300-01-13 - EQUINOXE, HUMERAL STEM PRIMARY, PRESS FIT 13MM 320-38-03 - EQUINOXE REVERSE 38MM HUMERAL LINER +2.5 320-10-00 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0 320-15-08 - SUP/POST AUG PLATE, R RS GLENOID BASEPLATE | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00159 | 2022-02-03 | 2021-04-08 | I, n, j, u, r, y | Manufacturer report | CC TIBIAL INSERT SZ 5, 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | CC TIBIAL INSERT SZ 5, 11MM | 208-25-11 |  | AS REPORTED, APPROXIMATELY 16.02 YEARS POST RIGHT SIDE TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO OSTEOLYSIS. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00158 | 2022-02-03 | 2022-01-14 | I, n, j, u, r, y | Manufacturer report | VANTAGE | TIBIAL PLATE MB SZ 4 RT | EXACTECH, INC. |  | 350-32-04 |  | IT WAS REPORTED THAT A YOUNG MALE PATIENT FORMED BONE CYSTS POST INITIAL TOTAL ANKLE REPLACEMENT. THERE WAS NO INFECTION PRESENT. REVISION HAD NOT YET OCCURRED. | PENDING EVALUATION. CONCOMITANT DEVICE(S): 350-02-03E, 6727809 - TALUS -RIGHT- SZ 3 - EU. 350-42-03, 6356286 - TIBIAL INSERT MB SZ 3 RT 6MM. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00156 | 2022-02-03 | 2021-01-13 | I, n, j, u, r, y | Manufacturer report | SPECIFIC DEVICE NOT REPORTED | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | SPECIFIC DEVICE NOT REPORTED | UNK |  | AS REPORTED, UNKNOWN TIME SINCE INITIAL TKA, THIS PATIENT UNDERWENT REVISION DUE TO AN UNKNOWN REASON. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00157 | 2022-02-03 | 2021-05-24 | I, n, j, u, r, y | Manufacturer report | SPECIFIC DEVICE NOT REPORTED | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | SPECIFIC DEVICE NOT REPORTED | 9999 |  | AS REPORTED, UNKNOWN TIME SINCE INITIAL TKA, THIS PATIENT UNDERWENT REVISION DUE TO AN UNKNOWN REASON. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00154 | 2022-02-03 | 2020-08-04 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 5, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 5, 11MM | 02-012-44-5009 |  | AS REPORTED, APPROXIMATELY 6.77 YEARS POST RIGHT SIDE TKA, THIS MALE PATIENT UNDERWENT REVISION OF THE INSERT DUE TO DEBRIS SYNOVITIS. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00153 | 2022-02-03 | 2021-06-17 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 3, 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 3, 13MM | 02-012-44-3013 |  | AS REPORTED, APPROXIMATELY 2.53 YEARS POST INITIAL LEFT SIDE TKA, THIS FEMALE PATIENT UNDERWENT REVISION DUE TO INSTABILITY. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00152 | 2022-02-03 | 2020-08-17 | I, n, j, u, r, y | Manufacturer report | TRULIANT TIB IMP PS INSERT SZ 4 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT TIB IMP PS INSERT SZ 4 11MM | 02-022-35-4011 |  | AS REPORTED, APPROXIMATELY 1 YEAR AND 8 MONTHS POST INITIAL RIGHT TKA, THIS MALE PATIENT UNDERWENT REVISION DUE TO INSTABILITY. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00151 | 2022-02-03 | 2021-01-05 | I, n, j, u, r, y | Manufacturer report | SPECIFIC DEVICE NOT REPORTED | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | SPECIFIC DEVICE NOT REPORTED | UNK |  | AS REPORTED, APPROXIMATELY 8 YEARS AND 4 MONTHS POST RIGHT SIDE TKA, THIS FEMALE PATIENT UNDERWENT REVISION DUE TO INSTABILITY. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00155 | 2022-02-03 | 2021-04-09 | I, n, j, u, r, y | Manufacturer report | PS TIBIAL INSERTS SZ 6, 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | PS TIBIAL INSERTS SZ 6, 13MM | 204-26-13 |  | AS REPORTED, APPROXIMATELY 11.43 YEARS POST TKA, THIS PATIENT UNDERWENT A REVISION FOR AN UNKNOWN REASON. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00168 | 2022-02-04 | 2022-01-12 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 46MM HUMERAL LINER +2.5 | EXACTECH, INC. | 320-46-03 | 320-46-03 |  | AS REPORTED, THIS (B)(6) MALE PATIENT'S LEFT SHOULDER HAD INSTABILITY AND NEEDED A REVISION. PATIENT APPEARED TO HAVE SHOULDER "CREEP" AND WAS DISLOCATING WITH INSTABILITY. A REVISION OCCURRED AND A 42 +4 GLENOSPHERE WAS PUT ON WITH A +10 ADAPTOR TRAY AND 42 +0 POLY. PATIENT WAS RELOCATED IN OFFICE AND THEN DISLOCATED AGAIN AND WAS SCHEDULED FOR A REVISION. NO DELAY TO SURGERY. DEVICE WAS THROWN AWAY BY HOSPITAL, UNABLE TO RETRIEVE. | PENDING EVALUATION. NO INFORMATION PROVIDED. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00169 | 2022-02-04 | 2021-08-11 | I, n, j, u, r, y | Manufacturer report | OPTETRAK | LOGIC TIBIA PS MOD INSRT SZ 3 11MM | EXACTECH, INC. | 02-012-35-3011 | 02-012-35-3011 |  | AS REPORTED, APPROXIMATELY 10.49 YEARS SINCE INITIAL IMPLANT, THIS FEMALE PATIENT WITH BMI 19.8 UNDERWENT A RIGHT SIDE TKA AND WAS IMPLANTED WITH AN EXACTECH UHMWPE INSERT, SERIAL # (B)(4). ON (B)(6) 2021, THE PATIENT UNDERWENT REVISION DUE TO A COMPONENT FAILURE. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. | DEVICE EVALUATED BY MANUFACTURER? PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00170 | 2022-02-04 | 2021-06-22 | I, n, j, u, r, y | Manufacturer report | OPTETRAK | CC TIBIAL INSERT SZ 4, 11MM | EXACTECH, INC. | 208-24-11 | 208-24-11 |  | AS REPORTED, APPROXIMATELY 11.60 YEARS SINCE INITIAL IMPLANT, THIS MALE PATIENT WITH BMI 23.8 UNDERWENT A LEFT SIDE TKA AND WAS IMPLANTED WITH AN EXACTECH UHMWPE INSERT, SERIAL # (B)(4). ON (B)(6) 2021, THE PATIENT UNDERWENT REVISION DUE TO OSTEOLYSIS. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00171 | 2022-02-04 | 2021-06-22 | I, n, j, u, r, y | Manufacturer report | OPTETRAK | LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 9MM | EXACTECH, INC. | 02-012-44-4009 | 02-012-44-4009 |  | AS REPORTED, APPROXIMATELY 4.01 YEARS SINCE INITIAL IMPLANT, THIS MALE PATIENT WITH BMI 34.6 UNDERWENT A LEFT SIDE TKA AND WAS IMPLANTED WITH AN EXACTECH UHMWPE INSERT, SERIAL # (B)(4). ON (B)(6) 2021, THE PATIENT UNDERWENT REVISION DUE TO PERSISTENT STIFFNESS. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00172 | 2022-02-04 | 2021-08-03 | I, n, j, u, r, y | Manufacturer report | OPTETRAK | LOGIC TIBIA IMPLANT PSC INSERT, SZ 1, 9MM | EXACTECH, INC. | 02-012-44-1009 | 02-012-44-1009 |  | AS REPORTED, APPROXIMATELY 8.05 YEARS SINCE INITIAL IMPLANT, THIS FEMALE PATIENT WITH BMI 24.9 UNDERWENT A LEFT SIDE TKA AND WAS IMPLANTED WITH AN EXACTECH UHMWPE INSERT, SERIAL # (B)(4). ON (B)(6) 2021, THE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00173 | 2022-02-04 | 2021-06-30 | I, n, j, u, r, y | Manufacturer report | OPTETRAK | LOGIC TIBIA PS MOD INSRT SZ 2 9MM | EXACTECH, INC. | 02-012-35-2009 | 02-012-35-2009 |  | AS REPORTED, APPROXIMATELY10.45 YEARS SINCE INITIAL IMPLANT, THIS FEMALE PATIENT WITH BMI 27.1 UNDERWENT A RIGHT SIDE TKA AND WAS IMPLANTED WITH AN EXACTECH UHMWPE INSERT, SERIAL # (B)(4). ON (B)(6) 2021, THE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00175 | 2022-02-04 | 2020-11-09 | I, n, j, u, r, y | Manufacturer report | OPTETRAK | LOGIC TIBIA PS MOD INSRT SZ 4 15MM | EXACTECH, INC. | 02-012-35-4015 | 02-012-35-4015 |  | AS REPORTED, APPROXIMATELY 9.91 YEARS SINCE INITIAL IMPLANT, THIS FEMALE PATIENT WITH BMI 32.1 UNDERWENT A RIGHT SIDE TKA AND WAS IMPLANTED WITH AN EXACTECH UHMWPE INSERT, SERIAL # (B)(4). ON (B)(6) 2020, THE PATIENT UNDERWENT REVISION DUE TO A PERIPROSTHETIC FRACTURE. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. | SECTION H10: (H3) BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN, MANUFACTURING, OR PATIENT RELATED ISSUES. THE CAUSE OF THE FRACTURE AND SUBSEQUENT REVISION CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY IS RELATED TO THE PATIENT¿S UNDERLYING CONDITION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00176 | 2022-02-04 | 2021-06-02 | I, n, j, u, r, y | Manufacturer report | OPTETRAK | LOGIC TIBIA IMPLANT PSC INSERT, SZ 2, 9MM | EXACTECH, INC. | 02-012-44-2009 | 02-012-44-2009 |  | AS REPORTED, APPROXIMATELY 8.84 YEARS SINCE INITIAL IMPLANT, THIS FEMALE PATIENT WITH BMI 25.7 UNDERWENT A RIGHT SIDE TKA AND WAS IMPLANTED WITH AN EXACTECH UHMWPE INSERT, SERIAL # (B)(4). ON (B)(6) 2021, THE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. | SECTION H10: (H3) THE REVISION REPORTED WAS LIKELY THE RESULT OF AN INSUFFICIENT BOND BETWEEN THE IMPLANT AND THE BONE, WHICH LED TO ASEPTIC (NON-INFECTED) LOOSENING. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICES WERE NOT AVAILABLE FOR EVALUATION AND PRE-REVISION X-RAYS WERE NOT PROVIDED. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00177 | 2022-02-04 | 2021-06-17 | I, n, j, u, r, y | Manufacturer report | OPTETRAK | LOGIC TIBIA PS MOD INSRT SZ 5 13MM | EXACTECH, INC. | 02-012-35-5013 | 02-012-35-5013 |  | AS REPORTED, APPROXIMATELY 3.86 YEARS SINCE INITIAL IMPLANT, THIS MALE PATIENT (B)(6) UNDERWENT A LEFT SIDE TKA AND WAS IMPLANTED WITH AN EXACTECH UHMWPE INSERT, SERIAL # (B)(4). ON (B)(6) 2021, THE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00178 | 2022-02-04 | 2021-05-10 | I, n, j, u, r, y | Manufacturer report | OPTETRAK | KNEE INSERT | EXACTECH, INC. | UNK | UNK |  | AS REPORTED, THIS PATIENT UNDERWENT A TKA. ON (B)(6) 2021, THE PATIENT UNDERWENT A REVISION FOR REASONS THAT ARE UNKNOWN. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00179 | 2022-02-04 | 2022-01-16 | I, n, j, u, r, y | Manufacturer report | OPTETRAK HI-FLEX TIBIAL INSERT SZ 5 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | OPTETRAK HI-FLEX TIBIAL INSERT SZ 5 9MM | 244-25-09 |  | AS REPORTED, THIS (B)(6) MALE PATIENT HAS A 6-MONTH HISTORY OF RIGHT KNEE PAIN. PATIENT HAD A REVISION OF RIGHT KNEE COMPONENTS DUE TO PAIN AND INSTABILITY. PRIOR TO PATIENT REVISION, ON A PAIN SCALE, THE KNEE WAS RATED A 5/10 - 9/10 AND IS INTERMITTENT. THE PAIN WAS PARTICULARLY BAD WITH STANDING AND WALKING. THE PAIN WAS LOCATED ON THE INSIDE OF AND AT THE FRONT OF THE KNEE. THE PAIN WAS DESCRIBED AS SHARP. PATIENT COMPLAINED OF SWELLING AND INSTABILITY. AS OF MOST RECENTLY, THE PAIN WAS GETTING WORSE. HE UNDERWENT A RIGHT PRIMARY TKA (B)(6) 2009. HE INITIALLY HAD A VERY SMOOTH RECOVERY UNTIL ABOUT 5- OR 6-YEARS POST-OP WHEN HE DEVELOPED A SEVERE BAKER'S CYST. HE HAS HAD THESE DRAINED MULTIPLE TIMES INCLUDING A SURGICAL I&D, HOWEVER IT RETURNS AFTER BEING DRAINED. AS A RESULT, HE HAS HAD ONGOING PAIN WITH WEIGHTBEARING ACTIVITIES, PAIN AND INSTABILITY. THIS HAS SIGNIFICANTLY IMPACTED HIS QUALITY OF LIFE AND HAS AFFECTED HIS ABILITY TO PERFORM DAY TO DAY ACTIVITIES. HE NOTES THAT HIS SYMPTOMS HAVE SIGNIFICANTLY WORSENED, AS HE HAS BEEN BARELY ABLE TO WALK FOR THE PAST 1 MONTH. | CONCOMITANT MEDICAL PRODUCTS OPTETRAK HI-FLEX TIBIAL INSERT SZ 5 9MM (CAT# 244-25-09); OPTETRAK ASY HI-FLEX PS CEM FEM, SZ 5, RIGHT (CAT# 244-03-05); THREE PEG PATELLA 38MM (CAT# 200-02-38). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00167 | 2022-02-04 | 2022-01-11 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | HUMERAL STEM PRIMARY, PRESS FIT 11MM | EXACTECH, INC. | 300-01-11 | 300-01-11 |  | AS REPORTED, APPROXIMATELY 5 YEAR POST OP INITIAL RTSA, THIS (B)(6) PATIENT WAS REVISED DUE TO THE HUMERAL STEM BECOMING LOOSE. THE HUMERAL STEM SHOULD HAVE BEEN CEMENTED BUT WASN¿T. SURGEON PERFORMED A REVISION TSR AND EXCHANGED THE STEM AND HUMERAL TRAY AND CEMENTED THEM IN PLACE. HE REMOVED THE 38MM GLENOSPHERE AND CHANGED TO A 42MM GLENOSPHERE AND NOTED THAT THE VERSION AND INCLINATION OF THE BASEPLATE LOOKED VERY ANTEVERTED AND SUPERIOR TILTED. FINAL REDUCTION WAS MADE WITH A +2.5MM 42 LINER. NO SURGICAL DELAY. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETUNING DUE TO THE HOSPITAL DISCARDING THEM. | PENDING EVALUATION . CONCOMITANT DEVICE(S): 320-15-01, (B)(4) - EQ REV GLENOID PLATE; 320-15-05, (B)(4) - EQ REV LOCKING SCREW; 320-01-38, (B)(4) - EQUINOXE REVERSE 38MM GLENOSPHERE; 320-38-00, (B)(4) - EQUINOXE REVERSE 38MM HUMERAL LINER +0; 320-01-38, (B)(4) - EQUINOXE REVERSE 38MM GLENOSPHERE; 320-20-00, (B)(4) - EQ REVERSE TORQUE DEFINING SCREW KIT; 320-10-00, (B)(4) - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00174 | 2022-02-04 | 2021-06-01 | I, n, j, u, r, y | Manufacturer report | OPTETRAK | LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 13MM | EXACTECH, INC. | 02-012-44-4013 | 02-012-44-4013 |  | AS REPORTED, APPROXIMATELY 4.08 YEARS SINCE INITIAL IMPLANT, THIS MALE PATIENT WITH BMI 30.4 UNDERWENT A RIGHT SIDE TKA AND WAS IMPLANTED WITH AN EXACTECH UHMWPE INSERT, SERIAL # (B)(4). ON (B)(6) 2021, THE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2021-10075 | 2022-02-04 | 2018-06-10 | I, n, j, u, r, y | Manufacturer report | EXACTECH | HEAD TRIAL UNKNOWN | EXACTECH, INC. |  | 143-36-XX |  | IT WAS REPORTED FROM US THAT A FEMORAL HEAD TRAIL WILL DISASSOCIATE FROM THE NECK AND EITHER REMAIN IN THE SHELL OR FLOAT ONTO THE JOINT SPACE AND INTO THE PELVIS. THERE HAVE BEEN MULTIPLE ATTEMPTS TO GAIN INFORMATION ABOUT THIS EVENT. THERE HAS BEEN NO RELIABLE INFORMATION OR EVENT INFORMATION TO CONFIRM THIS CLAIM OF DEVICES FLOATING INTO JOINTS OR THE PELVIS. THE PRODUCTS WERE RECEIVED FROM THE REPRESENTATIVE FOR ANALYSIS; HOWEVER, THE EVENT CLAIM OF PRODUCT DISASSEMBLY/MISASSEMBLY COULD NOT BE CONFIRMED. THIS FDA MDR IS BEING FILED IN AN ABUNDANCE OF CAUTION DUE TO THE REPORTED COMMENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE(S) RECEIVED FOR ANALYSIS. ANALYSIS COMPLETED BY SH. VISUAL EVALUATION CONDITION/FINDINGS: SCRATCHES FOUND ON THE RETURNED DEVICES APPEAR CONSISTENT WITH MARKINGS CREATED THROUGH REGULAR USE. THE DEFORMATION INDICATES THE TRIAL HEADS ARE BEING COMPLETELY ASSEMBLED TO THE FEMORAL STEM. ALL INNER DIAMETERS WERE MEASURED TO ENSURE THE DEVICES MET DRAWING SPECIFICATIONS. IT WAS FOUND ALL WERE WITHIN THE SPECIFIED DIMENSION RANGE. ALL TRIALS WERE ASSEMBLED TO A NOVATION ELEMENT FEMORAL STEM. EACH HEAD FIT APPROPRIATELY ONTO THE NECK OF THE STEM AND REQUIRED ADEQUATE FORCE TO BE REMOVED. DESIGN-RELATED ISSUES: THE DESIGN OF THE FEMORAL HEAD TRIALS (143-36-00, 143-36-03, 143-36-07, 143-36-10, 143-36-93) HAVE BEEN IN THE FIELD SINCE 2012. EXACTECH IS AWARE OF 1 RELATED COMPLAINT REPORT (B)(4) INVOLVING DISENGAGEMENT OF THE FEMORAL HEAD TRIALS FROM THE FEMORAL NECK. THE SALES DATA FOR ALL HIP STEMS WAS USED TO CALCULATE A COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINT REPORTS INVOLVING PARTS FROM THE FOLLOWING LIST OF LOTS. LOT 85106002 OF 50 PIECES, 2017. LOT 82501001 OF 50 PIECES, 2017. LOT 73391014 OF 50 PIECES, 2016. LOT 107216028 OF 50 PIECES, 2018. LOT 72546019 OF 50 PIECES, 2016. LOT 84598006 OF 25 PIECES, 2017. LOT 73391012 OF 50 PIECES, 2016. LOT 107216029 OF 50 PIECES, 2018. LOT 88007002 OF 50 PIECES, 2017. LOT 23517003 OF 30 PIECES, 2007. LOT 97008006 OF 25 PIECES, 2017. LOT 107216031 OF 25 PIECES, 2018. LOT 107216032 OF 50 PIECES, 2018. LOT 88007003 OF 50 PIECES, 2017. LOT 82226011 OF 50 PIECES, 2017. THEREFORE, THIS DOES NOT APPEAR TO BE A MANUFACTURING ISSUE. MOST LIKELY UNDERLYING CAUSE: THE REPORTED EVENT IN EXPERIENCE (B)(4) COULD NOT BE CONFIRMED TO BE DEVICE RELATED. IFU (B)(4): INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. AFTER REVIEW OF LABELING AND IFU (B)(4) REV. R, ONLY QUALIFIED SURGEONS KNOWLEDGEABLE IN ANATOMY, BIOMECHANICS, AND RECONSTRUCTIVE SURGERY SHOULD USE THESE DEVICES. THE SURGEON MUST BE FULLY KNOWLEDGEABLE ABOUT ALL ASPECTS OF THE EXACTECH HIP SYSTEM SURGICAL TECHNIQUE AND USE THESE IMPLANTS IN ACCORDANCE WITH THE RESPECTIVE INDICATIONS AND CONTRAINDICATIONS. IN ADDITION, THE SURGEON SHOULD BE FULLY KNOWLEDGEABLE ABOUT THE COMPATIBILITY OF SYSTEM COMPONENTS AND USE EACH DEVICE ACCORDINGLY. FINALLY, THE SURGEON MUST BE TRAINED ACCORDING TO THE PROPER USE OF THE SYSTEM INSTRUMENTATION, WHICH INCLUDES TRIAL DEVICES. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES. THIS FDA MDR IS BEING FILED IN AN ABUNDANCE OF CAUTION DUE TO THE REPORTED COMMENT OF DEVICES FLOATING INTO JOINT SPACE/THE PELVIS. DEVICES ARE TO BE INSPECTED FOR DAMAGE AND WEAR BEFORE USE AND THAT ONLY SURGEONS WITH THE KNOWLEDGE OF THE DEVICES AND INSTRUMENTATION ARE TO USE EXACTECH PRODUCTS. THE EVENT CLAIM OF PRODUCT DISASSEMBLY/MISASSEMBLY COULD NOT BE CONFIRMED BY THE ANALYSIS OF DEVICES RETURNED BY THE EVENT REPORTER. AN INVESTIGATION WAS CONDUCTED: THE REPORTED EVENT COULD NOT BE CONFIRMED TO BE DEVICE RELATED. | N | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2021-10083 | 2022-02-04 | 2017-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | HANDLE RETROVERSION | EXACTECH, INC. |  | 301-03-10 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC REVISION SURGERY, WHILE USING THE MALLET IN AN UPWARD MOTION, THE RETROVERSION OFFSET HANDLE BROKE AT THE POINT OF ATTACHMENT TO THE STEM INSERTER. THE CEMENT (A COMPETITOR) DID NOT FIXATE SUPERIORLY, RESULTING IN THE REMOVAL OF THE STEM AND CEMENT. THE SURGEON HAD TO REMOVE THE STEM USING THE MALLET WITH GREATER THAN USUAL FORCE. THOUGHT THE CASE DID RESULT IN LONGER TIME IN THE OR, IT WAS NOT THE RESULT OF THE BROKEN HANDLE, BUT RATHER DUE TO THE CEMENT NOT PROPERLY FIXATING. THERE WAS NO REPORT OF ADVERSE EVENT TO THE PATIENT. THE REPORT OF THE BROKEN HANDLE DID NOT CAUSE OR CONTRIBUTE TO SERIOUS INJURY TO THE PATIENT. THE AGENCY IS INACTIVE AND NO FURTHER INFORMATION WAS AVAILABLE ABOUT THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE COMPLAINT PRODUCT WAS NOT RECEIVED FOR ANALYSIS. THE REPORTED DEVICE MALFUNCTION OF BROKEN/NON-FUNCTIONAL WAS NOT CONFIRMED. THE FREQUENCY OF OCCURRENCE RANKING SCALE IS VERY LOW; THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. NO INFORMATION WAS PROVIDED REGARDING THE SERIAL OR LOT INFORMATION SO MANUFACTURING INFORMATION COULD NOT BE REVIEWED. THE BROKEN RETROVERSION HANDLE REPORTED WAS LIKELY THE RESULT OF IMPACTING THE STRIKE SURFACE OF THE BROACH HANDLE WITH THE THREADS OF THE RETROVERSION HANDLE PARTIALLY ENGAGED, ALLOWING FOR THE APPLIED FORCE TO BE TRANSMITTED TO THE THREADED TIP OF THE DEVICE, LEADING TO ULTIMATE FAILURE. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE COMPONENT WAS NOT RETURNED FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY DESIGN ISSUES, NOR DID IT LEAD TO PATIENT IMPACT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN RETROVERSION HANDLE REPORTED WAS LIKELY THE RESULT OF IMPACTING THE STRIKE SURFACE OF THE BROACH HANDLE WITH THE THREADS OF THE RETROVERSION HANDLE PARTIALLY ENGAGED, ALLOWING FOR THE APPLIED FORCE TO BE TRANSMITTED TO THE THREADED TIP OF THE DEVICE, LEADING TO ULTIMATE FAILURE. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE COMPONENT WAS NOT RETURNED FOR EVALUATION. ADDITIONAL INFORMATION- SHOULDER REVISION SURGERY DUE TO A GREATER TUBEROSITY FRACTURE CAPTURED IN MFR # 1038671-2017-00574. | N | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2021-10082 | 2022-02-04 |  | I, n, j, u, r, y | Manufacturer report | EXACTECH | TRULIANT LOCKING FEMORAL IMPACTOR | EXACTECH, INC. |  | 02-029-19-1000 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE INSTRUMENTATION. THE LOCKING FEMORAL IMPACTOR HAD TWO PIECES COME UNSCREWED FROM THE IMPACTOR DURING SURGERY. BOTH PIECES WERE RECOVERED. NO PIECES WERE LEFT IN THE PATIENT. THERE WAS NO DELAY. THE PATIENT'S HEALTH WAS STABLE LEAVING THE OR. THE DEVICE WAS NOT RECEIVED FOR ANALYSIS. THE REPRESENTATIVE WAS PRESENT AT THE TIME OF SURGERY. ATTEMPTS WERE MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REPRESENTATIVE, BUT NO RESPONSE WAS RECEIVED. NO ADDITIONAL INFORMATION HAS BEEN RECEIVED FROM THE CONTACTS ABOUT THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS ((B)(4)). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS NOT AVAILABLE FOR ANALYSIS. EMAIL FROM AH (B)(6) 2021. INVESTIGATION # AND DESCRIPTION: CAPA (B)(4): INVESTIGATION REVEALED THAT THE FAILURE OF THE CAP SCREW WAS DIRECTLY RELATED TO INADEQUATE WELDING BETWEEN THE CAP SCREW AND THE WASHER SUBCOMPONENT APPLIED DURING MANUFACTURING ASSEMBLY. THERE WAS ONLY A MINIMAL SPOT WELD APPLIED WHICH WAS VULNERABLE DURING IMPACTION USE. A REVIEW OF PRINT SHOWED THE WELD REQUIREMENT WAS CALLED OUT AS "TACK WELD", WHICH WAS NOT SPECIFIC ENOUGH AND LEFT ROOM FOR INTERPRETATION. FINDINGS: A SCREW AND WASHER WERE FOUND AT THE END OF THE CASE. THE DISASSEMBLY OF THE DEVICE WAS LIKELY DUE TO INADEQUATE WELDING OF THE SCREW AS COVERED IN CAPA(B)(4)-09. THE DEVICE AND DISASSEMBLED PARTS WERE NOT RETURNED TO EXACTECH AT THE TIME OF THIS EVALUATION. THE COMPLAINT WILL BE REOPENED IF THE DEVICES ARE RETURNED. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: THE REPORTED COMPLAINT WAS LIKELY THE RESULT OF THE CAP SCREW NOT BEING PROPERLY WELDED TO THE BODY OF THE DEVICE RESULTING IN THE SCREW LOOSENING AND FALLING OFF WITH THE ASSOCIATED WASHER. CORRECTIVE ACTION TAKEN: THE COMMITTEE HAS DETERMINED THAT THE OPERATIVE TECHNIQUE WILL BE UPDATED, AND A FIELD NOTICE WILL BE SENT TO THE AGENTS AND REPRESENTATIVES NOTIFYING THEM OF THE UPDATE. ONCE THE DESIGN CHANGE IS VALIDATED AND REPLACEMENT DEVICES ARE AVAILABLE, THE DEVICES WILL BE REMOVED FROM THE FIELD. REFER TO (B)(4) FOR MOST UP TO DATE INFORMATION. EXACTECH IFU (B)(4) STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO DESIGN ISSUES. THE PIECES WERE RECOVERED, NOTHING WAS LEFT IN THE PATIENT. AN INVESTIGATION WAS CONDUCTED: THE REPORTED COMPLAINT WAS LIKELY THE RESULT OF THE CAP SCREW NOT BEING PROPERLY WELDED TO THE BODY OF THE DEVICE RESULTING IN THE SCREW LOOSENING AND FALLING OFF WITH THE ASSOCIATED WASHER. INFORMATION WAS NOT PROVIDED. EVENT DATE NOT REPORTED. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00162 | 2022-02-04 | 2021-09-09 | I, n, j, u, r, y | Manufacturer report | OPTETRAK | LOGIC TIBIA IMPLANT PSC INSERT, SZ 3, 9MM | EXACTECH, INC. | 02-012-44-3009 | 02-012-44-3009 |  | AS REPORTED, APPROXIMATELY 7.62 YEARS SINCE INITIAL IMPLANT, THIS MALE PATIENT WITH BMI 29.3) UNDERWENT A RIGHT SIDE TKA AND WAS IMPLANTED WITH AN EXACTECH UHMWPE INSERT, SERIAL # (B)(4). ON (B)(6) 2021, THE PATIENT UNDERWENT REVISION DUE TO OSTEOLYSIS. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. | PENDING EVALUATION. NO INFORMATION PROVIDED. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00164 | 2022-02-04 | 2021-03-24 | I, n, j, u, r, y | Manufacturer report | SPECIFIC DEVICE NOT REPORTED | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED, POLYMER/METAL | EXACTECH, INC. | 208-22-09 | UNK |  | AS REPORTED, APPROXIMATELY 8.81 YEARS SINCE INITIAL IMPLANT, THIS FEMALE PATIENT WITH BMI 36.8 UNDERWENT A RIGHT SIDE TKA AND WAS IMPLANTED WITH AN EXACTECH UHMWPE INSERT, SERIAL # (B)(4). ON (B)(6) 2021, THE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. | DEVICE EVALUATED BY MFR: PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00165 | 2022-02-04 | 2019-09-17 | I, n, j, u, r, y | Manufacturer report | OPTETRAK | CC TIBIAL INSERT SZ 5, 18MM | EXACTECH, INC. | 208-25-18 | 208-25-18 |  | AS REPORTED, APPROXIMATELY 9.68 YEARS SINCE INITIAL IMPLANT, THIS MALE PATIENT WITH BMI 32.1 UNDERWENT A RIGHT SIDE TKA AND WAS IMPLANTED WITH AN EXACTECH UHMWPE INSERT, SERIAL # (B)(4). ON (B)(6) 2019, THE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00166 | 2022-02-04 | 2021-07-22 | I, n, j, u, r, y | Manufacturer report | OPTERTRAK LOGIC TIBIA IMPLANT PSC INSERT, SZ 5, 11MM | LOGIC TIBIA IMPLANT PSC INSERT, SZ 5, 11MM | EXACTECH, INC. | 208-25-18 | 02-012-44-5011 |  | AS REPORTED, APPROXIMATELY 9.29 YEARS SINCE INITIAL IMPLANT, THIS MALE PATIENT WITH BMI 28.64 UNDERWENT A RIGHT SIDE TKA AND WAS IMPLANTED WITH AN EXACTECH UHMWPE INSERT, SERIAL # (B)(4). ON (B)(6) 2021, THE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. | PENDING EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2021-10078 | 2022-02-04 | 2016-09-30 | I, n, j, u, r, y | Manufacturer report | EXACTECH | TALAR TRIAL | EXACTECH, INC. |  | 351-03-02 |  | IT WAS REPORTED THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE TALAR TRIAL. THERE WERE METAL SHAVINGS ON THE TALUS WHEN USING THE NEW TALAR TRIALS. THE SALES REPRESENTATIVE WAS PRESENT AT THE SURGERY. THE ISSUE WAS RESOLVED BY GETTING A NEW TALAR TRIAL. THERE WAS NO ADVERSE OUTCOME TO THE PATIENT. ACCORDING TO ENGINEERS ON THE ANKLE TEAM, THE SUPPLIER WAS NOT "BREAKING ALL EDGES" DURING MANUFACTURING WHICH MAY HAVE CAUSED THE REPORTED ISSUE. INFORMATION FROM CRC (B)(4): REPORT OF METAL DEBRIS AROUND THE TALAR DRILL HOLES DURING PREPARATION OF THE TALUS PEGS. THE SURGEON CLEANED THE AREA WITH PULSE LAVAGE AND CONFIRMED IT WAS REMOVED. MULTIPLE ATTEMPTS WERE MADE FOR INFORMATION DURING THE INITIAL INVESTIGATION OF THIS COMPLAINT IN 2017. INACTIVE AGENCY WITH NO HOSPITAL/OTHER CONTACT INFORMATION FOR THE EVENT. NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS ((B)(4)). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE(S) NOT AVAILABLE FOR ANALYSIS. ENGINEERING EVAL COMPLETED BY MB ON 11 MAY 2018. DESIGN-RELATED ISSUES: THE DESIGN OF THE TALAR TRIAL HAS BEEN IN THE FIELD SINCE 2015. EXACTECH HAS NOT RECEIVED ANY OTHER COMPLAINT REPORTS INVOLVING TALAR TRIALS SINCE 2015. BECAUSE THE TALAR TRIAL IS USED IN TOTAL ANKLE SURGERIES, SALES DATA FOR TIBIAL COMPONENT WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF 0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: NO INFORMATION REGARDING THE SERIAL NUMBER OR MANUFACTURING LOT WAS PROVIDED. HOWEVER, ACCORDING TO ENGINEERS ON THE ANKLE TEAM, THE SUPPLIER WAS NOT "BREAKING ALL EDGES" DURING MANUFACTURING WHICH MAY HAVE CAUSED THE REPORTED ISSUE. A REVIEW OF THE PROCESS FAILURE MODE EFFECTS ANALYSIS (PFMEA) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE PFMEA FOR THIS TALAR TRIAL, (B)(4)PFMEA REV A, WAS REVIEWED. THE RISK IS CAPTURED ON PAGE 13. THE METAL SHAVINGS REPORTED IN EXPERIENCE (B)(4)WERE LIKELY DUE TO THE SUPPLIER NOT "BREAKING ALL EDGES" IN ACCORDANCE WITH THE PRINT. CRC (B)(4): WEAR IN TOTAL ANKLE REPLACEMENTS IS A KNOWN FAILURE MECHANISM. RETRIEVAL STUDIES HAVE DOCUMENTED DEBRIS TRAPPED AT THE BONE INTERFACE OF THESE COMPONENTS AS WELL AS IN THE ARTICULAR SURFACE (TM(B)(4). HOWEVER, NO RETRIEVAL STUDIES INDICATED THE PRESENCE OF STAINLESS STEELS COMMONLY USED IN SURGICAL INSTRUMENTS. STUDIES THAT IDENTIFIED THE DEBRIS EITHER IN THE BONE OR AT THE ARTICULAR SURFACE FOUND THAT THEY WERE LIKELY FROM THE IMPLANT (HA, TITANIUM, POLYETHYLENE). GIVEN THE PUBLISHED EVIDENCE OF DEBRIS IN THE JOINT, THE LIKELIHOOD THAT DEBRIS FROM THIS INSTRUMENT WOULD ACCELERATE OR INCREASE THIS MECHANISM IS CONSIDERED LOW. OSTEOLYSIS AROUND ANKLE IMPLANTS IS COMMONLY FOUND IN THE LITERATURE. IT CAN MANIFEST AS BONE CYSTS OR BONE LOSS AT THE IMPLANT-BONE INTERFACE. IN CASES WHERE THIS LYSIS IS PROGRESSIVE, IT CAN REQUIRE SURGICAL INTERVENTION TO FILL VOIDS OR REVISE COMPONENTS. AT THIS TIME THERE IS NO EVIDENCE THAT DEBRIS GENERATED BY INSTRUMENTATION INTRAOPERATIVELY HAS CAUSED EARLY IMPLANT FAILURE. RETRIEVAL STUDIES INDICATE THE IMPLANT MATERIAL IS MOST COMMONLY FOUND IN THE ARTICULAR SURFACES AND SURROUNDING BONE. THUS, THE SEVERITY IS CONSIDERED MINOR. PULSE LAVAGE IS USED THROUGHOUT THE TOTAL ANKLE PROCEDURE TO CLEAR ANY DEBRIS GENERATED AT THE TIME OF SURGERY. THE PRESENCE OF METAL DEBRIS GENERATED WHILE DRILLING THE TALUS HOLES IS EXPECTED TO BE CLEARED BY THIS TECHNIQUE. THIS IS SUPPORTED BY RETRIEVAL STUDIES THAT HAVE DOCUMENTED MATERIALS FOUND IN THE INTERFACE ARE NOT STAINLESS STEEL, BUT RATHER HYDROXY-APATITE, TITANIUM OR POLYETHYLENE. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. THERE WERE METAL SHAVINGS ON THE TALUS WHEN USING THE NEW TALAR TRIALS. METAL DEBRIS ON THE INSTRUMENT CAN GET INTO THE JOINT ARTICULATION OR THE BONE IMPLANT INTERFACE. THE SURGEON CLEANED THE AREA WITH PULSE LAVAGE AND CONFIRMED IT WAS REMOVED. AN INVESTIGATION WAS CONDUCTED: THE METAL SHAVINGS REPORTED WERE LIKELY DUE TO THE SUPPLIER NOT "BREAKING ALL EDGES" IN ACCORDANCE WITH THE PRINT. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2021-10074 | 2022-02-04 | 2017-12-18 | I, n, j, u, r, y | Manufacturer report | EXACTECH | TRULIANT FEMORAL IMPACTOR HEAD | EXACTECH, INC. |  | 02-029-19-1000 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE LEFT KNEE WAS PREPPED USING TRULIANT KNEE INSTRUMENTS. SHORTLY AFTER IMPACTION OF THE FEMORAL TRIAL, THE SURGEON'S PA NOTED A SMALL PIECE OF METAL IN THE WOUND. THIS WAS RETRIEVED AND NOTED TO BE A SMALL WASHER. AS THE STAFF WAS TRYING TO FIGURE OUT FROM WHERE THE WASHER HAD COME, THE PA ALSO NOTED A SECOND PIECE OF METAL IN THE WOUND. THIS WAS RETRIEVED AND FOUND TO BE A VERY SMALL SCREW. A PHOTO WAS TAKEN AND SENT IMMEDIATELY TO THE COMPANY FOR CONFIRMATION OF WHAT THE TWO PIECES MIGHT BE FROM. IT WAS BELIEVED THAT THE PARTS CAME FROM THE TRULIANT FEMORAL IMPACTOR HEAD AS IT HAD JUST BEEN USED, THIS WAS CONFIRMED FROM THE PHOTO SENT TO THE COMPANY. A SECOND FEMORAL IMPACTOR HEAD WAS OBTAINED AND USED FOR THE REMAINDER OF THE CASE. X-RAYS WERE TAKEN IN THE O.R. TO ENSURE NO OTHER PIECES OF METAL HAD FALLEN OFF THE HANDLE OR ANY INSTRUMENT. X-RAYS WERE CLEAR FOR ANY OTHER METAL PARTS. THE CASE PROCEEDED WITHOUT FURTHER DELAY OR INCIDENT. THE AGENT WAS PRESENT DURING THE SURGERY. THERE WAS A DELAY TO SURGERY WITH NO PATIENT ADVERSE EVENT. INACTIVE AGENCY WITH NO HOSPITAL/OTHER CONTACT INFORMATION FOR THE EVENT. NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE AVAILABLE FOR ANALYSIS. ENGINEERING EVAL COMPLETED BY AD ON (B)(6) 2018 DESIGN-RELATED ISSUES: THE DESIGN OF THE TRULIANT FEMORAL LOCKING IMPACTOR HEAD HAS BEEN IN THE FIELD SINCE 2016. EXACTECH HAS RECEIVED 5 COMPLAINT REPORTS INVOLVING TRULIANT FEMORAL LOCKING IMPACTOR HEADS SINCE 2016. BECAUSE THE TRULIANT FEMORAL LOCKING IMPACTOR HEAD IS USED TO IMPACT FEMORAL COMPONENTS, SALES DATA FOR TRULIANT FEMORAL COMPONENTS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. ACCORDING TO INFORMATION PROVIDED FOR THE CORRECTION AND REMOVAL COMMITTEE MEETING ((B)(4)) AND THE CORRESPONDING CORRECTIVE ACTION ((B)(4)), A WELDING REQUIREMENT WAS NOT ADEQUATELY SPECIFIED. REFER TO (B)(4) AND (B)(4) FORFURTHER INFORMATION. MFG-RELATED ISSUES: THIS ISSUE WAS NOT FOUND TO BE MANUFACTURING RELATED. REFER TO (B)(4) AND (B)(4) FOR ADDITIONAL INFORMATION. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS TRULIANT FEMORAL LOCKING IMPACTOR HEAD, 750-2014-006-RMR REV B, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 35. THE DISASSEMBLY OF THE WASHER AND SCREW REPORTED IN EXPERIENCE (B)(4) WAS LIKELY THE RESULT OF INADEQUATE WELDING BETWEEN THE SCREW AND THE WASHER SUBCOMPONENT APPLIED DURING MANUFACTURING ASSEMBLY. REFER TO (B)(4) AND (B)(4) FOR ADDITIONAL INFORMATION. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. EXACTECH IFU 700-096-004 STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. PARTS OF THE DEVICE WERE REMOVED FROM THE PATIENTS. AN INVESTIGATION WAS CONDUCTED: THE DISASSEMBLY OF THE WASHER AND SCREW REPORTED WAS LIKELY THE RESULT OF INADEQUATE WELDING BETWEEN THE SCREW AND THE WASHER SUBCOMPONENT APPLIED DURING MANUFACTURING ASSEMBLY. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00163 | 2022-02-04 | 2021-09-09 | I, n, j, u, r, y | Manufacturer report | OPTETRAK | THREE PEG PATELLA 35MM | EXACTECH, INC. | 200-02-35 | 02-012-35-4015 |  | AS REPORTED, APPROXIMATELY 10.09 YEARS SINCE INITIAL IMPLANT, THIS MALE PATIENT WITH BMI 22.7 UNDERWENT A LEFT SIDE TKA AND WAS IMPLANTED WITH AN EXACTECH UHMWPE INSERT, SERIAL # (B)(4). ON (B)(6) 2021, THE PATIENT UNDERWENT REVISION DUE TO ASEPTIC LOOSENING. NO ADDITIONAL INFORMATION IS AVAILABLE OR FORTHCOMING. | DEVICE EVALUATED BY MFR: PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2021-10077 | 2022-02-04 | 2016-09-30 | I, n, j, u, r, y | Manufacturer report | EXACTECH | TALAR TRIAL | EXACTECH, INC. |  | 351-03-01 |  | IT WAS REPORTED THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE TALAR TRIAL. THERE WERE METAL SHAVINGS ON THE TALUS WHEN USING THE NEW TALAR TRIALS. THE SALES REPRESENTATIVE WAS PRESENT AT THE SURGERY. THE ISSUE WAS RESOLVED BY GETTING A NEW TALAR TRIAL. THERE WAS NO ADVERSE OUTCOME TO THE PATIENT. ACCORDING TO ENGINEERS ON THE ANKLE TEAM, THE SUPPLIER WAS NOT "BREAKING ALL EDGES" DURING MANUFACTURING WHICH MAY HAVE CAUSED THE REPORTED ISSUE. INFORMATION FROM (B)(4): REPORT OF METAL DEBRIS AROUND THE TALAR DRILL HOLES DURING PREPARATION OF THE TALUS PEGS. THE SURGEON CLEANED THE AREA WITH PULSE LAVAGE AND CONFIRMED IT WAS REMOVED. MULTIPLE ATTEMPTS WERE MADE FOR INFORMATION DURING THE INITIAL INVESTIGATION OF THIS COMPLAINT IN 2017. INACTIVE AGENCY WITH NO HOSPITAL/OTHER CONTACT INFORMATION FOR THE EVENT. NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE(S) NOT AVAILABLE FOR ANALYSIS. ENGINEERING EVAL COMPLETED BY MB ON 11 MAY 2018. DESIGN-RELATED ISSUES: THE DESIGN OF THE TALAR TRIAL HAS BEEN IN THE FIELD SINCE 2015. EXACTECH HAS NOT RECEIVED ANY OTHER COMPLAINT REPORTS INVOLVING TALAR TRIALS SINCE 2015. BECAUSE THE TALAR TRIAL IS USED IN TOTAL ANKLE SURGERIES, SALES DATA FOR TIBIAL COMPONENT WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4). THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: NO INFORMATION REGARDING THE SERIAL NUMBER OR MANUFACTURING LOT WAS PROVIDED. HOWEVER, ACCORDING TO ENGINEERS ON THE ANKLE TEAM, THE SUPPLIER WAS NOT "BREAKING ALL EDGES" DURING MANUFACTURING WHICH MAY HAVE CAUSED THE REPORTED ISSUE. A REVIEW OF THE PROCESS FAILURE MODE EFFECTS ANALYSIS (PFMEA) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE PFMEA FOR THIS TALAR TRIAL, (B)(4) REV A, WAS REVIEWED. THE RISK IS CAPTURED ON PAGE 13. THE METAL SHAVINGS REPORTED IN EXPERIENCE (B)(4) WERE LIKELY DUE TO THE SUPPLIER NOT "BREAKING ALL EDGES" IN ACCORDANCE WITH THE PRINT. (B)(4): WEAR IN TOTAL ANKLE REPLACEMENTS IS A KNOWN FAILURE MECHANISM. RETRIEVAL STUDIES HAVE DOCUMENTED DEBRIS TRAPPED AT THE BONE INTERFACE OF THESE COMPONENTS AS WELL AS IN THE ARTICULAR SURFACE ((B)(4)). HOWEVER, NO RETRIEVAL STUDIES INDICATED THE PRESENCE OF STAINLESS STEELS COMMONLY USED IN SURGICAL INSTRUMENTS. STUDIES THAT IDENTIFIED THE DEBRIS EITHER IN THE BONE OR AT THE ARTICULAR SURFACE FOUND THAT THEY WERE LIKELY FROM THE IMPLANT (HA, TITANIUM, POLYETHYLENE). GIVEN THE PUBLISHED EVIDENCE OF DEBRIS IN THE JOINT, THE LIKELIHOOD THAT DEBRIS FROM THIS INSTRUMENT WOULD ACCELERATE OR INCREASE THIS MECHANISM IS CONSIDERED LOW. OSTEOLYSIS AROUND ANKLE IMPLANTS IS COMMONLY FOUND IN THE LITERATURE. IT CAN MANIFEST AS BONE CYSTS OR BONE LOSS AT THE IMPLANT-BONE INTERFACE. IN CASES WHERE THIS LYSIS IS PROGRESSIVE, IT CAN REQUIRE SURGICAL INTERVENTION TO FILL VOIDS OR REVISE COMPONENTS. AT THIS TIME THERE IS NO EVIDENCE THAT DEBRIS GENERATED BY INSTRUMENTATION INTRAOPERATIVELY HAS CAUSED EARLY IMPLANT FAILURE. RETRIEVAL STUDIES INDICATE THE IMPLANT MATERIAL IS MOST COMMONLY FOUND IN THE ARTICULAR SURFACES AND SURROUNDING BONE. THUS, THE SEVERITY IS CONSIDERED MINOR. PULSE LAVAGE IS USED THROUGHOUT THE TOTAL ANKLE PROCEDURE TO CLEAR ANY DEBRIS GENERATED AT THE TIME OF SURGERY. THE PRESENCE OF METAL DEBRIS GENERATED WHILE DRILLING THE TALUS HOLES IS EXPECTED TO BE CLEARED BY THIS TECHNIQUE. THIS IS SUPPORTED BY RETRIEVAL STUDIES THAT HAVE DOCUMENTED MATERIALS FOUND IN THE INTERFACE ARE NOT STAINLESS STEEL, BUT RATHER HYDROXY-APATITE, TITANIUM OR POLYETHYLENE. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. THERE WERE METAL SHAVINGS ON THE TALUS WHEN USING THE NEW TALAR TRIALS. METAL DEBRIS ON THE INSTRUMENT CAN GET INTO THE JOINT ARTICULATION OR THE BONE IMPLANT INTERFACE. THE SURGEON CLEANED THE AREA WITH PULSE LAVAGE AND CONFIRMED IT WAS REMOVED. AN INVESTIGATION WAS CONDUCTED: THE METAL SHAVINGS REPORTED WERE LIKELY DUE TO THE SUPPLIER NOT "BREAKING ALL EDGES" IN ACCORDANCE WITH THE PRINT. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2021-10079 | 2022-02-04 | 2016-09-30 | I, n, j, u, r, y | Manufacturer report | EXACTECH | TALAR TRIAL | EXACTECH, INC. |  | 351-03-03 |  | IT WAS REPORTED THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE TALAR TRIAL. THERE WERE METAL SHAVINGS ON THE TALUS WHEN USING THE NEW TALAR TRIALS. THE SALES REPRESENTATIVE WAS PRESENT AT THE SURGERY. THE ISSUE WAS RESOLVED BY GETTING A NEW TALAR TRIAL. THERE WAS NO ADVERSE OUTCOME TO THE PATIENT. ACCORDING TO ENGINEERS ON THE ANKLE TEAM, THE SUPPLIER WAS NOT "BREAKING ALL EDGES" DURING MANUFACTURING WHICH MAY HAVE CAUSED THE REPORTED ISSUE. INFORMATION FROM (B)(4): REPORT OF METAL DEBRIS AROUND THE TALAR DRILL HOLES DURING PREPARATION OF THE TALUS PEGS. THE SURGEON CLEANED THE AREA WITH PULSE LAVAGE AND CONFIRMED IT WAS REMOVED. MULTIPLE ATTEMPTS WERE MADE FOR INFORMATION DURING THE INITIAL INVESTIGATION OF THIS COMPLAINT IN 2017. INACTIVE AGENCY WITH NO HOSPITAL/OTHER CONTACT INFORMATION FOR THE EVENT. NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE(S) NOT AVAILABLE FOR ANALYSIS. ENGINEERING EVAL COMPLETED BY (B)(6) ON 11 MAY 2018. DESIGN-RELATED ISSUES: THE DESIGN OF THE TALAR TRIAL HAS BEEN IN THE FIELD SINCE 2015. EXACTECH HAS NOT RECEIVED ANY OTHER COMPLAINT REPORTS INVOLVING TALAR TRIALS SINCE 2015. BECAUSE THE TALAR TRIAL IS USED IN TOTAL ANKLE SURGERIES, SALES DATA FOR TIBIAL COMPONENT WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4). THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: NO INFORMATION REGARDING THE SERIAL NUMBER OR MANUFACTURING LOT WAS PROVIDED. HOWEVER, ACCORDING TO ENGINEERS ON THE ANKLE TEAM, THE SUPPLIER WAS NOT "BREAKING ALL EDGES" DURING MANUFACTURING WHICH MAY HAVE CAUSED THE REPORTED ISSUE. A REVIEW OF THE PROCESS FAILURE MODE EFFECTS ANALYSIS (PFMEA) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE PFMEA FOR THIS TALAR TRIAL, (B)(4) PFMEA REV A, WAS REVIEWED. THE RISK IS CAPTURED ON PAGE 13. THE METAL SHAVINGS REPORTED IN EXPERIENCE (B)(6) WERE LIKELY DUE TO THE SUPPLIER NOT "BREAKING ALL EDGES" IN ACCORDANCE WITH THE PRINT. (B)(4): WEAR IN TOTAL ANKLE REPLACEMENTS IS A KNOWN FAILURE MECHANISM. RETRIEVAL STUDIES HAVE DOCUMENTED DEBRIS TRAPPED AT THE BONE INTERFACE OF THESE COMPONENTS AS WELL AS IN THE ARTICULAR SURFACE ((B)(4)). HOWEVER, NO RETRIEVAL STUDIES INDICATED THE PRESENCE OF STAINLESS STEELS COMMONLY USED IN SURGICAL INSTRUMENTS. STUDIES THAT IDENTIFIED THE DEBRIS EITHER IN THE BONE OR AT THE ARTICULAR SURFACE FOUND THAT THEY WERE LIKELY FROM THE IMPLANT (HA, TITANIUM, POLYETHYLENE). GIVEN THE PUBLISHED EVIDENCE OF DEBRIS IN THE JOINT, THE LIKELIHOOD THAT DEBRIS FROM THIS INSTRUMENT WOULD ACCELERATE OR INCREASE THIS MECHANISM IS CONSIDERED LOW. OSTEOLYSIS AROUND ANKLE IMPLANTS IS COMMONLY FOUND IN THE LITERATURE. IT CAN MANIFEST AS BONE CYSTS OR BONE LOSS AT THE IMPLANT-BONE INTERFACE. IN CASES WHERE THIS LYSIS IS PROGRESSIVE, IT CAN REQUIRE SURGICAL INTERVENTION TO FILL VOIDS OR REVISE COMPONENTS. AT THIS TIME THERE IS NO EVIDENCE THAT DEBRIS GENERATED BY INSTRUMENTATION INTRAOPERATIVELY HAS CAUSED EARLY IMPLANT FAILURE. RETRIEVAL STUDIES INDICATE THE IMPLANT MATERIAL IS MOST COMMONLY FOUND IN THE ARTICULAR SURFACES AND SURROUNDING BONE. THUS, THE SEVERITY IS CONSIDERED MINOR. PULSE LAVAGE IS USED THROUGHOUT THE TOTAL ANKLE PROCEDURE TO CLEAR ANY DEBRIS GENERATED AT THE TIME OF SURGERY. THE PRESENCE OF METAL DEBRIS GENERATED WHILE DRILLING THE TALUS HOLES IS EXPECTED TO BE CLEARED BY THIS TECHNIQUE. THIS IS SUPPORTED BY RETRIEVAL STUDIES THAT HAVE DOCUMENTED MATERIALS FOUND IN THE INTERFACE ARE NOT STAINLESS STEEL, BUT RATHER HYDROXY-APATITE, TITANIUM OR POLYETHYLENE. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. THERE WERE METAL SHAVINGS ON THE TALUS WHEN USING THE NEW TALAR TRIALS. METAL DEBRIS ON THE INSTRUMENT CAN GET INTO THE JOINT ARTICULATION OR THE BONE IMPLANT INTERFACE. THE SURGEON CLEANED THE AREA WITH PULSE LAVAGE AND CONFIRMED IT WAS REMOVED. AN INVESTIGATION WAS CONDUCTED: THE METAL SHAVINGS REPORTED WERE LIKELY DUE TO THE SUPPLIER NOT "BREAKING ALL EDGES" IN ACCORDANCE WITH THE PRINT. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2021-10073 | 2022-02-04 | 2018-12-11 | I, n, j, u, r, y | Manufacturer report | EXACTECH | LINER DRIVER HEAD, 32MM | EXACTECH, INC. |  | 141-01-32 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGICAL CASE THE SURGEON EXPERIENCED AN ISSUE WITH THE LINER DRIVER HEAD. THE 32MM LINER IMPACTOR BROKE. THIS DID NOT CAUSE ANY DELAY TO SURGERY AND THE PATIENT WAS NOT ADVERSELY AFFECTED. ALL PIECES WERE ACCOUNTED FOR AND THE WOUND SITE WAS WASHED OUT. THE PATIENT WAS STABLE AS THEY LEFT THE OPERATING ROOM. INACTIVE AGENCY WITH NO HOSPITAL/OTHER CONTACT INFORMATION FOR THE EVENT. NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RETURNED IN TWO PIECES. ANALYSIS COMPLETED ON (B)(6) 2019 BY NF. VISUAL EVALUATION CONDITIONS/FINDINGS: THE FRACTURED SURFACES APPEAR CONSISTENT WITH CRACK INITIATION THAT LIKELY INITIATED OVER TIME FROM THE CENTER OF THE DEVICE, PROPAGATION, AND ULTIMATE BRITTLE FRACTURE OF THE DRIVER HEAD. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE AND CONSISTENT WITH YEARS OF USE. DESIGN RELATED ISSUES: THE DESIGN OF THE LINER DRIVER HEAD HAS BEEN IN THE FIELD SINCE 2005. EXACTECH IS AWARE OF 2 OTHER SIMILAR COMPLAINT REPORTS INVOLVING BROKEN LINER DRIVER HEADS SINCE 2008. BECAUSE THE LINER DRIVER HEAD IS USED IN TOTAL HIP SURGERIES, SALES DATA FOR ACETABULAR LINERS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4). THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. (B)(4). THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS LINER DRIVER HEAD, 750-2003-022-RMR-INSTRUMENTS REV F, WAS REVIEWED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS " VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE BROKEN INSTRUMENT REPORTED IN EXPERIENCE (B)(4) WAS LIKELY THE RESULT OF A CRACK INITIATION, THAT INITIATED OVER TIME FROM THE CENTER OF THE DEVICE WHERE THE TIP OF THE LINER DRIVER HANDLE MEETS THE LINER DRIVER HEAD, PROPAGATION, AND ULTIMATE BRITTLE FRACTURE OF THE DRIVER HEAD. IFU (B)(4): INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. ALL PIECES WERE ACCOUNTED FOR FROM THE BROKEN LINER DRIVER HEAD AND THE WOUND SITE WAS WASHED OUT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF A CRACK INITIATION, THAT INITIATED OVER TIME FROM THE CENTER OF THE DEVICE WHERE THE TIP OF THE LINER DRIVER HANDLE MEETS THE LINER DRIVER HEAD, PROPAGATION, AND ULTIMATE BRITTLE FRACTURE OF THE DRIVER HEAD. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2021-10081 | 2022-02-04 | 2017-12-29 | I, n, j, u, r, y | Manufacturer report | EXACTECH | TRULIANT TIBIAL INSERT HANDLE | EXACTECH, INC. |  | 02-029-29-1000 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE INSTRUMENTATION. DURING ROUTINE TOTAL KNEE REPLACEMENTS, THE TIBIAL INSERT HANDLE STUCK IN BOTH THE TIBIAL INSERT TRIAL AND TIBIAL TRAY TRIAL. MOST NOTICEABLY, THE HANDLE STUCK HARD IN THE TIBIAL TRAY TRIAL. AFTER THE TRIAL WAS PINNED INTO PLACE ON THE TIBIA, THE HANDLE STUCK WHEN THE SURGEON TRIED TO REMOVE IT TO THE POINT THAT IT LIFTED THE PINNED TRAY TRIAL OFF THE TIBIA. THE TRAY TRIAL HAD TO BE REPINED TO THE TIBIA. THERE WAS NO SIGNIFICANT DELAY TO SURGERY, BUT IT WAS NOTED THAT THE TWO PRONGS OF THE TIBIAL INSERT HANDLE WERE BENT AND SCUFFED. THE PATIENT WAS STABLE LEAVING THE O.R. THE DEVICE WAS RECEIVED FOR ANALYSIS. INACTIVE AGENCY WITH NO HOSPITAL/OTHER CONTACT INFORMATION FOR THE EVENT. NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE COMPLAINT PRODUCT WAS RECEIVED FOR ANALYSIS. THE REPORTED CONDITION OF THE HANDLE GETTING STUCK IN THE TIBIAL TRIAL AND TIBIAL INSERT WAS CONFIRMED. THE FAILURE MODE HAD BEEN PREVIOUSLY INVESTIGATED UNDER (B)(4) TRULIANT TIBIAL INSERT HANDLE REDESIGN. UNDER THE CAPA IT WAS FOUND THAT THE PINS OF THE INSTRUMENT WERE DISASSOCIATING FROM THE MAIN BODY OF THE HANDLE. THIS WAS DUE TO DEFICIENCIES IN MANUFACTURING IN WHICH THERE WAS AN INSUFFICIENT LASER WELD AND USE OF A SLIP FIT VS THE SPECIFIED PRESS FIT REDUCED THE EFFICACY TO THE PIN- TO -MAIN BODY ADHESION. DESIGN IMPROVEMENTS WERE IMPLEMENTED TO BOLSTER THE PIN TO MAIN BODY INTERFACE. THE DEVICE IN THIS COMPLAINT WAS MANUFACTURED AND/OR USED PRIOR TO IMPLEMENTATION OF CORRECTIVE ACTIONS. NO FURTHER ACTION IS REQUIRED. EXACTECH IFU 700-096-004 STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO MANUFACTURING ISSUES. NO PATIENT ADVERSE EVENT WAS REPORTED, THE TRAY TRAIL HAD TO BE REPINED. THE FAILURE MODE HAD BEEN PREVIOUSLY INVESTIGATED UNDER (B)(4) TRULIANT TIBIAL INSERT HANDLE REDESIGN. UNDER THE CAPA IT WAS FOUND THAT THE PINS OF THE INSTRUMENT WERE DISASSOCIATING FROM THE MAIN BODY OF THE HANDLE. THIS WAS DUE TO DEFICIENCIES IN MANUFACTURING IN WHICH THERE WAS AN INSUFFICIENT LASER WELD AND USE OF A SLIP FIT VS THE SPECIFIED PRESS FIT REDUCED THE EFFICACY TO THE PIN- TO -MAIN BODY ADHESION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2021-10076 | 2022-02-04 | 2018-06-27 | I, n, j, u, r, y | Manufacturer report | EXACTECH | SHOULDER GPS REVERSE DRILL KIT | EXACTECH, INC. |  | 531-20-00 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE DRILL. DURING A GPS CASE IT WAS NOTICED THAT AFTER DRILLING THE HOLES FOR THE GLENOID COMPRESSION LOCKING SCREWS - THE REVERSE GLENOID COLOR-CODED DRILL (3.2MM) WAS LOSING MANY OF ITS COLORED BANDS. BITS OF COLORED PLASTIC WERE VISIBLE IN THE OPERATING SITE. THE SURGEON REMOVED ALL PIECES. NO FURTHER INCIDENT TO THE PATIENT. PATIENT'S HEALTH WAS STABLE LEAVING THE OR. THIS 3.2MM DRILL WAS PART OF A GPS LOAN KIT SENT TO THE HOSPITAL JUST FOR THIS CASE. THE AGENT WAS PRESENT AT SURGERY. THERE WAS A DELAY TO SURGERY WITH NO PATENT ADVERSE EVENT. DEVICE NOT RETURNING. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT FOLLOWING MULTIPLE ATTEMPTS. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE NOT AVAILABLE FOR EVALUATION. ENGINEERING EVAL COMPLETED BY AD ON 16 OCT 2018. THE MANUFACTURER OF THIS DRILL BIT, SYMMETRY MEDICAL, HAS PROVIDED CYTOTOXICITY TESTING FOR EVERY PAINT COLOR. REFER TO ETS-0229 AND TM-2007-331 FOR DETAILS. DESIGN-RELATED ISSUES: THE DESIGN OF THE 3.2MM DRILL (531-15-07) THAT IS PART OF THE GPS REVERSE DRILL KIT (531-20-00) HAS BEEN IN THE FIELD SINCE 2015. EXACTECH HAS RECEIVED 3 SIMILAR COMPLAINT REPORTS FOR COSMETIC (PAINT CHIPPING) ISSUES INVOLVING 4 INDIVIDUAL SHOULDER GPS DRILL BITS. THIS DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: MANUFACTURING DATA COULD NOT BE REVIEWED BECAUSE THE MANUFACTURING LOT INFORMATION WAS NOT PROVIDED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW" AND THE RISK IS CAPTURED IN THE RMR. THE CHIPPING PAINT REPORTED IN EXPERIENCE C2018-461 WAS LIKELY THE RESULT OF DEGRADATION OF THE EPOXY PAINT OVER TIME DUE TO NORMAL USE AND PROCESSING. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY:APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. BITS OF COLORED PLASTIC WERE VISIBLE IN THE OPERATING SITE; THE SURGEON REMOVED ALL PIECES. AN INVESTIGATION WAS CONDUCTED: THE CHIPPING PAINT REPORTED WAS LIKELY THE RESULT OF DEGRADATION OF THE EPOXY PAINT OVER TIME DUE TO NORMAL USE AND PROCESSING. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2021-10080 | 2022-02-04 | 2016-09-30 | I, n, j, u, r, y | Manufacturer report | EXACTECH | TALAR TRIAL | EXACTECH, INC. |  | 351-03-04 |  | IT WAS REPORTED THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE TALAR TRIAL. THERE WERE METAL SHAVINGS ON THE TALUS WHEN USING THE NEW TALAR TRIALS. THE SALES REPRESENTATIVE WAS PRESENT AT THE SURGERY. THE ISSUE WAS RESOLVED BY GETTING A NEW TALAR TRIAL. THERE WAS NO ADVERSE OUTCOME TO THE PATIENT. ACCORDING TO ENGINEERS ON THE ANKLE TEAM, THE SUPPLIER WAS NOT "BREAKING ALL EDGES" DURING MANUFACTURING WHICH MAY HAVE CAUSED THE REPORTED ISSUE. INFORMATION FROM (B)(4): REPORT OF METAL DEBRIS AROUND THE TALAR DRILL HOLES DURING PREPARATION OF THE TALUS PEGS. THE SURGEON CLEANED THE AREA WITH PULSE LAVAGE AND CONFIRMED IT WAS REMOVED. MULTIPLE ATTEMPTS WERE MADE FOR INFORMATION DURING THE INITIAL INVESTIGATION OF THIS COMPLAINT IN 2017. INACTIVE AGENCY WITH NO HOSPITAL/OTHER CONTACT INFORMATION FOR THE EVENT. NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE(S) NOT AVAILABLE FOR ANALYSIS. ENGINEERING EVAL COMPLETED BY MB ON (B)(6) 2018. DESIGN-RELATED ISSUES: THE DESIGN OF THE TALAR TRIAL HAS BEEN IN THE FIELD SINCE 2015. EXACTECH HAS NOT RECEIVED ANY OTHER COMPLAINT REPORTS INVOLVING TALAR TRIALS SINCE 2015. BECAUSE THE TALAR TRIAL IS USED IN TOTAL ANKLE SURGERIES, SALES DATA FOR TIBIAL COMPONENT WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: NO INFORMATION REGARDING THE SERIAL NUMBER OR MANUFACTURING LOT WAS PROVIDED. HOWEVER, ACCORDING TO ENGINEERS ON THE ANKLE TEAM, THE SUPPLIER WAS NOT "BREAKING ALL EDGES" DURING MANUFACTURING WHICH MAY HAVE CAUSED THE REPORTED ISSUE. A REVIEW OF THE PROCESS FAILURE MODE EFFECTS ANALYSIS (PFMEA) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE PFMEA FOR THIS TALAR TRIAL, (B)(4), WAS REVIEWED. THE METAL SHAVINGS REPORTED IN EXPERIENCE (B)(4) WERE LIKELY DUE TO THE SUPPLIER NOT "BREAKING ALL EDGES" IN ACCORDANCE WITH THE PRINT. (B)(4): WEAR IN TOTAL ANKLE REPLACEMENTS IS A KNOWN FAILURE MECHANISM. RETRIEVAL STUDIES HAVE DOCUMENTED DEBRIS TRAPPED AT THE BONE INTERFACE OF THESE COMPONENTS AS WELL AS IN THE ARTICULAR SURFACE ((B)(4)). HOWEVER, NO RETRIEVAL STUDIES INDICATED THE PRESENCE OF STAINLESS STEELS COMMONLY USED IN SURGICAL INSTRUMENTS. STUDIES THAT IDENTIFIED THE DEBRIS EITHER IN THE BONE OR AT THE ARTICULAR SURFACE FOUND THAT THEY WERE LIKELY FROM THE IMPLANT (HA, TITANIUM, POLYETHYLENE). GIVEN THE PUBLISHED EVIDENCE OF DEBRIS IN THE JOINT, THE LIKELIHOOD THAT DEBRIS FROM THIS INSTRUMENT WOULD ACCELERATE OR INCREASE THIS MECHANISM IS CONSIDERED LOW. OSTEOLYSIS AROUND ANKLE IMPLANTS IS COMMONLY FOUND IN THE LITERATURE. IT CAN MANIFEST AS BONE CYSTS OR BONE LOSS AT THE IMPLANT-BONE INTERFACE. IN CASES WHERE THIS LYSIS IS PROGRESSIVE, IT CAN REQUIRE SURGICAL INTERVENTION TO FILL VOIDS OR REVISE COMPONENTS. AT THIS TIME THERE IS NO EVIDENCE THAT DEBRIS GENERATED BY INSTRUMENTATION INTRAOPERATIVELY HAS CAUSED EARLY IMPLANT FAILURE. RETRIEVAL STUDIES INDICATE THE IMPLANT MATERIAL IS MOST COMMONLY FOUND IN THE ARTICULAR SURFACES AND SURROUNDING BONE. THUS, THE SEVERITY IS CONSIDERED MINOR. PULSE LAVAGE IS USED THROUGHOUT THE TOTAL ANKLE PROCEDURE TO CLEAR ANY DEBRIS GENERATED AT THE TIME OF SURGERY. THE PRESENCE OF METAL DEBRIS GENERATED WHILE DRILLING THE TALUS HOLES IS EXPECTED TO BE CLEARED BY THIS TECHNIQUE. THIS IS SUPPORTED BY RETRIEVAL STUDIES THAT HAVE DOCUMENTED MATERIALS FOUND IN THE INTERFACE ARE NOT STAINLESS STEEL, BUT RATHER HYDROXY-APATITE, TITANIUM OR POLYETHYLENE. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. THERE WERE METAL SHAVINGS ON THE TALUS WHEN USING THE NEW TALAR TRIALS. METAL DEBRIS ON THE INSTRUMENT CAN GET INTO THE JOINT ARTICULATION OR THE BONE IMPLANT INTERFACE. THE SURGEON CLEANED THE AREA WITH PULSE LAVAGE AND CONFIRMED IT WAS REMOVED. AN INVESTIGATION WAS CONDUCTED: THE METAL SHAVINGS REPORTED WERE LIKELY DUE TO THE SUPPLIER NOT "BREAKING ALL EDGES" IN ACCORDANCE WITH THE PRINT. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00180 | 2022-02-07 | 2022-01-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SHLDR GPS IMPACT HANDLE | ORTHOPEDIC MANUAL SURGICAL INSTRUMENT | EXACTECH, INC. | SHLDR GPS IMPACT HANDLE | 531-07-05 |  | IT WAS REPORTED THAT THE BALL BEARING AND SPRING FROM THE GPS GLENOID IMPACTOR HANDLE DISLODGED INSIDE OF PATIENT WHEN THE SURGEON IMPACTED THE GLENOID BASEPLATE. THE SURGEON BECAME AWARE WHEN THE SPRING WAS LOCATED DURING THE CASE. POST OP AP X RAY TAKEN SHOWS NO BALL BEARING LEFT IN THE PATIENT. THERE WAS LESS THAN A 5 MINUTE SURGICAL DELAY. THE PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00188 | 2022-02-08 | 2017-12-27 | I, n, j, u, r, y | Manufacturer report | LOGIC FEMORAL PS CEM RIGHT SZ 3.5 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC FEMORAL PS CEM RIGHT SZ 3.5 | 02-010-01-0335 |  | AS REPORTED, APPROXIMATELY FIVE YEARS POST INITIAL RIGHT SIDE TKA, THIS PATIENT HAD A REVISION DUE TO OSTEOLYSIS. NO ADDITIONAL INFORMATION RECEIVED. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00187 | 2022-02-08 | 2019-10-09 | I, n, j, u, r, y | Manufacturer report | LOGIC FEMORAL PS CEM RIGHT SZ 2.5 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC FEMORAL PS CEM RIGHT SZ 2.5 | 02-010-01-0325 |  | AS REPORTED, APPROXIMATELY FOUR YEARS POST LEFT SIDE TKA, THIS (B)(6) Y/O FEMALE PATIENT HAD A REVISION DUE TO OSTEOLYSIS. | (H3) THE REVISION REPORTED WAS LIKELY THE RESULT OF OSTEOLYSIS POSSIBLY RELATED TO WEAR OF THE TIBIAL INSERT, WHICH ALLOWED FOR THE FEMORAL COMPONENT AND TIBIAL TRAY TO LOOSEN FROM THE BONE. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE WEAR OF THE TIBIAL INSERT WAS NOT REPORTED, NO IMAGES OF THE REVISED COMPONENTS WERE PROVIDED, AND THE DEVICES WERE NOT RETURNED TO EXACTECH FOR EVALUATION. THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE REPORTED EVENT OF "PROSTHESIS WEAR¿ IS ASSOCIATED WITH MATERIAL DAMAGE TO A SURFACE, USUALLY INVOLVING PROGRESSIVE LOSS OR DISPLACEMENT OF MATERIAL, DUE TO RELATIVE MOTION BETWEEN THAT SURFACE AND A CONTACTING SUBSTANCE OR SUBSTANCES. FURTHERMORE, THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE EVENT OF "LOOSENING¿ IS ASSOCIATED WITH DESTRUCTION OR DISAPPEARANCE OF BONE TISSUE LIKELY RELATED TO WEAKENED INTEGRATION OF AN IMPLANT AT THE BONE-IMPLANT INTERFACE. SECTION D1: BRAND NAME SECTION D4: MODEL NUMBER, CATALOG NUMBER, SERIAL NUMBER, EXPIRATION DATE, UNIQUE IDENTIFIER (UDI)#. SECTION D10: CONCOMITANT MEDICAL PRODUCTS. SECTION H4: DEVICE MANUFACTURE DATE. SECTION D10: CONCOMITANT MEDICAL PRODUCTS: LOGIC TIBIA PS MOD INSRT SZ 2.5 9MM (CAT#: 02-012-35-2509 / SERIAL#: (B)(6). LGC TIBIAL FIT TRAY CEM SZ 2.5F / 2.5T (CAT#: 02-012-45-2525 / SERIAL#: (B)(6). THREE PEG PATELLA 32MM (CAT#: 200-02-32 / SERIAL#: (B)(6). | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00186 | 2022-02-08 | 2022-01-19 | I, n, j, u, r, y | Manufacturer report | LOGIC FEMORAL PS CEM LEFT SZ 3 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC FEMORAL PS CEM LEFT SZ 3 | 02-010-01-0230 |  | AS REPORTED, APPROXIMATELY SIX YEARS POST INITIAL LEFT SIDE TKA, THIS MALE PATIENT VISITED HIS SURGEON AND IS PENDING A REVISION DUE TO OSTEOLYSIS. NO ADDITIONAL INFORMATION AVAILABLE. | CONCOMITANT MEDICAL PRODUCTS: LOGIC TIBIA PS MOD INSRT SZ 3 9MM (CAT# 02-012-35-3009); LGC TIBIAL FIT TRAY CEM SZ 3F / 2T (CAT# 02-012-45-3020); THREE PEG PATELLA 32MM (CAT# 200-02-32); FLUTED STEM EXTENSION 25L X 14 MM (CAT# 204-34-02); TIBIAL STEM EXT. SCREW (CAT# 204-70-00). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00185 | 2022-02-08 | 2021-11-10 | I, n, j, u, r, y | Manufacturer report | LOGIC FEMORAL PS CEM LEFT SZ 2 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC FEMORAL PS CEM LEFT SZ 2 | 02-010-01-0220 |  | AS REPORTED, APPROXIMATELY EIGHT YEARS POST BILATERAL TKA, THIS (B)(6) FEMALE PATIENT HAD A GROSSLY LOOSE FEMORAL COMPONENT WITH SEVERE OSTEOLYSIS OF DISTAL FEMUR AND OSTEOLYSIS OF PROXIMAL TIBIAL WITH PARTIALLY LOOSE TIBIAL COMPONENT. PATIENT HAD LEFT TKA REVISION REQUIRING USE OF FEMORAL AUGMENTS AND TIBIAL CONE. | CONCOMITANT MEDICAL PRODUCTS: LOGIC TIBIA PS MOD INSERT SZ 2 13MM (CAT# 02-012-35-2013 / SERIAL# (B)(4)). LGC TIBIAL FIT TRAY CEM SZ 2F / 2T (CAT# 02-012-45-2020 / SERIAL# (B)(4)). THREE PEG PATELLA 32MM (CAT# 200-02-32 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00182 | 2022-02-08 | 2022-01-04 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | CAGE GLENOID SMALL, ALPHA | EXACTECH, INC. | 314-13-02 | 314-13-02 |  | AS REPORTED, APPROXIMATELY 10 MONTHS POST OF INITIAL LEFT TSA, THIS (B)(6) Y/O FEMALE PATIENT WAS REVISED DUE TO THE GLENOID COMPONENT BECOMING LOOSE. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETUNING DUE TO HOSPITAL POLICY. | DEVICE EVALUATED BY MANUFACTURER? PENDING EVALUATION. CONCOMITANT DEVICE(S): 300-10-15, (B)(4) - EQUINOXE REPLICATOR PLATE 1.5MM O/S; 300-20-02, (B)(4) - EQUINOX SQUARE TORQUE DEFINE SCREW DRIVE KIT; 310-01-41, (B)(4) - EQUINOXE, HUMERAL HEAD SHORT, 41MM (ALPHA). | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00183 | 2022-02-08 | 2021-06-21 | I, n, j, u, r, y | Manufacturer report | SPECIFIC DEVICE NOT REPORTED | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED, POLYMER/METAL | EXACTECH, INC. | LOGIC FEMORAL PS CEM LEFT SZ 2.5 | UNK |  | AS REPORTED, APPROXIMATELY 6 YEARS POST INITIAL LEFT SIDE TKA, THIS (B)(6) FEMALE PATIENT HAD A REVISION DUE TO SEVERE OSTEOLYSIS AROUND FEMORAL COMPONENT OF BOTH CONDYLES AND ANTERIOR FLANGE AND ALSO SEVERE OSTEOLYSIS OF PROXIMAL TIBIA. | CONCOMITANT MEDICAL PRODUCTS: LOGIC TIBIA PS MOD INSRT SZ 2.5 13MM (CAT# 02-012-35-2513 / SERIAL# (B)(4), LGC TIBIAL FIT TRAY CEM SZ 2.5F / 2.5T (CAT# 02-012-45-2525 / SERIAL# (B)(4) THREE PEG PATELLA 29MM (CAT# 200-02-29 / SERIAL# (B)(4), FLUTED STEM EXTENSION 25L X 14 MM (CAT# 204-34-02 / SERIAL# (B)(4) TIBIAL STEM EXT. SCREW (CAT# 204-70-00 / SERIAL# (B)(4). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00181 | 2022-02-08 | 2021-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EQUINOXE | MIDDLE SEGMENT TRIAL 25 | EXACTECH, INC. |  | 309-10-25 |  | AS REPORTED, THE TRIAL BROKE DURING USE IN SURGERY. ONE OF THE FLANGES BROKE OFF DURING THE TRIAL PROCESS. THERE WAS NO SURGICAL DELAY OR PROLONGATION AND NO PARTS OR PIECES THAT FELL INTO THE PATIENT¿S WOUND. THE PRODUCT IS AVAILABLE FOR RETURN. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. AFTER NEW INFORMATION WAS RECEIVED FROM THE ENGINEERS DURING THEIR INVESTIGATION, THIS DEVICE HAS BEEN DETERMINED TO BE REPORTABLE DUE TO THE NATURE OF THE BREAKAGE. | DEVICE EVALUATED BY MFR: PENDING EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00189 | 2022-02-08 | 2021-12-15 | I, n, j, u, r, y | Manufacturer report | LOGIC CR TIB INSERT SLOPE++, SZ 3, 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC CR TIB INSERT SLOPE++, SZ 3, 13MM | 02-012-49-3013 |  | AS REPORTED, APPROXIMATELY 5 YEARS POST INITIAL LEFT SIDE TKA, THIS (B)(6) FEMALE PATIENT EXPERIENCED PROGRESSIVE INSTABILITY OVER TIME THAT BROUGHT HER BACK TO THE CLINIC. POLY WEAR OVER TIME OF BOTH COMPARTMENTS. POSTERIOR WEAR BILATERALLY BUT WORSE MEDIALLY. IT WAS NOT PUT IN TOO TIGHT POSTERIORLY AND THE KNEE WAS WELL BALANCED. WHEN SURGEON TOOK POLY OUT AND REPLACED IT THERE WAS NO POSTERIOR TIGHTNESS. THERE WAS NO PARTS OR PIECES FELL INTO THE PATIENT WOUND SITE. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICE WILL BE RETURN. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00184 | 2022-02-08 | 2022-01-19 | I, n, j, u, r, y | Manufacturer report | LOGIC FEMORAL PS CEM LEFT SZ 2.5 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC FEMORAL PS CEM LEFT SZ 2.5 | 02-010-01-0225 |  | AS REPORTED, APPROXIMATELY SIX YEARS POST INITIAL TKA, THIS PATIENT VISITED HIS SURGEON AND IS PENDING A REVISION DUE TO OSTEOLYSIS. NO ADDITIONAL INFORMATION AVAILABLE. | CONCOMITANT MEDICAL PRODUCTS: LOGIC TIBIA PS MOD INSRT SZ 2.5 11MM (CAT# 02-012-35-2511 / SERIAL# (B)(4)) LGC TIBIAL FIT TRAY CEM SZ 2.5F / 2.5T (CAT# 02-012-45-2525 / SERIAL#: (B)(4)) THREE PEG PATELLA 32MM (CAT# 200-02-32 / SERIAL# (B)(4)) FLUTED STEM EXTENSION 25L X 14 MM (CAT# 204-34-02 / SERIAL# (B)(4)) TIBIAL STEM EXT. SCREW (CAT# 204-70-00 / SERIAL# (B)(4)) ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00190 | 2022-02-08 | 2022-01-19 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 3 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 3 13MM | 02-012-35-3013 |  | AS REPORTED, APPROXIMATELY SEVEN YEARS POST RIGHT TKA, THIS (B)(6) Y/O FEMALE PATIENT HAD A REVISION DUE TO OSTEOLYSIS. PATIENT WAS REVISED TO KNEE TO STRYKER. THERE WAS NO BREAKAGE OF DEVICE. THERE WAS NO SURGICAL DELAY/PROLONGATION. THE DEVICE IS AVAILABLE FOR EVALUATION. | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION RECEIVED THE FOLLOWING SECTIONS G3, G6, H2, H3 AND H6 HAVE BEEN UPDATED ACCORDINGLY. (H3) THE REVISION REPORTED WAS LIKELY THE RESULT OF SEVERE WEAR OF THE POLYETHYLENE SECONDARY TO COMPONENT ALIGNMENT AND KNEE INSTABILITY AND THE SUBSEQUENT PROSTHESIS WEAR MAY HAVE CONTRIBUTED TO PARTICLE-INDUCED OSTEOLYSIS. THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE REPORTED EVENT OF "OSTEOLYSIS¿ IS ASSOCIATED WITH DESTRUCTION OR DISAPPEARANCE OF BONE TISSUE LIKELY RELATED TO WEAKENED INTEGRATION OF AN IMPLANT AT THE BONE-IMPLANT INTERFACE; FURTHERMORE, THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE REPORTED EVENT OF "PROSTHESIS WEAR¿ IS ASSOCIATED WITH MATERIAL DAMAGE TO A SURFACE, USUALLY INVOLVING PROGRESSIVE LOSS OR DISPLACEMENT OF MATERIAL, DUE TO RELATIVE MOTION BETWEEN THAT SURFACE AND A CONTACTING SUBSTANCE OR SUBSTANCES. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00192 | 2022-02-09 | 2022-01-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | A+ BROACH HANDLE RIGHT | ORTHOPEDIC MANUAL SURGICAL INSTRUMENT | EXACTECH, INC. | A+ BROACH HANDLE RIGHT | 167-02-00 |  | AS REPORTED, DURING IMPACTION, THE SPRING ON THE BROACH HANDLE WAS BROKEN BY SURGEON. THIS SITUATION CAME TO NO SURPRISE BECAUSE THIS HAS HAPPENED TO SURGEON ON NUMEROUS OCCASIONS USING A PLUS SIZED MALLET WHILE BROACH FEMUR. THIS INCIDENT CAUSED NO DELAY, THE SURGICAL SITE WAS CONFIRMED CLEAN BY SURGEON AND THE PROCEDURE CARRIED ON AS PLANNED. PATIENT IS STABLE CONDITION. THE DEVICE WILL BE RETURN. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | N | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00191 | 2022-02-09 | 2021-10-14 | I, n, j, u, r, y | Manufacturer report | NV CROWN CUP CLSTR HOLE 48MM GROUP 1 | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV CROWN CUP CLSTR HOLE 48MM GROUP 1 | 180-01-48 |  | AS REPORTED BY THE LEGAL BRIEF, APPROXIMATELY FIVE YEARS POST INITIAL THA, THIS FEMALE PATIENT EXPERIENCE SEVERE PAIN IN HER RIGHT HIP, AND WAS SUBSEQUENTLY DIAGNOSED WITH LOOSENING OF THE HIP SYSTEM AND POLYETHYLENE WEAR OF THE ACETABULAR COMPONENT. PATIENT HAD SUSTAINED SEVERE AND PERMANENT INJURIES; SHE WAS REQUIRED TO UNDERGO MULTIPLE PERMANENT AND PAINFUL REVISION SURGERIES, INCLUDING THE REMOVAL AND RE-IMPLANTATION OF A REPLACEMENT PROSTHETIC HIP. SHE HAS REQUIRED MANY MONTHS OF HOSPITALIZATION, PHYSICAL THERAPY, AND OTHER CARE. SHE HAS AND WILL CONTINUE TO BE BEDRIDDEN. SHE HAS AND WILL CONTINUE TO EXPERIENCE SEVER AND PERMANENT PAIN; AND SHE HAS AND WILL CONTINUE TO SUFFER SEVERE AND PERMANENT COMPLICATIONS AND OTHER INJURIES TO THE RIGHT HIP. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00193 | 2022-02-10 | 2022-01-19 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 42MM HUMERAL LINER +2.5 | EXACTECH, INC. | 320-42-03 | 320-42-03 |  | AS REPORTED, THIS (B)(6) MALE PATIENT'S RIGHT SHOULDER WAS REVISED DUE TO THE 42 POLY DISASSOCIATED FROM THE TRAY. THE PATIENT FELT SOMETHING POP WHEN LIFTING THEIR ARM. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00194 | 2022-02-10 | 2022-04-19 | I, n, j, u, r, y | Manufacturer report | NV GXL LINR, NTRL, 32MM ID, GROUP 2 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LINR, NTRL, 32MM ID, GROUP 2 CUPS | 130-32-52 |  | AS REPORTED BY THE LEGAL BRIEF, THIS MALE PATIENT IMPLANTS FAILED FOR AN UNKNOWN REASON. NO ADDITIONAL INFORMATION AVAILABLE AND FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | CONCOMITANT MEDICAL PRODUCTS: BONE SCREW 6.5MM DIA X 20MM LONG (CAT# 120-65-20 / (B)(4)), COCR FEM HEAD 32MM +3.5 OFFSET 12/14 (CAT# 142-32-03 / (B)(4)), NOVATION ELEMENT RO S/O COL SZ 12 (CAT# 164-13-12 / (B)(4)), INTEGRIP CC, CLUSTER 54MM, G2 (CAT# 186-01-54 / (B)(4)). AS REPORTED BY THE LEGAL BRIEF, THIS MALE PATIENT IMPLANTS FAILED FOR AN UNKNOWN REASON. NO ADDITIONAL INFORMATION AVAILABLE AND FORTHCOMING. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. UPON REVIEW, THERE IS NO ALLEGATION AGAINST THE DEVICE AS THE REASON FOR REVISION WAS NOT REPORTED. THE MOST LIKELY CAUSE OF THE REPORTED EVENT IS UNKNOWN AT THIS TIME. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00195 | 2022-02-14 | 2021-08-01 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 42MM HUMERAL LINER +0 | EXACTECH, INC. | 320-42-00 | 320-42-00 |  | AS REPORTED, APPROXIMATELY 3 YEARS POST OP LEFT TSA, THIS (B)(6) MALE PATIENT WAS REVISED DUE TO SWELLING. SWELLING STARTED WITHOUT INQUIRY AROUND (B)(6) 2021. PREVIOUSLY PLACED ROTATOR CUFF ANCHOR MIGRATED. THE CASE REPORT FORM INDICATES THIS EVENT IS DEFINITELY NOT RELATED TO DEVICES OR PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | PENDING EVALUATION. CONCOMITANT DEVICE(S): 300-30-12, EQUINOXE PRESERVE STEM 12MM, 320-10-00, EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0, 320-02-42, RS EXPANDED GLENOSPHERE 42MM, +4MM OFFSET, 320-15-01, EQ REV GLENOID PLATE. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00196 | 2022-02-14 | 2022-01-25 | I, n, j, u, r, y | Manufacturer report | TRULIANT PS CEM FEM PS CEM RIGHT SZ 4.5 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT PS CEM FEM PS CEM RIGHT SZ 4.5 | 02-020-11-0345 |  | AS REPORTED, APPROXIMATELY TWO YEARS POST INITIAL TKA, THIS (B)(6) MALE PATIENT HAD A REVISION DUE TO SIGNIFICANT PAIN IN WEIGHT-BEARING AND MOVEMENT. THE SURGEON DIAGNOSIS FAILED RIGHT KNEE DUE TO INSTABILITY. THERE WAS NO RELATED TO BREAKAGE OF THE DEVICE. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICES ARE NOT AVAILABLE FOR EVALUATION DUE TO SURGEON WANTED TO KEEP IMPLANTS. | CONCOMITANT MEDICAL PRODUCTS: TRULIANT TIB IMP PS INSERT SZ 4.5 11MM (CAT# 02-022-35-4511 / SERIAL# (B)(4)). TRULIANT TIB FIT TRAY CEM SZ 4.5F/3.5T (CAT# 02-022-45-4535 / SERIAL# (B)(4)). LOGIC STEM EXT 14MM X 25MM (CAT# 02-012-60-1425 / SERIAL# (B)(4)). TIBIAL STEM EXT. SCREW (CAT# 204-70-00 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00197 | 2022-02-14 | 2022-01-26 | I, n, j, u, r, y | Manufacturer report | LOGIC FEMORAL PS CEM LEFT SZ 5 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC FEMORAL PS CEM LEFT SZ 5 | 02-010-01-0250 |  | AS REPORTED, APPROXIMATELY SIX YEARS POST INITIAL TKA, THIS (B)(6) MALE PATIENT COMPLAINED OF HAVING PAIN. PATIENT HAD A REVISION DUE TO FEMUR WAS LOOSE. THERE WAS NO BREAKAGE OF THE DEVICE, THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICE WILL BE RETURN. | CONCOMITANT MEDICAL PRODUCTS: LOGIC TIBIA PS MOD INSRT SZ 5 11MM (CAT# 02-012-35-5011 / SERIAL# (B)(46)). LGC TIBIAL FIT TRAY CEM SZ 5F / 4T (CAT# 02-012-45-5040 / SERIAL# (B)(4)). THREE PEG (B)(6),(CAT# 200-02-38 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00198 | 2022-02-15 | 2022-01-21 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 42MM HUMERAL LINER +2.5 | EXACTECH, INC. | 320-42-03 | 320-42-03 |  | AS REPORTED, APPROXIMATELY 2 YEARS POST OP INITIAL LEFT TSA, THIS (B)(6) Y/O MALE PATIENT WAS REVISED DUE TO INFECTION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETUNING DUE TO HOSPITAL POLICY. | DEVICE EVALUATED BY MANUFACTURER? PENDING EVALUATION. CONCOMITANT MEDICAL DEVICE(S): 320-15-01, (B)(4) - EQ REV GLENOID PLATE 320-20-26, (B)(4) - EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 26MM; 320-20-26, (B)(4) - EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 26MM; 320-20-30, (B)(4) - EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 30MM; 320-20-18, (B)(4) - EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 18MM; 320-15-05, (B)(4) - EQ REV LOCKING SCREW; 300-01-11, (B)(4) - EQUINOXE, HUMERAL STEM PRIMARY, PRESS FIT 11MM; 320-20-00, (B)(4) - EQ REVERSE TORQUE DEFINING SCREW KIT; 320-10-00, (B)(4) - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0; 320-01-42, (B)(4) - EQUINOXE REVERSE 42MM GLENOSPHERE. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00202 | 2022-02-16 | 2022-01-04 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | GLENOSPHERE 38MM | EXACTECH, INC. | 320-06-38 | 320-06-38 |  | A REPORTED, DURING THE PRIMARY IMPLANT PROCEDURE OF THE R SHOULDER, THIS (B)(6) FEMALE PATIENT EXPERIENCED A CORACOID TIP FRACTURE. SURGEON PERFORMED A SUTURE FIXATION. GPS WAS USED DURING THIS CASE, HOWEVER IT IS UNKNOWN WHERE THE GPS PATIENT TRACKER WAS ATTACHED. THE CASE REPORT FORM INDICATES THIS EVENT IS DEFINITELY RELATED TO DEVICES AND/OR PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | SECTION H10: (H3) BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN OR MANUFACTURING ISSUES. THE MOST LIKELY CAUSE OF THE REPORTED EVENT WAS A FRACTURE WHICH WAS LIKELY CAUSED DURING IMPLANTATION OF THE SHOULDER. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00199 | 2022-02-16 | 2022-01-18 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 38MM HUMERAL LINER +2.5 | EXACTECH, INC. | 320-38-03 | 320-38-03 |  | AS REPORTED, APPROXIMATELY 1.5 YEARS POST OP INITIAL RRTSA, THIS (B)(6) Y/O MALE PATIENT WAS REVISED DUE TO AN INFECTION. ALL IMPLANTS WERE REMOVED AND A SPACER WAS IMPLANTED. THERE WAS NO SURGICAL DELAY. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETUNING DUE TO HOSPITAL POLICY. | SECTION H10: (H3) BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN OR MANUFACTURING ISSUES. THE CAUSE OF SURGICAL REVISION FOR INFECTION CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00200 | 2022-02-16 | 2022-01-27 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 42MM HUMERAL LINER +0 | EXACTECH, INC. | 320-42-00 | 320-42-00 |  | AS REPORTED, APPROXIMATELY 2 MONTHS POST OP INITIAL RTSA, THIS (B)(6) MALE PATIENT WAS REVISED DUE TO AN INFECTION. THERE WAS A 15-30 MIN DELAY, THERE WAS NO ADVERSE EVENTS AS A RESULT. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETUNING DUE TO HOSPITAL POLICY, COMPONENTS WERE IN VERY GOOD SHAPE WITH NIL WEAR. | SECTION H10: (H3) BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN OR MANUFACTURING ISSUES. THE CAUSE OF SURGICAL REVISION FOR INFECTION CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION. THESE DEVICES ARE USED FOR TREATMENT NOT DIAGNOSIS. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00201 | 2022-02-16 | 2022-01-27 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | HUMERAL TRAY +15MM | EXACTECH, INC. | 320-10-15 | 320-10-15 |  | AS REPORTED, THIS (B)(6) MALE PATIENT'S LEFT SHOULDER WAS REVISED DUE TO INSTABILITY. SURGEON REMOVED IN-SITU HUMERAL ADAPTER TRAY, LINER AND GLENOID SPHERE, AND REPLACED WITH LARGER DIAMETER SPHERE AND NEW HUMERAL TRAY AND POLY LINER. NO NEGATIVE OUTCOMES AFTER SURGERY. | SECTION H10: (H3) BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN OR MANUFACTURING ISSUES. THE CAUSE OF THE SUBSEQUENT REVISION CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION. SECTION H11: THE FOLLOWING SECTIONS HAVE CORRECTED INFORMATION: (D1) BRAND NAME: HUMERAL LINER (D2B) COMMON DEVICE NAME: PROSTHESIS, SHOULDER, NON-CONSTRAINED, METAL/POLYMER CEMENTED. (D4) PLEASE DISREGARD ALL INFO IN D4, HUMERAL LINER INFORMATION IS UNKNOWN. (D6B) PLEASE DISREGARD, UNKNOWN. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00203 | 2022-02-17 | 2021-01-26 | I, n, j, u, r, y | Manufacturer report | VANTAGE | TIBIAL INSERT MB SZ 3 LT 6MM | EXACTECH, INC. |  | 350-41-03 |  | AS REPORTED, APPROXIMATELY 10MONTHS POST OP INITIAL LEFT TSA, THIS (B)(6) FEMALE PATIENT WAS REVISED DUE TO THE APPEARANCE OF POSSIBLE BONY DEFECTS/CYSTS IN THE X-RAYS TAKEN IN A POST OP CONTROL FOR THE INITIAL SURGERY IN (B)(6) 2021. DURING THE SURGERY, THE SURGEON NOTICED MARKS/WEAR ON THE SURFACE OF THE TIBIAL INSERT. DURING EXPLANATION OF THE TIBIAL INSERT THE SURGEON RECOGNIZED MARKS OR SIGNS OF WEAR ON THE LOWER, CURVED SURFACE OF THE INSERT. THIS SURFACE IS IN CONTACT WITH THE CURVED UPPER SURFACE OF THE USED TALAR IMPLANT. THE SURGEON DECIDED TO CHANGE THIS INSERT TO A NEW ONE OF SAME SIZE. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETUNING. | SECTION H10: (H3) THE REVISION REPORTED WAS LIKELY THE RESULT OF THE APPEARANCE OF POSSIBLE BONY DEFECTS/CYSTS ON POST-OPERATIVE X-RAYS. IT IS UNLIKELY THAT THE MINIMAL TIBIAL INSERT WEAR OBSERVED ON THE IMAGES PROVIDED CONTRIBUTED TO THE NEED FOR THE REVISION SURGERY. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE TIBIAL INSERT WAS NOT RETURNED TO EXACTECH FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00204 | 2022-02-18 | 2021-05-10 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | CAGE GLENOID MEDIUM, ALPHA | EXACTECH, INC. | 314-13-03 | 314-13-03 |  | AS REPORTED, APPROXIMALLY 1.5 YEARS POST OF THE INITIAL TSA, THIS (B)(6) MALE PATIENT IS EXPERIENCING ASEPTIC GLENOID LOOSENING AND DISASSOCIATION OF THE POLYETHYLENE. PATIENT REFUSES TO GET CARDIAC CLEARANCE TO PERFORM REVISION SURGERY. THE EVENT OUTCOME IS CONTINUING. THE CASE REPORT FORM INDICATES THIS EVENT IS DEFINITELY RELATED TO DEVICES AND UNLIKELY RELATED TO PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | DEVICE EVALUATED BY MFR: PENDING EVALUATION CONCOMITANT DEVICE(S): 300-62-02 - STEMLESS HUMERAL COMP INTEGRIP, CAGE, SIZE 2 AND 310-61-44 - STEMLESS HUMERAL HEAD 44MM X 17MM X ALPHA. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00208 | 2022-02-19 | 2022-01-31 | I, n, j, u, r, y | Manufacturer report | LOGIC CR TIB INSERT STD, SZ 2.5, 9 MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED | EXACTECH, INC. | LOGIC CR TIB INSERT STD, SZ 2.5, 9 MM | 02-012-47-2509 |  | AS REPORTED, APPROXIMATELY TWO YEARS POST INITIAL RIGHT SIDE TKA, THIS (B)(6) Y/O FEMALE PATIENT HAD A REVISION DUE TO INCREASE WEAR OF THE POLYETHYLENE INSERT. THERE WAS NO BREAKAGE OF THE DEVICE. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICE IS AVAILABLE FOR EVALUATION. | CONCOMITANT MEDICAL PRODUCTS: LOGIC CR FEMORAL POR, LEFT, SZ 2.5 (CAT# 02-010-04-0225 / SERIAL#: (B)(4)). LGC TIBIAL FIT TRAY CEM SZ 2.5F / 5T (CAT# 02-012-45-2525 / SERIAL# : (B)(4)). THREE PEG PATELLA 35MM (CAT# 200-02-35 / SERIAL#: (B)(4)). LOGIC CR FEMORAL POR, RIGHT, SZ 2.5 (CAT# 02-010-04-0325 / SERIAL#: (B)(4)). LGC TIBIAL FIT TRAY CEM SZ 2.5F / 5T (CAT# 02-012-45-2525 / SERIAL#: (B)(4)). LOGIC TIBIAL INSERT STD, SZ 2.5, 9 MM (CAT# 02-012-47-2509 / SERIAL#: (B)(4)). THREE PEG PATELLA 35MM (CAT# 200-02-35 / SERIAL#: (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00207 | 2022-02-19 | 2022-01-22 | I, n, j, u, r, y | Manufacturer report | LOGIC TIB INSERT IMPL CRC, SZ 4, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIB INSERT IMPL CRC, SZ 4, 9MM | 02-012-51-4009 |  | AS REPORTED, APPROXIMATELY FOURTEEN MONTHS POST INITIAL RIGHT SIDE TKA, THIS (B)(6) MALE PATIENT HAD A REVISION DUE TO INFECTION. PATIENT HAD PAIN AND SWELLING. THERE WAS NO BREAKAGE OF THE DEVICE. THERE WAS LESS THAN 5 MINUTES SURGICAL DELAY/PROLONGATION. NO ADVERSE EVENT. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION. REP WAS UNABLE TO OBTAIN X-RAYS OR PHOTOS. THE DEVICE IS AVAILABLE FOR EVALUATION. | CONCOMITANT MEDICAL PRODUCTS: LGC TIBIAL FIT TRAY CEM SZ 4F / 4T (CAT# 02-012-45-4040 / (B)(4)), LOGIC CR FEMORAL CEM, LEFT, SZ 4 (CAT# 02-010-03-0240 / (B)(4)), THREE PEG PATELLA 35MM (CAT# 200-02-35 / (B)(4)), THREE PEG PATELLA 35MM (CAT# 200-02-35 / (B)(4)), LOGIC TIB INSERT IMPL CRC, SZ 4, 13MM (CAT# 02-012-51-4013 / (B)(4)), LGC TIBIAL FIT TRAY CEM SZ 4F / 4T (CAT# 02-012-45-4040 / (B)(4)), LOGIC CR FEMORAL CEM, RIGHT, SZ 4 (CAT# 02-010-03-0340 / (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00205 | 2022-02-19 | 2022-01-27 | I, n, j, u, r, y | Manufacturer report | NV GXL LNR, 10 DEG FACE, 32MM ID, GRP 1 CUP | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LNR, 10 DEG FACE, 32MM ID, GRP 1 CUP | 138-32-51 |  | AS REPORTED, APPROXIMATELY TEN YEARS POST INITIAL THA, THIS FEMALE PATIENT'S POLY BROKE. PATIENT WAS REVISED TO AN XLE LINER. THERE WERE NO BROKEN PIECES FELL IN THE PATIENT WOUND SITE. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PHOTOS AND X-RAY WERE UNABLE TO OBTAIN. DEVICE IS NOT AVAILABLE FOR EVALUATION AS DEVICE WAS KEPT BY HOSPITAL AND THROWN AWAY. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00206 | 2022-02-19 | 2022-01-28 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 3.5 15MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 3.5 15MM | 02-012-35-3515 |  | AS REPORTED, APPROXIMATELY FOUR YEARS POST INITIAL TKA, THIS (B)(6) MALE PATIENT COMPLAINED OF PERSISTENT PAIN ASSOCIATED WITH TOTAL KNEE. PATIENT HAD A REVISION OF TOTAL KNEE. THE LINER WAS EXCHANGED. THERE WAS NO BREAKAGE OF THE DEVICE. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PICTURES OF LINER ARE AVAILABLE. THE DEVICE IS NOT AVAILABLE FOR EVALUATION AS THEY WERE LOST/DISPOSED OF DURING STERILIZATION PROCESS. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00209 | 2022-02-21 | 2022-02-01 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 2.5, 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 2.5, 11MM | 02-012-44-2511 |  | AS REPORTED, APPROXIMATELY SEVEN YEARS POST INITIAL LEFT SIDE TKA, THIS (B)(6) PATIENT PRESENTED TO SURGEON WITH SIGNIFICANT POLY WEAR WITH SIGNIFICANT OSTEOLYSIS. SURGEON PERFORMED A TOTAL REVISION KNEE. THERE WAS NO BREAKAGE OF THE DEVICE. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PHOTOS AND X-RAYS RECEIVED. THE DEVICE WILL BE RETURN. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00210 | 2022-02-21 | 2022-01-31 | I, n, j, u, r, y | Manufacturer report | NV GXL LINER NEUTRAL, 32MM ID GROUP 1 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LINER NEUTRAL, 32MM ID GROUP 1 CUPS | 130-32-51 |  | AS REPORTED, APPROXIMATELY FIVE YEARS POST INITIAL RIGHT SIDE THA, THIS FEMALE PATIENT HAD HER RIGHT FEMORAL HEAD AND LINER REMOVED DUE TO OSTEOLYSIS. THE PATIENT RECEIVED A NEW G1 32MM LIPPED XLE LINER AND A 32+7 BIOLOX HEAD. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. AS OF NOTE: "PER SALES REP MOST OF THE DAMAGE TO THE RIM OF THE POLY LINER WAS FROM REMOVAL". X-RAYS RECEIVED. THE DEVICE WILL BE RETURN. | CONCOMITANT MEDICAL PRODUCTS: COCR FEM HEAD 32MM +3.5 OFFSET 12/14 (CAT# 142-32-03 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00215 | 2022-02-22 | 2022-02-01 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REV COMPRESS SCREW LCK CAP KIT, 4.5 X 42MM | EXACTECH, INC. | 320-20-42 | 320-20-42 |  | AS REPORTED, APPROXIMAL 7 MONTHS POST OP LEFT SHOULDER REPLACEMENT, THIS (B)(6) FEMALE PATIENT WAS REVISED FOR A SECOND TIME DUE TO PAIN AND INSTABILITY. THERE WAS LOOSENING OF THE GLENOID AND HUMERAL COMPONENTS WITH IMPLANT BREAKAGE (ONE OF THE BASEPLATE SCREWS). PATIENT HAS POOR HEALTH. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. SURGEON WAS HAPPY WITH THE REVISION RESULT. DEVICES WILL NOT BE RETUNING DUE TO THE HOSPITAL DISPOSING OF THEM. | PENDING EVALUATION CONCOMITANT DEVICE(S): EQUINOXE REVERSE ADAPTER PLATE TRAY +5 - 320-10-05, 6230354, EQUINOXE REVERSE 38MM GLENOSPHERE - 320-01-38, 6747870, EQUINOXE REVERSE 38MM HUMERAL LINER +2.5 - 320-38-03, S130656, EQ REV LOCKING SCREW - 320-15-05, 6945504, EQ REVERSE TORQUE DEFINING SCREW KIT - 320-20-00, 6926688, EQUINOXE, HUMERAL PRIMARY, PRESS FIT 9MM - 300-01-09, 5824106, EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 30MM - 320-20-30, S270960, EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 30MM - 320-20-30, S210714, EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 34MM - 320-20-34, S213651, EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 34MM - 320-20-34, S222422, EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 42MM - 320-20-42, S117060, EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 42MM - 320-20-42, 6287728. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00216 | 2022-02-22 | 2022-02-01 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 42MM HUMERAL LINER +2.5 | EXACTECH, INC. | 320-42-03 | 320-42-03 |  | AS REPORTED, THIS MALE PATIENT'S RIGHT SHOULDER WAS REVISED DUE TO DISLOCATION. ALL IMPLANTS WERE INTACT. THE SURGEON UPSIZED THE TRAY FROM ZERO TO PLUS 5AND UPSIZED LINER FROM 0MM TO A 2.5MM LINER. THERE WAS NO DELAY. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICE IS NOT RETURNING DUE TO HOSPITAL POLICY. | PENDING EVALUATION: CONCOMITANT DEVICE(S): 320-01-00, 11049505 - TRAY, 320-06-42, 02180725 - GLENOSPHERE 42MM. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00214 | 2022-02-22 | 2022-02-01 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | RS GLENOID PLATE EXT CAG +10MM CAGE PEG | EXACTECH, INC. | 320-15-06 | 320-15-06 |  | AS REPORTED, APPROXIMAL 7 MONTHS POST OP LEFT SHOULDER REPLACEMENT, THIS (B)(6) PATIENT WAS REVISED FOR A SECOND TIME DUE TO PAIN AND INSTABILITY. THERE WAS LOOSENING OF THE GLENOID AND HUMERAL COMPONENTS WITH IMPLANT BREAKAGE (ONE OF THE BASEPLATE SCREWS). PATIENT HAS POOR HEALTH. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. SURGEON WAS HAPPY WITH THE REVISION RESULT. DEVICES WILL NOT BE RETUNING DUE TO THE HOSPITAL DISPOSING OF THEM. | PENDING EVALUATION: CONCOMITANT MEDICAL PRODUCTS: EQUINOXE REVERSE ADAPTER PLATE TRAY +5 - 320-10-05, (B)(4). EQUINOXE REVERSE 38MM GLENOSPHERE - 320-01-38, (B)(4). EQUINOXE REVERSE 38MM HUMERAL LINER +2.5 - 320-38-03, (B)(4). EQ REV LOCKING SCREW - 320-15-05, (B)(4). EQ REVERSE TORQUE DEFINING SCREW KIT - 320-20-00, (B)(4). EQUINOXE, HUMERAL PRIMARY, PRESS FIT 9MM - 300-01-09, (B)(4). EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 30MM - 320-20-30, (B)(4). EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 30MM - 320-20-30, (B)(4). EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 34MM - 320-20-34, (B)(4). EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 34MM - 320-20-34, (B)(4). EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 42MM - 320-20-42, (B)(4). EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 42MM - 320-20-42, (B)(4). | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00217 | 2022-02-22 | 2022-01-19 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 38MM GLENOSPHERE | EXACTECH, INC. | 320-01-38 | 320-01-38 |  | AS REPORTED, APPROXIMATELY 9 MONTHS POST OF THE INITIAL RSTA, THIS MALE PATIENT WAS REVISED DUE TO INFECTION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICE WILL NOT BE RETUNING DUE TO THE HOSPITAL. | PENDING EVALUATION. CONCOMITANT DEVICE(S): EQUINOXE REVERSE SHOULDER SUPERIOR AUGMENT GLENOID PLATE 10° - 320-15-02, (B)(4), EQUINOXE REVERSE SHOULDER COMPRESSION SCREW/ LOCKING CAP KIT 4.5 X 30MM SCREW, BLUE - 320-20-30, (B)(4), EQUINOXE REVERSE SHOULDER COMPRESSION SCREW/ LOCKING CAP KIT 4.5 X 38MM SCREW, GREEN - 320-20-38, (B)(4), EQUINOXE REVERSE SHOULDER COMPRESSION SCREW/ LOCKING CAP KIT 4.5 X 30MM SCREW, BLUE - 320-20-30 (B)(4), EQUINOXE REVERSE SHOULDER GLENOSPHERE LOCKING SCREW - 320-15-05, (B)(4), EQUINOXE HUMERAL STEM PRIMARY, PRESS-FIT - 300-01-13, (B)(4), EQUINOXE REVERSE SHOULDER HUMERAL ADAPTER TRAY +0MM - 320-10-00, (B)(4), EQUINOXE REVERSE SHOULDER FIXED ANGLE TORQUE DEFINING SCREW KIT - 320-20-00, (B)(4), EQUINOXE REVERSE SHOULDER 38MM HUMERAL LINER +0MM - 320-38-00, (B)(4), EQUINOXE REVERSE SHOULDER MODULAR DRILL BIT KIT, DIAMETER: 2.0MM, 3.2MM - 321-20-00, (B)(4). | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00212 | 2022-02-22 | 2022-01-31 | I, n, j, u, r, y | Manufacturer report | ACUMATCH GXL 15DEG LINER 32MM SZ H | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | ACUMATCH GXL 15DEG LINER 32MM SZ H | 132-32-28 |  | AS REPORTED, APPROXIMATELY SIXTEEN YEARS POST INITIAL THA, THIS MALE PATIENT HAD A REVISION DUE TO OSTEOLYSIS. THERE IS EVIDENCE OF POLY WEAR. PATIENT CURRENT STATUS IS STABLE. THE DEVICES ARE NOT AVAILABLE FOR EVALUATION. | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION RECEIVED THE FOLLOWING SECTIONS D10, G3, G6, H2, H3 AND H6 HAVE BEEN UPDATED ACCORDINGLY. SECTION D10: CONCOMITANT MEDICAL PRODUCT, 12/14 ZIRCONIA HEAD 32MM +3.5MM NECK (CAT# 148-32-03 / SERIAL# (B)(6)). (H3) THE REVISION REPORTED WAS LIKELY THE RESULT OF OSTEOLYSIS WHICH MAY HAVE BEEN DUE TO PATIENT CONDITIONS AND/OR THE REPORTED PROSTHESIS WEAR. THE WEAR WAS LIKELY DUE TO BEING IMPLANTED FOR APPROXIMATELY 15.5 YEARS. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE COMPONENT WAS NOT RETURNED FOR EVALUATION. THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE REPORTED EVENT OF "OSTEOLYSIS¿ IS ASSOCIATED WITH DESTRUCTION OR DISAPPEARANCE OF BONE TISSUE LIKELY RELATED TO WEAKENED INTEGRATION OF AN IMPLANT AT THE BONE-IMPLANT INTERFACE. FURTHERMORE, THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE REPORTED EVENT OF "PROSTHESIS WEAR¿ IS ASSOCIATED WITH MATERIAL DAMAGE TO A SURFACE, USUALLY INVOLVING PROGRESSIVE LOSS OR DISPLACEMENT OF MATERIAL, DUE TO RELATIVE MOTION BETWEEN THAT SURFACE AND A CONTACTING SUBSTANCE OR SUBSTANCES. SECTION D1: BRAND NAME SECTION D4: MODEL NUMBER, CATALOG NUMBER, SERIAL NUMBER, EXPIRATION DATE, UNIQUE IDENTIFIER (UDI) # SECTION G4: PMA/510(K)NUMBER. SECTION H4: DEVICE MANUFACTURE DATE. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00211 | 2022-02-22 | 2022-02-01 | I, n, j, u, r, y | Manufacturer report | CR TIBIAL INSERT SZ 5, 13MM, SLOPE ++ | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | CR TIBIAL INSERT SZ 5, 13MM, SLOPE ++ | 200-75-13 |  | AS REPORTED, APPROXIMATELY NINE YEARS POST INITIAL TKA, THIS 70 Y/O MALE PATIENT HAD A REVISION OF A CR TIBIAL POLY INSERT. KNEE WAS OPENED AND DEBRIDED AND WASHED OUT, CR TIBIAL POLY 13MM ++ WAS EXPLANTED AND NEW POLY 15MM NEUTRAL INSERT WAS IMPACTED INTO TIBIAL TRAY. KNEE WAS CLOSED. THERE WAS NO BREAKAGE TO DEVICE. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. X-RAYS RECEIVED. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO HOSPITAL POLICY. | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION RECEIVED THE FOLLOWING SECTIONS B5, G3, G6, H2 AND H6 HAVE BEEN UPDATED ACCORDINGLY. SECTION B5: ADDITIONAL INFORMATION RECEIVED INDICATES THAT PATIENT WAS HAVING PAIN AND COULD HEAR METAL ON METAL SOUND DURING GAIT. THERE WAS EVIDENCE OF POLY WEAR. ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00213 | 2022-02-22 | 2022-02-03 | I, n, j, u, r, y | Manufacturer report | LOGIC CR TIB INSERT SLOPE++, SZ 2, 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC CR TIB INSERT SLOPE++, SZ 2, 13MM | 02-012-49-2013 |  | AS REPORTED, APPROXIMATELY FOUR YEARS POST INITIAL TKA, THIS FEMALE PATIENT WAS REVISED TO A PERSONA REVISION KNEE FOR BOTH THE FEMUR AND TIBIA DUE TO PAIN, WEAR AND OSTEOLYSIS. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION. PHOTOS ATTACHED. THE DEVICE IS AVAILABLE FOR EVALUATION. | CONCOMITANT MEDICAL PRODUCTS: LOGIC CR FEMORAL CEM, LEFT SZ 2 (CAT# 02-010-03-0220 / SERIAL#:(B)(4)). LGC TIBIAL FIT TRAY CEM SZ 2F / 3T (CAT# 02-012-45-2030 / SERIAL#:(B)(4)). ADVANCED PATELLA 29M 3 PEG IMPLANT (CAT# 200-07-29 / SERIAL#:(B)(4)). 3" DRILL BIT, MOD. HEX 2 PACK (CAT# 201-78-89 / SERIAL#:(B)(4)). THREADED PIN SIZE 2.3 COLLARED 2PK (CAT# 521-78-23 / SERIAL#:(B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00218 | 2022-02-22 | 2022-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EQUINOXE | HUM LNR IMPACT TIP 42M | EXACTECH, INC. |  | 321-07-42 |  | IT WAS REPORTED THAT DURING IMPACTION, THE RESIDENT IMPACTED THE HUMERAL LINER TRIAL INTO THE HUMERAL ADAPTER TRAY TO GET IT SEATED, UTILIZING A 42MM HUMERAL LINER IMPACTOR TIP. UPON IMPACTION, THE TOP OF THE IMPACTOR TIP FRACTED OFF. THERE WAS NO SURGICAL DELAY. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICE TO BE RETURNED. | PENDING EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00220 | 2022-02-23 | 2022-02-02 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | SHOULDER COMPONENTS | EXACTECH, INC. |  | UNK |  | IT WAS REPORTED THAT THIS PATIENT HAD A REVISION OF THE LEFT SHOULDER DUE TO INFECTION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETUNING. | SECTION H10: (H3) BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN OR MANUFACTURING ISSUES. THE CAUSE OF SURGICAL REVISION FOR INFECTION CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00219 | 2022-02-23 | 2022-02-01 | I, n, j, u, r, y | Manufacturer report | VANTAGE | TIBIAL INSERT MB SZ 3 RT 8MM | EXACTECH, INC. |  | 350-42-23 |  | IT WAS REPORT THAT A PATIENT HAD A REVISION OF THE RIGHT ANKLE DUE TO PAIN. THE SURGEON SAW INFLAMMATION ON THE LATERAL SIDE OF THE MALLEOLUS AND EXCHANGED THE INSERT. HE ALSO FUSED THE FIBULA AND DEVIATED THE CALCANEUS DURING THE SURGERY. THERE WERE NO TRACES OF WEAR ON THE EXPLANT. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETUNING. | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00221 | 2022-02-24 | 2022-01-23 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | COMPRESSION SCREW | EXACTECH, INC. |  | UNK |  | AS REPORTED, APPROXIMATELY 3 YEARS POST OP THE INITIAL RSTA, THIS (B)(6) MALE PATIENT WAS REVISED DUE TO UNKNOWN REASONS. PATIENT DOES NOW KNOW WHEN IMPLANT BROKE. THE CAGE OF THE BASEPLATE AND A SCREW BROKE IN THE SCAPULA. BROKEN IMPLANTS WERE REMOVED AND REVISED TO A REVERSE. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETURNING. | DEVICE EVALUATED BY MFR: PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00222 | 2022-02-24 | 2022-01-23 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | SUP/POST AUG PLATE, R RS GLENOID BASEPLATE | EXACTECH, INC. | SPECIFIC DEVICE NOT REPORTED | 320-15-08 |  | AS REPORTED, APPROXIMATELY 3 YEARS POST OP THE INITIAL RSTA, THIS (B)(6) Y/O MALE PATIENT WAS REVISED DUE TO UNKNOWN REASONS. PATIENT DOES NOW KNOW WHEN IMPLANT BROKE. THE CAGE OF THE BASEPLATE AND A SCREW BROKE IN THE SCAPULA. BROKEN IMPLANTS WERE REMOVED AND REVISED TO A REVERSE. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETUNING. | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00230 | 2022-02-28 | 2019-07-15 | I, n, j, u, r, y | Manufacturer report | EXACTECH | TRULIANT TIB TAMP GUIDE | EXACTECH, INC. |  | 02-029-22-2000 |  | IT WAS REPORTED FROM THE US THAT DURING A TOTAL KNEE ARTHROPLASTY SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE INSTRUMENTATION. THE TIBIAL TAMP GUIDE HANDLE BUTTON THAT LOCKS THE HANDLE POPPED OFF WHILE IMPACTING THE PUNCH INTO THE TIBIA. ALL PIECES WERE CONFIRMED REMOVED FROM THE WOUND AND THE PATIENT. THERE WAS A DELAY, BUT THE PATIENT WAS NOT ADVERSELY AFFECTED. THERE IS NO REPORTED PATIENT INJURY DUE TO THE DEVICE PIECE RETRIEVAL FROM THE SURGICAL SITE. THE REP WAS PRESENT AT THE TIME OF SURGERY. THE DEVICE WAS NOT RECEIVED FOR ANALYSIS. ATTEMPTS WERE MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REP, BUT NO RESPONSES WERE RECEIVED. NO OTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS ABOUT THE DEVICE/EVENT/PATIENT. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO MANUFACTURING ISSUES. THERE IS NO REPORTED PATIENT INJURY DUE TO THE DEVICE PIECE RETRIEVAL FROM THE SURGICAL SITE. PER (B)(4), THE ROOT CAUSE FOR THE BUTTON DISENGAGEMENT REPORTED IN (B)(4) WAS LIKELY THE RESULT OF INSUFFICIENT OVERLAP BETWEEN THE SNAP RING AND THE LIP OF THE HOUSING TO PREVENT THE OCCURRENCE OF DISENGAGEMENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS NOT AVAILABLE FOR ANALYSIS. ENGINEERING ANALYSIS COMPLETED BY NF ON (B)(6) 2020. (B)(4). FINDINGS: THE BUTTON COMPONENT (02-029-222-2005) OF THE TIBIAL TAMP GUIDE DISENGAGED FROM THE INSTRUMENT (02-029-22-2000). ENGINEERING PRINTS OF 02-029-122-2000 WERE REVIEWED. THE OVERLAP BETWEEN THE SNAP RING¿S OD AND THE ID OF THE HOUSING LIP IS .008 FOR THE WORST-CASE SCENARIO. ALTHOUGH THE OVERLAP IS POSITIVE, THE SMALL AMOUNT COULD CAUSE DISENGAGEMENT. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: PER (B)(4), THE ROOT CAUSE FOR THE BUTTON DISENGAGEMENT REPORTED IN (B)(4) WAS LIKELY THE RESULT OF INSUFFICIENT OVERLAP BETWEEN THE SNAP RING AND THE LIP OF THE HOUSING TO PREVENT THE OCCURRENCE OF DISENGAGEMENT. CORRECTIVE ACTION TAKEN: PER (B)(4), THE COMMITTEE HAS DETERMINED THAT THE DESIGN WILL BE UPDATED, TESTED, AND IMPLEMENTED, AFTER WHICH A STAGED MARKET WITHDRAWAL WAS CONDUCTED. PER (B)(4): A STOCK AUDIT WAS PERFORMED FOR THE DEVICES LOCATED IN 01 (COC). SEE (B)(4) FOR DETAILS. THE DESIGN WAS UPDATED TO INCREASE THE OVERLAP OF THE SNAP RING AND THE HOUSING LIP COMPONENTS. SEE (B)(4) FOR DETAILS. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO MANUFACTURING ISSUES. THERE IS NO REPORTED PATIENT INJURY DUE TO THE DEVICE PIECE RETRIEVAL FROM THE SURGICAL SITE. PER (B)(4), THE ROOT CAUSE FOR THE BUTTON DISENGAGEMENT REPORTED IN CASE (B)(4) WAS LIKELY THE RESULT OF INSUFFICIENT OVERLAP BETWEEN THE SNAP RING AND THE LIP OF THE HOUSING TO PREVENT THE OCCURRENCE OF DISENGAGEMENT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00223 | 2022-02-28 | 2019-01-01 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 3 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 3 9MM | 02-012-35-3009 |  | AS REPORTED BY THE EXACTECH KNEE REPLACEMENT STUDY, APPROXIMATELY SEVEN YEARS POST INITIAL LEFT SIDE TKA, THIS (B)(6) FEMALE PATIENT EXPERIENCED POLYETHYLENE WEAR WITH EVIDENCE OF OSTEOLYSIS. PATIENT LEFT KNEE MEDIAL POLYETHYLENE TIBIAL WEAR EVIDENT ON RADIOGRAPHS TAKEN ON (B)(6) 2022. PATIENT REPORTS GRADUAL PROGRESSION OF PAIN AND SWELLING TO LEFT KNEE FOR 2-3 YEARS. THE EVENT IS ONGOING. | CONCOMITANT MEDICAL PRODUCTS: LOGIC FEMORAL PS CEM LEFT SZ 3 (CAT# 02-010-01-0230), LGC TIBIAL FIT TRAY CEM SZ 3F / 3T (CAT# 02-012-45-3030), THREE PEG PATELLA 32MM (CAT# 200-02-32). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00224 | 2022-02-28 | 2022-02-07 | I, n, j, u, r, y | Manufacturer report | NV GXL LNR, +5LAT, 36MM G3-56/58MM CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LNR, +5LAT, 36MM G3-56/58MM CUPS | 136-36-53 |  | AS REPORTED, THIS PATIENT HAD A REVISION DUE TO PREMATURE POLYETHYLENE WEAR AND OSTEOLYSIS. NO ADDITIONAL INFORMATION AVAILABLE. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00225 | 2022-02-28 | 2022-02-03 | I, n, j, u, r, y | Manufacturer report | TRULIANT PS POR FEM PS POR LEFT SZ 5 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT PS POR FEM PS POR LEFT SZ 5 | 02-020-12-0250 |  | AS REPORTED, APPROXIMATELY TWO YEARS POST INITIAL LEFT SIDE TKA, THIS MALE PATIENT WAS BROUGHT BACK WITH LEFT KNEE STARTUP PAIN. THE POLY WAS REMOVED. THE FEMUR WAS CHECKED AND FOUND TO BE LOOSE. THE TIBIA WELL FIXED. THE PATIENT WAS REVISED WITH A SIZE 5 LOGIC CC FEMUR WITH A 24X120 STEM, A 2MM OFFSET COUPLER, AND A SIZE 5 15MM INSERT. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. SALES REP WAS UNABLE TO OBTAIN PHOTOS OR X-RAYS. THE DEVICES WILL BE RETURN. | CONCOMITANT MEDICAL PRODUCTS: TRULIANT TIB IMP PSC INSERT SZ 5, 13MM (CAT# 02-022-44-5013 / (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00226 | 2022-02-28 | 2022-02-07 | I, n, j, u, r, y | Manufacturer report | GXL NUETRAL LINER, G0 28MM ID | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | COCR FEM HEAD 28MM +3.5 OFFSET 12/14 | 130-28-50 |  | REVISION OP REPORT RECEIVED VIA THE LEGAL DEPARTMENT- REVISION LEFT HIP PROCEDURE ON (B)(6) 2022. PRE & POST OPERATIVE DIAGNOSIS: FAILED TOTAL HIP ARTHROPLASTY SURGERY PERFORMED: LEFT HIP REVISION ARTHROPLASTY OF THE ACETABULAR AND FEMORAL COMPONENT. THE SURGEON NOTICED SIGNIFICANT ACETABULAR WEAR AND CYSTIC CHANGES IN THE ACETABULUM. THE PATIENT TOLERATED THE PROCEDURE WELL WITH NO COMPLICATIONS. SHE WAS TAKEN BACK TO THE POSTANESTHESIA RECOVERY ROOM IN STABLE FASHION WITH SUCCESSFUL LEFT ANTERIOR HIP REVISION ARTHROPLASTY OF BOTH THE ACETABULAR AND FEMORAL COMPONENTS. THERE IS NO OTHER INFORMATION AVAILABLE. | THIS FOLLOW-UP REPORT IS BEING SUBMITTED TO RELAY ADDITIONAL AND/OR CORRECTED INFORMATION. THE FOLLOWING SECTIONS WERE UPDATED/CORRECTED: H6. MDR SECTION CODES UPDATED/CORRECTED: A, D, G. THE MOST LIKELY CAUSE FOR THE REVISION REPORTED DUE TO EARLY PROSTHESIS WEAR AND OSTEOLYSIS IS A COMBINATION OF THE RISK FACTORS FOR GXL: SUCH AS USE ERROR, IMPLANT POSITIONING, IMPLANT SIZE SELECTION, AND PATIENT FACTORS (FITNESS FOR SURGERY, BIOMECHANICS, ACTIVITY LEVEL AND LOCAL TISSUE OXIDATION POTENTIAL). IF ANY FURTHER INFORMATION IS OBTAINED THAT WOULD CHANGE OR ALTER ANY INFORMATION PROVIDED, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00227 | 2022-02-28 | 2021-06-16 | I, n, j, u, r, y | Manufacturer report | LOGIC FEMORAL PS CEM LEFT SZ 2 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC FEMORAL PS CEM LEFT SZ 2 | 02-010-01-0220 |  | AS REPORTED BY THE EXACTECH KNEE REPLACEMENT STUDY, APPROXIMATELY FIVE YEARS POST INITIAL TKA, THIS (B)(6) PATIENT EXPERIENCED ANTERIOR KNEE PAIN. PATIENT HAD A GRADUAL ONSET LEFT ANTERIOR KNEE PAIN AND "LOCKING" ON RISING FROM A CHAIR. WAITS A MINUTE AND CAN MOBILIZE NORMALLY WITHOUT PAIN. QUADS WEAKNESS DETECTED GRADE 4+. THE EVENT IS ONGOING. THIS EVENT IS POSSIBLY RELATED TO THE DEVICE AND TO THE PROCEDURE. | CONCOMITANT MEDICAL PRODUCTS. LOGIC TIBIA PS MOD INSRT SZ 2 9MM (CAT# 02-012-35-2009). LGC TIBIAL FIT TRAY CEM SZ 2F / 1T (CAT# 02-012-45-2010). THREE PEG PATELLA 32MM (CAT# 200-02-32). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00228 | 2022-02-28 | 2022-02-08 | I, n, j, u, r, y | Manufacturer report | ACUMATCH 0 DEGREE LINER 28MM 54/56MM | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | ACUMATCH 0 DEGREE LINER 28MM 54/56MM | 130-28-08 |  | AS REPORTED, PATIENT HAD A REVISION FOR AN UNKNOWN REASON. THIS WAS IMPLANTED IN (B)(6) 2001. SURGEON USED OUR M-SERIES FOR YEARS AND WAS PART OF THE DESIGN TEAM. PHOTOS AND X-RAYS RECEIVED. NO ADDITIONAL INFORMATION AVAILABLE. | PATIENT ORIGINAL IMPLANTS WAS IN (B)(6) 2001. ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00229 | 2022-02-28 | 2017-01-10 | I, n, j, u, r, y | Manufacturer report | OPTETRAK ASY HI-FLEX PS CEM FEM, SZ 3, RIGHT | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | OPTETRAK ASY HI-FLEX PS CEM FEM, SZ 3, RIGHT | 244-03-03 |  | AS REPORTED, THIS (B)(6) FEMALE PATIENT, WEIGHT (B)(6)., APPROXIMATELY FIVE YEARS POST INITIAL RIGHT SIDE TKA, PATIENT HAD A REVISION OF THE FEMORAL AND TIBIAL COMPONENTS, LATERAL RETINACULAR RELEASE, AND LATERAL PATELLA FACETECTOMY. | CONCOMITANT MEDICAL PRODUCTS: OPTETRAK HI-FLEX TIBIAL INSERT SZ 3 15MM (CAT# 244-23-15 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-02
1038671-2022-00232 | 2022-03-01 | 2021-06-19 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | HUMERAL LINER, 40MM, +2.5 | EXACTECH, INC. | 320-40-03 | 320-40-03 |  | THIS REVISION WAS FOUND DURING THE INVESTIGATION OF (B)(4). REASON FOR THE REVISION IS UNKNOWN. | PENDING EVALUATION. CONCOMITANT MEDICAL PRODUCTS: 320-32-40, 5993404 - EXPANDED GLENOSPHERE, 40MM, FOR SMALL REVERSE. 320-35-01, 6066265 - SMALL GLENOID PLATE. 300-01-15, 6113495 - EQUINOXE, HUMERAL STEM PRIMARY, PRESS FIT 15MM. 320-20-34, 6171767 - EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 34MM. 320-10-00, 6203171 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0. 320-15-05, 6217400 - EQ REV LOCKING SCREW. 320-20-00, 6224120 - EQ REVERSE TORQUE DEFINING SCREW KIT. 320-20-30, 6225658 - EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 30MM. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00231 | 2022-03-01 | 2019-10-16 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 38MM HUMERAL LINER +0 | EXACTECH, INC. | 320-38-00 | 320-38-00 |  | THIS REVISION WAS FOUND DURING THE INVESTIGATION OF (B)(4). REASON FOR THE REVISION IS UNKNOWN. | PENDING EVALUATION . CONCOMITANT MEDICAL PRODUCTS: 320-15-01, 5222680 - EQ REV GLENOID PLATE. 320-01-38, 5237307 - EQUINOXE REVERSE 38MM GLENOSPHERE. 320-10-00, 5253917 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0. 320-20-26, 5280360 - EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 26MM. 320-20-22, 5285155 - EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 22MM. 320-20-26, 5285237 - EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 26MM. 320-20-22, 5288013 - EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 22MM. 320-20-00, 5288766 - EQ REVERSE TORQUE DEFINING SCREW KIT. 320-15-05, 5300560 - EQ REV LOCKING SCREW. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00233 | 2022-03-02 | 2022-02-11 | I, n, j, u, r, y | Manufacturer report | TRULIANT TIBIAL TRAY | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT TIBIAL TRAY | UNKNOWN |  | AS REPORTED, SURGEON REPORTED THAT HE MIGHT POSSIBLY HAVE A PATIENT WITH A LOGIC POROUS TIBIAL TRAY THAT IS LOOSE. HE WILL CONTINUE TO FOLLOW THE PATIENT AND INFORMED US IF THEY REQUIRE ANY REVISION SURGERY. THE X-RAY SHOWS A RADIOLUCENT LINE BUT HE EXPLAINED SOMETIMES THESE ARE ASYMPTOMATIC AND RESOLVE. NO ADDITIONAL INFORMATION AVAILABLE. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00235 | 2022-03-08 | 2022-02-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | VANTAGE | MOBILE GAP CHECK SZ 3-4 | EXACTECH, INC. |  | 351-10-10 |  | IT WAS REPORTED THAT THE GAP CHECK TOOL BROKE DURING SURGICAL USE. THE SURGEON WAS PERFORMING THE SURGERY WITHOUT SUPPORT OF A SALES REP. NO PARTS OR PIECES OF THE DEVICE FELL INTO THE PATIENT WOUND SITE. THERE WAS NO SURGICAL DELAY. DEVICE TO BE RETURNED. | PENDING EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00239 | 2022-03-08 | 2019-06-13 | I, n, j, u, r, y | Manufacturer report | EXACTECH | UNIVERSAL STEM EXTRACTOR | EXACTECH, INC. |  | 109-00-00 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE INSTRUMENTATION. WHEN ATTEMPTING TO REMOVE THE PREVIOUSLY IMPLANTED [COMPETITOR] LEFT HIP STEM, THE STEM EXTRACTOR BROKE OFF THE SLAP HAMMER. ALL ITEMS WERE RETRIEVED, AND NOTHING WAS LEFT IN THE PATIENT. THE STAFF GOT A SECOND SLAP HAMMER AND ATTACHED THE THREADED ADAPTER. THE SURGEON SUCCESSFULLY REMOVED THE STEM, HOWEVER, THE THREADED ADAPTER GOT STUCK IN THE STEM. THE JAMMED DEVICE IS NOT REPORTABLE, IT IS NOT LIKELY TO CAUSE INJURY. THE PATIENT¿S HIP WAS INFECTED [COMPETITOR DEVICE], SO THE STEM AND ADAPTER WERE DISCARDED TOGETHER BY THE FACILITY. THERE WAS NO DELAY, AND THE PATIENT WAS NOT ADVERSELY AFFECTED. PATIENT LEFT THE ROOM STABLE. THE SLAP HAMMER AND STEM EXTRACTOR WERE RETURNED FOR ANALYSIS. THE REPRESENTATIVE WAS PRESENT AT THE TIME OF SURGERY. ATTEMPTS WERE MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REPRESENTATIVE, BUT NO RESPONSES WERE RECEIVED. THERE HAS BEEN NO OTHER INFORMATION RECEIVED FOR THE DEVICES OR EVENT FROM THE CONTACTS. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. ENGINEER EVALUATION: THE DEVICES WERE AVAILABLE FOR ANALYSIS. DESIGN-RELATED ISSUES: THE DESIGNS OF THE SLAP HAMMER AND STEM EXTRACTOR HAVE BEEN IN THE FIELD SINCE 2001 AND 2003, RESPECTIVELY. EXACTECH IS AWARE OF ONE SIMILAR COMPLAINT REPORT INVOLVING BROKEN SLAP HAMMERS SINCE 2001. BECAUSE BOTH THE SLAP HAMMER AND STEM EXTRACTOR ARE USED IN TOTAL HIP SURGERIES, SALES DATA FOR FEMORAL STEMS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF LESS THAN 0.5%. THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: (B)(4). THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE DEVICE WILL BE CAPTURED IN THE SCOPE LIST OF THE RMR. THE SLAP HAMMER FRACTURE REPORTED WAS LIKELY THE RESULT OF LONG-TERM, RIGOROUS USE, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE THREADS INSIDE THE STEM EXTRACTOR. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. CONCLUSION: DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THE STEM EXTRACTOR BROKE OFF THE SLAP HAMMER; ALL ITEMS WERE RETRIEVED; NOTHING WAS LEFT IN THE PATIENT. THE SLAP HAMMER FRACTURE REPORTED WAS LIKELY THE RESULT OF LONG-TERM, RIGOROUS USE, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE THREADS INSIDE THE STEM EXTRACTOR. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00240 | 2022-03-08 | 2021-09-30 | I, n, j, u, r, y | Manufacturer report | NV GXL LINR, NTRL, 36MM ID, GROUP 3 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LINR, NTRL, 36MM ID, GROUP 3 CUPS | 130-36-53 |  | AS REPORTED BY THE LEGAL BRIEF, APPROXIMATELY EIGHT YEARS POST INITIAL LEFT SIDE THA, THIS MALE PATIENT PRESENTED TO SURGEON EXPERIENCING PAIN AND SEVERE OSTEOLYSIS. A REVISION SURGERY WAS NECESSARY DUE TO A FAILED POLYETHYLENE ACETABULAR LINER RESULTING IN SEVERE OSTEOLYSIS. DURING THE REVISION SURGERY, SURGEON NOTED THAT CUP HAD SEVERE OSTEOLYSIS AROUND THE SUPERIOR LATERAL REGIONS AND THERE WAS SEVERE AMOUNT OF OSTEOLYSIS ALONG THE POSTERIOR WALL/SUPERIOR LATERAL ACETABULUM AND ALONG THE ISCHIUM. NO ADDITIONAL INFORMATION IS AVAILABLE. THE DEVICES ARE NOT AVAILABLE FOR EVALUATION. | CONCOMITANT MEDICAL PRODUCTS BIOLOX DELTA FEMORAL HEAD 36MM OD, +0MM (CAT# 170-36-00 / (B)(4)), NV CROWN CUP CLSTR HOLE 56MM GROUP 3 (CAT# 180-01-56 / (B)(4)), 6.5MM ACETABULAR BONE SCREW 30MM (CAT# 122-65-30 / (B)(4)), ELEMENT-STEM, COLLARLESS W/HA, STD OFFSET, SZ 14 (CAT# 164-01-14 / (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00241 | 2022-03-08 | 2022-02-15 | I, n, j, u, r, y | Manufacturer report | BIOLOX DELTA FEMORAL HEAD 36MM OD, +0MM | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | BIOLOX DELTA FEMORAL HEAD 36MM OD, +0MM | 170-36-00 |  | AS REPORTED, APPROXIMATELY FIVE YEARS POST INITIAL RIGHT SIDE THA, THIS MALE PATIENT WAS REVISED DUE TO OSTEOLYSIS AND THE STEM WAS ALSO SUBSIDED. PATIENT RECEIVED DEPUY IMPLANTS. THERE WAS NO BREAKAGE OF DEVICE AND NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. SALES REP WAS UNABLE TO OBTAIN PHOTOS OR X-RAYS. THE DEVICES ARE NOT AVAILABLE FOR EVALUATION DUE TO HOSPITAL KEEP THE IMPLANTS. | CONCOMITANT MEDICAL PRODUCTS: NV GXL LNR, +5LAT, 36MM G2-52/54MM CUPS (CAT# 136-36-52), ALTEON 6.5MM SCREW, 20MM (CAT# 180-65-20) , INTEGRIP CC, CLUSTER 52MM, G2 (CAT# 186-01-52) ,AND BIOLOX DELTA FEMORAL HEAD 36MM OD, +0MM (CAT# 170-36-00). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00234 | 2022-03-08 | 2021-11-25 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | HUMERAL STEM PRIMARY, PRESS FIT 11MM | EXACTECH, INC. | 300-01-11 | 300-01-11 |  | AS REPORTED, APPROXIMATELY 3.5 YEARS POST OP INITIAL TSA, THIS (B)(6) PATIENT WAS EXPERIENCING PAIN DUE TO ASCETIC HUMERAL LOOSENING SEEM ON X-RAYS. PATIENT WAS OFFERED REVISION SURGERY, BUT DECLINED AT THIS STAGE. PATIENT WANTS TO WAIT. THE CASE REPORT FORM INDICATES THIS EVENT IS POSSIBLY RELATED TO DEVICES AND/OR PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | PENDING EVALUATION . CONCOMITANT DEVICE(S): 310-01-50 - EQUINOXE, HUMERAL HEAD SHORT, 50MM (BETA), 300-10-15 - EQUINOXE REPLICATOR PLATE 1.5MM O/S, AND 314-02-14 - EQUINOXE GLENOID, PEGGED BETA, LARGE. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00236 | 2022-03-08 | 2022-02-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | VANTAGE | LP TALAR IMPACTOR INSERT | EXACTECH, INC. |  | 351-07-04 |  | IT WAS REPORTED THAT THE TALAR IMPACTOR INSERT BROKE DURING SURGICAL USE. THE SURGEON WAS PERFORMING THE SURGERY WITHOUT SUPPORT OF A SALES REP. NO PARTS OR PIECES OF THE DEVICE FELL INTO THE PATIENT WOUND SITE. THERE WAS NO SURGICAL DELAY. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICE TO BE RETURNED. | PENDING EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00250 | 2022-03-09 | 2022-02-07 | I, n, j, u, r, y | Manufacturer report | NV GXL LINR, NTRL, 36MM ID, GROUP 3 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LINR, NTRL, 36MM ID, GROUP 3 CUPS | 130-36-53 |  | AS REPORTED, APPROXIMATELY TWELVE YEARS POST RIGHT THA, THIS MALE PATIENT NEEDED NEW POLY. CUP AND POLY WERE REMOVED. SMITH AND NEPHEW DUAL MOBILITY WAS IMPLANTED. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. X-RAYS AND PHOTOS RECEIVED. THE DEVICE WILL BE RETURN. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00249 | 2022-03-09 | 2022-02-04 | I, n, j, u, r, y | Manufacturer report | NV GXL LINR, NTRL, 36MM ID, GROUP 4 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LINR, NTRL, 36MM ID, GROUP 4 CUPS | 130-36-54 |  | AS REPORTED, APPROXIMATELY TWELVE YEARS POST INITIAL RIGHT THA, THIS (B)(6) MALE PATIENT NEEDED A HEAD AND POLY SWAP. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PHOTO RECEIVED. THE DEVICE WILL BE RETURN. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00252 | 2022-03-09 | 2022-02-21 | I, n, j, u, r, y | Manufacturer report | OPTETRAK ASY HI-FLEX PS CEM FEM, SZ 2, RIGHT | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | OPTETRAK ASY HI-FLEX PS CEM FEM, SZ 2, RIGHT | 244-03-02 |  | AS REPORTED, APPROXIMATELY 10 YEARS POST INITIAL RIGHT TKA, THIS (B)(6) FEMALE PATIENT WAS BROUGHT BACK WITH RIGHT KNEE PAIN. BONE SCAN SUGGESTED LOOSENING TO BOTH THE FEMORAL AND TIBIAL COMPONENTS. THE INSERT WAS REMOVED WITH SIGNS OF WEAR. THE TIBIA WAS CHECKED AND FOUND TO BE LOOSE AND REMOVED. THE FEMUR CHECKED AND REMOVED AS WELL. DUE TO THE LARGE AMOUNT OF BONE LOSS ON THE FEMUR, THE PATIENT RECEIVED A COMPETITOR'S DISTAL FEMORAL KNEE REPLACEMENT. THERE WAS NO BREAKAGE OF DEVICE. THERE WAS A 45 MINUTE SURGICAL DELAY/PROLONGATION. THIS TURNED INTO A MORE COMPLICATED SURGERY THAN EXPECTED DUE TO OSTEOLYSIS TO THE FEMUR AND TIBIA. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. X-RAYS AND PHOTOS RECEIVED. THE DEVICES WILL BE RETURN. AS PER SURGEON: "THIS TURNED INTO A MORE COMPLICATED SURGERY THAN EXPECTED DUE TO OSTEOLYSIS TO THE FEMUR AND TIBIA." | CONCOMITANT MEDICAL PRODUCTS: OPTETRAK HI-FLEX TIBIAL INSERT, SZ 2, 11MM (CAT# 244-22-11 / SERIAL# (B)(4)) AND CEMENTED FINNED TIB. TRA SZ 2F/2T (CAT# 200-04-22 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00251 | 2022-03-09 | 2022-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TIBIAL INSERT TRIAL SZ 5 15 MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TIBIAL INSERT TRIAL SZ 5 15 MM | 201-25-15 |  | AS REPORTED, DURING IMPACTION, THE SURGEON WAS ATTEMPTING TO TRIAL A TIBIAL INSERT. HE WAS HAVING SOME DIFFICULTY SEATING THE INSERT, SO HE TRIED IMPACTING IT WITH THE APPROPRIATE INSTRUMENT AND A SMALL PIECE BROKE OFF. ONE PIECE CAME OFF DURING IMPACTION AND WAS RETRIEVED. PATIENT WAS REVISED TO CR TIBIAL INSERT, SIZE 5, 15MM. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION. PHOTO RECEIVED. THE DEVICE WILL BE RETURN. AS OF NOTE: "PATIENT REVISION REPORTED UNDER MFR# 1038671-2022-00211." | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00248 | 2022-03-09 | 2022-05-04 | I, n, j, u, r, y | Manufacturer report | NV GXL LINER NEUTRAL, 32MM ID GROUP 1 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LINER NEUTRAL, 32MM ID GROUP 1 CUPS | 130-32-51 |  | AS REPORTED BY LEGAL DEPARTMENT, APPROXIMATELY FOUR YEARS POST INITIAL LEFT SIDE THA, THIS FEMALE PATIENT VISITED SURGEON DUE TO FOR THE PAST 2 YEARS, PATIENT HAS BEEN COMPLAINING OF SEVERE PAIN AND UNABLE TO WALK FOR A LONG TIME. SURGEON TOLD PATIENT THAT THE X-RAYS SHOWED SOME DAMAGE ON THE LT-HIP. PATIENT MADE AN APPOINTMENT WITH HER NEW ORTHO TO DISCUSS THE SURGERY. NO ADDITIONAL INFORMATION AVAILABLE. | CONCOMITANT MEDICAL PRODUCTS: COCR FEM HEAD 32MM +0 OFFSET 12/14 (CAT# 142-32-00/ SERIAL# (B)(4)). NOVATION ELEMENT RO S/O COL SZ 12 (CAT# 164-13-12 / SERIAL# (B)(4)). ALTEON 6.5MM SCREW, 20MM (CAT# 180-65-20 / SERIAL# (B)(4)). INTEGRIP CC, CLUSTER 50MM, G1 (CAT# 186-01-50 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00244 | 2022-03-09 | 2022-02-09 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 46MM HUMERAL LINER +2.5 | EXACTECH, INC. | 320-46-03 | 320-46-03 |  | AS REPORTED, APPROXIMATELY 1.5 YEARS POST OP INITIAL LEFT TSA, THIS (B)(6) PATIENT WAS REVISED DUE TO DISLOCATION OF JOINT WHICH LEAD TO DISLODGING OF THE POLY. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETUNING DUE TO THE HOSPITAL DISPOSING OF THEM. | PENDING EVALUATION . CONCOMITANT DEVICE(S): 320-01-46, (B)(4), EQUINOXE REVERSE 46MM GLENOSPHERE. 320-15-03, (B)(4), RS GLENOID PLATE L POST AUG. 8 DEG, LEFT, 300-30-11, (B)(4), EQUINOXE PRESERVE STEM 11MM. 320-15-05, (B)(4), EQ REV LOCKING SCREW. 320-10-00, (B)(4), EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0. 320-46-03, (B)(4), EQUINOXE REVERSE 46MM HUMERAL LINER +2.5. 320-20-00, (B)(4), EQ REVERSE TORQUE DEFINING SCREW KIT. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00247 | 2022-03-09 | 2022-02-15 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE TRAY ADAPTER PLATE TRAY +5 | EXACTECH, INC. | 320-10-05 | 320-10-05 |  | AS REPORTED, APPROXIMATELY 9.5 YEARS POST OP INITIAL LEFT TSA, THIS MALE PATIENT WAS REVISED DUE TO A RADIAL PATTERN ON THE TRAY AND LINER AND POSSIBLE MICRO MOTION. SURGEON EXPLANTED THE TRAY, LINER, AND GLENOSPHERE. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL BE RETURNING. | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00246 | 2022-03-09 | 2022-02-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ERGO BROACH HANDLE | BROACH HANDLE | EXACTECH, INC. |  | 301-05-01 |  | IT WAS REPORTED THAT DURING IMPACTION A LITTLE SLIVER OF METAL BROKE OFF ON THE LEFT SIDE OF THE BROACH HANDLE BETWEEN 0 AND 10 DEGREES OF VERSION. THE BROACH HANDLE WORKS BUT LOSING THAT PIECE OF METAL HAS CREATED A SHARP EDGE. NO PARTS OR PIECES FELL INTO THE PATIENT¿S WOUND SITE. THERE WAS NO SURGICAL DELAY. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. | PENDING EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00245 | 2022-03-09 | 2022-02-08 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE TRAY ADAPTER PLATE TRAY +5 | EXACTECH, INC. | 320-10-05 | 320-20-00 |  | AS REPORTED, APPROXIMATELY 9 MONTHS POST OF INITIAL LEFT TSA, THIS (B)(6) PATIENT WAS REVISED DUE TO THE TRAY DISLOCATING FROM THE STEM AND WAS ¿FLOATING¿ INSIDE THE JOINT SPACE. THE +5 TRAY, POLY AND TORQUE SCREW WORKED ITS WAY LOSE INSITU. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES TO BE RETURNED. | PENDING EVALUATION. CONCOMITANT DEVICE(S): 320-15-01, (B)(4), EQ REV GLENOID PLATE. 320-01-42, (B)(4), EQUINOXE REVERSE 42MM GLENOSPHERE. 300-01-13, (B)(4), EQUINOXE, HUMERAL STEM PRIMARY. PRESS FIT 13MM. 320-15-05, (B)(4), EQ REV LOCKING SCREW. 320-20-00, (B)(4), EQ REVERSE TORQUE DEFINING SCREW KIT. 320-42-03, (B)(4). EQUINOXE REVERSE 42MM HUMERAL LINER +2.5. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00243 | 2022-03-09 | 2022-02-15 | I, n, j, u, r, y | Manufacturer report | TRULIANT TIB IMP PSC INSERT SZ 1.5, 15MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT TIB IMP PSC INSERT SZ 1.5, 15MM | 02-022-44-1515 |  | AS REPORTED, APPROXIMATELY TWO YEARS POST INITIAL RIGHT SIDE TKA, THIS FEMALE PATIENT WAS HAVING RIGHT KNEE PATELLA DISLOCATIONS. THE POLY INSERT WAS REMOVED AND THE FEMORAL PROSTHESIS WAS REMOVED. THE PATIENT RECEIVED A SIZE 1 RIGHT LOGIC CC FEMUR,12X80 CEMENTED STEM, AND A 13MM POLY INSERT. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. X-RAYS RECEIVED. THE DEVICES ARE NOT AVAILABLE FOR EVALUATION DUE TO THE HOSPITAL SENT OFF IMPLANTS WITH TISSUE. | CONCOMITANT MEDICAL PRODUCTS. TRULIANT PS CEM FEM PS CEM RIGHT SZ 1.5 (CAT# 02-020-11-0315 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00242 | 2022-03-09 | 2021-12-29 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 4 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 4 13MM | 02-012-35-4013 |  | AS REPORTED, APPROXIMATELY SIX YEARS POST INITIAL LEFT SIDE TKA, THIS MALE PATIENT HAD A REVISED POLY DURING PROCEDURE. PATIENT REASON FOR REVISION WAS TO RESURFACE THE PATELLA. THERE WAS EVIDENCE OF POLY WEAR. PATIENT WAS REVISED TO A TRULIANT POLY. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. SALES REP WAS UNABLE TO OBTAIN PHOTOS/X-RAYS. THE DEVICE IS AVAILABLE FOR EVALUATION. | CONCOMITANT MEDICAL PRODUCTS: LOGIC FEMORAL PS CEM LEFT SZ 4 (CAT# 02-010-01-0240 / SERIAL# (B)(4). LGC TIBIAL FIT TRAY CEM SZ 4F / 5T (CAT# 02-012-45-4050 / SERIAL# (B)(4). 1-4137 SRYKER SYS 6 90X13X1.19 (CAT# 203-96-43 / SERIAL# (B)(4). THREADED PIN SIZE 3.0 COLLARLESS 2PK (CAT# 521-78-32 / SERIAL3 (B)(4). THREADED PIN SIZE 3.0 COLLARLESS 2PK (CAT# 521-78-32 / SERIAL# (B)(4). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00253 | 2022-03-11 | 2022-01-13 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 38MM HUMERAL LINER +0 | EXACTECH, INC. | 320-38-00 | 320-38-00 |  | AS REPORTED, APPROXIMATELY 4 YEARS POST OP THE INITIAL LEFT TSA, THIS (B)(6) FEMALE PATIENT WAS REVISED DUE TO DISLOCATION. THE PATIENT WAS DISLOCATING, SURGEON UPSIZED COMPONENTS FROM 38MM TO 42MM. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETUNING DUE TO HOSPITAL POLICY. | PENDING EVALUATION. CONCOMITANT DEVICE(S): 320-01-38, (B)(4) - EQUINOXE REVERSE 38MM GLENOSPHERE AND 320-15-05, (B)(4) - EQ REV LOCKING SCREW. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00238 | 2022-03-11 | 2022-02-14 | I, n, j, u, r, y | Manufacturer report | NV GXL LNR, LIPPED, 36MM ID GROUP 4 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LNR, LIPPED, 36MM ID GROUP 4 CUPS | 132-36-54 |  | AS REPORTED, APPROXIMATELY ELEVEN YEARS POST INITIAL RIGHT SIDE THA, THIS MALE PATIENT WAS REVISED AS HE WAS HAVING HIP PAIN AND MRI SHOWED FLUID COLLECTION IN THE PELVIS FROM THE LINER WEAR. HEAD LINER EXCHANGE TO A NOVATION XLE 36MM LINER WAS PERFORMED. HIP WAS STABLE WITH RANGE OF MOTION (ROM) TESTING. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION. X-RAYS AND PHOTOS ATTACHED. DEVICES ARE NOT AVAILABLE FOR EVALUATION DUE TO HOSPITAL DOES NOT RELEASE THE IMPLANT TO THE REP. | CONCOMITANT MEDICAL PRODUCTS: BONE SCREW 6.5MM DIA X 25MM LONG (CAT# 120-65-25 / SERIAL# (B)(4)) . PF STEM TAPERED PLASMA SZ 16 (CAT# 160-00-16 / SERIAL# (B)(4)). BIOLOX DELTA FEMORAL HEAD 36MM OD, +0MM (CAT# 170-36-00 / SERIAL# (B)(4)). NV CROWN CUP CLSTR HOLE 60MM GROUP 4 (CAT# 180-01-60 / SERIAL# (B)(4)). PLATELET RICH PLASMA KIT WITH SPRAY TIPS (CAT# 620-00-02 / SERIAL# (B)(4)). ACCELERATE CONC. SYS REP BY 620-12-02 (CAT# 620-11-02 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00237 | 2022-03-11 | 2022-02-10 | I, n, j, u, r, y | Manufacturer report | COCR FEM HEAD 32MM +7 OFFSET 12/14 | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | COCR FEM HEAD 32MM +7 OFFSET 12/14 | 142-32-07 |  | AS REPORTED, APPROXIMATELY 3.5 YEARS POST INITIAL THA, THIS MALE PATIENT WAS BROUGHT BACK FOR RIGHT HIP OSTEOLYSIS. THE PATIENT WAS COMPLAINING OF PAIN, AND IT WAS EVIDENT IN THE X RAYS. THE FEMUR WAS DISLOCATED. THE FEMORAL HEAD REMOVED. THE POLY LINER REMOVED USING A BONE SCREW. THE ACETABULAR SCREW WAS REMOVED. THE CUP WAS FOUND TO BE WELL FIXED. A NEW XLE LINER WAS IMPLANTED, 36MM ID LIPPED. THE STEM WAS CHECKED FOUND TO BE LOOSE. IT WAS REMOVED AND A 19X245 MONOBLOC WAS IMPLANTED WITH A 36+3.5 BIOLOX HEAD. THERE WAS NO BREAKAGE OF DEVICE AND NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. X-RAYS ATTACHED. THE DEVICES WILL BE RETURN. | CONCOMITANT MEDICAL PRODUCTS NV GXL LINR, NTRL, 32MM ID, GROUP 2 CUPS (CAT# 130-32-52 / SERIAL# (B)(4)). ALTEON 6.5MM SCREW, 20MM (CAT# 180-65-20 / SERIAL# (B)(4)). ALT HA S CLR EXT SZ 6 (CAT# 190-31-06 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00254 | 2022-03-11 | 2022-02-21 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 38MM HUMERAL LINER +0 | EXACTECH, INC. | 320-38-00 | 320-38-00 |  | AS REPORTED, APPROXIMATELY 2 MONTH POST OP INITIAL TSA, THIS (B)(6) FEMALE PATIENT WAS REVISED DUE TO DISLOCATION. SURGEON UPSIZED THE GLENOSPHERE AND ADDED CONSTRAINED POLY. PATIENT HAD NO TUBEROSITES WHICH MAY HAVE CONTRIBUTED TO DISLOCATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETUNING DUE TO THE HOSPITAL DISPOSING OF THEM. | CONCOMITANT MEDICAL PRODUCTS: 320-15-01, 7074945 - EQ REV GLENOID PLATE. 320-01-38, 7131229 - EQUINOXE REVERSE 38MM GLENOSPHERE. 304-22-11, 6725412 - 10.5MM PLATFORM FX STEM RIGHT. 320-20-22, S297638 - EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 22MM. 321-20-00, 6984003 - EQUINOXE REVERSE SHOULDER DRILL KIT. 320-15-05, 7121475 - EQ REV LOCKING SCREW. 320-20-00, 7122557 - EQ REVERSE TORQUE DEFINING SCREW KIT. 320-10-00, 7052746 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00255 | 2022-03-11 | 2022-02-20 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 42MM HUMERAL LINER +0 | EXACTECH, INC. | 320-42-00 | 320-42-00 |  | AS REPORTED, APPROXIMATELY 2 YEARS POST OP THE INITIAL LEFT TSA, THIS (B)(6) MALE PATIENT'S WAS REVISED DUE TO AN INFECTION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICE WILL NOT BE RETUNING DUE TO THE HOSPITAL DISPOSING OF THEM. | PENDING EVALUATION: CONCOMITANT DEVICE(S): 320-01-42, (B)(4) - EQUINOXE REVERSE 42MM GLENOSPHERE, 300-01-15, (B)(4) - EQUINOXE, HUMERAL STEM PRIMARY, PRESS FIT 15MM, 320-20-00, (B)(4) - EQ REVERSE TORQUE DEFINING SCREW KIT, 320-15-01, (B)(4) - EQ REV GLENOID PLATE, 320-15-05, (B)(4) - EQ REV LOCKING SCREW, AND 320-10-00, (B)(4) - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00256 | 2022-03-11 | 2022-02-21 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE TRAY ADAPTER PLATE TRAY +5 | EXACTECH, INC. | 320-10-05 | 320-10-05 |  | AS REPORTED, THIS (B)(6) MALE PATIENT'S RIGHT SHOULDER WAS REVISED DUE DISLOCATION. THE LINER DISASSOCIATED FROM THE HUMERAL TRAY. SURGEON REMOVED THE PRESERVE STEM, LINER AND TRAY. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES ARE RETURNING. | PENDING EVALUATION : CONCOMITANT DEVICE(S): UNKNOWN - 46 HUMERAL LINER AND 300-30-06, (B)(4), EQUINOXE PRESERVE STEM 6MM. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00261 | 2022-03-15 | 2022-02-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRU DISTAL FEM CUT BLOCK | ORTHOPEDIC MANUAL SURGICAL INSTRUMENT | EXACTECH, INC. | TRU DISTAL FEM CUT BLOCK | 02-029-11-8000 |  | AS REPORTED, DURING SURGICAL USE, THE PIN BECAME LODGED IN THE CUTTING BLOCK AND BROKE OFF IN THE PIN REMOVER AS WELL. THERE WAS NO BREAKAGE OF A DEVICE. PATIENT REQUIRED PROLONGED HOSPITALIZATION AS A DIRECT RESULT OF THIS ISSUE. THE SALES REP WAS UNABLE TO OBTAIN PHOTOS/X-RAYS. THE DEVICE IS AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00260 | 2022-03-15 | 2022-02-01 | I, n, j, u, r, y | Manufacturer report | NV GXL LINER LIPPED 32MM ID, GROUP 1 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LINER LIPPED 32MM ID, GROUP 1 CUPS | 132-32-51 |  | AS REPORTED BY LEGAL TEAM, APPROXIMATELY FOUR YEARS POST THA, THIS FEMALE PATIENT NEED A REVISION OF HET HIP. PATIENT DOES NOT HAVE A SCHEDULED REVISION SURGERY YET. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00258 | 2022-03-15 | 2022-02-23 | I, n, j, u, r, y | Manufacturer report | TRULIANT TIB IMP CR INS STD SZ 5, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT TIB IMP CR INS STD SZ 5, 9MM | 02-022-47-5009 |  | AS REPORTED, APPROXIMATELY TWO MONTHS POST INITIAL TKA, THIS MALE PATIENT RIGHT KNEE WAS WASHED OUT FOR SUSPECTED INFECTION. CULTURES WERE SENT TO LAB. A NEW SIZE 5 9MM INSERT WAS PLACED. PATIENT WAS LAST KNOW TO BE IN STABLE CONDITION. PHOTOS RECEIVED. THE EXPLANT NOT AVAILABLE FOR RETURN. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00257 | 2022-03-15 | 2022-02-21 | I, n, j, u, r, y | Manufacturer report | EQUINOXE REVERSE 46MM HUMERAL LINER +0 | PROSTHESIS, SHOULDER, NON-CONSTRAINED, METAL/POLYMER CEMENTED | EXACTECH, INC. | 320-46-00 | 320-46-00 |  | IT WAS REPORTED THAT THIS (B)(6) MALE PATIENT'S RIGHT SHOULDER WAS REVISED DUE TO HIS HUMERAL LINER BECOMING DISASSOCIATED FROM THE HUMERAL TRAY SOMETIME AFTER PRIMARY SURGERY. SURGEON REMOVED THE STEM, TRAY, AND LINER. DATE OF REVISION IS UNKNOWN AND INITIAL IMPLANT DATE IS UNKNOWN. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL BE RETURNING. | PENDING EVALUATION. CONCOMITANT MEDICAL PRODUCTS: 320-10-00, 66378004 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0 15MM STEM - S/L 03113741. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00259 | 2022-03-15 | 2018-06-11 | I, n, j, u, r, y | Manufacturer report | NV GXL LINER LIPPED 32MM ID, GROUP 1 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LINER LIPPED 32MM ID, GROUP 1 CUPS | 132-32-51 |  | AS REPORTED BY LEGAL TEAM, APPROXIMATELY ONE MONTH POST INITIAL THA, THIS FEMALE PATIENT HAD A REVISION DUE TO INFECTION. NO ADDITIONAL INFORMATION AVAILABLE. | CONCOMITANT MEDICAL PRODUCTS: COCR FEM HEAD 32MM +7 OFFSET 12/14 (CAT# 142-32-07 / SERIAL# (B)(4) ) ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00262 | 2022-03-16 | 2022-02-25 | I, n, j, u, r, y | Manufacturer report | NV GXL LINER LIPPED 36MM ID GROUP 2 | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LINER LIPPED 36MM ID GROUP 2 | 132-36-52 |  | AS REPORTED, APPROXIMATELY 4 YEARS POST INITIAL LEFT THA, SURGEON REVISED A GX OUTLINER OF THIS (B)(6) MALE PATIENT THAT WAS HAVING HIP PAIN. PATIENT WAS REVISED TO AN NOVATION XLE 10 DEGREE FACE CHANGING LINER GROUP 2 36MM WITH A BIOLOX ADAPTER SLEEVE +3.5 AND BIOLOX OPTION FEMORAL HEAD 36MM. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. X-RAYS RECEIVED. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO IMPLANTS WERE IMMEDIATELY TAKEN FROM OR TO A WAITING THIRD-PARTY TO EXAMINE AND DETERMINE THE POLY AND HEAD. | CONCOMITANT MEDICAL PRODUCTS: BIOLOX DELTA FEMORAL HEAD 36MM +3.5 (CAT# 170-36-03 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00263 | 2022-03-17 | 2022-02-25 | I, n, j, u, r, y | Manufacturer report | ACETABULAR COMPOENTS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | ACETABULAR COMPOENTS | UNKNOWN |  | AS REPORTED, THE PATIENT VISITED SURGEON FOR POSSIBLE REVISION. FEMORAL HEAD MATERIAL IS COBALT CHROME. THE INDEX SURGERY WAS (B)(6) 2015, X-RAYS RECEIVED. NO ADDITIONAL INFORMATION RECEIVED. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00264 | 2022-03-17 | 2021-10-11 | I, n, j, u, r, y | Manufacturer report | LOGIC CR TIB INSERT SLOPE++, SZ 1, 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC CR TIB INSERT SLOPE++, SZ 1, 13MM | 02-012-49-1013 |  | AS REPORTED BY LEGAL BRIEF, APPROXIMATELY 5 YEARS POST LEFT TKA, THIS (B)(6) FEMALE PATIENT UNDERWENT LEFT TOTAL KNEE ARTHROPLASTY REVISION FOR PROTHESIS WEAR. THE PATIENT STARTED HAVING PAIN AND ASSOCIATED DYSFUNCTION APPROXIMATELY ABOUT A YEAR AGO AND WORKUP ENSUED AND WAS NEGATIVE FOR INFECTION THROUGH ASPIRIN AS WELL AS LAB WORK, THEN CT SCANS SHOWED SOME LYTIC LESIONS CONCERNING FOR PARTICLE DISEASE. PATIENT WAS REVISED TO COMPETITOR. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | CONCOMITANT MEDICAL PRODUCTS: LOGIC CR FEMORAL CEM, LEFT SZ 1 (CAT# 02-010-03-0210 / SERIAL# (B)(4)); LGC TIBIAL FIT TRAY CEM SZ 1F / 1T (CAT# 02-012-45-1010 / SERIAL# (B)(4)); THREE PEG PATELLA 29MM (CAT# 200-02-29 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00265 | 2022-03-17 | 2019-06-12 | I, n, j, u, r, y | Manufacturer report | LOGIC FEMORAL PS CEM RIGHT SZ 2.5 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC FEMORAL PS CEM RIGHT SZ 2.5 | 02-010-01-0325 |  | AS REPORTED BY THE LEGAL TEAM, APPROXIMATELY 6.5 YEARS POST INITIAL RIGHT TKA, THE HUSBAND OF THE (B)(6) FEMALE PATIENT REPORTED THAT THE PATIENT¿S RIGHT KNEE WAS INITIALLY IMPLANTED AND PATIENT SEEMED TO RECOVER WELL FROM THE PROCEDURE, EXCEPT THAT THE SCAR WAS LARGE. THE PATIENT STARTED EXPERIENCING PAIN AND SWELLING FOR ¿AWHILE¿ BEFORE SHE WAS RETURNED FOR A REVISION WITH COMPETITOR¿S DEVICES. IT WAS NOTICED DURING THE REVISION THAT THE PATIENT¿S FEMORAL COMPONENT WAS LOOSENING. THE WOUND WAS CULTURED FOR INFECTION AND CAME BACK NEGATIVE. THERE WAS NO BREAKAGE OF DEVICE AND NO SURGICAL DELAY/PROLONGATION. THE PATIENT EXPERIENCED A LONG, SEVERELY PAINFUL RECOVERY AND STILL IS NOT AMBULATORY, SHE USES A WHEELCHAIR OR A WALKER AND ¿SEVERE SCARRING¿. ¿SHE IS TERRIFIED OF THE THOUGHT OF HAVING ANOTHER SURGERY.¿ SALES REP WAS UNABLE TO OBTAIN PHOTOS/X-RAYS. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | CONCOMITANT MEDICAL PRODUCTS: LOGIC TIBIA IMPLANT PSC INSERT, SZ 2.5, 11MM (CAT# 02-012-44-2511 / SERIAL# (B)(4)) ,LGC TIBIAL FIT TRAY CEM SZ 2.5F / 2.5T (CAT# 02-012-45-2525 / SERIAL#(B)(4)) , AND THREE PEG PATELLA 32MM (CAT# 200-02-32 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00266 | 2022-03-17 | 2018-08-08 | I, n, j, u, r, y | Manufacturer report | TRULIANT TIB IMP PS INSERT SZ 4 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT TIB IMP PS INSERT SZ 4 9MM | 02-022-35-4009 |  | AS REPORTED BY THE EXACTECH TRULIANT KNEE CLINICAL STUDY, APPROXIMATELY 14 MONTHS POST INITIAL TKA, THIS (B)(6) MALE PATIENT HAD A REVISION OF TIBIAL INSERT DUE TO WORSENING PAIN, DETERIORATING RANGE OF MOTION, PROGRESSIVE DEFORMITY, AND SIGNIFICANT INTERFERENCE WITH FUNCTION THE ONLY OPTION FOR PAIN CONTROL AND FUNCTIONAL RESTORATION. PER CLINICAL STUDY THIS EVENT IS UNLIKELY RELATED TO THE DEVICE OR THE PROCEDURE. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00267 | 2022-03-17 | 2022-02-28 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 2 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 2 11MM | 02-012-35-2011 |  | AS REPORTED, THIS FEMALE PATIENT HAD A REVISION OF RECALLED POLY INSERT DUE TO THERE WAS ACCELERATED DELAMINATING OF POLY. THE PATIENT PRESENTED WITH AN EFFUSED AND PAINFUL KNEE. POLY REVISION REQUESTED BY SURGEON. THERE WERE PARTS/PIECES THAT FELL ON THE PATIENT WOUND SITE. SURGEON REMOVED ALL THE PARTS/PIECES FROM THE PATIENT. PATIENT WAS REVISED TO A NEW INSERT. THERE WAS NO SURGICAL DELAY/PROLONGATION. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO THE PATIENT WANTED TO KEEP IT. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00268 | 2022-03-17 | 2022-03-01 | I, n, j, u, r, y | Manufacturer report | LOGIC CR TIB INSERT STD, SZ 2.5, 9 MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC CR TIB INSERT STD, SZ 2.5, 9 MM | 02-012-47-2509 |  | AS REPORTED, APPROXIMATELY EIGHT MONTHS POST INITIAL RIGHT TKA, THIS (B)(6) FEMALE PATIENT HAD A REVISION DUE TO PROSTHESIS WEAR, PLUS ARTHROFIBROSIS. THERE WAS NO BREAKAGE TO THE DEVICE AND NO SURGICAL DELAY/PROLONGATION. SALES REP WAS UNABLE TO OBTAIN PHOTOS/X-RAYS. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICE WILL BE RETURN. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00270 | 2022-03-19 | 2022-02-14 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | HUM LNR IMPACT TIP 42M | EXACTECH, INC. |  | 321-07-42 |  | IT WAS REPORTED THAT DURING SURGICAL USE THE SURGEON IMPACTED THE POLY WHICH LED TO CRACKING OF THE 38 AND 42 IMPACTOR TIPS, THIS ALSO BENT THE IMPACTOR HANDLE. WHEN THE SURGEON WAS IMPACTING DEFINITIVE POLY INTO THE TRAY USING THE 38MM IMPACTOR TIP (321-07-38), THE TIP IMPACTOR FRACTURED. THE SURGEON IMPACTED WITH THE 42MM IMPACTOR TIP (321-07-42) AND THE IMPACTOR WAS FRACTURED AND THE HANDLE BENT (321-07-05). EVENT IS NOT ASSOCIATED WITH A REVISION OF EXACTECH IMPLANTS. PIECES HAD TO BE REMOVED FROM THE MALE PATIENT'S WOUND SITE. THERE WAS A 5-15 MIN SURGICAL DELAY, NO ADVERSE EVENTS DUE TO THIS PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICE TO BE RETURNED. | PENDING EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00269 | 2022-03-19 | 2022-02-14 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | HUM LNR IMPACT TIP 38M | EXACTECH, INC. |  | 321-07-38 |  | IT WAS REPORTED THAT DURING SURGICAL USE THE SURGEON IMPACTED THE POLY WHICH LED TO CRACKING OF THE 38 AND 42 IMPACTOR TIPS, THIS ALSO BENT THE IMPACTOR HANDLE. WHEN THE SURGEON WAS IMPACTING DEFINITIVE POLY INTO THE TRAY USING THE 38MM IMPACTOR TIP (321-07-38), THE TIP IMPACTOR FRACTURED. THE SURGEON IMPACTED WITH THE 42MM IMPACTOR TIP (321-07-42) AND THE IMPACTOR WAS FRACTURED AND THE HANDLE BENT (321-07-05). EVENT IS NOT ASSOCIATED WITH A REVISION OF EXACTECH IMPLANTS. PIECES HAD TO BE REMOVED FROM THE MALE PATIENT'S WOUND SITE. THERE WAS A 5-15 MIN SURGICAL DELAY, NO ADVERSE EVENTS DUE TO THIS PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICE TO BE RETURNED. | SECTION H10: (H3) BASED ON HISTORICAL COMPLAINT INVESTIGATIONS AND THE RETURNED DEVICES, THE FRACTURED HUMERAL LINER IMPACTOR TIPS REPORTED WERE LIKELY THE RESULT OF A STRESS RISER INTRODUCED DURING IMPACTION, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE. ADDITIONALLY, THE DEVICES MAY HAVE EXPERIENCED MATERIAL DEGRADATION IF STERILIZED BEYOND THE RECOMMENDED PARAMETERS LISTED IN THE REPROCESSING INSTRUCTION WHICH COULD HAVE LED TO THE OBSERVED FAILURE. | Y | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00271 | 2022-03-21 | 2021-12-16 | I, n, j, u, r, y | Manufacturer report | TRULIANT TIB IMP CRC INSERT SZ 3.5, 17MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT TIB IMP CRC INSERT SZ 3.5, 17MM | 02-022-51-3517 |  | AS REPORTED BY THE TRULIANT KNEE CLINICAL STUDY; APPROXIMATELY FOURTEEN DAYS POST INITIAL TKA, THIS (B)(6) FEMALE PATIENT CAME TO FOLLOW UP APPT WITH SWOLLEN KNEE. FLUID WAS DRAINED. PATIENT WAS THEN TRANSPORTED TO THE EMERGENCY DEPARTMENT FOR A SYNCOPAL EVENT. IRRIGATION AND DEBRIDEMENT WITH POLY EXCHANGE WAS PERFORMED A DAY AFTER. CULTURES WERE TAKEN INTRA-OPERATIVELY SHOWN POSITIVE FOR INFECTION. PATIENT STARTED ON 6 WEEKS ANTIBIOTICS. PER CLINICAL STUDY THIS EVENT IS NOT RELATED TO THE DEVICE BUT RELATED TO THE PROCEDURE. | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION RECEIVED THE FOLLOWING SECTIONS G3, G6, H2, H3, H6 AND H9 HAVE BEEN UPDATED ACCORDINGLY. (H3) AS REPORTED BY THE TRULIANT KNEE CLINICAL STUDY; APPROXIMATELY FOURTEEN DAYS POST INITIAL TKA, THIS 61 Y/O FEMALE PATIENT CAME TO FOLLOW UP APPT WITH SWOLLEN KNEE. FLUID WAS DRAINED. PATIENT WAS THEN TRANSPORTED TO THE EMERGENCY DEPARTMENT FOR A SYNCOPAL EVENT. IRRIGATION AND DEBRIDEMENT WITH POLY EXCHANGE WAS PERFORMED A DAY AFTER. CULTURES WERE TAKEN INTRA-OPERATIVELY SHOWN POSITIVE FOR INFECTION. PATIENT STARTED ON 6 WEEKS ANTIBIOTICS. THE EVENT IS NOT RELATED TO THE DEVICE BUT RELATED TO THE PROCEDURE. NO ADDITIONAL INFORMATION. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN ISSUES. THE CAUSE OF SURGICAL REVISION FOR INFECTION CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION. THESE DEVICES ARE USED FOR TREATMENT NOT DIAGNOSIS. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00273 | 2022-03-23 | 2022-02-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EQUINOXE | GLND SCR DEPTH GAGE | EXACTECH, INC. |  | 321-15-09 |  | IT WAS REPORTED THAT DURING SURGICAL USE, THE DEPTH GAGE COLOR WAS FADING. IT WAS NOTICED BY THE SURGEON WHEN DRILLING THE BASEPLATE SCREW HOLES FOR AN (B)(6) FEMALE PATIENT. THERE WAS NO DELAY. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICE TO BE RETURNED. | PENDING EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00272 | 2022-03-23 | 2022-02-21 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 42MM GLENOSPHERE | EXACTECH, INC. | 320-01-42 | 320-01-42 |  | AS REPORTED, SIX DAYS AFTER INITIAL LEFT TSA, THIS (B)(6) OLD MALE PATIENT WAS REVISED DUE TO THE POST OP SCAN REVEALING THAT THE SPHERE WAS NOT PROPERLY SEATED. SURGEON IMPLANTED THE GLENOSPHERE INCORRECTLY. THE SAME SIZED COMPONENTS WERE USED. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETUNING. | PENDING EVALUATION. CONCOMITANT MEDICAL PRODUCTS: 320-15-05, 7134580 - EQ REV LOCKING SCREW; 320-42-00, S269246 - EQUINOXE REVERSE 42MM HUMERAL LINER +0. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00274 | 2022-03-23 | 2022-02-16 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | 42MMMM LNR TRL +0 | EXACTECH, INC. |  | 321-42-00 |  | IT WAS REPORTED THAT DURING REMOVAL/WITHDRAWAL, THE 42MM TRIAL LINER WAS REMOVED FROM THE MALE PATIENT'S RIGHT SHOULDER AFTER TRAILING AND SOME RESIDUAL YELLOW PARTICLES WERE LEFT IN THE PATIENT. THE YELLOW RESIDUAL WAS ONLY FOUND ON THE ADAPTER TRAY. THE SURGEON USED A SMALL FORCEPS AND FLUID TO REMOVE THE PIECES. THE SURGICAL SITE WAS CLEAN CLOSE UP. THERE WAS NO SURGICAL DELAY. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICE IS RETURNING FOR EVALUATION. REVISION STATED IN EXPERIENCE REPORT WAS CAPTURED UNDER 1038671-2022-00200 / CASE- (B)(4). | DEVICE EVALUATED BY MFR: PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00276 | 2022-03-24 | 2019-12-09 | I, n, j, u, r, y | Manufacturer report | EXACTECH | EQUINOXE CAGE GLENOID L, POST AUG, LEFT | EXACTECH, INC. |  | 314-13-24 |  | IT WAS REPORTED FROM THE US THAT A PATIENT HAD ORIGINAL LEFT TOTAL SHOULDER REPLACEMENT PERFORMED ON (B)(6) 2019. THIS PATIENT WAS BROUGHT IN FOR REVISION SURGERY BECAUSE THERE WAS A SUSPECTED ROTATOR CUFF TEAR AND ALSO, THEY HAD LIMITED RANGE OF MOTION. THE SURGEON DISSECTED DOWN TO THE JOINT AND DISLOCATED THE SHOULDER. HE THEN USED THE HEAD "POPPER" TOOL TO GET THE HUMERAL HEAD OFF. HE THEN USED THE SCREWDRIVER TO REMOVE THE SCREW HOLDING ON THE REPLICATOR PLATE. HE ALSO THEN REMOVED THE REPLICATOR PLATE LEAVING THE STEM IN PLACE. HE THEN TURNED HIS ATTENTION TO THE GLENOID AND WHEN HE GOT IT PROPERLY EXPOSED THE PLASTIC PART OF THE CAGED GLENOID FELL RIGHT AWAY FROM THE PATIENTS GLENOID. TWO PERIPHERAL PEGS AND THE CENTER PEG REMAINED BEHIND IN THE PATIENT. THE SURGEON THEN ATTEMPTED TO REMOVE THESE PEGS WITH THE APPROPRIATE TOOL FROM THE TRAY. AFTER SOME TIME, HE WAS ABLE TO RETRIEVE THEM ALL. AFTER ALL OF THE PREVIOUS IMPLANTS, EXCEPT FOR THE STEM WERE REMOVED HE BEGAN WORKING TO PREPARE THE GLENOID FOR AN RSA. HE IMPLANTED AN 8-DEGREE POSTERIOR AUGMENTED PLATE WITH 4 SCREWS AND LOCKING CAPS. HE THEN ATTACHED THE 42MM GLENOSPHERE WITH LOCKING SCREW. FOLLOWING THIS THE SURGEON THEN TRIALED AND DECIDED TO GO WITH A +0-ADAPTER TRAY AND +0 42MM LINER. THE SURGEON IMPLANTED THE FINAL COMPONENTS AND BEGAN CLOSURE OF THE PATIENT. THE PATIENT LEFT THE OR STABLE AND IS EXPECTED TO HAVE A GOOD OUTCOME. THE DEVICES WERE NOT RECEIVED AS THE HOSPITAL DOES NOT ALLOW BIO-CONTAMINATED IMPLANTS TO LEAVE THE OR. NO OTHER INFORMATION HAS BEEN PROVIDED ABOUT THE EVENT BY THE CONTACTS. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS NOT AVAILABLE FOR ANALYSIS. ENGINEERING ANALYSIS COMPLETED ON (B)(6) 2020. DESIGN-RELATED ISSUES: THE DESIGN OF THIS EQUINOXE CAGE GLENOID HAS BEEN IN THE FIELD SINCE 2011. EXACTECH IS AWARE OF 22 COMPLAINT REPORTS INVOLVING A REVISION RELATED TO THE DISASSOCIATION OF A CAGE GLENOID SINCE 2011. THE SALES DATA FOR ALL CAGE GLENOID COMPONENTS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF RECEIVING ANY OTHER COMPLAINT REPORTS INVOLVING PARTS FROM THE CAGE GLENOID MANUFACTURING LOTS. THE MANUFACTURING RECORDS FOR THIS DEVICE ONLY EXIST IN PAPER FORM AND ARE STORED IN CARDBOARD BOXES AT IRON MOUNTAIN. DUE TO RISKS ASSOCIATED WITH COVID-19 TRANSMISSION, THE PHYSICAL MANUFACTURING RECORDS FOR THIS COMPLAINT WILL NOT BE REVIEWED. EXACTECH WILL CONTINUE TO MONITOR THIS COMPLAINT TYPE AND, IF WARRANTED, WILL ESCALATE ANY POSITIVE TRENDS TO THE CORRECTIONS AND REMOVAL COMMITTEE (CRC) FOR ADDITIONAL ACTIONS. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS EQUINOXE CAGE GLENOID, (B)(4), WAS REVIEWED. THE RISK IS CAPTURED IN LINES 18 AND 67. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR. THE REVISION REPORTED WAS LIKELY THE RESULT OF A ROTATOR CUFF TEAR. THE REASON FOR THE REPORTED PEG DISASSOCIATION COULD NOT BE DETERMINED BUT MAY HAVE BEEN THE RESULT OF DRILLING THE PERIPHERAL PEG HOLES AT AN ANGLE, NOT FULLY SEATING THE CAGE GLENOID COMPONENT AT THE TIME OF IMPLANTATION, ECCENTRIC LOADING OF THE GLENOID COMPONENT IF THE REPORTED ROTATOR CUFF TEAR RESULTED IN THE HUMERAL HEAD SHIFTING SUPERIORLY, A TRAUMATIC EVENT, AND/OR THE SURGEON IMPARTING A FORCE TO THE GLENOID POLYETHYLENE BODY DURING THE REVISION SURGERY. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE COMPONENT WAS NOT RETURNED FOR EVALUATION AND X-RAYS WERE NOT PROVIDED. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THE REPORTED PEG DISASSOCIATION WAS RESOLVED BY REMOVAL OF ALL PIECES BY THE SURGEON. AN INVESTIGATION WAS CONDUCTED: THE REVISION REPORTED WAS LIKELY THE RESULT OF A ROTATOR CUFF TEAR. THE REASON FOR THE REPORTED PEG DISASSOCIATION COULD NOT BE DETERMINED BUT MAY HAVE BEEN THE RESULT OF DRILLING THE PERIPHERAL PEG HOLES AT AN ANGLE, NOT FULLY SEATING THE CAGE GLENOID COMPONENT AT THE TIME OF IMPLANTATION, ECCENTRIC LOADING OF THE GLENOID COMPONENT IF THE REPORTED ROTATOR CUFF TEAR RESULTED IN THE HUMERAL HEAD SHIFTING SUPERIORLY, A TRAUMATIC EVENT, AND/OR THE SURGEON IMPARTING A FORCE TO THE GLENOID POLYETHYLENE BODY DURING THE REVISION SURGERY. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE COMPONENT WAS NOT RETURNED FOR EVALUATION AND X-RAYS WERE NOT PROVIDED. EQUINOX SQUARE TORQUE DEFINE SCREW DRIVE KIT (CAT# 300-20-02 / SN#:(B)(4)); EQUINOXE, HUMERAL HEAD SHORT, 53MM (BETA) (CAT# 310-01-53 / SN#:(B)(4)); - EQUINOXE REPLICATOR PLATE 1.5MM O/S (CAT# 300-10-15 / SN#:(B)(4)). | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00275 | 2022-03-24 | 2022-03-01 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | HUMERAL STEM PRIMARY, PRESS FIT 13MM | EXACTECH, INC. | 300-01-13 | 300-01-13 |  | AS REPORTED, APPROXIMATELY 6 YEARS POST OP INITIAL LEFT TSA, THIS (B)(6) MALE PATIENT WAS REVISED DUE TO THE STEM BEING SUBSIDED. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETUNING DUE TO HOSPITAL POLICY. | CONCOMITANT MEDICAL PRODUCTS: 320-01-42, 3854482 - EQUINOXE REVERSE 42MM GLENOSPHERE. 320-10-00, 4195950 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0. 320-15-05, 4181502 - EQ REV LOCKING SCREW. 320-20-00, 4213897 - EQ REVERSE TORQUE DEFINING SCREW KIT. 320-42-03, 3912095 - EQUINOXE REVERSE 42MM HUMERAL LINER +2.5. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
MW5108532 | 2022-03-25 | 2021-08-15 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | R3 52 MM ACETABULAR COMPONENT | PROSTHESIS, HIP, CONSTRAINED, CEMENTED OR UNCEMENTED, METAL/POLYMER | SMITH & NEPHEW, INC. |  |  |  | A TOTAL HIP REPLACEMENT (RIGHT HIP) WAS PERFORMED (B)(6) 2020 BY DR. (B)(6) PER THE OPERATIVE REPORT, THE PROSTHESIS WAS COMPRISED OF A "SMITH & NEPHEW R3 52 MM ACETABULAR COMPONENT WITH A 36 MM ID DELTA CERAMIC ACETABULAR LINER AND A MICROPORT ORTHOPEDICS SIZE 13R LONG VARUS NECK FEMORAL COMPONENT WITH A 36 SHORT DELTA CERAMIC FEMORAL HEAD COMPONENT." AROUND (B)(6) 2021 THERE WAS A SQUEAKING NOISE FROM THE CERAMIC-ON CERAMIC PROSTHESIS WHEN I WOULD PUT MY RIGHT HIP IN FULL FLEXION. I SPOKE TO MY PHYSICIAN AND PT ABOUT THE NOISE, AND THEY RECOMMENDED AN IN PERSON MEETING WITH THE SURGEON. I SAW THE SURGEON IN PERSON ON (B)(6) 2022. X-RAYS WERE PERFORMED THE MORNING OF THE EXAMINATION; X-RAYS DID NOT INDICATE ANY PROBLEM WITH THE PROSTHESIS. DR. (B)(6) ATTRIBUTES THE NOISE TO "EDGE LOADING"; HOWEVER, SINCE THE PROSTHESIS IS STILL IMPLANTED, THERE ISN'T ANY CONCLUSION AS TO WHY THERE WOULD BE EDGE LOADING EXCEPT THAT MY DEGREE OF FLEXION MAY BE BEYOND THE RANGE OF MOTION OF THE PROSTHESIS. HE COMMENTED THAT FREQUENT EDGE LOADING WAS LIKELY "LEVERING THE ARTIFICIAL JOINT" AND THAT PROBLEMS COULD RESULT FROM REPEATED LEVERING. HIS SUGGESTION WAS TO CONTINUE WITH ACTIVITY BUT TO BACK OFF ON ANY ACTIVITY THAT CAUSED THE SQUEAK/EDGE LOADING. FDA SAFETY REPORT ID # (B)(4). |  | N | Y |  |  | 2022-03
1038671-2022-00277 | 2022-03-25 | 2018-11-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | GXL NUETRAL LINER, G0 28MM ID | EXACTECH, INC. |  | 130-28-50 |  | IT WAS REPORTED FROM OUS THAT A DOCTOR HAD IDENTIFIED A CASE OF PREMATURE POLY WEAR IN A PATIENT THAT WAS IMPLANTED ON (B)(6) 2013. THE PATIENT DID NOT WANT TO HAVE A REVISION, AS OF (B)(6) 2021 IT IS UNKNOWN ABOUT THE PATIENT¿S CONDITION OR REVISION STATUS. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DESIGN-RELATED ISSUES: EXACTECH IS AWARE OF 61 COMPLAINT REPORTS INVOLVING HIP COMPONENTS DUE TO PROTHESIS WEAR (11 NON-REVISIONS AND 50 REVISIONS) SINCE 2008. THE FREQUENCY OF OCCURRENCE RANKING SCALE IS VERY LOW; THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT. ACCORDING TO THE SALES DATA, ALL PIECES OF THE 42-PIECE LOT HAVE BEEN SOLD. THE DEVICE HISTORY RECORD FOR THE ACETABULAR LINER WAS REVIEWED, AND ALL PARTS WERE ACCEPTED WITH CONFORMANCE TO THE DEVICE SPECIFICATIONS. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE PROTHESIS WEAR IS LISTED IN THE HIP IMPLANT PRODUCT LABELING, THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR. THE PROTHESIS WEAR REPORTED IN EXPERIENCE C2018-756 WAS LIKELY THE RESULT OF POLYETHYLENE WEAR ACCELERATED BY EDGE LOADING OR NECK IMPINGEMENT. HOWEVER, THE UNDERLYING CAUSE OF PROTHESIS WEAR CANNOT BE DETERMINED BECAUSE THE PATIENT WAS NOT REVISED, AND THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. AS OF (B)(6) 2021, IT IS UNKNOWN IF THE 84 YO PATIENT HAD A SURGICAL REVISION. AN INVESTIGATION WAS CONDUCTED; THE PROTHESIS WEAR REPORTED WAS LIKELY THE RESULT OF POLYETHYLENE WEAR ACCELERATED BY EDGE LOADING OR NECK IMPINGEMENT. HOWEVER, THE UNDERLYING CAUSE OF PROTHESIS WEAR CANNOT BE DETERMINED BECAUSE THE PATIENT WAS NOT REVISED, AND THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. WHEN THE COMPANY IS NOTIFIED OF PATIENT REVISION THE FILE WILL BE UPDATED AND ASSESSED FOR MDR SERIOUS INJURY. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00278 | 2022-03-25 | 2022-02-23 | I, n, j, u, r, y | Manufacturer report | M-SERIES STEM | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | M-SERIES STEM | UNK |  | AS REPORTED, THIS (B)(6) FEMALE PATIENT HAD BILATERAL HIP REPLACEMENTS IN 2008/2009. IN 2012 THE PATIENTS HIPS WERE REVISED AND WITH THE M-SERIES AND NOVATION CROWN CUP WITH GXL. THE RIGHT HIP BECAME VERY PAINFUL FOR THE PATIENT WHICH IS WHAT LED TO THIS REVISION. SURGEON REMOVED A LOT OF BLACK TISSUE, INDICATING CLASSIC SIGNS OF METALOSIS. HE WAS UNABLE TO DETERMINE THE ORIGIN OF THE METALOSIS AS THE INTERFACE BETWEEN THE STEM AND THE METAPHYSEAL COMPONENT WERE COLD WELDED. SURGEON ALSO MADE NOTE OF THE POLY WEAR AND SUBSEQUENT OSTEOLYSIS IN THE PELVIS. THE SURGEON DECIDED TO KEEP THE NOVATION CROWN CUP AND THEN IMPLANTED A NOVATION XLE POLY. THERE WAS NO BREAKAGE OF THE DEVICE AND NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THIS EVENT. IMAGES RECEIVED. THE DEVICES ARE NOT AVAILABLE FOR EVALUATION DUE TO HOSPITAL POLICY WILL NOT ALLOW FOR THE RETURN OF EXPLANTS. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00280 | 2022-03-25 | 2022-03-02 | I, n, j, u, r, y | Manufacturer report | STEM EXTENSION W/SLOT 120L X18 MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | STEM EXTENSION W/SLOT 120L X18 MM | 204-38-12 |  | AS REPORTED, THIS (B)(6) MALE PATIENT HAD A REVISION TKA. SURGEON KEPT TIBIAL IMPLANTS AND REVISED TO CCK TRULIANT. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. SALES REP WAS UNABLE TO OBTAIN PHOTOS/X-RAYS. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO HOSPITAL POLICY. | H3: AS REPORTED, THIS 52 Y/O OBESE MALE PATIENT HAD A REVISION TKA. SURGEON KEPT TIBIAL IMPLANTS AND REVISED TO CCK TRULIANT. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. SALES REP WAS UNABLE TO OBTAIN PHOTOS/X-RAYS. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO HOSPITAL POLICY. NO ADDITIONAL INFORMATION PROVIDED AFTER MULTIPLE ATTEMPTS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN OR MANUFACTURING ISSUES. THE CAUSE OF THE SUBSEQUENT REVISION CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION. THESE DEVICES ARE USED FOR TREATMENT NOT DIAGNOSIS. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00281 | 2022-03-25 | 2022-02-28 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 11MM | 02-012-44-4011 |  | AS REPORTED, APPROXIMATELY SEVEN YEARS POST INITIAL LEFT TKA, THIS (B)(6) MALE PATIENT HAD A SORE KNEE AND UNDER X-RAY IT WAS DEEMED TO BE OSTEOLYTIC ALL COMPONENTS WERE LOOSE. PATIENT HAD AN OSTEOLYTIC REACTION TO POLYETHENE WEAR. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. X-RAYS RECEIVED. THE DEVICE IS NOT AVAILABLE FOR EVALUATION AS THEY WERE DISPOSED BY HOSPITAL. | CONCOMITANT MEDICAL PRODUCTS: POSTERIAR STABILIZED CEMENTED SIZE 4 LEFT (CAT# 02-010-02-0240 / (B)(4)), LGC TIBIAL FIT TRAY CEM SZ 4F / 4T (CAT# 02-012-45-4040 / (B)(4)), THREE PEG PATELLA 32MM (CAT# 200-02-32 / (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00282 | 2022-03-25 | 2008-04-15 | I, n, j, u, r, y | Manufacturer report | CC FEMORAL SZ 4 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | CC FEMORAL SZ 4 | 208-01-04 |  | AS REPORTED BY THE EXACTECH KNEE REPLACEMENT STUDY, APPROXIMATELY THIRTEEN DAYS POST INITIAL TKA, THIS (B)(6) MALE EXPERIENCED RIGHT INCISION WOUND. PATIENT LOWER CALF CELLULITIS AND RIGHT INCISION WOUND INCREASED REDNESS. THE EVENT WAS RESOLVED WITH MEDICATION. THIS EVENT IS NOT RELATED TO THE DEVICE OR THE PROCEDURE. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00284 | 2022-03-25 | 2022-03-04 | I, n, j, u, r, y | Manufacturer report | LOGIC FEMORAL PS CEM LEFT SZ 3 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC FEMORAL PS CEM LEFT SZ 3 | 02-010-01-0230 |  | AS REPORTED, APPROXIMATELY SIX YEARS POST INITIAL LEFT TKA, THIS (B)(6) FEMALE PATIENT WAS HAVING PAIN AND DISCOMFORT FOR SOMETIME. SURGEON REVISED THE KNEE TO A COMPETITOR'S DEVICE. THE DIAGNOSIS WAS A NICKEL ALLERGY. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. SALES REP WAS NOT PRESENT FOR THE SURGERY AND UNABLE TO OBTAIN PHOTOS/X-RAYS. THE DEVICES ARE NOT AVAILABLE FOR EVALUATION AS HOSPITAL DID NOT KEEP THE DEVICES. | CONCOMITANT MEDICAL PRODUCTS: LOGIC TIBIA PS MOD INSRT SZ 3 9MM (CAT# 02-012-35-3009 / SERIAL# (B)(4)). LGC TIBIAL FIT TRAY CEM SZ 3F / 2T (CAT# 02-012-45-3020 / SERIAL# (B)(4)). THREE PEG PATELLA 32MM (CAT# 200-02-32 / SERIAL# (B)(4)). FLUTED STEM EXTENSION 25L X 14 MM (CAT# 204-34-02 / SERIAL# (B)(4)). TIBIAL STEM EXT. SCREW (CAT# 204-70-00 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00285 | 2022-03-26 | 2022-03-07 | I, n, j, u, r, y | Manufacturer report | LOGIC RBK INSERT SZ 3.5, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC RBK INSERT SZ 3.5, 9MM | 02-012-38-3509 |  | AS REPORTED, APPROXIMATELY TWO YEARS POST INITIAL TKA, THIS (B)(6) FEMALE PATIENT HAD A REVISION DUE TO STAPH INFECTION. THERE WAS NO RELATED TO BREAKAGE OF DEVICE AND NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. SALES REP WAS UNABLE TO OBTAIN PHOTOS/X-RAYS. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO DEVICES WERE DISPOSED BY HOSPITAL. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00283 | 2022-03-26 | 2022-03-03 | I, n, j, u, r, y | Manufacturer report | ALT HA S CLR EXT SZ 8 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | ALT HA S CLR EXT SZ 8 | 190-31-08 |  | AS REPORTED, APPROXIMATELY FIVE YEARS POST INITIAL LEFT THA, THIS MALE PATIENT WAS BROUGHT BACK FOR LEFT HIP PAIN. X RAYS SHOWED SIGNS OF POTENTIAL LOOSENING OF THE STEM. THE HEAD WAS REMOVED. THE STEM WAS CHECKED AND FOUND TO BE LOOSE. THE STEM WAS REMOVED. ATTENTION TAKEN TO THE ACETABULAR COMPONENT. THE POLY LINER WAS REMOVED WITH USE OF A 45MM 6.5 BONE SCREW. THE CUP WAS CHECKED AND FOUND TO BE WELL FIXED. A NEW G2 LIPPED 36MM XLE LINER WAS IMPLANTED. A 21X245 STEM WAS IMPLANTED WITH A 36+7 BIOLOX HEAD. THERE WERE NO PARTS OR PIECES FALL INTO THE PATIENT WOUND SITE. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. X-RAY RECEIVED. THE DEVICE WILL BE RETURN. | CONCOMITANT MEDICAL PRODUCTS: G2 NEUTRAL 32MM GXL LINER (CAT# 130-32-52 / SERIAL# (B)(4)), 32+3.5 COCR HEAD (CAT# 142-32-03 / SERIAL# (B)(4)), 20MM 6.5 BONE SCREW (CAT# 180-65-20 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00279 | 2022-03-26 | 2022-03-01 | I, n, j, u, r, y | Manufacturer report | TRULIANT PS CEM FEM PS CEM LEFT SZ 4.5 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT PS CEM FEM PS CEM LEFT SZ 4.5 | 02-020-11-0245 |  | AS REPORTED, APPROXIMATELY THREE YEARS POST INITIAL LEFT TKA, THIS (B)(6) MALE PATIENT COMPLAINED OF PAIN. LOOSE FEMUR. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PHOTOS RECEIVED. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO CHAIN OF CUSTODY FOR HOSPITAL AND PATIENT. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00297 | 2022-03-28 | 2020-03-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MED, LEFT, 16 POST AUG UHMWPE GLENOID | PROSTHESIS, SHOULDER, NON-CONSTRAINED, METAL/POLYMER CEMENTED | EXACTECH, INC. |  | 314-06-23 |  | APPROXIMATELY 4 YEAR(S), 1 MONTH(S) AND 26 DAY(S) POST-OPERATIVE OF A LEFT TSA, IT WAS REPORTED VIA CLINICAL STUDY THAT THE PATIENT EXPERIENCED ASEPTIC GLENOID LOOSENING. LOOSE GLENOID, NEEDS REVISION. ACCORDING TO THE INFORMATION PROVIDED, THE PATIENT HAS DEGENERATIVE ARTHRITIS IN THEIR LEFT GLENOHUMERAL JOINT WITH POSTERIOR EROSION. THE PATIENT UNDERWENT A STANDARD TOTAL REVISION, AND THE OUTCOME OF THIS EVENT IS NOW CONSIDERED RESOLVED. THE CLINICAL REPORT INDICATES THAT THIS EVENT IS DEFINITELY NOT RELATED TO THE DEVICE(S) AND DEFINITELY RELATED TO THE PROCEDURE. THIS WAS REPORTED THROUGH CLINICAL TRIAL STUDY DATA COLLECTION ACTIVITIES AND NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. | THIS FOLLOW-UP REPORT IS BEING SUBMITTED TO RELAY ADDITIONAL AND/OR CORRECTED INFORMATION. THE REASON FOR THE REVISION REPORTED CANNOT BE CONFIRMED FROM THE INFORMATION PROVIDED BUT MAY BE THE RESULT OF GLENOID LOOSENING. POTENTIAL CONTRIBUTIONS OF MANUFACTURING, USER, AND PATIENT-RELATED CONSIDERATIONS TO THE EVENT COULD NOT BE ASSESSED AS THE DEVICES WERE NOT AVAILABLE FOR EVALUATION AND IMAGES, RADIOGRAPHS, RELEVANT CLINICAL INFORMATION, AND PRODUCT INFORMATION WERE NOT PROVIDED. IF ANY FURTHER INFORMATION IS OBTAINED THAT WOULD CHANGE OR ALTER ANY INFORMATION PROVIDED, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00296 | 2022-03-28 | 2020-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | ALT HA S CLR EXT SZ 7 | EXACTECH, INC. |  | 190-31-07 |  | IT WAS REPORTED FROM THE US THAT A PATIENT HAS AN ALTEON HA IN WHICH LUCENCIES ARE VISIBLE AROUND THE FEMORAL STEM, INDICATING STEM LOOSENING. X-RAY IMAGES WERE PROVIDED. THE PATIENT HAS OCCASIONAL PAIN. THE STEM HAS NOT YET BEEN REVISED (WHEN REPORTED IN (B)(6) 2020). IT WAS REPORTED THAT THE SURGEON WAS PLANNING TO SEE THE PATIENT IN 4 MONTHS FOR TREATMENT PLANS ((B)(6) 2020). THE MASTER COMPLAINTS LIST WAS REVIEWED, AND THIS DEVICE HAS NOT BEEN REVISED AS OF 30 APRIL 2021. EBI WAS SEARCHED AND THERE ARE NO RESULTS FOR A REVISION FOR THIS PART NUMBER OR THE SURGEON. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. DEVICE REMAINS IMPLANTED. PAIN DUE TO FEMORAL STEM LOOSENING. THE COLLAR OF THE FEMORAL STEM IS CONTACTING BONE. IT IS LIKELY THAT DUE TO THE CONTACT OF THE COLLAR WITH THE CALCAR BONE, THE IMPLANT DID NOT OBTAIN A GOOD INITIAL ML MECHANICAL PRESS FIT AT THE TIME OF IMPLANTATION. THIS SITUATION CAN POTENTIALLY OCCUR SECONDARY TO OVER BROACHING OF THE FEMUR. THERE APPEARS TO BE A GAP BETWEEN THE IMPLANT AND BONE ON THE PROXIMAL LATERAL STEM. PER DISCUSSION WITH THE HIP TEAM, PATIENT MAY HAVE INCREASED THEIR ACTIVITY LEVEL BEFORE OSSEOINTEGRATION HAD TIME TO OCCUR AND MICROMOTION FROM REPEATED EARLY LOAD BEARING RESULTED IN COMPACTION OF THE BONE ON THE MEDIAL SIDE OF THE FEMORAL CANAL AND SUBSEQUENT GAPPING ON THE LATERAL SIDE, HOWEVER THIS CANNOT BE CONFIRMED AS NO ADDITIONAL DETAILS WERE PROVIDED. ADDITIONALLY, OVER-BROACHING THE FEMORAL CANAL TO OBTAIN CONTACT BETWEEN THE COLLAR OF THE FEMORAL STEM AND CALCAR BONE CAN DIMINISH THE QUALITY OF THE DISTAL PRESS FIT BETWEEN THE IMPLANT AND BONE. DESIGN-RELATED ISSUES: AS OF 4/13/2020, EXACTECH IS AWARE OF 42 SIMILAR COMPLAINT REPORTS INVOLVING REVISION OF HIP COMPONENTS DUE TO LOOSENING SINCE 2008. THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 16 UNITS, WHICH HAS BEEN IN THE FIELD SINCE 2017. THE MANUFACTURING RECORDS FOR THIS DEVICE ONLY EXIST IN PAPER FORM AND ARE STORED IN CARDBOARD BOXES AT (B)(6). DUE TO RISKS ASSOCIATED WITH COVID-19 TRANSMISSION, THE PHYSICAL MANUFACTURING RECORDS FOR THIS COMPLAINT WILL NOT BE REVIEWED. EXACTECH WILL CONTINUE TO MONITOR THIS COMPLAINT TYPE AND, IF WARRANTED, WILL ESCALATE ANY POSITIVE TRENDS TO THE CORRECTIONS AND REMOVAL COMMITTEE (CRC) FOR ADDITIONAL ACTIONS. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE LOOSENING IS LISTED IN THE PRODUCT LABELING, THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RACR. THE LOOSENING REPORTED IS LIKELY THE RESULT OF FAILURE TO OBTAIN INITIAL OSSEOINTEGRATION BETWEEN THE BONE AND FEMORAL STEM. THIS MAY HAVE BEEN A COMBINATION OF SURGICAL TECHNIQUE, PATIENT REHABILITATION, OR OTHER CONTRIBUTING FACTORS; HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE IS NOT AVAILABLE FOR EVALUATION AND NO OTHER FIRST-HAND DETAILS WERE PROVIDED. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. ¿ CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. DEVICE SPECIFIC RISKS - FRACTURE, MIGRATION, LOOSENING, SUBLUXATION, OR DISLOCATION OF THE PROSTHESIS OR ANY OF ITS COMPONENTS, ANY OF WHICH MAY REQUIRE A SECOND SURGICAL INTERVENTION OR REVISION. EXCESSIVE WEAR OF THE IMPLANT COMPONENTS SECONDARY TO IMPINGEMENT OF COMPONENTS OR DAMAGE OF ARTICULAR SURFACES. METAL SENSITIVITY REACTIONS OR OTHER ALLERGIC/HISTOLOGICAL REACTIONS TO IMPLANT MATERIALS. POSSIBLE DETACHMENT OF THE COATING(S) ON THE COMPONENTS, POTENTIALLY LEADING TO INCREASED DEBRIS PARTICLES. ADVERSE EVENTS ASSOCIATED WITH THE USE OF BONE CEMENT. DISASSOCIATION OF MODULAR COMPONENTS. OSTEOLYSIS. UNINTENDED BONE FRACTURES. LIMB LENGTH DISCREPANCY. DECREASED RANGE OF MOTION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THE PATIENT HAS PAIN AND NOTED FEMORAL LOOSENING, THERE HAS NOT BEEN A REPORTED REVISION, NOR HAS A REVISION BEEN IDENTIFIED WITH EBI AND MASTER COMPLAINT RESEARCH. DEVICE SPECIFIC RISKS - FRACTURE, MIGRATION, LOOSENING, SUBLUXATION, OR DISLOCATION OF THE PROSTHESIS OR ANY OF ITS COMPONENTS, ANY OF WHICH MAY REQUIRE A SECOND SURGICAL INTERVENTION OR REVISION. AN INVESTIGATION WAS CONDUCTED; THE LOOSENING REPORTED IS LIKELY THE RESULT OF FAILURE TO OBTAIN INITIAL OSSEOINTEGRATION BETWEEN THE BONE AND FEMORAL STEM. THIS MAY HAVE BEEN A COMBINATION OF SURGICAL TECHNIQUE, PATIENT REHABILITATION, OR OTHER CONTRIBUTING FACTORS; HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE IS NOT AVAILABLE FOR EVALUATION AND NO OTHER FIRST-HAND DETAILS WERE PROVIDED. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00295 | 2022-03-28 | 2020-02-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | PRM STM INSRTR EXTRCT | EXACTECH, INC. |  | 301-07-10 |  | IT WAS REPORTED FROM OUS THAT WHEN A LOAN KIT WAS RETURNED IT WAS NOTED THAT THE INSERTER TIP HAD BEEN DAMAGED ON THE THREADS. ENGINEERING FINDINGS: THE RETURNED DEVICE HAS TOOL MARKS THAT APPEAR CONSISTENT WITH NORMAL USE. THE BLACK PLASTIC SPHERE ON THE TIP OF THE IMPACTOR IS FRACTURED AT THE START OF THE THREADED SHAFT. THE METAL KNOB THAT MATES WITH THE HUMERAL STEM WAS NOT RETURNED WITH THE DEVICE. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED FOR ANALYSIS. FINDINGS: THE RETURNED DEVICE HAS TOOL MARKS THAT APPEAR CONSISTENT WITH NORMAL USE. THE BLACK PLASTIC SPHERE ON THE TIP OF THE IMPACTOR IS FRACTURED AT THE START OF THE THREADED SHAFT. THE METAL KNOB THAT MATES WITH THE HUMERAL STEM WAS NOT RETURNED WITH THE DEVICE. A REVIEW OF THE RISK MANAGEMENT FILES WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RISK ASSESSMENT AND CONTROLS REPORT (RACR) FOR THIS PRIMARY STEM INSERTER/EXTRACTOR, RACR-00009 REV- WAS REVIEWED. THE HAZARDOUS SITUATION IS CAPTURED; HOWEVER, THE RISK IS BEING UPDATED UNDER DCRSP-19-541 TO REFLECT A LESSER HARM OF ¿NONE¿. AS PREVIOUSLY INVESTIGATED, THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF A STRESS RISER WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE SECONDARY TO YEARS OF NORMAL SURGICAL USE FOR IMPACTION AND SUBSEQUENT REPROCESSING AND STERILIZATION CYCLES. IFU 700-096-181: INSTRUMENT INSPECTION VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE WAS NO REPORTED PATIENT INVOLVEMENT. AS PREVIOUSLY INVESTIGATED, THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF A STRESS RISER WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE SECONDARY TO YEARS OF NORMAL SURGICAL USE FOR IMPACTION AND SUBSEQUENT REPROCESSING AND STERILIZATION CYCLES. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00294 | 2022-03-28 | 2020-02-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRU TIB INSERT TRIAL TOP PS, SZ 1.5 | EXACTECH, INC. | TRU TIB INSERT TRIAL TOP PS, SZ 1.5 | 02-029-25-3515 |  | IT WAS REPORTED FROM THE US THAT WHILE PUTTING INSTRUMENTS BACK TOGETHER FOLLOWING A CASE IT WAS REALIZED THE SPRING INSIDE THE TRIAL WAS MISSING. THE AGENT CONFIRMED THAT NO PIECES WERE LEFT IN THE PATIENT. THE AGENCY IS INACTIVE, NO FURTHER INFORMATION IS AVAILABLE. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE RECEIVED FOR ANALYSIS. VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR 10X OPTICAL UNLESS OTHERWISE SPECIFIED)- UPON VISUAL INSPECTION, THE BAL-SEAL SPRING APPEARS TO BE MISSING FROM THE RETURNED DEVICE. THIS IS CONSISTENT WITH THE REPORTED EVENT. THE MISSING SUBCOMPONENT IS LIKELY THE RESULT OF THE SPRING BECOMING DEFORMED, EITHER DURING A PRIOR USE OR TO REMOVE DURING REPROCESSING, WHICH MAY HAVE ALLOWED THE SPRING TO DISASSEMBLE DURING THE CASE AND SUBSEQUENTLY BECOME MISPLACED. THE SALES REPRESENTATIVE CONFIRMED THAT NO PIECES WERE LEFT IN THE PATIENT. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE AND CONSISTENT WITH SURGICAL USE. DURING A PREVIOUS COMPLAINT INVESTIGATION, A CORRECTION AND REMOVAL COMMITTEE MEETING WAS HELD TO DETERMINE THE NEED FOR FIELD ACTION REGARDING MISSING BAL-SEAL SPRINGS FROM THE TRULIANT TIBIAL INSERT TRIAL TOPPERS. THE COMMITTEE DETERMINED THAT THERE WAS INSUFFICIENT EVIDENCE TO TAKE FIELD ACTION AT THAT TIME. REFER TO (B)(4). DESIGN-RELATED ISSUES: THE DESIGN OF THIS TRULIANT PS TIBIAL INSERT TRIAL TOPPER HAS BEEN IN THE FIELD SINCE 2017. EXACTECH IS AWARE OF 3 COMPLAINT REPORTS INVOLVING 5 DISASSEMBLED BAL-SEAL SPRINGS FROM THE TIBIAL INSERT TRIAL TOPPERS SINCE 2017. THE SALES DATA FOR ALL TRULIANT TIBIAL INSERTS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THE DESIGN OF THE TOPPER HAS BEEN UPDATED TO INCLUDE A MORE ROBUST BAL-SEAL SPRING. REFER TO (B)(4). AS DOCUMENTED IN (B)(4), PREVIOUS INVESTIGATIONS FOUND THAT THERE WAS A MAJOR RISK OCCURRING WHEN A CANNULA-CLEANING BRISTLE BRUSH IS USED TO CLEAN THE HOLE WHERE THE BAL-SEAL SPRING IS LOCATED. ADDITIONALLY, THE TEST SUMMARY SUMMARIZES THE ASSESSMENT RESULTS OF THE NEW BAL-SEAL SPRING. IT WAS FOUND TO DEMONSTRATE IMPROVED RESISTANCE TO FAILURE COMPARED TO THE OLD BAL-SEAL SPRING AND THE INCREASED REMOVAL FORCE WITH THE MATING DEVICE WAS DEEMED ACCEPTABLE BY SURGEON USERS. THE COMPLAINT DEVICE WAS MANUFACTURED PRIOR TO THESE DESIGN CHANGES. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF (B)(4) UNITS, WHICH HAVE BEEN IN THE FIELD SINCE 2018. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING-RELATED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, THE RISK IS CAPTURED IN THE RMR AND RACR. THE MISSING SUBCOMPONENT REPORTED WAS LIKELY THE RESULT OF THE SPRING BECOMING DEFORMED, EITHER DURING A PRIOR USE OR IN AN ATTEMPT TO REMOVE DURING REPROCESSING, WHICH MAY HAVE ALLOWED THE SPRING TO DISASSEMBLE DURING THE CASE AND SUBSEQUENTLY BECOME MISPLACED. THE DESIGN OF THE BAL-SEAL SPRING HAS BEEN UPDATED. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. THE MISSING SUBCOMPONENT WAS VERIFIED BY THE AGENT AS NOT BEING LEFT IN THE PATIENT. AN INVESTIGATION WAS CONDUCTED; THE MISSING SUBCOMPONENT REPORTED WAS LIKELY THE RESULT OF THE SPRING BECOMING DEFORMED, EITHER DURING A PRIOR USE OR IN AN ATTEMPT TO REMOVE DURING REPROCESSING, WHICH MAY HAVE ALLOWED THE SPRING TO DISASSEMBLE DURING THE CASE AND SUBSEQUENTLY BECOME MISPLACED. THE DESIGN OF THE BAL-SEAL SPRING HAS BEEN UPDATED. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00293 | 2022-03-28 | 2020-01-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | HEAD PUSHER, MOD, LOWPRO | EXACTECH, INC. |  | 01-001-03-0001 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE END OF THE HIP FEMORAL HEAD IMPACTOR HAS CRACKED (BROKEN). THERE WAS NOTHING LEFT IN THE WOUND. THE AGENCY IS INACTIVE, NO FURTHER INFORMATION IS AVAILABLE. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICES RECEIVED FOR ANALYSIS. VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR 10X OPTICAL UNLESS OTHERWISE SPECIFIED) THE FRACTURED SURFACES OF THE IMPACTOR TIP APPEAR CONSISTENT WITH CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE BLACK PLASTIC COMPONENT. THE CRACK LIKELY INITIATED AT THE TIP OF THE SCREW THREADS. ONE OF THE FEMORAL HEAD IMPACTORS WAS RETURNED WITHOUT THE IMPACTOR TIP. ALL OTHER ASPECTS OF THE DEVICES APPEAR UNREMARKABLE AND CONSISTENT WITH YEARS OF USE. PER THE EXPERIENCE REPORT, ALL PIECES WERE RECOVERED FROM BOTH BROKEN IMPACTORS. DESIGN-RELATED ISSUES: THE DESIGN OF THIS FEMORAL HEAD PUSHER, LOW PROFILE/MODIFIED FEMORAL HEAD IMPACTOR HAS BEEN IN THE FIELD SINCE 2013. EXACTECH IS AWARE OF 6 TOTAL COMPLAINTS INVOLVING 9 FEMORAL HEAD PUSHERS SINCE 2013. THE ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF (B)(4) UNITS (ASIDE FROM THE 2 REPORTED FOR THIS CASE), WHICH HAVE BEEN IN THE FIELD SINCE 2015. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR. THE BROKEN DEVICES REPORTED WERE LIKELY THE RESULT OF STRESS RISERS INTRODUCED DURING IMPACTION WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICES. IFU 700-096-181: INSTRUMENT INSPECTION, VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. IT WAS REPORTED THAT ¿THERE WAS NOTHING LEFT IN THE WOUND¿, NO PATIENT ADVERSE EVENT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN DEVICES WERE LIKELY THE RESULT OF STRESS RISERS INTRODUCED DURING IMPACTION WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICES. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00292 | 2022-03-28 | 2020-01-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | FIT TRAY ADJ, TIBIAL TAMP HEAD, NET SHAPE | EXACTECH, INC. |  | 213-75-01 |  | IT WAS REPORTED FROM OUS THAT THE ¿TAMP HEAD BROKEN AFTER SURGERY¿. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: (B)(4). FINDINGS: PER (B)(4), IT WAS DETERMINED THAT A POTENTIAL CAUSE OF THIS FAILURE IS RELATED TO MISUSE OF THE DEVICE BY RETROGRADE IMPACTION. THE DEVICE RETURNED APPEARS TO HAVE A BRITTLE FRACTURE OF THE CONNECTION FEATURE OF THE TAMP HEAD. CORRECTIVE ACTION TAKEN: PER (B)(4), THE DESIGN OF THE TAMP HEAD WAS MODIFIED TO ALTER THE RADIUS BETWEEN THE SHAFT AND THE COLLAR HEAD IN ORDER TO REDUCE THE STRESS CONCENTRATION ON THE UNDERSIDE OF THE COLLAR LIP. VALIDATION TESTING WAS PERFORMED TO CONFIRM NO INTERFERENCE WITH THE ABILITY OF THE INSTRUMENT MATING FUNCTION WITH THE NEW KEEL HEAD DESIGN. SEE (B)(4). ROOT CAUSE TESTING WAS PERFORMED PER (B)(4) ON THE ORIGINAL DESIGN (REV A) AND THE CURRENT DESIGN (REV B). THE CURRENT DESIGN PASSED TESTING FOR RETROGRADE IMPACTION WHERE THE ORIGINAL DESIGN FAILED. ALL SURGICAL TECHNIQUES IN WHICH THE TAMP HEAD IS USED WERE UPDATED TO INCLUDE A PRECAUTION REGARDING RETROGRADE IMPACTION. A FIELD NOTIFICATION WAS ISSUED ON PROPER USE AND POTENTIAL MISUSE OF THE INSTRUMENTATION. THE RISK MANAGEMENT REPORT, (B)(4) INSTRUMENTS, WAS UPDATED TO INCLUDE RISKS ASSOCIATED WITH FRACTURE OF THE MODULAR PORTION OF THE FIT TAMP HEAD. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: THE ROOT CAUSE OF THE TIBIAL TAMP HEAD FRACTURES WAS DETERMINED TO BE A COMBINATION OF THE DESIGN AND USER MISUSE. IFU (B)(4): INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. THE ¿TAMP HEAD BROKEN AFTER SURGERY¿, THERE WAS NO REPORTED PATIENT ADVERSE EVENT. AN INVESTIGATION WAS CONDUCTED; THE ROOT CAUSE OF THE TIBIAL TAMP HEAD FRACTURES PER (B)(4) WAS DETERMINED TO BE A COMBINATION OF THE DESIGN AND USER MISUSE. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00298 | 2022-03-28 | 2020-03-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | SPECIAL IMPACTOR TIP | EXACTECH, INC. |  | SK-025605 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THAT WHILE IMPACTING THE TALAR COMPONENT, THE IMPACTOR PLASTIC BROKE IN HALF. NONE OF THE PIECES FELL INTO THE SURGICAL SITE. THIS RESULTED IN JUST A FEW SECONDS OF PROLONGATION AS WE HAD TO SET UP THE STANDARD TALLER IMPACT TO USE. THERE WAS NO INJURY /CLINICAL CONSEQUENCE/ADVERSE EVENT TO THE PATIENT. THE PATIENT HAD NO ISSUES AND FOLLOWED THE SURGEON¿S GUIDELINES FOR TOTAL ANKLE PATIENTS. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE AVAILABLE FOR EVALUATION. VISUAL EVALUATION CONDITION/FINDINGS: THE FRACTURED SURFACES OF THE RETURNED IMPACTOR TIP APPEAR CONSISTENT WITH A BRITTLE FRACTURE OF THE PLASTIC. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE AND CONSISTENT WITH YEARS OF USE. DESIGN-RELATED ISSUES: THE DESIGN OF THE SPECIAL TALUS IMPACTOR SK-025623, OF WHICH THIS IS A SUBCOMPONENT, HAS BEEN IN THE FIELD SINCE 2018. REFER TO TM-2018-1014 AND TM-2019-0600 FOR DETAILS. EXACTECH IS NOT AWARE OF ANY SIMILAR COMPLAINT REPORTS INVOLVING BROKEN IMPACTOR TIPS SINCE THE DEVICE WAS RELEASED. BASED ON THE MOST-LIKELY UNDERLYING CAUSE, THE ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 3 INITIAL SPECIAL TALUS IMPACTOR OR 3 SPECIAL IMPACTOR TIPS, WHICH HAVE BEEN IN THE FIELD SINCE 2018 AND 2019 RESPECTIVELY. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF TM-2019-0600 WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED. THE RISK THAT THE ¿RADEL COMPONENT DOES NOT REMAIN FUNCTIONAL AFTER USE¿ IS INCLUDED FOR THIS DEVICE. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE RISK IS INCLUDED IN THE TECHNICAL MEMO. MOST LIKELY CAUSE: THE BROKEN SPECIAL IMPACTOR TIP REPORTED IS MOST LIKELY THE RESULT OF A STRESS RISER INTRODUCED DURING IMPACTION WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. NONE OF THE PIECES FROM THE BROKEN DEVICE FELL INTO THE SURGICAL SITE. THERE WERE A FEW SECONDS OF SURGICAL PROLONGATION AS THE SET THE STANDARD TALLER IMPACT TO USE. THERE WAS NO INJURY /CLINICAL CONSEQUENCE/ADVERSE EVENT TO THE PATIENT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN SPECIAL IMPACTOR TIP REPORTED IS MOST LIKELY THE RESULT OF A STRESS RISER INTRODUCED DURING IMPACTION WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00291 | 2022-03-28 | 2019-12-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | QUICK IMPACTOR RADEL BARNARD | EXACTECH, INC. |  | CSK-8439 |  | IT WAS REPORTED FROM THE US THAT THE IMPACTOR TIPS WORK WELL BUT EVENTUALLY BREAK UPON IMPACT. THE AGENT BELIEVE THAT THIS IS DUE TO AUTOCLAVING THE INSTRUMENTS AND THE HEAT WEAKENING THEM OVERTIME. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE(S) NOT AVAILABLE FOR EVALUATION. HOSPITAL-SPECIFIC REPROCESSING INFORMATION WAS REQUESTED FROM THE SALES REPRESENTATIVE; HOWEVER, HE WAS UNABLE TO OBTAIN DETAILS REGARDING DETERGENTS, TEMPERATURE, AND TIME. THEREFORE, CONFIRMATION OF ADHERENCE TO IFU 700-096-124 REV D, REPROCESSING INSTRUCTIONS FOR REUSABLE SURGICAL INSTRUMENTATION, COULD NOT BE CONFIRMED. DESIGN-RELATED ISSUES: AS OF 6/24/2020, EXACTECH IS NOT AWARE OF ANY PREVIOUS COMPLAINT REPORTS ON THIS SPECIAL INSTRUMENT; HOWEVER, THERE ARE 57 TOTAL COMPLAINTS INVOLVING FRACTURED IMPACTOR TIPS FOR STANDARD SHOULDER INSTRUMENTATION SINCE 2011. EXACTECH IS AWARE OF 11 TOTAL COMPLAINTS INVOLVING A FRACTURED RADEL® SPHERE ON THE STANDARD INSTRUMENTATION FOR INSERTING THE HUMERAL STEM, THE PRIMARY STEM INSERTER/EXTRACTOR (CN: 301-07-10). SINCE IMPACTOR TIPS AND THE PRIMARY STEM INSERTER/EXTRACTOR ARE USED DURING BOTH ANATOMIC AND REVERSE SHOULDER ARTHROPLASTY SURGERIES, SALES DATA FOR ALL HUMERAL STEMS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: A TOTAL OF 29 SPECIAL BARNARD IMPACTOR TIPS WERE PRODUCED IN THREE SEPARATE LOTS (9 IN 2011 UNDER TM-2011-060, 10 IN 2012 UNDER TM-2012-521, AND 10 IN 2017 UNDER TM-2017-0131 AND DSR217046). THIS COMPLAINT INVOLVES MULTIPLE PARTS, HOWEVER, THE RELATIONSHIP TO A SPECIFIC MANUFACTURING LOT HAS NOT BEEN ESTABLISHED AS THE DSR/LOT INFORMATION WAS NOT PROVIDED. THEREFORE, THE CONTRIBUTION OF MANUFACTURING RELATED ISSUES TO THE REPORTED FAILURE CANNOT BE ASSESSED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE TECHNICAL MEMOS. THE MOST LIKELY UNDERLYING CAUSE FOR THE FRACTURED SPECIAL QUICK IMPACTOR REPORTED IS A STRESS-RISER BEING INTRODUCED DURING IMPACTION, SUBSEQUENTLY LEADING TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE. THE CONTRIBUTIONS OF REPROCESSING TO THE MATERIAL FAILURE CANNOT BE DETERMINED BECAUSE NO DEVICES WERE RETURNED FOR EVALUATION AND REPROCESSING INFORMATION WAS NOT PROVIDED. IFU 700-096-181: INSTRUMENT INSPECTION :VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY DESIGN ISSUES. THERE IS NO REPORTED PATIENT IMPACT. NO SPECIFIC CASE IS MENTIONED BY THE AGENT. AN INVESTIGATION WAS CONDUCTED; THE MOST LIKELY UNDERLYING CAUSE FOR THE FRACTURED SPECIAL QUICK IMPACTOR REPORTED IS A STRESS-RISER BEING INTRODUCED DURING IMPACTION, SUBSEQUENTLY LEADING TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE. THE CONTRIBUTIONS OF REPROCESSING TO THE MATERIAL FAILURE CANNOT BE DETERMINED BECAUSE NO DEVICES WERE RETURNED FOR EVALUATION AND REPROCESSING INFORMATION WAS NOT PROVIDED. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00290 | 2022-03-28 | 2020-01-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRUL FEM TRIAL EXTRACTOR | EXACTECH, INC. |  | 02-029-19-1200 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE FEMORAL TRIAL EXTRACTOR BROKE. IT WAS STATED THAT NO PIECES FELL INTO THE SURGICAL SITE. NO DELAY TO SURGERY AND NO ADVERSE EVENT TO PATIENT. THERE WERE NO NOTED PROBLEMS FOR THE PATIENT IN THE OVERALL PROCEDURE OUTCOME/RESULT. IN CASE OF MULTI-PIECE CONSTRUCTION, IT WAS FOUND THAT THE WELD ON THE EXTRACTOR BODY WAS BREAKING. THIS ALLOWED THE COMPONENT PARTS TO DISASSEMBLE. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | H11. AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THIS MEDWATCH IS BEING RESCINDED: UPON FURTHER REVIEW OF REPORTABILITY- IF THE MALFUNCTION WERE TO RECUR, IT IS NOT LIKELY TO CAUSE OR CONTRIBUTE TO SERIOUS INJURY OR DEATH. THE DISASSEMBLED PARTS ARE NOT LIKELY TO FALL INTO THE WOUND SITE OR INTERFERE WITH THE SURGICAL IMPLANT PROCEDURE. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00289 | 2022-03-28 | 2013-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | UHMWPE POST AUG GLENOID LARGE, RIGHT | EXACTECH, INC. |  | 314-02-34 |  | IT WAS REPORTED FROM THE EQUINOXE SHOULDER STUDY THAT THE PATIENT DEVELOPED A LOOSE GLENOID IMPLANT NOTED ON X-RAY AT YEAR 2 POST-OP (2013) SUBJECT DID NOT BECOME SYMPTOMATIC UNTIL AROUND YEAR 7. THE PATIENT HAS BEEN ADVISED TO HAVE A REVISION. NO REVISION REPORTED, DEVICE REMAINS IMPLANTED. NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. CONCOMITANT MEDICAL PRODUCTS: CAT #300-01-13 - EQUINOXE, HUMERAL STEM PRIMARY, PRESS FIT 13MM; SN (B)(4). CAT #300-10-45 - EQUINOXE REPLICATOR PLATE 4.5MM O/S; SN (B)(4). CAT #300-20-02 - EQUINOX SQUARE TORQUE DEFINE SCREW DRIVE KIT; SN (B)(4). CAT #310-01-47 - EQUINOXE, HUMERAL HEAD SHORT, 47MM (BETA); SN (B)(4). DEVICE IS IMPLANTED. DESIGN-RELATED ISSUES: EXACTECH IS AWARE OF 58 COMPLAINT REPORTS DUE TO LOOSENING/GLENOID LOOSENING INVOLVING AN ANATOMIC GLENOID COMPONENT SINCE 2008. SALES DATA FOR GLENOID COMPONENTS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINT REPORTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 17 UNITS. THE MANUFACTURING RECORDS FOR THIS DEVICE ONLY EXIST IN PAPER FORM AND ARE STORED IN CARDBOARD BOXES AT IRON MOUNTAIN. DUE TO RISKS ASSOCIATED WITH COVID-19 TRANSMISSION, THE PHYSICAL MANUFACTURING RECORDS FOR THIS COMPLAINT WILL NOT BE REVIEWED. EXACTECH WILL CONTINUE TO MONITOR THIS COMPLAINT TYPE AND, IF WARRANTED, WILL ESCALATE ANY POSITIVE TRENDS TO THE CORRECTIONS AND REMOVAL COMMITTEE (CRC) FOR ADDITIONAL ACTIONS. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE LOOSENING IS LISTED IN THE PRODUCT LABELING, THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE LOOSENING REPORTED WAS LIKELY THE RESULT OF AN INSUFFICIENT BOND BETWEEN THE IMPLANT AND THE BOND WHICH LED TO ASEPTIC (NON-INFECTED) LOOSENING OF THE GLENOID COMPONENT. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. POSTOPERATIVE COMPONENT LOOSENING IS A KNOWN COMPLICATION ASSOCIATED WITH TOTAL SHOULDER ARTHROPLASTY. PER COMPLICATIONS OF TOTAL SHOULDER ARTHROPLASTY (1), POSTOPERATIVE COMPONENT LOOSENING WAS THE NUMBER ONE MOST FREQUENT COMPLICATION OF TOTAL SHOULDER ARTHROPLASTY, WITH A PREVALENCE OF 39% AND 6.3% OF ALL SHOULDERS. THERE ARE ANYWHERE FROM 5% TO 42% REVISION RATES (UNRESTRAINED AND RESTRAINED IMPLANTS). LOOSENING OF THE HUMERAL AND GLENOID COMPONENTS IS A COMMON EVENT, ACCOUNTING FOR 39% OF COMPLICATIONS. FACTORS THAT HAVE BEEN ASSOCIATED WITH GLENOID STABILITY HAVE INCLUDED THE PRESERVATION OF THE SUBCHONDRAL PLATE, CONCENTRIC REAMING OF THE GLENOID, BIOMATERIAL DESIGN AND SELECTION, AND GLENOHUMERAL MISMATCH OF THE PROSTHETIC. BOHASHI, K. I., ET AL. COMPLICATIONS OF TOTAL SHOULDER ARTHROSCOPY. JOURNAL OF BONE AND JOINT SURGERY, 88:2279-2292, OCTOBER 2006. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. REVISIONS OR SURGICAL INTERVENTIONS ARE A KNOWN COMPLICATION FOUND IN JOINT REPLACEMENTS. THERE ARE DEVICE SPECIFIC RISKS OF FRACTURE, MIGRATION, LOOSENING, SUBLUXATION, OR DISLOCATION OF THE PROTHESIS OR ANY OF ITS COMPONENTS, AND ANY OF WHICH MAY REQUIRE A SECOND SURGICAL INTERVENTION OR REVISION. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES. REVISIONS OR SURGICAL INTERVENTIONS ARE A KNOWN COMPLICATION FOUND IN JOINT REPLACEMENTS. THERE ARE DEVICE SPECIFIC RISKS OF FRACTURE, MIGRATION, LOOSENING. THE PATIENT WAS ADVISED TO HAVE A REVISION. AN INVESTIGATION WAS CONDUCTED; THE LOOSENING REPORTED WAS LIKELY THE RESULT OF AN INSUFFICIENT BOND BETWEEN THE IMPLANT AND THE BOND WHICH LED TO ASEPTIC (NON-INFECTED) LOOSENING OF THE GLENOID COMPONENT. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00288 | 2022-03-28 | 2019-11-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TIBIAL PLATE FB SZ 3 LT | EXACTECH, INC. | TIBIAL PLATE FB SZ 3 LT | 350-11-03 |  | IT WAS REPORTED FROM THE VANTAGE TOTAL ANKLE STUDY, THAT A PATIENT EXPERIENCED POSSIBLE LOOSENING NOTED AT 2-YEAR FOLLOW UP X-RAY. ACTION TAKEN: MEDICATION. THE CASE REPORT FORM INDICATES THIS EVENT IS POSSIBLY RELATED TO THE DEVICES AND RELATED TO THE PROCEDURE. THE PATIENT IS ALSO SUFFERING PAIN ALONG POSTERIOR TIBIAL TENDON. NO ADDITIONAL INFORMATION IS AVAILABLE. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. CAT#350-21-13; SN NOT REPORTED; PRODUCT DESCRIPTION -FIXED BEARING LINER COMPONENT AND CAT# 350-01-03; SN NOT REPORTED; PRODUCT DESCRIPTION-TALAR COMPONENT. DEVICE REMAINS IMPLANTED. DESIGN-RELATED ISSUES: THE DESIGN OF THE VANTAGE ANKLE HAS BEEN IN THE FIELD SINCE 2015. SINCE 2015, EXACTECH HAS RECEIVED 6 COMPLAINT REPORTS ANKLE COMPONENTS DUE TO LOOSENING. SALES DATA FOR ALL TIBIAL ANKLE COMPONENTS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE DETAILED BELOW. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: THE TIBIAL IMPLANT IS A PACKAGING LOT COMPRISED OF TWO SPLIT MANUFACTURING LOTS OF 17 AND 23 UNITS. THE TALAR IMPLANT IS A PACKAGING LOT COMPRISED OF TWO SPLIT MANUFACTURING LOTS OF 23 AND 22 UNITS. EXACTECH IS NOT AWARE OF ANY COMPLAINTS INVOLVING PARTS FROM THESE MANUFACTURING LOTS, WHICH HAVE BEEN IN THE FIELD SINCE 2017. THE MANUFACTURING RECORDS FOR THIS DEVICE ONLY EXIST IN PAPER FORM AND ARE STORED IN CARDBOARD BOXES AT IRON MOUNTAIN. DUE TO RISKS ASSOCIATED WITH COVID-19 TRANSMISSION, THE PHYSICAL MANUFACTURING RECORDS FOR THIS COMPLAINT WILL NOT BE REVIEWED. EXACTECH WILL CONTINUE TO MONITOR THIS COMPLAINT TYPE AND, IF WARRANTED, WILL ESCALATE ANY POSITIVE TRENDS TO THE CORRECTIONS AND REMOVAL COMMITTEE (CRC) FOR ADDITIONAL ACTIONS. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE LOOSENING IS LISTED IN THE PRODUCT LABELING, THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR. THE LOOSENING REPORTED WAS LIKELY THE RESULT OF AN INSUFFICIENT BOND BETWEEN THE IMPLANT AND THE BONE, WHICH LED TO POSSIBLE ASEPTIC (NON-INFECTED) LOOSENING. HOWEVER, THIS CANNOT BE CONFIRMED AS THE PATIENT HAS NOT BEEN REVISED. PROPERLY FUNCTIONING IMPLANTS DEPEND ON THEIR APPROPRIATE FIXATION TO THE BONE; FIXATION IS USUALLY ACHIEVED BY CEMENTING THE IMPLANT ONTO THE BONE. SOME SURGEONS PREFER TO USE BIOLOGIC (NON-CEMENTED) FIXATION. ALTHOUGH IMPLANTS ARE FIRMLY FIXED AT THE INITIAL ANKLE REPLACEMENT SURGERY, THEY MAY BECOME LOOSE OVER TIME. THE CAUSE CAN BE RESIDUAL BONEY MALALIGNMENT OF THE FOOT OR POOR FUNCTIONAL BALANCE AND ROM OF THE JOINT SECONDARY TO MUSCULO-TENDON CONTRACTURES OR IMBALANCES BETWEEN ANTAGONISTIC MUSCLE GROUPS. EITHER OF THESE CAN INCREASE CONTACT PRESSURES AT THE ARTICULAR SURFACES OF THE COMPONENTS. ADDITIONALLY, FRICTION CAUSED BY THE JOINT SURFACES RUBBING AGAINST EACH OTHER WEARS AWAY THE SURFACES OF THE IMPLANT, CREATING TINY PARTICLES THAT ACCUMULATE AROUND THE JOINT. IN A PROCESS CALLED ASEPTIC (NONINFECTED) LOOSENING, THE BOND OF THE IMPLANT TO THE BONE IS DESTROYED BY THE BODY'S ATTEMPT TO DIGEST THE WEAR PARTICLES. DURING THIS PROCESS, NORMAL BONE IS ALSO DIGESTED (A CONDITION CALLED OSTEOLYSIS), WHICH CAN WEAKEN OR EVEN FRACTURE THE BONE. WHEN THE PROSTHESIS BECOMES LOOSE, THE PATIENT MAY EXPERIENCE PAIN, CHANGE IN ALIGNMENT, SUBSIDENCE, OR INSTABILITY. COMPONENT LOOSENING IS THE MOST COMMON MODE OF FAILURE OF ANKLE IMPLANTS(1,2). 1. ESCUDERO MI, LE V, BARAHONA M, ET AL. 2019. TOTAL ANKLE ARTHROPLASTY SURVIVAL AND RISK FACTORS FOR FAILURE. FOOT ANKLE INT. 40(9):997¿1006. 2. PENNER M, DAVIS WH, WING K, ET AL. 2018. THE INFINITY TOTAL ANKLE SYSTEM: EARLY CLINICAL RESULTS WITH 2- TO 4-YEAR FOLLOW-UP. FOOT ANKLE SPEC :(B)(4). IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THE STUDY PATIENT IS REPORTED TO BE SUFFERING PAIN ALONG THE POSTERIOR TIBIAL TENDON AND POSSIBLE LOOSENING NOTED AT 2-YEAR FOLLOW UP X-RAY. LOOSENING IS A KNOWN CAUSE FOR ORTHOPEDIC REVISION SURGERY. AN INVESTIGATION WAS CONDUCTED; THE LOOSENING REPORTED WAS LIKELY THE RESULT OF AN INSUFFICIENT BOND BETWEEN THE IMPLANT AND THE BONE, WHICH LED TO POSSIBLE ASEPTIC (NON-INFECTED) LOOSENING. HOWEVER, THIS CANNOT BE CONFIRMED AS THE PATIENT HAS NOT BEEN REVISED. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00287 | 2022-03-28 | 2019-08-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | HEAD TRIAL, 12/14 36 +0 | EXACTECH, INC. |  | 143-36-00 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE TRIAL HEAD CRACKED IN HALF WHEN REMOVING FROM THE HIP. THE DEVICE WAS NOT RETURNED AS IT WAS DISPOSED OF AT THE HOSPITAL; THEREFORE, NOT AVAILABLE FOR EVALUATION. NO PIECES FELL INTO THE SURGICAL SITE AND THERE WAS NO PATIENT INJURY, ADVERSE EVENT, OR CLINICAL CONSEQUENCE TO THE PATIENT. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS NOT RETURNED AS IT WAS DISPOSED OF AT THE HOSPITAL; THEREFORE, NOT AVAILABLE FOR EVALUATION. DESIGN-RELATED ISSUES: THE DESIGN OF THE 12/14 FEMORAL HEAD TRIAL, O-RING HAS BEEN IN THE FIELD SINCE 2007. EXACTECH IS AWARE OF 6 COMPLAINT REPORTS INVOLVING BROKEN HEAD TRIALS SINCE 2008. SALES DATA FOR ALL FEMORAL HEADS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4). THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: MANUFACTURING DATA COULD NOT BE REVIEWED BECAUSE THE SERIAL/LOT INFORMATION WAS NOT PROVIDED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) AND RISK ASSESSMENT AND CONTROLS REPORT (RACR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR AND RACR FOR THIS FEMORAL HEAD TRIAL, (B)(4), WERE REVIEWED. THE RISK IS CAPTURED, HOWEVER THE RACR IS BEING UPDATED UNDER (B)(4) TO CAPTURE AN APPROPRIATE HAZARDOUS SITUATION FOR THE REPORTED EVENT. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, THE RISK IS CAPTURED IN THE RMR AND RACR, AND THE HAZARDOUS SITUATION WILL BE UPDATED. THE BROKEN FEMORAL HEAD TRIAL REPORTED WAS LIKELY THE RESULT OF AN EXPOSURE TO IMPACTION FORCES AND SUBSEQUENT STERILIZATION CYCLES WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE WHEN REMOVING THE HEAD TRIAL. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. ¿ CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE WAS NO PATIENT ADVERSE EVENT, NONE OF THE DEVICE FRAGMENTS FELL INTO THE SURGICAL SITE. AN INVESTIGATION WAS CONDUCTED; THE BROKEN FEMORAL HEAD TRIAL REPORTED WAS LIKELY THE RESULT OF AN EXPOSURE TO IMPACTION FORCES AND SUBSEQUENT STERILIZATION CYCLES WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE WHEN REMOVING THE HEAD TRIAL. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00286 | 2022-03-28 | 2019-02-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LOGIC CR TRIAL TIB INST STD, SZ 3, 9MM | EXACTECH, INC. |  | 02-013-47-3009 |  | IT WAS REPORTED FROM THE OUS THAT THE FRONT FACE OF THE TRIAL HAS FRACTURED AND WAS DISCOURSED DURING SURGERY. THIS DID NOT CAUSE ANY ISSUES DURING SURGERY. THE DEVICE WAS DESTROYED BY THE HOSPITAL. NO PIECES FELL INTO THE SURGICAL SITE. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. ADDITIONAL MFR NARRATIVE INFORMATION WAS NOT PROVIDED FOR: PT. INFO. ANALYSIS COMPLETED ON 01/23/2020. DESIGN-RELATED ISSUES: THE DESIGN OF THE LOGIC TIBIAL INSERT TRIAL HAS BEEN IN THE FIELD SINCE 2009. EXACTECH IS AWARE OF 3 OTHER SIMILAR COMPLAINT REPORTS INVOLVING TIBIAL INSERTS TRIALS DUE TO DEFORMATION SINCE 2009. (B)(4). THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN-RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF (B)(4) UNITS, WHICH HAVE BEEN IN THE FIELD SINCE 2014. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING-RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISKS WERE INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS LOGIC TIBIAL INSERT TRIALS, 750-2006-012-RMR- INSTRUMENTS REV A, WAS REVIEWED. THE RISKS ARE CAPTURED IN LINES 20 AND 33. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE BROKEN DEVICE REPORTED IN EXPERIENCE (B)(4) WAS LIKELY THE RESULT OF YEARS OF NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES WHERE THE TIBIAL INSERT TRIAL MAY HAVE COME IN CONTACT WITH SHARP SURGICAL OBJECTS. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION; VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THE FRACTURED DEVICE DID NOT CAUSE A PATIENT ADVERSE EVENT, NO PIECES FELL INTO THE SURGICAL SITE. AN INVESTIGATION WAS CONDUCTED; THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF YEARS OF NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES WHERE THE TIBIAL INSERT TRIAL MAY HAVE MET SHARP SURGICAL OBJECTS. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10033 | 2022-03-29 | 2018-04-19 | I, n, j, u, r, y | Manufacturer report | EXACTECH | REAMER HEAD, MODULAR REVERSE, PILOT TIP, 46MM | EXACTECH, INC. |  | 321-25-46 |  | IT WAS REPORTED FROM THE US THAT WHILE REAMING THE GLENOID, THE PILOT TIP OF THE REAMER BROKE OFF IN THE BONE. THE TIP WAS EASILY RETRIEVED FROM THE BONE BY THE SURGEON. SURGERY OUTCOME WAS GOOD, AND ALL BROKEN PARTS WERE RETRIEVED FROM THE PATIENT WITH NO INJURY. PATIENT'S HEALTH WAS STABLE LEAVING THE OR. ISSUE WAS RESOLVED BY PULLING THE BROKEN PILOT TIP OUT OF THE BONE. THE AGENT WAS PRESENT AT THE TIME OF SURGERY. NO ADDITIONAL INFORMATION IS AVAILABLE. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. MFG DATE: 01/03/2011 BASED ON THE INVESTIGATION CONDUCTED UNDER (B)(4), IT WAS DETERMINED THAT THE USER INSTRUCTION (E.G. THE SURGICAL TECHNIQUE) WAS NOT ADEQUATE TO INFORM THE SURGEON OF THE APPROPRIATE USE OF THE REAMER. THE FORESEEABLE MISUSE OF THE GLENOID REAMER THAT COULD RESULT IN FRACTURE OF THE PILOT TIP ARE: IF THE USER APPLIES A BENDING MOMENT TO THE PILOT TIP DURING REAMING, WHICH WOULD REAM THE GLENOID ASYMMETRICALLY AND APPLY A TORQUE TO THE PILOT TIP, POTENTIALLY LEADING TO FRACTURE. IF THE USER APPLIES A BENDING MOMENT TO THE PILOT TIP BY USING THE GLENOID REAMER TO RETRACT THE HUMERAL HEAD TO HELP GAIN EXPOSURE AND ACCESS TO THE GLENOID, WHICH WOULD APPLY A TORQUE TO THE PILOT TIP, POTENTIALLY LEADING TO FRACTURE. THE RISK DOCUMENTATION WAS LACKING SPECIFIC FAILURES FOR THIS TYPE OF DEFECT FOR THESE INSTRUMENTS. IN ADDITION, THE SURGICAL TECHNIQUE WAS NOT PROVIDING INFORMATION TO THE USER OF THE INSTRUMENT THAT REAMING OFF AXIS CAN DAMAGE OR BREAK THE INSTRUMENTS. ADDITIONALLY, THE SURGICAL TECHNIQUE DID NOT PROVIDE ADEQUATE INSTRUCTION TO THE USER. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. THE TIP WAS EASILY RETRIEVED FROM THE BONE BY THE SURGEON. AN INVESTIGATION WAS CONDUCTED UNDER (B)(4); THE MOST LIKELY ROOT CAUSE WAS THE RISK DOCUMENTATION WAS LACKING SPECIFIC FAILURES FOR THIS TYPE OF DEFECT FOR THESE INSTRUMENTS. IN ADDITION, THE SURGICAL TECHNIQUE WAS NOT PROVIDING INFORMATION TO THE USER OF THE INSTRUMENT THAT REAMING OFF AXIS CAN DAMAGE OR BREAK THE INSTRUMENTS. ADDITIONALLY, THE SURGICAL TECHNIQUE DID NOT PROVIDE ADEQUATE INSTRUCTION TO THE USER. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10021 | 2022-03-29 | 2018-09-05 | I, n, j, u, r, y | Manufacturer report | EXACTECH | MODULAR REVERSE SHOULDER PILOT REAMER HEAD - 38MM | EXACTECH, INC. |  | 321-25-38 |  | IT WAS REPORTED FROM OUS THAT DURING A PRIMARY ORTHOPEDIC SURGERY THE MODULAR REVERSE SHOULDER PILOT REAMER HEAD BROKE. THE AGENT WAS PRESENT AT TIME OF SURGERY. NO DELAY TO SURGERY. THE PATIENT'S HEALTH WAS STABLE LEAVING THE OPERATING ROOM. INFORMATION RECEIVED ON (B)(6) 2021: "THE INSTRUMENT COULD HAVE BEEN BROKEN DURING THE SURGERY. HOWEVER, THE BREAKAGE HAS BEEN NOTICED AT THE RECOVERY OF THE INSTRUMENTATION-SET BY OUR LOGISTICS DEPT., AFTER THE CLEANING REALIZED BY THE STERILIZATION TEAM OF THE HOSPITAL. NO INFORMATION HAS BEEN RECEIVED FROM THE HOSPITAL REGARDING THIS ISSUE AND THE AGENCY DIDN'T RECEIVE ANY REQUEST FOR COMPLEMENTARY INFORMATION REGARDING THIS ISSUE AS WELL." MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. ENG EVAL COMPLETED ON (B)(6) 2018. CAPA: (B)(4). FINDINGS/ROOT CAUSE: BASED ON THE INVESTIGATION CONDUCTED UNDER CAPA (B)(4), IT WAS DETERMINED THAT THE USER INSTRUCTION WAS NOT ADEQUATE TO INFORM THE SURGEON OF THE APPROPRIATE USE OF THE REAMER. THE FORESEEABLE MISUSE OF THE GLENOID REAMER THAT COULD RESULT IN FRACTURE OF THE PILOT TIP ARE: IF THE USER APPLIES A BENDING MOMENT TO THE PILOT DIP DURING REAMING, WHICH WOULD REAM THE GLENOID ASYMMETRICALLY AND APPLY A TORQUE TO THE PILOT TIP, POTENTIALLY LEADING TO FRACTURE. IF THE USER APPLIES A BENDING MOMENT TO THE PILOT TIP BY USING THE GLENOID REAMER TO RETRACT THE HUMERAL HEAD TO HELP GAIN EXPOSURE AND ACCESS TO THE GLENOID, WHICH WOULD APPLY A TORQUE TO THE PILOT TIP, POTENTIALLY LEADING TO FRACTURE. CORRECTIVE ACTION TAKEN: AS PART OF CAPA (B)(4), THE FOLLOWING CORRECTIVE ACTIONS WERE TAKEN: THE SURGICAL TECHNIQUE WAS UPDATED TO ADVISE SURGEONS TO AVOID APPLYING A BENDING TORQUE TO THE PILOT TIP REAMER OR USING THE REAMER TO RETRACT THE HUMERAL HEAD AS THIS MAY CAUSE FRACTURE OF THE PILOT TIP. ((B)(4)). UPDATE THE RMR TO INCLUDE THE RISK OF PILOT TIP REAMER FRACTURE AS A RESULT OF BENDING OR ITS FORESEEABLE MISUSE AS A RETRACTOR. ((B)(4)). PER (B)(4), A FIELD ADVISORY NOTICE WAS SENT TO THE USERS NOTIFYING THEM OF THE ISSUE AND ALERTING THEM TO THE CHANGE IN THE OPERATIVE TECHNIQUE (EXACTECH RECALL NUMBER 1038671-04/07/2017-003-C, FDA RECALL NUMBERS Z-2663-2017 THROUGH Z-2668-2017, MAY 2017). THE RECALL WAS TERMINATED BY THE FDA ON (B)(6) 2019. CAPA (B)(4) WAS INITIATED TO TRACK REPORTED EVENTS THROUGH 2021 AND RECONVENE THE CRC IF A THRESHOLD OF >35 TOTAL EVENTS OR >3 EVENTS WITH REPORTED PATIENT HARM BE MET. AS A RESULT OF MONITORING UNDER (B)(4) ((B)(6) 2019) AND CAPA (B)(4), IT WAS NOTED THAT THE THRESHOLD OF EVENTS FOR 2020 HAS BEEN EXCEEDED. THIS ISSUE HAS BEEN RE-ESCALATED FOR EVALUATION. AS DETERMINED BY (B)(4) AND (B)(4), THE RISK FALLS WITHIN THE ACCEPTABLE LEVEL. THE COMMITTEE HAS DETERMINED THAT NO ADDITIONAL FIELD ACTION IS REQUIRED AT THIS TIME. EXACTECH WILL CONTINUE MONITOR THE EVENTS THROUGH CAPA (B)(4), WHICH WILL DRIVE ANY FURTHER HHE UPDATES AND/OR CRC ESCALATION. (B)(4): INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT IMPACT. BASED ON THE INVESTIGATION CONDUCTED UNDER CAPA (B)(4), IT WAS DETERMINED THAT THE USER INSTRUCTION WAS NOT ADEQUATE TO INFORM THE SURGEON OF THE APPROPRIATE USE OF THE REAMER. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10001 | 2022-03-29 | 2017-02-20 | I, n, j, u, r, y | Manufacturer report | EXACTECH | FEMORAL FINISHING GUIDE SIZE 4 | EXACTECH, INC. |  | 201-51-14 |  | IT WAS REPORTED THAT DURING AN ORTHOPEDIC KNEE SURGERY THE SURGEON EXPERIENCED THE SIZE 4 FEMORAL RESECTION BLOCK, THE TWO SHORT METAL PINS THAT FIX THE BLOCK TO THE BONE SHEARED OFF COMPLETELY. THE SALES REPRESENTATIVE WAS NOT PRESENT AT THE SURGERY AND REPORTED NO DEVICE FRAGMENTS FELL INTO THE SURGICAL SITE. THE SURGERY CONTINUED AND WAS COMPLETED WITHOUT ADDITIONAL ISSUES. THERE WAS NO DELAY OF SURGERY AND THERE WAS NO PATIENT INJURY, ADVERSE EVENT, OR CLINICAL CONSEQUENCE TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE REPORTER RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE IN QUESTION WAS RETURNED TO EXACTECH FOR ANALYSIS. DESIGN RELATED ISSUES: THIS DESIGN HAS BEEN IN THE FIELD SINCE 1995. AT THE TIME OF THIS ANALYSIS, 31 MAY 2018, EXACTECH HAD RECEIVED 13 SIMILAR COMPLAINTS INVOLVING 13 FEMORAL FINISHING GUIDES SINCE 2008. THE TOTAL CALCULATE OCCURRENCE RATE OF <0.5% IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE; THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MANUFACTURING: EXACTECH HAD NOT RECEIVED ANY OTHER COMPLAINT INVOLVING PARTS FROM THIS MANUFACTURING LOT WHICH HAS BEEN IN THE FIELD SINCE 2006. THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. RISK MANAGEMENT: A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED AND FOUND THE OCCURRENCE RATE IS BELOW THE PREDICTED THRESHOLD. THE RMR FOR TIS FEMORAL FINISHING GUIDE, 750-1992-001-RMR REV - AND THE RISK IS CAPTURED IN LINE 13. EXACTECH IFU 700-096-004 STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED INSTRUMENT MALFUNCTION IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO ANY REPORTED PATIENT IMPACT. AN INVESTIGATION WAS CONDUCTED THAT THE MOST LIKELY ROOT CAUSE OF THE DISASSEMBLY REPORTED IN THIS EVENT WAS LIKELY DUE TO A HIGH ENERGY EVENT, SUCH AS DROPPING THE INSTRUMENT ONTO A HARD SURFACE WHICH CAUSED THE PEGS TO DETACH FROM THE REST OF THE DEVICE. | Y | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10024 | 2022-03-29 | 2017-11-20 | I, n, j, u, r, y | Manufacturer report | EXACTECH | FEMORAL HEAD TRIAL | EXACTECH, INC. |  | 143-40-93 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE HEAD TRIAL. THE HEAD TRIAL GOT STUCK ON AN ELEMENT FEMORAL STEM AND THE WHOLE SYSTEM HAD TO BE REMOVED. ANOTHER SET WAS OPENED. THE PATIENT WAS FINE LEAVING THE OPERATING ROOM. THERE WAS A DELAY, BUT THE PATIENT WAS NOT ADVERSELY AFFECTED. THE REP WAS PRESENT AT THE TIME OF SURGERY. NO PATIENT DEMOGRAPHICS WERE REPORTED. THE HEAD TRIAL WAS RETURNED ON 12/13/2017. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT FOLLOWING MULTIPLE ATTEMPTS. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS AVAILABLE FOR ANALYSIS. ENGINEERING ANALYSIS COMPLETED BY AD ON 9/24/2018. DESIGN-RELATED ISSUES: THE DESIGN OF THE FEMORAL HEAD TRIAL HAS BEEN IN THE FIELD SINCE 2011. EXACTECH HAS RECEIVED 4 COMPLAINT REPORTS INVOLVING 6 FEMORAL HEAD TRIALS DUE TO DIFFICULT ASSEMBLY/DISASSEMBLY SINCE 2008. BECAUSE FEMORAL HEAD TRIALS ARE USED IN TOTAL HIP SURGERIES, SALES DATA FOR FEMORAL HEADS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4). THIS IS CONSIDERED "'VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH HAS NOT RECEIVED ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT, WHICH HAS BEEN IN THE FIELD SINCE 2012. HOWEVER, THIS PART WAS FOUND TO BE OUTSIDE THE TOLERANCE RANGE. THIS WAS CONFIRMED VIA SPECIAL INSPECTION. A SUPPLIER CORRECTIVE ACTION REQUEST (SCAR) HAS BEEN ISSUED BY THE PRODUCT TEAM. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS FEMORAL HEAD TRIAL, 750-2011-012-RMR REV A, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 7. THE DISASSEMBLY DIFFICULTY REPORTED IN EXPERIENCE (B)(4) WAS LIKELY THE RESULT OF THE HEAD TRIAL BEING MANUFACTURED OUTSIDE THE ALLOWABLE TOLERANCE RANGE. THIS WAS CONFIRMED VIA SPECIAL INSPECTION. A SUPPLIER CORRECTIVE ACTION REQUEST HAS BEEN ISSUED BY THE PRODUCT TEAM. EXACTECH IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO DESIGN ISSUES. THE HEAD TRIAL GOT STUCK ON AN ELEMENT FEMORAL STEM AND THE WHOLE SYSTEM HAD TO BE REMOVED. THERE WAS NO PATIENT ADVERSE EVENT. AN INVESTIGATION WAS CONDUCTED: THE DISASSEMBLY DIFFICULTY REPORTED WAS LIKELY THE RESULT OF THE HEAD TRIAL BEING MANUFACTURED OUTSIDE THE ALLOWABLE TOLERANCE RANGE. THIS WAS CONFIRMED VIA SPECIAL INSPECTION. A SUPPLIER CORRECTIVE ACTION REQUEST HAS BEEN ISSUED BY THE PRODUCT TEAM. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10026 | 2022-03-29 |  | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRI-DRIVE CANNULATED SHAFT | EXACTECH, INC. |  | 315-25-00 |  | IT WAS REPORTED FROM THE US THAT INSTRUMENTS BROKE DURING A CASE. THE INSTRUMENTS WERE RETURNED TO THE COMPANY WITH A NOTE STATING THEY BROKE DURING A CASE. NO ADDITIONAL INFORMATION WAS PROVIDED. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. THE AGENCY IS INACTIVE, NO FURTHER INFORMATION IS AVAILABLE RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICES WERE RECEIVED FOR ANALYSIS. ENG EVAL COMPLETED BY (B)(4). VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR 10X OPTICAL UNLESS OTHERWISE SPECIFIED) THE FRACTURED SURFACES APPEAR TO BE CONSISTENT WITH FATIGUE CRACK PROPAGATION OVER YEARS OF NORMAL USE. ALTHOUGH THE PIN IS STILL ASSEMBLED, THE WELD HOLDING THE PIN IN PLACE APPEARS TO BE CRACKED. THE CRACK AT THE WELD OF THE MATING FEATURE APPEARS CONSISTENT WITH DEGRADATION OF THE LASER WELD OVER TIME. THE TIP APPEARS TO HAVE DISASSEMBLED. THE RETAINING RING IS MEANT TO HOLD THE SLEEVE AND SPRING COMPONENTS IN PLACE BUT APPEARS TO HAVE DISLODGED. THE USER LIKELY HELD THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY AND ALLOWED THE TIP TO DISASSEMBLE. THE BALL BEARING WAS NOT RETURNED. A COMPLAINT COORDINATOR REACHED OUT TO THE AGENCY ON NOVEMBER 14TH, 2018 REGARDING THE BALL BEARING AND NO RESPONSE WAS RECEIVED AS OF JANUARY 11TH, 2019. DESIGN-RELATED ISSUES: THE DESIGN OF THE EQUINOXE REVERSE SHOULDER IMPACTOR HANDLE AND TRI-DRIVE CANNULATED HANDLE HAVE BEEN IN THE FIELD SINCE 2006 AND 2009, RESPECTIVELY. EXACTECH IS AWARE OF 6 OTHER SIMILAR COMPLAINT REPORTS INVOLVING EQUINOXE REVERSE SHOULDER IMPACTOR HANDLES AND 42 OTHER SIMILAR COMPLAINTS INVOLVING TRI-DRIVE CANNULATED HANDLE SINCE 2008. THE FREQUENCY OF OCCURRENCE RANKING SCALE IS "VERY LOW", THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. (B)(4). THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. THE RMR FOR THIS IMPACTOR HANDLE, 750-2005-021-RMR-INSTRUMENTS REV M, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 15. THE RMR FOR THIS TRI-DRIVE CANNULATED HANDLE, 750-2007-026-RMR-INSTRUMENTS REV J, WAS REVIEWED. THE RISK IS NOT CAPTURED. THIS DOCUMENT WILL BE UPDATED TO INCLUDE THIS FAILURE. REFER TO (B)(4) FOR DETAILS. CORRECTIVE ACTION FOR THE IMPACTOR HANDLE IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. CORRECTIVE ACTION FOR THE TRIDRIVE CANNULATED SHAFT IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RMR WILL BE UPDATED TO CAPTURE THE RISK. THE BROKEN IMPACTOR HANDLE REPORTED IN EXPERIENCE (B)(4) WAS LIKELY THE RESULT OF YEARS OF NORMAL USE, PROCESSING, AND HANDLING, WHICH LED TO WELD FRACTURING AT THE STRIKE PLATE-HANDLE INTERFACE. THE DISASSEMBLED TRI-DRIVE CANNULATED HANDLE REPORTED WAS LIKELY THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY AND ALLOWED THE TIP TO DISASSEMBLE. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE WAS NO PATIENT ADVERSE EVENT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN IMPACTOR HANDLE REPORTED WAS LIKELY THE RESULT OF YEARS OF NORMAL USE, PROCESSING, AND HANDLING, WHICH LED TO WELD FRACTURING AT THE STRIKE PLATE-HANDLE INTERFACE. THE DISASSEMBLED TRI-DRIVE CANNULATED HANDLE REPORTED WAS LIKELY THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY AND ALLOWED THE TIP TO DISASSEMBLE. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. EVENT DATE NOT REPORTED. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10005 | 2022-03-29 | 2017-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LARGE REAMER | EXACTECH, INC. |  | 315-35-14 |  | IT WAS REPORTED THAT DURING AN ORTHOPEDIC SHOULDER SURGERY THE SURGEON ATTEMPTED TO USE THE CANNULATED SYSTEM TO SET K-WIRE. THE WIRE BOUND IN THE GUIDE BECAUSE THEY COULD NOT GET STRAIGHT ON ACCESS (K-WIRE AND GUIDE PART NUMBERS UNKNOWN). THE SURGEON BROKE THE LARGE REAMER DURING USE BECAUSE OF LEVERING AGAINST THE HUMERUS (315-35-14) AND BROKE THE X-LARGE REAMER (315-35-15). THE SALES REPRESENTATIVE WAS PRESENT AT THE SURGERY AND REPORTED NO DEVICE FRAGMENTS FELL INTO THE SURGICAL SITE. THE SURGERY CONTINUED AND WAS COMPLETED WITHOUT ADDITIONAL ISSUES. THERE WAS A MINOR DELAY OF SURGERY AND THERE WAS NO REPORTED PATIENT INJURY, ADVERSE EVENT, OR CLINICAL CONSEQUENCE TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE REPORTER RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICES IN QUESTION WERE NOT RETURNED TO EXACTECH FOR ANALYSIS. THE BROKEN INSTRUMENTS REPORTED WERE LIKELY THE RESULT OF INADEQUATE EXPOSURE OF THE JOINT, PREVENTING "STRAIGHT ON" ACCESS FOR THE K-WIRE, GUIDE, AND REAMERS. EXACTECH IFU 700-096-004 STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THE DEVICES REPORTED IN THIS EVENT ARE USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED INSTRUMENTS MALFUNCTION IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO ANY REPORTED PATIENT IMPACT. AN INVESTIGATION WAS CONDUCTED AND CONCLUDED THAT THE LIKELY CAUSE OF THE ISSUES WAS THE RESULT OF INADEQUATE EXPOSURE OF THE JOINT, PREVENTING "STRAIGHT ON" ACCESS FOR THE K-WIRE, GUIDE, AND REAMERS. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10012 | 2022-03-29 | 2017-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LPI IMPACTOR HANDLE | EXACTECH, INC. |  | 213-65-00 |  | IT WAS REPORTED FROM OUS THAT DURING A SURGICAL PROCEDURE THE INSTRUMENT BROKE UPON IMPACT OF THE FEMUR. NO PIECES OF THE DEVICE FELL INTO THE SURGICAL SITE. THERE WAS NO DELAY OF SURGERY OR INJURY/CLINICAL CONSEQUENCES TO THE PATIENT. IT WAS REPORTED THAT "THE PATIENT IS FINE". THE AGENT WAS PRESENT AT THE TIME OF SURGERY. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. (B)(4) WAS INITIATED ON 03/02/2018 TO EVALUATE THIS ISSUE. AN INVESTIGATION OF COMPLAINTS RELATED TO LPI HANDLE - BROKEN NON-FUNCTIONAL, SPECIFICALLY RELATED TO CRACKS/BREAKAGE FOR THE PLASTIC HANDLE. THE LPI HANDLE CONSISTS OF A METAL STRIKE PLATE, METAL END PIECE WHICH ENGAGES WITH MATING INSTRUMENTS AND IMPLANTS, AND A PLASTIC CENTRAL HANDLE. THE METAL END PIECES ARE ASSEMBLED TO THE PLASTIC HANDLE VIA THERMAL ASSEMBLY WHICH INVOLVES HEATING THE METAL AND PRESSING THE HOT METAL PIECES INTO THE HANDLE. THE METAL PIECES HAVE CYLINDRICAL BOSSES WHICH FEATURE A PRESS -FIT WITH A FEMALE HOLE ON THE PLASTIC HANDLE. THE LPI HANDLE ASSEMBLES WITH TWO END PIECES FOR USE THREE TIMES DURING THE SURGICAL PROCEDURE: 1) IMPACTION OF THE FEMORAL TRIAL 2) IMPACTION OF THE TIBIAL TRAY IMPLANT AND 3) IMPACTION OF THE FEMORAL IMPLANT. AVAILABLE DEVICES FROM PREVIOUSLY REPORTED COMPLAINTS WERE EVALUATED. COMPLAINT DEVICES APPEARED TO HAVE SIMILAR FAILURE MODES, AND EITHER WERE CRACKED OR FRACTURED. CRACKS WERE CONSISTENT WITH STARTING INSIDE THE PLASTIC HANDLE AT THE BASE OF ONE OF THE TWO METAL ATTACHMENT FEATURES AND SPREADING OUTWARD. CRACKS WERE EITHER LINEAR, ALONG THE HANDLE AXIS OR HORIZONTAL, NEARLY IN LINE WITH THE ORIGIN. THE REVIEW OF THE LPI IMPACTOR HANDLES (213-65-00) CITED IN COMPLAINTS FOR BROKEN OR CRACKED HANDLES SHOWED A CONSISTENT FAILURE MODE. A REVIEW OF THE DEVICE ORDER HISTORY DID NOT SHOW A TREND IN THOSE LOTS WITH COMPLAINT DEVICES, NOR DID A REVIEW OF DESIGN CHANGES SHOW CHANGES MADE TO THE DESIGN WHICH COULD CONTRIBUTE TO THE COMPLAINT. BASED ON THE INVESTIGATION, THE MOST LIKELY CAUSE OF THE REPORTED EVENTS IS RELATED TO HIGH IMPACTION FORCES REQUIRED DURING IMPLANT SURGERY. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT IMPACT. THE HANDLE BROKE DURING IMPACTION. AN INVESTIGATION WAS CONDUCTED; THE MOST LIKELY CAUSE OF THE REPORTED EVENT IS RELATED TO HIGH IMPACTION FORCES REQUIRED DURING IMPLANT SURGERY. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10015 | 2022-03-29 | 2017-11-22 | I, n, j, u, r, y | Manufacturer report | EXACTECH | HUMERAL LINER IMPACTOR TIP ,38MM | EXACTECH, INC. |  | 321-07-38 |  | IT WAS REPORTED FROM THE US THAT THE SURGEON WAS IMPACTING THE HUMERAL LINER INTO THE HUMERAL TRAY IN THE USUAL AND CUSTOMARY MANNER. ON THE FIRST STRIKE OF THE MALLET, THE IMPACTOR TIP BROKE INTO TWO PARTS. IT WAS REPORTED THAT NO PIECES WERE LEFT IN THE PATIENT. THE ISSUE WAS RESOLVED BY USING THE 42 MM IMPACTOR HEAD. IT WAS IMPACTED ONCE AND CONFIRMED THAT THE IMPLANT WAS PROPERLY SEATED. THE AGENT WAS PRESENT AT SURGERY WITH NO ADVERSE EVENT TO THE PATIENT. MULTIPLE ATTEMPTS HAVE BEEN MADE FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. NO ADDITIONAL ATTEMPTS WILL BE MADE FOR INFORMATION REGARDING THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED FOR ANALYSIS. COMPLETED ON (B)(6) 2018 BY MB DESIGN-RELATED ISSUES: THE DESIGN OF THE HUMERAL LINER IMPACTOR HAS BEEN IN THE FIELD SINCE 2006. EXACTECH HAS RECEIVED 11 SIMILAR COMPLAINT REPORTS INVOLVING 15 IMPACTOR TIPS SINCE 2008. BECAUSE THE HUMERAL LINER IMPACTOR TIP IS USED IN REVERSE TOTAL SHOULDER SURGERIES, SALES DATA FOR HUMERAL LINERS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THE DESIGN OF THIS INSTRUMENT HAS SINCE BEEN UPDATED. REFER TO (B)(4). MFG-RELATED ISSUES: EXACTECH HAS NOT RECEIVED ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT, WHICH HAS BEEN IN THE FIELD SINCE 2007. THIS INSTRUMENT WAS MANUFACTURED PRIOR TO THE DESIGN BEING UPDATED. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING-RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS HUMERAL LINER IMPACTOR TIP, 750-2005-021-RMR (B)(4), WAS REVIEWED. THE RISK IS CAPTURED IN 17. MOST LIKELY CAUSE: THE BROKEN IMPACTOR TIP REPORTED IN EXPERIENCE (B)(6) WAS LIKELY THE RESULT OF A STRESS RISER INTRODUCED DURING IMPACTION, WHICH LED TO CRACK INITIATION, PROPAGATION AND ULTIMATE FRACTURE OF THE DEVICE. THE DESIGN HAS SINCE BEEN UPDATED. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. THIS EVENT DID NOT LEAD TO PATIENT ADVERSE EVENT, IT WAS REPORTED THAT NO PIECES WERE LEFT IN THE PATIENT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN IMPACTOR TIP REPORTED WAS LIKELY THE RESULT OF A STRESS RISER INTRODUCED DURING IMPACTION, WHICH LED TO CRACK INITIATION, PROPAGATION AND ULTIMATE FRACTURE OF THE DEVICE. THE DESIGN HAS SINCE BEEN UPDATED. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10019 | 2022-03-29 | 2018-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRULIANT FEMORAL IMPACTOR HEAD LOCKING | EXACTECH, INC. |  | 02-029-19-1000 |  | IT WAS REPORTED THAT DURING SETTING UP FOR AN ORTHOPEDIC KNEE SURGERY THE TRULIANT FEMORAL IMPACTOR HEAD LOCKING THAT WOULD NOT LOCK THE TAMP HEAD DUE TO BEING BROKEN. THE SALES REPRESENTATIVE WAS PRESENT AT THE SURGERY AND REPORTED NO DEVICE FRAGMENTS FELL INTO THE SURGICAL SITE AS THE EVENT OCCURRED DURING INSTRUMENT SET UP. THE SURGERY CONTINUED AND WAS COMPLETED WITHOUT ADDITIONAL ISSUES. THERE WAS NO DELAY OF SURGERY AND THERE WAS NO PATIENT INJURY, ADVERSE EVENT, OR CLINICAL CONSEQUENCE TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE REPORTER RELATED TO THIS EVENT. THE AGENT/AGENCY IS NO LONGER ACTIVE WITH EXACTECH. NO ADDITIONAL INFORMATION IS AVAILABLE TO FURTHER INVESTIGATE THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE INSTRUMENT WAS RETURNED OR ANALYSIS AND AFTER EVALUATION IT WAS FOUND THE CAP SCREW AND WASHER WERE MISSING FROM THE ASSEMBLY. ALL OTHER ASPECTS OF THE DEVICE APPEARED UNREMARKABLE. CAPA(B)(4): INVESTIGATION REVEALED THAT THE FAILURE OF THE CAP SCREW WAS DIRECTLY RELATED TO INADEQUATE WELDING BETWEEN THE CAP SCREW AND THE WASHER SUBCOMPONENT APPLIED DURING MANUFACTURING ASSEMBLY. THERE WAS ONLY A MINIMAL SPOT WELD APPLIED WHICH WAS VULNERABLE DURING IMPACTION USE. A REVIEW OF THE PRINT SHOWED THE WELD REQUIREMENT WAS CALLED OUT AS "TACK WELD", WHICH WAS NOT SPECIFIC ENOUGH AND LEFT ROOM FOR INTERPRETATION. EXACTECH IFU 700-096-004 STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED INSTRUMENT MALFUNCTION IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO ANY PATIENT IMPACT. AN INVESTIGATION WAS CONDUCTED THAT THE MOST LIKELY ROOT CAUSE OF THE DEVICE BREAKAGE REPORTED IN THIS EVENT WAS LIKELY DUE TO NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES WHICH LED TO THE WELD BREAKING. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00302 | 2022-03-29 | 2022-03-08 | I, n, j, u, r, y | Manufacturer report | LOGIC CR TIB INSERT SLOPE++, SZ 3, 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC CR TIB INSERT SLOPE++, SZ 3, 11MM | 02-012-49-3011 |  | AS REPORTED, APPROXIMATELY NINE YEARS POST INITIAL RIGHT TKA, THIS (B)(6) MALE PATIENT'S POLY HAD SEVERE POSTERIOR WEAR. WAS REVISED TO SIZE 3 CRC 15MM POLY. PATIENT STABLE AFTER IMPLANTATION. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. IMAGES RECEIVED. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO HOSPITAL POLICY. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10031 | 2022-03-29 | 2018-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LOGIC PS TIBIA, MODULAR INSERT TRIAL, SIZE 3, 9MM | EXACTECH, INC. |  | 02-013-35-3009 |  | IT WAS REPORTED FROM OUS THAT DURING A PRIMARY ORTHOPEDIC SURGICAL PROCEDURE THE TRIAL CHIPPED. THERE WERE NO FRAGMENTS OR PIECES THAT FELL INTO THE SURGICAL SITE. IT WAS REPORTED THAT THE PATIENT WAS STABLE LEAVING THE OR. THE AGENT WAS PRESENT AT THE TIME OF SURGERY. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED FOR ANALYSIS. THE COMPLAINED PRODUCT DESCRIPTION OF CHIPPED IS CONFIRMED. MFG DATE: 10/02/2014. VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR 10X OPTICAL) THE SCRATCHES AND DEFORMATION APPEAR CONSISTENT WITH MEETING THE TIBIAL TRAY TRIAL. THE LOCATIONS THAT ARE SEVERELY DEFORMED INDICATE THE TRIAL MAY HAVE NOT BEEN PROPERLY ALIGNED WITH THE TRAY TRIAL WHEN THE STABILITY ASSESSMENT WAS PERFORMED OR CAME LOOSE IN THE PROCESS. THESE MARKINGS MAY HAVE ALSO BEEN CREATED WHILE ATTEMPTING TO INSERT THE TRIAL WITH EXCESSIVE FORCES. THE FEMORAL TRIAL COMPONENT MAY HAVE PARTIALLY OBSTRUCTED THE INSERTION OF THE INSERT TRIAL RESULTING IN CONTACT FORCES BETWEEN THE SHARPER EDGES OF THE TRAY TRIAL AND INSERT TRIAL BEFORE BEING PROPERLY SEATED. THE FREQUENCY OF OCCURRENCE RANKING SCALE IS VERY LOW; THEREFORE, THIS DOES NOT APPEAR TO BE DESIGN-RELATED. EXACTECH IS NOT AWARE OF RECEIVING ANY OTHER COMPLAINT REPORTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF (B)(4) PIECES WHICH HAS BEEN IN THE FIELD SINCE 2014. EXACTECH COMPLAINT DATA FROM 2014 TO 2018 INVOLVING THE REPORTED FAILURE FOR THIS FAMILY OF DEVICES WAS REVIEWED. THE INVESTIGATIONS HAVE NOT IDENTIFIED ANY EVIDENCE THAT A MANUFACTURING ISSUE CAUSED OR CONTRIBUTED TO THIS REPORTED ISSUE. THEREFORE, THIS DOES NOT APPEAR TO BE MANUFACTURING-RELATED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE IS "VERY LOW" AND THE RISK IS CAPTURED IN THE RISK MANAGEMENT REPORT. THE DAMAGED TIBIAL INSERT TRIAL REPORTED IN EXPERIENCE (B)(4) WAS LIKELY THE RESULT OF THE USE OF EXCESSIVE FORCE WHEN ATTEMPTING TO INSERT AND ALIGN THE TIBIAL INSERT TRIAL ON THE TIBIAL TRAY TRIAL. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. THE CHIPPED PART OF THE TRIAL DID NOT FALL INTO THE SURGICAL SITE. AN INVESTIGATION WAS CONDUCTED; THE DAMAGED TIBIAL INSERT TRIAL WAS LIKELY THE RESULT OF THE USE OF EXCESSIVE FORCE WHEN ATTEMPTING TO INSERT AND ALIGN THE TIBIAL INSERT TRIAL ON THE TIBIAL TRAY TRIAL. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00304 | 2022-03-29 | 2019-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRUL FEM TRIAL EXTRACTOR | EXACTECH, INC. |  | 02-029-19-1200 |  | IT WAS REPORTED FROM THE US THAT THE FEMORAL TRIAL EXTRACTOR BROKE DURING SURGERY. THE DEVICE WAS RETURNED. ATTEMPTS WERE MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REP, BUT NO RESPONSES WERE RECEIVED. THERE IS NO OTHER INFORMATION PROVIDED ABOUT THIS DEVICE/EVENT FROM THE CONTACTS. | H10. AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THIS MEDWATCH IS BEING RESCINDED: UPON FURTHER REVIEW OF REPORTABILITY- IF THE MALFUNCTION WERE TO RECUR, IT IS NOT LIKELY TO CAUSE OR CONTRIBUTE TO SERIOUS INJURY OR DEATH. THE DISASSEMBLED PARTS ARE NOT LIKELY TO FALL INTO THE WOUND SITE OR INTERFERE WITH THE SURGICAL IMPLANT PROCEDURE. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00305 | 2022-03-29 | 2022-03-07 | I, n, j, u, r, y | Manufacturer report | TRAPEZOID TIBIAL TRAY SZ 5F/4T | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRAPEZOID TIBIAL TRAY SZ 5F/4T | 204-04-54 |  | AS REPORTED, APPROXIMATELY ELEVEN YEARS POST INITIAL TKA, THIS (B)(6) MALE PATIENT HAD A REVISION SURGERY DUE TO THE TIBIAL TRAY WAS CRACKED DUE TO ASYMMETRIC LOADING. THERE WAS EVIDENCE OF POLY WEAR, PITTING AND OXIDATION AND GROSS WEAR THROUGH POLY INTRAOP. PREOP, PATIENT HAD A BROKEN TIBIAL TRAY WITH FAILURE OF TIBIA TRAY IN VARUS AND LARGE MEDIAL TIBIAL OSSEOLYRIC DEFECT. THE DEVICE IS NOT AVAILABLE FOR EVALUATION AS HE REQUESTED IT AND HAS IN HIS POSSESSION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00306 | 2022-03-29 | 2022-02-07 | I, n, j, u, r, y | Manufacturer report | 2.5 TRULIANT CRC POLYETHYLEN | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | 2.5 TRULIANT CRC POLYETHYLEN | UNKNOWN |  | AS REPORTED, THIS (B)(6) FEMALE PATIENT HAD A REVISION DUE TO THE RECALLED POLYETHYLENE WAS RETRIEVED AND REPLACED WITH A NEW COMPONENT. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICE WILL BE RETURN. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00307 | 2022-03-29 | 2019-04-10 | I, n, j, u, r, y | Manufacturer report | EXACTECH | MDLR RV RMR STRT PLT TIP | EXACTECH, INC. |  | 321-25-01 |  | IT WAS REPORTED FROM OUS THAT DURING A PRIMARY ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE PILOT TIP REAMER. THE SURGEON WAS FOLLOWING THE GPS INSTRUCTIONS, ROTATED THE REAMER AND THE PILOT BROKE. NO PIECES WERE LEFT IN THE PATIENT. IT WAS NOT INDICATED IF ANY PIECES/PARTS HAD TO BE RETRIEVED FROM THE PATIENT. ¿DURING REAMER GLENOID WITH THE PRIMARY NUBBED REAMER, FOLLOWING THE ON-SCREEN GUIDANCE TO REAM THE PROPER ORIENTATION ACCORDING TO THE PLAN, THE CENTRAL PEG BROKE. VERY LIKELY THE CAUSE WAS AN EXCESSIVE FORCE AND OFF AXIS. NO COMPLICATION FOR PATIENT. ONLY A MINIMAL DELAY IN SURGICAL WORKFLOW.¿ NO PROBLEMS NOTED FOR THE PATIENT AT POST OP FOLLOW UP VISIT. THERE WERE 2 REPRESENTATIVES AT THE SURGERY. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED BY THE CONTACTS FOR THIS EVENT. (B)(4). | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE AVAILABLE FOR EVALUATION. (B)(4), CAPA (B)(4), (B)(4), AND (B)(4). FINDINGS: BASED ON THE INVESTIGATION CONDUCTED UNDER CAPA (B)(4), IT WAS DETERMINED THAT THE USER INSTRUCTION (E.G. THE SURGICAL TECHNIQUE) WAS NOT ADEQUATE TO INFORM THE SURGEON OF THE APPROPRIATE USE OF THE REAMER. THE FORESEEABLE MISUSE OF THE GLENOID REAMER THAT COULD RESULT IN FRACTURE OF THE PILOT TIP ARE: IF THE USER APPLIES A BENDING MOMENT TO THE PILOT TIP DURING REAMING, WHICH WOULD REAM THE GLENOID ASYMMETRICALLY AND APPLY A TORQUE TO THE PILOT TIP, POTENTIALLY LEADING TO FRACTURE. IF THE USER APPLIES A BENDING MOMENT TO THE PILOT TIP BY USING THE GLENOID REAMER TO RETRACT THE HUMERAL HEAD TO HELP GAIN EXPOSURE AND ACCESS TO THE GLENOID, WHICH WOULD APPLY A TORQUE TO THE PILOT TIP, POTENTIALLY LEADING TO FRACTURE. CORRECTIVE ACTION TAKEN: AS PART OF CAPA(B)(4), THE FOLLOWING CORRECTIVE ACTIONS WERE TAKEN: THE SURGICAL TECHNIQUE WAS UPDATED TO ADVISE SURGEON TO AVOID APPLYING A BENDING TORQUE TO THE PILOT TIP REAMER OR USING THE REAMER TO RETRACT THE HUMERAL HEAD AS THIS MAY CAUSE FRACTURE OF THE PILOT TIP (B)(4). UPDATE THE RMR TO INCLUDE THE RISK OF PILOT TIP REAMER FRACTURE AS A RESULT OF BENDING OR ITS FORESEEABLE MISUSE AS A RETRACTOR (B)(4). ADDITIONALLY, PER (B)(4), A RECALL COMMUNICATION WAS SENT TO USERS NOTIFYING THEM OF THE ISSUE AND ALERTING THEM TO THE CHANGE IN THE OPERATIVE TECHNIQUE (Z-2663 THROUGH 2668-2017, MAY 2017). MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: BASED ON CAPA (B)(4), THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE REAMER DURING USE, WHICH LED TO BRITTLE FRACTURE OF THE PILOT TIP FEATURE. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. THE PILOT TIP BROKE, IT WAS NOT REPORTED IF ANY PIECES/PARTS HAD TO BE RETRIEVED FROM THE PATIENT. NO PIECES/PARTS WERE LEFT IN THE PATIENT. AN INVESTIGATION WAS CONDUCTED: BASED ON CAPA (B)(4), THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE REAMER DURING USE, WHICH LED TO BRITTLE FRACTURE OF THE PILOT TIP FEATURE. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00308 | 2022-03-29 | 2022-02-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | VANTAGE | PUNCH LINER-FB-SZ 2-8MM | EXACTECH, INC. |  | 351-15-22 |  | AS REPORTED, DURING SURGICAL USE, THE SURGEON USED THE SIZE 2 8MM POLY TO SET ALIGNMENT WITH THE TALUS AND TIBIA TRIALS IN THE 74 YEAR OLD MALE PATIENT. SURGEON PULLED THE POLY OUT WITH A RONJOUR AND THE POLY BROKE. ALL PARTS AND PIECES THAT FALL INTO PATIENT WOUND SITE WERE REMOVED BY THE SURGEON. THE PIECE OF THE POLY WAS REMOVED. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. SALES REP WAS UNABLE TO OBTAIN PHOTOS/X-RAYS. THE DEVICE WILL BE RETURN. | DEVICE EVALUATED BY MFR: PENDING EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00309 | 2022-03-29 | 2022-03-07 | I, n, j, u, r, y | Manufacturer report | LOGIC TIB INSERT IMPL CRC, SZ 5, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED, POLYMER/METAL | EXACTECH, INC. | LOGIC TIB INSERT IMPL CRC, SZ 5, 9MM | 02-012-51-5009 |  | AS REPORTED, THIS MALE PATIENT HAD A REVISION DUE TO THE IMPLANT WAS ONE OF THE RECALLED POLYS. SURGEON REPLACED WORN POLY WITH A NEW POLY. THE POLY WAS SEVERELY DELAMINATED AND PITTED. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICE WILL BE RETURN. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00310 | 2022-03-29 | 2022-03-10 | I, n, j, u, r, y | Manufacturer report | CR TIBIAL INSERT SZ 2, 11MM, SLOPE ++ | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | CR TIBIAL INSERT SZ 2, 11MM, SLOPE ++ | 200-72-11 |  | AS REPORTED, APPROXIMATELY ELEVEN YEARS POST INITIAL TKA, THIS FEMALE PATIENT HAS A SCHEDULED REVISION TO BE DONE FOR AN UNKNOWN REASON. NO ADDITIONAL INFORMATION IS AVAILABLE. | CONCOMITANT MEDICAL PRODUCTS: THREE PEG PATELLA 32MM (CAT# 200-02-32 / SERIAL# (B)(4)). CEMENTED FINNED TIB. TRA SZ 2F/2T (CAT# 200-04-22 / SERIAL# (B)(4)). OPTETRAK ASY,CR CEMENTED FEMORAL, SZ 2 (CAT# 230-03-02 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00311 | 2022-03-29 | 2022-03-03 | I, n, j, u, r, y | Manufacturer report | LOGIC TIB INSERT IMPL CRC, SZ 2.5, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIB INSERT IMPL CRC, SZ 2.5, 9MM | 02-012-51-2509 |  | AS REPORTED BY THE LEGAL TEAM, APPROXIMATELY SEVEN YEARS POST INITIAL TKA, THIS (B)(6) FEMALE PATIENT IS EXPECTED TO HAVE A REVISION SURGERY ON (B)(6) 2022 IN ORDER TO REPLACE THE ERODED EXACTECH ARTIFICIAL CARTILAGE (PLASTIC TIBIAL INSERT). THE PATIENT WAS INITIALLY DIAGNOSED WITH OSTEOPENIA AND SINCE HER INITIAL SURGERY EXPERIENCED NUMEROUS SYMPTOMS, INCLUDING PAIN, SWELLING, DIFFICULTY WALKING, AND WAKING AT NIGHT DUE TO KNEE PAIN. THESE SYMPTOMS STARTED A COUPLE OF YEARS AFTER THE SURGERY. THE SYMPTOMS HAVE SLOWLY AND STEADILY WORSENED. PATIENT MENTIONED THAT AT FIRST, "SHE FOUND THAT SHE COULD NOT WALK ON PAVEMENT WITHOUT PAIN AND SWELLING THAT LASTED AFTERWARD FOR HOURS. THEN, SHE BEGAN TO HAVE THE SAME SYMPTOMS AFTER WALKING ON SMOOTH DIRT PATHS IN THE WOODS". | (H3) THE REVISION REPORTED WAS LIKELY THE RESULT OF THE PATIENT¿S PAIN, WHICH MAY HAVE BEEN DUE TO PROSTHESIS WEAR OF THE TIBIAL INSERT RELATED TO THIRD BODY WEAR AND/OR PATIENT-RELATED CONDITIONS. THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE REPORTED EVENT OF "PROSTHESIS WEAR¿ IS ASSOCIATED WITH MATERIAL DAMAGE TO A SURFACE, USUALLY INVOLVING PROGRESSIVE LOSS OR DISPLACEMENT OF MATERIAL, DUE TO RELATIVE MOTION BETWEEN THAT SURFACE AND A CONTACTING SUBSTANCE OR SUBSTANCES. | Y | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00312 | 2022-03-29 | 2019-08-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | HEAD TRIAL, 12/14 36 +0 | EXACTECH, INC. |  | 143-36-00 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE INSTRUMENTATION. THE TRIAL HEAD (36MM +0 TRIAL HEAD) CRACKED IN HALF WHEN REMOVING FROM HIP. THERE WAS DELAY TO THE SURGERY. THERE WAS NO REPORTED PATIENT INJURY. THE REPRESENTATIVE WAS PRESENT AT THE TIME OF THE SURGERY. THE DEVICE WAS NOT RETURNED AS IT WAS DISPOSED OF BY THE CUSTOMER. ATTEMPTS WERE MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REPRESENTATIVE, BUT NO RESPONSES WERE RECEIVED. NO OTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS ABOUT THE DEVICE/EVENT/PATIENT. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT RELATED TO THE BROKEN TRIAL. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS NOT AVAILABLE FOR ANALYSIS. ENGINEERING ANALYSIS COMPLETED BY (B)(4) ON 02/03/2020. DESIGN-RELATED ISSUES: THE DESIGN OF THE 12/14 FEMORAL HEAD TRIAL, O-RING HAS BEEN IN THE FIELD SINCE 2007. EXACTECH IS AWARE OF 6 COMPLAINT REPORTS INVOLVING BROKEN HEAD TRIALS SINCE 2008. SALES DATA FOR ALL FEMORAL HEADS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4). THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: MANUFACTURING DATA COULD NOT BE REVIEWED BECAUSE THE SERIAL/LOT INFORMATION WAS NOT PROVIDED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) AND RISK ASSESSMENT AND CONTROLS REPORT (RACR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR AND RACR FOR THIS FEMORAL HEAD TRIAL, RMR-00027 REV - AND RACR-00030 REV -, WERE REVIEWED. THE RISK IS CAPTURED, HOWEVER THE RACR IS BEING UPDATED UNDER (B)(4) TO CAPTURE AN APPROPRIATE HAZARDOUS SITUATION FOR THE REPORTED EVENT. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, THE RISK IS CAPTURED IN THE RMR AND RACR, AND THE HAZARDOUS SITUATION WILL BE UPDATED. THE BROKEN FEMORAL HEAD TRIAL REPORTED IN (B)(4) WAS LIKELY THE RESULT OF AN EXPOSURE TO IMPACTION FORCES AND SUBSEQUENT STERILIZATION CYCLES WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE WHEN REMOVING THE HEAD TRIAL. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT RELATED TO THE BROKEN TRIAL. AN INVESTIGATION WAS CONDUCTED: THE BROKEN FEMORAL HEAD TRIAL REPORTED WAS LIKELY THE RESULT OF AN EXPOSURE TO IMPACTION FORCES AND SUBSEQUENT STERILIZATION CYCLES WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE WHEN REMOVING THE HEAD TRIAL. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. INFORMATION WAS NOT PROVIDED FOR: PATIENT AGE, GENDER, WEIGHT, ETHNICITY AND RACE. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10020 | 2022-03-29 | 2018-06-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LOGIC PS TIBIA, MODULAR INSERT TRIAL, SIZE 3, 13MM | EXACTECH, INC. |  | 02-013-35-3013 |  | IT WAS REPORTED FROM OUS THAT DURING AN INITIAL IMPLANT PROCEDURE THE INSERT TRIAL BROKE DURING INSERTION OF THE IMPLANT. THE SURGEON DID THE LAST RANGE OF MOTION VERIFICATION AFTER INSERTION OF THE FEMORAL AND TIBIAL IMPLANTS, WHEN THE INSERT TRIAL WAS REMOVED IT BROKE. THE DEVICE BROKE INTO 2 LARGE PIECES AND WAS BEING HELD BY THE HANDLE, NOTHING FELL INTO THE SURGICAL SITE. THE SURGEON CONFIRMED THAT ALL BROKEN PIECES WERE ACCOUNTED FOR. THE PATIENT'S HEALTH WAS STABLE LEAVING THE OR. THE ISSUE WAS RESOLVED BY USING ANOTHER INSTRUMENT. THERE WAS NO DELAY OF THE PROCEDURE. THE AGENT WAS PRESENT AT THE TIME OF SURGERY. PATIENT INFORMATION WILL NOT BE PROVIDED DUE TO CITED NATIONAL PRIVACY LAWS. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. MFG DATE: 12/29/2012. (B)(4). CAPA (B)(4) -INVESTIGATION STATES: THE CAUSE OF THIS FAILURE COULD BE REPROCESSING OF THESE PARTS. IN ORDER TO UNDERSTAND WHY THESE DEVICES ARE FAILING IN SINGAPORE AND KOREA, A QUESTIONNAIRE WAS SENT TO THE SALES FORCE IN KOREA AND SINGAPORE TO INQUIRE ABOUT TKA TRIALING, GENERAL USE, AND STERILIZATION/REPROCESSING TECHNIQUES. ACCORDING TO AN EMAIL FROM THE SALES MANAGER AT EXACTECH KOREA, THESE INSTRUMENTS ARE BEING STERILIZED FOR 40-50 MINUTES AT 134 DEGREES (CELSIUS). THESE REPROCESSING METHODS ARE OUTSIDE OF EXACTECH'S RECOMMENDED STERILIZATION INSTRUCTIONS PER 700-096-057 - REPROCESSING INSTRUCTIONS FOR REUSABLE SURGICAL INSTRUMENTS, WHICH STATES PARAMETERS FOR STERILIZATION. (B)(4) IS THE INITIAL SCOPE ASSESSMENT. ROOT CAUSE: THE REPROCESSING OF THESE INSTRUMENTS EXCEEDS THE RECOMMENDED PARAMETERS LISTED IN THE REPROCESSING INSTRUCTIONS. THE MATERIAL DEGRADATION FROM EXCESSIVE STERILIZATION OF THESE INSTRUMENTS HAS LIKELY LED TO THE OBSERVED FAILURE. NO ADDITIONAL ACTION REQUIRED. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. AN INVESTIGATION WAS CONDUCTED UNDER CAPA(B)(4); THE REPROCESSING OF THESE INSTRUMENTS EXCEEDS THE RECOMMENDED PARAMETERS LISTED IN THE REPROCESSING INSTRUCTIONS. THE MATERIAL DEGRADATION FROM EXCESSIVE STERILIZATION OF THESE INSTRUMENTS HAS LIKELY LED TO THE OBSERVED FAILURE. NO ADDITIONAL ACTION REQUIRED. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00303 | 2022-03-29 | 2019-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRUL FEM TRIAL EXTRACTOR | EXACTECH, INC. |  | 02-029-19-1200 |  | IT WAS REPORTED FROM THE US THAT THE FEMORAL TRIAL EXTRACTOR BROKE DURING SURGERY. THE DEVICE WAS RETURNED. ATTEMPTS WERE MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REP, BUT NO RESPONSES WERE RECEIVED. THERE IS NO OTHER INFORMATION PROVIDED ABOUT THIS DEVICE/EVENT FROM THE CONTACTS. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS AVAILABLE FOR ANALYSIS. (B)(4) AND CAPA (B)(4). FINDINGS: IN THE CASE OF MULTI-PIECE CONSTRUCTION, IT WAS FOUND THAT THE WELD ON THE EXTRACTOR BODY WAS BREAKING. THIS ALLOWED THE COMPONENT PARTS TO DISASSEMBLE. THE ENGINEERING PRINT FOR THE FEMORAL TRIAL EXTRACTOR (02-029-219-1202 REV A) ALLOWED FOR ¿SINGLE OR MULTIPLE PIECE CONSTRUCTION¿ (NOTE 2). HOWEVER, THE PRINT DID NOT SPECIFY HOW TO WELD/ASSEMBLE THE DEVICE IF MULTIPLE-PIECE CONSTRUCTION WAS TO BE USED. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: BASED ON THE CAPA AND CRC, THE BROKEN DEVICE REPORTED IN (B)(4) WAS LIKELY THE RESULT OF AN INSUFFICIENT WELD ON 02-029-219-1202 REV A, WHICH LED TO WELD FRACTURE AND COMPONENT DISASSEMBLY. CORRECTIVE ACTION TAKEN: PARTS WILL NOT BE ORDERED UNLESS THEY ARE MANUFACTURED AS A SINGLE PIECE CONSTRUCTION. REFER TO CAPA (B)(4). IFU 700-096-181: INSTRUMENT INSPECTION : VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO MANUFACTURING ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. THE DEVICE CAN FRACTURE AND DISASSEMBLE. BASED ON (B)(4) AND CAPA (B)(4), THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF AN INSUFFICIENT WELD ON 02-029-219-1202 REV A, WHICH LED TO WELD FRACTURE AND COMPONENT DISASSEMBLY. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10022 | 2022-03-29 |  | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | IMPACTOR HANDLE | EXACTECH, INC. |  | 321-07-05 |  | IT WAS REPORTED FROM THE US THAT INSTRUMENTS BROKE DURING A CASE. THE INSTRUMENTS WERE RETURNED TO THE COMPANY WITH A NOTE STATING THEY BROKE DURING A CASE. NO ADDITIONAL INFORMATION WAS PROVIDED. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. THE AGENCY IS INACTIVE, NO FURTHER INFORMATION IS AVAILABLE RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICES WERE RECEIVED FOR ANALYSIS. ENG EVAL COMPLETED BY (B)(4). VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR 10X OPTICAL UNLESS OTHERWISE SPECIFIED) THE FRACTURED SURFACES APPEAR TO BE CONSISTENT WITH FATIGUE CRACK PROPAGATION OVER YEARS OF NORMAL USE. ALTHOUGH THE PIN IS STILL ASSEMBLED, THE WELD HOLDING THE PIN IN PLACE APPEARS TO BE CRACKED. THE CRACK AT THE WELD OF THE MATING FEATURE APPEARS CONSISTENT WITH DEGRADATION OF THE LASER WELD OVER TIME. THE TIP APPEARS TO HAVE DISASSEMBLED. THE RETAINING RING IS MEANT TO HOLD THE SLEEVE AND SPRING COMPONENTS IN PLACE BUT APPEARS TO HAVE DISLODGED. THE USER LIKELY HELD THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY AND ALLOWED THE TIP TO DISASSEMBLE. THE BALL BEARING WAS NOT RETURNED. A COMPLAINT COORDINATOR REACHED OUT TO THE AGENCY ON NOVEMBER 14TH, 2018 REGARDING THE BALL BEARING AND NO RESPONSE WAS RECEIVED AS OF JANUARY 11TH, 2019. DESIGN-RELATED ISSUES: THE DESIGN OF THE EQUINOXE REVERSE SHOULDER IMPACTOR HANDLE AND TRI-DRIVE CANNULATED HANDLE HAVE BEEN IN THE FIELD SINCE 2006 AND 2009, RESPECTIVELY. EXACTECH IS AWARE OF 6 OTHER SIMILAR COMPLAINT REPORTS INVOLVING EQUINOXE REVERSE SHOULDER IMPACTOR HANDLES AND 42 OTHER SIMILAR COMPLAINTS INVOLVING TRI-DRIVE CANNULATED HANDLE SINCE 2008. THE FREQUENCY OF OCCURRENCE RANKING SCALE IS "VERY LOW", THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS AWARE OF 2 OTHER COMPLAINTS INVOLVING PARTS FROM THE TRI-DRIVE CANNULATED HANDLE MANUFACTURING LOT OF (B)(4) UNITS, WHICH HAS BEEN IN THE FIELD SINCE 2012. THE DEVICE HISTORY RECORD WAS REVIEWED FOR THE TRI-DRIVE CANNULATED HANDLE, AND ALL PARTS WERE ACCEPTED WITH CONFORMANCE TO THE DEVICE SPECIFICATIONS. EXACTECH IS NOT AWARE OF ANY COMPLAINTS INVOLVING PARTS FROM THE IMPACTOR HANDLE MANUFACTURING LOT OF (B)(4) UNITS, WHICH HAS BEEN IN THE FIELD SINCE 2013. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. THE RMR FOR THIS IMPACTOR HANDLE, 750-2005-021-RMR-INSTRUMENTS REV M, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 15. THE RMR FOR THIS TRI-DRIVE CANNULATED HANDLE, 750-2007-026-RMR-INSTRUMENTS REV J, WAS REVIEWED. THE RISK IS NOT CAPTURED. THIS DOCUMENT WILL BE UPDATED TO INCLUDE THIS FAILURE. REFER TO (B)(4) FOR DETAILS. CORRECTIVE ACTION FOR THE IMPACTOR HANDLE IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. CORRECTIVE ACTION FOR THE TRIDRIVE CANNULATED SHAFT IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RMR WILL BE UPDATED TO CAPTURE THE RISK. THE BROKEN IMPACTOR HANDLE REPORTED IN EXPERIENCE (B)(4) WAS LIKELY THE RESULT OF YEARS OF NORMAL USE, PROCESSING, AND HANDLING, WHICH LED TO WELD FRACTURING AT THE STRIKE PLATE-HANDLE INTERFACE. THE DISASSEMBLED TRI-DRIVE CANNULATED HANDLE REPORTED WAS LIKELY THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY AND ALLOWED THE TIP TO DISASSEMBLE. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE WAS NO PATIENT ADVERSE EVENT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN IMPACTOR HANDLE REPORTED WAS LIKELY THE RESULT OF YEARS OF NORMAL USE, PROCESSING, AND HANDLING, WHICH LED TO WELD FRACTURING AT THE STRIKE PLATE-HANDLE INTERFACE. THE DISASSEMBLED TRI-DRIVE CANNULATED HANDLE REPORTED WAS LIKELY THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY AND ALLOWED THE TIP TO DISASSEMBLE. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. EVENT DATE NOT REPORTED. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10028 | 2022-03-29 | 2017-06-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LOGIC TIBIA PS INSERT, SIZE 2.5, 9MM | EXACTECH, INC. |  | 02-013-35-2509 |  | IT WAS REPORTED THAT DURING AN ORTHOPEDIC KNEE SURGERY THE INSERT TRIAL WAS CHIPPED. IT WAS NOT REPORTED IF THE SALES REPRESENTATIVE WAS PRESENT AT THE SURGERY. THE SURGERY CONTINUED AND WAS COMPLETED WITHOUT ADDITIONAL ISSUES. THERE WAS NO REPORTED DELAY OF SURGERY AND THERE WAS NO REPORTED PATIENT INJURY, ADVERSE EVENT, OR CLINICAL CONSEQUENCE TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE REPORTER RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DESIGN-RELATED ISSUES: THE DESIGN OF THIS TIBIAL INSERT TRIAL HAS BEEN IN THE FIELD SINCE 2009. (B)(4). MFG-RELATED ISSUES: EXACTECH HAS RECEIVED 4 OTHER COMPLAINT REPORTS INVOLVING A PART FROM THIS MANUFACTURING LOT. THE MANUFACTURING LOT CONSISTED OF (B)(4) PIECES AND HAS BEEN IN THE FIELD SINCE 2014. A CORRECTIONS AND REMOVAL COMMITTEE MEETING HAS BEEN HELD TO ADDRESS THIS HIGH MANUFACTURING LOT OCCURRENCE. REFER TO (B)(4). A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS INSERT TRIAL, 750-2006-012-RMR REV A, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 33. THE CHIPPING REPORTED IN EXPERIENCE (B)(4) WAS LIKELY THE RESULT OF AGGRESSIVE USE WHICH LED TO SCRATCHES, DEFORMATION, AND EVENTUAL CHIPPING OF THE ULTEM MATERIAL. EXACTECH IFU 700-096-004 STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED INSTRUMENT MALFUNCTION IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO ANY PATIENT IMPACT. AN INVESTIGATION WAS CONDUCTED THAT THE MOST LIKELY ROOT CAUSE OF THE DEVICE BREAKAGE REPORTED IN THIS EVENT WAS THE RESULT OF AGGRESSIVE USE WHICH LED TO SCRATCHES, DEFORMATION, AND EVENTUAL CHIPPING OF THE ULTEM MATERIAL. ATTEMPTS HAVE BEEN MADE FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. NO ADDITIONAL ATTEMPTS WILL BE MADE FOR INFORMATION REGARDING THIS EVENT. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10034 | 2022-03-29 | 2018-08-15 | I, n, j, u, r, y | Manufacturer report | EXACTECH | MDLR RVS RMR 38MM CNLTD | EXACTECH, INC. |  | 321-35-38 |  | IT WAS REPORTED FOR OUS THAT THE REP RECEIVED AN E-MAIL ON 8/15/2018 SAYING THAT A REVERSE GLENOID TRAY HAD BEEN RETURNED FROM THE STERILIZATION DEPARTMENT WITH THE PILOT TIP SNAPPED OFF OF THE 38MM REAMER - THE TRAY HAD THEREFORE BEEN TAKEN OUT OF CIRCULATION. THE DAMAGED REAMER HAD NOT BEEN NOTICED/REPORTED WHEN THE KIT WAS LAST USED FOR SURGERY AND HAD ONLY BEEN NOTED AFTER WASHING. THE TEAM LEADER HAS CONFIRMED THAT SHE HAS SPOKEN TO THE SURGEONS THAT LAST USED THE KIT TO HIGHLIGHT THE ISSUE AND NEITHER HAD ANY CONCERNS BUT SAID THEY WOULD CHECK THE X-RAYS AS A PRECAUTION. NOTHING ADVERSE HAS SINCE BEEN REPORTED. NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE RECEIVED FOR ANALYSIS. ENG EVAL COMPLETED ON (B)(6) 2018. CAPA(B)(6) AND (B)(4). FINDINGS/ROOT CAUSE: BASED ON THE INVESTIGATION CONDUCTED UNDER CAPA (B)(4), IT WAS DETERMINED THAT THE USER INSTRUCTION WAS NOT ADEQUATE TO INFORM THE SURGEON OF THE APPROPRIATE USE OF THE REAMER. THE FORESEEABLE MISUSE OF THE GLENOID REAMER THAT COULD RESULT IN FRACTURE OF THE PILOT TIP ARE: IF THE USER APPLIES A BENDING MOMENT TO THE PILOT DIP DURING REAMING, WHICH WOULD REAM THE GLENOID ASYMMETRICALLY AND APPLY A TORQUE TO THE PILOT TIP, POTENTIALLY LEADING TO FRACTURE. IF THE USER APPLIES A BENDING MOMENT TO THE PILOT TIP BY USING THE GLENOID REAMER TO RETRACT THE HUMERAL HEAD TO HELP GAIN EXPOSURE AND ACCESS TO THE GLENOID, WHICH WOULD APPLY A TORQUE TO THE PILOT TIP, POTENTIALLY LEADING TO FRACTURE. CORRECTIVE ACTION TAKEN AS PART OF CAPA (B)(4), THE FOLLOWING CORRECTIVE ACTIONS WERE TAKEN: 1) THE SURGICAL TECHNIQUE WAS UPDATED TO ADVISE SURGEONS TO AVOID APPLYING A BENDING TORQUE TO THE PILOT TIP REAMER OR USING THE REAMER TO RETRACT THE HUMERAL HEAD AS THIS MAY CAUSE FRACTURE OF THE PILOT TIP. (DCRSP- 17-135, -136, -137, -138). 2) UPDATE THE RMR TO INCLUDE THE RISK OF PILOT TIP REAMER FRACTURE AS A RESULT OF BENDING OR ITS FORESEEABLE MISUSE AS A RETRACTOR. (DCRSP-17-172). PER (B)(4) A FIELD ADVISORY NOTICE WAS SENT TO THE USERS NOTIFYING THEM OF THE ISSUE AND ALERTING THEM TO THE CHANGE IN THE OPERATIVE TECHNIQUE (EXACTECH RECALL NUMBER 1038671-04/07/2017-003-C, FDA RECALL NUMBERS Z-2663-2017 THROUGH Z-2668-2017, (B)(6) 2017). THE RECALL WAS TERMINATED BY THE FDA ON (B)(6) 2019. CAPA (B)(4) WAS INITIATED TO TRACK REPORTED EVENTS THROUGH 2021 AND RECONVENE THE CRC IF A THRESHOLD OF >35 TOTAL EVENTS OR >3 EVENTS WITH REPORTED PATIENT HARM BE MET. AS A RESULT OF MONITORING UNDER (B)(4) (APRIL 2019) AND CAPA (B)(4), IT WAS NOTED THAT THE THRESHOLD OF EVENTS FOR 2020 HAS BEEN EXCEEDED. THIS ISSUE HAS BEEN RE-ESCALATED FOR EVALUATION. AS DETERMINED BY HHE2020-09-11-01 (B)(4) AND (B)(4), THE RISK FALLS WITHIN THE ACCEPTABLE LEVEL. THE COMMITTEE HAS DETERMINED THAT NO ADDITIONAL FIELD ACTION IS REQUIRED AT THIS TIME. EXACTECH WILL CONTINUE MONITOR THE EVENTS THROUGH (B)(4), WHICH WILL DRIVE ANY FURTHER HHE UPDATES AND/OR CRC ESCALATION. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. THE PATIENTS THAT HAD THE DEVICE USED ON THEM WERE CHECKED BY X-RAY TO ENSURE THAT NO PIECES WERE IN SITU. BASED ON THE INVESTIGATION CONDUCTED UNDER CAPA (B)(6), IT WAS DETERMINED THAT THE USER INSTRUCTION WAS NOT ADEQUATE TO INFORM THE SURGEON OF THE APPROPRIATE USE OF THE REAMER. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10032 | 2022-03-29 | 2017-12-22 | I, n, j, u, r, y | Manufacturer report | EXACTECH | REAMER HEAD, MODULAR REVERSE, PILOT TIP, 38MM | EXACTECH, INC. |  | 321-25-38 |  | IT WAS REPORTED FROM AUSTRALIA THAT A BROKEN MODULAR PILOT TIP REAMER WAS NOTICED TO HAVE A BROKEN TIP. THE MODULAR PILOT DEVICE WAS FOUND DAMAGED IN CSSD DEPARTMENT AFTER SURGERY. IT IS STATED THAT THERE WERE NO DELAYS TO SURGERY AND NO PIECES WERE LEFT IN PATIENT. THE PATIENT WAS STABLE LEAVING THE OPERATING ROOM. NO ADDITIONAL INFORMATION IS AVAILABLE ON THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. MFG DATE: 11/30/15 DEVICE WAS RETURNED FOR ANALYSIS. (B)(4) - PILOT TIP REVERSE REAMER BREAKAGE. BASED ON THE INVESTIGATION CONDUCTED UNDER (B)(4), IT WAS DETERMINED THAT THE USER INSTRUCTION (E.G. THE SURGICAL TECHNIQUE) WAS NOT ADEQUATE TO INFORM THE SURGEON OF THE APPROPRIATE USE OF THE REAMER. THE FORESEEABLE MISUSE OF THE GLENOID REAMER THAT COULD RESULT IN FRACTURE OF THE PILOT TIP ARE: IF THE USER APPLIES A BENDING MOMENT TO THE PILOT TIP DURING REAMING, WHICH WOULD REAM THE GLENOID ASYMMETRICALLY AND APPLY A TORQUE TO THE PILOT TIP, POTENTIALLY LEADING TO FRACTURE. IF THE USER APPLIES A BENDING MOMENT TO THE PILOT TIP BY USING THE GLENOID REAMER TO RETRACT THE HUMERAL HEAD TO HELP GAIN EXPOSURE AND ACCESS TO THE GLENOID, WHICH WOULD APPLY A TORQUE TO THE PILOT TIP, POTENTIALLY LEADING TO FRACTURE. THE RISK DOCUMENTATION WAS LACKING SPECIFIC FAILURES FOR THIS TYPE OF DEFECT FOR THESE INSTRUMENTS. IN ADDITION, THE SURGICAL TECHNIQUE WAS NOT PROVIDING INFORMATION TO THE USER OF THE INSTRUMENT THAT REAMING OFF AXIS CAN DAMAGE OR BREAK THE INSTRUMENTS. ADDITIONALLY, THE SURGICAL TECHNIQUE DID NOT PROVIDE ADEQUATE INSTRUCTION TO THE USER. AS PART OF (B)(4), THE FOLLOWING CORRECTIVE ACTIONS WERE TAKEN: DCR-SP-17-135, -136, -137, -138 SURGICAL TECHNIQUE GUIDE UPDATE AND Z-2663 THROUGH 2668-2017 FOR CHANGE IN THE OP TECH. THE DEVICE REPORTED IN THIS COMPLAINT WAS MANUFACTURED AND/OR USED AFTER IMPLEMENTATION OF CORRECTIVE ACTIONS. THE ISSUE WAS RE-ESCALATED UNDER (B)(4). THE CRC DETERMINED THAT NO FURTHER ACTION WAS REQUIRED AT THIS TIME AS THIS ISSUE WILL BE ADDRESSED VIA THE EQUINOXE ERGO INSTRUMENTATION REDESIGN PROJECT. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. REPORTED REAMER PILOT TIP HAD THE TIP BROKEN OFF. BASED ON THE INVESTIGATION UNDER CAPA2017-12, THE RISK DOCUMENTATION WAS LACKING SPECIFIC FAILURES FOR THIS TYPE OF DEFECT FOR THESE INSTRUMENTS. IN ADDITION, THE SURGICAL TECHNIQUE WAS NOT PROVIDING INFORMATION TO THE USER OF THE INSTRUMENT THAT REAMING OFF AXIS CAN DAMAGE OR BREAK THE INSTRUMENTS. ADDITIONALLY, THE SURGICAL TECHNIQUE DID NOT PROVIDE ADEQUATE INSTRUCTION TO THE USER. THE ISSUE WAS RE-ESCALATED UNDER (B)(4). THE CRC DETERMINED THAT NO FURTHER ACTION WAS REQUIRED AT THIS TIME AS THIS ISSUE WILL BE ADDRESSED VIA THE EQUINOXE ERGO INSTRUMENTATION REDESIGN PROJECT. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00299 | 2022-03-29 | 2019-07-22 | I, n, j, u, r, y | Manufacturer report | EXACTECH | TRULIANT PS CEM FEM PS CEM RIGHT SZ 4 | EXACTECH, INC. |  | 02-020-11-0340 |  | IT WAS REPORTED FROM THE US THAT A PATIENT HAD A SURGICAL REVISION DUE TO A NICKEL ALLERGY REPORTED IN MFR 1038671-2019-00404. DURING THE REVISION SURGERY, THE SURGEON, REMOVED THE FEMORAL COMPONENT AND REPORTED NO CEMENT WAS ADHERED TO THE FEMUR. ACCORDING TO THE INFORMATION PROVIDED, A CEMENT MANTLE WAS LEFT ON THE FEMUR. THE SURGEON REPORTED HE IS WORRIED THERE IS A PROBLEM WITH THE UNDERSIDE COATING ON THE FEMORAL COMPONENT. THE TRACKING INFORMATION INDICATES THAT THE EXPLANT WAS DELIVERED, HOWEVER THE DEVICE WAS NOT RECEIVED BY THE COMPLAINTS DEPARTMENT. THEREFORE, THE DEVICE CANNOT BE ANALYZED, AND THE FAILURE CANNOT BE CONFIRMED. THERE WAS NO DELAY TO SURGERY. INACTIVE AGENCY WITH NO HOSPITAL/OTHER CONTACT INFORMATION FOR THE EVENT. NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS NOT AVAILABLE FOR ANALYSIS. THE TRACKING INFORMATION INDICATES THAT THE EXPLANT WAS DELIVERED, HOWEVER THE DEVICE WAS NOT RECEIVED BY THE COMPLAINTS DEPARTMENT. THE DEVICE CANNOT BE ANALYZED, AND THE FAILURE CANNOT BE CONFIRMED. BASED ON PREVIOUS SIMILAR INVESTIGATIONS, FEMORAL COMPONENT IDENTIFIED WITH MINIMAL, OR NO CEMENT ADHERED TO THE UNDERSIDE MAY BE THE RESULT OF FEMORAL LOOSENING. PROPERLY FUNCTIONING IMPLANTS DEPEND ON THEIR APPROPRIATE FIXATION TO THE BONE; FIXATION IS USUALLY ACHIEVED BY CEMENTING THE IMPLANT ONTO THE BONE. ALTHOUGH IMPLANTS ARE FIRMLY FIXED AT THE INITIAL KNEE REPLACEMENT SURGERY, THEY MAY BECOME LOOSE OVER TIME. FRICTION CAUSED BY THE JOINT SURFACES RUBBING AGAINST EACH OTHER WEARS AWAY THE SURFACES OF THE IMPLANT, CREATING TINY PARTICLES THAT ACCUMULATE AROUND THE JOINT. IN A PROCESS CALLED ASEPTIC (NON-INFECTED) LOOSENING, THE BOND OF THE IMPLANT TO THE BONE IS DESTROYED BY THE BODY'S ATTEMPT TO DIGEST THE WEAR PARTICLES. DURING THIS PROCESS, NORMAL BONE IS ALSO DIGESTED (A CONDITION CALLED OSTEOLYSIS), WHICH CAN WEAKEN OR EVEN FRACTURE THE BONE. WHEN THE PROSTHESIS BECOMES LOOSE, THE PATIENT MAY EXPERIENCE PAIN, CHANGE IN ALIGNMENT, OR INSTABILITY. DESIGN-RELATED ISSUES: AS OF (B)(6) 2020, EXACTECH IS AWARE OF 264 COMPLAINT REPORTS INVOLVING REVISIONS OF KNEE COMPONENTS DUE TO FEMORAL OR NON-SPECIFIC LOOSENING (31 FEMORAL, 233 NON-SPECIFIC) SINCE 2008. SALES DATA FOR FEMORAL COMPONENTS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 14 UNITS. ACCORDING TO THE SALES DATA, ALL PIECES OF THIS LOT HAVE BEEN SOLD AND NO OTHER REVISIONS HAVE BEEN REPORTED TO EXACTECH, AS OF (B)(6) 2020. THE MANUFACTURING RECORDS FOR THIS DEVICE ONLY EXIST IN PAPER FORM AND ARE STORED IN CARDBOARD BOXES AT IRON MOUNTAIN. DUE TO RISKS ASSOCIATED WITH COVID-19 TRANSMISSION, THE PHYSICAL MANUFACTURING RECORDS FOR THIS COMPLAINT WILL NOT BE REVIEWED. EXACTECH WILL CONTINUE TO MONITOR THIS COMPLAINT TYPE AND, IF WARRANTED, WILL ESCALATE ANY POSITIVE TRENDS TO THE CORRECTIONS AND REMOVAL COMMITTEE (CRC) FOR ADDITIONAL ACTIONS. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE LOOSENING IS LISTED IN THE PRODUCT LABELING (IFU 700-096-160 (B)(4)), THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR AND RACR. THE REVISION REPORTED WAS LIKELY THE RESULT OF ASEPTIC (NONINFECTED) FEMORAL LOOSENING. THE LACK OF BONE CEMENT ADHERED TO THE FEMORAL COMPONENT UPON REMOVAL MAY HAVE BEEN THE RESULT OF THE CEMENTING TECHNIQUE USED OR ENVIRONMENTAL CONDITIONS IN THE OPERATING ROOM DURING IMPLANTATION, OR MICROMOTION OF THE FEMORAL COMPONENT RELATIVE TO THE CEMENT MANTLE IN SITU. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION, AND THERE WERE NO DETAILS REGARDING CEMENTING TECHNIQUE OR ENVIRONMENTAL CONDITIONS PROVIDED. IFU- AS PART OF THE PREOPERATIVE ASSESSMENT, THE SURGEON MUST ENSURE THAT NO BIOLOGICAL, BIOMECHANICAL, OR OTHER FACTORS EXIST THAT MIGHT ADVERSELY AFFECT THE SURGERY AND/OR POSTOPERATIVE PERIOD. REVISIONS OR SURGICAL INTERVENTIONS ARE A KNOWN COMPLICATION FOUND IN JOINT REPLACEMENTS. IT IS KNOWN THAT DEVICE SPECIFIC RISKS CONSIST OF: FRACTURE, MIGRATION, LOOSENING, SUBLUXATION, OR DISLOCATION OF THE PROTHESIS OR ANY OF ITS COMPONENTS, ANY OF WHICH MAY REQUIRE A SECOND SURGICAL INTERVENTION OR REVISION. AS PART OF THE PREOPERATIVE ASSESSMENT, THE SURGEON MUST ENSURE THAT NO BIOLOGICAL, BIOMECHANICAL, OR OTHER FACTORS EXIST THAT MIGHT ADVERSELY AFFECT THE SURGERY AND/OR POSTOPERATIVE PERIOD. REVISIONS OR SURGICAL INTERVENTIONS ARE A KNOWN COMPLICATION FOUND IN JOINT REPLACEMENTS IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THE REVISION REPORTED WAS DUE TO THE PATIENT NICKEL ALLERGY IN MFR 1038671-2019-00404. DURING THE REVISION THE SURGEON REMOVED THE FEMORAL COMPONENT AND REPORTED NO CEMENT WAS ADHERED TO THE FEMUR. ASEPTIC (NONINFECTED) FEMORAL LOOSENING. THE LACK OF BONE CEMENT ADHERED TO THE FEMORAL COMPONENT UPON REMOVAL MAY HAVE BEEN THE RESULT OF THE CEMENTING TECHNIQUE USED OR ENVIRONMENTAL CONDITIONS IN THE OPERATING ROOM DURING IMPLANTATION, OR MICROMOTION OF THE FEMORAL COMPONENT RELATIVE TO THE CEMENT MANTLE IN SITU. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION, AND THERE WERE NO DETAILS REGARDING CEMENTING TECHNIQUE OR ENVIRONMENTAL CONDITIONS PROVIDED. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00300 | 2022-03-29 | 2022-03-03 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REPLICATOR PLATE 1.5MM O/S | EXACTECH, INC. | 300-10-15 | 300-10-15 |  | AS REPORTED, APPROXIMATELY 7 MONTHS POST OP THE INITIAL RTSA, THIS (B)(6) FEMALE PATIENT WAS REVISED TO A REVERSE FROM A HEMI. REASON HAS NOT BEEN REPORTED. THERE WAS NO SURGICAL DELAY. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETUNING, HOSPITAL DISPOSED OF DEVICES. | PENDING EVALUATION: CONCOMITANT DEVICE(S): 300-20-02, 6359582 - EQUINOX SQUARE TORQUE DEFINE SCREW DRIVE KIT. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00301 | 2022-03-29 | 2022-03-09 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | HUMERAL STEM PRIMARY, PRESS FIT 11MM | EXACTECH, INC. | 300-01-11 | 300-01-11 |  | AS REPORTED, APPROXIMATELY 3 YEARS POST OP THE INITIAL LEFT TSA, THIS (B)(6) MALE PATIENT'S WAS REVISED DUE A PAINFUL SHOULDER WITH RADIOGRAPHIC LOOSE STEM. HUMERAL LINER HAD DISASSOCIATED, CREATING METALLOSIS AND OSTEOLYSIS. ALL IMPLANTS WERE REMOVED. THERE WAS NOT DELAY. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETUNING, HOSPITAL DISPOSED OF IMPLANTS. | PENDING EVALUATION. CONCOMITANT DEVICE(S): 320-15-04, (B)(4) - RS GLENOID PLATE POST AUG, 8 DEG, RIGHT, 320-15-05, (B)(4) - EQ REV LOCKING SCREW, 320-01-42, (B)(4) - EQUINOXE REVERSE 42MM GLENOSPHERE, 320-10-00, (B)(4) - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0, 320-20-00, (B)(4) - EQ REVERSE TORQUE DEFINING SCREW KIT, 320-42-03, (B)(4) - EQUINOXE REVERSE 42MM HUMERAL LINER +2.5,AND 320-42-00, (B)(4) - EQUINOXE REVERSE 42MM HUMERAL LINER +0. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10006 | 2022-03-29 | 2017-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | HIP INSTRUMENT | EXACTECH, INC. |  | SK-000527-2 |  | IT WAS REPORTED BY THE EXACTECH REP THAT "THE INSTRUMENT IS BROKEN. ALL PIECES WERE RETURNED. THERE WAS NO ADVERSE ACTION TO THE PATIENT, SINCE THEY WERE ABLE TO GRAB ANOTHER HANDLE, IT WAS OKAY. (B)(4) WAS PRESENT AT THE SURGERY. THERE WAS ONLY A 1-2-MINUTE DELAY TO GRAB ANOTHER HANDLE AND ALL PIECES OF THE BROKEN INSTRUMENT WERE ACCOUNTED FOR. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DESIGN-RELATED ISSUES: THE DESIGN OF THE NOVATION A+ 30° BROACH HANDLE HAS BEEN IN THE FIELD SINCE 2013. EXACTECH IS NOT AWARE OF ANY OTHER RELATED COMPLAINTS DATING BACK TO 2014. (B)(4). THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 3 UNITS, WHICH HAVE BEEN IN THE FIELD SINCE 2014. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING-RELATED. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE AND CONSISTENT WITH YEARS OF USE. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS NOVATION A+ 30° BROACH HANDLE, 750-2002-036A-RMR-INSTRUMENTS REV H, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 12. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. THE FAILURE REPORTED WAS MOST LIKELY THE RESULT OF EXCESSIVE FORCES APPLIED TO THE RELEASE LATCH LEADING TO FAILURE OF THE ROLL PIN. THE DESIGN OF THIS DEVICE WAS CHANGED IN (B)(4). BASED ON TM-2013-264, TESTING DEMONSTRATED THAT THE SMALLER ROLL PIN INCREASED THE FRACTURE RESISTANCE OF THE LATCH PIN UNDER RIGOROUS IMPACT LOADING CONDITIONS. HOWEVER, THE TEST ALSO SHOWED THAT WITH EXCESSIVE FORCE THE ROLL PIN CAN FAIL. EXACTECH IFU 700-096-004 STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED INSTRUMENT MALFUNCTION IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO ANY REPORTED PATIENT IMPACT. AN INVESTIGATION WAS CONDUCTED THAT THE MOST LIKELY ROOT CAUSE OF THE DISASSEMBLY REPORTED IN THIS EVENT WAS LIKELY DUE TO EXCESSIVE FORCE APPLIED TO THE RELEASE LATCH LEADING TO FAILURE OF THE ROLL PIN. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10027 | 2022-03-29 | 2017-02-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | EXTRA-LARGE REAMER | EXACTECH, INC. |  | 315-35-15 |  | IT WAS REPORTED THAT DURING AN ORTHOPEDIC SHOULDER SURGERY THE SURGEON ATTEMPTED TO USE THE CANNULATED SYSTEM TO SET K-WIRE. THE WIRE BOUND IN THE GUIDE BECAUSE THEY COULD NOT GET STRAIGHT ON ACCESS (K-WIRE AND GUIDE PART NUMBERS UNKNOWN). THE SURGEON BROKE THE LARGE REAMER DURING USE BECAUSE OF LEVERING AGAINST THE HUMERUS (315-35-14) AND BROKE THE X-LARGE REAMER (315-35-15). THE SALES REPRESENTATIVE WAS PRESENT AT THE SURGERY AND REPORTED NO DEVICE FRAGMENTS FELL INTO THE SURGICAL SITE. THE SURGERY CONTINUED AND WAS COMPLETED WITHOUT ADDITIONAL ISSUES. THERE WAS A MINOR DELAY OF SURGERY AND THERE WAS NO REPORTED PATIENT INJURY, ADVERSE EVENT, OR CLINICAL CONSEQUENCE TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE REPORTER RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICES IN QUESTION WERE NOT RETURNED TO EXACTECH FOR ANALYSIS. THE BROKEN INSTRUMENTS REPORTED WERE LIKELY THE RESULT OF INADEQUATE EXPOSURE OF THE JOINT, PREVENTING "STRAIGHT ON" ACCESS FOR THE K-WIRE, GUIDE, AND REAMERS. EXACTECH IFU 700-096-004 STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THE DEVICES REPORTED IN THIS EVENT ARE USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED INSTRUMENTS MALFUNCTION IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO ANY REPORTED PATIENT IMPACT. AN INVESTIGATION WAS CONDUCTED AND CONCLUDED THAT THE LIKELY CAUSE OF THE ISSUES WAS THE RESULT OF INADEQUATE EXPOSURE OF THE JOINT, PREVENTING "STRAIGHT ON" ACCESS FOR THE K-WIRE, GUIDE, AND REAMERS. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10023 | 2022-03-29 | 2017-02-03 | I, n, j, u, r, y | Manufacturer report | EXACTECH | POSTERIOR AUGMENT CAGE GLENOID | EXACTECH, INC. |  | 314-13-XX |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SHOULDER SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH GLENOID CAGE. THE SURGEON ATTEMPTED TO INSERT THE POSTERIOR AUGMENT CAGE GLENOID, BUT COULD NOT GET IT SEATED. CONTINUAL IMPACTION WITHOUT SUCCESS AND LEFT THE IMPLANT SITTING PROUD. WHEN TRYING TO REDUCE THE HUMERAL HEAD, THE PLASTIC DISENGAGED FROM THE CAGE OF THE GLENOID. THERE WAS A DELAY TO SURGERY DUE TO THE ISSUE WITH NO EFFECT TO THE PATIENT. ALL PIECES OF THE DEVICE WERE ACCOUNTED FOR. NOTHING WAS LEFT IN THE WOUND SITE. THE STAFF USED ANOTHER GLENOID CAGE. PATIENT WAS DOING WELL AT THE END OF THE SURGERY. THE DEVICE WAS NOT RECEIVED FOR ANALYSIS. PARTS AND LOT/SN WERE NOT PROVIDED. PATIENT DEMOGRAPHIC INFORMATION NOT PROVIDED. INACTIVE AGENCY WITH NO HOSPITAL/OTHER CONTACT INFORMATION FOR THE EVENT. NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE(S) NOT AVAILABLE FOR ANALYSIS. PART/LOT/SN WERE NOT PROVIDED. ENGINEERING EVAL BY AD, (B)(6) 2017. THE GLENOID DISASSEMBLY REPORTED IN EXPERIENCE (B)(6) WAS LIKELY THE RESULT OF INADEQUATE EXPOSURE OF THE JOINT, LEADING TO SEATING DIFFICULTY AND THE IMPLANT BEING LEFT PROUD, WHICH ALLOWED FOR THE POLYETHYLENE TO DISENGAGE FROM THE CAGE OF THE GLENOID UPON REDUCTION OF THE HUMERAL HEAD. EXACTECH IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED INSTRUMENT MALFUNCTION IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. WHEN TRYING TO REDUCE THE HUMERAL HEAD, THE PLASTIC DISENGAGED FROM THE CAGE OF THE GLENOID. ALL PARTS WERE ACCOUNTED FOR. AN INVESTIGATION WAS CONDUCTED: THE GLENOID DISASSEMBLY REPORTED WAS LIKELY THE RESULT OF INADEQUATE EXPOSURE OF THE JOINT, LEADING TO SEATING DIFFICULTY AND THE IMPLANT BEING LEFT PROUD, WHICH ALLOWED FOR THE POLYETHYLENE TO DISENGAGE FROM THE CAGE OF THE GLENOID UPON REDUCTION OF THE HUMERAL HEAD. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10008 | 2022-03-29 | 2017-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LPI IMPACTOR HANDLE | EXACTECH, INC. |  | 213-65-00 |  | IT WAS REPORTED THAT IT WAS REPORTED THAT DURING AN ORTHOPEDIC KNEE SURGERY THE LPI IMPACT HANDLE BROKE. THE SALES REPRESENTATIVE WAS NOT PRESENT AT THE SURGERY. IT WAS REPORTED NO DEVICE FRAGMENTS FELL INTO THE SURGICAL SITE. THE SURGERY CONTINUED AND WAS COMPLETED WITHOUT ADDITIONAL ISSUES. THERE WAS NO REPORTED DELAY OF SURGERY AND THERE WAS NO REPORTED PATIENT INJURY, ADVERSE EVENT, OR CLINICAL CONSEQUENCE TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE REPORTER RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE IN QUESTION WAS RETURNED TO EXACTECH FOR ANALYSIS. DESIGN-RELATED ISSUES: THE LPI IMPACTOR HANDLE HAS BEEN ON THE MARKET SINCE 2011. AT THE TIME OF THE ENGINEERING EVALUATION, (B)(6) 2019 EXACTECH WAS AWARE OF 35 COMPLAINTS INVOLVING THE HANDLE SINCE 2014. BECAUSE THIS DEVICE IS USED IMPLANT TIBIAL TRAYS, SALES DATA FOR TIBIAL TRAYS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS AWARE OF NO OTHER COMPLAINT REPORTS INVOLVING PARTS FROM THIS MANUFACTURING LOT. THE DEVICE HISTORY RECORD COULD NOT BE FOUND FOR REVIEW. THEREFORE, IT CANNOT BE DETERMINED WHETHER OR NOT THIS ISSUE IS MANUFACTURING-RELATED. RISK MANAGEMENT: A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESH OLD. THE RMR FOR THIS LPI IMPACTOR HANDLE, 750-2003-019- RMR-INSTRUMENTS REV-, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 23. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE MISSING SPRING FROM THE LPI IMPACTOR HANDLE REPORTED WAS LIKELY THE RESULT OF PULLING THE TIBIAL TRAY OR TIBIAL TRIALS OFF OF THE IMPACTOR HANDLE AT AN ANGLE AND/OR WITH MORE FORCE THAN NECESSARY. EXACTECH IFU 700-096-004 STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED INSTRUMENT MALFUNCTION IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO ANY REPORTED PATIENT IMPACT. AN INVESTIGATION WAS CONDUCTED THAT THE MOST LIKELY ROOT CAUSE OF THE MISSING SPRING FROM THE LPI IMPACTOR HANDLE REPORTED WAS THE RESULT OF PULLING THE TIBIAL TRAY OR TIBIAL TRIALS OFF OF THE IMPACTOR HANDLE AT AN ANGLE AND/OR WITH MORE FORCE THAN NECESSARY. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10009 | 2022-03-29 | 2017-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LPI IMPACTOR HANDLE | EXACTECH, INC. |  | 213-65-00 |  | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED INSTRUMENT MALFUNCTION IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO ANY REPORTED PATIENT IMPACT. AN INVESTIGATION WAS CONDUCTED AND CONCLUDED THAT THE MOST LIKELY CAUSE OF THE LPI IMPACTOR HANDLE DEFORMED SPRING WAS THE RESULT OF ATTEMPTING TO REMOVE THE BAL-SPRING DURING CLEANING, LEADING TO THE DEFORMATION OF THE SPRING COMPONENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE IN QUESTION WAS RETURNED TO EXACTECH FOR ANALYSIS. DESIGN-RELATED ISSUES: THE DESIGN OF THIS LPI IMPACTOR HANDLE HAS BEEN IN THE FIELD SINCE 2009. AT THE TIME OF THE (B)(4). THE DEVICE HISTORY RECORD WAS REVIEWED, AND ALL PARTS WERE ACCEPTED WITH CONFORMANCE TO PRINT SPECIFICATIONS. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. RISK MANAGEMENT: A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESH OLD. THE RMR FOR THIS LPI IMPACTOR HANDLE, 750-2003-019- RMR-INSTRUMENTS REV-, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 23. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE MISSING SPRING FROM THE LPI IMPACTOR HANDLE DEFORMED SPRING WAS THE RESULT OF ATTEMPTING TO REMOVE THE BAL-SPRING DURING CLEANING, LEADING TO THE DEFORMATION OF THE SPRING COMPONENT. EXACTECH IFU 700-096-004 STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED INSTRUMENT MALFUNCTION IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO ANY REPORTED PATIENT IMPACT. AN INVESTIGATION WAS CONDUCTED AND CONCLUDED THAT THE MOST LIKELY CAUSE OF THE LPI IMPACTOR HANDLE DEFORMED SPRING WAS THE RESULT OF ATTEMPTING TO REMOVE THE BAL-SPRING DURING CLEANING, LEADING TO THE DEFORMATION OF THE SPRING COMPONENT. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10004 | 2022-03-29 | 2017-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LPI IMPACTOR HANDLE | EXACTECH, INC. |  | 213-65-00 |  | IT WAS REPORTED THAT IT WAS REPORTED THAT DURING AN ORTHOPEDIC KNEE SURGERY THE LPI IMPACT HANDLE BROKE. THE SALES REPRESENTATIVE WAS NOT PRESENT AT THE SURGERY. IT WAS REPORTED NO DEVICE FRAGMENTS FELL INTO THE SURGICAL SITE. THE SURGERY CONTINUED AND WAS COMPLETED WITHOUT ADDITIONAL ISSUES. THERE WAS NO REPORTED DELAY OF SURGERY AND THERE WAS NO REPORTED PATIENT INJURY, ADVERSE EVENT, OR CLINICAL CONSEQUENCE TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE REPORTER RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE IN QUESTION WAS RETURNED TO EXACTECH FOR ANALYSIS. DESIGN-RELATED ISSUES: THE LPI IMPACTOR HANDLE HAS BEEN ON THE MARKET SINCE 2011. (B)(4). THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS AWARE OF 3 OTHER COMPLAINT REPORTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF (B)(4) UNITS AND THE RESULTS FROM THE INSPECTION OF ONE OF THOSE 3 DEVICES SHOWED THAT DRAWING ZONES 213-365-01 A-3 AND 213-365-01 A-2 WERE OUTSIDE SPECIFICATIONS. HOWEVER, A REVIEW OF THE DEVICE HISTORY RECORD (DHR) SHOWED THAT THE NAMED DEVICES WERE ACCEPTED WITH CONFORMANCE TO THE DEVICE REQUIREMENTS. RISK MANAGEMENT: A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESH OLD. THE RMR FOR THIS LPI IMPACTOR HANDLE, 750-2003-019-RMR-INSTRUMENTS REV-, WAS REVIEWED. THE RISK IS CAPTURED IN LINES 18 AND 23. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE BROKEN LPI IMPACTOR HANDLE REPORTED IN EXPERIENCE (B)(4) WAS LIKELY THE RESULT OF DEFORMATION FROM MULTIPLE IMPACTIONS OVER THE COURSE OF OVER 6 YEARS OF USE. EXACTECH IFU 700-096-004 STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED INSTRUMENT MALFUNCTION IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO ANY REPORTED PATIENT IMPACT. AN INVESTIGATION WAS CONDUCTED THAT THE MOST LIKELY ROOT CAUSE OF THE DISASSEMBLY REPORTED IN THIS EVENT WAS LIKELY THE RESULT OF DEFORMATION FROM MULTIPLE IMPACTIONS OVER THE COURSE OF OVER 6 YEARS OF USE. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10003 | 2022-03-29 | 2017-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LPI IMPACTOR HANDLE | EXACTECH, INC. |  | 213-65-00 |  | IT WAS REPORTED THAT IT WAS REPORTED THAT DURING AN ORTHOPEDIC KNEE SURGERY THE LPI IMPACT HANDLE BROKE. THE SALES REPRESENTATIVE WAS NOT PRESENT AT THE SURGERY. IT WAS REPORTED NO DEVICE FRAGMENTS FELL INTO THE SURGICAL SITE. THE SURGERY CONTINUED AND WAS COMPLETED WITHOUT ADDITIONAL ISSUES. THERE WAS NO REPORTED DELAY OF SURGERY AND THERE WAS NO REPORTED PATIENT INJURY, ADVERSE EVENT, OR CLINICAL CONSEQUENCE TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE REPORTER RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE IN QUESTION WAS RETURNED TO EXACTECH FOR ANALYSIS. DESIGN-RELATED ISSUES: THE DESIGN OF THIS LPI IMPACTOR HANDLE HAS BEEN IN THE FIELD SINCE 2009. (B)(4). THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT, WHICH HAS BEEN IN THE FIELD SINCE 2011. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING-RELATED. RISK MANAGEMENT: A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESH OLD. THE RMR FOR THIS LPI IMPACTOR HANDLE, 750-2003-019- RMR-INSTRUMENTS REV-, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 23. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE DEFORMED DEVICE REPORTED WAS LIKELY THE RESULT OF MULTIPLE APPLICATIONS OF HIGH FORCE DURING NORMAL USE. EXACTECH IFU 700-096-004 STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED INSTRUMENT MALFUNCTION IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO ANY REPORTED PATIENT IMPACT. AN INVESTIGATION WAS CONDUCTED THAT THE MOST LIKELY ROOT CAUSE OF THE DISASSEMBLY REPORTED IN THIS EVENT WAS LIKELY DUE TO MULTIPLE SURGERIES, THE LPI IMPACTOR BREAKS DUE TO MULTIPLE APPLICATIONS OF HIGH FORCE DURING NORMAL USE. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10000 | 2022-03-29 | 2017-01-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LOGIC FIT TRAY ADJ TIBIAL PUNCH HANDLE | EXACTECH, INC. |  | 213-65-00 |  | IT WAS REPORTED FROM EXACTECH KOREA THAT DURING AN ORTHOPEDIC KNEE REPLACEMENT SURGICAL PROCEDURE THE WELDED PARTS OF THE LPI HANDLE WERE BROKEN DURING AN IMPLANT PROCEDURE. ANOTHER HANDLE WAS USED TO FINISH THE PROCEDURE. THE SALES REPRESENTATIVE WAS PRESENT AT THE SURGERY AND REPORTED NO DEVICE FRAGMENTS FELL INTO THE SURGICAL SITE. THERE WAS A SLIGHT DELAY IN THE SURGERY, BUT THE PATIENT DID NOT EXPERIENCE ANY ADVERSE EFFECTS. NO ADDITIONAL INFORMATION TO FURTHER INVESTIGATE THIS EVENT WAS PROVIDED BY THE CONTACTS. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED FOR ANALYSIS. THE FAILURE MODE IS THE SAME AS INVESTIGATED IN PREVIOUS EVALUATION (B)(4) WHICH DETERMINED THAT, DUE TO THE LOW RATE OF OCCURRENCE, FIELD ACTION IS NOT REQUIRED AT THIS TIME (SEPTEMBER 25, 2018). DESIGN: THE LPI HANDLE ASSEMBLES WITH TWO END PIECES FOR USE THREE TIMES DURING THE SURGICAL PROCEDURE: 1) IMPACTION OF THE FEMORAL TRIAL 2) IMPACTION OF THE TIBIAL TRAY IMPLANT AND 3) IMPACTION OF THE FEMORAL IMPLANT. MANUFACTURING: AVAILABLE DEVICES FROM PREVIOUSLY REPORTED COMPLAINTS WERE EVALUATED. COMPLAINT DEVICES APPEARED TO HAVE SIMILAR FAILURE MODES, AND EITHER WERE CRACKED OR FRACTURED. CRACKS WERE CONSISTENT WITH STARTING INSIDE THE PLASTIC HANDLE AT THE BASE OF ONE OF THE TWO METAL ATTACHMENT FEATURES AND SPREADING OUTWARD. CRACKS WERE EITHER LINEAR, ALONG THE HANDLE AXIS OR HORIZONTAL, NEARLY IN LINE WITH THE ORIGIN. THE REVIEW OF THE LPI IMPACTOR HANDLES (213-65-00) CITED IN COMPLAINTS FOR BROKEN OR CRACKED HANDLES SHOWED A CONSISTENT FAILURE MODE. A REVIEW OF THE DEVICE ORDER HISTORY DID NOT SHOW A TREND IN THOSE LOTS WITH COMPLAINT DEVICES, NOR DID A REVIEW OF DESIGN CHANGES SHOW CHANGES MADE TO THE DESIGN WHICH COULD CONTRIBUTE TO THE COMPLAINT. RISK MANAGEMENT: A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) AND RISK ASSESSMENT AND CONTROLS REPORT (RACR) WAS CONDUCTED TO CONFIRM THE RISK WAS INCLUDED AND THE RATE OF OCCURRENCE IS BELOW THE PREDICTED THRESHOLD. PER 750-2009-023-RMR-INSTRUMENT, REV B - DEVICE FRACTURE/BREAKAGE (LINE ITEM 18) CAN LEAD TO A SURGICAL DELAY. THE PREDICTED SEVERITY OF THIS HARM IS MODERATE (3), DEFINED AS "ANY INSTRUMENT-RELATED HAZARD THAT COULD RESULT IN MARKED INTRA-OPERATIVE INCONVENIENCE TO THE SURGEON AND STAFF", AND THE PREDICTED OCCURRENCE IS LOW (2), DEFINED AS LESS THAN 1%. AN ANALYSIS WAS PERFORMED FOR THE COMPLAINT RATE OF HANDLE BREAKAGE BASED ON YEAR OF HANDLE MANUFACTURE. THE ANALYSIS SHOWED THAT THE COMPLAINT RATE IS VERY LOW WITH NO UPWARD TREND. THE OVERALL OCCURRENCE RATE WAS DETERMINED TO BE 0.030%. BASED ON THE AVAILABLE DATA, THE DEVICE IS PERFORMING WITHIN ITS PREDICTED RISK. MOST LIKELY CAUSE: BASED ON THE INVESTIGATION THE MOST LIKELY CAUSE OF THE REPORTED EVENT IS RELATED TO HIGH IMPACTION FORCES APPLIED TO THE INSTRUMENT DURING THE SURGICAL PROCEDURE. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. IT WAS REPORTED THAT THE LPI HANDLE OF THE TIBIAL IMPACTOR BROKE DURING IMPACTION DURING A KNEE REPLACEMENT SURGERY. A REVIEW OF THE DEVICE HISTORY DID NOT SHOW A TREND IN THOSE LOTS WITH COMPLAINT DEVICES, NOR DID A REVIEW OF DESIGN CHANGES SHOW ANY CHANGES MADE TO THE DESIGN WHICH COULD HAVE CONTRIBUTED TO THE COMPLAINT. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING OR DESIGN ISSUES, NOR DID IT LEAD TO ANY ADVERSE PATIENT IMPACT. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10029 | 2022-03-29 | 2017-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRULIANT TIBIAL INSERT TRIAL/TRAY ADAPTER, SIZE 4 | EXACTECH, INC. |  | 02-029-25-4040 |  | IT WAS REPORTED FROM THE US THAT THE SURGEON INSERTED A SIZE 4 TIBIAL INSERT TRIAL/TRAY ADAPTER INTO A SIZE 5F/4T TIBIAL FIT TRAY. HE THEN INSERTED A HOHMANN RETRACTOR TO REMOVE THE ADAPTER PLATE. THE HOHMANN TIP MADE CONTACT WITH THE FLEXURE ON THE ADAPTER PLATE AND THE FLEXURE DISASSOCIATED FROM THE MAIN BODY OF THE PLATE. THE SALES REPRESENTATIVE WAS PRESENT AT THE SURGERY. THERE WAS NO DELAY OF SURGERY AND THERE WAS NO PATIENT INJURY, ADVERSE EVENT, OR CLINICAL CONSEQUENCE TO THE PATIENT. MULTIPLE EMAILS HAVE BEEN SENT. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE REPORTER RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS NOT RETURNED. ENG EVAL COMPLETED ON (B)(6) 2018 BY MB. THE ROOT CAUSE ASSOCIATED WITH CAPA (B)(4) RESIDES IN THE MISUSE/ MISTREATMENT OF THE ADAPTOR PLATES OR CANNOT BE IDENTIFIED DUE TO INSUBSTANTIAL EVIDENCE (REPROCESSING RELATED COMPLAINTS, (B)(4)). THE CAPA TEAM HAS DETERMINED, HOWEVER, THAT USE CONSIDERATIONS INDICATE THAT THE FLEXURE DESIGN IS SUSCEPTIBLE TO DAMAGE FROM SAID MISUSE/MISTREATMENT. DESIGN: A REVIEW OF THE TOLERANCE STACK-UP PERFORMED IN TM-2017-0417 WAS CONDUCTED. AS PART OF THIS INVESTIGATION, THE TOLERANCE STACK-UP BETWEEN THE TRULIANT TIBIAL TRIAL ADAPTOR PLATES AND THE CORRESPONDING TRULIANT AND LOGIC TIBIAL TRAYS WAS REVISITED TO CONFIRM IT HAD BEEN ADEQUATELY PERFORMED. THE REVIEW WAS FOUND TO BE ACCURATE. THERE IS NO NEW INFORMATION THAT HAS BEEN FOUND THAT WOULD ALTER THE RESULTS OR METHODS OF SAID ANALYSIS. IN ADDITION, ATTEMPTS WERE MADE IN THE LABORATORY TO REPLICATE FAILURE MODES WITHOUT SUCCESS. THE DESIGN FUNCTIONS AS INTENDED UNDER WORST CASE CONDITIONS. MANUFACTURING REVIEW OF THE DEVICE DHR CONCLUDED THAT ALL PARTS WERE MADE WITHIN SPECIFICATION. USE CONSIDERATIONS. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. AN INVESTIGATION WAS CONDUCTED; THE ROOT CAUSE ASSOCIATED WITH CAPA (B)(4) RESIDES IN THE MISUSE/ MISTREATMENT OF THE ADAPTOR PLATES OR CANNOT BE IDENTIFIED DUE TO INSUBSTANTIAL EVIDENCE. THE CAPA TEAM HAS DETERMINED, HOWEVER, THAT USE CONSIDERATIONS INDICATE THAT THE FLEXURE DESIGN IS SUSCEPTIBLE TO DAMAGE FROM SAID MISUSE/MISTREATMENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10014 | 2022-03-29 | 2017-08-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LOGIC PS TIBIA, MODULAR INSERT TRIAL, SIZE 2.5, 9MM | EXACTECH, INC. |  | 02-013-35-2509 |  | IT WAS REPORTED FROM KOREA THAT THE INSERT TRIAL WAS DESTROYED DURING THE REDUCTION AND ROM TESTING OF THE INSERT. NO BROKEN PIECES WERE LEFT IN THE PATIENT. NO PARTS/PIECES OF THE DEVICE FELL INTO THE SURGICAL SITE. A SECOND SET OF INSTRUMENTS WAS USED. NO PROBLEMS WITH THE PATIENT, BUT THE SURGEON COMPLAINED ABOUT THE PRODUCT. THE AGENT WAS PRESENT AT SURGERY AND THERE WAS NO DELAY. NO PATIENT DEMOGRAPHIC WITH BE PROVIDED DUE TO SITED PRIVACY LAWS. ATTEMPTS HAVE BEEN MADE FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. NO ADDITIONAL ATTEMPTS WILL BE MADE FOR INFORMATION REGARDING THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RETURNED. MFG DATE: 09/03/2014. CAPA (B)(4) INVESTIGATION STATES: THE CAUSE OF THIS FAILURE COULD BE REPROCESSING OF THESE PARTS. IN ORDER TO UNDERSTAND WHY THESE DEVICES ARE FAILING IN SINGAPORE AND KOREA, A QUESTIONNAIRE WAS SENT TO THE SALES FORCE IN KOREA AND SINGAPORE TO INQUIRE ABOUT TKA TRIALING, GENERAL USE, AND STERILIZATION/REPROCESSING TECHNIQUES. ACCORDING TO AN EMAIL FROM THE SALES MANAGER AT EXACTECH KOREA, THESE INSTRUMENTS ARE BEING STERILIZED FOR 40-50 MINUTES AT 134 DEGREES (CELSIUS). THESE REPROCESSING METHODS ARE OUTSIDE OF EXACTECH'S RECOMMENDED STERILIZATION INSTRUCTIONS PER 700-096-057 - REPROCESSING INSTRUCTIONS FOR REUSABLE SURGICAL INSTRUMENTS, WHICH STATES PARAMETERS FOR STERILIZATION. (B)(4) IS THE INITIAL SCOPE ASSESSMENT. ROOT CAUSE: THE REPROCESSING OF THESE INSTRUMENTS EXCEEDS THE RECOMMENDED PARAMETERS LISTED IN THE REPROCESSING INSTRUCTIONS. THE MATERIAL DEGRADATION FROM EXCESSIVE STERILIZATION OF THESE INSTRUMENTS HAS LIKELY LED TO THE OBSERVED FAILURE. NO ADDITIONAL ACTION REQUIRED. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE WAS NO PATIENT ADVERSE EVENT. AN INVESTIGATION WAS CONDUCTED UNDER CAPA (B)(4). THE REPROCESSING OF THESE INSTRUMENTS EXCEEDS THE RECOMMENDED PARAMETERS LISTED IN THE REPROCESSING INSTRUCTIONS. THE MATERIAL DEGRADATION FROM EXCESSIVE STERILIZATION OF THESE INSTRUMENTS HAS LIKELY LED TO THE OBSERVED FAILURE. NO ADDITIONAL ACTION REQUIRED. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10030 | 2022-03-29 | 2017-06-29 | I, n, j, u, r, y | Manufacturer report | EXACTECH | REAMER STARTER, MODULAR REVERSE, PILOT TIP | EXACTECH, INC. |  | 321-25-01 |  | IT WAS REPORTED BY ORTHO REPS (B)(4) THAT WHEN DR. (B)(6) WAS REAMING WITH THE PILOT TIP REAMER FOR THE GLENOID, THE TIP BROKE OFF. THE REAMER HAD CONTACTED THE RETRACTORS THAT WERE BEING USED. THE HOSPITAL IS KEEPING THE INSTRUMENT TO EVALUATE BEFORE RETURNING IT TO THE AGENT. FLOURO WAS USED AND THE TIP OF THE REAMER WAS NOT SEEN WITHIN THE SURGICAL SITE. THE CANNULATED REAMER WAS USED TO FINISH. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS NOT RETURNED TO EXACTECH FOR ANALYSIS. DESIGN-RELATED ISSUES: THE DESIGN OF THE REVERSE PILOT TIP REAMER HAS BEEN IN THE FIELD SINCE 2009. AT THE TIME OF THIS EVALUATION, 09 MAY 2018, EXACTECH HAD RECEIVED 23 SIMILAR COMPLAINT REPORTS INVOLVING (B)(4) DEVICES SINCE 2009. BECAUSE THE REVERSE PILOT TIP REAMER IS USED IN REVERSE TOTAL SHOULDER SURGERIES, SALES DATA FOR GLENOSPHERE COMPONENTS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4). THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. A CORRECTIONS AND REMOVAL COMMITTEE (CRC) MEETING HAS BEEN HELD TO DISCUSS HOW TO FIX THIS ISSUE. REFER TO CRC (B)(4) FOR DETAILS. A CAPA WAS ALSO OPENED TO ADDRESS THE ISSUE. THIS CAPA RESULTED IN A FIELD ADVISORY NOTICE. REFER TO CAPA (B)(4). MFG-RELATED ISSUES: NO INFORMATION REGARDING SERIAL NUMBER OR MANUFACTURING LOT WAS PROVIDED. RISK MANAGEMENT: A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS REVERSE PILOT TIP REAMER, 750-2007-026-RMR REV H, WAS REVIEWED. THE RISK IS CAPTURED IN ROW 17. NO CORRECTIVE ACTION IS REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE BROKEN PILOT TIP REAMER REPORTED IN (B)(6) WAS LIKELY THE RESULT OF THE TIP OF THE DEVICE CONTACTING THE RETRACTORS BEING USED DURING THE REAMING STEP. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. EXACTECH IFU 700-096-004 STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED INSTRUMENT MALFUNCTION IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO ANY PATIENT IMPACT. AN INVESTIGATION WAS CONDUCTED THAT THE MOST LIKELY ROOT CAUSE OF THE DEVICE BREAKAGE REPORTED IN THIS EVENT WAS DUE TO THE TIP OF THE DEVICE CONTACTING THE RETRACTORS BEING USED DURING THE REAMING STEP. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. ATTEMPTS HAVE BEEN MADE FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. NO ADDITIONAL ATTEMPTS WILL BE MADE FOR INFORMATION REGARDING THIS EVENT. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10018 | 2022-03-29 | 2017-09-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LPI IMPACTOR HANDLE | EXACTECH, INC. |  | 213-65-00 · |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGICAL PROCEDURE THE LPI HANDLE BROKE. THE INSTRUMENT BROKE BY IMPACTING IT WITH A HAMMER. THERE WERE NO PIECES OF THE DEVICE THAT FELL INTO THE PATIENT. THE ISSUE WAS RESOLVED BY USING THE ALTERNATE INSTRUMENT. THERE WAS NO DELAY TO SURGERY AND NO ADVERSE CLINICAL EFFECT TO THE PATIENT. NO PATIENT INFORMATION WILL BE PROVIDED BY THE CONTACT. MULTIPLE ATTEMPTS WERE MADE FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED FOR ANALYSIS. ENG EVAL 11/15/2018 BY (B)(4). MFG DATE: 07/11/2011. AN INVESTIGATION OF COMPLAINTS RELATED TO LPI HANDLE - BROKEN NON-FUNCTIONAL, SPECIFICALLY RELATED TO CRACKS/BREAKAGE FOR THE PLASTIC HANDLE WAS CONDUCTED UNDER (B)(4). THE LPI HANDLE CONSISTS OF A METAL STRIKE PLATE, METAL END PIECE WHICH ENGAGES WITH MATING INSTRUMENTS AND IMPLANTS, AND A PLASTIC CENTRAL HANDLE. THE METAL END PIECES ARE ASSEMBLED TO THE PLASTIC HANDLE VIA THERMAL ASSEMBLY WHICH INVOLVES HEATING THE METAL AND PRESSING THE HOT METAL PIECES INTO THE HANDLE. THE METAL PIECES HAVE CYLINDRICAL BOSSES WHICH FEATURE A PRESS-FIT WITH A FEMALE HOLE ON THE PLASTIC HANDLE. THE LPI HANDLE ASSEMBLES WITH TWO END PIECES FOR USE THREE TIMES DURING THE SURGICAL PROCEDURE: 1) IMPACTION OF THE FEMORAL TRIAL 2) IMPACTION OF THE TIBIAL TRAY IMPLANT AND 3) IMPACTION OF THE FEMORAL IMPLANT. AVAILABLE DEVICES FROM PREVIOUSLY REPORTED COMPLAINTS WERE EVALUATED. COMPLAINT DEVICES APPEARED TO HAVE SIMILAR FAILURE MODES, AND EITHER WERE CRACKED OR FRACTURED. CRACKS WERE CONSISTENT WITH STARTING INSIDE THE PLASTIC HANDLE AT THE BASE OF ONE OF THE TWO METAL ATTACHMENT FEATURES AND SPREADING OUTWARD. CRACKS WERE EITHER LINEAR, ALONG THE HANDLE AXIS OR HORIZONTAL, NEARLY IN LINE WITH THE ORIGIN. THE REVIEW OF THE LPI IMPACTOR HANDLES (213-65-00) CITED IN COMPLAINTS FOR BROKEN OR CRACKED HANDLES SHOWED A CONSISTENT FAILURE MODE. A REVIEW OF THE DEVICE ORDER HISTORY DID NOT SHOW A TREND IN THOSE LOTS WITH COMPLAINT DEVICES, NOR DID A REVIEW OF DESIGN CHANGES SHOW CHANGES MADE TO THE DESIGN WHICH COULD CONTRIBUTE TO THE COMPLAINT. PER 750-2009-023-RMR-INSTRUMENT, REV B - DEVICE FRACTURE/BREAKAGE (LINE ITEM 18) CAN LEAD TO A SURGICAL DELAY. THE PREDICTED SEVERITY OF THIS HARM IS MODERATE (3), DEFINED AS "ANY INSTRUMENT - RELATED HAZARD THAT COULD RESULT IN MARKED INTRA-OPERATIVE INCONVENIENCE TO THE SURGEON AND STAFF", AND THE PREDICTED OCCURRENCE IS LOW (2), DEFINED AS LESS THAN 1%. (B)(4) WAS INITIATED ON 03/02/2018 TO EVALUATE THIS ISSUE. THE MOST LIKELY CAUSE OF THE REPORTED EVENTS IS RELATED TO HIGH IMPACTION FORCES REQUIRED DURING IMPLANT SURGERY. BASED ON THE AVAILABLE DATA, THE DEVICE IS PERFORMING WITHIN ITS PREDICTED RISK. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE WAS NO SURGICAL DELAY OR PATIENT IMPACT. AN INVESTIGATION WAS CONDUCTED UNDER (B)(4); THE MOST LIKELY CAUSE OF THE REPORTED EVENTS IS RELATED TO HIGH IMPACTION FORCES REQUIRED DURING IMPLANT SURGERY. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10011 | 2022-03-29 | 2017-05-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LPI IMPACTOR HANDLE | EXACTECH, INC. |  | 213-65-00 |  | IT WAS REPORTED THAT DURING AN ORTHOPEDIC KNEE SURGERY THE HANDLE SPRING OF THE LPI IMPACTOR HANDLE BROKE. THE SALES REPRESENTATIVE WAS PRESENT AT THE SURGERY. THE SURGERY CONTINUED AND WAS COMPLETED WITHOUT ADDITIONAL ISSUES. THERE WAS NO REPORTED DELAY OF SURGERY AND THERE WAS NO REPORTED PATIENT INJURY, ADVERSE EVENT, OR CLINICAL CONSEQUENCE TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE REPORTER RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED FOR ANALYSIS. PER (B)(4): AN INVESTIGATION OF COMPLAINTS RELATED TO LPI HANDLE - BROKEN NON-FUNCTIONAL, SPECIFICALLY RELATED TO CRACKS/BREAKAGE FOR THE PLASTIC HANDLE WAS CONDUCTED UNDER TM (B)(4). THE LPI HANDLE CONSISTS OF A METAL STRIKE PLATE, METAL END PIECE WHICH ENGAGES WITH MATING INSTRUMENTS AND IMPLANTS, AND A PLASTIC CENTRAL HANDLE. THE METAL END PIECES ARE ASSEMBLED TO THE PLASTIC HANDLE VIA THERMAL ASSEMBLY WHICH INVOLVES HEATING THE METAL AND PRESSING THE HOT METAL PIECES INTO THE HANDLE. THE METAL PIECES HAVE CYLINDRICAL BOSSES WHICH FEATURE A PRESS-FIT WITH A FEMALE HOLE ON THE PLASTIC HANDLE. THE LPI HANDLE ASSEMBLES WITH TWO END PIECES FOR USE THREE TIMES DURING THE SURGICAL PROCEDURE: 1) IMPACTION OF THE FEMORAL TRIAL 2) IMPACTION OF THE TIBIAL TRAY IMPLANT AND 3) IMPACTION OF THE FEMORAL IMPLANT. AVAILABLE DEVICES FROM PREVIOUSLY REPORTED COMPLAINTS WERE EVALUATED. COMPLAINT DEVICES APPEARED TO HAVE SIMILAR FAILURE MODES, AND EITHER WERE CRACKED OR FRACTURED. CRACKS WERE CONSISTENT WITH STARTING INSIDE THE PLASTIC HANDLE AT THE BASE OF ONE OF THE TWO METAL ATTACHMENT FEATURES AND SPREADING OUTWARD. CRACKS WERE EITHER LINEAR, ALONG THE HANDLE AXIS OR HORIZONTAL, NEARLY IN LINE WITH THE ORIGIN. THE REVIEW OF THE LPI IMPACTOR HANDLES (213-65-00) CITED IN COMPLAINTS FOR BROKEN OR CRACKED HANDLES SHOWED A CONSISTENT FAILURE MODE. A REVIEW OF THE DEVICE ORDER HISTORY DID NOT SHOW A TREND IN THOSE LOTS WITH COMPLAINT DEVICES, NOR DID A REVIEW OF DESIGN CHANGES SHOW CHANGES MADE TO THE DESIGN WHICH COULD CONTRIBUTE TO THE COMPLAINT. IFU 700-096-181 STATES: INSTRUMENT INSPECTION: · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFUS ALSO STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT IMPACT. AN INVESTIGATION WAS CONDUCTED AND CONCLUDED THAT THE MOST LIKELY CAUSE OF THE REPORTED EVENTS IS RELATED TO HIGH IMPACTION FORCES REQUIRED DURING IMPLANT SURGERY. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10016 | 2022-03-29 | 2017-11-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LPI LOCKING FEMORAL IMPACTOR | EXACTECH, INC. |  | 213-64-01 |  | IT WAS REPORTED FROM OUS THAT DURING THE SURGERY, THE INSTRUMENT WAS BROKEN DURING IMPACTING. A BROKEN PIECE FELL ON THE FLOOR NOT INTO THE PATIENT'S SURGICAL SITE. THE IMPLANT WAS ALREADY FIXATED BEFORE THE BREAKAGE OF THE IMPACTOR. THE AGENT WAS PRESENT AT SURGERY. THERE WAS NO DELAY TO THE SURGERY AND NO CLINICAL CONSEQUENCE TO THE PATIENT. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED FOR ANALYSIS. COMPLETED ON 03/07/2019 KN. MFG DATE: 06/15/2006. DESIGN-RELATED ISSUES: THE DESIGN OF THE LPI LOCKING FEMORAL IMPACTOR HAS BEEN IN THE FIELD SINCE 2005. EXACTECH HAS RECEIVED 2 COMPLAINT REPORTS INVOLVING LPI LOCKING FEMORAL IMPACTOR DEVICES THAT HAVE FRACTURED AT THE THREADED COMPONENT OF THE LOCKING JAW SINCE 2014. BECAUSE THE FEMORAL IMPACTOR IS USED IN TOTAL KNEE SURGERIES, SALES DATA FROM 2014 TO 2018 FOR FEMORAL IMPLANTS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF RECEIVING ANY COMPLAINT REPORTS INVOLVING ANOTHER PART FROM THIS MANUFACTURING LOT OF 25 PIECES WHICH HAVE BEEN IN THE FIELD SINCE 2016. EXACTECH COMPLAINT DATA FROM 2014 THROUGH 2018 INVOLVING THE REPORTED FAILURE FOR THIS FAMILY OF DEVICES WAS REVIEWED. THE INVESTIGATIONS FOR THOSE INCIDENCES HAVE NOT IDENTIFIED ANY EVIDENCE THAT A MANUFACTURING ISSUE CAUSED OR CONTRIBUTED TO THIS REPORTED ISSUE. THEREFORE, THIS DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS LOCKING FEMORAL IMPACTOR, 750-2003-019- RMR INSTRUMENTS REV-, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 18. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS " VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. THE MOST LIKELY CAUSE: THE BROKEN INSTRUMENT REPORTED IN EXPERIENCE C2017-656 WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE LOCKING JAW, WHICH LED TO THE PLASTIC DEFORMATION AND ULTIMATE FRACTURE OF THE SCREW THREADS. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES. THIS EVENT DID NOT LEAD TO PATIENT ADVERSE EVENT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE LOCKING JAW, WHICH LED TO THE PLASTIC DEFORMATION AND ULTIMATE FRACTURE OF THE SCREW THREADS. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10025 | 2022-03-29 | 2017-11-20 | I, n, j, u, r, y | Manufacturer report | EXACTECH | NOVATION ELEMENT FEMORAL STEM, COLLARLESS, HA, HIGH OFFSET, 12/14 TAPER, SIZE 10 | EXACTECH, INC. |  | 164-12-10 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE HEAD TRIAL. THE HEAD TRIAL GOT STUCK ON AN ELEMENT FEMORAL STEM AND THE WHOLE SYSTEM HAD TO BE REMOVED. ANOTHER SET WAS OPENED. THE PATIENT WAS FINE LEAVING THE OPERATING ROOM. THERE WAS A DELAY, BUT THE PATIENT WAS NOT ADVERSELY AFFECTED. THE REP WAS PRESENT AT THE TIME OF SURGERY. NO PATIENT DEMOGRAPHICS WERE REPORTED. THE HEAD TRIAL WAS RETURNED ON 12/13/2017. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT FOLLOWING MULTIPLE ATTEMPTS. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS AVAILABLE FOR ANALYSIS. ENGINEERING ANALYSIS COMPLETED BY AD ON 9/24/2018. DESIGN-RELATED ISSUES: THE DESIGN OF THE FEMORAL HEAD TRIAL HAS BEEN IN THE FIELD SINCE 2011. EXACTECH HAS RECEIVED 4 COMPLAINT REPORTS INVOLVING 6 FEMORAL HEAD TRIALS DUE TO DIFFICULT ASSEMBLY/DISASSEMBLY SINCE 2008. BECAUSE FEMORAL HEAD TRIALS ARE USED IN TOTAL HIP SURGERIES, SALES DATA FOR FEMORAL HEADS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "'VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH HAS NOT RECEIVED ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT, WHICH HAS BEEN IN THE FIELD SINCE 2012. HOWEVER, THIS PART WAS FOUND TO BE OUTSIDE THE TOLERANCE RANGE. THIS WAS CONFIRMED VIA SPECIAL INSPECTION. A SUPPLIER CORRECTIVE ACTION REQUEST (SCAR) HAS BEEN ISSUED BY THE PRODUCT TEAM. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS FEMORAL HEAD TRIAL, 750-2011-012-RMR REV A, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 7. THE DISASSEMBLY DIFFICULTY REPORTED IN EXPERIENCE (B)(6) WAS LIKELY THE RESULT OF THE HEAD TRIAL BEING MANUFACTURED OUTSIDE THE ALLOWABLE TOLERANCE RANGE. THIS WAS CONFIRMED VIA SPECIAL INSPECTION. A SUPPLIER CORRECTIVE ACTION REQUEST HAS BEEN ISSUED BY THE PRODUCT TEAM. EXACTECH IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO DESIGN ISSUES. THE HEAD TRIAL GOT STUCK ON AN ELEMENT FEMORAL STEM AND THE WHOLE SYSTEM HAD TO BE REMOVED. THERE WAS NO PATIENT ADVERSE EVENT. AN INVESTIGATION WAS CONDUCTED: THE DISASSEMBLY DIFFICULTY REPORTED WAS LIKELY THE RESULT OF THE HEAD TRIAL BEING MANUFACTURED OUTSIDE THE ALLOWABLE TOLERANCE RANGE. THIS WAS CONFIRMED VIA SPECIAL INSPECTION. A SUPPLIER CORRECTIVE ACTION REQUEST HAS BEEN ISSUED BY THE PRODUCT TEAM. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10010 | 2022-03-29 | 2017-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LPI IMPACTOR HANDLE | EXACTECH, INC. |  | 213-65-00 |  | IT WAS REPORTED THAT DURING AN ORTHOPEDIC KNEE SURGERY THE HANDLE SPRING OF THE LPI IMPACTOR HANDLE HAS BROKEN. THE SALES REPRESENTATIVE WAS NOT PRESENT AT THE SURGERY AND REPORTED NO DEVICE FRAGMENTS FELL INTO THE SURGICAL SITE. THE SURGERY CONTINUED AND WAS COMPLETED WITHOUT ADDITIONAL ISSUES. THERE WAS NO REPORTED DELAY OF SURGERY AND THERE WAS NO REPORTED PATIENT INJURY, ADVERSE EVENT, OR CLINICAL CONSEQUENCE TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE REPORTER RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED FOR ANALYSIS. DESIGN: THE LPI IMPACTOR HANDLE HAS BEEN ON THE MARKET SINCE 2011. EXACTECH IS AWARE OF (B)(4) OTHER COMPLAINTS INVOLVING A BROKEN LPI IMPACTOR HANDLE SINCE 2014. SINCE THIS INSTRUMENT IS USED TO IMPLANT TIBIAL TRAYS, THE SALES DATA FOR ALL TIBIAL TRAYS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4). THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE DETAILED BELOW. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MANUFACTURING: EXACTECH IS AWARE OF 3 OTHER COMPLAINT REPORTS INVOLVING PARTS FROM THIS MANUFACTURING LOT AND THE RESULTS FROM THE INSPECTION OF ONE OF THOSE 3 DEVICES SHOWED THAT DRAWING ZONES (B)(4) WERE OUTSIDE SPECIFICATIONS. A REVIEW OF THE DEVICE HISTORY RECORD SHOWED THAT THE NAMED DEVICES WERE ACCEPTED WITH CONFORMANCE TO THE DEVICE REQUIREMENTS; IN PARTICULAR, BOTH A FIT AND LOCK CHECK WERE SATISFIED. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. RISK MANAGEMENT: A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THE LPI INSTRUMENTS, (B)(4) INSTRUMENTS REV B, WAS REVIEWED. THE RISK IS CAPTURED IN LINES 18 AND 23 OF THE RMR. MOST LIKELY CAUSE: THE BROKEN LPI IMPACTOR HANDLE WAS LIKELY THE RESULT OF DEFORMATION FROM MULTIPLE IMPACTIONS OVER THE COURSE OF OVER 6 YEARS OF USE. IFU 700-096-181 STATES: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFUS ALSO STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT IMPACT. AN INVESTIGATION WAS CONDUCTED AND CONCLUDED THAT THE BROKEN LPI IMPACTOR HANDLE WAS LIKELY THE RESULT OF DEFORMATION FROM MULTIPLE IMPACTIONS OVER THE COURSE OF OVER 6 YEARS OF USE. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10013 | 2022-03-29 | 2017-06-19 | I, n, j, u, r, y | Manufacturer report | EXACTECH | REAMER STARTER, MODULAR REVERSE, PILOT TIP | EXACTECH, INC. |  | 321-25-01 |  | IT WAS REPORTED THAT THE REAMER TIP WAS PLACED IN THE 2MM DRILL HOLE THAT WAS DRILLED FOR THE REAMER, FLUSH AGAINST THE GLENOID AND THE REAMING SEQUENCE STARTED. AS THE REAMER CONTACTED THE BONE AND REAMING BEGAN, THE TORQUE OF THE REAMER SHEARED OFF. WE PROCEEDED TO A 38 REAMER AND FINISH REAMING. THE PATIENT RECEIVED A 38 GLENOSPHERE WITH GREAT RANGE OF MOTION AND FUNCTION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED FOR ANALYSIS. OTHER THAN THE BROKEN REAMER TIP, ALL OTHER ASPECTS OF THE INSTRUMENT APPEAR TO BE UNREMARKABLE. DESIGN: THIS DEVICE DESIGN HAS BEEN IN THE FIELD SINCE 2009. AT THIS TIME OF THE ENGINEERING ANALYSIS ON 30 JUL 2017, EXACTECH HAD RECEIVED 25 SIMILAR COMPLAINTS OF TIP BREAKAGE INVOLVING 27 INDIVIDUAL DEVICES WITH AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4). THIS RATE IS "VERY LOW" AND WAS ADDRESS IN CAPA (B)(4) WITH THE OPERATIVE GUIDE UPDATED TO INCLUDE A WARNING REGARDING OFF-AXIS REAMING. MANUFACTURING: NO OTHER COMPLAINTS HAD BEEN RECEIVED INVOLVING THIS LOT, CONTAINING (B)(4) UNITS. THIS LOT HAS BEEN IN THE FIELD SINCE 2015; THEREFORE, THIS ISSUE DOES NOT SEEM TO BE RELATED TO ANY MANUFACTURING - RELATED CAUSE. MOST LIKELY CAUSE: AN INVESTIGATION WAS CONDUCTED THAT THE MOST LIKELY ROOT CAUSE OF THE DEVICE BREAKAGE REPORTED IN THIS EVENT WAS LIKELY DUE TO OFF-AXIS REAMING, ALLOWING EXCESSIVE TORQUE ON THE TIP OF THE DEVICE WHICH LEAD TO THE BREAKAGE. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT IMPACT. AN INVESTIGATION WAS CONDUCTED; THIS EVENT WAS LIKELY DUE TO OFF-AXIS REAMING, ALLOWING EXCESSIVE TORQUE ON THE TIP OF THE DEVICE WHICH LEAD TO THE DEVICE BREAKAGE. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10007 | 2022-03-29 | 2017-03-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LPI IMPACTOR HANDLE | EXACTECH, INC. |  | 213-65-00 |  | IT WAS REPORTED THAT DURING AN ORTHOPEDIC KNEE SURGERY THE HANDLE SPRING OF THE LPI IMPACTOR HANDLE BROKE. THE SALES REPRESENTATIVE WAS NOT PRESENT AT THE SURGERY. IT WAS REPORTED THAT NO DEVICE FRAGMENTS FELL INTO THE SURGICAL SITE. THE SURGERY CONTINUED AND WAS COMPLETED WITHOUT ADDITIONAL ISSUES. THERE WAS NO DELAY OF SURGERY AND THERE WAS NO REPORTED PATIENT INJURY, ADVERSE EVENT, OR CLINICAL CONSEQUENCE TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE REPORTER RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE IN QUESTION WAS RETURNED TO EXACTECH FOR ANALYSIS. DESIGN-RELATED ISSUES: THE DESIGN OF THIS LPI IMPACTOR HANDLE HAS BEEN IN THE FIELD SINCE 2009. (B)(4). THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT, WHICH HAS BEEN IN THE FIELD SINCE 2016. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. RISK MANAGEMENT: A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS LPI IMPACTOR HANDLE, 750-2003-019- RMR-INSTRUMENTS REV-, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 23. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS " VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE DEFORMED DEVICE REPORTED IN EXPERIENCE (B)(4) WAS LIKELY THE RESULT OF ATTEMPTING TO REMOVE THE BAL-SEAL SPRING DURING CLEANING WHICH LED TO DEFORMATION OF THE SPRING COMPONENT. EXACTECH IFU 700-096-004 STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED INSTRUMENT MALFUNCTION IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO ANY PATIENT IMPACT. AN INVESTIGATION WAS CONDUCTED AND CONCLUDED THAT THE MOST LIKELY ROOT CAUSE OF THE DEFORMED DEVICE REPORTED IN EXPERIENCE (B)(4) WAS LIKELY THE RESULT OF ATTEMPTING TO REMOVE THE BAL-SEAL SPRING DURING CLEANING WHICH LED TO DEFORMATION OF THE SPRING COMPONENT. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-10017 | 2022-03-29 | 2017-12-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LPI PATELLA PREP HANDLE | EXACTECH, INC. |  | 213-60-00 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGICAL PROCEDURE A PATELLA PREP HANDLE BROKE. A PIN FROM THE HANDLE WAS FOUND ON THE GROUND AND THROWN AWAY BY THE HOSPITAL STAFF. THE SURGEON USED ANOTHER INSTRUMENT, PATELLA CLAMP, TO COMPLETE THE PROCEDURE. AN X-RAY WAS TAKEN TO ENSURE THAT NO PIECES WERE LEFT IN THE PATIENT. THERE WAS A SURGICAL DELAY WITH NO ADVERSE EVENT REPORTED FOR THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. THE AGENT AND AGENCY ARE NO LONGER ACTIVE WITH THE COMPANY; NO ADDITIONAL INFORMATION IS AVAILABLE. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE WAS RECEIVED FOR ANALYSIS. ENG ANALYSIS COMPLETED 03/15/2018 BY MB THE MISSING PIN IS ASSEMBLED USING A TACK WELD, WHICH APPEARS TO HAVE FAILED. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE AND CONSISTENT WITH YEARS OF USE. DESIGN: THE OCCURRENCE RATE IS LESS THAN 0.5%, "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THIS DOES NOT APPEAR TO BE DESIGN RELATED. MANUFACTURING: THERE HAVE NOT BEEN ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT WHICH HAS BEEN IN THE FIELD SINCE 2012. THIS DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. MOST LIKELY CAUSE: THE EVENT REPORTED ABOUT WAS LIKELY THE RESULT OF A FAILURE OF THE TACK WELD THAT HOLDS THE PIN IN PLACE, ALLOWING IT TO DISASSOCIATE FROM THE INSTRUMENT. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT IMPACT. AN X-RAY WAS TAKEN TO ENSURE THAT NO PIECES WERE LEFT IN THE PATIENT. AN INVESTIGATION WAS CONDUCTED; THE EVENT REPORTED ABOUT WAS LIKELY THE RESULT OF A FAILURE OF THE TACK WELD THAT HOLDS THE PIN IN PLACE, ALLOWING IT TO DISASSOCIATE FROM THE INSTRUMENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00335 | 2022-03-30 | 2018-11-06 | I, n, j, u, r, y | Manufacturer report | EXACTECH | MDLR RV RMR 42MM PLT TIP | EXACTECH, INC. |  | 321-25-42 |  | IT WAS REPORTED FROM THE OUS THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED A BROKEN PILOT TIP REAMER. THE TIGHT SHOULDER RETRACTED OFF THE HUMERAL HEAD AND THE 38MM TIP BROKE. THE SURGEON TRIED AGAIN AFTER SOME MORE RESECTION AND HAD TO USER THE BIGGER 42MM THE TIP BROKE AGAIN. THERE WERE NO PIECES THAT FELL INTO THE SURGICAL SITE, NO DELAY, AND NO ADVERSE EVENT TO THE PATIENT. THE AGENT WAS PRESENT AT THE TIME OF SURGERY. THE PATIENT'S HEALTH STABLE LEAVING THE OR. NO ADDITIONAL INFORMATION. 2 OF 2 MDRS. |  | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00326 | 2022-03-30 | 2019-02-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | DRILL GUIDE | EXACTECH, INC. |  | 181-00-05 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE DRILL GUIDE. THE INSTRUMENT BROKE WHILE ATTEMPTING TO DRILL THE SCREW HOLE THROUGH. THE INSERTED CUP WAS REMOVED AND ANOTHER DRILL GUIDE IN THE SET WAS USED. THERE WAS NO DELAY TO SURGERY. NO COMPROMISE TO PATIENT, WHO LEFT THE OPERATING ROOM IN NORMAL TIME. THE SHEARED SURFACES APPEAR CONSISTENT WITH APPLYING A BENDING MOMENT. THE REPRESENTATIVE WAS PRESENT AT THE TIME OF SURGERY. THE DEVICE WAS RETURNED. MULTIPLE ATTEMPTS WERE MADE FOR ADDITIONAL EVENT/PATIENT INFORMATION, NO RESPONSES. THERE IS NO ADDITIONAL INFORMATION FROM THE CONTACTS ABOUT THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED. THE DEVICE WAS AVAILABLE FOR ANALYSIS. FINDINGS: THE SHEARED SURFACES APPEAR CONSISTENT WITH APPLYING A BENDING MOMENT TO THE DRILL GUIDE WHILE USING A DRILL, CAUSING THE WELD BETWEEN THE DRILL GUIDE TIP AND THE DRILL GUIDE SHAFT TO BREAK. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: BASED ON PREVIOUS INVESTIGATIONS, THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE DRILL GUIDE WHILE USING A DRILL, WHICH LED TO FRACTURE OF THE WELD BETWEEN THE DRILL GUIDE TIP AND THE DRILL GUIDE SHAFT. RMR REVIEW: THE RISK IS CAPTURED. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. ¿ CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. THE SHEARED SURFACES APPEAR CONSISTENT WITH APPLYING A BENDING MOMENT. AN INVESTIGATION WAS CONDUCTED: BASED ON PREVIOUS INVESTIGATIONS, THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE DRILL GUIDE WHILE USING A DRILL, WHICH LED TO FRACTURE OF THE WELD BETWEEN THE DRILL GUIDE TIP AND THE DRILL GUIDE SHAFT. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00334 | 2022-03-30 | 2019-07-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | FCT GAP CHECK TL SZ 3-4 | EXACTECH, INC. |  | 351-10-13 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE INSTRUMENTATION. THE ANKLE PEG PUNCH BROKE AT JUNCTION BETWEEN ANKLE PEG PUNCH AND MODULAR HANDLE WHILE IMPACTING TIBIA. THE SURGEON PROCEEDED TO IMPACT THE TIBIA WITH ANKLE PEG PUNCH WITHOUT CONNECTION TO MODULAR HANDLE. THERE WAS A DELAY, BUT THE PATIENT WAS NOT ADVERSELY AFFECTED. NO PARTS/PIECES FELL INTO THE SURGICAL SITE. ¿NO PIECES WERE LEFT IN THE PATIENT.¿ CASE COMPLETED WITHOUT IT CAUSING ANY ADVERSE EFFECTS. THE REP WAS PRESENT AT THE TIME OF SURGERY. THE DEVICE WAS RECEIVED FOR ANALYSIS. ATTEMPT WAS MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REP AND A RESPONSE WAS RECEIVED. PER THE REP, THERE WAS A DELAY OF 5 MINUTES. THE DEVICE BROKE WHEN IMPACTING THE TIBIAL COMPONENT, ON AXIS BUT WITH VERY AGGRESSIVE STRIKES. NO PROBLEMS WERE NOTED AT POST-OP FOLLOW-UP. NO OTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS ABOUT THE DEVICE/EVENT/PATIENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. THE PROJECT TEAM REVIEWED PREVIOUS DURABILITY TESTING PERFORMED USING THE DEVICE ((B)(4)) AND FOUND THE USE DID NOT MEET CURRENT DURABILITY TESTING STANDARDS. ADDITIONAL DURABILITY TESTING WAS PERFORMED ON THE DEVICE UNDER (B)(4) USING MORE RIGOROUS IMPACTION FORCE AND HIGHER NUMBER OF CYCLES BUT WAS NOT ABLE TO REPLICATE THE FAILURE. RACR-0006 REV A FOR TOTAL ANKLE INSTRUMENTS RISK ID 9.6 RECOGNIZES THE RISK THAT A DEVICE IS NOT DESIGNED TO WITHSTAND CLINICALLY RELEVANT FORCES LEADING TO DEVICE MALFUNCTION/FRACTURE INTRAOPERATIVELY. THE CURRENTLY EXPECTED OCCURRENCE IS A ¿1¿ BASED ON MITIGATIONS IN PLACE. (B)(4). THE OTHER INCIDENT HAD NO KNOWN DATE BUT COULD HAVE OCCURRED IN 2019. IN THAT EVENT, THE RATE OF THIS FAILURE IS A ¿2¿ PER SOP 701-103-132 REV K. AFTER REVIEWING THE MECHANICAL TESTING (BOTH PREVIOUSLY PERFORMED AND TESTING PERFORMED AS PART OF THE INVESTIGATION), THE FAILURE OF THE DEVICE WAS MOST LIKELY CAUSED BY MISUSE. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: PER CAPA(B)(4), THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF IMPROPER USE OF THE DEVICE DURING SURGICAL PROCEDURE, AS SUPPORTED BY DURABILITY TESTING, REVIEW OF INSTRUMENTATION DESIGN, AND OPERATIVE TECHNIQUE. CORRECTIVE ACTION TAKEN: A COMPLAINT REVIEW FOR THE MODULAR IMPACTOR ARM (351-10-13) WILL BE PERFORMED IN APPROXIMATELY ONE MONTH. SPECIFICS OF MALLETS USED IN THE REPORTED COMPLAINTS WILL BE EVALUATED. PROPER USE OF IMPACTORS TO SITES REPORTING COMPLAINTS WILL BE COMMUNICATED. POTENTIAL TO REMOVE JUNCTION TO THE IMPACTOR HANDLE WILL BE RESEARCHED AND EVALUATED. [THIS SURGEON RECEIVED COMMUNICATION 15 DEC 2020]. THE CAPA IS STILL OPENED AT THE TIME OF THIS INVESTIGATION, AND CORRECTIVE ACTIONS ARE SUBJECT TO CHANGE. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. ¿ CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT, NO PIECES FELL INTO THE SURGICAL SITE. PER CAPA(B)(4), THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF IMPROPER USE OF THE DEVICE DURING SURGICAL PROCEDURE, AS SUPPORTED BY DURABILITY TESTING, REVIEW OF INSTRUMENTATION DESIGN, AND OPERATIVE TECHNIQUE. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00333 | 2022-03-30 | 2019-06-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRUL FEM TRIAL EXTRACTOR | EXACTECH, INC. |  | 02-029-19-1200 |  | IT WAS REPORTED FROM THE US THAT THE FEMORAL TRIAL EXTRACTOR BROKE DURING SURGERY. THE DEVICE WAS RETURNED. THE DEVICE CAN FRACTURE AND DISASSEMBLE. ATTEMPTS WERE MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REP, BUT NO RESPONSES WERE RECEIVED. THERE IS NO OTHER INFORMATION PROVIDED ABOUT THIS DEVICE/EVENT FROM THE CONTACTS. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. FINDINGS: IN THE CASE OF MULTI-PIECE CONSTRUCTION, IT WAS FOUND THAT THE WELD ON THE EXTRACTOR BODY WAS BREAKING. THIS ALLOWED THE COMPONENT PARTS TO DISASSEMBLE. THE ENGINEERING PRINT FOR THE FEMORAL TRIAL EXTRACTOR (02-029-219-1202 REV A) ALLOWED FOR ¿SINGLE OR MULTIPLE PIECE CONSTRUCTION¿ (NOTE 2). HOWEVER, THE PRINT DID NOT SPECIFY HOW TO WELD/ASSEMBLE THE DEVICE IF MULTIPLE-PIECE CONSTRUCTION WAS TO BE USED. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: BASED ON THE CAPA AND CRC, THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF AN INSUFFICIENT WELD ON 02-029-219-1202 REV A, WHICH LED TO WELD FRACTURE AND COMPONENT DISASSEMBLY. CORRECTIVE ACTION TAKEN: PARTS WILL NOT BE ORDERED UNLESS THEY ARE MANUFACTURED AS A SINGLE PIECE CONSTRUCTION. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. ¿ CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO MANUFACTURING ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. THE DEVICE CAN FRACTURE AND DISASSEMBLE. BASED ON (B)(4) AND CAPA(B)(4), THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF AN INSUFFICIENT WELD ON 02-029-219-1202 REV A, WHICH LED TO WELD FRACTURE AND COMPONENT DISASSEMBLY. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00332 | 2022-03-30 | 2019-04-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | MINI HEADED FIXATION PIN REPLACED BY 201-78-15 | EXACTECH, INC. |  | 201-78-05 |  | IT WAS REPORTED FROM OUS THAT THE HOLDING PIN HEADS WERE BROKEN. THE ISSUE WAS NOTICED AFTER THE SURGICAL PROCEDURE. THERE WAS NO DELAY TO SURGERY AND NO PATIENT ADVERSE EVENT REPORTED. THERE WERE NO PIECES THAT FELL INTO THE SURGICAL SITE; THEREFORE, NO PIECES WERE LEFT IN THE PATIENT. THE SURGERY WAS COMPLETED WITHOUT ANY ISSUES. NO ADDITIONAL INFORMATION IS AVAILABLE FROM THE CONTACTS ABOUT THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. FINDINGS: ONE MINI HEADED FIXATION PIN WAS REPORTED IN THE EXPERIENCE REPORT; HOWEVER, TWO PINS WERE RETURNED. BOTH PINS APPEAR TO HAVE THE SAME FAILURE, THEREFORE WERE BOTH EVALUATED IN THIS INVESTIGATION. THE FRACTURED HEADS OF THE MINI HEADED FIXATION PINS APPEAR CONSISTENT WITH APPLYING A BENDING MOMENT TO THE PROXIMAL HEAD OF THE PIN WITH THE PIN PULLER INSTRUMENT. TENSION WAS LIKELY APPLIED TO THE PIN PULLER AT AN ANGLE RATHER THAN AXIALLY AS INTENDED. MOST LIKELY UNDERLYING CAUSE: BASED ON PREVIOUS SIMILAR INVESTIGATIONS, THE BROKEN PINS REPORTED WERE LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE HEAD OF THE PIN WITH THE PIN PULLER, WHICH ALLOWED FOR THE PROXIMAL RIM TO FRACTURE. RMR REVIEW: THE RISK IS CAPTURED. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. ¿ CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES. NO PATIENT ADVERSE EVENT THERE WERE NO PARTS/PIECES THAT FELL INTO THE SURGICAL SITE. AN INVESTIGATION WAS CONDUCTED: BASED ON PREVIOUS SIMILAR INVESTIGATIONS, THE BROKEN PINS REPORTED WERE LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE HEAD OF THE PIN WITH THE PIN PULLER, WHICH ALLOWED FOR THE PROXIMAL RIM TO FRACTURE. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00331 | 2022-03-30 | 2019-05-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | GLND SKID | EXACTECH, INC. |  | 315-03-01 |  | IT WAS REPORTED FROM OUS THAT DURING A PRIMARY ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH INSTRUMENT. THE GLENOID SKID BROKE WHILE BEING USED BY THE SURGEON. ITEM WAS RETURNED TO THE AGENCY IN 2 PARTS. IT WAS STATED AS ¿UNKNOWN¿ IF ANY PARTS/PIECES FELL INTO THE SURGICAL SITE. IT WAS CONFIRMED THAT NO PIECES WERE LEFT IN THE PATIENT. THE AGENT WAS NOT PRESENT AT SURGERY. THERE WAS NO DELAY AND NO CLINICAL CONSEQUENCE TO THE PATIENT. NO PATIENT DEMOGRAPHICS WERE PROVIDED. THE PATIENT WENT HOME WITH NO COMPLICATIONS. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED BY THE CONTACTS ABOUT THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. VISUAL EVALUATION CONDITIONS/FINDINGS: THE BROKEN GLENOID SKID APPEARS CONSISTENT WITH APPLYING A BENDING MOMENT DURING USE. ALL OTHER ASPECTS OF THE DEVICES APPEAR UNREMARKABLE AND CONSISTENT WITH NORMAL USE. DESIGN-RELATED ISSUES: THE DESIGN OF THE GLENOID SKID HAS BEEN IN THE FIELD SINCE 2007. EXACTECH IS AWARE OF 4 SIMILAR COMPLAINT REPORTS INVOLVING THIS DEVICE SINCE 2007. BECAUSE THE GLENOID SKID IS NOT USED IN EVERY EQUINOXE REVERSE TOTAL SHOULDER SURGERY, THE NUMBER OF THESE DEVICES IN THE FIELD WAS COMPARED TO THE NUMBER OF A MORE COMMONLY USED DEVICE (38MM REVERSE PILOT TIP REAMER). FROM THIS COMPARISON, IT WAS ESTIMATED THAT THIS DEVICE IS USED IN APPROXIMATELY 30% OF REVERSE TOTAL SHOULDER SURGERIES. THEREFORE, 30% OF GLENOID BASEPLATE SALES WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4). THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF (B)(4) UNITS, WHICH HAVE BEEN IN THE FIELD SINCE 2018. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS GLENOID SKID, 750-2002-039-RMR-INSTRUMENT ADDITIONS REV G, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 17. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO ONE END OF THE INSTRUMENT, WHICH LED TO FRACTURE OF THE DEVICE. HOWEVER, NO INFORMATION REGARDING ITS USE DURING SURGERY WAS PROVIDED. IFU 700-096-181: INSTRUMENT INSPECTION. VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES. NO PATIENT ADVERSE EVENT, THE GLENOID SKID BROKE WHILE BEING USED BY THE SURGEON. NO PIECES WERE LEFT IN THE PATIENT, IT IS UNKNOWN IF ANY PIECES/PARTS FELL INTO THE SURGICAL SITE. AN INVESTIGATION WAS CONDUCTED: THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO ONE END OF THE INSTRUMENT, WHICH LED TO FRACTURE OF THE DEVICE. HOWEVER, NO INFORMATION REGARDING ITS USE DURING SURGERY WAS PROVIDED. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00330 | 2022-03-30 | 2019-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRUL FEM TRIAL EXTRACTOR | EXACTECH, INC. |  | 02-029-19-1200 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE INSTRUMENTATION. THE SURGEON WAS BACKSLAPPING WITH THE SLAP HAMMER ¿REALLY HARD, MAYBE A LITTLE OFF AXIS, AND THE TWO PRONGS CAME OFF THE MOUNT OF THE FEMORAL EXTRACTOR¿. THE SURGEON WAS ABLE TO EXTRACT THE BASE TRIAL COMPONENT, ¿AND WE HAD A BACK-UP IF HE WANTED TO TRIAL AGAIN¿. THE DEVICE WAS RECEIVED FOR ANALYSIS. THE REP WAS PRESENT AT THE TIME OF SURGERY. PER THE REPRESENTATIVE NO PIECES WERE IN THE SURGICAL SITE. THE CASE WAS DELAYED BY 5 MINUTES AS THEY KEPT TRYING TO REMOVE THE PART. NO INJURY OR ADVERSE EVENT. THE OUTCOME WAS ¿GOOD¿. THERE IS NO ADDITIONAL INFORMATION FROM THE CONTACTS FOR THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. FINDINGS: THE WELD ON 02-029-219-1202 REV A IN THE CASE OF MULTI-PIECE CONSTRUCTION WAS BREAKING. THIS ALLOWED THE COMPONENT PARTS TO DISASSEMBLE. THE PRINT DOES NOT SPECIFY A TYPE OF WELD TO BE USED IN THE CASE OF MULTI-PIECE CONSTRUCTION. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: BASED ON THE CAPA AND CRC, THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF AN INSUFFICIENT WELD ON 02-029-219-1202, WHICH LED TO WELD FRACTURE AND COMPONENT DISASSEMBLY. CORRECTIVE ACTION TAKEN: PARTS WILL NOT BE ORDERED UNLESS THEY ARE SINGLE PIECE CONSTRUCTION. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. ¿ CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO MANUFACTURING ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. BROKEN PIECES DID NOT FALL INTO THE SURGICAL SITE. BASED ON (B)(4) THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF AN INSUFFICIENT WELD ON 02-029-219-1202, WHICH LED TO WELD FRACTURE AND COMPONENT DISASSEMBLY. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00329 | 2022-03-30 | 2019-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LOGIC CR INSERT TRIAL SIZE 3.5, NEUTRAL, 9 MM | EXACTECH, INC. |  | 02-013-57-3509 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE TRIAL. THE TRAIL WAS DAMAGED BY A SURGEON USING A HOHMANN RETRACTOR TO REMOVE IT. NO ISSUE TO THE PATIENT OCCURRED DURING THIS CASE. THE REPRESENTATIVE WAS PRESENT. THE DEVICE WAS RETURNED. PER THE REPRESENTATIVE, NO PIECE OF THE TRIAL WAS LEFT IN OR RETRIEVED FROM THE PATIENT. THE TRIAL WAS REMOVED IN AN ENTIRE PIECE. NO DELAY TO SURGERY. NO PROBLEMS POST-OP. PATIENT DEMOGRAPHICS ARE NOT AVAILABLE. THERE IS NO ADDITIONAL INFORMATION ABOUT THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DESIGN-RELATED ISSUES: THE DESIGN OF THE LOGIC CR TIBIAL INSERT TRIAL HAS BEEN IN THE FIELD SINCE 2012. EXACTECH IS AWARE OF 6 COMPLAINT REPORTS INVOLVING A FRACTURED TIBIAL INSERT TRIAL EXTRACTION FEATURE SINCE 2012. SALES DATA FOR TIBIAL INSERTS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 14 UNITS, WHICH HAS BEEN IN THE FIELD SINCE 2014. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THESE LOGIC TIBIAL INSERT TRIALS, 750-2010-016-RMR-INSTRUMENTS REV -, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 13. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR. THE BROKEN TIBIAL INSERT TRIAL REPORTED WAS LIKELY THE RESULT OF APPLYING EXCESS FORCE TO THE EXTRACTION FEATURE WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FAILURE. IFU 700-096-181: INSTRUMENT INSPECTION. VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT DUE TO THE FRACTURED TRIAL. AN INVESTIGATION WAS CONDUCTED: THE BROKEN TIBIAL INSERT TRIAL REPORTED WAS LIKELY THE RESULT OF APPLYING EXCESS FORCE TO THE EXTRACTION FEATURE WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FAILURE. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00328 | 2022-03-30 | 2019-03-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | DRILL GUIDE | EXACTECH, INC. |  | 181-00-05 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE DRILL GUIDE. THE INSTRUMENT BROKE WHILE ATTEMPTING TO DRILL THE SCREW HOLE THROUGH. THE INSERTED CUP WAS REMOVED AND ANOTHER DRILL GUIDE IN THE SET WAS USED. THERE WAS NO DELAY TO SURGERY. THE PATIENT HAS A HISTORY OF LEFT TOTAL HIP ARTHROPLASTY. NO COMPROMISE TO PATIENT, WHO LEFT THE OPERATING ROOM IN NORMAL TIME. THE SHEARED SURFACES APPEAR CONSISTENT WITH APPLYING A BENDING MOMENT. THE REPRESENTATIVE WAS NOT PRESENT AT THE TIME OF SURGERY. THE DEVICE WAS RETURNED. MULTIPLE ATTEMPTS WERE MADE FOR ADDITIONAL EVENT/PATIENT INFORMATION, NO RESPONSES. THERE IS NO ADDITIONAL INFORMATION FROM THE CONTACTS ABOUT THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. FINDINGS: THE SHEARED SURFACES APPEAR CONSISTENT WITH APPLYING A BENDING MOMENT TO THE DRILL GUIDE WHILE USING A DRILL, CAUSING THE WELD BETWEEN THE DRILL GUIDE TIP AND THE DRILL GUIDE SHAFT TO BREAK. MOST LIKELY UNDERLYING CAUSE: BASED ON PREVIOUS INVESTIGATIONS, THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE DRILL GUIDE WHILE USING A DRILL, WHICH LED TO FRACTURE OF THE WELD BETWEEN THE DRILL GUIDE TIP AND THE DRILL GUIDE SHAFT. RMR REVIEW: THE RISK IS CAPTURED. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. ¿ CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. THE SHEARED SURFACES APPEAR CONSISTENT WITH APPLYING A BENDING MOMENT. AN INVESTIGATION WAS CONDUCTED: BASED ON PREVIOUS INVESTIGATIONS, THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE DRILL GUIDE WHILE USING A DRILL, WHICH LED TO FRACTURE OF THE WELD BETWEEN THE DRILL GUIDE TIP AND THE DRILL GUIDE SHAFT. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00327 | 2022-03-30 | 2019-03-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | DRILL GUIDE | EXACTECH, INC. |  | 181-00-05 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE DRILL GUIDE. THE INSTRUMENT BROKE WHILE ATTEMPTING TO DRILL THE SCREW HOLE THROUGH. THE INSERTED CUP WAS REMOVED AND ANOTHER DRILL GUIDE IN THE SET WAS USED. THERE WAS NO DELAY TO SURGERY. NO COMPROMISE TO PATIENT, WHO LEFT THE OPERATING ROOM IN NORMAL TIME. THE SHEARED SURFACES APPEAR CONSISTENT WITH APPLYING A BENDING MOMENT. THE REPRESENTATIVE WAS PRESENT AT THE TIME OF SURGERY. THE DEVICE WAS RETURNED. MULTIPLE ATTEMPTS WERE MADE FOR ADDITIONAL EVENT/PATIENT INFORMATION, NO RESPONSES. THERE IS NO ADDITIONAL INFORMATION FROM THE CONTACTS ABOUT THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. FINDINGS: THE SHEARED SURFACES APPEAR CONSISTENT WITH APPLYING A BENDING MOMENT TO THE DRILL GUIDE WHILE USING A DRILL, CAUSING THE WELD BETWEEN THE DRILL GUIDE TIP AND THE DRILL GUIDE SHAFT TO BREAK. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: BASED ON PREVIOUS INVESTIGATIONS, THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE DRILL GUIDE WHILE USING A DRILL, WHICH LED TO FRACTURE OF THE WELD BETWEEN THE DRILL GUIDE TIP AND THE DRILL GUIDE SHAFT. RMR REVIEW: THE RISK IS CAPTURED. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. ¿ CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. THE SHEARED SURFACES APPEAR CONSISTENT WITH APPLYING A BENDING MOMENT. AN INVESTIGATION WAS CONDUCTED: BASED ON PREVIOUS INVESTIGATIONS, THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE DRILL GUIDE WHILE USING A DRILL, WHICH LED TO FRACTURE OF THE WELD BETWEEN THE DRILL GUIDE TIP AND THE DRILL GUIDE SHAFT. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00325 | 2022-03-30 | 2019-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | FCT GAP CHECK TL SZ 3-4 | EXACTECH, INC. |  | 351-10-13 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC ANKLE SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE INSTRUMENT. DURING SURGERY THE ANKLE PEG PUNCH DID BREAK DURING IMPACTING THE CENTER PEG PUNCH. ¿IN HARD BONE, SOMETIMES IT¿S ALMOST IMPOSSIBLE TO IMPACT THE CENTER PEG PUNCH¿ AND EVEN THE ¿PERIPHERAL PEG PUNCH¿. THE SURGEON ¿NEEDED TO HIT SO HARD¿ THAT DURING THIS SURGERY THE ¿ANKLE PEG PUNCH¿ (BOWED ARM) BROKE. NO ISSUE WITH THE PATIENT. PATIENT LEFT THE OR WITHOUT ANY COMPLICATION. THE DEVICE WAS RETURNED. THE REPRESENTATIVE WAS NOT PRESENT AT THE TIME OF SURGERY. IMAGES WERE PROVIDED. MULTIPLE ATTEMPTS WERE MADE FOR ADDITIONAL INFORMATION FROM THE REPRESENTATIVE, BUT NO RESPONSES WERE RECEIVED. THERE IS NO ADDITIONAL FOR THIS EVENT FROM THE CONTACTS. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. THE DEVICE WAS AVAILABLE FOR ANALYSIS. EXACTECH¿S ENGINEERING MANAGER-LOWER EXTREMITIES, SPOKE WITH THE SURGEON, WHO REPORTED THAT HE IMPACTED THE BONE WITH THIS INSTRUMENT FOR APPROXIMATELY 30 MINUTES BEFORE IT FRACTURED. THE PROJECT TEAM REVIEWED PREVIOUS DURABILITY TESTING PERFORMED USING THE DEVICE (TR-2018-0869) AND FOUND THE USE DID NOT MEET CURRENT DURABILITY TESTING STANDARDS. ADDITIONAL DURABILITY TESTING WAS PERFORMED ON THE DEVICE UNDER (B)(4) USING MORE RIGOROUS IMPACTION FORCE AND HIGHER NUMBER OF CYCLES BUT WAS NOT ABLE TO REPLICATE THE FAILURE. RACR-0006 REV A FOR TOTAL ANKLE INSTRUMENTS RISK ID 9.6 RECOGNIZES THE RISK THAT A DEVICE IS NOT DESIGNED TO WITHSTAND CLINICALLY RELEVANT FORCES LEADING TO DEVICE MALFUNCTION/FRACTURE INTRAOPERATIVELY. THE CURRENTLY EXPECTED OCCURRENCE IS A ¿1¿ BASED ON MITIGATIONS IN PLACE. (B)(4). THE OTHER INCIDENT HAD NO KNOWN DATE BUT COULD HAVE OCCURRED IN 2019. IN THAT EVENT, THE RATE OF THIS FAILURE IS A ¿2¿ PER SOP (B)(4) REV K. AFTER REVIEWING THE MECHANICAL TESTING (BOTH PREVIOUSLY PERFORMED AND TESTING PERFORMED AS PART OF THE INVESTIGATION), THE FAILURE OF THE DEVICE WAS MOST LIKELY CAUSED BY MISUSE. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: PER CAPA(B)(4), THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF IMPROPER USE OF THE DEVICE DURING SURGICAL PROCEDURE, AS SUPPORTED BY DURABILITY TESTING, REVIEW OF INSTRUMENTATION DESIGN, AND OPERATIVE TECHNIQUE. CORRECTIVE ACTION TAKEN: A COMPLAINT REVIEW FOR THE MODULAR IMPACTOR ARM (351-10-13) WILL BE PERFORMED IN APPROXIMATELY ONE MONTH. SPECIFICS OF MALLETS USED IN THE REPORTED COMPLAINTS WILL BE EVALUATED. PROPER USE OF IMPACTORS TO SITES REPORTING COMPLAINTS WILL BE COMMUNICATED. [THE SURGEON/FACILITY RECEIVED COMMUNICATION ON 8 DEC 2020] POTENTIAL TO REMOVE JUNCTION TO THE IMPACTOR HANDLE WILL BE RESEARCHED AND EVALUATED. THE CAPA IS STILL OPENED AT THE TIME OF THIS INVESTIGATION, AND CORRECTIVE ACTIONS ARE SUBJECT TO CHANGE. REFER TO CAPA(B)(4) FOR THIS MOST UP TO DATE INFORMATION. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. ¿ CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. PER CAPA(B)(4), THE BROKEN INSTRUMENT REPORTED IN (B)(4) WAS LIKELY THE RESULT OF IMPROPER USE OF THE DEVICE DURING SURGICAL PROCEDURE, AS SUPPORTED BY DURABILITY TESTING, REVIEW OF INSTRUMENTATION DESIGN, AND OPERATIVE TECHNIQUE. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00318 | 2022-03-30 | 2019-01-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LOGIC TIBIA PS MOD INSRT TRIAL SZ 2.5 15MM | EXACTECH, INC. |  | 02-013-35-2515 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY A SURGEON EXPERIENCED AN ISSUE WITH THE TRIALS. ONE INSERT TRIAL BROKE AFTER THE OTHER TRIAL BROKE. WHEN THE SURGEON WAS TRYING TO TRIAL WITH 13,15MM OF SIZE 2.5 LOGIC INSERT TRIAL, 13MM WAS NOT FIT WITH THE PATIENT¿S GAP, SO HE EXTRACTED THE 13MM OF 2.5 INSERT TRIAL. THE 13MM TRIAL WAS BROKEN EVEN WHEN THE SURGEON WAS USING THE TIBIAL TRIAL INSERT HANDLE (231-04-03). THEN 15MM WAS BROKEN RIGHT AFTER TRIAL. THIS CAUSED A 10- MINUTE DELAY BUT HAD NO ADVERSE EFFECT ON THE PATIENT AND THE PATIENT LEFT THE OPERATING ROOM IN STABLE CONDITION. DEVICES WERE NOT RETURNED, THEY WERE LOST. NO PARTS/PIECES FELL INTO THE SURGICAL SITE. NO ADDITIONAL INFORMATION AVAILABLE. 1 OF 2 MDRS. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: WEIGHT, ETHNICITY, RACE. WHILE SOME TRIAL FAILURES WERE ¿EXPECTED¿ AND RELATED TO NORMAL USE AND WEAR OVER TIME, THERE ARE REPORTS THAT TRIALS HAVE BEEN ¿FOUND BROKEN UPON ARRIVAL¿ OR ¿BROKEN AFTER AUTOCLAVING.¿ EXAMINATION OF THE FRACTURED SURFACES REVEALED A BRITTLE FAILURE MECHANISM THAT INITIATES ON THE SURFACE OF THE INSERT IN VARIED LOCATIONS. THESE FAILURES DO NOT APPEAR TO BE ISOLATED TO ONE SIZE OR ITS ASSOCIATED MATERIAL COLOR. A REVIEW OF THE MANUFACTURING DATES, DEVICE DIMENSIONS, AND MATERIAL WAS PERFORMED. NO OBVIOUS CORRELATIONS OR ABNORMALITIES WERE FOUND. A VISIT TO THE MANUFACTURING SITE REVEALED NOTHING SIGNIFICANT. TO HELP ESTABLISH A ROOT CAUSE, IN-HOUSE TESTING WAS PERFORMED. THE FIRST TEST WAS TO EVALUATE THE EFFECT OF THERMAL QUENCHING AFTER AUTOCLAVE STERILIZATION ON THE PROPENSITY FOR CRACKING AND PROPAGATION FOR LOGIC INSERT TRIALS. THE SAMPLE TRIALS USED WERE EXPOSED TO TWENTY THERMAL CYCLES WHICH INCLUDED A THREE-MINUTE STERILIZATION CYCLE, A TWENTY- MINUTE DRYING PERIOD, AND A SUBSEQUENT QUENCHING IN AN ICE BATH. FOLLOWING TESTING, CRACK PROPAGATION WAS NOT OBSERVED ON ANY OF THE INSERT TRIALS. REFER TO TR-2018-0219 FOR FURTHER DETAILS. THE SAME SAMPLES WERE THEN USED TO EVALUATE THE EFFECT OF SIMULATED SURGEON USAGE AND REPROCESSING ON THE PROPENSITY OF CRACK INITIATION AND PROPAGATION FOR LOGIC INSERT TRIALS. DURING THIS STUDY, THE TRIALS WERE EXPOSED TO TWENTY THERMAL AND MECHANICAL CYCLES. EACH CYCLE CONSISTED OF AUTOCLAVE STERILIZATION AND COLD-WATER QUENCHING FOLLOWED BY INSERTION INTO A TIBIAL TRAY, LOADING TO 390 N, AND REMOVAL FROM THE TIBIAL TRAY UNDER APPROXIMATELY 100 N USING EITHER A LPI TIBIAL INSERT TRIAL HANDLE OR AN OSTEOTOME. FOLLOWING TESTING, CRACK PROPAGATION WAS NOT OBSERVED. REFER TO TR-2018-0677 FOR FURTHER DETAILS. MOST LIKELY CAUSE: A ROOT CAUSE HAS NOT BEEN ESTABLISHED AT THE TIME OF THIS INVESTIGATION. THE PERFORMED TESTING HAS BEEN UNABLE TO RECREATE THE REPORTED CRACKING. AS A RESULT OF CRC2018-01-22-02, (B)(4) WAS INITIATED TO FURTHER EVALUATE THE ROOT CAUSE. IFU 700-096-181: INSTRUMENT INSPECTION. VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES. NO PATIENT ADVERSE EVENT. THE TRIALS BROKE DURING SURGERY. NO PARTS/PIECES FELL INTO THE SURGICAL SITE. A ROOT CAUSE HAS NOT BEEN ESTABLISHED AT THE TIME OF THIS INVESTIGATION. THE PERFORMED TESTING HAS BEEN UNABLE TO RECREATE THE REPORTED CRACKING. AS A RESULT OF CRC2018-01-22-02, (B)(4) WAS INITIATED TO FURTHER EVALUATE THE ROOT CAUSE. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00323 | 2022-03-30 | 2019-02-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LPI IMPACTOR HANDLE | EXACTECH, INC. |  | 213-65-00 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE LPI IMPACTOR HANDLE. THE HANDLE BROKE WHEN IMPACTING THE FEMORAL. IT WAS A CLEAN BREAK AT THE WEAKEST PART OF THE HANDLE. THIS DID NOT CAUSE A DELAY TO SURGERY. THE PATIENT IS FINE. THE REP WAS PRESENT AT THE TIME OF SURGERY. NO PIECES FELL INTO THE SITE. NO ADDITIONAL INFORMATION HAS BEEN MADE AVAILABLE. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. THE DEVICE WAS AVAILABLE FOR ANALYSIS. FINDINGS: IT WAS REPORTED FROM THE FIELD THAT THE LPI IMPACTOR HANDLE BROKE DURING IMPACTION OF THE FEMORAL COMPONENT. PER (B)(4), A REVIEW OF THE LPI IMPACTOR HANDLES (213-65-00) CITED IN COMPLAINTS FOR BROKEN OR CRACKED HANDLES SHOWED A CONSISTENT FAILURE MODE. A REVIEW OF THE DEVICE ORDER HISTORY DID NOT SHOW A TREND IN THOSE LOTS WITH COMPLAINT DEVICES, NOR DID A REVIEW OF DESIGN CHANGES SHOW CHANGES MADE TO THE DESIGN WHICH COULD CONTRIBUTE TO THE COMPLAINT. THE OVERALL OCCURRENCE FOR BROKEN OR CRACKED PLASTIC HANDLES IS LOW. SIMILAR DEVICES ALSO HAVE LOW OCCURRENCE FOR BROKEN OR CRACKED HANDLES. REFER TO (B)(4) FOR ADDITIONAL DETAILS. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: PER (B)(4), THE BROKEN LPI IMPACTOR HANDLE REPORTED IN (B)(4) WAS LIKELY THE RESULT OF THE HIGH IMPACTION FORCES REQUIRED DURING IMPLANT SURGERIES. CORRECTIVE ACTION TAKEN: (B)(4) DETERMINED THAT DUE TO THE LOW RATE OF OCCURRENCE, FIELD ACTION IS NOT REQUIRED AT THIS TIME. EXACTECH WILL CONTINUE TO TRACK AND TREND FAILURES OF THIS NATURE AND WILL ESCALATE THE ISSUE AS REQUIRED BY SOP (B)(4) (COMPLAINT HANDLING) AND SOP (B)(4) (CORRECTIVE AND PREVENTATIVE ACTIONS). IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. ¿ CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. NO PARTS OR PIECES OF THE FRACTURED DEVICE FELL INTO THE SURGICAL SITE. PER (B)(4), THE BROKEN LPI IMPACTOR HANDLE REPORTED WAS LIKELY THE RESULT OF THE HIGH IMPACTION FORCES REQUIRED DURING IMPLANT SURGERIES. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00322 | 2022-03-30 | 2018-11-06 | I, n, j, u, r, y | Manufacturer report | EXACTECH | MDLR RV RMR 38MM PLT TIP | EXACTECH, INC. |  | 321-25-38 |  | IT WAS REPORTED FROM THE OUS THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED A BROKEN PILOT TIP REAMER. THE TIGHT SHOULDER RETRACTED OFF THE HUMERAL HEAD AND THE 38MM TIP BROKE. THE SURGEON TRIED AGAIN AFTER SOME MORE RESECTION AND HAD TO USER THE BIGGER 42MM THE TIP BROKE AGAIN. THERE WERE NO PIECES THAT FELL INTO THE SURGICAL SITE, NO DELAY, AND NO ADVERSE EVENT TO THE PATIENT. THE AGENT WAS PRESENT AT THE TIME OF SURGERY. THE PATIENT'S HEALTH STABLE LEAVING THE OR. NO ADDITIONAL INFORMATION. 1 OF 2 MDRS. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICES WERE RETURNED FOR ANALYSIS. BASED ON THE INVESTIGATION CONDUCTED UNDER CAPA (B)(4), IT WAS DETERMINED THAT THE USER INSTRUCTION WAS NOT ADEQUATE TO INFORM THE SURGEON OF THE APPROPRIATE USE OF THE REAMER. THE FORESEEABLE MISUSE OF THE GLENOID REAMER THAT COULD RESULT IN FRACTURE OF THE PILOT TIP ARE: IF THE USER APPLIES A BENDING MOMENT TO THE PILOT DIP DURING REAMING, WHICH WOULD REAM THE GLENOID ASYMMETRICALLY AND APPLY A TORQUE TO THE PILOT TIP, POTENTIALLY LEADING TO FRACTURE. IF THE USER APPLIES A BENDING MOMENT TO THE PILOT TIP BY USING THE GLENOID REAMER TO RETRACT THE HUMERAL HEAD TO HELP GAIN EXPOSURE AND ACCESS TO THE GLENOID, WHICH WOULD APPLY A TORQUE TO THE PILOT TIP, POTENTIALLY LEADING TO FRACTURE. CORRECTIVE ACTION TAKEN AS PART OF CAPA (B)(4), THE FOLLOWING CORRECTIVE ACTIONS WERE TAKEN: THE SURGICAL TECHNIQUE WAS UPDATED TO ADVISE SURGEONS TO AVOID APPLYING A BENDING TORQUE TO THE PILOT TIP REAMER OR USING THE REAMER TO RETRACT THE HUMERAL HEAD AS THIS MAY CAUSE FRACTURE OF THE PILOT TIP. (DCRSP-17-135, -136, -137, -138). UPDATE THE RMR TO INCLUDE THE RISK OF PILOT TIP REAMER FRACTURE AS A RESULT OF BENDING OR ITS FORESEEABLE MISUSE AS A RETRACTOR. (DCRSP-17-172). PER (B)(4), A FIELD ADVISORY NOTICE WAS SENT TO THE USERS NOTIFYING THEM OF THE ISSUE AND ALERTING THEM TO THE CHANGE IN THE OPERATIVE TECHNIQUE (EXACTECH RECALL NUMBER 1038671-04/07/2017-003-C, FDA RECALL NUMBERS Z-2663-2017 THROUGH Z-2668-2017, MAY 2017). THE RECALL WAS TERMINATED BY THE FDA ON (B)(6) 2019. CAPA (B)(4) WAS INITIATED TO TRACK REPORTED EVENTS THROUGH 2021 AND RECONVENE THE CRC IF A THRESHOLD OF >35 TOTAL EVENTS OR >3 EVENTS WITH REPORTED PATIENT HARM BE MET. AS A RESULT OF MONITORING UNDER (B)(4) ((B)(6) 2019) AND CAPA (B)(4), IT WAS NOTED THAT THE THRESHOLD OF EVENTS FOR 2020 HAS BEEN EXCEEDED. THIS ISSUE HAS BEEN RE-ESCALATED FOR EVALUATION. AS DETERMINED BY HHE2020-09-11-01 (B)(4) AND (B)(4), THE RISK FALLS WITHIN THE ACCEPTABLE LEVEL. THE COMMITTEE HAS DETERMINED THAT NO ADDITIONAL FIELD ACTION IS REQUIRED AT THIS TIME. EXACTECH WILL CONTINUE MONITOR THE EVENTS THROUGH CAPA (B)(4), WHICH WILL DRIVE ANY FURTHER HHE UPDATES AND/OR CRC ESCALATION. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT OR DELAY TO SURGERY. BOTH REAMER TIPS BROKE, NO PIECES FELL INTO THE SURGICAL SITE. BASED ON THE INVESTIGATION CONDUCTED UNDER CAPA (B)(4), IT WAS DETERMINED THAT THE USER INSTRUCTION WAS NOT ADEQUATE TO INFORM THE SURGEON OF THE APPROPRIATE USE OF THE REAMER. THE FORESEEABLE MISUSE OF THE GLENOID REAMER THAT COULD RESULT IN FRACTURE OF THE PILOT TIP ARE: IF THE USER APPLIES A BENDING MOMENT TO THE PILOT DIP DURING REAMING, WHICH WOULD REAM THE GLENOID ASYMMETRICALLY AND APPLY A TORQUE TO THE PILOT TIP, POTENTIALLY LEADING TO FRACTURE. IF THE USER APPLIES A BENDING MOMENT TO THE PILOT TIP BY USING THE GLENOID REAMER TO RETRACT THE HUMERAL HEAD TO HELP GAIN EXPOSURE AND ACCESS TO THE GLENOID, WHICH WOULD APPLY A TORQUE TO THE PILOT TIP, POTENTIALLY LEADING TO FRACTURE. | Y | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00321 | 2022-03-30 | 2018-10-16 | I, n, j, u, r, y | Manufacturer report | EXACTECH | MDLR RV RMR 38MM PLT TIP | EXACTECH, INC. |  | 321-25-38 |  | IT WAS REPORTED BY THE OUS THAT THE REAMER WAS RETURNED WITH A BROKEN TIP. THE PATIENT¿S HEALTH WAS STABLE LEAVING THE OPERATING ROOM. THERE WAS NO DELAY, AND NO PIECES WERE LEFT IN THE PATIENT. A REP WAS PRESENT AT THE TIME OF SURGERY. THE DEVICE WAS RECEIVED ON (B)(6) 2019. RESPONSE ON (B)(6) 2021 STATES THAT WERE IDENTIFIED DURING ROUTINE WAREHOUSE INSTRUMENT INSPECTION. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION.THE DEVICE WAS AVAILABLE FOR ANALYSIS. FINDINGS/ROOT CAUSE: BASED ON THE INVESTIGATION CONDUCTED UNDER CAPA(B)(4), IT WAS DETERMINED THAT THE USER INSTRUCTION WAS NOT ADEQUATE TO INFORM THE SURGEON OF THE APPROPRIATE USE OF THE REAMER. THE FORESEEABLE MISUSE OF THE GLENOID REAMER THAT COULD RESULT IN FRACTURE OF THE PILOT TIP ARE: IF THE USER APPLIES A BENDING MOMENT TO THE PILOT DIP DURING REAMING, WHICH WOULD REAM THE GLENOID ASYMMETRICALLY AND APPLY A TORQUE TO THE PILOT TIP, POTENTIALLY LEADING TO FRACTURE. IF THE USER APPLIES A BENDING MOMENT TO THE PILOT TIP BY USING THE GLENOID REAMER TO RETRACT THE HUMERAL HEAD TO HELP GAIN EXPOSURE AND ACCESS TO THE GLENOID, WHICH WOULD APPLY A TORQUE TO THE PILOT TIP, POTENTIALLY LEADING TO FRACTURE.CORRECTIVE ACTION TAKEN: AS PART OF CAPA(B)(4), THE FOLLOWING CORRECTIVE ACTIONS WERE TAKEN: THE SURGICAL TECHNIQUE WAS UPDATED TO ADVISE SURGEONS TO AVOID APPLYING A BENDING TORQUE TO THE PILOT TIP REAMER OR USING THE REAMER TO RETRACT THE HUMERAL HEAD AS THIS MAY CAUSE FRACTURE OF THE PILOT TIP. (B)(4). UPDATE THE RMR TO INCLUDE THE RISK OF PILOT TIP REAMER FRACTURE AS A RESULT OF BENDING OR ITS FORESEEABLE MISUSE AS A RETRACTOR. ((B)(4)). PER (B)(4), A FIELD ADVISORY NOTICE WAS SENT TO THE USERS NOTIFYING THEM OF THE ISSUE AND ALERTING THEM TO THE CHANGE IN THE OPERATIVE TECHNIQUE (EXACTECH RECALL NUMBER 1038671-04/07/2017-003-C, FDA RECALL NUMBERS Z-2663-2017 THROUGH Z-2668-2017, MAY 2017). THE RECALL WAS TERMINATED BY THE FDA ON JULY 26, 2019. CAPA(B)(4) WAS INITIATED TO TRACK REPORTED EVENTS THROUGH 2021 AND RECONVENE THE CRC IF A THRESHOLD OF >35 TOTAL EVENTS OR >3 EVENTS WITH REPORTED PATIENT HARM BE MET. AS A RESULT OF MONITORING UNDER (B)(4) ((B)(6) 2019) AND CAPA(B)(4), IT WAS NOTED THAT THE THRESHOLD OF EVENTS FOR 2020 HAS BEEN EXCEEDED. THIS ISSUE HAS BEEN RE-ESCALATED FOR EVALUATION. AS DETERMINED BY HHE2020-09-11-01 (B)(4) AND (B)(4), THE RISK FALLS WITHIN THE ACCEPTABLE LEVEL. THE COMMITTEE HAS DETERMINED THAT NO ADDITIONAL FIELD ACTION IS REQUIRED AT THIS TIME. EXACTECH WILL CONTINUE MONITOR THE EVENTS THROUGH CAPA(B)(4), WHICH WILL DRIVE ANY FURTHER HHE UPDATES AND/OR CRC ESCALATION.IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION.IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION.IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES.THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. THE REAMER WAS RETURNED WITH A BROKEN TIP.BASED ON THE INVESTIGATION CONDUCTED UNDER CAPA(B)(4), IT WAS DETERMINED THAT THE USER INSTRUCTION WAS NOT ADEQUATE TO INFORM THE SURGEON OF THE APPROPRIATE USE OF THE REAMER. THE FORESEEABLE MISUSE OF THE GLENOID REAMER THAT COULD RESULT IN FRACTURE OF THE PILOT TIP ARE: IF THE USER APPLIES A BENDING MOMENT TO THE PILOT DIP DURING REAMING, WHICH WOULD REAM THE GLENOID ASYMMETRICALLY AND APPLY A TORQUE TO THE PILOT TIP, POTENTIALLY LEADING TO FRACTURE. IF THE USER APPLIES A BENDING MOMENT TO THE PILOT TIP BY USING THE GLENOID REAMER TO RETRACT THE HUMERAL HEAD TO HELP GAIN EXPOSURE AND ACCESS TO THE GLENOID, WHICH WOULD APPLY A TORQUE TO THE PILOT TIP, POTENTIALLY LEADING TO FRACTURE. | Y | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00319 | 2022-03-30 | 2019-01-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LOGIC TIBIA PS MOD INSRT TRIAL SZ 2.5 13MM | EXACTECH, INC. |  | 02-013-35-2513 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY A SURGEON EXPERIENCED AN ISSUE WITH THE TRIALS. ONE INSERT TRIAL BROKE AFTER THE OTHER TRIAL BROKE. WHEN THE SURGEON WAS TRYING TO TRIAL WITH 13,15MM OF SIZE 2.5 LOGIC INSERT TRIAL, 13MM WAS NOT FIT WITH THE PATIENT¿S GAP, SO HE EXTRACTED THE 13MM OF 2.5 INSERT TRIAL. THE 13MM TRIAL WAS BROKEN EVEN WHEN THE SURGEON WAS USING THE TIBIAL TRIAL INSERT HANDLE (231-04-03). THEN 15MM WAS BROKEN RIGHT AFTER TRIAL. THIS CAUSED A 10- MINUTE DELAY BUT HAD NO ADVERSE EFFECT ON THE PATIENT AND THE PATIENT LEFT THE OPERATING ROOM IN STABLE CONDITION. DEVICES WERE NOT RETURNED, THEY WERE LOST. NO PARTS/PIECES FELL INTO THE SURGICAL SITE. NO ADDITIONAL INFORMATION AVAILABLE. 2 OF 2 MDRS. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. WHILE SOME TRIAL FAILURES WERE ¿EXPECTED¿ AND RELATED TO NORMAL USE AND WEAR OVER TIME, THERE ARE REPORTS THAT TRIALS HAVE BEEN ¿FOUND BROKEN UPON ARRIVAL¿ OR ¿BROKEN AFTER AUTOCLAVING.¿ EXAMINATION OF THE FRACTURED SURFACES REVEALED A BRITTLE FAILURE MECHANISM THAT INITIATES ON THE SURFACE OF THE INSERT IN VARIED LOCATIONS. THESE FAILURES DO NOT APPEAR TO BE ISOLATED TO ONE SIZE OR ITS ASSOCIATED MATERIAL COLOR. A REVIEW OF THE MANUFACTURING DATES, DEVICE DIMENSIONS, AND MATERIAL WAS PERFORMED. NO OBVIOUS CORRELATIONS OR ABNORMALITIES WERE FOUND. A VISIT TO THE MANUFACTURING SITE REVEALED NOTHING SIGNIFICANT. TO HELP ESTABLISH A ROOT CAUSE, IN-HOUSE TESTING WAS PERFORMED. THE FIRST TEST WAS TO EVALUATE THE EFFECT OF THERMAL QUENCHING AFTER AUTOCLAVE STERILIZATION ON THE PROPENSITY FOR CRACKING AND PROPAGATION FOR LOGIC INSERT TRIALS. THE SAMPLE TRIALS USED WERE EXPOSED TO TWENTY THERMAL CYCLES WHICH INCLUDED A THREE-MINUTE STERILIZATION CYCLE, A TWENTY- MINUTE DRYING PERIOD, AND A SUBSEQUENT QUENCHING IN AN ICE BATH. FOLLOWING TESTING, CRACK PROPAGATION WAS NOT OBSERVED ON ANY OF THE INSERT TRIALS. REFER TO (B)(4) FOR FURTHER DETAILS. THE SAME SAMPLES WERE THEN USED TO EVALUATE THE EFFECT OF SIMULATED SURGEON USAGE AND REPROCESSING ON THE PROPENSITY OF CRACK INITIATION AND PROPAGATION FOR LOGIC INSERT TRIALS. DURING THIS STUDY, THE TRIALS WERE EXPOSED TO TWENTY THERMAL AND MECHANICAL CYCLES. EACH CYCLE CONSISTED OF AUTOCLAVE STERILIZATION AND COLD-WATER QUENCHING FOLLOWED BY INSERTION INTO A TIBIAL TRAY, LOADING TO 390 N, AND REMOVAL FROM THE TIBIAL TRAY UNDER APPROXIMATELY 100 N USING EITHER A LPI TIBIAL INSERT TRIAL HANDLE OR AN OSTEOTOME. FOLLOWING TESTING, CRACK PROPAGATION WAS NOT OBSERVED. REFER TO (B)(4) FOR FURTHER DETAILS. MOST LIKELY CAUSE: A ROOT CAUSE HAS NOT BEEN ESTABLISHED AT THE TIME OF THIS INVESTIGATION. THE PERFORMED TESTING HAS BEEN UNABLE TO RECREATE THE REPORTED CRACKING. AS A RESULT OF (B)(4), CAPA(B)(4) WAS INITIATED TO FURTHER EVALUATE THE ROOT CAUSE. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. ¿ CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES. NO PATIENT ADVERSE EVENT. THE TRIALS BROKE DURING SURGERY. NO PARTS/PIECES FELL INTO THE SURGICAL SITE. A ROOT CAUSE HAS NOT BEEN ESTABLISHED AT THE TIME OF THIS INVESTIGATION. THE PERFORMED TESTING HAS BEEN UNABLE TO RECREATE THE REPORTED CRACKING. AS A RESULT OF (B)(4), CAPA(B)(4) WAS INITIATED TO FURTHER EVALUATE THE ROOT CAUSE. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00317 | 2022-03-30 | 2019-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRUL FEM TRIAL EXTRACTOR | EXACTECH, INC. |  | 02-029-19-1200 |  | IT WAS REPORTED FROM THE US THAT THE TRIAL EXTRACTOR BROKE DURING A TOTAL KNEE ARTHROPLASTY SURGERY. THE FEMUR TRIAL EXTRACTOR BROKE AT THE JUNCTION OF THE TWO PRONGS THAT HOOK INTO THE TRIAL WHILE BACKSLAPPING THE TRIAL OFF. THERE WAS A DELAY, BUT THE PATIENT WAS NOT ADVERSELY AFFECTED. THE DEVICES WERE NOT RETURNED FOR ANALYSIS. ATTEMPTS WERE MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REPRESENTATIVE, BUT NO RESPONSES WERE RECEIVED. THERE IS NO OTHER INFORMATION ABOUT THE DEVICE OR EVENT FROM THE CONTACTS. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS NOT AVAILABLE FOR ANALYSIS. ENGINEERING ANALYSIS COMPLETED BY (B)(4) ON 9/17/2019. (B)(4) AND CAPA(B)(4). FINDINGS: IN THE CASE OF MULTI-PIECE CONSTRUCTION, IT WAS FOUND THAT THE WELD ON THE EXTRACTOR BODY WAS BREAKING. THIS ALLOWED THE COMPONENT PARTS TO DISASSEMBLE. THE ENGINEERING PRINT FOR THE FEMORAL TRIAL EXTRACTOR (02-029-219-1202 REV A) ALLOWED FOR ¿SINGLE OR MULTIPLE PIECE CONSTRUCTION¿ (NOTE 2). HOWEVER, THE PRINT DID NOT SPECIFY HOW TO WELD/ASSEMBLE THE DEVICE IF MULTIPLE-PIECE CONSTRUCTION WAS TO BE USED. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: BASED ON THE CAPA AND CRC, THE BROKEN DEVICE REPORTED IN (B)(4) WAS LIKELY THE RESULT OF AN INSUFFICIENT WELD ON 02-029-219-1202 REV A, WHICH LED TO WELD FRACTURE AND COMPONENT DISASSEMBLY. CORRECTIVE ACTION TAKEN: PARTS WILL NOT BE ORDERED UNLESS THEY ARE MANUFACTURED AS A SINGLE PIECE CONSTRUCTION. REFER TO CAPA(B)(4). IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. ¿ CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO MANUFACTURING ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. THE DEVICE CAN FRACTURE AND DISASSEMBLE. BASED ON (B)(4) AND CAPA(B)(4), THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF AN INSUFFICIENT WELD ON 02-029-219-1202 REV A, WHICH LED TO WELD FRACTURE AND COMPONENT DISASSEMBLY. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. | N | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00316 | 2022-03-30 | 2022-01-01 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | GLENOSPHERE, 40MM | EXACTECH, INC. | 320-31-40 | 320-31-40 |  | AS REPORTED, POST OP SHOULDER SURGERY, THIS (B)(6) PATIENT REPORTED INSTABILITY AND SUBLUXATION. THE PATIENT HAS HAD ROUGHLY 6 SUBLUXATION EPISODES. THE CASE REPORT FORM INDICATES THIS EVENT IS POSSIBLY RELATED TO DEVICES AND DEFINITELY NOT RELATED TO PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. REVISION HAS NOT OCCURRED, CONTINUING DELTOID STRENGTHENING EXERCISES. OUTCOME IS CONTINUING. | CONCOMITANT DEVICE(S): 300-01-10 - EQUINOXE HUMERAL STEM PRIMARY PRESS FIT 10MM; 320-40-00 - HUMERAL LINER, 40MM, +0; 320-10-00 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0; 320-35-01 - SMALL GLENOID PLATE. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00315 | 2022-03-30 | 2020-03-01 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | GLENOSPHERE, 36MM | EXACTECH, INC. | 320-31-36 | 320-31-36 |  | AS REPORTED, APPROXIMATELY 1 YEAR POST OP INITIAL SHOULDER SURGERY, THIS (B)(6) FEMALE PATIENT PRESENTED A TRAUMATIC ACROMIAL FRACTURE AND UNDERWENT AN ACROMIONECTOMY ON (B)(6) 2021. THE CASE REPORT FORM INDICATES THIS EVENT IS UNLIKELY RELATED TO DEVICES AND DEFINITELY NOT RELATED TO PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. OUTCOME WAS CONSIDERED RESOLVED ON (B)(6) 2022. | CONCOMITANT DEVICE(S): 300-30-08 - EQUINOXE PRESERVE STEM 8MM; 320-36-00 - 36MM HUMERAL LINER +0 UNCONSTRAINED; 320-10-00 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0; 320-35-01 - SMALL GLENOID PLATE. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00314 | 2022-03-30 | 2021-11-22 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | GLENOSPHERE, 36MM | EXACTECH, INC. | 320-31-36 | 320-31-36 |  | AS REPORTED, APPROXIMATELY 2 YEARS POST OF SHOULDER SURGERY, THIS (B)(6) FEMALE PATIENT REPORTED INSTABILITY AND 2 SUBLUXATIONS FROM REACHING BEHIND HER CAR SEAT TO LIFT HER SMALL DOG. THE CASE REPORT FORM INDICATES THIS EVENT IS POSSIBLY RELATED TO DEVICES AND DEFINITELY NOT RELATED TO PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. REVISION HAS NOT OCCURRED, CONTINUING DELTOID STRENGTHENING EXERCISES. OUTCOME IS CONTINUING. | PENDING EVALUATION . CONCOMITANT DEVICE(S): 300-01-11 - EQUINOXE, HUMERAL STEM PRIMARY, PRESS FIT 11MM, 320-36-03 - 36MM HUMERAL LINER +2.5MM UNCONSTRAINED, 320-10-00 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0,AND 320-35-03 - SMALL POSTERIOR AUGMENT GLENOID PLATE, LEFT. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00313 | 2022-03-30 | 2020-08-10 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 42MM GLENOSPHERE | EXACTECH, INC. | 320-01-42 | 320-01-42 |  | AS REPORTED, APPROXIMATELY 2 YEARS POST OP SHOULDER SURGERY, THIS (B)(6) MALE PATIENT REPORTED INSTABILITY AND 2 SUBLUXATIONS OVER THE PAST YEAR, 4 TOTAL SINCE THE SURGERY, NONE REQUIRING FORMAL REDUCTION. THE CASE REPORT FORM INDICATES THIS EVENT IS POSSIBLY RELATED TO DEVICES AND PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. REVISION HAS NOT OCCURRED, CONTINUING DELTOID STRENGTHENING IN PHYSICAL THERAPY. OUTCOME IS CONTINUING. | CONCOMITANT DEVICE(S): 300-30-06 - EQUINOXE PRESERVE STEM 6MM; 320-42-00 - EQUINOXE REVERSE 42MM HUMERAL LINER +0; 320-10-00 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0; 320-15-01 - EQ REV GLENOID PLATE. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00324 | 2022-03-30 | 2019-01-03 | I, n, j, u, r, y | Manufacturer report | EXACTECH | MDLR RV RMR 38MM PLT TIP | EXACTECH, INC. |  | 321-25-38 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THERE WAS AN ISSUE WITH THE REAMER. THE TIP OF THE REAMER BROKE. THE PATIENT¿S HEALTH WAS STABLE LEAVING THE OR. NO PIECES WERE LEFT IN THE PATIENT. THE ISSUE WAS RESOLVED BY USING THE CANNULATED OPTION, THERE IS NO INDICATION THE DEVICE WAS USED. THERE WAS NO DELAY AND NO ADVERSE EFFECT TO THE PATIENT. THE DEVICE WAS RECEIVED ON 2/25/2019 ON RG452296. THE REPRESENTATIVE WAS PRESENT AT THE TIME OF SURGERY. TWO ATTEMPTS WERE MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REPRESENTATIVE, BUT NO RESPONSES WERE RECEIVED. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS AVAILABLE FOR ANALYSIS. BASED ON THE INVESTIGATION CONDUCTED UNDER CAPA (B)(4), IT WAS DETERMINED THAT THE USER INSTRUCTION WAS NOT ADEQUATE TO INFORM THE SURGEON OF THE APPROPRIATE USE OF THE REAMER. THE FORESEEABLE MISUSE OF THE GLENOID REAMER THAT COULD RESULT IN FRACTURE OF THE PILOT TIP ARE: IF THE USER APPLIES A BENDING MOMENT TO THE PILOT DIP DURING REAMING, WHICH WOULD REAM THE GLENOID ASYMMETRICALLY AND APPLY A TORQUE TO THE PILOT TIP, POTENTIALLY LEADING TO FRACTURE. IF THE USER APPLIES A BENDING MOMENT TO THE PILOT TIP BY USING THE GLENOID REAMER TO RETRACT THE HUMERAL HEAD TO HELP GAIN EXPOSURE AND ACCESS TO THE GLENOID, WHICH WOULD APPLY A TORQUE TO THE PILOT TIP, POTENTIALLY LEADING TO FRACTURE. CORRECTIVE ACTION TAKEN AS PART OF CAPA (B)(4), THE FOLLOWING CORRECTIVE ACTIONS WERE TAKEN: THE SURGICAL TECHNIQUE WAS UPDATED TO ADVISE SURGEONS TO AVOID APPLYING A BENDING TORQUE TO THE PILOT TIP REAMER OR USING THE REAMER TO RETRACT THE HUMERAL HEAD AS THIS MAY CAUSE FRACTURE OF THE PILOT TIP. ((B)(4)). UPDATE THE RMR TO INCLUDE THE RISK OF PILOT TIP REAMER FRACTURE AS A RESULT OF BENDING OR ITS FORESEEABLE MISUSE AS A RETRACTOR. ((B)(4)). PER (B)(4), A FIELD ADVISORY NOTICE WAS SENT TO THE USERS NOTIFYING THEM OF THE ISSUE AND ALERTING THEM TO THE CHANGE IN THE OPERATIVE TECHNIQUE (EXACTECH RECALL NUMBER 1038671-04/07/2017-003-C, FDA RECALL NUMBERS Z-2663-2017 THROUGH Z-2668-2017, MAY 2017). THE RECALL WAS TERMINATED BY THE FDA ON JULY 26, 2019. CAPA (B)(4) WAS INITIATED TO TRACK REPORTED EVENTS THROUGH 2021 AND RECONVENE THE CRC IF A THRESHOLD OF >35 TOTAL EVENTS OR >3 EVENTS WITH REPORTED PATIENT HARM BE MET. AS A RESULT OF MONITORING UNDER (B)(4) (APRIL 2019) AND CAPA (B)(4), IT WAS NOTED THAT THE THRESHOLD OF EVENTS FOR 2020 HAS BEEN EXCEEDED. THIS ISSUE HAS BEEN RE-ESCALATED FOR EVALUATION. AS DETERMINED BY (B)(4), THE RISK FALLS WITHIN THE ACCEPTABLE LEVEL. THE COMMITTEE HAS DETERMINED THAT NO ADDITIONAL FIELD ACTION IS REQUIRED AT THIS TIME. EXACTECH WILL CONTINUE MONITOR THE EVENTS THROUGH CAPA (B)(4), WHICH WILL DRIVE ANY FURTHER HHE UPDATES AND/OR CRC ESCALATION. IFU 700-096-181: INSTRUMENT INSPECTION. VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT FROM THE BROKEN TIP. BASED ON THE INVESTIGATION CONDUCTED UNDER CAPA (B)(4), IT WAS DETERMINED THAT THE USER INSTRUCTION WAS NOT ADEQUATE TO INFORM THE SURGEON OF THE APPROPRIATE USE OF THE REAMER. THE FORESEEABLE MISUSE OF THE GLENOID REAMER THAT COULD RESULT IN FRACTURE OF THE PILOT TIP ARE: IF THE USER APPLIES A BENDING MOMENT TO THE PILOT DIP DURING REAMING, WHICH WOULD REAM THE GLENOID ASYMMETRICALLY AND APPLY A TORQUE TO THE PILOT TIP, POTENTIALLY LEADING TO FRACTURE. IF THE USER APPLIES A BENDING MOMENT TO THE PILOT TIP BY USING THE GLENOID REAMER TO RETRACT THE HUMERAL HEAD TO HELP GAIN EXPOSURE AND ACCESS TO THE GLENOID, WHICH WOULD APPLY A TORQUE TO THE PILOT TIP, POTENTIALLY LEADING TO FRACTURE. | Y | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00339 | 2022-03-31 | 2022-03-11 | I, n, j, u, r, y | Manufacturer report | TRULIANT TIB IMP PS INSERT SZ 4.5 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT TIB IMP PS INSERT SZ 4.5 9MM | 02-022-35-4509 |  | AS REPORTED BY THE LEGAL TEAM, THIS MALE PATIENT COMPLAINTS OF PAIN AND ACHES IN RIGHT LEG, LIMPING. HE SAID HE HAS BEEN TELLING HIS SURGEON THAT FOR ABOUT A YEAR. PATIENT HAD X-RAYS DONE ABOUT A MONTH AGO AND SURGEON SAID THEY WERE FINE. PATIENT IS NOW WITH RECALL CONCERNED THAT SURGEON MAY NOT HAVE BEEN SEARCHING FOR THE RIGHT THING AND HE IS GOING TO SCHEDULE AN APPOINTMENT WITH HIS SURGEON FOR FOLLOW-UP. PATIENT ALSO SPOKE TO AN ATTORNEY. NO ADDITIONAL INFORMATION AVAILABLE. | CONCOMITANT MEDICAL PRODUCTS: TRULIANT PS CEM FEM PS CEM RIGHT SZ 4.5 (CAT# 02-020-11-0345 / SERIAL# (B)(4)). TRULIANT TIB FIT TRAY CEM SZ 4.5F/3.5T (CAT# 02-022-45-4535 / SERIAL# (B)(4)). ADVANCED PATELLA 35MM 3 PEG IMPLANT (CAT# 200-07-35 / SERIAL# (B)(4)). COLLAR FIXATION PIN 2PK 40MM, QUICK RELEASE (CAT# 201-78-82 / SERIAL# (B)(4)). 3" DRILL BIT, MOD. HEX 2 PACK (CAT# 201-78-89 / SERIAL# (B)(4)). 3" DRILL BIT, MOD. HEX 2 PACK (CAT# 201-78-89 / SERIAL# (B)(4)). THREADED PIN SIZE 2.6 COLLARLESS 2PK (CAT# 521-78-31 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00338 | 2022-03-31 | 2022-03-11 | I, n, j, u, r, y | Manufacturer report | NV GXL LINR, NTRL, 32MM ID, GROUP 2 CUPS | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | NV GXL LINR, NTRL, 32MM ID, GROUP 2 CUPS | 130-32-52 |  | AS REPORTED, THIS (B)(6) FEMALE PATIENT HAD AN EXACTECH THA DONE APPROXIMATELY 9 YEARS AGO. BY EXAMINATION OF THE X-RAYS, IT LOOKS TO BE AN EXACTECH NOVATION CROWN CUP, GXL LINER, ELEMENT STEM, UNKNOWN HEAD. THE PATIENT CAME TO SURGEON PRESENTING WITH MEDIALIZATION OF THE FEMORAL HEAD AND COMPONENTS INTO THE LINER AND CUP. IT IS ANTICIPATED THAT THERE IS ABNORMAL POLYETHYLENE WEAR AND POSSIBLE OSTEOLYSIS WITHIN AND AROUND THE COMPONENTS DUE TO THIS. THE SURGEON BOOKED THE PATIENT FOR A THA REVISION OF THE ACETABULAR CUP, POLY, AND HEAD. THE SURGEON REMOVED THE EXACTECG HEAD, LINER, AND CUP, USED DBM BONE GRAFT, AND THEN REPLACED THE ACETABULAR COMPONENT WITH A G7 MULTIHOLE ZB CUP WITH 4, 6.5MM SCREWS, LONGEVITY HIGHLY CROSS-LINKED POLY LINER - NEUTRAL 40MM, AND USED AN EXACTECH BIOLOX OPTION ADAPTER +3.5MM AND A 40MM BIOLOX OPTION FEMORAL HEAD. THERE WAS NO BREAKAGE TO DEVICE. THERE WAS MORE THAN 45 MIN DELAY/PROLONGATION. THE PATIENT WAS STABLE FOLLOWING THE REVISION PROCEDURE. AN X-RAY IS PROVIDED. A REPRESENTATIVE WAS NOT REQUIRED TO BE AT THE SURGERY. THE IMPLANTS REMOVED FROM THE PATIENT WERE DISCARDED BY THE FACILITY. CONCOMITANT MEDICAL PRODUCTS: NOVATION CROWN CUP. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00337 | 2022-03-31 | 2022-03-11 | I, n, j, u, r, y | Manufacturer report | FEMORAL COMPONENTS | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | FEMORAL COMPONENTS | UNKNOWN |  | AS REPORTED, APPROXIMATELY SEVEN YEARS POST INITIAL TKA, PATIENT VISITED SURGEON FOR SUSPECTED OSTEOLYSIS. THE FEMORAL BONE ON THE MEDIAL SIDE APPEARS TO SHOW OSTEOLYSIS. SURGEONS WORKING ON SCHEDULING THE PATIENT SOON FOR A REVISION. | THE KNEE WAS DONE IN 2013. ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00336 | 2022-03-31 | 2022-02-28 | I, n, j, u, r, y | Manufacturer report | LOGIC FEMORAL PS CEM RIGHT SZ 2.5 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC FEMORAL PS CEM RIGHT SZ 2.5 | 02-010-01-0325 |  | AS REPORTED, APPROXIMATELY SIX YEARS POST INITIAL RIGHT TKA, THIS FEMALE PATIENT POLY WAS BEING REVISED DUE TO LOOSE FEMUR. PATIENT HAD A LOOSE FEMUR AND SLIGHTLY WORN POLY. PATIENT WAS REVISED TO (B)(6). THERE WAS NO BREAKAGE OF A DEVICE AND NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION. SALES REP WAS UNABLE TO OBTAIN PHOTOS/X-RAYS. THE DEVICE WILL BE RETURN. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-03
1038671-2022-00320 | 2022-03-31 | 2019-07-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | MDLR RV RMR 38MM PLT TIP | EXACTECH, INC. |  | 321-25-38 |  | IT WAS REPORTED FROM INTERNAL EXACTECH THAT PILOT REAMER WAS FOUND WITH A BROKEN TIP WAS RETURNED WITH NO INFORMATION. THERE WAS NO REPORT FROM AN AGENCY OR FACILITY. NO ADDITIONAL INFORMATION IS AVAILABLE. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES. NO PATIENT ADVERSE EVENT HAS BEEN REPORTED FROM AN AGENCY OR FACILITY. AN INVESTIGATION WAS CONDUCTED: BASED ON CAPA (B)(4), THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE REAMER DURING USE, WHICH LED TO BRITTLE FRACTURE OF THE PILOT TIP FEATURE. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE(S) AVAILABLE FOR EVALUATION. EVALUATION COMPLETED ON 10/01/2019 BY NF. FINDINGS: BASED ON THE INVESTIGATION CONDUCTED UNDER CAPA (B)(4), IT WAS DETERMINED THAT THE USER INSTRUCTION (E.G. THE SURGICAL TECHNIQUE) WAS NOT ADEQUATE TO INFORM THE SURGEON OF THE APPROPRIATE USE OF THE REAMER. THE FORESEEABLE MISUSE OF THE GLENOID REAMER THAT COULD RESULT IN FRACTURE OF THE PILOT TIP ARE: 1) IF THE USER APPLIES A BENDING MOMENT TO THE PILOT TIP DURING REAMING, WHICH WOULD REAM THE GLENOID ASYMMETRICALLY AND APPLY A TORQUE TO THE PILOT TIP, POTENTIALLY LEADING TO FRACTURE. 2) IF THE USER APPLIES A BENDING MOMENT TO THE PILOT TIP BY USING THE GLENOID REAMER TO RETRACT THE HUMERAL HEAD TO HELP GAIN EXPOSURE AND ACCESS TO THE GLENOID, WHICH WOULD APPLY A TORQUE TO THE PILOT TIP, POTENTIALLY LEADING TO FRACTURE. MOST LIKELY CAUSE: BASED ON CAPA (B)(4), THE BROKEN DEVICE REPORTED IN CASE-(B)(4) WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE REAMER DURING USE, WHICH LED TO BRITTLE FRACTURE OF THE PILOT TIP FEATURE. CORRECTIVE ACTION TAKEN: AS PART OF CAPA (B)(4), THE FOLLOWING CORRECTIVE ACTIONS WERE TAKEN: 1. THE SURGICAL TECHNIQUE WAS UPDATED TO ADVISE SURGEON TO AVOID APPLYING A BENDING TORQUE TO THE PILOT TIP REAMER OR USING THE REAMER TO RETRACT THE HUMERAL HEAD AS THIS MAY CAUSE FRACTURE OF THE PILOT TIP (DCR-SP-17-135, -136, -137, -138). 2. UPDATE THE RMR TO INCLUDE THE RISK OF PILOT TIP REAMER FRACTURE AS A RESULT OF BENDING OR ITS FORESEEABLE MISUSE AS A RETRACTOR (DCR-SP-172). ADDITIONALLY, PER CRC2017-03-27-04, A RECALL COMMUNICATION WAS SENT TO USERS NOTIFYING THEM OF THE ISSUE AND ALERTING THEM TO THE CHANGE IN THE OPERATIVE TECHNIQUE (Z-2663 THROUGH 2668-2017, MAY 2017). IFU 700-096-181: INSTRUMENT INSPECTION. VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES. NO PATIENT ADVERSE EVENT HAS BEEN REPORTED FROM AN AGENCY OR FACILITY. AN INVESTIGATION WAS CONDUCTED: BASED ON CAPA (B)(4), THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE REAMER DURING USE, WHICH LED TO BRITTLE FRACTURE OF THE PILOT TIP FEATURE. | N | Y | OTHER |  | 2022-03
1038671-2022-00341 | 2022-04-01 | 2018-09-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRULIANT TIB TAMP GUIDE | EXACTECH, INC. |  | 02-029-22-2000 |  | IT WAS REPORTED FROM THE US THAT AFTER AN ORTHOPEDIC SURGERY THE TIBIAL TAMP GUIDE WAS FOUND IN TWO PIECES. IT WAS STATED THAT THE HANDLE APPEARED TO WORK PROPERLY DURING THE CASE. WHEN THE AGENT WENT TO PUT THE TRAYS BACK TOGETHER, THE HANDLE WAS IN TWO PIECES. NO PATIENT INVOLVEMENT AT THE TIME OF DISCOVERY. NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH (B)(4) HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PATIENT. ENGINEERING PRINTS OF 02-029-122-2000 WERE REVIEWED. THE OVERLAP BETWEEN THE SNAP RING¿S OD AND THE ID OF THE HOUSING LIP IS .008 FOR THE WORST-CASE SCENARIO. ALTHOUGH THE OVERLAP IS POSITIVE, THE SMALL AMOUNT COULD CAUSE DISENGAGEMENT ROOT CAUSE PER CAPA (B)(4), THE ROOT CAUSE FOR THE BUTTON DISENGAGEMENT WAS LIKELY THE RESULT OF INSUFFICIENT OVERLAP BETWEEN THE SNAP RING AND THE LIP OF THE HOUSING TO PREVENT THE OCCURRENCE OF DISENGAGEMENT. CORRECTIVE ACTION TAKEN: "PER (B)(4), THE COMMITTEE HAS DETERMINED THAT THE DESIGN WILL BE UPDATED, TESTED, AND IMPLEMENTED, AFTER WHICH A STAGED MARKET WITHDRAWAL WAS CONDUCTED. PER CAPA (B)(4): A STOCK AUDIT WAS PERFORMED FOR THE DEVICES LOCATED IN 01 (COC). SEE NPR-40446 AND NPR-45222 FOR DETAILS. DEFECTIVE DEVICES WERE DISPOSITIONED FOR REWORK. THE DESIGN WAS UPDATED TO INCREASE THE OVERLAP OF THE SNAP RING AND THE HOUSING LIP COMPONENTS. SEE DCRTC-18-0401 AND DCRTC-19-0446 FOR DETAILS. THE DEFECTED DEVICES RETURNED FROM THE FIELD ARE TO BE REWORKED. SEE STAGED MARKET WITHDRAWAL OF TRULIANT TIBIAL TAMP GUIDE DOCUMENTED IN THE CAPA FOR DETAILS." IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. NO PATIENT ADVERSE EVENT. PER CAPA (B)(4), THE ROOT CAUSE FOR THE BUTTON DISENGAGEMENT WAS LIKELY THE RESULT OF INSUFFICIENT OVERLAP BETWEEN THE SNAP RING AND THE LIP OF THE HOUSING TO PREVENT THE OCCURRENCE OF DISENGAGEMENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00340 | 2022-04-01 | 2022-03-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | VANTAGE | LNR TRIAL-MB-SZ 3-L-6MM | EXACTECH, INC. |  | 351-41-03 |  | IT WAS REPORTED THAT DURING SURGICAL USE, AFTER IMPACTING THE TIBIAL AND TALAR IMPLANTS, THE SAMPLE INSERT WAS INSERTED TO CHECK THE RANGE OF MOTION AND CORRECT HEIGHT SELECTION OF THE FINAL POLY INLAY. DURING REMOVAL WITH THE INSERTER HANDLE SCISSOR/CLAMP (351-01-10) IT WAS RECOGNIZED THAT IT WAS VERY HARD TO REMOVE THE INSERT FROM THE JOINT AND WHILE THE SURGEON CONTINUED TRYING, THE ANTERIOR HOLES FOR THE INSERTER HANDLE BROKE OUT. EVENT NOT ASSOCIATED WITH REVISION OF EXACTECH IMPLANTS. NO PIECES OF THE DEVICE FELL INTO THE PATIENT WOULD SITE. NO SURGICAL DELAY/PROLONGATION. FEMALE PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. NO OTHER PATIENT INFORMATION OR MEDICAL HISTORY REPORTED. DEVICE WILL BE RETURNING. | PENDING EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00342 | 2022-04-01 | 2022-03-08 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REV LOCKING SCREW | EXACTECH, INC. | 320-15-05 | 320-15-05 |  | AS REPORTED, APPROXIMATELY 5 MONTHS POST OP THE INITIAL RTSA, THIS FEMALE PATIENT WAS REVISED. THE GLENOSPHERE LOCKING SCREW WAS BROKEN CAUSING THE GLENOSPHERE TO DISASSOCIATE FROM THE GLENOID BASEPLATE. SURGEON REPLACED THE GLENOSPHERE, GLENOID, HUMERAL TRAY AND HUMERAL LINER. EVENT STARTED BY DOING NORMAL, EVERYDAY ACTIVITIES. FEMALE PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. NO OTHER PATIENT INFORMATION/MEDICAL HISTORY REPORTED. | PENDING EVALUATION. CONCOMITANT DEVICE(S): 320-35-02, 6960981 - SMALL SUPERIOR AUGMENT GLENOID PLATE; 320-31-36, 6917056 - GLENOSPHERE, 36MM. NO INFORMATION PROVIDED IN THE FOLLOWING SECTION(S). | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00343 | 2022-04-01 | 2018-10-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | MDLR CNLTD TRIDRV DRVR | EXACTECH, INC. |  | 315-25-00 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SPRING BROKE WHILE DRILLING PERIPHERAL HOLES. NO DEVICE PARTS OR PIECES FELL INTO THE SURGICAL SITE. THE PATIENT¿S HEALTH WAS STABLE LEAVING THE OR. THE ONLY EFFECT TO THE CASE WAS A DELAY IN TIME OF COMPLETION DUE TO GOING TO GRAB ANOTHER REPLICATOR PLATE. NO PROBLEMS WERE NOTED FOR THE PATIENT RELATED TO PROCEDURE OUTCOME/RESULT. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. CRC2019-04-05 ¿TRI-DRIVE MODULAR REAMER DRIVER DISASSEMBLY THE DESIGN OF THE TRI-DRIVE MODULAR REAMER DRIVER HAS BEEN IN THE FIELD SINCE 2009. EXACTECH IS AWARE OF 42 COMPLAINT REPORTS INVOLVING TIP DISASSEMBLY OF 44 TRI-DRIVES. BECAUSE THE TRI-DRIVE IS USED IN BOTH ANATOMIC AND REVERSE TOTAL SHOULDER ARTHROPLASTY, SALES DATA FOR ANATOMIC GLENOID COMPONENTS AND GLENOSPHERE COMPONENTS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE IN THE RMR. THE DISASSEMBLED TRI-DRIVE MODULAR REAMER DRIVER MAY RESULT IN A DELAY IN SURGERY TO FIND A NEW ONE. TYPICALLY, IN ANATOMIC SHOULDER CASES BOTH ANATOMIC AND REVERSE SETS ARE IN THE OR. IF THE TRI-DRIVE WERE TO BREAK, A SECOND TRI-DRIVE WOULD BE IN THE REVERSE SHOULDER SET. THE CRC DETERMINED THAT NO FURTHER ACTION WAS REQUIRED AT THIS TIME AS THIS ISSUE WILL BE ADDRESSED VIA THE EQUINOXE ERGO INSTRUMENTATION REDESIGN PROJECT. IN THE NEW EQUINOXE ERGO INSTRUMENTATION, THE TRI-DRIVE HANDLES WILL NO LONGER BE USED. SINCE THE DESIGN OF THE PILOT TIP REAMERS ARE CHANGING (ALL THE REAMERS ARE CANNULATED), THERE WILL NOW BE TWO DIFFERENT HANDLES: A PILOT TIP HANDLE AND A CANNULATED HANDLE, REPLACING THE TRI-DRIVE. THIS NEW ERGO INSTRUMENTATION DESIGN WILL BE IMPLEMENTED AS PART OF A PHASED ROLL-OUT, FIRST IN THE US AND THEN TO OTHER MARKETS, BEGINNING IN MID-2019. THESE BROKEN TRI-DRIVES WERE LIKELY THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING OR THE TRI-DRIVE CONTACTING A RETRACTOR, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY, ALLOWING THE TIP TO DISASSEMBLE. IFU 700-096-181: INSTRUMENT INSPECTION. VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT IMPACT. PER CRC2019-04-05; THESE BROKEN TRI-DRIVES WERE LIKELY THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING OR THE TRI-DRIVE CONTACTING A RETRACTOR, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY, ALLOWING THE TIP TO DISASSEMBLE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00346 | 2022-04-01 | 2022-01-27 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 42MM HUMERAL LINER +2.5 | EXACTECH, INC. | 320-42-03 | 320-42-03 |  | AS REPORTED, APPROXIMATELY 2.5 YEARS POST OP THE INITIAL RIGHT TSA, THIS (B)(6) MALE PATIENT'S WAS REVISED DUE THE HUMERAL LINER DISLODGING FROM THE HUMERAL ADAPTER TRAY. HE PATIENT REPORTEDLY REACHED BACKWARD TO PICK UP AN ITEM AROUND THE DATE LISTED ABOVE. THEN, THE PATIENT FELT PAIN. SURGEON CONDUCTED A REVISION SURGERY AND FOUND THE LINER TO BE DISLODGED FROM THE TRAY. THE HUMERAL TRAY, TORQUE SCREW, AND HUMERAL LINER WERE REMOVED AND REPLACED WITH NEW IMPLANTS. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. REVISION SURGERY SUCCESSFUL. DEVICES WERE KEPT BY THE HOSPITAL AS PART OF THE STANDARD PROTOCOL. | PENDING EVALUATION. CONCOMITANT DEVICE(S): 320-10-10 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +10 AND 320-20-00 - EQ REVERSE TORQUE DEFINING SCREW KIT. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00344 | 2022-04-01 | 2022-03-10 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | CAGE GLENOID MEDIUM, ALPHA | EXACTECH, INC. | 314-13-03 | 314-13-03 |  | AS REPORTED, APPROXIMATELY 3 YEARS POST OP THE INITIAL LEFT TSA, THIS (B)(6) YEAR OLD FEMALE PATIENT'S WAS REVISED DUE TO HER SHOULDER BECOMING UNSTABLE. SURGEON REVISED ANATOMIC SHOULDER IMPLANTS TO REVERSE. PATIENT WAS STABLE AFTER IMPLANTS. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. | PENDING EVALUATION. CONCOMITANT MEDICAL PRODUCTS: 300-10-45, (B)(4) - EQUINOXE REPLICATOR PLATE 4.5MM O/S; 300-20-02, (B)(4) - EQUINOX SQUARE TORQUE DEFINE SCREW DRIVE KIT; 310-01-44, (B)(4) - EQ HUMERAL HEAD SHORT 44M. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00345 | 2022-04-01 | 2022-03-11 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | HUMERAL STEM PRIMARY PRESS FIT 12MM | EXACTECH, INC. | 300-01-12 | 300-01-12 |  | AS REPORTED, APPROXIMATELY 3.5 YEARS POST OP INITIAL RIGHT TSA, THIS (B)(6) Y/O MALE PATIENT WAS REVISED DUE TO INCREASING PERSISTENT PAIN IN OPERATIVE SHOULDER. SURGEON EXPRESSED CONCERN THAT THE HUMERAL STEM WAS LOOSE. INTRAOPERATIVELY, THE STEM WAS FOUND TO HAVE SOME SLIGHT MOVEMENT. THERE WAS NO SIGNIFICANT SIGN OF INFECTION BUT CULTURES WERE TAKEN. ALL HUMERAL COMPONENTS WERE REMOVED & REVISED. THE HUMERAL STEM WAS CEMENTED THIS TIME. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WERE KEPT BY HOSPITAL. | CONCOMITANT DEVICE(S): 320-10-00 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0, 320-38-10 - EQUINOXE REVERSE 38MM HUMERAL CONST LINER +0, 320-20-00 - EQ REVERSE TORQUE DEFINING SCREW KIT. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00363 | 2022-04-04 | 2020-01-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRULIANT TIB TAMP GUIDE | EXACTECH, INC. |  | 02-029-22-2000 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THERE WAS A DISASSEMBLY OF AN INSTRUMENT. THE BLUE BUTTON ON THE DEVICE WAS MISSING DURING SURGERY. THE INSTRUMENT WAS USED BY HOLDING THE TIBIAL BROACH BY HAND DURING IMPACTION. THE SURGEON WAS ABLE TO USE THE INSTRUMENT IN THE NONFUNCTIONAL STATE. THE SURGERY NOR PATIENT COMPROMISED BY THE EVENT. THE BROACH WAS REMOVED WITH PLIERS AFTER IMPACTION. NO PARTS LEFT IN WOUND SITE. THERE WAS NO REPORTED PATIENT ADVERSE EVENT. NO INFORMATION ON THE OVERALL PATIENT OUTCOME OR PATIENT DEMOGRAPHICS. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4) HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. FINDINGS: ACCORDING TO THE EXPERIENCE REPORT, ¿BLUE BUTTON ON DEVICE WAS MISSING DURING SURGERY. INSTRUMENT WAS USED BY HOLDING TIBIAL BROACH BY HAND DURING IMPACTION. BROACH WAS REMOVED WITH PLIERS AFTER IMPACTION. BROKEN INSTRUMENT WILL BE RETRIEVED AND RETURNED. NO PARTS LEFT IN WOUND SITE.¿ ENGINEERING PRINTS OF 02-029-122-2000 WERE REVIEWED. THE OVERLAP BETWEEN THE SNAP RING¿S OD AND THE ID OF THE HOUSING LIP IS .008 FOR THE WORST-CASE SCENARIO. ALTHOUGH THE OVERLAP IS POSITIVE, THE SMALL AMOUNT COULD CAUSE DISENGAGEMENT. CORRECTIVE ACTION TAKEN: PER CRC2018-08-31-02, THE COMMITTEE DETERMINED THAT THE DESIGN WOULD BE UPDATED, TESTED, AND IMPLEMENTED, AFTER WHICH A STAGED MARKET WITHDRAWAL WAS CONDUCTED. PER CAPA2018-35: 1. A STOCK AUDIT WAS PERFORMED FOR THE DEVICES LOCATED IN 01 (COC). SEE NPR-40446 AND NPR-45222 FOR DETAILS. DEFECTIVE DEVICES WERE DISPOSITIONED FOR REWORK. 2. THE DESIGN WAS UPDATED TO INCREASE THE OVERLAP OF THE SNAP RING AND THE HOUSING LIP COMPONENTS. SEE DCRTC-18-0401 AND DCRTC-19-0446 FOR DETAILS. 3. THE DEFECTED DEVICES RETURNED FROM THE FIELD ARE TO BE REWORKED. SEE STAGED MARKET WITHDRAWAL OF TRULIANT TIBIAL TAMP GUIDE DOCUMENTED IN THE CAPA FOR DETAILS. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: PER CAPA2018-35, THE ROOT CAUSE FOR THE BUTTON DISENGAGEMENT REPORTED WAS LIKELY THE RESULT OF INSUFFICIENT OVERLAP BETWEEN THE SNAP RING AND THE LIP OF THE HOUSING TO PREVENT THE OCCURRENCE OF DISENGAGEMENT. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. ¿ CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. THERE WAS NO REPORTED PATIENT ADVERSE EVENT DUE TO THE MISSING BUTTON ON THE INSTRUMENT, NOTHING WAS LEFT IN THE PATIENT. PER CAPA2018-35, THE ROOT CAUSE FOR THE BUTTON DISENGAGEMENT REPORTED WAS LIKELY THE RESULT OF INSUFFICIENT OVERLAP BETWEEN THE SNAP RING AND THE LIP OF THE HOUSING TO PREVENT THE OCCURRENCE OF DISENGAGEMENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00364 | 2022-04-04 | 2020-01-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | MDLR CNLTD TRIDRV DRVR | EXACTECH, INC. |  | 315-25-00 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE TRIDRIVE DRIVER BROKE. WHILE REAMING THE GLENOID, THE COLLAR AND SPRING BROKE OFF THE HANDLE. THE AGENT THEN OPENED ANOTHER TRAY AND WHILE DRILLING THE CENTER HOLE, THE SAME THING HAPPENED ON THE SECOND HANDLE. ALL PIECES WERE RETRIEVED AND THERE WAS NO REAL DELAY TO SURGERY OTHER THAN WHILE OPENING ANOTHER TRAY. THERE WAS NO REPORTED PATIENT ADVERSE EVENT. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. 1 OF 2 FDA MDRS FOR THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. FINDINGS: AS PART OF THE CAPA INVESTIGATION, PREVIOUS COMPLAINT EVALUATIONS REGARDING TRI-DRIVES WERE REVIEWED. DURING ONE OF THESE PREVIOUS EVALUATIONS, FOUR JAMMED TRI-DRIVES WERE DISASSEMBLED AND CONFIRMED THE PRESENCE OF METAL WEAR PARTICLES INSIDE THE SLEEVE, MOST NOTABLY IN THE SLOT WITHIN WHICH THE RETAINING RING SEATS. IT WAS ALSO FOUND THAT THE RETAINING RING IS OFTEN DEFORMED. ADDITIONALLY, AN ATTEMPT WAS MADE TO RECREATE THE FAILURE MODE IN A LAB SETTING. IT WAS CONFIRMED THAT THE SLEEVE, SPRING, AND BALL BEARING CAN DISASSEMBLE FROM THE SHAFT AS A RESULT OF HOLDING THE SLEEVE STATIONARY DURING REAMING. THE PRIMARY CAUSE OF THE DISASSEMBLY WAS CONCLUDED TO BE DEFORMATION OF THE RETAINING RING, LEADING TO SUBSEQUENT DISENGAGEMENT FROM THE SLEEVE. FURTHERMORE, THE PRESENCE OF WEAR DEBRIS WAS OBSERVED, AND IS CONSISTENT WITH PREVIOUS COMPLAINT EVALUATIONS. CORRECTIVE ACTION TAKEN: THE CRC DETERMINED THAT NO FIELD ACTION WAS NECESSARY AT THAT TIME. THIS ISSUE WILL BE ADDRESSED VIA THE EQUINOXE ERGO INSTRUMENTATION REDESIGN PROJECT. IN THE NEW EQUINOXE ERGO INSTRUMENTATION, THE MODULAR CANNULATED TRI-DRIVE WILL NOT BE USED. THE ERGO INSTRUMENTATION WILL BE IMPLEMENTED AS PART OF A PHASED ROLL-OUT, FIRST IN THE UNITED STATES AND THEN TO OTHER MARKETS, BEGINNING IN MID-2019. FULL IMPLEMENTATION IS EXPECTED BY Q4 2020 (US) AND Q4 2021 (OUS). THE COMMITTEE WILL RECONVENE IF A THRESHOLD OF > 15 EVENTS OR = 1 EVENT WITH HARM IS REACHED. EVENT TRACKING WILL BE CAPTURED IN CAPA(B)(4). ACCORDING TO THE CAPA INVESTIGATION RESULTS, THE CAPA(B)(4) INVESTIGATION CAN BE CLOSED WITHOUT FURTHER CORRECTION OR CONTAINMENT OR CORRECTIVE OR PREVENTIVE ACTIONS. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: PER CAPA(B)(4), THE DISASSEMBLED TRI-DRIVE HANDLES REPORTED WERE LIKELY THE RESULT OF THE RETAINING RING RESPONSIBLE FOR MAINTAINING ASSEMBLY OF THE SLEEVE ON THE MODULAR CANNULATED TRI-DRIVE SHAFT BECOMING DEFORMED AS A RESULT OF HOLDING THE SLEEVE STATIONARY WHILE REAMING. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. ALL PIECES OF THE BROKEN INSTRUMENT WERE RETRIEVED AND THERE WAS NO REAL DELAY TO SURGERY. THERE WAS NO REPORTED PATIENT ADVERSE EVENT. PER CAPA(B)(4), THE DISASSEMBLED TRI-DRIVE HANDLES REPORTED WERE LIKELY THE RESULT OF THE RETAINING RING RESPONSIBLE FOR MAINTAINING ASSEMBLY OF THE SLEEVE ON THE MODULAR CANNULATED TRI-DRIVE SHAFT BECOMING DEFORMED AS A RESULT OF HOLDING THE SLEEVE STATIONARY WHILE REAMING. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00362 | 2022-04-04 | 2022-01-26 | I, n, j, u, r, y | Manufacturer report | TIBIAL INSERT SIZE 3 9++ | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TIBIAL INSERT SIZE 3 9++ | UNKNOWN |  | AS REPORTED, THIS FEMALE PATIENT HAD NARROWING OF THE JOINT SPACE AND SURGEON TOOK THE PATIENT BACK TO HAVE THE POLY EXCHANGED. PATIENT HAD A REVISION ON (B)(6) 2022. THERE WAS EVIDENCE OF POLY WEAR. THERE WERE NO PARTS OR PIECES FELL IN THE PATIENT WOUND SITE. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PATIENT IS DOING WELL. SALES REP WAS UNABLE TO OBTAIN PHOTOS/X-RAYS. THE DEVICE WILL BE RETURN. | (H3) THE REVISION REPORTED MAY HAVE BEEN THE RESULT OF POSTERIOR TIBIAL SLOPE, THE SLOPE ++ TIBIAL INSERT, AND PATIENT-RELATED CONDITIONS WHICH ALLOWED FOR EXCESSIVE POSTERIOR CONTACT BETWEEN THE FEMORAL COMPONENT AND THE TIBIAL INSERT AND LED TO THE REPORTED NARROWING OF THE JOINT SPACE AND WEAR OF THE POLYETHYLENE. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION AND PRE-REVISION X-RAYS WERE UNABLE TO BE OBTAINED. THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE REPORTED EVENT OF "PROSTHESIS WEAR¿ IS ASSOCIATED WITH MATERIAL DAMAGE TO A SURFACE, USUALLY INVOLVING PROGRESSIVE LOSS OR DISPLACEMENT OF MATERIAL, DUE TO RELATIVE MOTION BETWEEN THAT SURFACE AND A CONTACTING SUBSTANCE OR SUBSTANCES. SECTION D9: DEVICE AVAILABLE FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00366 | 2022-04-04 | 2020-02-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | EQUINOXE II BROACH HANDL | EXACTECH, INC. |  | 301-05-01 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SHOULDER SURGERY THE BROACH HANDLE BROKE, A SMALL PIECE OF METAL HIT THE OR FLOOR. THE SURGEON WAS BROACHING A SHOULDER USING THE PRESERVE BROACH. HE FINISHED BROACHING AND HANDED BACK THE BROACH TO THE TECH AND WHEN THE TECH PULLED THE TRIGGER TO RELEASE THE BROACH A SMALL PIECE OF METAL CAME OFF THE BROACH HANDLE AND LANDED ON THE FLOOR. IT WAS NOT BEING USED IN THE PATIENT AT THE TIME. AT THE COMPLETION OF THE CASE, THE AGENT LOOKED AT THE BROACH HANDLE AND FIGURED OUT THAT THE SMALL METAL PIECE CAME FROM THE BROACH HANDLE BETWEEN THE 0 AND 10 DEGREE HOLES ON THE LEFT SIDE OF THE BROACH. THE EVENT DID NOT LEAD TO INJURY OF THE PATIENT OR USER, AND NO PARTS OF A BROKEN DEVICE/DEVICE FRAGMENTS FELL IN THE PATIENT. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. RG507778 WAS RECEIVED ON 3/11/2020. HOWEVER; WAS LOST DURING TRANSFER TO COMPLAINTS. THIS COMPLAINT WILL BE INVESTIGATED AS NO PRODUCT RETURNED. DESIGN-RELATED ISSUES: THE DESIGN OF THE EQUINOXE BROACH HANDLE HAS BEEN IN THE FIELD SINCE 2017. EXACTECH IS AWARE OF 5 OTHER SIMILAR COMPLAINT REPORTS INVOLVING THIS DEVICE SINCE ITS RELEASE. FREQUENCY OF OCCURRENCE RANKING SCALE IS VERY LOW. HOWEVER; THE BRIDGE FEATURE WAS REMOVED FROM THE NEW ERGO BROACH HANDLE DESIGN, CN: (B)(4). CAPA (B)(4) HAS BEEN INITIATED TO ADDRESS THE DESIGN CHANGES RELATED TO THIS FAILURE MODE. MFG-RELATED ISSUES: EXACTECH IS AWARE OF 1 OTHER COMPLAINT INVOLVING PARTS FROM THIS MANUFACTURING LOT. THE DEVICE HISTORY RECORD (DHR) FOR THIS DEVICE WAS REVIEWED. THE DHR REVIEW SHOWED THAT A QUANTITY OF (B)(4) UNITS WAS PURCHASED IN 2019 UNDER PURCHASE ORDER (B)(4). NPR 42018 WAS ISSUED FOR THIS LOT. THE ENTIRE LOT WAS INITIALLY REJECTED ON 2/21/2019 DUE TO NOT BEING ABLE TO USE A PIN GAGE FOR INSPECTION AS STATED IN THE IRRS (QC301-05-01 REV C) LINE 11. THE IRRS WAS REVISED ON DCRTC-19-0185 TO REMOVE THE REQUIREMENT FOR INSPECTING THIS DIMENSION AS IT CANNOT BE INSPECTED ON THESE PARTS. THE DEVICES WERE DISPOSITIONED ¿USE AS IS.¿ HOWEVER, THREE OF THE 53 DEVICES WERE REJECTED AGAIN FOR UNRELATED MATING ISSUES AND SENT TO THE SUPPLIER FOR REWORK OR REPLACEMENT. THE SUPPLIER SHIPPED BACK THREE BROACH HANDLES ON 3/28/2019, WHICH WERE RE-INSPECTED ACCEPTED ON 4/17/2019. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿ AND THE POTENTIAL RISKS ARE CAPTURED IN THE RMR AND RACR. HOWEVER, CAPA (B)(4) HAS BEEN INITIATED TO ADDRESS THE DESIGN CHANGES RELATED TO THIS FAILURE MODE THE MOST LIKELY UNDERLYING CAUSE FOR THIS BROKEN BROACH HANDLE REPORTED COULD NOT BE DETERMINED AT THE TIME OF THIS INVESTIGATION. CONTRIBUTIONS OF INSTRUMENT REPROCESSING AND SURGICAL USE TO THE SEQUENCE OF CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE METAL BRIDGE COULD NOT BE DETERMINED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION; CORROSION; OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES; MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. A BROKEN PIECE OF THE INSTRUMENT LANDED ON THE OR FLOOR AFTER THE SURGEON HAD USED IT AND THE SCRUB TECH WAS TRYING TO RELEASE THE BROACH. THERE WAS NO REPORTED PATIENT ADVERSE EVENT. AN INVESTIGATION WAS CONDUCTED; THE MOST LIKELY UNDERLYING CAUSE FOR THIS BROKEN BROACH HANDLE REPORTED COULD NOT BE DETERMINED AT THE TIME OF THIS INVESTIGATION. CONTRIBUTIONS OF INSTRUMENT REPROCESSING AND SURGICAL USE TO THE SEQUENCE OF CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE METAL BRIDGE COULD NOT BE DETERMINED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00365 | 2022-04-04 | 2020-01-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | MDLR CNLTD TRIDRV DRVR | EXACTECH, INC. |  | 315-25-00 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE TRIDRIVE DRIVER BROKE. WHILE REAMING THE GLENOID, THE COLLAR AND SPRING BROKE OFF THE HANDLE. THE AGENT THEN OPENED ANOTHER TRAY AND WHILE DRILLING THE CENTER HOLE, THE SAME THING HAPPENED ON THE SECOND HANDLE. ALL PIECES WERE RETRIEVED AND THERE WAS NO REAL DELAY TO SURGERY OTHER THAN WHILE OPENING ANOTHER TRAY. THERE WAS NO REPORTED PATIENT ADVERSE EVENT. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. 2 OF 2 FDA MDRS FOR THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. FINDINGS: AS PART OF THE CAPA INVESTIGATION, PREVIOUS COMPLAINT EVALUATIONS REGARDING TRI-DRIVES WERE REVIEWED. DURING ONE OF THESE PREVIOUS EVALUATIONS, FOUR JAMMED TRI-DRIVES WERE DISASSEMBLED AND CONFIRMED THE PRESENCE OF METAL WEAR PARTICLES INSIDE THE SLEEVE, MOST NOTABLY IN THE SLOT WITHIN WHICH THE RETAINING RING SEATS. IT WAS ALSO FOUND THAT THE RETAINING RING IS OFTEN DEFORMED. ADDITIONALLY, AN ATTEMPT WAS MADE TO RECREATE THE FAILURE MODE IN A LAB SETTING. IT WAS CONFIRMED THAT THE SLEEVE, SPRING, AND BALL BEARING CAN DISASSEMBLE FROM THE SHAFT AS A RESULT OF HOLDING THE SLEEVE STATIONARY DURING REAMING. THE PRIMARY CAUSE OF THE DISASSEMBLY WAS CONCLUDED TO BE DEFORMATION OF THE RETAINING RING, LEADING TO SUBSEQUENT DISENGAGEMENT FROM THE SLEEVE. FURTHERMORE, THE PRESENCE OF WEAR DEBRIS WAS OBSERVED, AND IS CONSISTENT WITH PREVIOUS COMPLAINT EVALUATIONS. CORRECTIVE ACTION TAKEN: THE CRC DETERMINED THAT NO FIELD ACTION WAS NECESSARY AT THAT TIME. THIS ISSUE WILL BE ADDRESSED VIA THE EQUINOXE ERGO INSTRUMENTATION REDESIGN PROJECT. IN THE NEW EQUINOXE ERGO INSTRUMENTATION, THE MODULAR CANNULATED TRI-DRIVE WILL NOT BE USED. THE ERGO INSTRUMENTATION WILL BE IMPLEMENTED AS PART OF A PHASED ROLL-OUT, FIRST IN THE UNITED STATES AND THEN TO OTHER MARKETS, BEGINNING IN MID-2019. FULL IMPLEMENTATION IS EXPECTED BY Q4 2020 (US) AND Q4 2021 (OUS). THE COMMITTEE WILL RECONVENE IF A THRESHOLD OF > 15 EVENTS OR = 1 EVENT WITH HARM IS REACHED. EVENT TRACKING WILL BE CAPTURED IN CAPA(B)(4). ACCORDING TO THE CAPA INVESTIGATION RESULTS, THE CAPA(B)(4) INVESTIGATION CAN BE CLOSED WITHOUT FURTHER CORRECTION OR CONTAINMENT OR CORRECTIVE OR PREVENTIVE ACTIONS. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: PER CAPA(B)(4), THE DISASSEMBLED TRI-DRIVE HANDLES REPORTED WERE LIKELY THE RESULT OF THE RETAINING RING RESPONSIBLE FOR MAINTAINING ASSEMBLY OF THE SLEEVE ON THE MODULAR CANNULATED TRI-DRIVE SHAFT BECOMING DEFORMED AS A RESULT OF HOLDING THE SLEEVE STATIONARY WHILE REAMING. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. ALL PIECES OF THE BROKEN INSTRUMENT WERE RETRIEVED AND THERE WAS NO REAL DELAY TO SURGERY. THERE WAS NO REPORTED PATIENT ADVERSE EVENT. PER CAPA(B)(4), THE DISASSEMBLED TRI-DRIVE HANDLES REPORTED WERE LIKELY THE RESULT OF THE RETAINING RING RESPONSIBLE FOR MAINTAINING ASSEMBLY OF THE SLEEVE ON THE MODULAR CANNULATED TRI-DRIVE SHAFT BECOMING DEFORMED AS A RESULT OF HOLDING THE SLEEVE STATIONARY WHILE REAMING. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00361 | 2022-04-04 | 2022-03-08 | I, n, j, u, r, y | Manufacturer report | TRULIANT PS CEM FEM PS CEM RIGHT SZ 3 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT PS CEM FEM PS CEM RIGHT SZ 3 | 02-020-11-0330 |  | AS REPORTED BY THE LEGAL TEAM, APPROXIMATELY 3 YEARS POST INITIAL LEFT TKA, THIS FEMALE PATIENT IS ASKING TO BE COMPENSATED FOR PAIN AND SUFFERING. THERE IS NO INDICATION OF REVISION ON EITHER SIDE TO DATE. | CONCOMITANT MEDICAL PRODUCTS: TRULIANT TIB IMP PS INSERT SZ 3 11MM (CAT# 02-022-35-3011 / (B)(4). TRULIANT TIB FIT TRAY CEM SZ 3F / 3T (CAT# 02-022-45-3030 / (B)(4). ADVANCED PATELLA 35MM 3 PEG IMPLANT (CAT# 200-07-35 / (B)(4). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00354 | 2022-04-04 | 2018-11-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | MINI HEADED FIXATION PIN REPLACED BY 201-78-15 | EXACTECH, INC. |  | 201-78-05 |  | IT WAS REPORTED FROM OUS THAT DURING AN INITIAL ORTHOPEDIC SURGERY THE HOLDING PIN BROKE. THERE WERE NO PIECES OF THE BROKEN DEVICE LEFT IN THE PATIENT. THE PROCEDURE WAS COMPLETED WITH SPARE DEVICES. THE PATIENT WAS STABLE LEAVING THE OPERATING ROOM. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR L0X OPTICAL UNLESS OTHERWISE SPECIFIED: THE FRACTURED HEAD OF THE MINI HEADED FIXATION PIN APPEARS CONSISTENT WITH APPLYING A BENDING MOMENT TO THE PROXIMAL HEAD OF THE PIN WITH THE PIN PULLER INSTRUMENT. TENSION WAS LIKELY APPLIED TO THE PIN PULLER AT AN ANGLE RATHER THAN AXIALLY AS INTENDED. ALL OTHER ASPECTS OF THE DEVICES APPEAR UNREMARKABLE AND CONSISTENT WITH YEARS OF USE. DESIGN-RELATED ISSUES: THE DESIGN OF THE MINI HEADED FIXATION PIN HAS BEEN IN THE FIELD SINCE 1999. EXACTECH IS AWARE OF 5 COMPLAINT REPORTS INVOLVING 6 BROKEN MINI HEADED FIXATION PINS SINCE 2008. THE FREQUENCY OF OCCURRENCE RANKING SCALE IS VERY LOW; THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: MANUFACTURING RELATED DATA COULD NOT BE REVIEWED BECAUSE NO INFORMATION REGARDING THE MANUFACTURING LOTS OR SERIAL NUMBERS WERE PROVIDED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. THE BROKEN PIN REPORTED WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE HEAD OF THE PIN WITH THE PIN PULLER, WHICH ALLOWED FOR THE PROXIMAL RIM TO FRACTURE. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY DESIGN ISSUES. THERE WAS NO REPORTED PATIENT ADVERSE EVENT. NO DEVICE FRAGMENTS WERE LEFT IN THE PATIENT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN PIN REPORTED WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE HEAD OF THE PIN WITH THE PIN PULLER, WHICH ALLOWED FOR THE PROXIMAL RIM TO FRACTURE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00358 | 2022-04-04 | 2022-03-08 | I, n, j, u, r, y | Manufacturer report | TRULIANT PS CEM FEM PS CEM LEFT SZ 3.5 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT PS CEM FEM PS CEM LEFT SZ 3.5 | 02-020-11-0235 |  | AS REPORTED BY THE LEGAL TEAM, APPROXIMATELY 3 YEARS POST INITIAL RIGHT TKA, THIS FEMALE PATIENT IS ASKING TO BE COMPENSATED FOR PAIN AND SUFFERING. THERE IS NO INDICATION OF REVISION ON EITHER SIDE TO DATE. | CONCOMITANT MEDICAL PRODUCTS: TRULIANT TIB IMP PS INSERT SZ 3.5 11MM (CAT# 02-022-35-3511 / SERIAL# (B)(4)). TRULIANT TIB FIT TRAY CEM SZ 3.5F / 2.5T (CAT# 02-022-45-3525 / SERIAL# (B)(4)). THREE PEG PATELLA 38MM (CAT# 200-02-38 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00355 | 2022-04-04 | 2018-11-19 | I, n, j, u, r, y | Manufacturer report | EXACTECH | MINI HEADED FIXATION PIN REPLACED BY 201-78-15 | EXACTECH, INC. |  | 201-78-05 |  | IT WAS REPORTED FROM OUS THAT DURING AN INITIAL ORTHOPEDIC SURGERY THE HOLDING PIN BROKE. THERE WERE NO PIECES OF THE BROKEN DEVICE LEFT IN THE PATIENT. THE PROCEDURE WAS COMPLETED WITH SPARE DEVICES. THE PATIENT WAS STABLE LEAVING THE OPERATING ROOM. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION.H.10 INFORMATION WAS NOT PROVIDED FOR: A.5, A.6VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR L0X OPTICAL UNLESS OTHERWISE SPECIFIED) - THE FRACTURED HEAD OF THE MINI HEADED FIXATION PIN APPEARS CONSISTENT WITH APPLYING A BENDING MOMENT TO THE PROXIMAL HEAD OF THE PIN WITH THE PIN PULLER INSTRUMENT. TENSION WAS LIKELY APPLIED TO THE PIN PULLER AT AN ANGLE RATHER THAN AXIALLY AS INTENDED. ALL OTHER ASPECTS OF THE DEVICES APPEAR UNREMARKABLE AND CONSISTENT WITH YEARS OF USE.DESIGN-RELATED ISSUES: THE DESIGN OF THE MINI HEADED FIXATION PIN HAS BEEN IN THE FIELD SINCE 1999. EXACTECH IS AWARE OF 5 COMPLAINT REPORTS INVOLVING 6 BROKEN MINI HEADED FIXATION PINS SINCE 2008. THE FREQUENCY OF OCCURRENCE RANKING SCALE IS VERY LOW; THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED.MFG-RELATED ISSUES: MANUFACTURING RELATED DATA COULD NOT BE REVIEWED BECAUSE NO INFORMATION REGARDING THE MANUFACTURING LOTS OR SERIAL NUMBERS WERE PROVIDED.CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR.THE BROKEN PIN REPORTED WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE HEAD OF THE PIN WITH THE PIN PULLER, WHICH ALLOWED FOR THE PROXIMAL RIM TO FRACTURE.IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. ¿ CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION.IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION.IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES.THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY DESIGN ISSUES. THERE WAS NO REPORTED PATIENT ADVERSE EVENT. NO DEVICE FRAGMENTS WERE LEFT IN THE PATIENT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN PIN REPORTED WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE HEAD OF THE PIN WITH THE PIN PULLER, WHICH ALLOWED FOR THE PROXIMAL RIM TO FRACTURE. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00352 | 2022-04-04 | 2018-11-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | MINI HEADED FIXATION PIN REPLACED BY 201-78-15 | EXACTECH, INC. |  | 201-78-05 |  | IT WAS REPORTED FROM OUS THAT AFTER USE OF PINS IN AN ORTHOPEDIC SURGICAL CASE THE HEAD OF THE PINS WERE NOTED TO BE BROKEN. ALL PIECES OF THE PINS WERE ACCOUNTED FOR, NO PIECES FELL INTO THE SURGICAL SITE. THE PATIENT WAS STABLE AFTER SURGERY. THE REPRESENTATIVE WAS PRESENT AT TIME OF SURGERY. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. 1 OF 2 MDRS FOR THIS EVENT | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR L0X OPTICAL UNLESS OTHERWISE SPECIFIED) - THE FRACTURED HEADS OF THE PINS APPEAR CONSISTENT WITH APPLYING A BENDING MOMENT TO THE PROXIMAL HEADS WITH THE PIN PULLER INSTRUMENT. TENSION WAS LIKELY APPLIED TO THE PIN PULLER AT AN ANGLE RATHER THAN AXIALLY AS INTENDED. THE FRACTURED PIECES WERE NOT RETURNED WITH THE MAIN BODIES OF THE PINS. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE. DESIGN-RELATED ISSUES: THE DESIGN OF THE MINI HEADED FIXATION PIN HAS BEEN IN THE FIELD SINCE 1999. EXACTECH IS AWARE OF (B)(4) SIMILAR COMPLAINT REPORTS INVOLVING (B)(4) DEVICES SINCE 2008. THE FREQUENCY OF OCCURRENCE RANKING SCALE IS VERY LOW, THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: MANUFACTURING RELATED DATA COULD NOT BE REVIEWED BECAUSE NO INFORMATION REGARDING THE MANUFACTURING LOTS OR SERIAL NUMBERS WERE PROVIDED. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. THE BROKEN PINS REPORTED WERE LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE HEAD OF THE PINS WITH THE PIN PULLER, WHICH ALLOWED FOR THE PROXIMAL RIMS TO FRACTURE. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THESE DEVICES ARE USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY DESIGN ISSUES. THERE WAS NO REPORTED PATIENT ADVERSE EVENT DUE TO THE PIN HEAD FRAGMENTS. AN INVESTIGATION WAS CONDUCTED; THE BROKEN PINS REPORTED WERE LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE HEAD OF THE PINS WITH THE PIN PULLER, WHICH ALLOWED FOR THE PROXIMAL RIMS TO FRACTURE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00359 | 2022-04-04 | 2018-12-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | MDLR CNLTD TRIDRV DRVR | EXACTECH, INC. |  | 315-25-00 |  | IT WAS REPORTED FROM THE US THAT THE DEVICE WAS BROKEN. THE PIECE CONNECTED TO REAMER LOCKED, THEN DISENGAGED FROM REAMER. IT WAS REPORTED THAT NO PIECES WERE LEFT IN THE PATIENT. THE AGENT WAS PRESENT AT SURGERY. THERE WAS NO DELAY OR CLINICAL CONSEQUENCE TO THE PATIENT. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO, LOT#. DEVICE WAS NOT AVAILABLE FOR ANALYSIS. DESIGN-RELATED ISSUES: THE DESIGN OF THE TRI-DRIVE CANNULATED HANDLE HAS BEEN IN THE FIELD SINCE 2009. (B)(4). THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: MANUFACTURING DATA COULD NOT BE EVALUATED BECAUSE NO FURTHER INFORMATION REGARDING THE SERIAL NUMBERS OR MANUFACTURING LOT WAS PROVIDED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD, RISK IS CAPTURED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY UNDERLYING CAUSE: ACCORDING TO THE INFORMATION PROVIDED, THE SALES REPRESENTATIVE REPORTED, "BROKEN. PIECE CONNECTED TO REAMER LOCKED AND THEN DISENGAGED FROM REAMER." THE BROKEN TRI-DRIVE CANNULATED HANDLE REPORTED WAS LIKELY THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY AND ALLOWED THE TIP TO DISASSEMBLE. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE DEVICE WAS NOT AVAILABLE FOR EVALUATION AND NO FURTHER INFORMATION WAS PROVIDED. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES. THERE WAS NO REPORTED PATIENT ADVERSE EVENT. IT WAS REPORTED THAT NO PIECES WERE LEFT IN THE PATIENT. AN INVESTIGATION WAS CONDUCTED; ACCORDING TO THE INFORMATION PROVIDED, THE SALES REPRESENTATIVE REPORTED, "BROKEN. PIECE CONNECTED TO REAMER LOCKED AND THEN DISENGAGED FROM REAMER. "THE BROKEN TRI-DRIVE CANNULATED HANDLE REPORTED WAS LIKELY THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY AND ALLOWED THE TIP TO DISASSEMBLE. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE DEVICE WAS NOT AVAILABLE FOR EVALUATION AND NO FURTHER INFORMATION WAS PROVIDED. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00356 | 2022-04-04 | 2018-11-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRI-DRIVE MODULAR HANDLE | EXACTECH, INC. |  |  |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE TRI-DRIVER CAME APART WHEN INSERTING A REAMER. THE CASE CONTINUED AND ¿OUTCOME WAS GREAT¿. PATIENT¿S HEALTH WAS STABLE LEAVING THE OR. NO PIECES WERE LEFT IN THE PATIENT. NO DELAY TO SURGERY. AGENT WAS PRESENT AT TIME OF SURGERY. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO, PRODUCT INFO. (B)(4) ¿TRI-DRIVE MODULAR REAMER DRIVER DISASSEMBLY. FINDINGS: THE DESIGN OF THE TRI-DRIVE MODULAR REAMER DRIVER HAS BEEN IN THE FIELD SINCE 2009. (B)(4). THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE IN THE RMR. THE DISASSEMBLED TRI-DRIVE MODULAR REAMER DRIVER MAY RESULT IN A DELAY IN SURGERY TO FIND A NEW ONE. TYPICALLY, IN ANATOMIC SHOULDER CASES BOTH ANATOMIC AND REVERSE SETS ARE IN THE OR. IF THE TRI-DRIVE WERE TO BREAK, A SECOND TRI-DRIVE WOULD BE IN THE REVERSE SHOULDER SET. CORRECTIVE ACTION TAKEN: THE CRC DETERMINED THAT NO FURTHER ACTION WAS REQUIRED AT THIS TIME AS THIS ISSUE WILL BE ADDRESSED VIA THE EQUINOXE ERGO INSTRUMENTATION REDESIGN PROJECT. IN THE NEW EQUINOXE ERGO INSTRUMENTATION, THE TRI-DRIVE HANDLES WILL NO LONGER BE USED. SINCE THE DESIGN OF THE PILOT TIP REAMERS ARE CHANGING (ALL THE REAMERS ARE CANNULATED), THERE WILL NOW BE TWO DIFFERENT HANDLES: A PILOT TIP HANDLE AND A CANNULATED HANDLE, REPLACING THE TRI-DRIVE. THIS NEW ERGO INSTRUMENTATION DESIGN WILL BE IMPLEMENTED AS PART OF A PHASED ROLL-OUT, FIRST IN THE US AND THEN TO OTHER MARKETS, BEGINNING IN MID-2019. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. THERE IS NO INDICATION THAT THE DEVICE PIECES FELL INTO THE SURGICAL SITE. ¿NO PIECES WERE LEFT IN THE PATIENT.¿ (B)(4) ¿TRI-DRIVE MODULAR REAMER DRIVER DISASSEMBLY: THESE BROKEN TRI-DRIVES WERE LIKELY THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING OR THE TRI-DRIVE CONTACTING A RETRACTOR, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY, ALLOWING THE TIP TO DISASSEMBLE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00357 | 2022-04-04 | 2018-12-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | EQ SCR DRV HNDL | EXACTECH, INC. |  | 301-07-50 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY; THE SURGEON WAS USING THE SCREWDRIVER HANDLE ON THE BACK TABLE TO TIGHTEN THE REVERSE SHOULDER TORQUE SCREW ONTO THE HUMERAL ADAPTER TRAY AND THE TINES ON THE SCREWDRIVER BROKE. NO DAMAGE WAS DONE TO ANY COMPONENTS AND THE PATIENT PROCEEDED FORWARD WITHOUT INCIDENT. SUCCESSFUL OUTCOME. PATIENT¿S HEALTH WAS STABLE LEAVING THE OPERATING ROOM. NO PIECES WERE LEFT IN THE PATIENT. THE SALES REPRESENTATIVE WAS PRESENT AT THE SURGERY AND THERE WAS NO DELAY. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. VISUAL EVALUATION CONDITION/FINDINGS THE DAMAGED INSTRUMENT APPEARS CONSISTENT WITH CRACK INITIATION, PROPAGATION, AND BRITTLE FRACTURE OF THE INSTRUMENT MATERIAL. THE LOCATION AND NATURE OF THE FRACTURE APPEARS CONSISTENT WITH APPLYING A SIDE-TO-SIDE FORCE WHILE TRYING TO MANIPULATE THE ATTACHED DEVICE INTO A SPECIFIC POSITION DURING SURGERY. THE CRACK MAY HAVE INITIATED DURING A PREVIOUS USE AND GONE UNNOTICED. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE AND CONSISTENT WITH YEARS OF USE. DESIGN-RELATED ISSUES: THE DESIGN OF THE HUDSON SCREW- DRIVER HANDLE HAS BEEN IN THE FIELD SINCE 2005. EXACTECH IS AWARE OF 3 OTHER SIMILAR COMPLAINT REPORTS INVOLVING HUDSON DRIVER HANDLES SINCE 2008. (B)(4). THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF (B)(4) UNITS, WHICH HAS BEEN IN THE FIELD SINCE 2011. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING-RELATED A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS HUDSON SCREW-DRIVER HANDLE, 750-2002-039-RMR-INSTRUMENT ADDITIONS REV G, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 17. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS " VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE FRACTURED HUDSON SCREW -DRIVER HANDLE REPORTED WAS LIKELY THE RESULT OF CRACK INITIATION, LIKELY BY APPLYING A SIDE-TO-SIDE FORCE WHILE TRYING TO MANIPULATE THE ATTACHED INSTRUMENT INTO A SPECIFIC POSITION DURING A PREVIOUS USE, PROPAGATION, AND ULTIMATE BRITTLE FRACTURE OF THE TINES. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES. THERE WAS NO PATIENT ADVERSE EVENT, THE TINES BROKE ON THE BACK TABLE. AN INVESTIGATION WAS CONDUCTED; THE FRACTURED HUDSON SCREW -DRIVER HANDLE REPORTED WAS LIKELY THE RESULT OF CRACK INITIATION, LIKELY BY APPLYING A SIDE-TO-SIDE FORCE WHILE TRYING TO MANIPULATE THE ATTACHED INSTRUMENT INTO A SPECIFIC POSITION DURING A PREVIOUS USE, PROPAGATION, AND ULTIMATE BRITTLE FRACTURE OF THE TINES. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00367 | 2022-04-05 | 2022-03-15 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 38MM HUMERAL CONST LINER +0 | EXACTECH, INC. | 320-38-10 | 320-38-10 |  | AS REPORTED, APPROXIMATELY 3 YEARS POST OP THE INITIAL LEFT TSA, THIS (B)(6) MALE PATIENT WAS REVISED DUE TO DISSOCIATION OF THE LINER FROM THE TRAY WHILE LIFTING DOG CAGES. PATIENT WAS ON VACATION IN (B)(6) TWO WEEKS AGO. WHILE LIFTING DOG CAGES HE FELT A POP IN THE SHOULDER FOLLOWED BY PAIN AN INABILITY TO ROTATE THE ARM EASILY. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES HAVE BEEN RETURNED. | CONCOMITANT DEVICE(S): 320-15-05, 6053781 - EQ REV LOCKING SCREW; 320-01-38, 6068866 - EQUINOXE REVERSE 38MM GLENOSPHERE. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00368 | 2022-04-05 | 2022-03-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | VANTAGE | LP TALAR IMPACTOR INSERT | EXACTECH, INC. | 351-07-04 | 351-07-04 |  | SURGEON WAS IMPACTING THE TALUS AND WHEN HE WENT TO MALLET, THE BLACK TIP OF THE TALUS IMPACTOR BROKE OFF. THE TINY PIECE THAT BROKE OFF THE TIP WAS DISCARDED IN THE OR BUT I WAS ABLE TO KEEP THE MAIN PORTION OF THE BLACK TIP OFF THE TALUS IMPACTOR. PATIENT WAS NOT AFFECTED BY THIS. PATIENT NOT AFFECTED BY THIS. THERE WAS NO SURGICAL DELAY/PROLONGATION. NO PARTS OR PIECES FELL INTO THE PATIENT WOUND SITE. DEVICE TO BE RETURNED. | PENDING EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00360 | 2022-04-05 | 2019-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | PIN DRIVER, SHORT LOW PROFILE | EXACTECH, INC. |  | 521-78-11 |  | IT WAS REPORTED FROM OUS THAT A CONSIGNMENT INSTRUMENT 521-78-11 HAD A PIECE BREAK OFF IT DURING TOTAL KNEE REPLACEMENT SURGERY. NO PIECES FELL INTO THE SURGICAL SITE. THERE WERE NO PIECES LEFT IN THE PATIENT. THERE WAS NO DELAY TO THE SURGERY. THE PATIENT WAS DISCHARGED HOME WITHOUT COMPLICATIONS. THE DEVICE WAS DISPOSED OF BY THE HOSPITAL AND WILL NOT BE RETURNED. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DESIGN-RELATED ISSUES: THE DESIGN OF THE GPS PIN DRIVER HAS BEEN IN THE FIELD SINCE 2014. EXACTECH IS AWARE OF 4 SIMILAR COMPLAINT REPORTS INVOLVING BROKEN GPS PIN DRIVERS SINCE 2014. BECAUSE THE GPS PIN DRIVER HAS BEEN USED IN OVER 12,000 TOTAL KNEE SURGERIES USING THE GPS SYSTEM, SURGERY DATA FROM THE GPS WEB STATION USAGE FILE WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: MANUFACTURING DATA COULD NOT BE REVIEWED BECAUSE THE SERIAL/LOT INFORMATION WAS NOT PROVIDED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) AND RISK ASSESSMENT AND CONTROLS REPORT (RACR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR/RACR FOR THIS PIN DRIVER, RMR-00042 REV ¿ AND RACR-00056 REV -, WERE REVIEWED. THE RISK IS CAPTURED, HOWEVER THE RACR IS BEING UPDATED UNDER DCRSP-19-611 TO CAPTURE AN APPROPRIATE HARM FOR THIS HAZARDOUS SITUATION. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, THE RISK IS CAPTURED IN THE RMR AND RACR, AND THE HARM WILL BE UPDATED. MOST LIKELY CAUSE: THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE DRIVER WHILE INSERTING A PIN INTO BONE. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. ¿ CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY DESIGN ISSUES. THERE WAS NO PATIENT ADVERSE EVENT DUE TO THE FRACTURED DEVICE. AN INVESTIGATION WAS CONDUCTED; THE BROKEN INSTRUMENT WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE DRIVER WHILE INSERTING A PIN INTO BONE. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00369 | 2022-04-05 | 2019-05-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRULIANT TIB TAMP GUIDE | EXACTECH, INC. |  | 02-029-22-2000 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE TIBIAL PUNCH. THE SURGEON PUNCHED THE TIBIA THEN HANDED THE PUNCH BACK TO THE TECH. AS THE TECH WAS PUTTING THE PUNCH BACK IN THE TRAY IT FELL APART. NONE OF THE DEVICE PIECES WERE IN THE WOUND. THE WOUND WAS CLEANED OUT. ALL DEVICE PIECES WERE IDENTIFIED. THE PATIENT WAS STABLE AS THEY LEFT THE OPERATING ROOM. THERE WAS NO DELAY TO SURGERY OR REPORTED PATIENT ADVERSE EVENT. THE AGENT WAS PRESENT AT THE SURGERY. INACTIVE AGENCY WITH NO HOSPITAL/OTHER CONTACT INFORMATION FOR THE EVENT. NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: AGE OR DATE OF BIRTH, SEX, WEIGHT [76 NOT REPORTED KG OR POUNDS], ETHNICITY, RACE. FINDINGS: THE BUTTON COMPONENT (02-029-222-2005) OF THE TIBIAL TAMP GUIDE DISENGAGED FROM THE INSTRUMENT (02-029-22-2000). ENGINEERING PRINTS OF 02-029-122-2000 WERE REVIEWED. THE OVERLAP BETWEEN THE SNAP RING¿S OD AND THE ID OF THE HOUSING LIP IS .008 FOR THE WORST-CASE SCENARIO. ALTHOUGH THE OVERLAP IS POSITIVE, THE SMALL AMOUNT COULD CAUSE DISENGAGEMENT. CORRECTIVE ACTION: PER (B)(4), THE COMMITTEE HAS DETERMINED THAT THE DESIGN WILL BE UPDATED, TESTED, AND IMPLEMENTED, AFTER WHICH A STAGED MARKET WITHDRAWAL WILL BE CONDUCTED. IN THE MEANTIME, THE EXACTECH TEAM WILL CONTINUE TO MONITOR THIS INSTRUMENT FOR THE ISSUE. CAPA(B)(4) IS STILL ONGOING AT THE TIME OF THIS EVALUATION, AND CORRECTIVE ACTIONS ARE SUBJECT TO CHANGE. REFER TO CAPA(B)(4) FOR MOST UP TO DATE INFORMATION. COMMENTS: WHEN INSPECTING THE INSTRUMENT, MOISTURE WAS FOUND IN THE WINDOW OF THE RETURNED DEVICE. THIS MOISTURE IS ADDRESSED IN (B)(4). THE ROOT CAUSE FOR THE BUTTON DISENGAGEMENT, AS DETERMINED BY CAPA(B)(4), WAS LIKELY THE RESULT OF INSUFFICIENT OVERLAP BETWEEN THE SNAP RING AND THE LIP OF THE HOUSING TO PREVENT THE OCCURRENCE OF DISENGAGEMENT. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. NONE OF THE DEVICE PIECES WERE IN THE WOUND. THE WOUND WAS CLEANED OUT. ALL DEVICE PIECES WERE IDENTIFIED. THE ROOT CAUSE FOR THE BUTTON DISENGAGEMENT, AS DETERMINED BY CAPA(B)(4), WAS LIKELY THE RESULT OF INSUFFICIENT OVERLAP BETWEEN THE SNAP RING AND THE LIP OF THE HOUSING TO PREVENT THE OCCURRENCE OF DISENGAGEMENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00370 | 2022-04-05 | 2019-05-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | MDLR CNLTD TRIDRV DRVR | EXACTECH, INC. |  | 315-25-00 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE REMOVABLE TIP ON TRI-DRIVE MODULAR HANDLE. THE DEVICE WAS BROKEN DURING REAMING CAUSING AN INSTABILITY TO THE CONNECTION WITH CENTER PEG DRILL. THE SURGEON COULD NOT HELP SUSPENDING APPROXIMATELY 3 MINUTES FOR CHECKING THE SURGICAL SITE. THE SURGEON STARTED THE SURGERY WITH AN UNSTABLE CONNECTION WITH THE DEVICE. THE INCIDENT RESULTED IN A 3-MINUTE DELAY IN SURGERY. THERE WAS NO REPORTED ADVERSE EVENT. THE PATIENT WAS STABLE LEAVING THE OR. MULTIPLE ATTEMPTS WERE MADE FOR ADDITIONAL INFORMATION. NO OTHER INFORMATION HAS BEEN PROVIDED ABOUT THE EVENT FROM THE CONTACTS. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INITIAL REPORTER NAME AND ADDRESS: COUNTRY IS (B)(6). FINDINGS: THE DESIGN OF THE TRI-DRIVE MODULAR REAMER DRIVER HAS BEEN IN THE FIELD SINCE 2009. EXACTECH IS AWARE OF 45 COMPLAINT REPORTS INVOLVING TIP DISASSEMBLY OF 47 TRI-DRIVES. BECAUSE THE TRI-DRIVE IS USED IN BOTH ANATOMIC AND REVERSE TOTAL SHOULDER ARTHROPLASTY, SALES DATA FOR ANATOMIC GLENOID COMPONENTS AND GLENOSPHERE COMPONENTS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE IN THE RMR. THE DISASSEMBLED TRI-DRIVE MODULAR REAMER DRIVER MAY RESULT IN A DELAY IN SURGERY TO FIND A NEW ONE. TYPICALLY, IN ANATOMIC SHOULDER CASES BOTH THE ANATOMIC AND THE REVERSE SETS ARE IN THE OR. IF THE TRI-DRIVE WERE TO BREAK, A SECOND TRI-DRIVE WOULD BE IN THE REVERSE SHOULDER SET. CORRECTIVE ACTION TAKEN: THE CRC DETERMINED THAT NO FURTHER ACTION WAS REQUIRED AT THIS TIME AS THIS ISSUE WILL BE ADDRESSED VIA THE EQUINOXE ERGO INSTRUMENTATION REDESIGN PROJECT. IN THE NEW EQUINOXE ERGO INSTRUMENTATION, THE TRI-DRIVE HANDLES WILL NO LONGER BE USED. SINCE THE DESIGN OF THE PILOT TIP REAMERS ARE CHANGING (ALL THE REAMERS ARE CANNULATED), THERE WILL NOW BE TWO DIFFERENT HANDLES: A PILOT TIP HANDLE AND A CANNULATED HANDLE, REPLACING THE TRI-DRIVE. THIS NEW ERGO INSTRUMENTATION DESIGN WILL BE IMPLEMENTED AS PART OF A PHASED ROLL-OUT, FIRST IN THE US AND THEN TO THE OTHER MARKETS BEGINNING IN MID-2019. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: THIS BROKEN TRI-DRIVE REPORTED WAS LIKELY THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING OR THE TRI-DRIVE CONTACTING A RETRACTOR, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY, ALLOWING THE TIP TO DISASSEMBLE. IFU 700-096-181: INSTRUMENT INSPECTION : VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THE DEVICE CAME APART DURING REAMING, THE SURGEON HAD TO SEARCH THE SURGICAL SITE. THERE IS NO REPORTED PATIENT ADVERSE EVENT. AN INVESTIGATION WAS CONDUCTED: THIS BROKEN TRI-DRIVE REPORTED WAS LIKELY THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING OR THE TRI-DRIVE CONTACTING A RETRACTOR, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY, ALLOWING THE TIP TO DISASSEMBLE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00371 | 2022-04-05 | 2019-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRUL FEM IMPACTOR HEAD LOCKING | EXACTECH, INC. |  | 02-029-19-1000 |  | IT WAS REPORTED BY EXACTECH (B)(6) THAT THEY HAD A BROKEN INSTRUMENT THAT THEY WANTED REPLACED. THE DEVICE WAS RECEIVED ON 9/30/2019. ATTEMPTS WERE MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REP, AND A RESPONSE WAS RECEIVED. PER THE REP, THE EVENT WAS REPORTED BY A WAREHOUSE WORKER. THEY DO NOT KNOW WHICH HOSPITAL IT OCCURRED IN. THE BROKEN PART WAS RECEIVED IN THE WAREHOUSE FROM A LOANER SET AND IT WAS A NEW PART. THE AGENT DOES NOT HAVE THE EXACT EVENT DATE. NO PARTS/PIECES FELL INTO THE SURGICAL SITE OR WERE LEFT IN THE PATIENT. THERE WAS NO DELAY OR INJURY/CLINICAL CONSEQUENCE/ADVERSE EVENT TO THE PATIENT. THERE IS NO REPORTED PATIENT ADVERSE EVENT. THE FRACTURED DEVICE CAN ALLOW PIECES TO FALL INTO/NEAR SURGICAL SITES. THERE IS NO OTHER INFORMATION ABOUT THE DEVICE FROM THE CONTACTS. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: AGE OR DATE OF BIRTH, SEX, WEIGHT, ETHNICITY, AND RACE. FINDINGS: THE REPORTED FRACTURED IMPACTOR FACE APPEARS CONSISTENT WITH MATERIAL DEGRADATION DUE TO NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. MOST LIKELY UNDERLYING CAUSE: BASED ON SIMILAR PREVIOUS COMPLAINT INVESTIGATIONS, THE FRACTURED IMPACTOR FACE REPORTED WAS LIKELY THE RESULT OF MATERIAL DEGRADATION ASSOCIATED WITH YEARS OF NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. IFU 700-096-181: INSTRUMENT INSPECTION. VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. THE FRACTURED DEVICE CAN ALLOW PIECES TO FALL INTO/NEAR SURGICAL SITES. AN INVESTIGATION WAS CONDUCTED: BASED ON SIMILAR PREVIOUS COMPLAINT INVESTIGATIONS, THE FRACTURED IMPACTOR FACE REPORTED WAS LIKELY THE RESULT OF MATERIAL DEGRADATION ASSOCIATED WITH YEARS OF NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00372 | 2022-04-05 | 2019-07-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | MINI HEADED FIXATION PIN REPLACED BY 201-78-15 | EXACTECH, INC. |  | 201-78-05 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE INSTRUMENTATION. THE HOLDING PIN HEAD WAS BROKEN. THE BROKEN PIN HEAD IS MATERIAL FATIGUE/DEBRIS. THE SURGERY WAS COMPLETED, AND THE PATIENT IS FINE. NO PIECES WERE LEFT IN THE PATIENT. NEW INFORMATION 27 JUL 2021: DAMAGE WAS DETECTED AFTER THE SURGERY. NO PIECES FELL INTO THE SURGICAL SITE. THERE WAS NO SURGICAL DELAY OR PATIENT IMPACT. NO OTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS ABOUT THE DEVICE/EVENT/PATIENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. FINDINGS: THE FRACTURED HEAD OF THE MINI HEADED FIXATION PIN MAY HAVE BEEN THE RESULT OF APPLYING A BENDING MOMENT TO THE PROXIMAL HEAD OF THE PIN WITH THE PIN PULLER INSTRUMENT. TENSION WAS LIKELY APPLIED TO THE PIN PULLER AT AN ANGLE RATHER THAN AXIALLY AS INTENDED. THE PATIENT WAS REPORTED STABLE LEAVING THE OPERATING ROOM AND NO PIECES WERE LEFT IN THE PATIENT. MOST LIKELY UNDERLYING CAUSE: BASED ON PREVIOUS SIMILAR INVESTIGATIONS, THE BROKEN PIN REPORTED WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE HEAD OF THE PIN WITH THE PIN PULLER, WHICH ALLOWED FOR THE PROXIMAL RIM TO FRACTURE. RMR REVIEW: THE RMRS FOR THESE MINI HEADED FIXATION PINS, 750-2001-029-RMR-INSTRUMENTS REV D AND 750-2003-019-RMR-INSTRUMENTS REV B, WERE REVIEWED. THE RISK IS CAPTURED. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT, THE BROKEN PIN HEAD IS MATERIAL FATIGUE/DEBRIS. AN INVESTIGATION WAS CONDUCTED: BASED ON PREVIOUS SIMILAR INVESTIGATIONS, THE BROKEN PIN REPORTED WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE HEAD OF THE PIN WITH THE PIN PULLER, WHICH ALLOWED FOR THE PROXIMAL RIM TO FRACTURE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00373 | 2022-04-05 | 2019-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | CANNULATED REAMER DRIVER | EXACTECH, INC. |  | 315-50-10 |  | IT WAS REPORTED FROM THE US VIA THE ERGO SURVEY THAT THE NEW ERGO CANNULATED REAMER HANDLES DO NOT HOLD THE REAMER TIPS ON SECURE. AFTER THE SURGEON REAMS THE GLENOID THE REAMER STAYS ON THE K-WIRE AND DETACHES FROM THE REAMER HANDLE. THE SURGEON THEN HAS TO ASK FOR A CLAMP TO GRAB THE REAMER AND PULL IT OUT OF THE PATIENT OFF OF THE REAMER. THERE WERE DELAYS, BUT THE PATIENTS WERE NOT ADVERSELY AFFECTED. PATIENTS LEFT THE OR STABLE. THE REPRESENTATIVE WAS PRESENT AT THE TIME OF SURGERIES. THE DEVICE WAS NOT RECEIVED FOR ANALYSIS. INACTIVE AGENCY WITH NO HOSPITAL/OTHER CONTACT INFORMATION FOR THE EVENT. NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. FINDINGS: A REVIEW OF DESIGN PRINTS, DESIGN VERIFICATION, AND DESIGN VALIDATION, PER CAPA(B)(4), IDENTIFIED THE FOLLOWING: 1. LMC INTERFERENCE WAS NOT VERIFIED AS PART OF THE EQUINOXE ERGO TOLERANCE ANALYSIS REPORT ((B)(4)). THE DEVICE HISTORY RECORDS (DHRS) FOR THE DRIVERS AND REAMERS WERE REVIEWED AS PART OF CAPA(B)(4), AND ALL PARTS WERE ACCEPTED WITH CONFORMANCE TO THE DEVICE SPECIFICATIONS. THE ERGO PILOT RMR, 750-2016-002 REV ¿ DID NOT EXPLICITLY DEFINE THE FAILURE MODE OF THE REAMERS NOT BEING RETAINED IN THE DRIVERS. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: AS DETERMINED BY CAPA(B)(4), THE MATING ISSUE REPORTED IN (B)(4) WAS LIKELY THE RESULT OF NOT ADEQUATELY DEFINING THE ACCEPTANCE CRITERIA FOR THE LOCKING MECHANISM RESULTING IN NOT ENOUGH INTERFERENCE BETWEEN REAMERS AND REAMER DRIVERS. CORRECTIVE ACTION TAKEN: A CORRECTION AND REMOVAL COMMITTEE (CRC) MEETING WAS ALSO HELD TO ADDRESS THE NEED FOR FIELD ACTION. THE COMMITTEE DETERMINED THAT DUE TO THE LOW RISK TO HEALTH AND LOW RATE OF OCCURRENCE, NO FIELD ACTION IS NECESSARY AT THIS TIME. SEE (B)(4) FOR DETAILS. THE SYSTEMATIC ISSUES IN REGARD TO DESIGN CONTROLS NOT BEING FOLLOWED ARE BEING FURTHER ADDRESSED AS PART OF CAPA(B)(4). THE DIR WAS REVISED TO INCLUDE THE ACCEPTANCE CRITERIA FOR THE LOCKING MECHANISM. SEE (B)(4). NEW PART NUMBERS WERE CREATED AND INSTRUMENT DESIGN/DRAWING WERE REVISED. SEE (B)(4). A NEW ORDER FOR PARTS WERE RECEIVED ((B)(4)), AND A CADAVER VALIDATION LAB WAS COMPLETED WITH SURGEON USERS. VERIFICATION AND VALIDATION WILL BE PERFORMED ON THE NEW DESIGN, AND DESIGN TRANSFER AND FINAL DESIGN REVIEW WILL BE CONDUCTED. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO MANUFACTURING ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT WHEN THE SURGEON MUST PULL THE REAMER OUT OF THE PATIENT. AS DETERMINED BY CAPA(B)(4), THE MATING ISSUE REPORTED WAS LIKELY THE RESULT OF NOT ADEQUATELY DEFINING THE ACCEPTANCE CRITERIA FOR THE LOCKING MECHANISM RESULTING IN NOT ENOUGH INTERFERENCE BETWEEN REAMERS AND REAMER DRIVERS. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00374 | 2022-04-05 | 2019-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LOGIC TIBIA PS MOD INSRT TRIAL SZ 4 9MM | EXACTECH, INC. |  | 02-013-35-4009 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE TRIAL. WHILE TRIALING, THE INSERT CHIPPED AND WAS UNABLE TO BE USED AGAIN DURING THE CASE. THERE WAS NO DELAY, AND THE PATIENT WAS NOT ADVERSELY AFFECTED. THE REPRESENTATIVE WAS PRESENT AT THE TIME OF SURGERY. THE DEVICE WAS RETURNED. ATTEMPTS WERE MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REPRESENTATIVE, BUT NO RESPONSES WERE RECEIVED. THERE HAS BEEN NO OTHER INFORMATION PROVIDED BY THE CONTACTS FOR THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: AGE OR DATE OF BIRTH, SEX, WEIGHT, ETHNICITY,AND RACE. DESIGN-RELATED ISSUES: THE DESIGN OF THE LOGIC PS TIBIAL INSERT TRIAL HAS BEEN IN THE FIELD SINCE 2008. EXACTECH IS AWARE OF 49 SIMILAR COMPLAINT REPORTS INVOLVING 69 CRACKED/FRACTURED TIBIAL INSERT TRIALS (17 CRACKED; 52 FRACTURED) SINCE 2008. BECAUSE THE LOGIC PS TIBIAL INSERT TRIAL IS USED IN TOTAL KNEE SURGERIES, SALES DATA FOR TIBIAL INSERTS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4). THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF (B)(4) UNITS, WHICH HAVE BEEN IN THE FIELD SINCE 2009. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS LOGIC PS TIBIAL INSERT TRIAL, (B)(4) REV B, WAS REVIEWED. THE RISK IS CAPTURED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR. THE DAMAGED LOGIC PS TIBIAL INSERT TRIAL REPORTED WAS LIKELY THE RESULT OF EITHER YEARS OF NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES, THE DEVICE COMING IN CONTACT WITH SHARP SURGICAL TOOLS DURING TRIALING, OR A COMBINATION OF THE TWO. IFU 700-096-181: INSTRUMENT INSPECTION. VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT DUE TO THE FRAGMENT FROM THE TRIAL. AN INVESTIGATION WAS CONDUCTED: THE DAMAGED LOGIC PS TIBIAL INSERT TRIAL REPORTED WAS LIKELY THE RESULT OF EITHER YEARS OF NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES, THE DEVICE COMING IN CONTACT WITH SHARP SURGICAL TOOLS DURING TRIALING, OR A COMBINATION OF THE TWO. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00375 | 2022-04-05 | 2019-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRULIANT TIB TAMP GUIDE | EXACTECH, INC. |  | 02-029-22-2000 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE INSTRUMENTATION. THE BUTTON POPPED OFF TIBIAL TAMP GUIDE DURING IMPACTION. THERE WAS NO DELAY TO THE SURGERY. THE PATIENT WAS NOT NEGATIVELY AFFECTED. THERE WAS NO REPORTED PATIENT INJURY. NO PIECES WERE LEFT IN THE PATIENT. THE REPRESENTATIVE WAS NOT PRESENT AT THE TIME OF THE SURGERY. THE DEVICE WAS RETURNED. INACTIVE AGENCY WITH NO HOSPITAL/OTHER CONTACT INFORMATION FOR THE EVENT. NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI(B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: AGE OR DATE OF BIRTH,SEX, ETHNICITY, AND RACE. FINDINGS: THE BUTTON COMPONENT (02-029-222-2005) OF THE TIBIAL TAMP GUIDE DISENGAGED FROM THE INSTRUMENT (02-029-22-2000). ENGINEERING PRINTS OF 02-029-122-2000 WERE REVIEWED. THE OVERLAP BETWEEN THE SNAP RING¿S OD AND THE ID OF THE HOUSING LIP IS .008 FOR THE WORST-CASE SCENARIO. ALTHOUGH THE OVERLAP IS POSITIVE, THE SMALL AMOUNT COULD CAUSE DISENGAGEMENT. CORRECTIVE ACTION TAKEN: PER (B)(4), THE COMMITTEE HAS DETERMINED THAT THE DESIGN WILL BE UPDATED, TESTED, AND IMPLEMENTED, AFTER WHICH A STAGED MARKET WITHDRAWAL WAS CONDUCTED. PER CAPA(B)(4): 1. A STOCK AUDIT WAS PERFORMED FOR THE DEVICES LOCATED IN 01 (COC). SEE NPR-40446 AND NPR-45222 FOR DETAILS. 2. THE DESIGN WAS UPDATED TO INCREASE THE OVERLAP OF THE SNAP RING AND THE HOUSING LIP COMPONENTS. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO MANUFACTURING ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT DUE TO THE DISENGAGED DEVICE. PER CAPA(B)(4), THE ROOT CAUSE FOR THE BUTTON DISENGAGEMENT REPORTED WAS LIKELY THE RESULT OF INSUFFICIENT OVERLAP BETWEEN THE SNAP RING AND THE LIP OF THE HOUSING TO PREVENT THE OCCURRENCE OF DISENGAGEMENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00376 | 2022-04-05 | 2018-10-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRULIANT PIN PULLER SYRINGE STYLE | EXACTECH, INC. |  | 02-029-90-4100 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE TRULIANT PIN PULLER SYRINGE STYLE BROKE. THE SURGEON USED THE PIN PULLER WITH THE HEADED PIN IN IT TO PLACE IN THE TIBIA TRAY INTO THE BONE. THE BONE WAS SOFT, SO HE WAS ABLE TO JUST PUSH IT DOWN BY HAND. WHILE RELEASING THE PIN PULLER HE WENT TO GRAB THE NEXT PIN FROM THE TECH, HE PUSHED THE PIN PULLER TO EXTEND THE METAL GRABBING PRONGS AND ONE HAD BROKEN OFF AND FELL ONTO THE TABLE. HE WAS STILL ABLE TO GRAB THE NEXT PIN AS THREE PRONGS REMAINED. THE SURGERY DID NOT GET SLOWED DOWN BY THIS AND THE PATIENT WAS NOT AFFECTED. THE PATIENT LEFT THE OPERATING ROOM IN STABLE CONDITION. THE AGENCY IS INACTIVE, THERE IS NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. ONE OR MORE OF THE METAL PRONGS USED TO CLAMP A PIN HEAD ARE MISSING FROM THE RECEIVED TRULIANT SYRINGE PIN PULLERS. WHEN ANALYZING THE FRACTURE SURFACE FROM PREVIOUS DEVICE, THE FAILURE APPEARS CONSISTENT WITH BRITTLE FRACTURE. THIS TYPE OF FAILURE IS MOST LIKELY THE RESULT OF THE PRONGS EXPERIENCING EXCESSIVE FORCES. FURTHERMORE, IF THE PIN PULLER WAS USED OFF-AXIS RELATIVE TO THE BONE PIN, AN UNEVEN FORCE DISTRIBUTION MAY HAVE BEEN APPLIED TO THE METAL PRONGS AND COULD EXPLAIN WHY THEY FRACTURED FROM ONLY ONE SIDE. OVER TIME, THESE EXCESSIVE OFF-AXIS FORCES COULD POTENTIALLY LEAD TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE MATERIAL. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD; THE RISK IS CAPTURED. THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF THE PRONGS OF THE PIN PULLER EXPERIENCING EXCESSIVE FORCES DURING THE EXTRACTION OF BONE PINS. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. ¿ CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. THIS EVENT DID NOT LEAD TO PATIENT ADVERSE EVENT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF THE PRONGS OF THE PIN PULLER EXPERIENCING EXCESSIVE FORCES DURING THE EXTRACTION OF BONE PINS. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00382 | 2022-04-06 | 2018-08-08 | I, n, j, u, r, y | Manufacturer report | TRULIANT TIB IMP PS INSERT SZ 4 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT TIB IMP PS INSERT SZ 4 9MM | 02-022-35-4009 |  | AS REPORTED BY THE EXACTECH TRULIANT KNEE CLINICAL STUDY, APPROXIMATELY FOURTEEN MONTHS POST INITIAL TKA, THIS 41 Y/O MALE PATIENT HAD A REVISION DUE TO WORSENING PAIN, DETERIORATING RANGE OF MOTION, PROGRESSIVE DEFORMITY, AND SIGNIFICANT INTERFERENCE WITH FUNCTION THE ONLY OPTION FOR PAIN CONTROL AND FUNCTIONAL RESTORATION IS TOTAL KNEE REPLACEMENT REVISION. PER CLINICAL STUDY THIS REPORTED EVENT IS UNLIKELY RELATED TO THE DEVICE OR TOT THE PROCEDURE. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00380 | 2022-04-06 | 2022-03-16 | I, n, j, u, r, y | Manufacturer report | THREE PEG PATELLA 32MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | THREE PEG PATELLA 32MM | 200-02-32 |  | AS REPORTED, APPROXIMATELY FOURTEEN YEARS POST INITIAL LEFT TKA, THIS (B)(6) FEMALE PATIENT HAD A REVISION DUE TO PATELLA WAS LOOSE. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PHOTOS RECEIVED. THE DEVICES ARE NOT AVAILABLE FOR EVALUATION AS THEY WERE DISCARDED BY HOSPITAL. | CONCOMITANT MEDICAL PRODUCTS. PS TIBIAL INSERTS SZ 4, 9MM (CAT# 204-24-09 / (B)(4). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00381 | 2022-04-06 | 2021-01-01 | I, n, j, u, r, y | Manufacturer report | TRULIANT CR CEM FEM CR CEM LEFT SZ 4.5 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT CR CEM FEM CR CEM LEFT SZ 4.5 | 02-020-13-0245 |  | AS REPORTED BY THE EXACTECH TRULIANT KNEE CLINICAL STUDY, APPROXIMATELY THREE MONTHS POST INITIAL RTKA, THIS (B)(6) FEMALE PATIENT'S KNEE WAS STILL STIFF. SURGEON DID CLOSED MANIPULATION NINETEEN DAYS LATER AND PATIENT WILL FOLLOW UP WITH PHYSICAL THERAPY. PER CLINICAL STUDY THIS EVENT IS NOT RELATED TO THE DEVICE BUT DEFINITELY RELATED TO THE PROCEDURE. | CONCOMITANT MEDICAL PRODUCTS: TRULIANT TIB IMP CRC INSERT SZ 4.5, 11M (CAT# 02-022-51-4511) AND TRULIANT TIB FIT TRAY CEM SZ 4.5F/4.5T (CAT# 02-022-45-4545). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00378 | 2022-04-06 | 2022-03-17 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | 36MM GLENOSPHERE | EXACTECH, INC. | 320-01-36 | 320-01-36 |  | AS REPORTED, APPROXIMATELY 3 MONTHS POST OP THE INITIAL RIGHT TSA, THIS 60 Y/O FEMALE PATIENT WAS REVISED TO AN HRP FOR PROXIMAL BONE LOSS. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WERE DISPOSED OF BY THE HOSPITAL, WILL NOT BE RETURNING. | PENDING EVALUATION. CONCOMITANT MEDICAL PRODUCTS: 320-36-00, (B)(4) - 36MM HUMERAL LINER +0 UNCONSTRAINED; 320-10-00, (B)(4) - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0; 320-20-00, (B)(4) - EQ REVERSE TORQUE DEFINING SCREW KIT; 320-15-05, (B)(4) - EQ REV LOCKING SCREW. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00377 | 2022-04-06 | 2022-03-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | VANTAGE | LP TALAR IMPACTOR INSERT | EXACTECH, INC. | 351-07-04 | 351-07-04 |  | AS REPORTED, DURING IMPACTION OF THEN TALAR IMPLANT, THE TALAR INSERTER BROKE. NO PARTS OR PIECES FELL INTO THE PATIENTS WOUND SITE. NO SURGICAL DELAY PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICE TO BE RETURNED. | PENDING EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00379 | 2022-04-06 | 2022-03-09 | I, n, j, u, r, y | Manufacturer report | TRULIANT TIB IMP PSC INSERT SZ 4, 12MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT TIB IMP PSC INSERT SZ 4, 12MM | 02-022-44-4012 |  | AS REPORTED, PATIENT SAID THAT SINCE HIS SURGERY HIS KNEE HAS BEEN SWOLLEN AND HE HAS BEEN IN PAIN. HE SAID THAT HIS SURGEON SAID HE WILL NEED A REVISION BECAUSE HE IS ¿HIGHLY ALLERGIC TO NICKEL¿. NO ADDITIONAL INFORMATION AVAILABLE. | CONCOMITANT MEDICAL PRODUCTS: TRULIANT PS CEM FEM PS CEM LEFT SZ 4 (CAT# 02-020-11-0240). TRULIANT TIB FIT TRAY CEM SZ 4F / 4T (CAT# 02-022-45-4040). THREE PEG PATELLA 35MM (CAT# 200-02-35). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10038 | 2022-04-07 | 2019-10-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | DRILL GUIDE HANDLE | EXACTECH, INC. |  | 315-07-20 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE THREADED TIP BROKE OFF. THE AGENT WAS PRESENT AT SURGERY. THERE WAS NO DELAY TO THE CASE AND NO CLINICAL CONSEQUENCE TO THE PATIENT. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE RECEIVED FOR ANALYSIS. ANALYSIS COMPLETED ON (B)(6) 2020 BY NF. VISUAL EVALUATION CONDITION/FINDING: THE BROKEN THREADED TIP OF THIS DRILL GUIDE HANDLE APPEARS CONSISTENT WITH APPLYING AN OFF-AXIS TORQUE TO THE HANDLE DURING USE. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE AND CONSISTENT WITH YEARS OF USE. DESIGN-RELATED ISSUES: THE DESIGN OF THE DRILL GUIDE HANDLE HAS BEEN IN THE FIELD SINCE 2004. EXACTECH IS AWARE OF 5 SIMILAR COMPLAINT REPORTS INVOLVING FRACTURE OF THE THREADED TIP OF THE DRILL GUIDE HANDLE SINCE 2008. BECAUSE THE DRILL GUIDE HANDLE IS USED IN BOTH ANATOMIC AND REVERSE TOTAL SHOULDER SURGERIES, SALES DATA FOR HUMERAL STEMS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 60 UNITS, WHICH HAVE BEEN IN THE FIELD SINCE 2017. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING-RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) AND RISK ASSESSMENT AND CONTROLS REPORT (RACR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR AND RACR FOR THIS DRILL GUIDE HANDLE, (B)(4) - AND (B)(4) -, WERE REVIEWED. THE RISK IS CAPTURED, HOWEVER THE RACR IS BEING UPDATED UNDER DCRSP-19-541 TO CAPTURE AN APPROPRIATE HARM FOR THIS HAZARDOUS SITUATION. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", THE RISK IS CAPTURED IN THE RMR AND RACR, AND THE HARM WILL BE UPDATED. MOST LIKELY CAUSE: THE BROKEN INSTRUMENT REPORTED IN EXPERIENCE (B)(4) WAS LIKELY THE RESULT OF APPLYING AN OFF-AXIS TORQUE TO THE HANDLE DURING USE WHICH ALLOWED THE THREADS TO FRACTURE. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE WAS NO REPORTED PATIENT ADVERSE EVENT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF APPLYING AN OFF-AXIS TORQUE TO THE HANDLE DURING USE WHICH ALLOWED THE THREADS TO FRACTURE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10061 | 2022-04-07 | 2018-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRULIANT TIBIAL TAMP GUIDE | EXACTECH, INC. |  | 02-029-22-2000 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGICAL PROCEDURE THE SURGEON WAS IMPACTING THE TIBIAL TAMP GUIDE WHEN THE PUNCH RELEASE BUTTON MOVED OUT. IT DID NOT COMPLETELY FALL OUT UNTIL IT WAS PLACED ON THE BACK TABLE. THE STAFF HAD TO OPEN A 3ND TIBIAL PREP TRAY. THERE WAS NO SURGICAL DELAY WITH NO ADVERSE EFFECT TO THE PATIENT. THE PATIENT WAS DISCHARGED FROM THE OR IN STABLE CONDITION. THE AGENT WAS PRESENT DURING THE SURGERY. THE AGENCY IS INACTIVE, THERE IS NO FURTHER INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. MFG DATE: 11/20/2017. ENG EVAL COMPLETED ON 09/24/2018. PER INVESTIGATION CONDUCTED UNDER CAPA (B)(4), 9 COMPLAINTS INCLUDING 10 DEVICES WERE EVALUATED AND THE REPORTED CONDITION WAS CONFIRMED. INVESTIGATION CONFIRMED THAT ALL PARTS WERE MANUFACTURED TO SPECIFICATION. A REVIEW OF THE SPECIFICATIONS ON THE ENGINEERING PRINT REVEALED THAT THE OVERLAP BETWEEN THE SNAP RING'S OD (.720+.000/-.004) AND THE ID OF THE HOUSING LIP (.705+/-.003) IS ON THE LOW SIDE. THE ROOT CAUSE WAS DETERMINED TO BE THAT THE DESIGN OF THE SNAP RING THAT HOLDS THE BUTTON IN PLACE WAS NOT ROBUST. THE AMOUNT OF THE OVERLAP BETWEEN THE SNAP RING AND THE LIP OF THE HOUSING WAS ON THE LOW SIDE, WHICH INCREASES THE RISK OF DISENGAGEMENT. AN UPDATE TO THE DESIGN TO PREVENT RECURRENCE IS PLANNED. FROM THE HEALTH RISK INDEX (HRI) CONDUCTED FOR CRC-2018-08-31-02: THE BUTTON COMPONENT FELL OFF THE TRULIANT TIBIAL TAMP GUIDE DURING VARIOUS OCCASIONS (USAGE, CLEANING, TRANSPORTATION, ETC.). EXACTECH HAS RECEIVED 9 INCIDENT REPORTS AFTER MORE THAN 3000 CASES (OCCURRENCE RATE ~3%). THERE IS NO ANTICIPATED HARM TO THE PATIENT DUE TO THE DELAY IN SURGERY. PER THE TRILLIANT PRIMARY KNEE SYSTEMS OPERATIVE TECHNIQUE (LIT 712-35-31, REV A), THE TIBIAL TAMP GUIDE IS USED FOR TAMPING OF THE TIBIAL BASEPLATE TRIAL DURING KNEE IMPLANT SURGERY VIA IMPACTION AND RETRO-IMPACTION. THE DEVICE IFU (700-096-0160) WARNS THAT ONLY QUALIFIED SURGEONS KNOWLEDGEABLE IN ANATOMY, BIOMECHANICS, AND RECONSTRUCTIVE SURGERY SHOULD USE THESE DEVICES. THE SURGEON MUST BE FULLY KNOWLEDGEABLE ABOUT ALL ASPECTS OF THE TRULIANT FEMORAL COMPONENTS, TIBIAL INSERTS AND TIBIAL TRAYS SURGICAL TECHNIQUE AND USE THESE IMPLANTS IN ACCORDANCE WITH THE RESPECTIVE INDICATIONS AND CONTRAINDICATIONS SUMMARIZED IN THIS DOCUMENT. IN ADDITION, THE SURGEON SHOULD BE FULLY KNOWLEDGEABLE ABOUT THE COMPATIBILITY BETWEEN SYSTEM COMPONENTS AND USE ACCORDINGLY. FINALLY, THE SURGEON MUST BE TRAINED ON THE PROPER OPERATIVE TECHNIQUE INCLUDING THE PROPER USE OF SYSTEM INSTRUMENTATION. IT IS ESSENTIAL THAT THE PROPER SURGICAL TECHNIQUE BE USED WHEN SECURING LOCKING COMPONENTS. AGGRESSIVE IMPACTION TECHNIQUES AND FORCEFUL ASSEMBLY METHODS MAY DAMAGE COMPONENTS AND LEAD TO DEVICE FAILURE. REPEATED ASSEMBLY AND DISASSEMBLY OF MODULAR COMPONENTS COULD COMPROMISE THE STRENGTH OF LOCKING MECHANISMS. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. THE TIBIAL TAMP GUIDE BUTTON FELL OUT DURING A SURGICAL PROCEDURE. THERE WAS NO ADVERSE EVENT TO THE PATIENT. INVESTIGATION CONDUCTED UNDER CAPA (B)(4); THE ROOT CAUSE WAS DETERMINED TO BE THAT THE DESIGN OF THE SNAP RING THAT HOLDS THE BUTTON IN PLACE WAS NOT ROBUST. THE AMOUNT OF THE OVERLAP BETWEEN THE SNAP RING AND THE LIP OF THE HOUSING WAS ON THE LOW SIDE, WHICH INCREASES THE RISK OF DISENGAGEMENT. CORRECTIVE ACTIONS ARE ONGOING. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10044 | 2022-04-07 | 2017-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | CSA BROACH HANDLE | EXACTECH, INC. |  | CSK9064 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SPRING FELL OUT OF THE HUMERAL BROACH HANDLE RELEASE MECHANISM. THE BROACH HANDLE COULD NOT BE DISASSOCIATED FROM BROACH. THE AGENT WAS PRESENT AT SURGERY. THERE WAS A DELAY IN SURGERY WITH NO ADVERSE EVENT TO PATIENT. NO PATIENT DEMOGRAPHIC INFORMATION WAS NOT PROVIDED. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE NOT AVAILABLE FOR ANALYSIS. ENG EVAL COMPLETED BY CY. THE DEVICE WAS NOT AVAILABLE FOR EVALUATION AT THE TIME OF THIS INVESTIGATION. IF THE DEVICE IS RECEIVED IN. THE FUTURE, THE FILE WILL BE REOPENED. ACCORDING TO THE EXPERIENCE REPORT, THE SPRING FELL OUT OF THE. HUMERAL BROACH HANDLE RELEASE MECHANISM. BROACH HANDLE COULD NOT BE DISASSOCIATED FROM BROACH. DESIGN-RELATED ISSUES: THE DESIGN OF THE SPECIAL CSA BROACH HANDLE HAS BEEN IN THE FIELD SINCE 2017. THE FAILURE OF THIS INSTRUMENT APPEARS TO BE DIRECTLY RELATED TO THE FEATURE THAT MADE THIS A SPECIAL INSTRUMENT AND DIFFERENT FROM THE PARENT/PREDICATE INSTRUMENT. (B)(4) WAS INITIATED TO ENSURE THAT SPECIAL INSTRUMENTS ARE DESIGNED AND MANUFACTURED AFTER SUFFICIENT DESIGN CONTROL ACTIVITIES ARE COMPLETED AND DOCUMENTED. THE SPECIAL INSTRUMENT SOP IS BEING UPDATED AS PART OF THE CAPA TO IMPROVE THE CONNECTION BETWEEN INPUTS, PREDICATE DEVICES, RISK ASSESSMENT, AND POST-MARKET FEEDBACK. MFG-RELATED ISSUES: THE MANUFACTURING LOT INFORMATION FOR THE SPECIAL CSA BROACH HANDLE PIECE WAS REVIEWED. THIS BREAKAGE APPEARS TO BE RELATED TO THE DESIGN OF THE SPECIAL INSTRUMENT. THE INSTRUMENT WAS REWORKED AND IT WAS CONFIRMED THAT IT WAS REPAIRED AND OPERATED IN ACCORDANCE WITH THE REPAIR REQUIREMENTS BEFORE RETURNING TO THE SURGEON. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISKS WERE INCLUDED; THE RISK IS INCLUDED. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. THERE WAS NO PATIENT ADVERSE EVENT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN SPECIAL CSA BROACH HANDLE REPORTED WAS LIKELY THE RESULT OF THE DESIGN OF THE INSTRUMENT. CAPA 2018-33 HAS BEEN INITIATED TO ENSURE THAT SPECIAL INSTRUMENTS ARE DESIGNED AND MANUFACTURED AFTER SUFFICIENT DESIGN CONTROL ACTIVITIES ARE COMPLETED AND DOCUMENTED. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10037 | 2022-04-07 | 2017-08-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRULIANT TIBIAL INSERT TRIAL PS TOPPER | EXACTECH, INC. |  | 02-029-25-3550 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE TIBIAL INSERT TRIALS. THE BAL-SEAL SPRINGS WERE MISSING FROM THE TIBIAL INSERT TRIALS, IT ISN'T CLEAR WHEN THE DISENGAGEMENT HAPPENED. TWO X-RAYS WERE CHECKED, AND THE BAL-SPRINGS WERE NOT LEFT IN THE PATIENT. THE PATIENT HAD AN UNEVENTFUL RECOVERY. THERE WAS NO DELAY TO SURGERY. X-RAYS WERE SATISFACTORY WITH NO EVIDENCE OF ANY FOREIGN BODY. THE PATIENT WAS SEEN AT 6 WEEKS POST-OP AND WAS NOTED TO BE DOING WELL. NO ADVERSE PATIENT OUTCOME. PHOTOS WERE PROVIDED. PATIENT DEMOGRAPHIC INFORMATION WAS NOT PROVIDED. THE DEVICES WERE RECEIVED. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED BY THE CONTACTS ABOUT THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICES WERE AVAILABLE FOR ANALYSIS. ENGINEERING ANALYSIS COMPLETED BY AD ON 6/5/2018. DESIGN-RELATED ISSUES: OVER 560 CASES HAVE BEEN COMPLETED (AS OF NOVEMBER 20 I 7) WITH NO OTHER REPORTED INSTANCES. THEREFORE, THIS DOES NOT APPEAR TO BE RELATED TO THE INSTRUMENT DESIGN. MFG-RELATED ISSUES: OVER 560 CASES HAVE BEEN COMPLETED (AS OF NOVEMBER 2017) WITH NO OTHER REPORTED INSTANCES. THEREFORE, THIS DOES NOT APPEAR TO BE RELATED TO THE MANUFACTURING OF THE INSTRUMENT. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS INSTRUMENT, 750- 2014-006-RMR- INSTRUMENTS REV B, WAS REVIEWED. THE RISK IS CAPTURED IN 35. D. 9 B. THE BAL-SPRING DISENGAGEMENT REPORTED IN EXPERIENCE C2017-548 WAS LIKELY THE RESULT OF USING A TOOL TO REMOVE THE SPRING, POSSIBLY DURING CLEANING. THERE IS NO EVIDENCE THAT THE DISENGAGEMENT OCCURRED DURING NORMAL USE IN SURGERY. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. EXACTECH IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO ADVERSE EVENT REPORTED. THE DISENGAGEMENT OF THE BAL-SPRING WAS NOT KNOWN AT THE TIME OF SURGERY, X-RAYS WERE TAKEN TO ENSURE NO FOREIGN OBJECTS WERE IN THE PATIENT. AN INVESTIGATION WAS CONDUCTED: THE BAL-SPRING DISENGAGEMENT REPORTED WAS LIKELY THE RESULT OF USING A TOOL TO REMOVE THE SPRING, POSSIBLY DURING CLEANING. THERE IS NO EVIDENCE THAT THE DISENGAGEMENT OCCURRED DURING NORMAL USE IN SURGERY. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10053 | 2022-04-07 | 2018-02-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | COMBO TRAY ADJUSTABLE TIBIAL TAMP HEAD, NET SHAPE | EXACTECH, INC. |  | 213-75-01 |  | IT WAS REPORTED FROM THE US THAT WHEN THE SURGEON WAS PREPARING THE TIBIA, THE TAMP MALFUNCTIONED, LEADING TO IT GETTING STUCK INSIDE THE BONE. AFTER REMOVING IT WITH SOME PLIERS, THE TAMP PIECE WAS DAMAGED. THE AGENT WAS PRESENT AT SURGERY. THERE WAS NO DELAY TO THE SURGERY AND NO CLINICAL CONSEQUENCE TO THE PATIENT. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE WAS NOT RECEIVED FOR ANALYSIS. ENG EVAL COMPLETED BY KN ON 03/08/2019. (B)(4) BASED ON THE INVESTIGATION, IT WAS CONCLUDED THAT THE ROOT CAUSE OF THE TIBIAL TAMP HEAD FRACTURES WAS A COMBINATION OF THE DESIGN AND A USER MISUSE ISSUE. REVISION A OF THE TIBIAL TAMP HEAD HAS A SHARP CORNER PRESENT THAT COULD CREATE A STRESS CONCENTRATION THAT CONTRIBUTES TO THE BREAKAGE OF THE DEVICE. THE SURGICAL TECHNIQUE INSTRUCTS THAT USE OF THE EXTRACTION LEVER AND MAULDIN MULTI-TOOL SHOULD BE USED TO EXTRACT THE TIBIAL TAMP HEAD FROM THE BONE. THE SURGICAL TECHNIQUE DOES NOT SPECIFICALLY STATE THAT THE SURGEON SHOULD NOT HIT THE TAMP HANDLE/GUIDE BACKWARDS (RETRO-GRADE) TO REMOVE THE TAMP ASSEMBLY FROM THE BONE. AS A CORRECTION, EXACTECH HAS UPDATED THE SURGICAL TECHNIQUE TO CLARIFY THE PROPER TECHNIQUE AND INSTRUMENTATION TO REMOVE THE TAMP ASSEMBLY FROM THE BONE, I.E., TO INCLUDE A CAUTION STATEMENT ABOUT THE POTENTIAL FOR INSTRUMENT BREAKAGE IF THE TAMP HANDLE/GUIDE IS MISUSED BY IMPACTING IN RETRO-GRADE. A FIELD ACTION WAS SENT TO ALL CONSIGNEES CONCERNING THE UPDATE TO THE SURGICAL TECHNIQUE. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THE SURGICAL TECHNIQUE INSTRUCTS THAT USE OF THE EXTRACTION LEVER AND MAULDIN MULTI-TOOL SHOULD BE USED TO EXTRACT THE TIBIAL TAMP HEAD FROM THE BONE. THE SURGICAL TECHNIQUE DOES NOT SPECIFICALLY STATE THAT THE SURGEON SHOULD NOT HIT THE TAMP HANDLE/GUIDE BACKWARDS (RETRO-GRADE) TO REMOVE THE TAMP ASSEMBLY FROM THE BONE. AS A CORRECTION, EXACTECH HAS UPDATED THE SURGICAL TECHNIQUE TO CLARIFY THE PROPER TECHNIQUE AND INSTRUMENTATION TO REMOVE THE TAMP ASSEMBLY FROM THE BONE, I.E., TO INCLUDE A CAUTION STATEMENT ABOUT THE POTENTIAL FOR INSTRUMENT BREAKAGE IF THE TAMP HANDLE/GUIDE IS MISUSED BY IMPACTING IN RETRO-GRADE. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. THERE WAS NO PATIENT ADVERSE EVENT THE TAMP WAS REMOVED WITH PLIERS. AN INVESTIGATION WAS CONDUCTED; BASED ON THE INVESTIGATION, IT WAS CONCLUDED THAT THE ROOT CAUSE OF THE TIBIAL TAMP HEAD FRACTURES WAS A COMBINATION OF THE DESIGN AND A USER MISUSE ISSUE. AS A CORRECTION, EXACTECH HAS UPDATED THE SURGICAL TECHNIQUE TO CLARIFY THE PROPER TECHNIQUE AND INSTRUMENTATION TO REMOVE THE TAMP ASSEMBLY FROM THE BONE, I.E., TO INCLUDE A CAUTION STATEMENT ABOUT THE POTENTIAL FOR INSTRUMENT BREAKAGE IF THE TAMP HANDLE/GUIDE IS MISUSED BY IMPACTING IN RETRO-GRADE. A FIELD ACTION WAS SENT TO ALL CONSIGNEES CONCERNING THE UPDATE TO THE SURGICAL TECHNIQUE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10058 | 2022-04-07 |  | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRULIANT TIBIAL TAMP GUIDE | EXACTECH, INC. |  | 02-029-22-2000 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE LOCKING BUTTON ON THE TIBIAL TAMP GUIDE DISENGAGED FROM THE HANDLE. THE AGENT WAS ABLE TO PLACE THE BUTTON BACK IN. BUT AFTER A SECOND TIME IT WOULD NOT STAY ENGAGED WHEN THE TOP OF THE IMPACTOR WAS HIT. THIS ISSUE DID NOT HAVE ANY EFFECT ON THE CASE OF THE OUTCOME OF THE PATIENT. INACTIVE AGENCY, NO ADDITIONAL INFORMATION IS AVAILABLE. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. MFG DATE: 01/09/2018. ENG EVAL COMPLETED ON 90/24/2018. PER INVESTIGATION CONDUCTED UNDER CAPA2018-35, 9 COMPLAINTS INCLUDING 10 DEVICES WERE EVALUATED AND THE REPORTED CONDITION WAS CONFIRMED. INVESTIGATION CONFIRMED THAT ALL PARTS WERE MANUFACTURED TO SPECIFICATION. A REVIEW OF THE SPECIFICATIONS ON THE ENGINEERING PRINT REVEALED THAT THE OVERLAP BETWEEN THE SNAP RING'S OD (.720+.000/-.004) AND THE ID OF THE HOUSING LIP (.705+/-.003) IS ON THE LOW SIDE. AN UPDATE TO THE DESIGN TO PREVENT RECURRENCE IS PLANNED. FROM THE HEALTH RISK INDEX (HRI) CONDUCTED FOR CRC-2018-08-31-02: THE BUTTON COMPONENT FELL OFF THE TRULIANT TIBIAL TAMP GUIDE DURING VARIOUS OCCASIONS (USAGE, CLEANING, TRANSPORTATION, ETC.). EXACTECH HAS RECEIVED 9 INCIDENT REPORTS AFTER MORE THAN 3000 CASES (OCCURRENCE RATE ~3%). THERE IS NO ANTICIPATED HARM TO THE PATIENT DUE TO THE DELAY IN SURGERY. THE ROOT CAUSE WAS DETERMINED TO BE THAT THE DESIGN OF THE SNAP RING THAT HOLDS THE BUTTON IN PLACE WAS NOT ROBUST. THE AMOUNT OF THE OVERLAP BETWEEN THE SNAP RING AND THE LIP OF THE HOUSING WAS ON THE LOW SIDE, WHICH INCREASES THE RISK OF DISENGAGEMENT. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. AN INVESTIGATION WAS CONDUCTED UNDER CAPA2018-35; THE ROOT CAUSE WAS DETERMINED TO BE THAT THE DESIGN OF THE SNAP RING THAT HOLDS THE BUTTON IN PLACE WAS NOT ROBUST. THE AMOUNT OF THE OVERLAP BETWEEN THE SNAP RING AND THE LIP OF THE HOUSING WAS ON THE LOW SIDE, WHICH INCREASES THE RISK OF DISENGAGEMENT. CORRECTIVE ACTIONS ARE ONGOING. EVENT DATE NOT REPORTED. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10067 | 2022-04-07 | 2019-11-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | SPEC DUAL STRIKE PLATE BROACH HANDLE | EXACTECH, INC. |  | SK-025375 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SPECIAL BROACH HANDLE INSTRUMENT BROKE. IT WAS CONFIRMED THAT ALL PIECES WERE LOCATED OUTSIDE OF THE PATIENT. THERE WAS NO REPORTED PATIENT ADVERSE EVENT. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE RECEIVED FOR ANALYSIS. ENG EVAL COMPLETED ON 2020.03.05 BY AH. (B)(4) FINDINGS: IT WAS FOUND THAT THE SPECIAL DUAL STRIKE PLATE BROACH HANDLE (CN: SK-025375, LOT: DSR219060) CAN DISASSEMBLE DURING SURGERY. IT WAS REPORTED THAT ALL PIECES OF THE BROKEN DEVICE WERE RECOVERED AND NOT LEFT WITHIN THE PATIENT. AS DOCUMENTED IN CRC2019-12-05-04, THE DOWEL PIN RETAINING THE LEVER TO THE ACTUATOR CAN DISASSEMBLE CAUSING THE INTERNAL COMPONENTS TO DISASSEMBLE FROM THE BROACH HANDLE BODY. UPON INSPECTION OF THE RETURNED DEVICE, THE DISASSEMBLY OF THE DOWEL PIN WAS DUE TO A WELD FAILURE LIKELY AS A RESULT OF IMPACTION FORCES. ADDITIONALLY, THE SPECIAL BROACH HANDLE WAS FOUND TO BE OUT OF SPECIFICATION; THE DEPTH OF THE BODY CAVITY ON THE HANDLE IS TOO DEEP AND THE RETAINING PIN HAS NOT BEEN WELDED AS SPECIFIED ON THE ENGINEERING PRINT. REFER TO TM-2019- 0629 AND SK-025375 REV 3. THIS IS CONSISTENT WITH THE FINDINGS OF THE CRC AND CASE-2020- 00002002. CORRECTIVE ACTION TAKEN: IT WAS DETERMINED BY THE CORRECTION AND REMOVAL COMMITTEE THAT THE REMAINING DEVICES WILL BE REMOVED FROM THE FIELD. CAPA2019-80 WILL BE OPENED TO ADDRESS THE ISSUE. THIS COMPLAINT IS PRIOR TO IMPLEMENTATION OF CORRECTIVE ACTIONS. NO ADDITIONAL ACTION REQUIRED. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: AS DOCUMENTED IN CRC2019-12-05-04, THE BROKEN INSTRUMENT REPORTED IN CASE-2019-00001764 WAS LIKELY THE RESULT OF A COMBINATION OF A WELD FAILURE AS A RESULT OF IMPACTIONS FORCES, THE DOWEL PIN NOT BEING RETAINED WITHIN THE BODY OF THE BROACH HANDLE, AND THE OUT OF SPECIFICATION CONDITIONS WHICH ALLOWED FOR THE SUBSEQUENT DISASSEMBLY OF THE DEVICE. FORM NUMBER 701-104-752 REV: - EFFECTIVE DATE: 10/15/18 CONTROLLED DOCUMENT COMPANY CONFIDENTIAL PAGE 2 OF 3 IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. THERE WAS NO REPORTED PATIENT ADVERSE EVENT. AS DOCUMENTED IN CRC2019-12-05-04, THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF A COMBINATION OF A WELD FAILURE AS A RESULT OF IMPACTIONS FORCES, THE DOWEL PIN NOT BEING RETAINED WITHIN THE BODY OF THE BROACH HANDLE, AND THE OUT OF SPECIFICATION CONDITIONS WHICH ALLOWED FOR THE SUBSEQUENT DISASSEMBLY OF THE DEVICE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10035 | 2022-04-07 | 2017-08-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRULIANT TIBIAL INSERT TRIAL PS TOPPER | EXACTECH, INC. |  | 02-029-25-3535 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE TIBIAL INSERT TRIALS. THE BAL-SEAL SPRINGS WERE MISSING FROM THE TIBIAL INSERT TRIALS, IT ISN'T CLEAR WHEN THE DISENGAGEMENT HAPPENED. TWO X-RAYS WERE CHECKED, AND THE BAL-SPRINGS WERE NOT LEFT IN THE PATIENT. THE PATIENT HAD AN UNEVENTFUL RECOVERY. THERE WAS NO DELAY TO SURGERY. X-RAYS WERE SATISFACTORY WITH NO EVIDENCE OF ANY FOREIGN BODY. THE PATIENT WAS SEEN AT 6 WEEKS POST-OP AND WAS NOTED TO BE DOING WELL. NO ADVERSE PATIENT OUTCOME. PHOTOS WERE PROVIDED. PATIENT DEMOGRAPHIC INFORMATION WAS NOT PROVIDED. THE DEVICES WERE RECEIVED. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED BY THE CONTACTS ABOUT THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICES WERE AVAILABLE FOR ANALYSIS. ENGINEERING ANALYSIS COMPLETED BY AD ON 6/5/2018. DESIGN-RELATED ISSUES: OVER 560 CASES HAVE BEEN COMPLETED (AS OF NOVEMBER 20 I 7) WITH NO OTHER REPORTED INSTANCES. THEREFORE, THIS DOES NOT APPEAR TO BE RELATED TO THE INSTRUMENT DESIGN. MFG-RELATED ISSUES: OVER 560 CASES HAVE BEEN COMPLETED (AS OF NOVEMBER 2017) WITH NO OTHER REPORTED INSTANCES. THEREFORE, THIS DOES NOT APPEAR TO BE RELATED TO THE MANUFACTURING OF THE INSTRUMENT. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS INSTRUMENT, 750- 2014-006-RMR- INSTRUMENTS REV B, WAS REVIEWED. THE RISK IS CAPTURED IN 35. D. 9 B. THE BAL-SPRING DISENGAGEMENT REPORTED IN EXPERIENCE C2017-548 WAS LIKELY THE RESULT OF USING A TOOL TO REMOVE THE SPRING, POSSIBLY DURING CLEANING. THERE IS NO EVIDENCE THAT THE DISENGAGEMENT OCCURRED DURING NORMAL USE IN SURGERY. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. EXACTECH IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO ADVERSE EVENT REPORTED. THE DISENGAGEMENT OF THE BAL-SPRING WAS NOT KNOWN AT THE TIME OF SURGERY, X-RAYS WERE TAKEN TO ENSURE NO FOREIGN OBJECTS WERE IN THE PATIENT. AN INVESTIGATION WAS CONDUCTED: THE BAL-SPRING DISENGAGEMENT REPORTED WAS LIKELY THE RESULT OF USING A TOOL TO REMOVE THE SPRING, POSSIBLY DURING CLEANING. THERE IS NO EVIDENCE THAT THE DISENGAGEMENT OCCURRED DURING NORMAL USE IN SURGERY. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10054 | 2022-04-07 | 2018-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | REAMER, ACETABULAR, 52MM | EXACTECH, INC. |  | 123-00-52 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC INITIAL IMPLANT SURGICAL PROCEDURE THE ACETABULAR REAMER BROKE. THE AGENT WAS PRESENT AT THE TIME OF SURGERY. THE WELDING BROKE AND THE CROSS BAR SNAPPED OFF THE CUP REAMER. THE REAMER AND CROSS BAR WERE COMPLETELY REMOVED, NO PIECES WERE LEFT IN PATIENT. THERE WAS NO SURGICAL DELAY. THE AGENCY IS INACTIVE AND NO FURTHER INFORMATION CAN BE OBTAINED ON THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE ACTUAL DEVICE WAS RETURNED FOR ANALYSIS. ENG EVAL COMPLETED ON 10/03/2018 BY MB. THE BROKEN WELDS WERE PHOTOGRAPHED AND APPEAR TO BE CONSISTENT WITH EXCESS TORQUE BEING APPLIED EVENTUALLY LEADING THE HANDLE ATTACHED OF THE REAMER TO BREAK OFF FROM THE BLADES. THIS EVENT IS VERY LOW ON THE FREQUENCY OF OCCURRENCE RANKING SCALE; THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE A DESIGN RELATED ISSUE. THIS MANUFACTURING LOT NUMBER HAS BEEN IN THE FIELD FOR AT LEAST 10 YEARS AND THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. THE ACETABULAR REAMER IS USED IN TOTAL HIP SURGERIES, SALES DATA FOR THE ACETABULAR SHELL WAS USED TO CALCULATE APPROXIMATE COMPLAINT OCCURRENCE RATE OF <1.5%. EXACTECH IS NOT AWARE OF ANY COMPLAINTS INVOLVING COMPONENTS FROM THIS MANUFACTURING LOT. THE 4-DIGIT MANUFACTURING NUMBER ENGRAVED ON THE INSTRUMENTS IDENTIFIES THIS PART WAS MADE PRIOR TO 2009. THE BROKEN WELD REPORTED IS LIKELY THE RESULT OF EXCESS TORQUE BEING APPLIED WHEN USING THE REAMER HANDLE EVENTUALLY BREAKING THE WELD. IFU 700-096-181: INSTRUMENT INSPECTION VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. ACCORDING TO INFORMATION CONTAINED IN THE OPERATIVE TECHNICAL GUIDE (OP-TECH-GUIDE #711-02-32) THE HANDLE IS ASSEMBLED FIRST BEFORE USE, INDICATED AT THE POINT OF THE CROSS BAR WELD BREAK, THE CROSS BAR WOULD HAVE REMAINED ATTACHED TO THE REAMER HANDLE. ASSEMBLE THE ACETABULAR REAMER HANDLE, STRAIGHT, INTO THE POWER SOURCE. ASSEMBLE A HEMISPHERICAL ACETABULAR REAMER 4 TO 6MM SMALLER THAN THE SIZE TEMPLATED ONTO THE REAMER HANDLE. THE ACETABULAR REAMER SELECTED FOR INITIAL REAMING IS TYPICALLY 4 TO 6MM SMALLER THAN THE SIZE TEMPLATED. ACETABULAR REAMERS ARE DESIGNED TO MATCH THE HEMISPHERICAL SHAPE OF THE CORRESPONDING IMPLANT. REAMING DEPTH IS TO THE EDGE OF THE OPENING IN THE REAMER BASKET. AN INITIAL REAMING MAY BE DIRECTED MORE MEDIALLY, BUT SUBSEQUENT REAMING SHOULD BE DONE IN 45 DEGREES OF ABDUCTION AND 20 DEGREES OF FORWARD FLEXION (ANTEVERSION). IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. NO PIECES WERE LEFT IN THE PATIENT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN WELD REPORTED IS LIKELY THE RESULT OF EXCESS TORQUE BEING APPLIED WHEN USING THE REAMER HANDLE EVENTUALLY BREAKING THE WELD. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10056 | 2022-04-07 | 2018-03-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | REAMER HEAD, MODULAR PRIMARY, PILOT TIP, LARGE | EXACTECH, INC. |  | 315-25-14 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE REAMER BROKE WHILE REAMING THE GLENOID AFTER HITTING AGAINST RETRACTORS. THE CASE WENT WELL AND THERE WAS NO INJURY TO THE PATIENT. PATIENTS HEALTH WAS STABLE LEAVING THE OR. ISSUE WAS RESOLVED BY OPENING ANOTHER TRAY. THE SALES REPRESENTATIVE WAS PRESENT AT THE SURGERY. THERE WAS NO DELAY OF SURGERY AND THERE WAS NO PATIENT INJURY, ADVERSE EVENT, OR CLINICAL CONSEQUENCE TO THE PATIENT. MULTIPLE ATTEMPTS HAVE BEEN MADE FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED FOR ANALYSIS. COMPLETED ON: 05/22/2018 BY (B)(4). VISUAL EVALUATION CONDITION/FINDINGS: (CONDUCTED WITH A 7X OR 10X OPTICAL UNLESS OTHERWISE SPECIFIED) THE FRACTURES APPEAR CONSISTENT WITH BRITTLE FRACTURE LIKELY CAUSED BY COMING INTO CONTACT WITH RETRACTORS WHILE REAMING AS STATED IN THE EXPERIENCE REPORT. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE. DESIGN-RELATED ISSUES: THE DESIGN OF THE PRIMARY PILOT TIP REAMER HEAD HAS BEEN IN THE FIELD SINCE 2009. (B)(4). THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH HAS NOT RECEIVED ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT, WHICH HAS BEEN IN THE FIELD SINCE 2017. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED, AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS PRIMARY PILOT TIP REAMER HEAD, 750-2007-026-RMR REV H, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 17. NO CORRECTIVE ACTION IS REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE BROKEN REAMER HEAD REPORTED IN EXPERIENCE (B)(4) WAS LIKELY THE RESULT OF THE REAMER COMING INTO CONTACT WITH RETRACTORS WHILE REAMING. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES. THERE IS NO REPORT OF PATIENT ADVERSE EVENT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN REAMER HEAD REPORTED WAS LIKELY THE RESULT OF THE REAMER ENCOUNTERING RETRACTORS WHILE REAMING. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10066 | 2022-04-07 | 2019-10-31 | I, n, j, u, r, y | Manufacturer report | EXACTECH | 38 BOWTIE REAMER | EXACTECH, INC. |  | 321-55-38 |  | IT WAS REPORTED FROM THE US THAT THE "SURGEON PUT BOW TIE REAMER ON MODULAR HANDLE (ON DRILL NOT REAM). PUT IT OVER A K WIRE AND RAN IT DOWN TO THE GLENOID. STARTED TO REAM THE GLENOID AND THE PIECE THAT ATTACHES TO THE HANDLE SHEARED OFF. IT DID END UP NOTCHING THE GLENOID, BUT WE WERE ABLE TO FACILITATE THE BASEPLATE ON WITHOUT COMPLICATIONS." MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. NO DEVICE RECEIVED FOR ANALYSIS. ANALYSIS COMPLETED ON 04/10/2020 BY AH. DESIGN-RELATED ISSUES: EXACTECH IS AWARE OF 5 COMPLAINT REPORTS INVOLVING A FRACTURED REAMER SHAFT SINCE 2008. BECAUSE THIS DEVICE IS USED IN TOTAL REVERSE SHOULDER SURGERIES, SALES DATA FOR GLENOSPHERES WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. MFG-RELATED ISSUES: MANUFACTURING-RELATED INFORMATION COULD NOT BE REVIEWED BECAUSE NO INFORMATION REGARDING MANUFACTURING LOTS WAS PROVIDED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) AND THE RISK ASSESSMENT AND CONTROLS REPORT (RACR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR EQUINOXE NARROW REAMERS, RMR-00007 REV B AND RACR-00010 REV B, WERE REVIEWED. THE RISK IS CAPTURED IN RISK ID 5.3. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR AND RACR. THE BROKEN REAMER REPORTED IN (B)(4) WAS LIKELY THE RESULT OF APPLYING A FORCE GREATER THAN YIELD STRENGTH OF THE MATERIAL AND SUBSEQUENT STERILIZATION CYCLES WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FAILURE. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE COMPONENT WAS NOT AVAILABLE FOR EVALUATION. ADDITIONAL COMMENTS: ACCORDING TO THE EXPERIENCE REPORT, "SURGEON PUT BOW TIE REAMER ON MODULAR HANDLE (ON DRILL NOT REAM). PUT IT OVER A K WIRE AND RAN IT DOWN TO THE GLENOID. STARTED TO REAM THE GLENOID AND THE PIECE THAT ATTACHES TO THE HANDLE SHEARED OFF." IT WAS REPORTED THAT THIS HAS OCCURRED ON 3 SEPARATE OCCASIONS. THE OTHER TWO EVENTS ARE CAPTURE IN (B)(4) AND (B)(4). IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY DESIGN ISSUES. THE GLENOID WAS REPORTED TO BE NOTCHED DUE TO THIS EVENT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN REAMER REPORTED WAS LIKELY THE RESULT OF APPLYING A FORCE GREATER THAN YIELD STRENGTH OF THE MATERIAL AND SUBSEQUENT STERILIZATION CYCLES WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FAILURE. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE COMPONENT WAS NOT AVAILABLE FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10057 | 2022-04-07 | 2018-05-23 | I, n, j, u, r, y | Manufacturer report | EXACTECH | HUMERAL LINER IMPACTOR TIP, 38MM | EXACTECH, INC. |  | 321-07-38 |  | IT WAS REPORTED FROM OUS THAT DURING THE STEM TRIAL OF AN EQUINOXE SURGERY, THE TIP OF THE IMPACTOR WAS BROKEN INTO 2 PIECES WHEN IMPACTED. "THE DOCTOR STRUCK BALANCE, NOT STRONGLY." THE DOCTOR CONFIRMED THAT ALL BROKEN PIECES WERE ACCOUNTED FOR. THERE WAS NO DELAY TO SURGERY OR REPORTED PATIENT ADVERSE EVENT. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE COMPLAINT PRODUCT WAS RECEIVED FOR ANALYSIS. ENG EVAL COMPLETED ON 03/05/2019 BY (B)(4). VISUAL - THE DEVICE WAS RETURNED IN TWO PIECES. THE FRACTURE SURFACES APPEAR CONSISTENT WITH CRACK INITIATION AND PROPAGATION CAUSED BY A STRESS RISER. THE CIRCULAR WITNESS MARK SEEN IN THE CENTER OF THE DETACHED PIECE IS APPROXIMATELY 0.191 INCHES IN DIAMETER AND APPEARS CONSISTENT WITH ENCOUNTERING THE MATING IMPACTOR INSTRUMENT. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE AND CONSISTENT WITH USE REVIEW OF OTHER PREVIOUS COMPLAINTS - EXACTECH IS NOT AWARE OF RECEIVING ANY COMPLAINT REPORTS INVOLVING ANOTHER PART FROM THIS MANUFACTURING LOT OF (B)(4) PIECES, WHICH HAS BEEN IN THE FIELD SINCE 2018. EXACTECH COMPLAINT DATA FROM 2014 THROUGH 2018 INVOLVING THE REPORTED FAILURE FOR THIS FAMILY OF DEVICES WAS REVIEWED. DESIGN-RELATED: THE DESIGN OF THE HUMERAL LINER IMPACTOR TIP HAS BEEN IN THE FIELD SINCE 2007. EXACTECH IS AWARE OF 11 OTHER SIMILAR COMPLAINT REPORTS INVOLVING HUMERAL LINER IMPACTOR TIPS SINCE 2014. THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF RECEIVING ANY COMPLAINT REPORTS INVOLVING ANOTHER PART FROM THIS MANUFACTURING LOT OF (B)(4) PIECES, WHICH HAS BEEN IN THE FIELD SINCE 2018. EXACTECH COMPLAINT DATA FROM 2014 THROUGH 2018 INVOLVING THE REPORTED FAILURE FOR THIS FAMILY OF DEVICES WAS REVIEWED. THE INVESTIGATIONS FOR THOSE INCIDENCES HAVE NOT IDENTIFIED ANY EVIDENCE THAT A MANUFACTURING ISSUE CAUSED OR CONTRIBUTED TO THIS REPORTED ISSUE. THEREFORE, THIS DOES NOT APPEAR TO BE MANUFACTURING RELATED. CORRECTIVE ACTIONS ARE NOT RECOMMENDED, THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY ROOT CAUSE: THE BROKEN IMPACTOR TIP REPORTED IN EXPERIENCE (B)(4) WAS LIKELY THE RESULT OF A STRESS RISER INTRODUCED DURING IMPACTION, WHICH LED TO CRACK INITIATION, PROPAGATION AND ULTIMATE FRACTURE OF THE DEVICE. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING OR DESIGN ISSUES. THERE WAS NO REPORTED PATIENT ADVERSE EVENT. THE SURGEON CONFIRMED THAT ALL PIECES HAD BEEN ACCOUNTED FOR. AN INVESTIGATION WAS CONDUCTED; THE BROKEN IMPACTOR TIP REPORTED WAS LIKELY THE RESULT OF A STRESS RISER INTRODUCED DURING IMPACTION, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00385 | 2022-04-07 | 2022-03-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | LOCKING FEMORAL IMPACTOR 0-5 | ORTHOPEDIC MANUAL SURGICAL INSTRUMENT | EXACTECH, INC. | LOCKING FEMORAL IMPACTOR 0-5 | 213-64-01 |  | AS REPORTED, PART OF THE IMPACTOR BROKE OFF DURING THE IMPACTION OF THE FEMORAL COMPONENT. THE IMPLANT WAS ALREADY SEATED AND CAUSED NO DELAY TO SURGERY. THERE WAS NO PARTS OR PIECES FELL INTO THE PATIENT WOUND SITE. THE SURGICAL SITE WAS CONFIRMED CLEAN BY SURGEON AND THE PROCEDURE CARRIED ON WITH NO DELAY. PATIENT REMAINED IN STABLE CONDITION. SALES REP WAS UNABLE TO OBTAIN PHOTOS/X-RAYS. THE DEVICE WILL BE RETURN. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00387 | 2022-04-07 | 2021-02-24 | I, n, j, u, r, y | Manufacturer report | GXL POLYETHYLENE LINER 32MM DIAMETER HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | GXL POLYETHYLENE LINER 32MM DIAMETER HEAD | UNKNOWN |  | AS REPORTED BY THE LEGAL BRIEF, ON OR ABOUT 2014, PATIENT HAD INITIAL THA, HER HIP SYSTEM INCLUDED A 48 ANTERIOR GRIP CUP WITH A GXL POLYETHYLENE LINE OF A 32 TO ACCEPT A 32-MM DIAMETER HEAD. THE FEMORAL STEM INCLUDED A HA-COATED ELEMENT STEM WITH A 28+0 CERAMIC FEMORAL HEAD. APPROXIMATELY SEVEN YEARS POST INITIAL THA, PATIENT BEGAN TO EXPERIENCE SEVERE PAIN IN HER HIP, AND WAS SUBSEQUENTLY DIAGNOSED WITH FAILED POLYETHYLENE AND POLYETHYLENE WEAR ON THE DOME OF THE HIP SYSTEM. PATIENT WAS CAUSED TO UNDERGO A PAINFUL REVISION SURGERY ON OR ABOUT (B)(6) 2021. AS A DIRECT AND PROXIMATE RESULT, PATIENT WAS CAUSED TO SUSTAIN SERIOUS, SEVERE AND PAINFUL PERSONAL INJURIES, SUFFERING INTERNAL PHYSICAL PAIN, AND MENTAL ANGUISH; WAS COMPELLED TO SEEK MEDICAL CARE AND ATTENTION, INCURRING EXPENSES IN CONNECTION THEREWITH, IN AN ATTEMPT TO RECTIFY AND REMEDY HER CONDITION; PATIENT WAS CAUSED TO UNDERGO UNWARRANTED AND UNNECESSARY SURGICAL PROCEDURES; PATIENT WAS CAUSED TO EXPERIENCE SEVERE PHYSICAL AND PSYCHOLOGICAL PAIN, SUFFERING, REQUIRED SURGERY AND TREATMENT, INCURRED EXPENSES, SHORTENED LIFE-EXPECTANCY AND IN OTHER RESPECTS, WAS DAMAGED. | CONCOMITANT MEDICAL PRODUCTS: FEMORAL STEM INCLUDED A HA-COATED ELEMENT STEM WITH A 28+0 CERAMIC FEMORAL HEAD. ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10052 | 2022-04-07 | 2018-02-07 | I, n, j, u, r, y | Manufacturer report | EXACTECH | HUMERAL LINER IMPACTOR TIP, 38MM | EXACTECH, INC. |  | 321-07-38 |  | IT WAS REPORTED FROM OUS THAT DURING THE STEM TRAIL OF THE EQUINOXE SURGERY, THE TIP OF THE IMPACTOR WAS BROKEN INTO 2 PIECES WHEN THE SURGEON IMPACTED IT. "HE STRUCK A BALANCE AS USUAL, NOT STRONGLY." THE SURGEON CONFIRMED THE BROKEN PIECES DID NOT REMAIN IN THE PATIENT. THE DEVICES WERE DISCARDED BY THE HOSPITAL. THERE WAS NO DELAY TO THE SURGERY AND NO CLINICAL CONSEQUENCE TO THE PATIENT. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED FOR ANALYSIS. COMPLETED ON 06/26/2018 BY CY. VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR 10X OPTICAL UNLESS OTHERWISE SPECIFIED) THE FRACTURE SURFACE APPEAR CONSISTENT WITH FRACTURE INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE LIKELY DUE TO A STRESS RISER. THE DAMAGE TO THE CONTACT SURFACE APPEARS CONSISTENT WITH CONTACTING ANOTHER SURGICAL DEVICE DURING IMPACTION. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE. DESIGN-RELATED ISSUES: THE DESIGN OF THE HUMERAL LINER IMPACTOR TIP HAS BEEN IN THE FIELD SINCE 2007. EXACTECH HAS RECEIVED 11 SIMILAR COMPLAINT REPORTS INVOLVING HUMERAL LINER IMPACTOR TIPS SINCE 2014. THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH HAS NOT RECEIVED ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 60 PIECES, WHICH HAS BEEN IN THE FIELD SINCE 2017. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING-RELATED. EXACTECH COMPLAINT DATA FROM 2014 THROUGH 2018 INVOLVING THE REPORTED FAILURE FOR THIS FAMILY OF DEVICES WAS REVIEWED. THE INVESTIGATIONS FOR THOSE INCIDENCES HAVE NOT IDENTIFIED ANY EVIDENCE THAT A MANUFACTURING ISSUE CAUSED OR CONTRIBUTED TO THIS REPORTED ISSUE. THEREFORE, THIS DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS HUMERAL LINER IMPACTOR TIP, 750-2005-021-RMR-INSTRUMENT REV N, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 15. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE BROKEN DEVICE REPORTED IN EXPERIENCE C2018-125 WAS LIKELY THE RESULT OF A STRESS RISER INTRODUCED DURING IMPACTION, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE. IFU 700-096-181: INSTRUMENT INSPECTION, VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES. THERE WAS NO PATIENT ADVERSE EVENT. THE SURGEON CONFIRMED THE BROKEN PIECES DID NOT REMAIN IN THE PATIENT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF A STRESS RISER INTRODUCED DURING IMPACTION, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10064 | 2022-04-07 | 2017-05-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | GLENOID SKID | EXACTECH, INC. |  | 315-03-01 |  | IT WAS REPORTED THAT DURING AN ORTHOPEDIC SHOULDER SURGERY THE MODULAR CANNULATED TRI-DRIVE SPRING BROKE AND THE GLENOID SKID BROKE WHEN THE SURGEON CUT THROUGH THE ITEM AND APPLIED TOO MUCH LEVERAGE. THE SALES REPRESENTATIVE WAS PRESENT AT THE SURGERY AND REPORTED NO DEVICE FRAGMENTS FELL INTO THE SURGICAL SITE. THE SURGERY CONTINUED AND WAS COMPLETED WITHOUT ADDITIONAL ISSUES. THERE WAS NO DELAY OF SURGERY AND THERE WAS NO PATIENT INJURY, ADVERSE EVENT, OR CLINICAL CONSEQUENCE TO THE PATIENT. THE AGENCY/AGENT IS NO LONGER ACTIVE WITH EXACTECH. NO ADDITIONAL INFORMATION IS AVAILABLE FROM THE REPORTER RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE CANNULATED TRI-DRIVE WAS RETURNED TO EXACTECH FOR ANALYSIS. DESIGN-RELATED ISSUES: THIS DESIGN HAS BEEN IN THE FIELD SINCE 2009. EXACTECH HAS RECEIVED 38 SIMILAR COMPLAINT REPORTS INVOLVING 54 DEVICES SINCE ITS INTRODUCTION. BECAUSE THIS INSTRUMENT IS USED IN BOTH PRIMARY AND REVERSE TOTAL SHOULDER SURGERIES, THE SALES DATA FOR HUMERAL STEM COMPONENTS WERE USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4). THIS IS CONSIDERED VERY LOW ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH HAS RECEIVED 3 OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT CONTAINING 100 UNITS, WHICH HAS BEEN IN THE FIELD SINCE 2011. THE DEVICE HISTORY RECORD WAS REVIEWED. ALL INSTRUMENTS WERE ACCEPTED WITH CONFORMANCE TO THE DEVICE REQUIREMENTS. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING-RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS TRI-DRIVE MODULAR REAMER DRIVER, 750-2007-026-RMRINSTRUMENTS REV H, WAS REVIEWED. THE RISK IS NOT CAPTURED. THIS DOCUMENT WILL BE UPDATED TO INCLUDE THIS FAILURE. REFER TO DCRSP-18-340 FOR DETAILS. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS VERY LOW, AND THE RMR WILL BE UPDATED TO CAPTURE THIS RISK. THE DISASSEMBLED DRIVER REPORTED IN EXPERIENCE C2017-322 WAS LIKELY THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY AND ALLOWED THE TIP TO DISASSEMBLE. EXACTECH IFU 700-096-004 STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED INSTRUMENT MALFUNCTION IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO ANY PATIENT IMPACT. AN INVESTIGATION WAS CONDUCTED THAT THE MOST LIKELY ROOT CAUSE OF THE DEVICE BREAKAGE REPORTED IN THIS EVENT WAS THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY AND ALLOWED THE TIP TO DISASSEMBLE. THE AGENT/AGENCY IS NO LONGER ACTIVE WITH EXACTECH. NO ADDITIONAL INFORMATION IS AVAILABLE FROM THE CONTACTS RELATED TO THIS EVENT. NO ADDITIONAL ATTEMPTS WILL BE MADE FOR INFORMATION REGARDING THIS EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10045 | 2022-04-07 | 2017-05-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | DRIVER, MODULAR CANNULATED TRI-DRIVE | EXACTECH, INC. |  | 315-25-00 |  | IT WAS REPORTED THAT DURING AN ORTHOPEDIC SHOULDER SURGERY THE MODULAR CANNULATED TRI-DRIVE SPRING BROKE AND THE GLENOID SKID BROKE WHEN THE SURGEON CUT THROUGH THE ITEM AND APPLIED TOO MUCH LEVERAGE. THE SALES REPRESENTATIVE WAS PRESENT AT THE SURGERY AND REPORTED NO DEVICE FRAGMENTS FELL INTO THE SURGICAL SITE. THE SURGERY CONTINUED AND WAS COMPLETED WITHOUT ADDITIONAL ISSUES. THERE WAS NO DELAY OF SURGERY AND THERE WAS NO PATIENT INJURY, ADVERSE EVENT, OR CLINICAL CONSEQUENCE TO THE PATIENT. THE AGENCY/AGENT IS NO LONGER ACTIVE WITH EXACTECH. NO ADDITIONAL INFORMATION IS AVAILABLE FROM THE REPORTER RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE CANNULATED TRI-DRIVE WAS RETURNED TO EXACTECH FOR ANALYSIS. DESIGN-RELATED ISSUES: THIS DESIGN HAS BEEN IN THE FIELD SINCE 2009. (B)(4). THE DEVICE HISTORY RECORD WAS REVIEWED. ALL INSTRUMENTS WERE ACCEPTED WITH CONFORMANCE TO THE DEVICE REQUIREMENTS. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING-RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS TRI-DRIVE MODULAR REAMER DRIVER, 750-2007-026-RMRINSTRUMENTS REV H, WAS REVIEWED. THE RISK IS NOT CAPTURED. THIS DOCUMENT WILL BE UPDATED TO INCLUDE THIS FAILURE. REFER TO (B)(4) FOR DETAILS. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RMR WILL BE UPDATED TO CAPTURE THIS RISK. THE DISASSEMBLED DRIVER REPORTED IN EXPERIENCE (B)(4) WAS LIKELY THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY AND ALLOWED THE TIP TO DISASSEMBLE. EXACTECH IFU 700-096-004 STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED INSTRUMENT MALFUNCTION IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO ANY PATIENT IMPACT. AN INVESTIGATION WAS CONDUCTED THAT THE MOST LIKELY ROOT CAUSE OF THE DEVICE BREAKAGE REPORTED IN THIS EVENT WAS THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY AND ALLOWED THE TIP TO DISASSEMBLE. THE AGENT/AGENCY IS NO LONGER ACTIVE WITH EXACTECH. NO ADDITIONAL INFORMATION IS AVAILABLE FROM THE CONTACTS RELATED TO THIS EVENT. NO ADDITIONAL ATTEMPTS WILL BE MADE FOR INFORMATION REGARDING THIS EVENT. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00386 | 2022-04-07 | 2022-03-18 | I, n, j, u, r, y | Manufacturer report | PATELLA | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | PATELLA | UNKNOWN |  | AS REPORTED, SURGEON ASPIRATED PATIENT'S JOINT TODAY AND THE PATIENT IS APPARENTLY HAVING PAIN. PATIENT HAD HER KNEE DONE IN 2019. X-RAYS RECEIVED. NO ADDITIONAL INFORMATION RECEIVED. | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION RECEIVED THE FOLLOWING SECTIONS B5, G3, G6 AND H2 HAVE BEEN UPDATED ACCORDINGLY. SECTION B5: ADDITIONAL INFORMATION RECEIVED INDICATES THAT PATIENT NEEDS A PATELLA RESURFACING. THE PRIMARY WAS A TKA INDICATION WHERE THE PATELLA WAS NOT REPLACED. SURGEON PLAN IS TO RESURFACE PATELLA AND LOOK AT POLY INSERT. THERE IS NO SIGN OF POLY WEAR AT THIS POINT. REVISION SURGERY HAS NOT BEEN SCHEDULED YET. THE DEVICE WILL BE RETURN. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10065 | 2022-04-07 | 2019-10-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TALAR MILL BIT | EXACTECH, INC. |  | 351-05-10 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE MILL TOOL WAS NOT COMPLETELY INSERTED INTO THE CUTTING BLOCK WHEN THE DRILL MACHINE WAS TURNED ON. THIS CAUSED THE CUTTER HEAD TO TILT AND THUS SHEAR OFF. THE SURGERY WENT WELL. IT WAS CONFIRMED THAT NO PIECES FELL INTO THE SURGICAL SITE AND NO PIECES WERE LEFT IN THE PATIENT. IT WAS STATED THAT THE SURGERY WENT WELL WITHOUT ANY OTHER COMPLAINTS. THERE WERE NO DELAYS AND NO REPORTED PATIENT ADVERSE EVENT. THE RETURNED INSTRUMENT APPEARS TO BE MISSING THE TALAR MILL BIT COLLAR SUBCOMPONENT. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE RECEIVED FOR ANALYSIS. ENG EVAL COMPLETED ON 2020.05.18 BY NF. VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR 10X OPTICAL UNLESS OTHERWISE SPECIFIED). THE RETURNED INSTRUMENT APPEARS TO BE MISSING THE TALAR MILL BIT COLLAR SUBCOMPONENT. BASED ON THE INFORMATION PROVIDED BY THE SALES REPRESENTATIVE AND THE FRACTURE SURFACES OF THE SHAFT AND OF THE HUDSON CONNECTION THE FRACTURE APPEARS CONSISTENT WITH THE DEVICE BEING USED OFF-AXIS IN THE CUTTING BLOCK. THE DEFORMATION APPEARS CONSISTENT WITH NORMAL USE DURING MILLING. THE DEFORMATION APPEARS CONSISTENT WITH ASSEMBLING/DISASSEMBLING FROM A MATING DEVICE. THE DEFORMATION ON THE TWO SUBCOMPONENTS MAY ALSO BE THE RESULT OF MEETING SHARP, SURGICAL TOOLS DURING REPROCESSING. DESIGN-RELATED ISSUES: THE DESIGN OF THE TALUS MILL BIT HAS BEEN IN THE FIELD SINCE 2017. EXACTECH IS NOT AWARE OF ANY OTHER SIMILAR COMPLAINT REPORT INVOLVING BROKEN TALUS MILL BITS SINCE ITS INTRODUCTION. BECAUSE THE TALUS MILL BIT IS USED IN TOTAL ANKLE SURGERIES, SALES DATA FOR TALAR COMPONENTS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 50 UNITS, WHICH HAVE BEEN IN THE FIELD SINCE 2018. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING-RELATED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR AND RACR. THE BROKEN TALUS MILL BIT REPORTED IN CASE-2019-00001748 WAS LIKELY THE RESULT OF THE DEVICE BEING USED OFF-AXIS DUE TO NOT FULLY INSERTING THE MILL BIT INTO THE CUTTING BLOCK, AS STATED BY THE SALES REPRESENTATIVE. AT THE TIME OF THIS INVESTIGATION, THIS INCIDENT APPEARS TO BE AN ISOLATED EVENT. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THE RETURNED INSTRUMENT APPEARS TO BE MISSING THE TALAR MILL BIT COLLAR SUBCOMPONENT. THERE WAS NO REPORTED PATIENT ADVERSE EVENT, AND IT WAS CONFIRMED THAT PIECES FELL INTO THE SURGICAL SITE. AN INVESTIGATION WAS CONDUCTED; THE BROKEN TALUS MILL BIT REPORTED WAS LIKELY THE RESULT OF THE DEVICE BEING USED OFF-AXIS DUE TO NOT FULLY INSERTING THE MILL BIT INTO THE CUTTING BLOCK, AS STATED BY THE SALES REPRESENTATIVE. AT THE TIME OF THIS INVESTIGATION, THIS INCIDENT APPEARS TO BE AN ISOLATED EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10060 | 2022-04-07 | 2018-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRULIANT PIN PULLER, SYRINGE STYLE | EXACTECH, INC. |  | 02-029-90-4100 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON WAS TRYING TO PULL ONE OF THE PINS OUT OF THE TIBIA CUTTING BLOCK AND THE METAL RING AROUND THE PRONGS "POPPED" OFF. THE PATIENT WAS STABLE LEAVING THE OR. NO PIECES FELL INTO THE SURGICAL SITE. THERE WAS NO DELAY TO SURGERY THE PATIENT WAS NOT AFFECTED. NO PATIENT DEMOGRAPHIC INFORMATION IS AVAILABLE. THE PATIENT WAS NOTED TO BE FINE AT DISCHARGE FROM THE HOSPITAL. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. (B)(4) & CAPA (B)(4) ENG EVAL COMPLETED ON (B)(6) 2018. ACCORDING TO THE CAPA INVESTIGATION, THE DEVICE HISTORY RECORDS OF THE SUBJECT DEVICES WERE REVIEWED, AND NO ISSUES WERE OBSERVED. A DESIGN REVIEW SHOWED THAT THE BREAKAGE AREA OF THE COMPONENT IS RELATIVELY THIN WHICH DOES NOT OFFER GREAT SAFETY MARGIN WHEN THE INSTRUMENT IS SUBJECT TO OFF-AXIS LOADING. ALTHOUGH THE EVALUATION COULDN'T CONFIRM USER MISUSE, OFF-AXIS LOADING WHICH CAN BE GENERATED BY THE USER APPLYING A BENDING MOMENT TO THE DEVICE DURING PIN DRIVING/PULLING COULD CONTRIBUTE TO THE BREAKAGE. ROOT CASE BASED ON THE CAPA INVESTIGATION, THE BROKEN TRULIANT PIN PULLER REPORTED IN (B)(4) WAS LIKELY THE RESULT OF: THE BREAKAGE AREA OF THE COMPONENT BEING RELATIVELY THIN WHICH DOES NOT OFFER GREAT SAFETY MARGIN WHEN THE INSTRUMENT IS SUBJECT TO OFF-AXIS LOADING. OFF-AXIS LOADING WHEN THE USER APPLIES A BENDING MOMENT TO THE DEVICE DURING PIN DRIVING/PULLING. CORRECTIVE ACTION TAKEN: UPDATE DESIGN TO IMPROVE STRENGTH OF BREAKAGE AREA; UPDATE OPERATIVE TECHNIQUE TO INSTRUCT USERS TO AVOID USING THE INSTRUMENT OFF-AXIS. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT IMPACT. BASED ON CAPA (B)(4) INVESTIGATION, THE BROKEN TRULIANT PIN PULLER REPORTED IN (B)(4) WAS LIKELY THE RESULT OF: THE BREAKAGE AREA OF THE COMPONENT BEING RELATIVELY THIN WHICH DOES NOT OFFER GREAT SAFETY MARGIN WHEN THE INSTRUMENT IS SUBJECT TO OFF-AXIS LOADING. OFF-AXIS LOADING WHEN THE USER APPLIES A BENDING MOMENT TO THE DEVICE DURING PIN DRIVING/PULLING. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00383 | 2022-04-07 | 2022-02-10 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | HUMERAL STEM PRIMARY, PRESS FIT 13MM | EXACTECH, INC. | 300-01-13 | 300-01-13 |  | AS REPORTED, APPROXIMATELY 4.5 YEARS POST OP THE INITIAL TSA, THIS (B)(6) FEMALE PATIENT WAS REVISED DUE A MIDSHAFT HUMERAL FRACTURE. REVISED TO LONG STEM HUMERUS AND ORIF OF PERIPROSTHETIC HUMERAL FRACTURE. REASON FOR REVISION WAS NOT REPORTED. DESCRIPTION INFORMATION OF ADVERSE EVENT IS NOT AVAILABLE. THE CASE REPORT FORM INDICATES THIS EVENT IS UNLIKELY RELATED TO DEVICES OR PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. DEVICES WILL NOT BE RETUNING. | PENDING EVALUATION: CONCOMITANT DEVICE(S): 320-38-00 - EQUINOXE REVERSE 38MM HUMERAL LINER +0, 320-10-00 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0, 320-01-38 - EQUINOXE REVERSE 38MM GLENOSPHERE. NO INFORMATION PROVIDED IN THE FOLLOWING SECTION(S). | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10040 | 2022-04-07 | 2016-11-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TIBIAL PROTECTOR | EXACTECH, INC. |  | 351-17-01 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGICAL PROCEDURE THE TIBIAL PROTECTOR BROKE DURING TALAR IMPACTION. THE EVENT DID NOT CAUSE A DELAY TO SURGERY. THE PATIENT WAS NOT ADVERSELY AFFECTED, NO PIECE OF THE DEVICE WAS LEFT IN THE PATIENT. THE REPRESENTATIVE WAS PRESENT AT THE TIME OF SURGERY. IT WAS REPORTED THAT THE "INSTRUMENT WAS REENGINEERED". NO. ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS NOT RECEIVED FOR ANALYSIS. CAPA (B)(4) INVESTIGATED BROKEN TIBIAL PROTECTORS (351-17-01) DURING TALAR FINDINGS: IN CASES WHERE JOINT SPACE IS REDUCED, THE TIBIAL PROTECTOR COMPONENT CAN CONTACT THE TALAR IMPACTOR TIP DURING INTRAOPERATIVELY THREE TIMES AND EACH TIME THE PORTION THAT PROTECTED THE IMPLANT REMAINED IN PLACE BUT HAD TO BE TO GRAB WAS MISSING. ROOT CAUSE: THE TIBIAL PROTECTOR PLATE HAS AN ANTERIOR FLANGE TO ALLOW FOR MANUAL INSERTION OR REMOVAL OF THE DEVICE. IMPACTOR TIP (351-07-00) CAN CONTACT THE ANTERIOR FLANGE DURING IMPACTION OF THE TALAR IMPLANT. THIS PUTS A BENDING FRACTURE AFTER REPEATED IMPACTION. THE VALIDATION OF THE TIBIAL PROTECTOR PLATE WAS PERFORMED ON CADAVERIC LEGS, BUT NO SEVERE TISSUE CONTRACTURE WAS PRESENT. PROTECTED THE TIBIAL COMPONENTS, BUT WAS NOT STRUCK DURING USE. THEREFORE, THE CONTACT DURING IMPACTION BETWEEN GRIPPING THE COMPONENT RESULTING IN FRACTURE OF THE DEVICE WAS NOT IDENTIFIED. CORRECTIVE ACTION TAKEN: THE DESIGN OF THE TIBIAL PROTECTOR WAS MODIFIED TO REMOVE THE ANTERIOR FLANGE PORTION AND INSTRUMENTS WAS UPDATED TO INCLUDE FRACTURE OF THE TIBIAL PROTECTOR COMPONENT UNDER DCRSP-17-490. NO ADDITIONAL ACTION REQUIRED. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT IMPACT. THE TIBIAL PROTECTOR BROKE DURING TALAR IMPACTION. AN INVESTIGATION WAS CONDUCTED UNDER CAPA (B)(4); THE TIBIAL PROTECTOR PLATE HAS AN ANTERIOR FLANGE TO ALLOW FOR MANUAL INSERTION OR REMOVAL OF THE DEVICE. IN TIGHT ANKLES OR ANKLES WITH POOR EXPOSURE, THE TALAR IMPACTOR TIP CAN CONTACT THE ANTERIOR FLANGE DURING IMPACTION OF THE TALAR IMPLANT. THIS PUTS A BENDING FORCE ON THE ANTERIOR FLANGE AND LEADS TO A CRACK OR FRACTURE AFTER REPEATED IMPACTION. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10036 | 2022-04-07 | 2017-08-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRULIANT TIBIAL INSERT TRIAL PS TOPPER | EXACTECH, INC. |  | 02-029-25-3535 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE TIBIAL INSERT TRIALS. THE BAL-SEAL SPRINGS WERE MISSING FROM THE TIBIAL INSERT TRIALS, IT ISN'T CLEAR WHEN THE DISENGAGEMENT HAPPENED. TWO X-RAYS WERE CHECKED, AND THE BAL-SPRINGS WERE NOT LEFT IN THE PATIENT. THE PATIENT HAD AN UNEVENTFUL RECOVERY. THERE WAS NO DELAY TO SURGERY. X-RAYS WERE SATISFACTORY WITH NO EVIDENCE OF ANY FOREIGN BODY. THE PATIENT WAS SEEN AT 6 WEEKS POST-OP AND WAS NOTED TO BE DOING WELL. NO ADVERSE PATIENT OUTCOME. PHOTOS WERE PROVIDED. PATIENT DEMOGRAPHIC INFORMATION WAS NOT PROVIDED. THE DEVICES WERE RECEIVED. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED BY THE CONTACTS ABOUT THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICES WERE AVAILABLE FOR ANALYSIS. ENGINEERING ANALYSIS COMPLETED BY AD ON 6/5/2018. DESIGN-RELATED ISSUES: OVER 560 CASES HAVE BEEN COMPLETED (AS OF NOVEMBER 20 I 7) WITH NO OTHER REPORTED INSTANCES. THEREFORE, THIS DOES NOT APPEAR TO BE RELATED TO THE INSTRUMENT DESIGN. MFG-RELATED ISSUES: OVER 560 CASES HAVE BEEN COMPLETED (AS OF NOVEMBER 2017) WITH NO OTHER REPORTED INSTANCES. THEREFORE, THIS DOES NOT APPEAR TO BE RELATED TO THE MANUFACTURING OF THE INSTRUMENT. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS INSTRUMENT, 750- 2014-006-RMR- INSTRUMENTS REV B, WAS REVIEWED. THE RISK IS CAPTURED IN 35. D. 9 B. THE BAL-SPRING DISENGAGEMENT REPORTED IN EXPERIENCE C2017-548 WAS LIKELY THE RESULT OF USING A TOOL TO REMOVE THE SPRING, POSSIBLY DURING CLEANING. THERE IS NO EVIDENCE THAT THE DISENGAGEMENT OCCURRED DURING NORMAL USE IN SURGERY. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. EXACTECH IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO ADVERSE EVENT REPORTED. THE DISENGAGEMENT OF THE BAL-SPRING WAS NOT KNOWN AT THE TIME OF SURGERY, X-RAYS WERE TAKEN TO ENSURE NO FOREIGN OBJECTS WERE IN THE PATIENT. AN INVESTIGATION WAS CONDUCTED: THE BAL-SPRING DISENGAGEMENT REPORTED WAS LIKELY THE RESULT OF USING A TOOL TO REMOVE THE SPRING, POSSIBLY DURING CLEANING. THERE IS NO EVIDENCE THAT THE DISENGAGEMENT OCCURRED DURING NORMAL USE IN SURGERY. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10050 | 2022-04-07 | 2018-02-08 | I, n, j, u, r, y | Manufacturer report | EXACTECH | TRULIANT LOCKING FEMORAL IMPACTOR | EXACTECH, INC. |  | 02-029-19-1000 |  | IT WAS REPORTED FROM THE US THAT DURING A SURGICAL PROCEDURE A DEVICE WAS BROKEN. "THE GRAY PLASTIC PIECE THAT RESTS AGAINST THE FEMORAL TRIAL OR FEMORAL IMPLANT WAS BROKEN DURING IMPACTION ONTO THE BONE." THE PATIENT'S HEALTH STABLE LEAVING THE OPERATING ROOM. THERE WAS NO DELAY TO SURGERY AND NO PATIENT ADVERSE EFFECT. THE REPRESENTATIVE WAS PRESENT DURING THE SURGERY. IT WAS CONFIRMED THAT NO PIECES WERE LEFT IN THE PATIENT. THE ISSUE WAS RESOLVED BY USING THE MODULAR IMPACTOR. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE COMPLAINT PRODUCT WAS NOT RECEIVED FOR ANALYSIS. THE BROKEN/NON-FUNCTIONAL IMPACTOR CANNOT BE CONFIRMED. THE DESCRIPTION OF THE FAILURE IN THE EXPERIENCE REPORT INDICATES THE PLASTIC PORTION OF THE IMPACTOR THAT CONTACTS THE FEMORAL IMPLANT WAS CRACKED DURING IMPACTION. HOWEVER, THIS CANNOT BE PROVEN UNTIL THE DEVICES ARE RECEIVED BY EXACTECH. REVIEW OF PREVIOUS INVESTIGATED SIMILAR COMPLAINTS CONFIRMED THE REPORTED FAILURE AND REVEALED THAT THE RADEL MATERIAL ON THE IMPACTION SURFACE HAD EITHER CRACKED OR FRACTURED. REVIEW OF THE PREVIOUSLY INVESTIGATED COMPLAINT REVEALED NO DESIGN OR MANUFACTURING RELATED ISSUES ASSOCIATED WITH THIS REPORTED FAILURE. THE CAUSE OF THE REPORTED EVENT IS MOST LIKELY RELATED TO NORMAL WEAR AND TEAR OF THESE DEVICES USED UNDER IMPACTION FORCES. IFU 700-096-181: INSTRUMENT INSPECTION. VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. IT WAS CONFIRMED THAT NO PIECES WERE LEFT IN THE PATIENT. AN INVESTIGATION WAS CONDUCTED; THE CAUSE OF THE REPORTED EVENT IS MOST LIKELY RELATED TO NORMAL WEAR AND TEAR OF THESE DEVICES USED UNDER IMPACTION FORCES. | N | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10041 | 2022-04-07 | 2016-11-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TIBIAL PROTECTOR | EXACTECH, INC. |  | 351-17-01 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGICAL PROCEDURE THE TIBIAL PROTECTOR BROKE DURING TALAR IMPACTION. THERE WAS NO ADVERSE IMPACT TO PATIENT AND NO DELAY TO SURGERY. NO PIECES WERE LEFT IN THE PATIENT. THE REPRESENTATIVE WAS PRESENT DURING THE SURGERY. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS NOT RECEIVED FOR ANALYSIS. CAPA (B)(4) INVESTIGATED BROKEN TIBIAL PROTECTORS (351-17-01) DURING TALAR IMPACTION, AS WELL AS BROKEN TIBIAL ALIGNMENT GUIDES. FINDINGS: IN CASES WHERE JOINT SPACE IS REDUCED, THE TIBIAL PROTECTOR COMPONENT CAN CONTACT THE TALAR IMPACTOR TIP DURING USE. THE TIBIAL PROTECTOR PLATE HAS BROKEN INTRAOPERATIVELY THREE TIMES AND EACH TIME THE PORTION THAT PROTECTED THE IMPLANT REMAINED IN PLACE BUT HAD TO BE REMOVED USING A CURETTE OR HEMOSTAT BECAUSE THE REGION TO GRAB WAS MISSING. THE VALIDATION OF THE TIBIAL PROTECTOR PLATE WAS PERFORMED ON CADAVERIC LEGS, BUT NO SEVERE TISSUE CONTRACTURE WAS PRESENT. THEREFORE, THE TIBIAL PROTECTOR PLATE ADEQUATELY PROTECTED THE TIBIAL COMPONENTS, BUT WAS NOT STRUCK DURING USE. THEREFORE; THE CONTACT DURING IMPACTION BETWEEN THE TALAR IMPACTOR TIP AND THE ANTERIOR FLANGE MEANT FOR GRIPPING THE COMPONENT RESULTING IN FRACTURE OF THE DEVICE WAS NOT IDENTIFIED. CORRECTIVE ACTION TAKEN: THE DESIGN OF THE TIBIAL PROTECTOR WAS MODIFIED TO REMOVE THE ANTERIOR FLANGE PORTION AND UTILIZE THE SCISSOR INSERTERS. THE RMR FOR VANTAGE INSTRUMENTS WAS UPDATED TO INCLUDE FRACTURE OF THE TIBIAL PROTECTOR COMPONENT UNDER DCRSP-17-490. ROOT CAUSE: THE TIBIAL PROTECTOR PLATE HAS AN ANTERIOR FLANGE TO ALLOW FOR MANUAL INSERTION OR REMOVAL OF THE DEVICE. IN TIGHT ANKLES OR ANKLES WITH POOR EXPOSURE, THE TALAR IMPACTOR TIP (351- 07-00) CAN CONTACT THE ANTERIOR FLANGE DURING IMPACTION OF THE TALAR IMPLANT. THIS PUTS A BENDING FORCE ON THE ANTERIOR FLANGE AND LEADS TO A CRACK OR FRACTURE AFTER REPEATED IMPACTION. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT IMPACT. THE TIBIAL PROTECTOR BROKE DURING TALAR IMPACTION. AN INVESTIGATION WAS CONDUCTED UNDER CAPA (B)(4); THE TIBIAL PROTECTOR PLATE HAS AN ANTERIOR FLANGE TO ALLOW FOR MANUAL INSERTION OR REMOVAL OF THE DEVICE. IN TIGHT ANKLES OR ANKLES WITH POOR EXPOSURE, THE TALAR IMPACTOR TIP (351-07-00) CAN CONTACT THE ANTERIOR FLANGE DURING IMPACTION OF THE TALAR IMPLANT. THIS PUTS A BENDING FORCE ON THE ANTERIOR FLANGE AND LEADS TO A CRACK OR FRACTURE AFTER REPEATED IMPACTION. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10042 | 2022-04-07 | 2017-02-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | SPECIAL INSTRUMENT PATELLA GUIDE | EXACTECH, INC. |  | SSI-200-009471 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGICAL PROCEDURE THE PATELLA GUIDE BROKE. THE LOCKING WASHER FELL OFF, SO THE CUT WAS DONE FREE HANDED. NOTHING WAS LEFT IN THE PATIENT AND THERE WAS NO NEGATIVE OUTCOME TO PATIENT. WHEN ASKED IF THE BROKEN PATELLA GUIDE WOULD BE RETURNING FOR EVALUATION, THE CONTACT STATED THAT THE HANDLE WAS REPAIRED, AND THE SURGEON WILL CONTINUE USING IT. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE COMPLAINT PRODUCT WAS NOT RETURNED FOR ANALYSIS. THE COMPLAINT CONDITION OF A BROKEN SPECIAL. INSTRUMENT PATELLA GUIDE IS NOT CONFIRMED. MANUFACTURING-RELATED DATA COULD NOT BE REVIEWED. BECAUSE NO INFORMATION REGARDING THE MANUFACTURING LOTS WAS PROVIDED. HOWEVER, A PFMEA FOR KNEE INSTRUMENTS, 750-2012-027-PFMEA-INSTRUMENT KNEE REV-, COVERED ISSUES RELATED TO MANUFACTURING, INCLUDING INADEQUATE FASTENING THAT CAUSES DISASSEMBLY AND IMPROPER INSPECTION, BOTH OF WHICH POTENTIALLY CAUSED THE LOCKING WASHER TO FALL OFF. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE SPECIAL HANDLES ARE NOT CONTINUING TO BE MANUFACTURED, AND THE RISK IS CAPTURED IN THE RMR. ON (B)(6) 2017, WHEN ASKED IF THE BROKEN PATELLA GUIDE WOULD BE RETURNING TO EXACTECH FOR EVALUATION, THE CONTACT STATED THAT THE HANDLE WAS REPAIRED, AND THE SURGEON WILL CONTINUE USING IT. WHILE IT IS UNDETERMINED WHICH PART OF THE BROKEN PATELLA GUIDE (I.E., WHAT IS THE "LOCKING WASHER"?) REPORTED FELL OFF, THE PIECES MOST LIKELY TO FALL OFF WERE LIKELY DUE TO A BROKEN END PIECE AND/OR THE SURGEON UNSCREWING A PIECE PAST THE THREADS WHILE ADJUSTING THE PATELLA GUIDE. HOWEVER, THE LIKELY CAUSE CANNOT BE CONFIRMED BECAUSE THE INSTRUMENT WAS NOT RETURNED FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT IMPACT. THE LOCKING WASHER FELL OFF, SO THE CUT WAS DONE FREE HANDED, NOTHING WAS LEFT IN THE PATIENT. AN INVESTIGATION WAS CONDUCTED; WHILE IT IS UNDETERMINED WHICH PART OF THE BROKEN PATELLA GUIDE (I.E., WHAT IS THE "LOCKING WASHER"?) REPORTED FELL OFF, THE PIECES MOST LIKELY TO FALL OFF WERE LIKELY DUE TO A BROKEN END PIECE AND/OR THE SURGEON UNSCREWING A PIECE PAST THE THREADS WHILE ADJUSTING THE PATELLA GUIDE. HOWEVER, THE LIKELY CAUSE CANNOT BE CONFIRMED BECAUSE THE INSTRUMENT WAS NOT RETURNED FOR EVALUATION. THIS IS A SPECIAL INSTRUMENT MFG DATE IS 11/02/2016 OR 11/04/2016. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10043 | 2022-04-07 | 2017-05-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | HANDLE, MODULAR BROACH | EXACTECH, INC. |  | 301-03-01 |  | SUMMARY OF REPORTED INCIDENT: WHILE THE SURGEON WAS USING A MALLET ON THE MODULAR BROACH HANDLE TO BROACH THE HUMERUS WITH BROACH, THE SWIVEL PLATE AT THE TIP OF THE MODULAR BROACH HANDLE FELL OFF. BROACHING WAS COMPLETED WITHOUT ISSUE. PLATE WAS REMOVED. BROACHING CONTINUED AS IT WAS ALMOST COMPETED. BROACH HANDLE WAS REMOVED FROM THE SET AND IS BEING RETURNED FOR EVALUATION. NO DELAY. AGENT PRESENT AT TIME OF SURGERY. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED FOR ANALYSIS. COMPLETED ON 20 MAR 2018 BY AD. VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR L0X OPTICAL UNLESS OTHERWISE SPECIFIED) THE DETACHED FOOT APPEARS TO HAVE DISASSOCIATED FROM THE HANDLE ASSEMBLY. THE FOOT IS ATTACHED TO THE REST OF THE DEVICE BY TWO 316 STAINLESS STEEL DOWEL PINS. THE PINS APPEAR TO HAVE BACKED OUT, ALLOWING THE FOOT TO DETACH FROM THE ASSEMBLY. ONLY ONE DOWEL PIN WAS RETURNED. THE SALES REPRESENTATIVE CONFIRMED THAT ALL COMPONENTS WERE RETRIEVED FROM THE SURGICAL FIELD FOLLOWING THE SURGERY. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE AND CONSISTENT WITH YEARS OF USE. DESIGN-RELATED ISSUES: THIS DESIGN HAS BEEN IN THE FIELD SINCE 2004. EXACTECH HAS RECEIVED 10 SIMILAR COMPLAINT REPORTS INVOLVING THIS DEVICE SINCE 2008. BECAUSE THIS DEVICE IS USED IN BOTH PRIMARY AND REVERSE SHOULDER REPLACEMENTS, SALES DATA FOR HUMERAL STEMS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGNRELATED. MFG-RELATED ISSUES: EXACTECH HAS NOT RECEIVED ANY OTHER COMPLAINT REPORTS INVOLVING PARTS FROM THIS MANUFACTURING LOT, WHICH HAS BEEN IN THE FIELD SINCE 2012. THEREFORE; THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING-RELATED A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS MODULAR BROACH HANDLE, 750-2002-039-RMR REV F, WAS REVIEWED. MOST LIKELY CAUSE: THE BROKEN INSTRUMENT REPORTED IN EXPERIENCE C2017-233 WAS LIKELY THE RESULT OF THE DOWEL PINS BACKING OUT, ALLOWING FOR THE FOOT COMPONENT TO DETACH FROM THE HANDLE. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE WAS NO PATIENT ADVERSE EVENT RELATED TO THE DETACHED TIP. AN INVESTIGATION WAS CONDUCTED; THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF THE DOWEL PINS BACKING OUT, ALLOWING FOR THE FOOT COMPONENT TO DETACH FROM THE HANDLE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00384 | 2022-04-07 | 2022-03-18 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | CAGED GLENOID | EXACTECH, INC. |  | UNK |  | AS REPORTED, POST OP THE INITIAL RIGHT TSA, THIS (B)(6) MALE PATIENT'S WAS REVISED DUE LOOSENING. SURGEON REVISED AN ANATOMIC TOTAL SHOULDER TO A HEMI. HE REMOVED THE CAGED GLENOID, HUMERAL HEAD, AND REPLICATOR PLATE THEN REPLACED THEM. CENTER CAGE OF THE GLENOID DID NOT COME OUT WITH THE GLENOID. THE REMOVAL TOOL COULD NOT THREAD INTO THE CAGE BECAUSE IT HAD BROKEN. OSTEOTOMES WERE USED TO REMOVE THE CAGE AND HALF OF THE CAGE HAD BROKEN. SALES REPRESENTATIVE CLARIFIED THAT THERE WAS NO ISSUE WITH THE REMOVAL TOOL. THERE WAS A 5-15 MIN DELAY/PROLONGATION BUT NO ADVERSE EVENTS AS A RESULT. SURGEON WAS HAPPY WITH THE OUTCOME UPON LOOKING AT FINAL XRAY. XRAY OF PRE OP IMAGE IS PROVIDED. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETURNING, DOCTOR DID NOT RELEASE. | CONCOMITANT DEVICE(S): HUMERAL HEAD, REPLICATOR PLATE. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10002 | 2022-04-07 | 2017-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LPI IMPACTOR HANDLE | EXACTECH, INC. |  | 213-65-00 |  | IT WAS REPORTED THAT DURING AN ORTHOPEDIC KNEE SURGERY THE SURGEON EXPERIENCED THE SPRING IN THE LPI IMPACTOR HANDLE BROKE. THE SALES REPRESENTATIVE WAS NOT PRESENT AT THE SURGERY. IT WAS REPORTED THAT NO DEVICE FRAGMENTS FELL INTO THE SURGICAL SITE. THE SURGERY CONTINUED AND WAS COMPLETED WITHOUT ADDITIONAL ISSUES. THERE WAS NO DELAY OF SURGERY AND THERE WAS NO REPORTED PATIENT INJURY, ADVERSE EVENT, OR CLINICAL CONSEQUENCE TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE REPORTER RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE IN QUESTION WAS RETURNED TO EXACTECH FOR ANALYSIS. DESIGN-RELATED ISSUES: THE LPI IMPACTOR HANDLE HAS BEEN ON THE MARKET SINCE 2011. EXACTECH IS AWARE OF 37 COMPLAINTS INVOLVING THE HANDLE SINCE 2014. SINCE THIS INSTRUMENT IS USED TO IMPLANT TIBIAL TRAYS, THE SALES DATA FOR ALL TIBIAL TRAYS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE DETAILED BELOW. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS AWARE OF 2 OTHER COMPLAINT REPORTS INVOLVING DAMAGED SPRINGS ON PARTS FROM THIS MANUFACTURING LOT OF 52 UNITS. HOWEVER, A REVIEW OF THE DEVICE HISTORY RECORD (DHR) SHOWED THAT THE NAMED DEVICES WERE ACCEPTED WITH CONFORMANCE TO THE DEVICE REQUIREMENTS. THEREFORE, THIS DOES NOT APPEAR TO BE MANUFACTURING-RELATED. RISK MANAGEMENT: A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THE LPI INSTRUMENTS, 750-2003-019-RMR-INSTRUMENTS REV B, WAS REVIEWED. THE RISK IS CAPTURED IN LINES 18 AND 23 OF THE RMR. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE OUTSTRETCHED SPRING REPORTED WAS LIKELY THE RESULT OF REMOVING A TIBIAL TRAY/TRIAL AT AN ANGLE AFTER MULTIPLE IMPACTIONS THROUGHOUT THE LIFE OF THE DEVICE. EXACTECH IFU 700-096-004 STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED INSTRUMENT MALFUNCTION IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO ANY PATIENT IMPACT. AN INVESTIGATION WAS CONDUCTED AND CONCLUDED THAT THE MOST LIKELY ROOT CAUSE OF THE OUTSTRETCHED SPRING REPORTED WAS LIKELY THE RESULT OF REMOVING A TIBIAL TRAY/TRIAL AT AN ANGLE AFTER MULTIPLE IMPACTIONS THROUGHOUT THE LIFE OF THE DEVICE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10047 | 2022-04-07 | 2017-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | BASEPLATE INSERTER | EXACTECH, INC. |  | SK-003217 |  | IT WAS REPORTED FROM THE US DURING AN ORTHOPEDIC SURGERY THAT A SURGEON HAD AN ISSUE WITH THE BASEPLATE INSERTER. THE WELD CAME OFF ON THE BACK OF THE CUSTOM BASEPLATE INSERTERS. THE DOCTOR WAS STILL ABLE TO USE INSERTER AS INTENDED, SHE HAD TO USE PLIERS TO REMOVE IT. THE AGENT WAS PRESENT AT SURGERY AND THERE WAS NO DELAY.ATTEMPTS HAVE BEEN MADE FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. NO ADDITIONAL ATTEMPTS WILL BE MADE FOR INFORMATION REGARDING THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020. EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED FOR ANALYSIS. COMPLETED ON 20 JUN 2019 BY CY. MFG DATE: 12/18/2015 VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR 10X OPTICAL UNLESS OTHERWISE SPECIFIED) LOOSE STRIKE PLATE ON THE SPECIAL COUNTER TORQUE BASEPLATE INSERTER. THE LOOSE HANDLE ON THE SPECIAL COUNTER TORQUE BASEPLATE INSERTER APPEARS CONSISTENT WITH THE FAILURE OF THE WELD. THE FAILURE OF THE WELD APPEARS CONSISTENT WITH SIGNIFICANT MECHANICAL FORCES BEING APPLIED TO THE STRIKE PLATE AND LOSS OF MECHANICAL INTEGRITY OF THE DEVICE. THIS FAILURE APPEARS TO BE DESIGN RELATED. DESIGN-RELATED ISSUES: THE SPECIAL COUNTER TORQUE BASEPLATE INSERTER HAS BEEN IN THE FIELD SINCE 2013. THE FAILURE OF THIS INSTRUMENT APPEARS TO BE DIRECTLY RELATED TO THE FEATURE THAT MADE THIS A SPECIAL INSTRUMENT AND DIFFERENT FROM THE PARENT/PREDICATE INSTRUMENT. CAPA (B)(4) WAS INITIATED TO ENSURE THAT SPECIAL INSTRUMENTS ARE DESIGNED AND MANUFACTURED AFTER SUFFICIENT DESIGN CONTROL ACTIVITIES ARE COMPLETED AND DOCUMENTED. THE SPECIAL INSTRUMENT SOP IS BEING UPDATED AS PART OF THE CAPA TO IMPROVE THE CONNECTION BETWEEN INPUTS, PREDICATE DEVICES, RISK ASSESSMENT, AND POST -MARKET FEEDBACK. MFG-RELATED ISSUES: THE MANUFACTURING LOT INFORMATION FOR THE SPECIAL COUNTER TORQUE BASEPLATE INSERTER PIECE WAS REVIEWED. THIS BREAKAGE APPEARS TO BE RELATED TO THE DESIGN OF THE SPECIAL INSTRUMENT. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) FOR THE PREDICATE INSTRUMENT WAS CONDUCTED TO ENSURE THE RISKS WERE INCLUDED. THE RMR FOR THE GLENOID INSERTER, 750-2007-026-RMR-INSTRUMENTS REV L, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 18. MOST LIKELY UNDERLYING CAUSE: THE BROKEN SPECIAL COUNTER TORQUE BASEPLATE INSERTER REPORTED IN EXPERIENCE C2017-352 WAS LIKELY THE RESULT OF THE DESIGN OF THE INSTRUMENT. CAPA (B)(4) HAS BEEN INITIATED TO ENSURE THAT SPECIAL INSTRUMENTS ARE DESIGNED AND MANUFACTURED AFTER SUFFICIENT DESIGN CONTROL ACTIVITIES ARE COMPLETED AND DOCUMENTED. IFU 700-096-181: INSTRUMENT INSPECTION VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED INSTRUMENT MALFUNCTION IS RELATED TO ANY MANUFACTURING ISSUES. NOR DID IT LEAD TO ANY PATIENT IMPACT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN SPECIAL COUNTER TORQUE BASEPLATE INSERTER REPORTED WAS LIKELY THE RESULT OF THE DESIGN OF THE INSTRUMENT. CAPA (B)(4) HAS BEEN INITIATED TO ENSURE THAT SPECIAL INSTRUMENTS ARE DESIGNED AND MANUFACTURED AFTER SUFFICIENT DESIGN CONTROL ACTIVITIES ARE COMPLETED AND DOCUMENTED. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10039 | 2022-04-07 |  | I, n, j, u, r, y | Manufacturer report | EXACTECH | FIXED BEARING LINER TRIAL | EXACTECH, INC. |  | 321-21-43 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON HAD DIFFICULTY REMOVING THE ANKLE TRIAL, THE SURGEON USED THE OSTEOTOME AND BROKE THE PLASTIC. THERE WAS NO DELAY TO SURGERY AND NO ADVERSE EFFECT. NO PIECES LEFT IN THE WOUND SITE. THE REPORTER WAS UNSURE OF THE SURGERY DATE. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE COMPLAINT PRODUCT WAS RETURNED FOR ANALYSIS. THE CONDITION OF BEING BROKEN WAS CONFIRMED. VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR 10X OPTICAL) THE BROKEN FIXED BEARING LINER TRIAL APPEARS TO BE CONSISTENT WITH ATTEMPTING TO REMOVE THE TRIAL WITH AN OSTEOTOME AND BREAKING IT IN THE PROCESS. AFTER FIGURING OUT THE PROPER SIZE FOR THE IMPLANT, THE TRIAL IS MEANT TO BE REMOVED USING THE SCISSOR STYLE INSERTER HANDLE. THE FREQUENCY OF OCCURRENCE RANKING SCALE IS VERY LOW; THEREFORE, THIS DOES NOT APPEAR TO BE DESIGN RELATED. THE COMPANY IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 5 UNITS, WHICH HAS BEEN IN THE FIELD SINCE 2016. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING-RELATED. THE DAMAGED COMPONENT REPORTED IS THE RESULT OF USING AN OSTEOTOME TO REMOVE THE LINER TRIAL, EVENTUALLY LEADING TO THE FRACTURE OF THE MATING FEATURE. OSTEOTOMES ARE NOT THE INSTRUMENT INDICATED FOR REMOVAL OF LINER TRIALS IN THE OPERATIVE TECHNIQUE. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT IMPACT. THE SURGEON HAD DIFFICULTY REMOVING THE ANKLE TRIAL, THEY ENDED UP USING OSTEOTOME AND BROKE THE PLASTIC. AN INVESTIGATION WAS CONDUCTED; THE DAMAGED COMPONENT REPORTED IS THE RESULT OF USING AN OSTEOTOME TO REMOVE THE LINER TRIAL, EVENTUALLY LEADING TO THE FRACTURE OF THE MATING FEATURE. OSTEOTOMES ARE NOT THE INSTRUMENT INDICATED FOR REMOVAL OF LINER TRIALS IN THE OPERATIVE TECHNIQUE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10063 | 2022-04-07 | 2018-11-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LOCKING FEMORAL IMPACTOR | EXACTECH, INC. |  | SK-012228 |  | IT WAS REPORTED FROM THE US THAT DURING A BILATERAL TOTAL KNEE REPLACEMENT AN INSTRUMENT BROKE. AFTER BOTH KNEES HAD BEEN REPLACED, AND ALL COMPONENTS CEMENTED INTO PLACE, A SMALL PIECE OF BLACK PLASTIC WAS NOTED ON THE STERILE FIELD. THIS WAS NOTICED DURING CLOSURE OF THE LEFT KNEE. THE RIGHT KNEE WAS ALREADY CLOSED AND DRESSED. A BRIEF SEARCH FOR THE SOURCE OF THE PIECE OF BLACK PLASTIC REVEALED THAT THE BLACK POLY FROM THE HEAD OF THE LOCKING FEMORAL IMPACTOR HAD CRACKED DOWN ITS CENTER BETWEEN THE TWO PEGS THAT HOLD THE POLY IN PLACE ON THE HANDLE. A SMALL CHIP HAD CRACKED OFF THE CENTER CRACK AND FALLEN LOOSE. ALL PIECES WERE COLLECTED, WASHED, AND ACCOUNTED FOR. THE AGENT WAS PRESENT DURING THE SURGICAL PROCEDURE. THERE WAS NO DELAY TO SURGERY AND NO PATIENT ADVERSE EVENT. THE AGENCY IN INACTIVE AND NO MORE INFORMATION IS AVAILABLE FOR THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED FOR ANALYSIS. ENG EVAL ATTACHED JUN 2019 BY (B)(6) THE BROKEN IMPACTION SURFACE APPEARS CONSISTENT WITH MATERIAL DEGRADATION ASSOCIATED WITH NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. ALL OTHER ASPECTS OF THE DEVICES APPEAR UNREMARKABLE AND CONSISTENT WITH NORMAL USE. THE DESIGN OF THIS SPECIAL INSTRUMENT WAS DEVELOPED BY MODIFYING A PREDICATE LOCKING LPI FEMORAL IMPACTOR (CATALOG #: 213-64-01). THE FIRST REVISION (DSR215678) REPLACED THE SCREWS USED TO HOLD THE PLASTIC IMPACTION CUSHION TO THE METAL BASE WITH TWO SMOOTH METAL PEGS THAT ARE STABILIZED BY DOWEL PINS. REFER TO TM-2016-0015. THE SECOND MODIFICATION (DSR217362) REPLACED THE STOP ON THE END OF THE ADJUSTMENT SCREW WITH A WELD. REFER TO TM 2017 0680 SI. DESIGN-RELATED ISSUES: THE DESIGN OF THE LOCKING LPI FEMORAL IMPACTOR HAS BEEN IN THE FIELD SINCE 2005. EXACTECH IS AWARE OF 16 OTHER COMPLAINT REPORTS INVOLVING BROKEN FEMORAL IMPACTION SURFACES OF LOCKING FEMORAL IMPACTORS SINCE 2014. SINCE LOCKING FEMORAL IMPACTORS ARE USED DURING TOTAL KNEE SURGERIES, SALES DATA FOR FEMORAL COMPONENTS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH COMPLAINT DATA FROM 2014 TO 2018 INVOLVING THE REPORTED FAILURE FOR THIS FAMILY OF DEVICES WAS REVIEWED. THE INVESTIGATIONS HAVE NOT IDENTIFIED ANY EVIDENCE THAT A MANUFACTURING ISSUE CAUSED OR CONTRIBUTED TO THIS REPORTED ISSUE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY UNDERLYING CAUSE-THE BROKEN INSTRUMENT REPORTED IN EXPERIENCE C2018-075 WAS LIKELY THE RESULT OF MATERIAL DEGRADATION ASSOCIATED WITH NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. THE INSTRUMENT BROKE AND CAUSED A SMALL "CHIP" TO FALL INTO THE STERILE FIELD. AN INVESTIGATION WAS CONDUCTED; THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF MATERIAL DEGRADATION ASSOCIATED WITH NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. INFORMATION WAS NOT PROVIDED FOR: ETHNICITY, RACE. OF THE 2 DEVICES REPORTED THE SPECIFIC BROKEN DEVICE WAS NOT IDENTIFIED. (2 OF 2 COMPLAINTS RELATED TO C2018-075 CAT # SK-012228/ LOT # DSR215678). | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10049 | 2022-04-07 | 2018-01-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LOCKING FEMORAL IMPACTOR | EXACTECH, INC. |  | SK-012228 |  | IT WAS REPORTED FROM THE US THAT DURING A BILATERAL TOTAL KNEE REPLACEMENT AN INSTRUMENT BROKE. AFTER BOTH KNEES HAD BEEN REPLACED, AND ALL COMPONENTS CEMENTED INTO PLACE, A SMALL PIECE OF BLACK PLASTIC WAS NOTED ON THE STERILE FIELD. THIS WAS NOTICED DURING CLOSURE OF THE LEFT KNEE. THE RIGHT KNEE WAS ALREADY CLOSED AND DRESSED. A BRIEF SEARCH FOR THE SOURCE OF THE PIECE OF BLACK PLASTIC REVEALED THAT THE BLACK POLY FROM THE HEAD OF THE LOCKING FEMORAL IMPACTOR HAD CRACKED DOWN ITS CENTER BETWEEN THE TWO PEGS THAT HOLD THE POLY IN PLACE ON THE HANDLE. A SMALL CHIP HAD CRACKED OFF THE CENTER CRACK AND FALLEN LOOSE. ALL PIECES WERE COLLECTED, WASHED, AND ACCOUNTED FOR. THE AGENT WAS PRESENT DURING THE SURGICAL PROCEDURE. THERE WAS NO DELAY TO SURGERY AND NO PATIENT ADVERSE EVENT. THE AGENCY IN INACTIVE AND NO MORE INFORMATION IS AVAILABLE FOR THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED FOR ANALYSIS. ENG EVAL ATTACHED JUN 2019 BY (B)(4). THE BROKEN IMPACTION SURFACE APPEARS CONSISTENT WITH MATERIAL DEGRADATION ASSOCIATED WITH NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. ALL OTHER ASPECTS OF THE DEVICES APPEAR UNREMARKABLE AND CONSISTENT WITH NORMAL USE. THE DESIGN OF THIS SPECIAL INSTRUMENT WAS DEVELOPED BY MODIFYING A PREDICATE LOCKING LPI FEMORAL IMPACTOR (CATALOG #: 213-64-01). THE FIRST REVISION ((B)(4)) REPLACED THE SCREWS USED TO HOLD THE PLASTIC IMPACTION CUSHION TO THE METAL BASE WITH TWO SMOOTH METAL PEGS THAT ARE STABILIZED BY DOWEL PINS. REFER TO (B)(4). THE SECOND MODIFICATION ((B)(4)) REPLACED THE STOP ON THE END OF THE ADJUSTMENT SCREW WITH A WELD. REFER TO TM 2017 0680 SI. DESIGN-RELATED ISSUES: THE DESIGN OF THE LOCKING LPI FEMORAL IMPACTOR HAS BEEN IN THE FIELD SINCE 2005. (B)(4). MFG-RELATED ISSUES: EXACTECH COMPLAINT DATA FROM 2014 TO 2018 INVOLVING THE REPORTED FAILURE FOR THIS FAMILY OF DEVICES WAS REVIEWED. THE INVESTIGATIONS HAVE NOT IDENTIFIED ANY EVIDENCE THAT A MANUFACTURING ISSUE CAUSED OR CONTRIBUTED TO THIS REPORTED ISSUE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY UNDERLYING CAUSE-THE BROKEN INSTRUMENT REPORTED IN EXPERIENCE (B)(4) WAS LIKELY THE RESULT OF MATERIAL DEGRADATION ASSOCIATED WITH NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. THE INSTRUMENT BROKE AND CAUSED A SMALL "CHIP" TO FALL INTO THE STERILE FIELD. AN INVESTIGATION WAS CONDUCTED; THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF MATERIAL DEGRADATION ASSOCIATED WITH NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10051 | 2022-04-07 | 2018-01-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | FEMORAL IMPACTOR | EXACTECH, INC. |  | 201-64-00 |  | IT WAS REPORTED FROM THE US THAT DURING A REVISION PROCEDURE AN INSTRUMENT BROKE. THE SURGEON USED A FEMORAL IMPACTOR TO HELP REMOVE THE OLD PROSTHESIS DURING THIS PROCESS THE BLACK PLASTIC HEAD BROKE OFF. IT WAS STATED THAT NO PIECES WERE LEFT IN THE PATIENT. THE SURGEON USED THE FEMORAL IMPACTOR TO REMOVE THE OLD PROSTHESIS BECAUSE HE NEEDED TO FORCIBLY LOOSEN THE PROSTHESIS FROM THE CEMENT; HE DECIDED THIS WOULD BE BEST FOR THIS SITUATION. HE SUCCESSFULLY REMOVED THE PROSTHESIS AND FINISHED THE PROCEDURE. THE PATIENT'S HEALTH WAS STABLE LEAVING THE OR. THE AGENCY IS INACTIVE, NO ADDITIONAL INFORMATION IS AVAILABLE FOR THIS PATIENT OR EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE WAS RECEIVED FOR ANALYSIS. ENG ANALYSIS COMPLETED 07/30/2018 BY MB THE FRACTURED AND DEFORMED IMPACTION SURFACE APPEARS CONSISTENT WITH DEGRADATION OF THE PLASTIC MATERIAL DUE TO YEARS OF NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE AND CONSISTENT WITH YEARS OF USE. DESIGN: THE OCCURRENCE RATE IS LESS THAN (B)(4), "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THIS DOES NOT APPEAR TO BE DESIGN RELATED. MANUFACTURING: THERE HAVE NOT BEEN ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT, WHICH HAS BEEN IN THE FIELD SINCE 2000. THIS DOES NOT APPEAR TO BE MANUFACTURING RELATED. MOST LIKELY UNDERLYING CAUSE: THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF YEARS OF NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES, WHICH LED TO DEGRADATION AND FRACTURE OF THE BLACK, PLASTIC IMPACTION COMPONENT. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. IT WAS STATED THAT NO PIECES WERE LEFT IN THE PATIENT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF YEARS OF NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES, WHICH LED TO DEGRADATION AND FRACTURE OF THE BLACK, PLASTIC IMPACTION COMPONENT. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10059 | 2022-04-07 | 2019-06-04 | I, n, j, u, r, y | Manufacturer report | EXACTECH | RETROVERSION HNDL | EXACTECH, INC. |  | 301-03-10 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THAT A SURGEON HAD AN ISSUE WITH A RETROVERSION HANDLE. THE RETROVERSION HANDLE THREADED TIP BROKE. NOTHING WAS LEFT IN THE PATIENT. THE PATIENT WAS DISCHARGED HOME WITHOUT ANY COMPLICATIONS. NO DELAY TO SURGERY AND NO PATIENT ADVERSE EVENT. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT FOLLOWING MULTIPLE ATTEMPTS. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE(S) NOT AVAILABLE FOR EVALUATION. (B)(4). FINDINGS: IN THE PREVIOUS COMPLAINTS, THE THREADED PORTION OF THE HANDLE SHEARED OFF FROM THE REST OF THE HANDLE. THE FRACTURE SURFACE APPEARED CONSISTENT WITH BEING EXPOSED TO A SHEAR FORCE, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FAILURE OF THE DEVICE AT THE START OF THE THREADS. THE SAME FAILURE LIKELY OCCURRED IN THIS CASE. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS RETROVERSION HANDLE, 750-2002-039- RMR-INSTRUMENT ADDITIONS REV G, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 19. MOST LIKELY UNDERLYING CAUSE: THE BROKEN RETROVERSION HANDLE REPORTED IN EXPERIENCE CASE(B)(4) WAS LIKELY THE RESULT OF IMPACTING THE STRIKE SURFACE OF THE BROACH HANDLE WITH THE THREADS OF THE RETROVERSION HANDLE PARTIALLY ENGAGED, ALLOWING FOR THE APPLIED FORCE TO BE TRANSMITTED TO THE THREADED TIP OF THE DEVICE, LEADING TO ULTIMATE FAILURE. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THE RETROVERSION HANDLE THREADED TIP BROKE. NOTHING WAS LEFT IN THE PATIENT. AN INVESTIGATION WAS CONDUCTED: THE BROKEN RETROVERSION HANDLE REPORTED WAS LIKELY THE RESULT OF IMPACTING THE STRIKE SURFACE OF THE BROACH HANDLE WITH THE THREADS OF THE RETROVERSION HANDLE PARTIALLY ENGAGED, ALLOWING FOR THE APPLIED FORCE TO BE TRANSMITTED TO THE THREADED TIP OF THE DEVICE, LEADING TO ULTIMATE FAILURE. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10048 | 2022-04-07 | 2017-11-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | HUDSON CANNULATED SHAFT | EXACTECH, INC. |  | 315-26-00 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGICAL PROCEDURE THE "COLLETED" ATTACHMENT PORTION OF THE SHAFT BROKE/DISASSEMBLED WHILE THE SURGEON WAS ATTEMPTING TO REAM THE GLENOID BONE. THE AGENT STATED: "THE SHAFT "BROKE/BECAME DISASSEMBLED", TO MY RECOLLECTION, THE FOLLOWING PARTS WERE PASSED TO THE SURGICAL TECH. - THE MODULAR REAMER, THE SHAFT, THE "COLLET", A SPRING, A SMALL METAL BALL AND POSSIBLY ONE OTHER SEPARATE PIECE". THE AGENT WAS PRESENT AT SURGERY AND THERE WAS NO DELAY TO THE PROCEDURE. IT WAS REPORTED THAT NO PIECES WERE LEFT IN THE PATIENT. THE ISSUE WAS RESOLVED BY USING ANOTHER MODULAR CANNULATED TRI-DRIVE WHICH WAS PRESENT IN THE SET OF INSTRUMENTS. THERE WAS NO PATIENT INJURY, ADVERSE EVENT, OR CLINICAL CONSEQUENCE TO THE PATIENT. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED FOR ANALYSIS. ENG EVAL COMPLETED BY CY. MFG DATE: 01/10/2011. DESIGN-RELATED ISSUES: THE HUDSON CANNULATED SHAFT HAS BEEN IN THE FIELD SINCE 2013. EXACTECH IS AWARE OF 1 OTHER COMPLAINT REPORTS INVOLVING A HUDSON CANNULATED SHAFT REPORTED AS "BROKEN/ NON-FUNCTIONAL" SINCE 2014. SALES DATA FOR CAGE GLENOIDS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS AWARE OF RECEIVING ONE OTHER COMPLAINT REPORT INVOLVING ANOTHER PART FROM THE MANUFACTURING LOT OF 5 (HUDSON CANULATED SHAFT) PIECES THAT HAVE BEEN IN THE FIELD SINCE 2013. A REVIEW OF THE DEVICE HISTORY RECORD FOR THE HUDSON CANULATED SHAFT SHOWED THAT THE NAMED DEVICES WERE ACCEPTED WITH CONFORMANCE TO THE DEVICE REQUIREMENTS. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISKS WERE INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS HUDSON CANNULATED SHAFT, 750-2007-026-RMR-INSTRUMENTS REV K WAS REVIEWED. THE RISK IS CAPTURED IN LINE 18. MOST LIKELY CAUSE: THE BROKEN HANDLE REPORTED IN EXPERIENCE C2017-659 WAS LIKELY THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY AND ALLOWING THE TIP TO DISASSEMBLE. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES. THERE WAS NO PATIENT ADVERSE EVENT TO THE PATIENT. THE SMALL PIECES WERE PASSED OUT OF THE STERILE FIELD AND NO PARTS FELL INTO THE SURGICAL SITE. AN INVESTIGATION WAS CONDUCTED; THE BROKEN HANDLE REPORTED WAS LIKELY THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY AND ALLOWING THE TIP TO DISASSEMBLE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00391 | 2022-04-08 | 2022-03-18 | I, n, j, u, r, y | Manufacturer report | CC TIBIAL INSERT SZ 4, 15MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | CC TIBIAL INSERT SZ 4, 15MM | 208-24-15 |  | AS REPORTED, APPROXIMATELY FIVE YEARS POST INITIAL LEFT TKA, THIS MALE PATIENT HAD A REVISION DUE TO KNEE FAILED AND INSTABILITY. THERE WAS POLY WEAR AND OSTEOLYSIS PER VIEWING THE POLY AND SURGEON INPUT. PATIENT WAS REVISED TO COMPETITOR'S HINGE. THERE WAS NO BREAKAGE OF DEVICE. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PHOTOS AND X-RAYS RECEIVED. THE DEVICE WILL BE RETURN. | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION RECEIVED THE FOLLOWING SECTIONS D9, G3 AND H3 HAVE BEEN UPDATED ACCORDINGLY. ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT SECTION D1: BRAND NAME SECTION D4: MODEL NUMBER, CATALOG NUMBER, SERIAL NUMBER, EXPIRATION DATE, UNIQUE IDENTIFIER (UDI) #. SECTION H4: DEVICE MANUFACTURE DATE. SECTION D10: CONCOMITANT MEDICAL PRODUCTS. SCREW, OFFSET, SZ GG (CAT# 208-04-75 / SERIAL# (B)(6)). | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00390 | 2022-04-08 | 2022-03-21 | I, n, j, u, r, y | Manufacturer report | LOGIC CR TIB INSERT STD, SZ 5, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC CR TIB INSERT STD, SZ 5, 9MM | 02-012-47-5009 |  | AS REPORTED, APPROXIMATELY ONE YEAR POST INITIAL LEFT TKA, THIS(B)(6) MALE PATIENT HAD A REVISION DUE TO INCREASE WEAR OF THE POLYETHYLENE INSERT. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. SALES REP WAS UNABLE TO OBTAIN PHOTOS/X-RAYS. THE DEVICE WILL BE RETURN. | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION RECEIVED THE FOLLOWING SECTIONS D9, G3, G6, H2, H3, H6 AND H9 HAVE BEEN UPDATED ACCORDINGLY. (H3) THE REVISION REPORTED MAY HAVE BEEN THE RESULT OF MALIGNMENT BETWEEN THE FEMORAL AND TIBIAL COMPONENTS, THIRD BODY WEAR, PATIENT-RELATED CONDITIONS, OR ANY COMBINATION OF THESE POSSIBILITIES, WHICH LED TO POLYETHYLENE WEAR OF THE TIBIAL INSERT. ADDITIONALLY, A CONTRIBUTING FACTOR TO THE OBSERVED WEAR MAY HAVE BEEN THE RESULT OF BEING PACKAGED IN A NON-CONFORMING VACUUM BAG FOR MORE THAN FIVE YEARS. THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE REPORTED EVENT OF "PROSTHESIS WEAR¿ IS ASSOCIATED WITH MATERIAL DAMAGE TO A SURFACE, USUALLY INVOLVING PROGRESSIVE LOSS OR DISPLACEMENT OF MATERIAL, DUE TO RELATIVE MOTION BETWEEN THAT SURFACE AND A CONTACTING SUBSTANCE OR SUBSTANCES. | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00388 | 2022-04-08 | 2022-03-09 | I, n, j, u, r, y | Manufacturer report | EQUINOX | SQUARE TORQUE DEFINE SCREW DRIVE KIT | EXACTECH, INC. | 300-20-02 | 300-20-02 |  | IT WAS REPORTED THAT THIS PATIENT WAS REVISED DUE TO THE SURGEON IMPLANTING AN ANATOMIC SCREW IN A REVERSE TRAY. ANATOMIC TORQUE SCREWS WERE THEN USED WITH REVERSE HUMERAL TRAYS (ONE +10 AND ONE +15). | CONCOMITANT DEVICE(S): 320-42-00, S338526 - EQUINOXE REVERSE 42MM HUMERAL LINER +0. 320-10-15, 3777335 - HUMERAL TRAY +15MM. 300-01-07, 7160740 - EQUINOXE, HUMERAL STEM PRIMARY, PRESS FIT 7MM. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00389 | 2022-04-08 | 2022-03-09 | I, n, j, u, r, y | Manufacturer report | EQUINOX | SQUARE TORQUE DEFINE SCREW DRIVE KIT | EXACTECH, INC. | 300-20-02 | 300-20-02 |  | IT WAS REPORTED THAT THIS PATIENT WAS REVISED DUE TO THE SURGEON IMPLANTING AN ANATOMIC SCREW IN A REVERSE TRAY. ANATOMIC TORQUE SCREWS WERE THEN USED WITH REVERSE HUMERAL TRAYS (ONE +10 AND ONE +15). | CONCOMITANT DEVICE(S): 308-02-06, 6962186 - 6X120MM DISTAL STEM MODULAR CEMENTED. 308-05-17, 6486068 - DISTAL FIXATION RING HA 17.5. 308-08-00, 6842253 - XS PROX BODY +0. 308-15-01, 7037402 - TAPER LOCKING SCREW 0. 300-21-00, 6248129 - 0MM FIXED ANGLED KIT. 320-42-10, 5874247 - EQUINOXE REVERSE 42MM HUMERAL CONST LINER +0. 320-10-10, 7045815 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +10. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00412 | 2022-04-11 | 2022-01-13 | I, n, j, u, r, y | Manufacturer report | ACUMATCH GXL 15DEG LINER 28MM SZ G | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | ACUMATCH GXL 15DEG LINER 28MM SZ G | 132-28-27 |  | AS REPORTED, APPROXIMATELY SEVENTEEN YEARS POST INITIAL RIGHT THA, THIS (B)(6) MALE PATIENT PRESENTED WITH APPARENT HIP LINER THAT WEAR LEADING TO INSTABILITY. PATIENT HAS A SUCCESSFUL REVISION SURGERY. PATIENT HIP LINER AND FEMORAL HEAD WAS REMOVED AND REPLACED. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PHOTOS AND X-RAYS RECEIVED. THE DEVICE WILL BE RETURN. | (H3) THE REVISION REPORTED WAS LIKELY THE RESULT OF LONG-TERM PROSTHESIS WEAR SECONDARY TO BEING IMPLANTED FOR 17 YEARS. THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE REPORTED EVENT OF "PROSTHESIS WEAR¿ IS ASSOCIATED WITH MATERIAL DAMAGE TO A SURFACE, USUALLY INVOLVING PROGRESSIVE LOSS OR DISPLACEMENT OF MATERIAL, DUE TO RELATIVE MOTION BETWEEN THAT SURFACE AND A CONTACTING SUBSTANCE OR SUBSTANCES. | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00411 | 2022-04-11 | 2015-01-10 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 5 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 5 11MM | 02-012-35-5011 |  | AS REPORTED BY THE EXACTECH KNEE REPLACEMENT STUDY, APPROXIMATELY THREE YEARS POST INITIAL TKA, THIS (B)(6) MALE PATIENT EXPERIENCED ONGOING WORSENING PROGRESSIVELY, PAIN ANTERIORLY AND LATERALLY RIGHT KNEE. PATIENT HAD SIGNIFICANT SWELLING AND WARMTH, WORSENS WITH ACTIVITY AND LIMITING KNEE EXTENSION. PATIENT STARTED RESTRICTING HOBBIES SUCH AS DISTANCE WALKING AS MOBILITY HAS REDUCED. PATIENT WAS ADVISED ICE/ELEVATION FOR PAIN RELIEF AND NSAIDS EVENT IS RESOLVED. CLINICAL STUDY CASE REPORT FORM INDICATES THIS EVENT IS DEFINITELY RELATED TO THE DEVICE AND POSSIBLY RELATED TO THE PROCEDURE. | (H3) AS REPORTED BY THE EXACTECH KNEE REPLACEMENT STUDY, APPROXIMATELY THREE YEARS POST INITIAL TKA FOR OSTEOARTHRITIS, THIS 75 Y/O OVERWEIGHT MALE PATIENT EXPERIENCED ONGOING WORSENING PROGRESSIVELY, PAIN ANTERIORLY AND LATERALLY RIGHT KNEE. PATIENT HAD SIGNIFICANT SWELLING AND WARMTH, WORSENS WITH ACTIVITY AND LIMITING KNEE EXTENSION. PATIENT STARTED RESTRICTING HOBBIES SUCH AS DISTANCE WALKING AS MOBILITY HAS REDUCED. PATIENT WAS ADVISED ICE/ELEVATION FOR PAIN RELIEF AND NSAIDS EVENT IS RESOLVED. CLINICAL STUDY CASE REPORT FORM INDICATES THIS EVENT IS DEFINITELY RELATED TO THE DEVICE AND POSSIBLY RELATED TO THE PROCEDURE. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN OR MANUFACTURING ISSUES. THE CAUSE OF THE SUBSEQUENT PAIN CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION. THESE DEVICES ARE USED FOR TREATMENT NOT DIAGNOSIS. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00410 | 2022-04-11 | 2022-02-28 | I, n, j, u, r, y | Manufacturer report | LOGIC FEMORAL PS CEM RIGHT SZ 2.5 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC FEMORAL PS CEM RIGHT SZ 2.5 | 02-010-01-0325 |  | AS REPORTED, APPROXIMATELY ELEVEN YEARS POST INITIAL RIGHT TKA, THIS (B)(6) FEMALE PATIENT PRESENTED WITH INSTABILITY IN THE RIGHT KNEE. SURGEON REVISED THE FEMORAL COMPONENT TO A CONSTRAINED SIZE 2 RIGHT CONSTRAINED FEMUR WITH A 2/13MM TRANSITIONAL POLY. THERE WAS LITTLE OSTEOLYSIS. THERE WAS NO SURGICAL DELAY/PROLONGATION REPORTED. THE PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. IMAGES RECEIVED. SALES REP WAS UNABLE TO OBTAIN X-RAYS. THE DEVICE WILL BE RETURN. | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION RECEIVED THE FOLLOWING SECTIONS DATE REC¿D BY MFR, IF FOLLOW-UP, WHAT TYPE, DEVICE EVALUATED BY MFR AND ADVERSE EVENT PROBLEM HAVE BEEN UPDATED ACCORDINGLY. AS REPORTED, APPROXIMATELY ELEVEN YEARS POST INITIAL RIGHT TKA, THIS 80 Y/O FEMALE PATIENT PRESENTED WITH INSTABILITY IN THE RIGHT KNEE. SURGEON REVISED THE FEMORAL COMPONENT TO A CONSTRAINED SIZE 2 RIGHT CONSTRAINED FEMUR WITH A 2/13MM TRANSITIONAL POLY. THERE WAS LITTLE OSTEOLYSIS. THERE WAS NO SURGICAL DELAY/PROLONGATION REPORTED. THE PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. IMAGES ATTACHED. UNABLE TO OBTAIN X-RAYS. THE DEVICE WAS RETURNED. THERE IS NO ADDITIONAL INFORMATION. PATIENT MEDICAL HISTORY: OBESE, HIGH BLOOD PRESSURE, AND LTKA EXACTECH. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN OR MANUFACTURING ISSUES. THE CAUSE OF THE SUBSEQUENT REVISION CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION. THESE DEVICES ARE USED FOR TREATMENT NOT DIAGNOSIS. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00409 | 2022-04-11 | 2022-03-10 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 15MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 15MM | 02-012-44-4015 |  | AS REPORTED BY LEGAL TEAM, APPROXIMATELY THREE YEARS POST INITIAL LEFT TKA, THIS MALE PATIENT HAD A REVISION FOR AN UNKNOWN REASON. NO OTHER INFORMATION PROVIDED. | H3: AS REPORTED BY LEGAL TEAM, APPROXIMATELY THREE YEARS POST INITIAL LEFT TKA, THIS MALE PATIENT HAD A REVISION FOR AN UNKNOWN REASON. NO OTHER INFORMATION PROVIDED. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN ISSUES. THE CAUSE OF THE SUBSEQUENT REVISION CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION. THESE DEVICES ARE USED FOR TREATMENT NOT DIAGNOSIS. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00408 | 2022-04-11 | 2019-08-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | FIT TRAY ADJUSTABLE TIBIAL TAMP | EXACTECH, INC. |  | 213-75-00 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE INSTRUMENTATION. THE LPI FIT TIBIAL TEMP WAS BROKEN DURING IMPACTION. ALL PIECES WERE ACCOUNTED FOR, NOTHING LEFT IN PATIENT. NO PIECES/PARTS OF THE DEVICE FELL INTO THE SURGICAL SITE. THERE WAS NO DELAY TO SURGERY. THE PATIENT IS FINE. THERE WAS NO REPORTED PATIENT INJURY. THE REP WAS PRESENT AT THE TIME OF SURGERY. THE DEVICE WAS RETURNED. NO ADDITIONAL INFORMATION ABOUT THIS EVENT HAS BEEN PROVIDED BY THE CONTACTS. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PATIENT ETHNICITY AND RACE. DESIGN-RELATED ISSUES: THE ADJUSTABLE TIBIAL TAMP GUIDE HAS BEEN IN THE FIELD SINCE 2012. EXACTECH IS AWARE OF 3 SIMILAR COMPLAINT REPORTS INVOLVING BROKEN GUIDE BUTTONS ON THE COMBO TRAY ADJUSTABLE TIBIAL TAMP GUIDE SINCE 2012. BECAUSE THE TIBIAL TAMP GUIDE IS USED IN OPTETRAK LOGIC TOTAL KNEE SURGERIES, SALES DATA FOR OPTETRAK LOGIC TIBIAL FIT TRAYS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4). THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF (B)(4) UNITS, WHICH HAVES BEEN IN THE FIELD SINCE 2014. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE APPROPRIATE RISK MANAGEMENT REPORT (RMR) AND RISK ASSESSMENT AND CONTROLS REPORT (RACR) FOR THE ADJUSTABLE TIBIAL TAMP GUIDE WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. RMR-00039 REV- AND RACR-00048 REV- WERE REVIEWED. RACR-00048 REV- IS BEING UPDATED UNDER DCRSP- 19-527 TO INCLUDE THE APPROPRIATE RISK. CORRECTIVE ACTIONS ARE NOT REQUIRED THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK WILL BE CAPTURED IN THE UPDATED RACR. THE FRACTURE OF THE TIBIAL TAMP GUIDE BUTTON REPORTED MAY HAVE BEEN THE RESULT OF EITHER RETROGRADE IMPACTION OF THE TIBIAL TAMP GUIDE, NOT ENGAGING THE BUTTON PRIOR TO REMOVING THE MATING TAMP HEAD, OR A COMBINATION OF THE TWO WHICH LIKELY LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE FACE OF THE GUIDE BUTTON. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. NO PIECES/PARTS OF THE DEVICE FELL INTO THE SURGICAL SITE. AN INVESTIGATION WAS CONDUCTED: THE FRACTURE OF THE TIBIAL TAMP GUIDE BUTTON REPORTED MAY HAVE BEEN THE RESULT OF EITHER RETROGRADE IMPACTION OF THE TIBIAL TAMP GUIDE, NOT ENGAGING THE BUTTON PRIOR TO REMOVING THE MATING TAMP HEAD, OR A COMBINATION OF THE TWO WHICH LIKELY LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE FACE OF THE GUIDE BUTTON. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00407 | 2022-04-11 | 2022-03-21 | I, n, j, u, r, y | Manufacturer report | BIOLOX DELTA FEMORAL HEAD 36MM OD, -3.5MM | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | BIOLOX DELTA FEMORAL HEAD 36MM OD, -3.5MM | 170-36-93 |  | AS REPORTED, APPROXIMATELY ONE MONTHS POST INITIAL LEFT THA, THIS (B)(6) MALE PATIENT HAD A PERIPROSTHETIC HIP FRACTURE. IT IS UNKNOWN HOW IT HAPPENED. PATIENT WAS REVISED TO AN ALTEON MONOBLOC STEM AND BIOLOX DELTA HEAD WERE IMPLANTED. THERE WAS NO BREAKAGE OF DEVICES OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. SALES REP WAS UNABLE TO OBTAIN PHOTOS/X-RAYS. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | AS REPORTED, APPROXIMATELY ONE MONTH POST INITIAL LEFT THA, THIS 57 Y/O MALE PATIENT HAD A PERIPROSTHETIC HIP FRACTURE. IT IS UNKNOWN HOW IT HAPPENED. PATIENT WAS REVISED TO AN ALTEON MONOBLOC STEM AND BIOLOX DELTA HEAD WERE IMPLANTED. THERE WAS NO BREAKAGE OF DEVICES OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. SALES REP WAS UNABLE TO OBTAIN PHOTOS/X-RAYS. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. UPON REVIEW OF THE AVAILABLE INFORMATION, THERE IS NO EVIDENCE THAT THIS IS A DEVICE RELATED PROBLEM AND THERE IS NO ALLEGATION AGAINST THE DEVICE. IMPLANTATION OF A TOTAL JOINT COULD RESULT PAIN, FRACTURE, INFECTION, OR LOOSENING OF TOTAL JOINT HARDWARE. THE MOST LIKELY CAUSE OF THE REPORTED EVENT WAS A MALUNION OF A FRACTURE WHICH WAS LIKELY CAUSED BY OVER-REAMING DURING IMPLANTATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00406 | 2022-04-11 | 2019-08-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | MINI HEADED FIXATION PIN REPLACED BY 201-78-15 | EXACTECH, INC. |  | 201-78-05 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE INSTRUMENTATION. THE HOLDING PIN HEAD WAS BROKEN WHILE PULLING WITH THE PIN PULLER. ALL PIECES WERE ACCOUNTED FOR, NOTHING LEFT IN PATIENT. NO PIECES/PARTS OF THE DEVICE FELL INTO THE SURGICAL SITE. THERE WAS NO DELAY TO SURGERY. THE PATIENT IS FINE. THERE WAS NO REPORTED PATIENT INJURY. THE REP WAS PRESENT AT THE TIME OF SURGERY. THE DEVICE WAS RETURNED. NO ADDITIONAL INFORMATION ABOUT THIS EVENT HAS BEEN PROVIDED BY THE CONTACTS. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020. EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. FINDINGS: THE HEAD OF THE HOLDING PIN IN THIS CASE APPEARS TO HAVE A FRACTURE ON ONE SIDE. THE FRACTURED HEAD APPEARS CONSISTENT WITH APPLYING A BENDING MOMENT TO THE PROXIMAL HEAD OF THE PIN WITH THE PIN PULLER INSTRUMENT. TENSION WAS LIKELY APPLIED TO THE PIN PULLER AT AN ANGLE RATHER THAN AXIALLY AS INTENDED. PER THE EXPERIENCE REPORT, ALL PIECES WERE ACCOUNTED FOR AND NOTHING LEFT IN THE PATIENT. THERE WAS NO REPORTED PATIENT INJURY. MOST LIKELY UNDERLYING CAUSE: BASED UPON PREVIOUS SIMILAR INVESTIGATIONS, THE BROKEN PIN REPORTED WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE HEAD OF THE PIN WITH THE PIN PULLER WHICH ALLOWED THE PROXIMAL RIM TO FRACTURE. RMR REVIEW: THE CURRENT RMR AND RACR FOR THIS DEVICE, RMR-00039 REV- AND RACR-00051 REV- WERE REVIEWED. RACR-00051 WILL BE UPDATED UNDER DCRSP-19-527 TO APPROPRIATELY CAPTURE THE RISK. IFU 700-096-181: INSTRUMENT INSPECTION VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. NO PIECES/PARTS OF THE DEVICE FELL INTO THE SURGICAL SITE. AN INVESTIGATION WAS CONDUCTED: BASED UPON PREVIOUS SIMILAR INVESTIGATIONS, THE BROKEN PIN REPORTED WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE HEAD OF THE PIN WITH THE PIN PULLER WHICH ALLOWED THE PROXIMAL RIM TO FRACTURE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00405 | 2022-04-11 | 2019-08-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRUL FEM TRIAL EXTRACTOR | EXACTECH, INC. |  | 02-029-19-1200 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE INSTRUMENTATION. THEN THE FEMORAL TRIAL EXTRACTOR BROKE AS THE PHYSICIAN WAS REMOVING THE FEMORAL TRIAL. IT APPEARS THE SPRING BROKE. NOTHING WAS LEFT IN THE PATIENT. THE PATIENT WAS STABLE AS THEY LEFT THE OR. THIS DID NOT CAUSE A SIGNIFICANT DELAY TO THE CASE. IT WAS LESS THAN 5 MINUTES. THERE WAS NO REPORTED PATIENT INJURY. THE REP WAS PRESENT AT THE TIME OF THE SURGERY. THE DEVICES WERE NOT RECEIVED IN COMPLAINTS AND THE REP COULD NOT PROVIDE ANY TRACKING INFORMATION. INACTIVE AGENCY WITH NO HOSPITAL/OTHER CONTACT INFORMATION FOR THE EVENT. NO ADDITIONAL INFORMATION. | H11. AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION.THIS MEDWATCH IS BEING RESCINDED: UPON FURTHER REVIEW OF REPORTABILITY- IF THE MALFUNCTION WERE TO RECUR, IT IS NOT LIKELY TO CAUSE OR CONTRIBUTE TO SERIOUS INJURY OR DEATH. THE DISASSEMBLED PARTS ARE NOT LIKELY TO FALL INTO THE WOUND SITE OR INTERFERE WITH THE SURGICAL IMPLANT PROCEDURE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00403 | 2022-04-11 | 2022-03-10 | I, n, j, u, r, y | Manufacturer report | FEMORAL COMPONENTS | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | FEMORAL COMPONENTS | UNKNOWN |  | AS REPORTED, SURGEON REMOVED OLD POLY OF THIS (B)(6) MALE PATIENT AND REPLACED WITH NEW OF SAME SIZE. PATIENT TOLERATED WELL AND IS STABLE. PATIENT'S REMOVED IMPLANT WAS SAVED BY PHYSICIAN FOR PATIENT AT PATIENT'S REQUEST. | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION RECEIVED THE FOLLOWING SECTIONS G3, G6, H2, H3 AND H6 HAVE BEEN UPDATED ACCORDINGLY. (H3) AS REPORTED, SURGEON REMOVED OLD POLY OF THIS 66 Y/O MALE PATIENT AND REPLACED WITH NEW OF SAME SIZE. PATIENT TOLERATED WELL AND IS STABLE. THE REMOVED IMPLANT WAS SAVED BY PHYSICIAN FOR PATIENT AT PATIENT'S REQUEST. NO ADDITIONAL INFORMATION IS AVAILABLE. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN OR MANUFACTURING ISSUES. THE CAUSE OF THE SUBSEQUENT REVISION CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION. REASON NOT REPORTED. THESE DEVICES ARE USED FOR TREATMENT NOT DIAGNOSIS. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00402 | 2022-04-11 | 2019-03-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | MDLR CNLTD TRIDRV DRVR | EXACTECH, INC. |  | 315-25-00 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE CANNULATED TRI-DRIVE DRIVER. THE PROXIMAL PIN OF THE HANDLE BROKE DURING THE BEGINNING OF REAMING. NO INCIDENT TO PATIENT. THE DEVICE WAS RETURNED, BUT NOT RETURNED TO THE COMPLAINTS DEPARTMENT, AND WAS UNABLE TO BE LOCATED FOR ENGINEERING ANALYSIS. NEW INFORMATION RECEIVED 17 AUG 2021. THE INSTRUMENT COULD HAVE BEEN BROKEN DURING THE SURGERY. HOWEVER; THE BREAKAGE HAS BEEN NOTICED AT THE RECOVERY OF THE INSTRUMENTATION-SET BY OUR LOGISTICS DEPT. AFTER THE CLEANING REALIZED BY THE STERILIZATION TEAM OF THE HOSPITAL. NO INFORMATION HAS BEEN RECEIVED FROM THE HOSPITAL REGARDING THIS ISSUE AND [THE AGENCY] DIDN¿T RECEIVE ANY REQUEST FOR COMPLEMENTARY INFORMATION REGARDING THIS ISSUE AS WELL. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED ABOUT THE EVENT FROM THE CONTACTS. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DESIGN-RELATED ISSUES: THE DESIGN OF THE TRI-DRIVE HANDLE HAS BEEN IN THE FIELD SINCE 2009. EXACTECH IS AWARE OF 10 SIMILAR COMPLAINT REPORTS INVOLVING REPORTED BROKEN TRI-DRIVE HANDLES DUE TO PIN DISASSOCIATION SINCE 2009. BECAUSE THE TRI-DRIVE HANDLE IS USED IN BOTH ANATOMIC AND REVERSE TOTAL SHOULDER SURGERIES, SALES DATA FOR HUMERAL STEMS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THESE MANUFACTURING LOTS OF (B)(4) UNITS, WHICH HAVE BEEN IN THE FIELD SINCE 2016 (74360036) AND 2017 (9201001). THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) AND RISK ASSESSMENT AND CONTROLS REPORT (RACR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR AND RACR FOR THESE TRI-DRIVE HANDLES, RMR-00007 REV ¿ AND RACR-00010 REV -, WERE REVIEWED. THE RISK IS CAPTURED IN RISK ID 9.5. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR AND RACR. THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF THE DOWEL PIN DISASSOCIATING FROM THE ASSEMBLY, WHICH MAY HAVE ALLOWED THE REAR CONNECTION COMPONENT TO DISASSEMBLE. HOWEVER; THIS CANNOT BE CONFIRMED AS THE DEVICES WERE NOT AVAILABLE FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT FROM THE BROKEN INSTRUMENT. AN INVESTIGATION WAS CONDUCTED: THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF THE DOWEL PIN DISASSOCIATING FROM THE ASSEMBLY, WHICH MAY HAVE ALLOWED THE REAR CONNECTION COMPONENT TO DISASSEMBLE. HOWEVER; THIS CANNOT BE CONFIRMED AS THE DEVICES WERE NOT AVAILABLE FOR EVALUATION | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00404 | 2022-04-11 | 2019-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | MDLR CNLTD TRIDRV DRVR | EXACTECH, INC. |  | 315-25-00 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE CANNULATED TRI-DRIVE DRIVER. THE PROXIMAL PIN OF THE HANDLE BROKE DURING THE BEGINNING OF REAMING. NO INCIDENT TO PATIENT. THE DEVICE WAS RETURNED, BUT NOT RETURNED TO THE COMPLAINTS DEPARTMENT, AND WAS UNABLE TO BE LOCATED FOR ENGINEERING ANALYSIS. NEW INFORMATION RECEIVED 17 AUG 2021: ¿THE INSTRUMENT COULD HAVE BEEN BROKEN DURING THE SURGERY. HOWEVER, THE BREAKAGE HAS BEEN NOTICED AT THE RECOVERY OF THE INSTRUMENTATION-SET BY OUR LOGISTICS DEPT., AFTER THE CLEANING REALIZED BY THE STERILIZATION TEAM OF THE HOSPITAL. NO INFORMATION HAS BEEN RECEIVED FROM THE HOSPITAL REGARDING THIS ISSUE AND [THE AGENCY] DIDN¿T RECEIVE ANY REQUEST FOR COMPLEMENTARY INFORMATION REGARDING THIS ISSUE AS WELL.¿ NO ADDITIONAL INFORMATION HAS BEEN PROVIDED ABOUT THE EVENT FROM THE CONTACTS. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO.; UNABLE TO LOCATE FOR EVALUATION. DESIGN-RELATED ISSUES: THE DESIGN OF THE TRI-DRIVE HANDLE HAS BEEN IN THE FIELD SINCE 2009. (B)(4). THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR AND RACR. THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF THE DOWEL PIN DISASSOCIATING FROM THE ASSEMBLY, WHICH MAY HAVE ALLOWED THE REAR CONNECTION COMPONENT TO DISASSEMBLE. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICES WERE NOT AVAILABLE FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT FROM THE BROKEN INSTRUMENT. AN INVESTIGATION WAS CONDUCTED: THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF THE DOWEL PIN DISASSOCIATING FROM THE ASSEMBLY, WHICH MAY HAVE ALLOWED THE REAR CONNECTION COMPONENT TO DISASSEMBLE. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICES WERE NOT AVAILABLE FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00400 | 2022-04-11 | 2019-02-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | MDLR CNLTD TRIDRV DRVR | EXACTECH, INC. |  | 315-25-00 |  | IT WAS REPORTED BY EXACTECH AUSTRALIA THAT THE LOCKING MECHANISM HAS BECOME UNATTACHED AND DISMANTLED FROM THE TRI-DRIVER. THIS CONDITION WAS FOUND DURING SURGERY. NO PIECES WERE LEFT IN THE PATIENT. THERE WAS NO DELAY TO SURGERY OR ADVERSE EFFECT TO THE PATIENT. THE PATIENT¿S HEALTH WAS STABLE LEAVING THE OR. IMAGES PROVIDED. THE REPRESENTATIVE WAS PRESENT AT THE TIME OF SURGERY. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. TWO ATTEMPTS WERE MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REP, AND A RESPONSE WAS RECEIVED. PER EXACTECH AUSTRALIA ON 10 AUG 2021, NO ADDITIONAL INFORMATION IS AVAILABLE AS THE BROKEN INSTRUMENT WAS IDENTIFIED DURING A WAREHOUSE CHECK. NO ADDITIONAL INFORMATION IS AVAILABLE. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. FINDINGS: AS PART OF THE CAPA INVESTIGATION, PREVIOUS COMPLAINT EVALUATIONS REGARDING TRI-DRIVES WERE REVIEWED. DURING ONE OF THESE PREVIOUS EVALUATIONS, FOUR JAMMED TRI-DRIVES WERE DISASSEMBLED AND CONFIRMED THE PRESENCE OF METAL WEAR PARTICLES INSIDE THE SLEEVE, MOST NOTABLY IN THE SLOT WITHIN WHICH THE RETAINING RING SEATS. IT WAS ALSO FOUND THAT THE RETAINING RING IS OFTEN DEFORMED. ADDITIONALLY, AN ATTEMPT WAS MADE TO RECREATE THE FAILURE MODE IN A LAB SETTING. IT WAS CONFIRMED THAT THE SLEEVE, SPRING, AND BALL BEARING CAN DISASSEMBLE FROM THE SHAFT AS A RESULT OF HOLDING THE SLEEVE STATIONARY DURING REAMING. THE PRIMARY CAUSE OF THE DISASSEMBLY WAS CONCLUDED TO BE DEFORMATION OF THE RETAINING RING, LEADING TO SUBSEQUENT DISENGAGEMENT FROM THE SLEEVE. FURTHERMORE, THE PRESENCE OF WEAR DEBRIS WAS OBSERVED, AND IS CONSISTENT WITH PREVIOUS COMPLAINT EVALUATIONS. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: PER CAPA(B)(4), THE DISASSEMBLED TRI-DRIVE HANDLE REPORTED WAS LIKELY THE RESULT OF THE RETAINING RING RESPONSIBLE FOR MAINTAINING ASSEMBLY OF THE SLEEVE ON THE MODULAR CANNULATED TRI-DRIVE SHAFT BECOMING DEFORMED AS A RESULT OF HOLDING THE SLEEVE STATIONARY WHILE REAMING. CORRECTIVE ACTION TAKEN: THE CRC DETERMINED THAT NO FIELD ACTION WAS NECESSARY AT THAT TIME. THIS ISSUE WILL BE ADDRESSED VIA THE EQUINOXE ERGO INSTRUMENTATION REDESIGN PROJECT. IN THE NEW EQUINOXE ERGO INSTRUMENTATION, THE MODULAR CANNULATED TRI-DRIVE WILL NOT BE USED. THE ERGO INSTRUMENTATION WILL BE IMPLEMENTED AS PART OF A PHASED ROLL-OUT, FIRST IN THE UNITED STATES AND THEN TO OTHER MARKETS, BEGINNING IN MID-2019. FULL IMPLEMENTATION IS EXPECTED BY Q4 2020 (US) AND Q4 2021 (OUS). THE COMMITTEE WILL RECONVENE IF A THRESHOLD OF > 15 EVENTS OR = 1 EVENT WITH HARM IS REACHED. EVENT TRACKING WILL BE CAPTURED IN CAPA(B)(4). ACCORDING TO THE CAPA INVESTIGATION RESULTS, THE CAPA(B)(4) INVESTIGATION CAN BE CLOSED WITHOUT FURTHER CORRECTION OR CONTAINMENT OR CORRECTIVE OR PREVENTIVE ACTIONS. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. PER CAPA(B)(4), THE DISASSEMBLED TRI-DRIVE HANDLE REPORTED WAS LIKELY THE RESULT OF THE RETAINING RING RESPONSIBLE FOR MAINTAINING ASSEMBLY OF THE SLEEVE ON THE MODULAR CANNULATED TRI-DRIVE SHAFT BECOMING DEFORMED AS A RESULT OF HOLDING THE SLEEVE STATIONARY WHILE REAMING. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00399 | 2022-04-11 | 2019-02-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | A+ BROACH HANDLE LEFT | EXACTECH, INC. |  | 167-01-00 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE INSTRUMENT. THE LEFT A+ BROACH HANDLE BROKE. THE HOLDING PIN FELL OUT ONTO THE FLOOR. ALL PIECES WERE ACCOUNTED FOR. NOTHING WAS NOTED TO HAVE FALLEN INTO THE WOUND SITE AND ¿NO PIECES WERE LEFT IN THE WOUND SITE¿. ONLY A SLIGHT DELAY WITH NO ADVERSE EFFECT TO PATIENT. THE PATIENT LEFT THE ROOM IN STABLE CONDITION. THE DEVICES WERE RETURNED. THE REPRESENTATIVE WAS PRESENT AT THE TIME OF SURGERY. TWO ATTEMPTS WERE MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REP, BUT NO RESPONSES WERE RECEIVED. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DESIGN-RELATED ISSUES: THE DESIGN OF THE NOVATION ANTERIOR BROACH HANDLE HAS BEEN IN THE FIELD SINCE 2009. EXACTECH IS AWARE OF 8 SIMILAR COMPLAINT REPORTS INVOLVING BROKEN NOVATION ANTERIOR BROACH HANDLES SINCE 2009. BECAUSE THE NOVATION ANTERIOR BROACH HANDLE IS USED IN TOTAL HIP SURGERIES, SALES DATA FOR FEMORAL STEMS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THIS NOVATION ANTERIOR BROACH HANDLE WAS MANUFACTURED IN 2009. THE DESIGN OF THE DEVICE HAS SINCE BEEN UPDATED (B)(4) TO REPLACE THE PIVOT PIN THAT IS RETAINED WITH A SPIRAL RETAINING RING WITH A PIVOT RIVET THAT IS MECHANICALLY DEFORMED. ENDURANCE TESTING WAS PERFORMED ON THIS MECHANICALLY DEFORMED (SWAGED) DESIGN AND IS DOCUMENTED IN (B)(4). MFG-RELATED ISSUES: EXACTECH IS AWARE OF 4 OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 30 UNITS, WHICH HAVE BEEN IN THE FIELD SINCE 2009, BEFORE THE DESIGN WAS UPDATED. THIS DEVICE WAS MANUFACTURED BEFORE THE DESIGN WAS UPDATED. THE DEVICE HISTORY RECORD WAS REVIEWED AS PART OF ANOTHER COMPLAINT FOR THIS DEVICE, AND ALL PARTS WERE ACCEPTED WITH CONFORMANCE TO THE DEVICE SPECIFICATIONS. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS NOVATION ANTERIOR BROACH HANDLE, (B)(4) REV C, WAS REVIEWED. THE RISK IS CAPTURED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR. THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF WEAR DUE TO YEARS OF NORMAL USE, WHICH LED TO DISASSEMBLY OF THE LEVER PIVOT PIN, THE STAINLESS STEEL RETAINING RING, THE COMPRESSION SPRING, AND THE BROACH LOCKING PIN. THE DESIGN HAS BEEN UPDATED TO REPLACE THE LEVER PIVOT PIN AND STAINLESS STEEL RETAINING RING WITH A LEVER PIVOT RIVET TO IMPROVE THE ROBUSTNESS OF THE DESIGN. IFU 700-096-181: INSTRUMENT INSPECTION : VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. THE LEFT A+ BROACH HANDLE BROKE. THE HOLDING PIN FELL OUT ONTO THE FLOOR. THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF WEAR DUE TO YEARS OF NORMAL USE, WHICH LED TO DISASSEMBLY OF THE LEVER PIVOT PIN, THE STAINLESS STEEL RETAINING RING, THE COMPRESSION SPRING, AND THE BROACH LOCKING PIN. THE DESIGN HAS BEEN UPDATED TO REPLACE THE LEVER PIVOT PIN AND STAINLESS STEEL RETAINING RING WITH A LEVER PIVOT RIVET TO IMPROVE THE ROBUSTNESS OF THE DESIGN. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00398 | 2022-04-11 | 2019-01-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LOCKING FEMORAL IMPACTOR 0-5 | EXACTECH, INC. |  | 213-64-01 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE IMPACTOR. DURING TOTAL KNEE ARTHROPLASTY THE LPI LOCKING FEMORAL IMPACTOR BROKE. THIS EVENT DID NOT CAUSE A DELAY TO SURGERY NOR DID IT AFFECT THE PATIENT. THE PATIENT WAS STABLE AS THEY LEFT THE OPERATING ROOM. PER THE REPRESENTATIVE, FROM THEIR RECOLLECTION, ALL PIECES WERE ACCOUNTED FOR AND NOTHING WAS LEFT IN THE WOUND. NO INDICATION OR REPORT THAT PIECES WERE IN THE WOUND. THE DEVICE WAS RETURNED. INACTIVE AGENCY WITH NO HOSPITAL/OTHER CONTACT INFORMATION FOR THE EVENT. NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. FINDINGS: THE DISASSEMBLED IMPACTOR FACE APPEARS TO HAVE LIKELY FRACTURED DURING IMPACTION. THE RETURNED FRACTURED IMPACTOR FACE APPEARS CONSISTENT WITH MATERIAL DEGRADATION ASSOCIATED WITH YEARS OF NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. MOST LIKELY UNDERLYING CAUSE: THE FRACTURED IMPACTOR FACE REPORTED WAS LIKELY THE RESULT OF MATERIAL DEGRADATION ASSOCIATED WITH YEARS OF NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. RMR REVIEW: THE RMR FOR THIS LPI LOCKING FEMORAL IMPACTOR, (B)(4) INSTRUMENTS REV B, WAS REVIEWED. THE RISK IS CAPTURED. IFU 700-096-181: INSTRUMENT INSPECTION : VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. ALL PIECES WERE ACCOUNTED FOR AND NOTHING WAS LEFT IN THE WOUND. AN INVESTIGATION WAS CONDUCTED: THE FRACTURED IMPACTOR FACE REPORTED WAS LIKELY THE RESULT OF MATERIAL DEGRADATION ASSOCIATED WITH YEARS OF NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00397 | 2022-04-11 | 2019-01-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRULIANT PATELLAR DRILL 3-PEG | EXACTECH, INC. |  | 02-029-32-4000 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGICAL PROCEDURE A SURGEON EXPERIENCED AN ISSUE WITH A 3-PEG PATELLAR DRILL. THE DRILL GOT CAUGHT AND STRIPPED IN THE DRILL GUIDE. THE AGENTS DIDN¿T THINK THERE IS ANY PROBLEM WITH EITHER INSTRUMENT, THE RESIDENT JUST STARTED DRILLING AT THE WRONG ANGLE AND SHREDDED THE METAL. NO DELAY AND NO REPORTED ADVERSE EVENT. THE PATIENT WAS STABLE LEAVING THE OPERATING ROOM. AGENT WAS NOT PRESENT AT TIME OF SURGERY. INACTIVE AGENCY WITH NO HOSPITAL/OTHER CONTACT INFORMATION FOR THE EVENT. NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE TRULIANT PATELLAR DRILL (02-029-32-4000) BECOMES JAMMED IN THE TRULIANT ADVANCED PATELLA DRILL GUIDE (02-029-32-3000) AND THE TRULIANT STANDARD PATELLA DRILL GUIDES (02-029-32-20**). AT LEAST ONE HOLE OF ONE DRILL GUIDE WAS FOUND TO BE UNDERSIZED. TESTING WAS PERFORMED, AND IT WAS FOUND THAT THE TRULIANT ADVANCED PATELLA DRILL GUIDE AND TRULIANT STANDARD PATELLA DRILL GUIDES DO NOT LAST THROUGH AS MANY CYCLES OF MISUSE (OFF-AXIS DRILLING) AS EXPECTED. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: PER (B)(4), CAPA (B)(4)3, AND (B)(4), THE DEFORMED TRULIANT PATELLAR DRILL AND TRULIANT STANDARD PATELLA DRILL GUIDE REPORTED WERE LIKELY THE RESULT OF DRILLING OFF-AXIS. HOWEVER; FURTHER INVESTIGATION AND TESTING WILL OCCUR AS PART OF CAPA (B)(4). IFU 700-096-181: INSTRUMENT INSPECTION VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THE RESIDENT WAS DRILLING AT THE WRONG ANGLE AND SHREDDED THE METAL OF THE DRILL, NO DELAY, AND NO REPORTED ADVERSE EVENT. PER (B)(4), CAPA (B)(4), AND (B)(4), THE DEFORMED TRULIANT PATELLAR DRILL AND TRULIANT STANDARD PATELLA DRILL GUIDE REPORTED WERE LIKELY THE RESULT OF DRILLING OFF-AXIS. HOWEVER; FURTHER INVESTIGATION AND TESTING WILL OCCUR AS PART OF CAPA (B)(4). | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00396 | 2022-04-11 | 2019-01-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | MDLR CNLTD TRIDRV DRVR | EXACTECH, INC. |  | 315-25-00 |  | IT WAS REPORTED BY THE US THAT DURING A REVERSE SHOULDER ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH TRI-DRIVER. THE TRI-DRIVER WAS NOT ABLE TO HOLD ONTO THE DRILL BIT AS THE LOCKING MECHANISM SEIZED. THE PATIENT¿S HEALTH WAS STABLE LEAVING THE OR. NO PIECES WERE LEFT IN THE PATIENT. THE ENGINEERING EVALUATION DETAILS THE ABSENCE OF THE BALL BEARING FROM THE TRI-DRIVE. AGENT WAS PRESENT DURING SURGERY. THERE WAS NO DELAY TO SURGERY AND NO ADVERSE EFFECT TO PATIENT. INACTIVE AGENCY WITH NO HOSPITAL/OTHER CONTACT INFORMATION FOR THE EVENT. NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4) HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. VISUAL EVALUATION CONDITIONS/FINDINGS: THE SLEEVE OF THE DEVICE IS JAMMED AND CANNOT BE REMOVED. THIS APPEARS CONSISTENT WITH FOREIGN MATERIAL OR METAL WEAR PARTICLES BEING STUCK IN THE SLEEVE CAUSING IT TO JAM. THE BALL BEARING WAS NOT RETURNED WITH THE DEVICE; HOWEVER, THE SALES REPRESENTATIVE CONFIRMED THE PATIENT WAS STABLE LEAVING THE OPERATING ROOM AND NO PIECES WERE LEFT IN THE PATIENT. THE DEFORMATION MARKS ON THE SLEEVE HANDLE AND THE TRI­ DRIVE CONNECTION APPEARS CONSISTENT WITH COMING IN CONTACT WITH OTHER TOOLS WHILE BEING USED DURING SURGERY. ALL OTHER ASPECTS OF THE DEVICES APPEAR UNREMARKABLE AND CONSISTENT WITH YEARS OF USE. DESIGN RELATED ISSUES: THE DESIGN OF THE TRI-DRIVE CANNULATED HANDLE HAS BEEN IN THE FIELD SINCE 2009. EXACTECH IS AWARE OF 9 SIMILAR COMPLAINT REPORTS INVOLVING 16 TRI­ DRIVE CANNULATED HANDLE DUE TO ASSEMBLY DIFFICULTIES SINCE 2009. BECAUSE THE TRI­ DRIVE CANNULATED HANDLE IS USED IN ANATOMIC AND REVERSE TOTAL SHOULDERS SURGERIES, SALES DATA FOR HUMERAL STEMS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS AWARE OF 3 OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 40 UNITS, WHICH HAVE BEEN IN THE FIELD SINCE 2015. THE DEVICE HISTORY RECORD FOR THIS LOT OF TRI-DRIVE CANNULATED HANDLES WAS REVIEWED DURING A PREVIOUS COMPLAINT INVESTIGATION, AND ALL PARTS WERE ACCEPTED WITH CONFORMANCE TO THE DEVICE SPECIFICATION. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS TRI-DRIVE CANNULATED HANDLE, 750-2007-026-RMR-INSTRUMENTS REV K, WAS REVIEWED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE ASSEMBLY DIFFICULTY REPORTED WAS LIKELY THE RESULT OF BONE FRAGMENT OR OTHER FOREIGN MATERIAL BECOMING LODGED IN THE MECHANISM, PREVENTING THE SLEEVE FROM PROPERLY RETRACTING AND BECOMING JAMMED. IFU 700-096-181: INSTRUMENT INSPECTION VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT, THERE WERE NO PIECES LEFT IN THE PATIENT. THE BALL BEARING WAS MISSING FROM THE DEVICE. AN INVESTIGATION WAS CONDUCTED; THE ASSEMBLY DIFFICULTY REPORTED WAS LIKELY THE RESULT OF BONE FRAGMENT OR OTHER FOREIGN MATERIAL BECOMING LODGED IN THE MECHANISM, PREVENTING THE SLEEVE FROM PROPERLY RETRACTING AND BECOMING JAMMED. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00395 | 2022-04-11 | 2019-01-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | T-10 SCREW DRIVER | EXACTECH, INC. |  | 341-01-38 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE SCREWDRIVER. ON SCREWING IN USING THE T10 SCREWDRIVER THE VERY END OF THE SCREWDRIVER TIP BROKE OFF. NO NEGATIVE OUTCOME OF SURGERY THERE WAS A SPARE ONE WHICH WE HAD ON THE SET DUE TO A PREVIOUS REQUEST FROM THE SURGEON TO HAVE 2 X T-10. THE AGENT MET WITH THE SURGEON TO FOLLOW UP ON THE INCIDENT. THE SURGEON FEELS THAT BY NATURE OF THE STAR SHAPED SCREW DRIVER IT IS WEAK. HE DID NOT PUT ANY UNDUE PRESSURE ONTO THE SCREW DRIVER. THE AGENT STATED THE PROBLEM AROSE DURING SCREW INSERTION AND NOT FINAL TIGHTENING WHEN A TORQUE LIMITER WAS USED. IT PREDOMINANTLY AFFECTED THE LOWER HEAD SCREWS (CALCAR SCREWS) WHICH HAVE TO PENETRATE THICKER LATERAL CORTICAL BONE AND THE VERY TOP OF THE SHAFT. IT WAS CONFIRMED THAT THERE WERE NO PIECES LEFT IN THE PATIENT. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT FOLLOWING MULTIPLE ATTEMPTS. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. TWO COMPLAINTS OF A BROKEN EQUINOXE T-10 DRIVER WERE REPORTED TO EXACTECH IN JANUARY 2019. BOTH INCIDENTS INVOLVED THE SAME SURGEON AND HAD THE SAME REPORTED FAILURE. ONE OF THE DRIVERS WAS RETURNED TO EXACTECH AND IS CAPTURED. BASED ON COMMUNICATIONS, THE T-10 DRIVER BEING EVALUATED IN THIS COMPLAINT REPORTEDLY ¿CAME BACK IN THE FX TRAY [THAT WAS] COLLECTED/SWAPPED FROM HOSPITAL AND IS NOT BE AVAILABLE FOR EVALUATION. IT WAS REPORTED THAT THE PROBLEM WITH THE DRIVER OCCURRED ¿DURING SCREW INSERTION AND NOT FINAL TIGHTENING WHEN A TORQUE LIMITER WAS USED. BASED ON THIS DESCRIPTION, IT IS LIKELY THAT THE TORQUE LIMITER WAS NOT BEING USED WHEN THE DEVICE FRACTURED. DESIGN-RELATED ISSUES: THE DESIGN OF THE EQUINOXE FRACTURE PLATE T-10 DRIVER HAS BEEN IN THE FIELD SINCE 2010. EXACTECH IS AWARE OF 4 COMPLAINT REPORTS INVOLVING A BROKEN EQUINOXE FRACTURE PLATE DRIVER SINCE 2010. SALES DATA FOR EQUINOXE FRACTURE PLATES WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 40 UNITS, WHICH HAS BEEN IN THE FIELD SINCE 2016. THEREFORE; THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. MOST LIKELY CAUSE: THE FRACTURED DRIVER REPORTED WAS LIKELY THE RESULT OF NOT USING A TORQUE LIMITER, LEADING TO A TORQUE GREATER THAN THE YIELD STRENGTH OF THE MATERIAL BEING APPLIED TO THE DRIVER TIP AND SUBSEQUENTLY ULTIMATE FAILURE. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE COMPONENT WAS NOT RETURNED FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES. THE TIP OF THE SCREWDRIVER BROKE OFF. IT WAS CONFIRMED THAT THERE WERE NO PIECES LEFT IN THE PATIENT. NO REPORTED PATIENT ADVERSE EVENT. AN INVESTIGATION WAS CONDUCTED: THE FRACTURED DRIVER REPORTED WAS LIKELY THE RESULT OF NOT USING A TORQUE LIMITER, LEADING TO A TORQUE GREATER THAN THE YIELD STRENGTH OF THE MATERIAL BEING APPLIED TO THE DRIVER TIP AND SUBSEQUENTLY ULTIMATE FAILURE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00394 | 2022-04-11 | 2019-01-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | T-10 SCREW DRIVER | EXACTECH, INC. |  | 341-01-38 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE SCREWDRIVER; THE TIP BROKE OFF. THIS ISSUE DELAYED SURGERY SLIGHTLY AS THEY HAD TO SEARCH A NUMBER OF OTHER INSTRUMENT SETS IN THE HOSPITAL FOR A SUITABLE SCREWDRIVER. THERE WAS NO NEGATIVE OUTCOME TO THE PATIENT. NO REPORTED PATIENT ADVERSE EVENT. PATIENT¿S HEALTH WAS STABLE LEAVING THE OPERATING ROOM. THE AGENT WAS NOT PRESENT AT THE TIME OF SURGERY. THE AGENT STATED THE PROBLEM [..] AROSE DURING SCREW INSERTION AND NOT FINAL TIGHTENING WHEN A TORQUE LIMITER WAS USED. IT PREDOMINANTLY AFFECTED THE LOWER HEAD SCREWS (CALCAR SCREWS) WHICH HAVE TO PENETRATE THICKER LATERAL CORTICAL BONE AND THE VERY TOP OF THE SHAFT. IT WAS CONFIRMED THAT THERE WERE NO PIECES LEFT IN THE PATIENT. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT FOLLOWING MULTIPLE ATTEMPTS. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. VISUAL EVALUATION CONDITIONS/FINDINGS: THE BROKEN TIP APPEARS CONSISTENT WITH APPLYING A CLOCKWISE TORQUE GREATER THAN THE YIELD STRENGTH OF THE MATERIAL, LEADING TO SLIGHT PLASTIC DEFORMATION IN THE COUNTERCLOCKWISE DIRECTION, AND EVENTUALLY, ULTIMATE FAILURE. IT WAS REPORTED THAT THE PROBLEM OCCURRED ¿DURING SCREW INSERTION AND NOT FINAL TIGHTENING WHEN A TORQUE LIMITER WAS USED.¿ BASED ON THIS DESCRIPTION, IT IS LIKELY THAT THE TORQUE LIMITER WAS NOT BEING USED WHEN THE DEVICE FRACTURED AND IS LIKELY THE REASON EXCESS TORQUE WAS APPLIED TO THE DRIVER TIP. DESIGN-RELATED ISSUES: THE DESIGN OF THE EQUINOXE FRACTURE PLATE T-10 DRIVER HAS BEEN IN THE FIELD SINCE 2010. (B)(4). THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF (B)(4) UNITS, WHICH HAS BEEN IN THE FIELD SINCE 2015. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, THE RISK IS CAPTURED IN THE RMR AND RACR, AND THE RACR WILL BE UPDATED TO CAPTURE THE APPROPRIATE HAZARDOUS SITUATION AND HARM FOR THIS REPORTED FAILURE. MOST LIKELY CAUSE: THE FRACTURED DRIVER REPORTED WAS LIKELY THE RESULT OF NOT USING A TORQUE LIMITER, LEADING TO A TORQUE GREATER THAN THE YIELD STRENGTH OF THE MATERIAL BEING APPLIED TO THE DRIVER TIP AND SUBSEQUENTLY ULTIMATE FAILURE. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES. THE TIP OF THE SCREWDRIVER BROKE OFF. IT WAS CONFIRMED THAT THERE WERE NO PIECES LEFT IN THE PATIENT. NO REPORTED PATIENT ADVERSE EVENT. AN INVESTIGATION WAS CONDUCTED: THE FRACTURED DRIVER REPORTED WAS LIKELY THE RESULT OF NOT USING A TORQUE LIMITER, LEADING TO A TORQUE GREATER THAN THE YIELD STRENGTH OF THE MATERIAL BEING APPLIED TO THE DRIVER TIP AND SUBSEQUENTLY ULTIMATE FAILURE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00393 | 2022-04-11 | 2019-01-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRU TIBIAL ADAPTOR PLATE TRIAL/IMPLANT, SIZE 3.5 | EXACTECH, INC. |  | 02-029-25-4035 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE TIBIAL ADAPTOR PLATE. IT WAS FOUND THAT THERE WAS A BROKEN TRIAL AFTER THE TEST OF KNEE JOINT REDUCTION AND RANGE OF MOTION WAS CONDUCTED. THERE WAS NO DELAY TO SURGERY. NO REPORTED PATIENT ADVERSE EVENT. THERE IS NO PATIENT INFORMATION REPORTED. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT FOLLOWING MULTIPLE ATTEMPTS. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: A.2, A.3, A.4, A.5, A.6. COUNTRY: (B)(6). (B)(4) VARIOUS INSTANCES OF TRULIANT TIBIAL TRIAL ADAPTOR PLATE CENTRAL FLEXURE DISASSOCIATED FROM THE PLATES. FINDINGS: DURING THE INVESTIGATION, A REVIEW OF THE PROCEDURE USED FOR THE INITIAL TOLERANCE STACK-UP BETWEEN THE TRULIANT TIBIAL TRAY ADAPTER AND THE CORRESPONDING TRULIANT AND LOGICAL TIBIAL TRAYS WAS PERFORMED AND DEEMED ADEQUATE. ATTEMPTS TO REPLICATE THE FAILURE MODES MENTIONED IN SIMILAR COMPLAINTS WERE UNSUCCESSFUL UNDER WORST CASE CONDITIONS. A REVIEW OF THE DHR CONCLUDED THAT ALL PARTS WERE MADE WITHIN SPECIFICATION. WHEN CONSIDERING USE, IT WAS FOUND THAT THE DESIGN OF THE INSTRUMENT TRAY CASES MAY ALLOW THE ADAPTER PLATES TO BECOME DISLODGED FROM THEIR DESIGNATED POSITION AND BECOME DAMAGED. HUMAN ERROR MAY ALSO BE A CAUSE OF FAILURE SUCH AS USING THE INCORRECT SIZE ADAPTER PLATE OR IMPACTING THE FLEXURE. THE DESIGN LENDS TO BEING SUSCEPTIBLE TO BREAKAGE IF USED INCORRECTLY DUE TO THE USE OF ULTEM MATERIAL WITH A THINNER ASPECT UNDER LOAD. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: THE ROOT CAUSE ASSOCIATED WITH (B)(4) RESIDES IN THE MISUSE/MISTREATMENT OF THE ADAPTOR PLATES OR CANNOT BE IDENTIFIED DUE TO INSUBSTANTIAL EVIDENCE (REPROCESSING RELATED COMPLAINTS, (B)(4). THE CAPA TEAM HAS DETERMINED, HOWEVER, THAT USE CONSIDERATIONS INDICATE THAT THE FLEXURE DESIGN IS SUSCEPTIBLE TO DAMAGE FROM SAID MISUSE/MISTREATMENT. CORRECTIVE ACTION TAKEN: THE DESIGN OF THE ADAPTER PLATE HAS BEEN CHANGED. REFERENCE (B)(4). THE DESIGN OF THE INSTRUMENT TRAY HAS ALSO BEEN UPDATED. REFERENCE (B)(4). IFU 700-096-181: INSTRUMENT INSPECTION VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. NO PATIENT ADVERSE EVENT REPORTED FROM THE BROKEN DEVICE. THE ROOT CAUSE ASSOCIATED WITH CAPA2017-53 RESIDES IN THE MISUSE/MISTREATMENT OF THE ADAPTOR PLATES OR CANNOT BE IDENTIFIED DUE TO INSUBSTANTIAL EVIDENCE. THE CAPA TEAM HAS DETERMINED, HOWEVER, THAT USE CONSIDERATIONS INDICATE THAT THE FLEXURE DESIGN IS SUSCEPTIBLE TO DAMAGE FROM SAID MISUSE/MISTREATMENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00392 | 2022-04-11 | 2018-12-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | MDLR CNLTD TRIDRV DRVR | EXACTECH, INC. |  | 315-25-00 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE GLENOID REAMER. THE SURGEON WAS REAMING THE GLENOID AND THE REAMER BECAME UNATTACHED TO THE HANDLE. THEY WERE ABLE TO FIT IT. OPENED A BACKUP AND CONTINUED WITHOUT INCIDENT. ONLY A 3-MINUTE DELAY THERE WAS NO ADVERSE EVENT TO THE PATIENT. NO PIECES FELL INTO THE SURGICAL SITE. ALL PIECES WERE ACCOUNTED FOR WITHOUT THE DOCTOR HAVING TO REMOVE ANY OF THE PIECES. THE PATIENT OVERALL PROCEDURE OUTCOME/RESULT WAS NOTED TO BE GOOD. NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. THE DESIGN OF THE TRI-DRIVE MODULAR REAMER DRIVER HAS BEEN IN THE FIELD SINCE 2009. (B)(4). THE DISASSEMBLED TRI-DRIVE MODULAR REAMER DRIVER MAY RESULT IN A DELAY IN SURGERY TO FIND A NEW ONE. TYPICALLY, IN ANATOMIC SHOULDER CASES BOTH ANATOMIC AND REVERSE SETS ARE IN THE OR. IF THE TRI-DRIVE WERE TO BREAK, A SECOND TRI-DRIVE WOULD BE IN THE REVERSE SHOULDER SET. THE CRC DETERMINED THAT NO FURTHER ACTION WAS REQUIRED AT THIS TIME AS THIS ISSUE WILL BE ADDRESSED VIA THE EQUINOXE ERGO INSTRUMENTATION REDESIGN PROJECT. IN THE NEW EQUINOXE ERGO INSTRUMENTATION, THE TRI-DRIVE HANDLES WILL NO LONGER BE USED. SINCE THE DESIGN OF THE PILOT TIP REAMERS ARE CHANGING (ALL THE REAMERS ARE CANNULATED), THERE WILL NOW BE TWO DIFFERENT HANDLES: A PILOT TIP HANDLE AND A CANNULATED HANDLE, REPLACING THE TRI-DRIVE. THIS NEW ERGO INSTRUMENTATION DESIGN WILL BE IMPLEMENTED AS PART OF A PHASED ROLL-OUT, FIRST IN THE US AND THEN TO OTHER MARKETS, BEGINNING IN MID-2019. THESE BROKEN TRI-DRIVES WERE LIKELY THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING OR THE TRI-DRIVE CONTACTING A RETRACTOR, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY, ALLOWING THE TIP TO DISASSEMBLE. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT OR SURGICAL DELAY. NO PIECES FELL INTO THE SURGICAL SITE. ALL PIECES WERE ACCOUNTED FOR WITHOUT THE DOCTOR HAVING TO REMOVE ANY OF THE PIECES. THE PATIENT OVERALL PROCEDURE OUTCOME/RESULT WAS NOTED TO BE GOOD. AN INVESTIGATION WAS CONDUCTED; THESE BROKEN TRI-DRIVES WERE LIKELY THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING OR THE TRI-DRIVE CONTACTING A RETRACTOR, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY, ALLOWING THE TIP TO DISASSEMBLE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00401 | 2022-04-11 | 2019-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | MINI HEADED FIXATION PIN REPLACED BY 201-78-15 | EXACTECH, INC. |  | 201-78-05 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE PIN. IT WAS REALIZED AFTER USE THAT THE PIN WAS BROKEN. THE PIN WAS DISCARDED AND WILL NOT BE RETURNING. THE REPRESENTATIVE WAS PRESENT AT THE TIME OF SURGERY. AN ATTEMPT WAS MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REPRESENTATIVE AND A RESPONSE WAS RECEIVED. PER THE REP, NO PARTS/PIECES OF THE DEVICE WERE REMOVED FROM THE SURGICAL SITE. NO DELAY TO SURGERY AND THERE WAS NO INJURY/CLINICAL CONSEQUENCE/ADVERSE EVENT TO THE PATIENT. THERE HAS BEEN NO ADDITIONAL INFORMATION PROVIDED FOR THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DESIGN-RELATED ISSUES: THE DESIGN OF THE MINI HEADED FIXATION PIN HAS BEEN IN THE FIELD SINCE 1999. EXACTECH IS AWARE OF 6 COMPLAINT REPORTS INVOLVING 7 BROKEN MINI HEADED FIXATION PINS SINCE 2008. BECAUSE THE MINI HEADED FIXATION PIN IS THE SAME PIN CONTAINED IN THE MINI HOLDING PIN SHARP POINT 4 PACK (201-78-15), INVENTORY NUMBERS FOR BOTH DEVICES WERE USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: MANUFACTURING-RELATED DATA COULD NOT BE REVIEWED BECAUSE NO INFORMATION REGARDING THE MANUFACTURING LOTS OR SERIAL NUMBERS WERE PROVIDED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS MINI HEADED FIXATION PIN, (B)(4) INSTRUMENTS REV D, WAS REVIEWED. THE RISK IS CAPTURED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR. THE BROKEN PIN REPORTED WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE HEAD OF THE PIN WITH THE PIN PULLER, WHICH ALLOWED FOR THE PROXIMAL RIM TO FRACTURE. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION :VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT DUE TO THE BROKEN PIN. AN INVESTIGATION WAS CONDUCTED: THE BROKEN PIN REPORTED WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE HEAD OF THE PIN WITH THE PIN PULLER, WHICH ALLOWED FOR THE PROXIMAL RIM TO FRACTURE. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00413 | 2022-04-12 | 2022-03-18 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | CAGE GLENOID LARGE, BETA | EXACTECH, INC. | 314-13-14 | 314-13-14 |  | IT WAS REPORTED THAT DURING REMOVAL, THE SURGEON FOUND THE PATIENT LEFT SHOULDER IMPLANT TO BE DAMAGED. THE GLENOID WAS MISSING A CHUNK. SURGEON ONLY REMOVED THE GLENOID FROM THE PATIENT. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICE WILL NOT BE RETUNING, SURGEON KEPT IMPLANT. | SECTION H10: (H3) THE CAUSE OF THE REVISION REPORTED CANNOT BE DETERMINED FROM THE INFORMATION PROVIDED, BUT MAY HAVE BEEN RELATED TO THE CAGE GLENOID COMPONENT, WHICH WAS REPORTEDLY FOUND TO BE ¿MISSING A CHUNK¿ DURING THE REVISION SURGERY. IT IS UNCLEAR IF THE GLENOID WAS EXCESSIVELY WORN, THE POLYETHYLENE FRACTURED, OR IF THE CAGE/PERIPHERAL PEGS DISASSOCIATED FROM THE POLYETHYLENE BODY. THE EXACT EVENTS AND ORDER OF EVENTS CANNOT BE CONFIRMED BECAUSE NO IMAGES OF THE EXPLANTED DEVICE OR RADIOGRAPHS WERE PROVIDED AND THE REVISED COMPONENT WAS NOT RETURNED TO EXACTECH FOR EVALUATION. SECTION H11: *THE FOLLOWING SECTIONS HAVE CORRECTED INFORMATION: (H6) COMPONENT CODE: 4755, PART/COMPONENT/SUB-ASSEMBLY TERM NOT APPLICABLE. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00414 | 2022-04-12 | 2022-03-23 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | PRESERVE STEM 6MM | EXACTECH, INC. | 300-30-06 | 300-30-06 |  | AS REPORTED, DURING STEM IMPACTION IN A L TOTAL SHOULDER PROCEDURE ON A (B)(6) FEMALE PATIENT, THERE WAS A CALCAR FRACTURE TO THE L SHOULDER. IT WAS REPAIRED WITH A CERCLAGE AND THE PROCEDURE WAS COMPLETED. THE PATIENT HAS A HISTORY OF OSTEOARTHRITIS. THERE IS NO DEVICE TO BE RETURNED. THE CASE REPORT FORM INDICATES THIS EVENT IS POSSIBLY RELATED TO DEVICES AND/OR POSSIBLY RELATED TO PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | SECTION H10: (H3) THE CALCAR FRACTURE WAS REPORTED TO HAVE OCCURRED DURING THE IMPLANT IMPLEMENTATION. IT LIKELY OCCURRED WHILE PREPARING THE SHOULDER FOR AND MAY HAVE BEEN RELATED TO THE FEED RATE/PRESSURE OF THE CALCAR PLANER, REAMER ROTATIONAL SPEED, OR THE PATIENT¿S BONE QUALITY/GEOMETRY. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00415 | 2022-04-12 | 2021-08-20 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REV GLENOID PLATE | EXACTECH, INC. | 320-15-01 | 320-15-01 |  | AS REPORTED, APPROXIMATELY 5 MONTHS POST OF THE INITIAL TSA, THIS (B)(6) MALE PATIENT HAS AN ACROMIAL FRACTURE (TYPE 1). HE LIFTED A HEAVY SUITCASE AWAY FROM HIS BODY AND FELT PAIN, NO REVISION HAS OCCURRED. THE CASE REPORT FORM INDICATES THIS EVENT IS DEFINITELY RELATED TO DEVICE AND UNLIKELY RELATED TO PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. OUTCOME IS CONTINUING, NO ACTION TAKEN AT THIS TIME. | UPON REVIEW OF THE AVAILABLE INFORMATION, THERE IS NO EVIDENCE THAT THIS IS A DEVICE RELATED PROBLEM AND THERE IS NO ALLEGATION AGAINST THE DEVICE. IMPLANTATION OF A TOTAL JOINT COULD RESULT PAIN, INFECTION, OR LOOSENING OF TOTAL JOINT HARDWARE. THE MOST LIKELY CAUSE OF THE REPORTED EVENT WAS A MALUNION OF A FRACTURE WHICH WAS LIKELY CAUSED BY OVER-REAMING DURING IMPLANTATION. CONCOMITANT DEVICE(S): 300-30-12 - EQUINOXE PRESERVE STEM 12MM, 320-38-03 - EQUINOXE REVERSE 38MM HUMERAL LINER +2.5, 320-10-00 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0, AND 320-15-01 - EQ REV GLENOID PLATE. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00420 | 2022-04-15 | 2022-01-03 | I, n, j, u, r, y | Manufacturer report | RBK PS TIBIAL INSERT SZ 3, 09MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | RBK PS TIBIAL INSERT SZ 3, 09MM | 264-23-09 |  | AS REPORTED, THIS (B)(6) Y/O MALE PATIENT HAD A REVISION TO RETRIEVED PE EXPLANTS THAT SHOWED IMPORTANT WEAR AND DELAMINATION. THERE WAS MASSIVE BONE LOSS OSTEOLYSIS DUE TO WEAR PARTICLES. PATIENT WAS REVISED TO OPTETRAK CC WITH TIBIAL AND FEMORAL WEDGES AS WELL AS BONE GRAFT. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. SALES REP. WAS UNABLE TO OBTAIN X-RAYS/PHOTOS. THE DEVICE WILL BE RETURN. | SECTION D10: CONCOMITANT MEDICAL PRODUCTS OPTETRAK ASY HF PS POR FEM, SZ 3, L (CAT# 246-02-03 / SERIAL# (B)(6)); THREE PEG PATELLA, 35MM (CAT# 200-02-35). (H3) THE REVISION REPORTED WAS LIKELY THE RESULT OF JOINT LAXITY/INSTABILITY, WHICH LED TO UNEVEN FORCE DISTRIBUTION ACROSS THE JOINT AND PROSTHESIS WEAR OF THE TIBIAL INSERT AND PATELLA. THE OSTEOLYSIS REPORTED WAS LIKELY THE RESULT OF THE PROSTHESIS WEAR. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE NO X-RAYS WERE PROVIDED. THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE REPORTED EVENT OF "PROSTHESIS WEAR¿ IS ASSOCIATED WITH MATERIAL DAMAGE TO A SURFACE, USUALLY INVOLVING PROGRESSIVE LOSS OR DISPLACEMENT OF MATERIAL, DUE TO RELATIVE MOTION BETWEEN THAT SURFACE AND A CONTACTING SUBSTANCE OR SUBSTANCES. FURTHERMORE, THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE REPORTED EVENT OF "INSTABILITY¿ IS ASSOCIATED WITH UNDESIRABLE STABILITY OF THE JOINT OR IMPLANT CONSTRUCT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00419 | 2022-04-15 | 2015-02-06 | I, n, j, u, r, y | Manufacturer report | PS TIBIAL INSERTS SZ 2, 18MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | PS TIBIAL INSERTS SZ 2, 18MM | 204-22-18 |  | AS REPORTED BY LEGAL TEAM, APPROXIMATELY 3.5 YEARS POST INITIAL LEFT TKA, THIS FEMALE PATIENT HAD A REVISION FOR AN UNKNOWN REASON. PATIENT WAS REVISED TO EXACTECH PARTS THAT ALSO CONFIRMED TO BE PART OF THE RECALL. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00421 | 2022-04-15 | 2022-03-24 | I, n, j, u, r, y | Manufacturer report | TRULIANT PS CEM FEM PS CEM RIGHT SZ 5 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT PS CEM FEM PS CEM RIGHT SZ 5 | 02-020-11-0350 |  | AS REPORTED, APPROXIMATELY FOUR MONTHS POST INITIAL RIGHT TKA, THIS (B)(6) Y/O MALE PATIENT HAD A REVISION FOR AN UNKNOWN REASON AND SURGEON WAS NOT SURE IF THE POLY WAS PART OF THE RECALL. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. IMAGE RECEIVED. THE DEVICE WILL BE RETURN. | CONCOMITANT MEDICAL PRODUCTS: TRULIANT TIB FIT TRAY CEM SZ 5F / 5T (CAT# 02-022-45-5050 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00417 | 2022-04-15 | 2022-03-25 | I, n, j, u, r, y | Manufacturer report | TRAPEZOID TIBIAL TRAY SZ 5F/4T | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRAPEZOID TIBIAL TRAY SZ 5F/4T | 204-04-54 |  | AS REPORTED, APPROXIMATELY THIRTEEN YEARS POST INITIAL LEFT TKA, THIS (B)(6) PATIENT WAS HAVING PAINFUL KNEE AND CAME TO SURGEON WITH DISCOMFORT AND SWELLING. UPON EVALUATION SURGEON FELT A LINER EXCHANGE WOULD HELP. INTRA OPERATIVELY SURGEON FOUND METALLOSIS AND A MAN HOLE COVER THAT CAME DISCONNECTED FROM THE TIBIA. IT WAS STUCK BETWEEN THE LINER AND PROXIMAL TIBIA TRAY. ALSO THE PATELLA HAD SOME WEAR TO IT. THERE WAS NO OSTEOLYSIS. SO SURGEON REPLACED IT. SURGEON PLACED A 5/11MM PS AND REVISED THE PATELLA TO A 38MM. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICE WILL BE RETURN. | CONCOMITANT MEDICAL PRODUCTS: THREE PEG PATELLA 35MM (CAT# 200-02-35 / SERIAL# (B)(4)) AND PS TIBIAL INSERTS SZ 5, 9MM (CAT# 204-25-09 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00416 | 2022-04-15 | 2022-03-24 | I, n, j, u, r, y | Manufacturer report | ALT HA S CLR STD SZ 3 | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | ALT HA S CLR STD SZ 3 | 190-30-03 |  | AS REPORTED, APPROXIMATELY FOUR YEARS POST INITIAL LEFT THA, THIS (B)(6) FEMALE PATIENT WAS BROUGHT BACK FOR LEFT HIP PAIN DUE TO A LOOSE STEM. THE HIP WAS DISLOCATED, THE FEMORAL HEAD WAS REMOVED, ATTENTION TURNED TO THE CUP, THE LINER WAS REMOVED WITH THE HELP OF AN OSTEOTOME, THE CUP WAS CHECKED AND FOUND TO BE WELL FIXED. A NEW XLE G1 LIPPED LINER WAS IMPLANTED. THE STEM WAS CHECKED, FOUND TO BE LOOSE AND REMOVED. A NEW STEM AND HEAD WAS THEN IMPLANTED. PATIENT WAS REVISED TO XLE G1 32MM LIPPED INSERT, A STD 7 TAPER WEDGE STEM AND A 32++0 BIOLOX HEAD. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICE WILL BE RETURN. | CONCOMITANT MEDICAL PRODUCTS: NV GXL LINER NEUTRAL, 32MM ID GROUP 1 CUPS (CAT# 130-32-51 / (B)(4)); BIOLOX DELTA FEMORAL HEAD 32MM OD, +0MM (CAT# 170-32-00 / (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00418 | 2022-04-15 | 2021-02-03 | I, n, j, u, r, y | Manufacturer report | TRULIANT TIB IMP PS INSERT SZ 4 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED, | EXACTECH, INC. | TRULIANT TIB IMP PS INSERT SZ 4 9MM | 02-022-35-4009 |  | AS REPORTED BY LEGAL TEAM, APPROXIMATELY FOUR YEARS POST INITIAL RIGHT TKA, THIS FEMALE PATIENT HAD A REVISION OF RIGHT-KNEE AND PLACEMENT OF TEMPORARY SPACER. THEN, THREE MONTHS LATER PATIENT HAD A RIGHT-REVISION TOTAL KNEE ARTHROPLASTY ON (B)(6) 2021. PATIENT ALSO HAD MANIPULATION UNDER ANESTHESIA (MUA) TWO MONTHS AFTER RIGHT KNEE REVISION ARTHROPLASTY. PATIENT IS AMBULATING WITH A WALKER, STILL EXPERIENCING SOME STIFFNESS BUT HAS NO PAIN THIS TIME. THE PATIENT HAS NO MAJOR COMPLAINTS AND IS DOING WELL OVERALL. | CONCOMITANT MEDICAL PRODUCTS: TRULIANT PS CEM FEM PS CEM RIGHT SZ 4 (CAT# 02-020-11-0340), TRULIANT TIB IMP PSC INSERT SZ 4, 11MM (CAT# 02-022-44-4011), TRULIANT TIB FIT TRAY CEM SZ 4F / 4T (CAT# 02-022-45-4040), ADVANCED PATELLA 35MM 3 PEG IMPLANT (CAT# 200-07-35), HOLDING PIN MINI SHARP POINT 4 PK (CAT# 201-78-15), COLLAR FIXATION PIN 2PK 40MM, QUICK RELEASE (CAT# 201-78-82), 3" DRILL BIT, MOD. HEX 2 PACK (CAT# 201-78-89). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00349 | 2022-04-18 | 2018-10-29 | I, n, j, u, r, y | Manufacturer report | EXACTECH | RESURFACING CAGE TAMP | EXACTECH, INC. |  | 313-05-10 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGICAL PROCEDURE A DOWL PIN DISASSEMBLED FROM THE RESURFACING CAGE TAMP SHAFT. THE SURGEON HAD A RESURFACING, AFTER USING THE TAMP FOR THE CAGE IT WAS NOTED THAT THE LOCKING CATCH WAS NOT THERE, AND THE INSTRUMENT KEPT DETACHING AS A RESULT. THE SURGEON USED A PIN TO PLACE IN THE HOLE WHERE THE CATCH WOULD BE TO BE ABLE TO REMOVE THE TAMP FROM THE PATIENT. THE PATIENT REQUIRED X RAY TO DETERMINE WHETHER THE MISSING LOCK ON THE INSTRUMENT CAME DETACHED WHILST IN THE PATIENT, THERE WAS NOTHING FOUND ON THE X RAY AND THE SURGEON WAS SATISFIED THAT THE MISSING CATCH WAS NOT IN THE PATIENT. THERE WAS A DELAY TO SURGERY WITH NO PATIENT ADVERSE EVENT. THE AGENT WAS PRESENT AT THE TIME OF SURGERY. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4) HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PATIENT AGE, GENDER, ETHNICITY AND RACE. VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR 10X OPTICAL UNLESS OTHERWISE SPECIFIED) THE RETURNED RESURFACING CAGE TAMP APPEARS TO BE MISSING THE DOWEL PIN. THIS IS LIKELY DUE TO DEGRADATION OF THE LASER WELD OVER TIME. THE PIN WAS NOT RETURNED FOR EVALUATION. DESIGN-RELATED ISSUES: THE DESIGN OF THE RESURFACING CAGE TAMP HAS BEEN IN THE FIELD SINCE 2014. EXACTECH IS NOT AWARE OF ANY SIMILAR COMPLAINT REPORTS INVOLVING BROKEN RESURFACING CAGE TAMPS SINCE 2014. THE FREQUENCY OF OCCURRENCE RANKING SCALE IS ¿VERY LOW¿. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 20 UNITS, WHICH HAVE BEEN IN THE FIELD SINCE 2015. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING-RELATED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR. THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF DEGRADATION OF THE LASER WELD OVER TIME, WHICH CAUSED THE DOWEL PIN TO DISASSEMBLE FROM RESURFACING CAGE TAMP SHAFT. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. AN X-RAY WAS TAKEN OF THE PATIENT TO ENSURE THAT THE MISSING PIN WAS NOT RETAINED IN THE PATIENT. THERE WAS NO PATIENT ADVERSE EVENT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF DEGRADATION OF THE LASER WELD OVER TIME, WHICH CAUSED THE DOWEL PIN TO DISASSEMBLE FROM RESURFACING CAGE TAMP SHAFT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00350 | 2022-04-18 | 2018-09-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | SHLDR GPS MODULAR DRIVER | EXACTECH, INC. |  | 531-25-00 |  | IT WAS REPORTED FROM THE US THAT DURING AN ANATOMIC SHOULDER SURGERY THE MODULAR DRIVER TIP SHEARED OFF WHILE SURGEON WAS REAMING GLENOID. THERE WERE NO PARTS OR PIECES FELL INTO THE SURGICAL SITE. THERE WAS NO DELAY OR ADVERSE EVENT TO THE PATIENT. THE PATIENT¿S HEALTH WAS STABLE LEAVING THE OR. PATIENT DEMOGRAPHIC INFORMATION WAS NOT PROVIDED. THE AGENT WAS PRESENT AT THE TIME OF SURGERY. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR L0X OPTICAL UNLESS OTHERWISE SPECIFIED) THE FRACTURE SURFACES APPEAR CONSISTENT WITH THE POWER CONNECTION FEATURE SHEARING OFF FROM THE MAIN SHAFT DURING GLENOID REAMING. THIS MAY HAVE OCCURRED DUE TO A POSTERIORLY ANGLED GLENOID, WHICH NECESSITATED HOLDING THE HANDLE STATIONARY WHILE ROTATING THE POWER SOURCE POSTERIORLY. THE ILLEGIBLE PART MARKINGS APPEAR CONSISTENT WITH NORMAL USE OF THE DEVICE. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE. DESIGN-RELATED ISSUES: THE DESIGN OF THE SHOULDER GPS MODULAR DRIVER HAS BEEN IN THE FIELD SINCE 2017. EXACTECH IS NOT AWARE OF ANY OTHER SIMILAR COMPLAINT REPORTS INVOLVING THIS DEVICE SINCE ITS INTRODUCTION. THE FREQUENCY OF OCCURRENCE RANKING SCALE IS CONSIDERED "VERY LOW". THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: MANUFACTURING RELATED DATA COULD NOT BE EVALUATED BECAUSE NO INFORMATION REGARDING THE MANUFACTURING LOT WAS PROVIDED. THE MANUFACTURER HAS BEEN NOTIFIED OF THIS EVENT. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE POWER SOURCE DURING GLENOID REAMING, WHICH LED TO THE POWER CONNECTION FEATURE SHEARING OFF FROM THE MAIN SHAFT. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY DESIGN ISSUES. THERE WERE NO PARTS OR PIECES FELL INTO THE SURGICAL SITE. THERE WAS NO DELAY OR ADVERSE EVENT TO THE PATIENT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE POWER SOURCE DURING GLENOID REAMING, WHICH LED TO THE POWER CONNECTION FEATURE SHEARING OFF FROM THE MAIN SHAFT. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00351 | 2022-04-18 | 2018-11-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | OFFSET CUP IMPACTOR | EXACTECH, INC. |  | 121-01-02 |  | IT WAS REPORTED FROM JAPAN THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED ISSUES WITH THE OFFSET CUP IMPACTOR. THE CLUSTER HOLDS IMPLANT, 48MM, THE DEVICE WAS ATTACHED AND IMPACTED. AFTER IMPACTION, THE SURGEON COULD NOT REMOVE THE OFFSET CUP IMPACTOR IN QUESTION SMOOTHLY FROM THE IMPLANT BECAUSE THE RELEASE LEVER ¿WAS TOO HARD UNDER USUAL STRENGTH¿. ¿SO, THE DOCTOR TRIED SEVERAL TIMES AND FINALLY HE PUT ALL HIS STRENGTH INTO HIS ARM, NOT LIKE THE USUAL WAY. HE WAS ABLE TO REMOVE IT FROM THE IMPLANT, BUT THE SCREW AND SPRING FALL OFF AT THE SAME TIME. THE CAUSE AND EFFECT ARE UNCERTAIN. ALL PIECES WERE ACCOUNTED FOR AND THIS DID NOT DELAY SURGERY. THE PATIENT¿S HEALTH WAS STABLE LEAVING THE OPERATING ROOM.¿ THE REPRESENTATIVE WAS PRESENT AT THE TIME OF SURGERY. NO PATIENT DEMOGRAPHIC INFORMATION WAS PROVIDED. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4) HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR PATIENT. VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR L0X OPTICAL UNLESS OTHERWISE SPECIFIED) - THE DAMAGE TO THE DEVICE APPEARS CONSISTENT WITH USING PLYERS OR A SIMILAR TOOL TO UNSCREW THE SCREW FROM THE MAIN ASSEMBLY OF THE DEVICE. THIS SCREW IS NOT INTENDED TO BE REMOVED AND IS INTENDED TO CONTAIN THE SPRING AND BALL. THE LOOSENESS OF THE SCREW FROM PREVIOUS REMOVAL MAY HAVE CONTRIBUTED TO THE DISASSEMBLY OF THE SCREW SPRING, THE BALL. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE. DESIGN-RELATED ISSUES: THE DESIGN OF THE OFFSET CUP IMPACTOR HAS BEEN IN THE FIELD SINCE 2006. EXACTECH IS NOT AWARE OF ANY OTHER SIMILAR COMPLAINT REPORTS INVOLVING THIS DEVICE DUE SINCE 2008. THE FREQUENCY OF OCCURRENCE RANKING SCALE IS VERY LOW; THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT, WHICH HAS BEEN IN THE FIELD SINCE 2016. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS " VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. THE DISASSEMBLY DIFFICULTY REPORTED WAS LIKELY THE RESULT OF METAL BINDING IN THE JOINTS OF THE DEVICE, WHICH LED TO THE MECHANISM BEING DIFFICULT TO MOVE. THE DISASSEMBLY OF COMPONENTS OF THE DEVICE WAS LIKELY THE RESULT OF PREVIOUS REMOVAL OR LOOSENING OF THE SCREW COMPONENT, WHICH ALLOWED FOR IT TO DISASSEMBLE IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THE SPRING AND BALL ¿FELL¿ OUT OF THE INSTRUMENT, ALL PIECES WERE ACCOUNTED FOR. THIS EVENT DID NOT LEAD TO PATIENT ADVERSE EVENT. AN INVESTIGATION WAS CONDUCTED; THE DISASSEMBLY DIFFICULTY REPORTED WAS LIKELY THE RESULT OF METAL BINDING IN THE JOINTS OF THE DEVICE, WHICH LED TO THE MECHANISM BEING DIFFICULT TO MOVE. THE DISASSEMBLY OF COMPONENTS OF THE DEVICE WAS LIKELY THE RESULT OF PREVIOUS REMOVAL OR LOOSENING OF THE SCREW COMPONENT, WHICH ALLOWED FOR IT TO DISASSEMBLE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00422 | 2022-04-18 | 2019-05-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | A+ BROACH HANDLE LEFT | EXACTECH, INC. |  | 167-01-00 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE INSTRUMENTATION. THE BROACH HANDLE BROKE WHEN IT WAS IMPACTED BY THE SURGEON. THERE WAS NO DELAY TO SURGERY AND NO PATIENT ADVERSE EVENT. THE PATIENT WAS NOT AFFECTED IN ANY WAY. THE REPRESENTATIVE WAS PRESENT AT THE TIME OF SURGERY. THE DEVICE WAS RECEIVED FOR ANALYSIS. ATTEMPTS WERE MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REPRESENTATIVE, BUT NO RESPONSES WERE RECEIVED. NO OTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS ABOUT THE DEVICE/EVENT/PATIENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS AVAILABLE FOR ANALYSIS. ENGINEERING ANALYSIS COMPLETED BY NF ON 9/12/2019. DESIGN-RELATED ISSUES: THE DESIGN OF THE NOVATION ANTERIOR BROACH HANDLE HAS BEEN IN THE FIELD SINCE 2009. EXACTECH IS AWARE OF 9 SIMILAR COMPLAINT REPORTS INVOLVING BROKEN NOVATION ANTERIOR BROACH HANDLES SINCE 2009. BECAUSE THE NOVATION ANTERIOR BROACH HANDLE IS USED IN TOTAL HIP SURGERIES, SALES DATA FOR FEMORAL STEMS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THIS NOVATION ANTERIOR BROACH HANDLE WAS MANUFACTURED IN 2010. THE DESIGN OF THE DEVICE HAS SINCE BEEN UPDATED (DCRTC-12-0106 AND DCRTC-14-0171) TO REPLACE THE PIVOT PIN THAT IS RETAINED WITH A SPIRAL RETAINING RING WITH A PIVOT RIVET THAT IS MECHANICALLY DEFORMED. ENDURANCE TESTING WAS PERFORMED ON THIS MECHANICALLY DEFORMED (SWAGED) DESIGN AND IS DOCUMENTED IN TM-2012-477. ADDITIONALLY, ENDURANCE TESTING WAS PERFORMED ON A SPECIAL A+ STRAIGHT BROACH HANDLE WITH A PROTOTYPE DESIGN CHANGE TO THE RELEASE LATCH AND LATCH PIN. RESULTS OF THIS TEST WERE USED TO DETERMINE THE NEED FOR AN UPDATED ROLL PIN, RELEASE LATCH, AND LATCH PIN. SEE TM-2013-264 FOR DETAILS. MFG-RELATED ISSUES: EXACTECH IS AWARE OF 2 OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 20 UNITS, WHICH HAVE BEEN IN THE FIELD SINCE 2010. THIS DEVICE WAS MANUFACTURED BEFORE THE DESIGN WAS UPDATED. THE DEVICE HISTORY RECORD WAS REVIEWED AS PART OF COMPLAINT (B)(4) FOR THIS DEVICE, AND ALL PARTS WERE ACCEPTED WITH CONFORMANCE TO THE DEVICE SPECIFICATIONS. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS NOVATION ANTERIOR BROACH HANDLE, (B)(4) REV C, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 9. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR. THE BROKEN DEVICE REPORTED IN EXPERIENCE (B)(4) WAS LIKELY THE RESULT OF WEAR DUE TO YEARS OF NORMAL USE, WHICH LED TO DISASSEMBLY OF THE LEVER PIVOT PIN, THE STAINLESS-STEEL RETAINING RING, THE COMPRESSION SPRING, AND THE BROACH LOCKING PIN. THE DESIGN HAS BEEN UPDATED TO REPLACE THE LEVER PIVOT PIN AND STAINLESS-STEEL RETAINING RING WITH A LEVER PIVOT RIVET TO IMPROVE THE ROBUSTNESS OF THE DESIGN. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT DUE TO THE DISASSEMBLY OF THE DEVICE WHEN IT WAS IMPACTED BY THE SURGEON. AN INVESTIGATION WAS CONDUCTED: THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF WEAR DUE TO YEARS OF NORMAL USE, WHICH LED TO DISASSEMBLY OF THE LEVER PIVOT PIN, THE STAINLESS-STEEL RETAINING RING, THE COMPRESSION SPRING, AND THE BROACH LOCKING PIN. THE DESIGN HAS BEEN UPDATED TO REPLACE THE LEVER PIVOT PIN AND STAINLESS-STEEL RETAINING RING WITH A LEVER PIVOT RIVET TO IMPROVE THE ROBUSTNESS OF THE DESIGN. | N | Y |  |  | 2022-04
1038671-2022-00423 | 2022-04-18 | 2019-07-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LPI PATELLA PREP HANDLE | EXACTECH, INC. |  | 213-60-00 |  | IT WAS REPORTED FROM THE OUS THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE INSTRUMENT. THE LOCKING PIN IS MISSING FROM THE PATELLA CLAMP. IT WAS CONFIRMED THAT THE LOCKING PIN DID NOT FALL INTO THE SURGICAL SITE, THEREFORE, WAS NOT LEFT IN THE PATIENT. THE PATIENT¿S HEALTH WAS STABLE LEAVING THE OPERATING ROOM. NEW INFORMATION PROVIDED 4 AUG 2021: LOCKING PIN MISSING FROM THE PATELLA CLAMP, IDENTIFIED WHEN LOAN KIT WAS RETURNED FROM THE HOSPITAL. NO ADDITIONAL INFORMATION IS AVAILABLE. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE(S) ARE AVAILABLE FOR EVALUATION. EVALUATION COMPLETED ON BY 12/06/2019 BY (B)(4). FINDINGS: THE MISSING DOWEL PIN IS INTENDED TO SECURE THE SPRING COMPONENT TO THE HANDLE LINK. ACCORDING TO THE PART DRAWING, (B)(4), THE PIN IS SECURED WITH A LASER TACK WELD IN AN ISOLATED AREA. THE WELD APPEARS TO HAVE FAILED IN THIS ASSEMBLY. THE DOWEL PIN WAS NOT RETURNED TO EXACTECH. HOWEVER; ACCORDING TO THE INFORMATION PROVIDED, NO PIECES WERE LEFT IN THE PATIENT. MOST LIKELY CAUSE: BASED ON PREVIOUS SIMILAR INVESTIGATIONS, THE MISSING PIN REPORTED IN (B)(4) WAS LIKELY THE RESULT OF A FAILURE OF THE TACK WELD THAT HOLDS THE PIN IN PLACE, ALLOWING IT TO DISASSOCIATE FROM THE INSTRUMENT. RMR REVIEW: THE RMR AND RACR FOR THIS LPI PATELLA PREP HANDLE, (B)(4) WERE REVIEWED. THE RISK IS CAPTURED IN (B)(4). IFU 700-096-181: INSTRUMENT INSPECTION VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES. NO PATIENT ADVERSE EVENT REPORT DUE TO THE MISSING DOWEL PIN. AN INVESTIGATION WAS CONDUCTED: BASED ON PREVIOUS SIMILAR INVESTIGATIONS, THE MISSING PIN REPORTED WAS LIKELY THE RESULT OF A FAILURE OF THE TACK WELD THAT HOLDS THE PIN IN PLACE, ALLOWING IT TO DISASSOCIATE FROM THE INSTRUMENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00424 | 2022-04-18 | 2019-09-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | EQ REV LOCKING SCREW | EXACTECH, INC. |  | 320-15-05 |  | IT WAS REPORTED FROM THE US THAT THE PATIENT HAD A REVISION DUE TO A LINER DISASSOCIATION FROM THE HUMERAL ADAPTER TRAY REPORTED IN MFR 1038671-2019-00501. DURING THE REVISION THE SURGEON NOTED THAT THE GLENOSPHERE SCREW WAS LOOSE AND BACKING OUT ALLOWING THE GLENOSPHERE TO LOOSEN FROM THE GLENOID PLATE. IN ADDITION, HE NOTED THAT THE TORQUE SCREW WOULD NOT BACK OUT AND THE HEX HEAD WAS STRIPPED. THERE HAS BEEN NO OTHER INFORMATION PROVIDED BY THE CONTACTS FOR THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS AVAILABLE FOR ANALYSIS. ENGINEERING ANALYSIS COMPLETED BY (B)(4) ON 6/11/2020. DESIGN-RELATED ISSUES: THE DESIGN OF THE EQUINOXE REVERSE SHOULDER SYSTEM HAS BEEN IN THE FIELD SINCE 2007. EXACTECH IS AWARE OF 31 COMPLAINT REPORTS INVOLVING THE GLENOSPHERE DISASSOCIATING FROM THE GLENOID BASEPLATE (DISASSEMBLY, SCREW BREAKAGE, ETC.) SINCE 2008. SALES DATA FOR ALL GLENOSPHERE LOCKING SCREWS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4). THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THESE MANUFACTURING LOTS OF (B)(4) UNITS (GLENOSPHERE) AND (B)(4) UNITS (GLENOSPHERE LOCKING SCREW). THE MANUFACTURING RECORDS FOR THESE DEVICES ONLY EXIST IN PAPER FORM AND ARE STORED IN CARDBOARD BOXES AT IRON MOUNTAIN. DUE TO RISKS ASSOCIATED WITH COVID-19 TRANSMISSION, THE PHYSICAL MANUFACTURING RECORDS FOR THIS COMPLAINT WILL NOT BE REVIEWED. EXACTECH WILL CONTINUE TO MONITOR THIS COMPLAINT TYPE AND, IF WARRANTED, WILL ESCALATE ANY POSITIVE TRENDS TO THE CORRECTIONS AND REMOVAL COMMITTEE (CRC) FOR ADDITIONAL ACTIONS. REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS GLENOSPHERE AND GLENOSPHERE LOCKING SCREW, (B)(4), WAS REVIEWED. THE RISK IS CAPTURED IN LINE 36 AND 65. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE LOOSENING OF THE PROSTHESIS AND DISASSOCIATION OF MODULAR COMPONENTS ARE LISTED IN THE PRODUCT LABELING (IFU 700-096-060 REV P), THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR. THE GLENOSPHERE LOCKING SCREW MALFUNCTION REPORTED IN (B)(4) WAS LIKELY THE RESULT OF NOT PROPERLY ALIGNING THE GLENOSPHERE RELATIVE TO THE GLENOID BASEPLATE DURING THE INITIAL SURGERY. THIS MISALIGNMENT LIKELY PROHIBITED THE GLENOSPHERE FROM FULLY SEATING AND THE GLENOSPHERE LOCKING SCREW FROM FULLY ENGAGING THE GLENOID BASEPLATE, RESULTING IN GLENOSPHERE LOOSENING AND SUBSEQUENTLY DISASSEMBLING FROM THE GLENOID BASEPLATE. IFU 700-096-181: INSTRUMENT INSPECTION VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THE PATIENT WAS REVISED REPORTED IN MFR 1038671-2019-00501. AN INVESTIGATION WAS CONDUCTED: THE GLENOSPHERE LOCKING SCREW MALFUNCTION REPORTED WAS LIKELY THE RESULT OF NOT PROPERLY ALIGNING THE GLENOSPHERE RELATIVE TO THE GLENOID BASEPLATE DURING THE INITIAL SURGERY. THIS MISALIGNMENT LIKELY PROHIBITED THE GLENOSPHERE FROM FULLY SEATING AND THE GLENOSPHERE LOCKING SCREW FROM FULLY ENGAGING THE GLENOID BASEPLATE, RESULTING IN GLENOSPHERE LOOSENING AND SUBSEQUENTLY DISASSEMBLING FROM THE GLENOID BASEPLATE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00432 | 2022-04-19 | 2018-12-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LOGIC CC FEMORAL SIZE 4, RIGHT | EXACTECH, INC. |  | 02-010-06-0340 |  | IT WAS REPORTED FROM THE US THAT DURING A REVISION OF COMPETITOR COMPONENTS TO EXACTECH COMPONENTS; THE SURGEON EXPERIENCED AN ISSUE WITH THE IMPLANT PACKAGING. DURING THE CASE THERE WAS A PACKAGING ISSUE WITH THE FEMORAL COMPONENT. THE PRODUCT WAS OPENED ON THE STERILE FIELD. A SCRUB TECH NOTICED A HOLE IN THE PACKAGING. THE PRODUCT WAS PASSED OFF AND TABLE RE-DRAPED. AFTER EXAMINATION OF THE PACKAGING, IT WAS FOUND THAT BOTH CLEAR PACKAGES HAD A HOLE IN THEM. THE OUTER BOX WAS WRAPPED IN PLASTIC (NOT COMPROMISED) AND THE BOX DIDN¿T HAVE ANY DAMAGE OR HOLE IN IT (NOT COMPROMISED). THE REPORTER STATED, ¿THE HOLE IN THE STERILE CONTAINERS HAD TO HAVE OCCURRED BEFORE PRODUCT WAS PUT INTO THE BOX AND WRAPPED IN PLASTIC. THE END RESULT WAS SATISFACTORY WITH NO HARM TO THE PATIENT.¿ THERE WAS A DELAY TO SURGERY, BUT THE PATIENT WAS NOT ADVERSELY AFFECTED. THE DEVICE WAS RECEIVED IN COMPLAINTS FOR ANALYSIS ON (B)(6) 2019. THE COMPROMISED STERILITY OF THE DEVICE POSES POTENTIAL FOR PATIENT INFECTION. NO ADDITIONAL INFORMATION HAS BEEN MADE AVAILABLE. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS AVAILABLE FOR ANALYSIS. ENGINEERING ANALYSIS COMPLETED BY MB ON (B)(6) 2019. DESIGN-RELATED ISSUES: EXACTECH IS AWARE OF 18 COMPLAINT REPORTS INVOLVING 19 FEMORAL COMPONENTS DUE TO INNER PACKAGING DAMAGE SINCE 2008. SALES DATA FOR FEMORAL COMPONENTS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT. A SIBLING DEVICE WAS EVALUATED, AND ALL INNER PACKAGING WAS INTACT. SEE VISUAL EVALUATION SECTION ABOVE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS LOGIC CC FEMORAL COMPONENT, 750-2011-032-RMR-IMPLANTS REV E, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 13. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR. A CRC WAS HELD REGARDING THIS ISSUE. SEE (B)(4) NO FIELD ACTION WILL BE TAKEN AT THIS TIME. THE INNER PACKAGING DAMAGE REPORTED IN EXPERIENCE (B)(4) WAS LIKELY THE RESULT OF A HIGH ENERGY EVENT DURING TRANSPORTATION OR HANDLING OF THE DEVICE PACKAGING, WHICH LED TO THE FEMORAL FLANGE PUNCTURING OR WEDGING PAST THE INNER FOAM SHEET AND PUNCTURING BOTH TYVEK LIDS. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. ALL PACKAGING FOR INSTRUMENTS AND IMPLANTS SHOULD BE CLOSELY INSPECTED FOR DAMAGE AND STERILE INTEGRITY. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. THE LOSS OF IMPLANT STERILITY POSES A THREAT OF PATIENT INFECTION. ALL PACKAGING FOR INSTRUMENTS AND IMPLANTS SHOULD BE CLOSELY INSPECTED FOR DAMAGE AND STERILE INTEGRITY. AN INVESTIGATION WAS CONDUCTED: THE INNER PACKAGING DAMAGE REPORTED WAS LIKELY THE RESULT OF A HIGH ENERGY EVENT DURING TRANSPORTATION OR HANDLING OF THE DEVICE PACKAGING, WHICH LED TO THE FEMORAL FLANGE PUNCTURING OR WEDGING PAST THE INNER FOAM SHEET AND PUNCTURING BOTH LIDS. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00431 | 2022-04-19 | 2020-01-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | SPECIFIC DEVICE NOT REPORTED | EXACTECH, INC. |  | 9999 |  | IT WAS REPORTED FROM THE US THAT INFORMATION WAS RECEIVED FROM A CLINICIAN MEETING OF ALTEON USERS. THE FOLLOWING INFORMATION IS ALL THAT WAS REPORTED. SLIGHT SUBSIDENCE OBSERVED IN 1 MONTH FOLLOW-UP VISIT. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE(S) NOT AVAILABLE FOR EVALUATION. ENG EVAL COMPLETED BY NF ON (B)(6) 2020. DESIGN-RELATED ISSUES: EXACTECH IS AWARE OF 19 COMPLAINT REPORTS DUE TO FEMORAL STEM SUBSIDENCE (7 REVISED, 12 NON-REVISED) SINCE 2008. OF THE 19 COMPLAINTS, 2 COMPLAINT INVOLVED ALTEON HA FEMORAL STEMS. MFG-RELATED ISSUES: MANUFACTURING DATA COULD NOT BE REVIEWED BECAUSE THE SERIAL/LOT INFORMATION WAS NOT PROVIDED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE MIGRATION OF THE PROSTHESIS IS LISTED IN THE PRODUCT LABELING (IFU 700-096-018 REV U), THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR. THE FEMORAL STEM SUBSIDENCE REPORTED MAY HAVE BEEN THE RESULT OF EITHER UNDER SIZING THE FEMORAL STEM COMPARED TO THE PATIENT¿S ANATOMY OR INSUFFICIENT BONY SUPPORT FOR THE IMPLANT. HOWEVER, THIS CANNOT BE CONFIRMED AS THE PATIENT HAD NOT BEEN REVISED AT THE TIME OF THIS INVESTIGATION; THEREFORE, THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. IT IS NOTED IN IFU 700-096-018 THAT ONLY QUALIFIED SURGEONS ARE TO USE THESE PRODUCTS: WHO ARE FULLY KNOWLEDGEABLE ABOUT ALL ASPECTS OF THE SURGICAL TECHNIQUE AND USE OF THESE IMPLANTS, HAVE FULL KNOWLEDGE ABOUT THE SYSTEM COMPATIBILITY¿S, AND MUST BE FULLY TRAINED TO PROPERLY USE THE SYSTEM AND INSTRUMENTATION. ALSO, AS PART OF THE PREOPERATIVE ASSESSMENT, THE SURGEON MUST ENSURE THAT NO BIOLOGICAL, BIOMECHANICAL, OR OTHER FACTORS EXIST THAT MIGHT ADVERSELY AFFECT THE SURGERY AND/OR POSTOPERATIVE PERIOD. BONE QUALITY MUST BE CONSIDERED TO ENSURE THAT THE PROSTHESES DOES NOT SUBSIDE OR MIGRATE WITHIN THE FEMORAL CANAL OR ACETABULUM; FRACTURE OF HOST BONE SHOULD ALSO BE CONSIDERED. DEVICE SPECIFIC RISKS - FRACTURE, MIGRATION, LOOSENING, SUBLUXATION, OR DISLOCATION OF THE PROSTHESIS OR ANY OF ITS COMPONENTS, ANY OF WHICH MAY REQUIRE A SECOND SURGICAL INTERVENTION OR REVISION. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THE PATIENT HAS DEVICE SUBSIDENCE WITHOUT REPORTED REVISION. DEVICE SPECIFIC RISKS INCLUDED THE FOLLOWING: FRACTURE, MIGRATION, LOOSENING, SUBLUXATION, OR DISLOCATION OF THE PROSTHESIS OR ANY OF ITS COMPONENTS, ANY OF WHICH MAY REQUIRE A SECOND SURGICAL INTERVENTION OR REVISION. AN INVESTIGATION WAS CONDUCTED; THE FEMORAL STEM SUBSIDENCE REPORTED MAY HAVE BEEN THE RESULT OF EITHER UNDER SIZING THE FEMORAL STEM COMPARED TO THE PATIENT¿S ANATOMY OR INSUFFICIENT BONY SUPPORT FOR THE IMPLANT. HOWEVER, THIS CANNOT BE CONFIRMED AS THE PATIENT HAD NOT BEEN REVISED AT THE TIME OF THIS INVESTIGATION; THEREFORE, THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00430 | 2022-04-19 | 2020-01-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | SPECIFIC DEVICE NOT REPORTED | EXACTECH, INC. |  | 9999 |  | IT WAS REPORTED FROM THE US THAT INFORMATION WAS RECEIVED FROM A CLINICIAN MEETING OF ALTEON USERS. THE FOLLOWING INFORMATION IS ALL THAT WAS REPORTED. COMPACTION BROACHES USED, SUBSIDENCE 2 WEEKS AFTER SURGERY OF ALTEON HA STEM. PATIENT HAS NOT BEEN REVISED AT THIS TIME. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE(S) NOT AVAILABLE FOR EVALUATION. ENG EVAL COMPLETED BY (B)(4) ON 2020.04.02. DESIGN-RELATED ISSUES: EXACTECH IS AWARE OF 19 COMPLAINT REPORTS DUE TO FEMORAL STEM SUBSIDENCE (7 REVISED, 12 NON-REVISED) SINCE 2008. OF THE 19 COMPLAINTS, 2 COMPLAINT INVOLVED ALTEON HA FEMORAL STEMS. MFG-RELATED ISSUES: MANUFACTURING DATA COULD NOT BE REVIEWED BECAUSE THE SERIAL/LOT INFORMATION WAS NOT PROVIDED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE MIGRATION OF THE PROSTHESIS IS LISTED IN THE PRODUCT LABELING (IFU 700-096-018 REV U), THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR. THE FEMORAL STEM SUBSIDENCE REPORTED IN (B)(4) MAY HAVE BEEN THE RESULT OF EITHER UNDER SIZING THE FEMORAL STEM COMPARED TO THE PATIENT¿S ANATOMY OR INSUFFICIENT BONY SUPPORT FOR THE IMPLANT. HOWEVER, THIS CANNOT BE CONFIRMED AS THE PATIENT HAD NOT BEEN REVISED AT THE TIME OF THIS INVESTIGATION; THEREFORE, THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. IT IS NOTED IN IFU 700-096-018 THAT ONLY QUALIFIED SURGEONS ARE TO USE THESE PRODUCTS: WHO ARE FULLY KNOWLEDGEABLE ABOUT ALL ASPECTS OF THE SURGICAL TECHNIQUE AND USE OF THESE IMPLANTS, HAVE FULL KNOWLEDGE ABOUT THE SYSTEM COMPATIBILITY¿S, AND MUST BE FULLY TRAINED TO PROPERLY USE THE SYSTEM AND INSTRUMENTATION. ALSO, AS PART OF THE PREOPERATIVE ASSESSMENT, THE SURGEON MUST ENSURE THAT NO BIOLOGICAL, BIOMECHANICAL, OR OTHER FACTORS EXIST THAT MIGHT ADVERSELY AFFECT THE SURGERY AND/OR POSTOPERATIVE PERIOD. BONE QUALITY MUST BE CONSIDERED TO ENSURE THAT THE PROSTHESES DOES NOT SUBSIDE OR MIGRATE WITHIN THE FEMORAL CANAL OR ACETABULUM; FRACTURE OF HOST BONE SHOULD ALSO BE CONSIDERED. DEVICE SPECIFIC RISKS - FRACTURE, MIGRATION, LOOSENING, SUBLUXATION, OR DISLOCATION OF THE PROSTHESIS OR ANY OF ITS COMPONENTS, ANY OF WHICH MAY REQUIRE A SECOND SURGICAL INTERVENTION OR REVISION. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THE PATIENT HAS DEVICE SUBSIDENCE WITHOUT REPORTED REVISION. DEVICE SPECIFIC RISKS INCLUDED THE FOLLOWING: FRACTURE, MIGRATION, LOOSENING, SUBLUXATION, OR DISLOCATION OF THE PROSTHESIS OR ANY OF ITS COMPONENTS, ANY OF WHICH MAY REQUIRE A SECOND SURGICAL INTERVENTION OR REVISION. AN INVESTIGATION WAS CONDUCTED; THE FEMORAL STEM SUBSIDENCE REPORTED MAY HAVE BEEN THE RESULT OF EITHER UNDER SIZING THE FEMORAL STEM COMPARED TO THE PATIENT¿S ANATOMY OR INSUFFICIENT BONY SUPPORT FOR THE IMPLANT. HOWEVER, THIS CANNOT BE CONFIRMED AS THE PATIENT HAD NOT BEEN REVISED AT THE TIME OF THIS INVESTIGATION. THEREFORE; THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00429 | 2022-04-19 | 2019-05-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | ERGO REAMER | EXACTECH, INC. |  |  |  | IT WAS REPORTED FROM THE US VIA THE ERGO SURVEY THAT WHEN THE SURGEON WAS USING THE REAMERS, THE REAMERS KEPT COMING OFF THE HANDLE WHILE ON THE GLENOID. THE SURGEON WOULD THEN HAVE TO PULL THEM OUT BY HAND, AND THEN REATTACH IT TO THE HANDLE BEFORE STARTING AGAIN. NO PIECES WERE LEFT IN THE PATIENT. THE PATIENT WAS DISCHARGED HOME WITH NO COMPLICATIONS. THE REP WAS PRESENT AT THE TIME OF SURGERY. TWO ATTEMPTS TO REQUEST ADDITIONAL INFORMATION WERE MADE DURING REMEDIATION AND NO RESPONSES WERE RECEIVED. NO ADDITIONAL INFORMATION RELATED TO THE EVENT HAS BEEN PROVIDED. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO.; DEVICE INFO. FINDINGS: A REVIEW OF DESIGN PRINTS, DESIGN VERIFICATION, AND DESIGN VALIDATION, PER CAPA(B)(4), IDENTIFIED THE FOLLOWING: 1. LMC INTERFERENCE WAS NOT VERIFIED AS PART OF THE EQUINOXE ERGO TOLERANCE ANALYSIS REPORT ((B)(4)). THE DEVICE HISTORY RECORDS (DHRS) FOR THE DRIVERS AND REAMERS WERE REVIEWED AS PART OF CAPA(B)(4), AND ALL PARTS WERE ACCEPTED WITH CONFORMANCE TO THE DEVICE SPECIFICATIONS. THE ERGO PILOT RMR, 750-2016-002 REV ¿ DID NOT EXPLICITLY DEFINE THE FAILURE MODE OF THE REAMERS NOT BEING RETAINED IN THE DRIVERS. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: AS DETERMINED BY CAPA(B)(4), THE MATING ISSUE REPORTED IN (B)(4) WAS LIKELY THE RESULT OF NOT ADEQUATELY DEFINING THE ACCEPTANCE CRITERIA FOR THE LOCKING MECHANISM RESULTING IN NOT ENOUGH INTERFERENCE BETWEEN REAMERS AND REAMER DRIVERS. CORRECTIVE ACTION TAKEN: A CORRECTION AND REMOVAL COMMITTEE (CRC) MEETING WAS ALSO HELD TO ADDRESS THE NEED FOR FIELD ACTION. THE COMMITTEE DETERMINED THAT DUE TO THE LOW RISK TO HEALTH AND LOW RATE OF OCCURRENCE, NO FIELD ACTION IS NECESSARY AT THIS TIME. SEE (B)(4) FOR DETAILS. THE SYSTEMATIC ISSUES IN REGARD TO DESIGN CONTROLS NOT BEING FOLLOWED ARE BEING FURTHER ADDRESSED AS PART OF CAPA(B)(4). THE DIR WAS REVISED TO INCLUDE THE ACCEPTANCE CRITERIA FOR THE LOCKING MECHANISM. SEE (B)(4). NEW PART NUMBERS WERE CREATED, AND INSTRUMENT DESIGN/DRAWING WERE REVISED. SEE (B)(4). A NEW ORDER FOR PARTS WERE RECEIVED ((B)(4)), AND A CADAVER VALIDATION LAB WAS COMPLETED WITH SURGEON USERS. VERIFICATION AND VALIDATION WILL BE PERFORMED ON THE NEW DESIGN, AND DESIGN TRANSFER AND FINAL DESIGN REVIEW WILL BE CONDUCTED. THE CAPA IS STILL OPENED AT THE TIME OF THIS INVESTIGATION. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. THIS DOES NOT APPEAR TO BE A MANUFACTURING ISSUE. THERE IS NO REPORTED PATIENT ADVERSE EVENT. THE REAMERS KEPT COMING OFF THE HANDLE WHILE ON THE GLENOID. THE SURGEON WOULD THEN HAVE TO PULL THEM OUT BY HAND, AND THEN REATTACH IT TO THE HANDLE BEFORE STARTING AGAIN. NO PIECES WERE LEFT IN THE PATIENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00426 | 2022-04-19 | 2022-03-30 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | CAGE GLENOID MEDIUM, ALPHA | EXACTECH, INC. | 314-13-03 | 314-13-03 |  | AS REPORTED, APPROXIMATELY 4 YEARS POST OP THE INITIAL LEFT TSA, THIS (B)(6) FEMALE PATIENT WAS REVISED DUE TO PAIN. IN SURGERY, THE GLENOID WAS PULLED OUT WITHOUT AN RESISTANCE. SURGEON ALSO HAD TO TAKE OUT THE PRESERVE STEM AND CEMENT A REGULAR STEM. HE CHECKED FOR INFECTION, BUT THERE WAS NONE. SHOULDER WAS REVISED TO AN EXTENDED CAGE. THERE WAS SIGNIFICANT BONE LOSS ON THE HUMERUS AND GLENOID. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. | PENDING EVALUATION. CONCOMITANT DEVICE(S): 300-30-08, 84829002 - EQUINOXE PRESERVE STEM 8MM HUMERAL HEAD; REPLICATOR PLATE 300-11-00, 99472002 - EQUINOXE TORQUE DEFINING SCREWS. NO INFORMATION PROVIDED IN THE FOLLOWING SECTION(S). | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00428 | 2022-04-19 | 2022-03-30 | I, n, j, u, r, y | Manufacturer report | VANTAGE | LP TALAR IMPACTOR INSERT | EXACTECH, INC. | 351-07-04 | 351-07-04 |  | IT WAS REPORTED THAT DURING IMPACTION, THE JOINT WAS VERY STIFF AND THE SURGEON HAD TO HIT THE TALAR IMPACTOR VERY HARD TO PUSH THE FINAL IMPLANT IN. WHEN THE FIRST TALAR IMPACTOR PLASTIC PROTECTION BROKE, WE HAD ANOTHER ON THAT WAS ON SITE TO USE BUT IT BROKE AS WELL. AFTER REMOVING THE BROKEN PIECE AND CLEANING OUT THE PATIENT, WE COULD STILL USE THE TALAR IMPACTOR PLASTIC PROTECTION TO COMPLETE THE SURGERY WITH NO FURTHER COMPLICATIONS. THERE WAS A SMALL SURGICAL DELAY THAT WAS LESS THAN 5 MINUTES. THE PATIENT DID NOT EXPERIENCE ANY ADVERSE EVENT AS A RESULT OF THE SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. | PENDING EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00425 | 2022-04-19 | 2022-03-24 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 42MM HUMERAL CONST LINER +2.5 | EXACTECH, INC. | 320-42-13 | 320-42-13 |  | AS PREVIOUSLY REPORTED, THIS (B)(6) MALE PATIENT HAS HAD 2 PREVIOUS REVISIONS WITH ANTIBIOTIC SPACERS PUT IN PLACE DURING 2021. OVERTIME. THIS PATIENT HAD INITIAL REVERSE SHOULDER IMPLANTED ON (B)(6) 2021, WHICH WAS PARTIALLY REVISED, DUE TO INFECTION ON (B)(6) 2021. EVERYTHING WAS THEN PULLED ON (B)(6) 2021 AND AN INTERSPACE ANTIBIOTIC SPACER WAS IMPLANTED. THAT SPACER WAS THEN PULLED OUT ON (B)(6) 2021 AND A NEW SPACER WAS IMPLANTED. CURRENTLY, THE 2ND ANTIBIOTIC SPACER WAS REMOVED ON (B)(6) 2022 AND PATIENT HAD A FULL REVERSE SHOULDER IMPLANTED. DUE TO ANOTHER INFECTION, ALL OF THOSE IMPLANTS WERE REMOVED AND THE SAME SIZES WERE IMPLANTED ON (B)(6) 2022. THIS PATIENT HAS HAD PREVIOUS REVISIONS, CAPTURED IN MDR#'S 1038671-2021-00150 AND 1038671-2021-00207. | PENDING EVALUATION. CONCOMITANT DEVICE(S): 300-30-14, (B)(4) - EQUINOXE PRESERVE STEM 14MM; 320-10-00, (B)(4) - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0; 320-15-04, (B)(4) - RS GLENOID PLATE R POST AUG, 8 DEG, RIGHT; 320-15-05, (B)(4) - EQ REV LOCKING SCREW; 320-20-00, (B)(4) - EQ REVERSE TORQUE DEFINING SCREW KIT; 320-20-30, (B)(4) - EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 30MM; 320-20-38, (B)(4) - EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 38MM ;320-20-46,(B)(4)- EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 46MM; 320-08-42, (B)(4) - GLENOSPHERE EXP 42MM +4MM OFFSET. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00427 | 2022-04-19 | 2022-03-30 | I, n, j, u, r, y | Manufacturer report | VANTAGE | LP TALAR IMPACTOR INSERT | EXACTECH, INC. | 351-07-04 | 351-07-04 |  | IT WAS REPORTED THAT DURING IMPACTION, THE JOINT WAS VERY STIFF AND THE SURGEON HAD TO HIT THE TALAR IMPACTOR VERY HARD TO PUSH THE FINAL IMPLANT IN. WHEN THE FIRST TALAR IMPACTOR PLASTIC PROTECTION BROKE, WE HAD ANOTHER ON THAT WAS ONSITE TO USE BUT IT BROKE AS WELL. AFTER REMOVING THE BROKEN PIECE AND CLEANING OUT THE PATIENT, WE COULD STILL USE THE TALAR IMPACTOR PLASTIC PROTECTION TO COMPLETE THE SURGERY WITH NO FURTHER COMPLICATIONS. THERE WAS A SMALL SURGICAL DELAY THAT WAS LESS THAN 5 MINUTES. THE PATIENT DID NOT EXPERIENCE ANY ADVERSE EVENT AS A RESULT OF THE SURGICAL DELAY/ PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. | CONCOMITANT MEDICAL PRODUCTS: PENDING EVALUATION, "NO INFORMATION PROVIDED." | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00439 | 2022-04-20 | 2022-03-30 | I, n, j, u, r, y | Manufacturer report | TRULIANT TIB IMP PSC INSERT SZ 4, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT TIB IMP PSC INSERT SZ 4, 9MM | 02-022-44-4009 |  | AS REPORTED, APPROXIMATELY TEN MONTHS POST INITIAL RIGHT TKA, THIS (B)(6) MALE PATIENT WAS EXPERIENCING KNEE PAIN. SURGEON DECIDED TO REVISE THE FEMORAL AND TIBIAL COMPONENTS. THE PATIENT WAS REVISED TO EXACTECH KNEE REVISION COMPONENTS. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. SALES REP. WAS UNABLE TO OBTAIN PHOTOS OR X-RAYS. THE DEVICES ARE NOT AVAILABLE FOR EVALUATION DUE TO IMPLANTS HAZARDOUS FOR BEING IN PATIENTS' BODY. | CONCOMITANT MEDICAL PRODUCTS: TRULIANT PS CEM FEM PS CEM RIGHT SZ 4 (CAT# 02-020-11-0340 / (B)(4)); TRULIANT TIB FIT TRAY CEM SZ 4F / 3T (CAT# 02-022-45-4030 / (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00441 | 2022-04-20 | 2022-03-31 | I, n, j, u, r, y | Manufacturer report | FEMORAL COMPONENTS | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | FEMORAL COMPONENTS | UNKNOWN |  | AS REPORTED BY LEGAL TEAM, PATIENT IS SCHEDULE TO HAVE REVISION FOR AN UNKNOW REASON. NO ADDITIONAL INFORMATION AVAILABLE. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00440 | 2022-04-20 | 2022-02-03 | I, n, j, u, r, y | Manufacturer report | +5MM 36MM LATERALIZED GXL LINER | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | +5MM 36MM LATERALIZED GXL LINER | UNKNOWN |  | AS REPORTED, THIS (B)(6) FEMALE PATIENT LINER WAS BROKEN INTO SEVERAL PIECES. SURGEON REMOVED ALL PIECES FROM THE PATIENT. SURGEON DO NOT KNOW WHEN THE LINER BROKE. THE CUP, SCREW, AND LINER WERE EXPLANTED. THIS HIP WAS A HYBRID, EXACTECH COMPONENTS IN THE ACETABULUM AND ZIMMER COMPONENTS IN THE FEMUR. BECAUSE IT WAS A HYBRID, THE DOCTOR CHOSE TO REVISE THE EXPLANTED PORTION TO COMPETITORS TO MATCH STEM THAT WAS ALREADY IN THE PATIENT'S FEMUR. THERE WERE NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. X-RAYS RECEIVED. THE DEVICES ARE NOT AVAILABLE FOR EVALUATION. | CONCOMITANT MEDICAL PRODUCTS: 52MM CROWN CLUSTER CUP WITH INTEGRIP; 30MM BONE SCREW. ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00437 | 2022-04-20 | 2022-03-21 | I, n, j, u, r, y | Manufacturer report | LOGIC CR TIB INSERT SLOPE++, SZ 3, 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC CR TIB INSERT SLOPE++, SZ 3, 13MM | 02-012-49-3013 |  | AS REPORTED, APPROXIMATELY 3.5 YEARS POST INITIAL RIGHT TKA, THIS (B)(6) MALE PATIENT HAD A TOTAL KNEE REPLACEMENT WHICH FAILED. PATIENT REASON FOR THE REVISION WAS POLY FAILURE. THERE IS LOTS OF POLY WEAR AND IT HAD A SANDPAPER TEXTURE TO IT. IT WAS DISCOLORED AND YELLOW IN SPOTS. THERE WAS OSTEOLYSIS. SURGEON SAW THE PATIENT IN CLINIC AND RECOGNIZED A FAILING COMPONENT. PATIENT REVISED TO COMPETITOR'S DEVICE. THERE WAS NO PARTS OR PIECES FELL INTO THE PATIENT WOUND SITE. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. SALES REP WAS UNABLE TO OBTAIN PHOTOS/X-RAYS. THE DEVICE WILL BE RETURN. | CONCOMITANT MEDICAL PRODUCTS: LOGIC CR FEMORAL CEM, RIGHT, SZ 3 (CAT# 02-010-03-0330 / SERIAL# (B)(4)). LGC TIBIAL FIT TRAY CEM SZ 3F / 4T (CAT# 02-012-45-3040 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00438 | 2022-04-20 | 2022-03-31 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 3.5 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 3.5 9MM | 02-012-35-3509 |  | AS REPORTED, APPROXIMATELY TEN YEARS POST INITIAL TKA, THIS FEMALE PATIENT HAD A POLY SWAP AND PATELLA SWAP REVISION SURGERY. PATIENT WAS HAVING PAIN IN THE KNEE. THERE WAS SOME POLY WEAR AND OSTEOLYSIS WAS PRESENT. PATIENT WAS REVISED WITH A NEW POLY AND PATELLA. THERE WAS NO BREAKAGE OF DEVICES OR SURGICAL DELAY/PROLONGATION. SURGERY WENT WELL AND THERE IS NO POST-OP INFO. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. SALES REP WAS UNABLE TO OBTAIN X-RAYS/PHOTOS. THE DEVICES ARE NOT AVAILABLE FOR EVALUATION AS THERE WERE THROWN AWAY BY HOSPITAL. | CONCOMITANT MEDICAL PRODUCTS: THREE PEG PATELLA 35MM (CAT# 200-02-35 / (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00436 | 2022-04-20 | 2022-03-30 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 1, 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 1, 11MM | 02-012-44-1011 |  | AS REPORTED, APPROXIMATELY 2.5 YEARS POST INITIAL LEFT TKA, THIS (B)(6) Y/O FEMALE PATIENT HAD A REVISION DUE TO CHRONIC HYDRARTHROSIS OCCURRING AFTER SOMEWHAT PROLONGED WALKS. PATIENT HAD PE ISSUE AND PATIENT PAINFUL WHILE WALKING. THERE WAS EVIDENCE OF POLY WEAR. PATIENT WAS SLIGHTLY LAX, SURGEON EXPLAINED TO PATIENT, UNTIL PROVEN OTHERWISE, THAT SYNOVECTOMY IS NECESSARY WITH REPLACEMENT OF A SLIGHTLY THICKER INSERT (13 VS 11MM). THE WEAR MARKS ON THE CENTER CAM FOOT WERE REAL BUT RELATIVELY UNOBTRUSIVE. THE SYNOVIAL FLUID WAS NOT VERY ABUNDANT AND IT WAS PERFECTLY CLEAR. PATIENT WAS REVISED TO EXACTECH DEVICES. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION REPORTED. PHOTOS AND X-RAYS RECEIVED, THE DEVICE WILL BE RETURN. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00435 | 2022-04-20 | 2018-08-08 | I, n, j, u, r, y | Manufacturer report | TRULIANT TIB IMP PS INSERT SZ 4 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT TIB IMP PS INSERT SZ 4 9MM | 02-022-35-4009 |  | AS REPORTED BY EXACTECH TRULIANT KNEE CLINICAL STUDY, APPROXIMATELY FOURTEEN MONTHS POST INITIAL TKA, THIS (B)(6) Y/O MALE PATIENT HAD A REVISION DUE TO WORSENING PAIN, DETERIORATING RANGE OF MOTION, PROGRESSIVE DEFORMITY, AND SIGNIFICANT INTERFERENCE WITH FUNCTION THE ONLY OPTION FOR PAIN CONTROL AND FUNCTIONAL RESTORATION IS TOTAL KNEE REPLACEMENT REVISION. PER CLINICAL STUDY CASE REPORT FORM INDICATES THIS EVENT IS UNLIKELY RELATED TO DEVICES OR PROCEDURE. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00434 | 2022-04-20 | 2022-03-31 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE TRAY ADAPTER PLATE TRAY +0 | EXACTECH, INC. | NA | 320-10-00 |  | AS REPORTED, APPROXIMATELY 5 YEARS POST OP THE INITIAL RIGHT TSA, THIS (B)(6) MALE PATIENT WAS REVISED. TWO YEARS AGO, THE 46MM POLY DISASSOCIATED FROM THE ADAPTOR TRAY AND SLOWLY MIGRATED TO THE SCAPULA. OVER TIME, THE 46MM GLENOSPHERE IMPACTED ON THE ADAPTER TRAY CREATING METALOSIS. ON (B)(6) 2021, THE PATIENT REPORTED INCREASED PAIN TO HIS LOCAL ORTHOPEDIC SURGEON. ON (B)(6) 2022, THE PAIN INCREASED QUICKLY. PATIENT IS A FARMER WITH HIGH PAIN TOLERANCE. THERE WAS A 15-30 MINUTE DELAY/PROLONGATION. | PENDING EVALUATION. CONCOMITANT DEVICE(S): 300-01-15, (B)(4) - EQUINOXE, HUMERAL STEM PRIMARY, PRESS FIT 15MM; 320-01-46, (B)(4) - EQUINOXE REVERSE 46MM GLENOSPHERE; 320-15-05, (B)(4) - EQ REV LOCKING SCREW; 320-20-00, (B)(4) - EQ REVERSE TORQUE DEFINING SCREW KIT; 320-46-03, (B)(4)- EQUINOXE REVERSE 46MM HUMERAL LINER +2.5. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00433 | 2022-04-20 | 2021-02-15 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | SHOULDER COMPONENTS | EXACTECH, INC. |  | UNK |  | AS REPORTED, POST OP THE INITIAL TSA, THIS PATIENT WAS REVISED DUE INFECTION. A FULL HRP IMPLANT ASSEMBLY WAS REMOVED DUE TO A PATIENT HAVING AN INFECTION. EXACTECH PRODUCT WAS FUNCTIONING PROPERLY. THERE WAS NO SURGICAL DELAY/PROLONGATION. NO OTHER PATIENT INFORMATION/MEDICAL HISTORY REPORTED. | PENDING EVALUATION. NO INFORMATION PROVIDED IN THE FOLLOWING SECTION(S). | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00442 | 2022-04-21 | 2022-03-31 | I, n, j, u, r, y | Manufacturer report | EQUINOXE REVERSE 46MM HUMERAL LINER +0 | PROSTHESIS, SHOULDER, NON-CONSTRAINED, METAL/POLYMER CEMENTED | EXACTECH, INC. | 320-46-00 | 320-46-00 |  | AS REPORTED, APPROXIMATELY ONE YEAR POST OP THE INITIAL TSA, THIS MALE PATIENT'S WAS REVISED DUE TO HUMERAL LINER DISASSOCIATION. PATIENT WAS PERFORMING OVERHEAD ACTIVITY 2 WEEKS AGO AND FELT SOMETHING IN HIS SHOULDER. PRESENTED TO CLINIC WITH PAIN AND X-RAY SHOWING POSSIBLE LINER DISASSOCIATION. HE WAS REVISED YESTERDAY BY SURGEON TO A NEW O TRAY AND 46 MM LINER OF SAME SIZE. NO FURTHER INFORMATION AVAILABLE. EXPLANT IS NOT GIVEN TO REP WITHOUT PATIENT PERMISSION. PATIENT LAST KNOWN TO BE IN STABLE CONDITION. | PENDING EVALUATION: CONCOMITANT DEVICE(S): 320-10-00, 05117281 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00443 | 2022-04-21 | 2022-03-31 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 38MM HUMERAL LINER +2.5 | EXACTECH, INC. | 320-38-03 | 320-38-03 |  | AS REPORTED, THIS (B)(6) MALE PATIENT'S RIGHT TSA WAS REVISED DUE TO HIS SHOULDER BECOMING UNSTABLE. SURGEON DID A POLY EXCHANGE WITH A LARGER PROXIMAL HUMERAL PLATE. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETUNING DUE TO HOSPITAL POLICY. | CONCOMITANT DEVICE(S): 320-20-22, S335829 - EQ REV COMPRESS SCREW LCK CAP KIT 4.5MM X 22 MM. 320-15-05, 7246636 - EQ REV LOCKING SCREW. 320-20-00, 7151171 - EQ REVERSE TORQUE DEFINING SCREW KIT. 320-08-38, 7123648 - GLENOSPHERE EXP 38MM +4MM OFFSET. 320-10-10, 6886305 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +10. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00444 | 2022-04-21 | 2022-03-30 | I, n, j, u, r, y | Manufacturer report | THREE PEG PATELLA 38MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | N/A | 200-02-38 |  | AS REPORTED, APPROXIMATELY SIX YEARS POST RIGHT TKA, THIS (B)(6) PATIENT WAS EXPERIENCING ANTERIOR KNEE PAIN. PATIENT HAD A REVISION TO REMOVED LOOSE PATELLA, LOOSE TIBIA COMPONENT AND EXCHANGED POLY. DURING SURGERY SURGEON NOTICED FRAGMENTED, LOOSE PATELLA AND SOME CEMENT PARTICLE. TIBIA WAS APPARENT LOOSE, THERE WAS POLY WEAR AND SYNOVITIS. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. SALES REP WAS UNABLE TO OBTAIN PHOTOS/X-RAYS. THE DEVICES WILL BE RETURN. | CONCOMITANTS PRODUCTS: LOGIC TIBIA PS MOD INSRT SZ 4 11MM (CAT# 02-012-35-4011 / (B)(4)); LGC TIBIAL FIT TRAY CEM SZ 4F / 4T (CAT# 02-012-45-4040 / (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00445 | 2022-04-21 | 2022-04-01 | I, n, j, u, r, y | Manufacturer report | LOGIC CR TIB INSERT STD, SZ 2, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC CR TIB INSERT STD, SZ 2, 9MM | 02-012-47-2009 |  | AS REPORTED, APPROXIMATELY 4.5 YEARS POST INITIAL RIGHT TKA, THIS PATIENT PRESENTED IN CLINIC AFTER RECEIVING RECALL LETTER WITH PAIN, SWELLING, AND X-RAY INDICATION OF MEDIAL POLY WEAR. MAGNETIC RESONANCE IMAGING (MRI) DID NOT SHOW LOOSENING BUT DID SHOW SYNOVITIS. SURGEON PERFORMED REVISION. SURGEON USED A OSTEOTOME TO REMOVED THE OLD PINS WET. THE PATIENT¿S INSERT WAS EXCHANGED AND KNEE WASHED OUT. THE KNEE WAS STABLE WITH GOOD RANGE OF MOTION (ROM) WITH SAME SIZE POLY 9MM. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION. SURGEON COMMENTED THAT THE INSERT DID NOT LOOK EXTREMELY WORN BUT DID HAVE SOME PITTING AND INDENTION. X-RAY RECEIVED. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO PATIENT WILL BE COLLECTING HER OLD POLY FOR EVALUATION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00446 | 2022-04-21 | 2022-04-01 | I, n, j, u, r, y | Manufacturer report | TRULIANT TIB IMP PS INSERT SZ 6 12MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT TIB IMP PS INSERT SZ 6 12MM | 02-022-35-6012 |  | AS REPORTED BY LEGAL TEAM, THIS MALE PATIENT HAD A KNEE REPLACEMENT SURGERY LAST YEAR. PATIENT HAS DEVELOPED SERIOUS SYMPTOMS SINCE THEN, INCLUDING PAIN, PERSISTENT SWELLING, AND THE KNEE ¿CLICKING.¿ HIS MOBILITY IS NOT GOOD AND IT¿S VERY PAINFUL TO WALK. PATIENT RECEIVED THE RECALL LETTER AND ACTUALLY FELT RELIEVED BECAUSE HE¿S HOPING THAT IT SHOWS THERE¿S A CLEAR ANSWER TO HIS ISSUES. PATIENT HAS NO SCHEDULE REVISION AT THIS TIME. NO ADDITIONAL INFORMATION AVAILABLE. | CONCOMITANT MEDICAL PRODUCTS: TRULIANT TIB IMP PS INSERT SZ 6 12MM (CAT# 02-022-35-6012 / SERIAL# (B)(4)). LOGIC STEM EXT 14MM X 25MM (CAT# 02-012-60-1425 / SERIAL# (B)(4)). TRULIANT PS CEM FEM PS CEM LEFT SZ 6 (CAT# 02-020-11-0260 / SERIAL# (B)(4)). TRULIANT TIB FIT TRAY CEM SZ 6F / 5T (CAT# 02-022-45-6050 / SERIAL# (B)(4)). TIBIAL STEM EXT. SCREW (CAT# 204-70-00 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00447 | 2022-04-21 | 2022-02-08 | I, n, j, u, r, y | Manufacturer report | STEM EXTENSION 22MM X 80 L. | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | STEM EXTENSION 22MM X 80 L. | 204-46-08 |  | AS REPORTED, APPROXIMATELY 13.5 YEARS POST TKA, THIS 80 Y/O MALE PATIENT EXPERIENCED INCREASING PAIN AND LEFT LEG NO LONGER ABLE TO BEAR WEIGHT. RADIOLOGICALLY SHOWED PERIPROSTHETIC FRACTURE WITH FRACTURE OF THE FEMORAL STEM AT THE ADAPTER/STEM TRANSITION. PATIENT HAD A REVISION AND THE EXPLANT WAS PRESERVED. NO ADDITIONAL INFORMATION AVAILABLE. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00448 | 2022-04-21 | 2022-04-01 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 4 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 4 9MM | 02-012-35-4009 |  | AS REPORTED, APPROXIMATELY SEVEN YEARS POST INITIAL LEFT TKA, THIS (B)(6) Y/O FEMALE PATIENT WAS REVISED DUE TO INFECTION. THERE WAS A SMALL POSTERIOR BEARING SURFACE SECTION OF THE INSERT THAT LOOKS WHITER THAN THE REST OF THE POLY. PATIENT WAS REVISED TO A SIZE 4X12MM PSC INSERT. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS IN GOOD CONDITION FOLLOWING THE INSERT SWAP AND WASH OUT. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO SURGEON RETAINED POLY. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00449 | 2022-04-21 | 2022-04-05 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 4 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 4 11MM | 02-012-35-4011 |  | AS REPORTED, APPROXIMATELY 8.5 YEARS POST INITIAL LEFT TKA, THIS (B)(6) MALE PATIENT COMPLAINED THAT DURING THE LAST FEW YEARS HE HAD DISCOMFORT AND PAIN WHILE WALKING. PATIENT HAD A REVISION AND THERE WAS SIGNIFICANT AMOUNT OF WEAR AND DESTRUCTION OF THE LINER, HOWEVER, THERE WAS NOT A SIGNIFICANT AMOUNT OF OSTEOLYSIS. THE PATIENT WAS REVISED TO A SIZE 4 LEFT CONDYLAR CONSTRAINED FEMUR WITH THE 22X120 SPLINED STEM AND A LARGE 42 FEMORAL CONE. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PHOTOS OF AFFECTED PRODUCT PROVIDED. THE DEVICES WILL BE RETURN. | CONCOMITANT PRODUCTS: LEFT SIZE 4 PS FEMUR (CAT# 02-010-01-0240 / SERIAL# (B)(4) ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00458 | 2022-04-25 | 2019-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRUL FEM IMPACTOR HEAD LOCKING | EXACTECH, INC. |  | 02-029-19-1000 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE INSTRUMENTATION. THE FEMORAL IMPACTOR BROKE OFF DURING IMPACTION OF FEMORAL COMPONENT. THE SECONDARY FEMORAL IMPACTOR WAS USED TO FINISH CASE. THERE WAS NO DELAY. THE DEVICE WAS RECEIVED FOR ANALYSIS. THE REPRESENTATIVE WAS PRESENT AT THE TIME OF SURGERY. ATTEMPT WAS MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REPRESENTATIVE AND A RESPONSE WAS RECEIVED. PER THE REPRESENTATIVE, THERE WAS NO ADVERSE OR CLINICAL CONSEQUENCES TO THE PATIENT. NO PARTS FELL INTO THE SURGICAL SITE. NO OTHER INFORMATION HAS BEEN RECEIVED ABOUT THE PATIENT OR EVENT FROM THE CONTACTS. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE REPORTED FRACTURED IMPACTOR FACE APPEARS CONSISTENT WITH MATERIAL DEGRADATION DUE TO NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. MOST LIKELY UNDERLYING CAUSE: BASED ON SIMILAR PREVIOUS COMPLAINT INVESTIGATIONS, THE FRACTURED IMPACTOR FACE REPORTED WAS LIKELY THE RESULT OF MATERIAL DEGRADATION ASSOCIATED WITH YEARS OF NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. RMR REVIEW: THE RMR FOR THIS TRULIANT FEMORAL LOCKING IMPACTOR HEAD, 750-2014-006-RMR-INSTRUMENTS (B)(4) WAS REVIEWED. THE RISK IS CAPTURED. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. ¿ CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. NO PARTS/PIECES FELL INTO THE SURGICAL SITE FROM THE BROKEN DEVICE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00457 | 2022-04-25 | 2019-06-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRUL FEM IMPACTOR HEAD LOCKING | EXACTECH, INC. |  | 02-029-19-1000 |  | IT WAS REPORTED INTERNALLY THAT BROKEN DEVICES WERE LOCATED IN 05BULK. THE DEVICES WERE FOUND TO HAVE DISASSEMBLED SCREW CAP AND WASHERS AND FRACTURED FACES. THERE IS NO ADDITIONAL INFORMATION ON THE DEVICES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. THERE IS NO REPORTED PATIENT IMPACT OR HOW THE DEVICES WERE USED. THE PART NUMBER IS 02-029-19-1000, AND THE LOT NUMBERS ARE 89989003 AND 82387001. THIS IS 2 OF 2 MDRS FOR THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. THE REPORTED FRACTURED IMPACTOR FACES REPORTED APPEAR CONSISTENT WITH MATERIAL DEGRADATION DUE TO NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. THE RMR FOR THESE TRULIANT FEMORAL LOCKING IMPACTOR HEADS, 750-2014-006-RMR-INSTRUMENTS REV C, WAS REVIEWED. THE RISK IS CAPTURED. MOST LIKELY UNDERLYING CAUSE: BASED ON SIMILAR PREVIOUS COMPLAINT INVESTIGATIONS, THE FRACTURED IMPACTOR FACES REPORTED ARE LIKELY THE RESULT OF MATERIAL DEGRADATION ASSOCIATED WITH NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. BROKEN CAP SCREW: (B)(4) AND CAPA(B)(4) 05BULK: (B)(4) AND CAPA(B)(4). BROKEN CAP SCREW: BOTH OF THE FOUND DEVICES WERE MISSING THEIR CAP SCREW. SINCE THE INTRODUCTION OF THE TRULIANT FEMORAL LOCKING IMPACTOR HEAD IN 2017, THERE HAS BEEN 17 COMPLAINT REPORTS INVOLVING TRULIANT FEMORAL LOCKING IMPACTOR HEADS (CATALOG #: 02-029-19-1000). OF THESE COMPLAINT REPORTS, 9 WERE DUE TO DISASSEMBLY OF THE CAP SCREW AND WASHER. INVESTIGATIONS REVEALED THAT THE FAILURE OF THE CAP SCREW WAS DIRECTLY RELATED TO INADEQUATE WELDING BETWEEN THE CAP SCREW AND THE WASHER SUBCOMPONENT APPLIED DURING MANUFACTURING ASSEMBLY. THERE WAS ONLY A MINIMAL SPOT WELD APPLIED WHICH WAS VULNERABLE DURING IMPACTION USE. A REVIEW OF THE PRINT SHOWED THE WELD REQUIREMENT WAS CALLED OUT AS A ¿TACK WELD¿, WHICH WAS NOT SPECIFIC ENOUGH AND LEFT ROOM FOR INTERPRETATION. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: THE ROOT CAUSE OF THE DISASSEMBLED SCREW CAP AND WASHER, AS DETERMINED BY CAPA(B)(4), WAS LIKELY THE RESULT OF THE WELDING REQUIREMENT NOT BEING ADEQUATELY SPECIFIED. BASED ON SIMILAR PREVIOUS COMPLAINT INVESTIGATIONS, THE FRACTURED IMPACTOR FACES REPORTED ARE LIKELY THE RESULT OF MATERIAL DEGRADATION ASSOCIATED WITH NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. THERE IS NO REPORTED PATIENT IMPACT OR HOW THE DEVICES WERE USED. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00456 | 2022-04-25 | 2022-04-07 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 4 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 4 13MM | 02-012-35-4013 |  | AS REPORTED BY LEGAL TEAM, APPROXIMATELY SIX YEARS POST INITIAL LEFT TKA, THIS 62 Y/O MALE PATIENT HAS BEEN EXPERIENCING PAIN AND TIGHTNESS IN BOTH KNEES LEFT KNEE WAS REVISED FOR POTENTIAL WEAR AND RIGHT KNEE TO BE REVISED IN THE FUTURE. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00455 | 2022-04-25 | 2019-06-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRUL FEM IMPACTOR HEAD LOCKING | EXACTECH, INC. |  | 02-029-19-1000 |  | IT WAS REPORTED INTERNALLY THAT BROKEN DEVICES WERE LOCATED IN 05BULK. THE DEVICES WERE FOUND TO HAVE DISASSEMBLED SCREW CAP AND WASHERS AND FRACTURED FACES. THERE IS NO ADDITIONAL INFORMATION ON THE DEVICES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. THERE IS NO REPORTED PATIENT IMPACT OR HOW THE DEVICES WERE USED. THE PART NUMBER IS 02-029-19-1000, AND THE LOT NUMBERS ARE 89989003 AND 82387001. THIS IS 1 OF 2 MDRS. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. THE REPORTED FRACTURED IMPACTOR FACES REPORTED APPEAR CONSISTENT WITH MATERIAL DEGRADATION DUE TO NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. THE RMR FOR THESE TRULIANT FEMORAL LOCKING IMPACTOR HEADS, 750-2014-006-RMR-INSTRUMENTS REV C, WAS REVIEWED. THE RISK IS CAPTURED. MOST LIKELY UNDERLYING CAUSE: BASED ON SIMILAR PREVIOUS COMPLAINT INVESTIGATIONS, THE FRACTURED IMPACTOR FACES REPORTED ARE LIKELY THE RESULT OF MATERIAL DEGRADATION ASSOCIATED WITH NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. BROKEN CAP SCREW: (B)(4) AND CAPA(B)(4) 05BULK: (B)(4) AND CAPA(B)(4) BROKEN CAP SCREW: BOTH OF THE FOUND DEVICES WERE MISSING THEIR CAP SCREW. SINCE THE INTRODUCTION OF THE TRULIANT FEMORAL LOCKING IMPACTOR HEAD IN 2017, THERE HAS BEEN 17 COMPLAINT REPORTS INVOLVING TRULIANT FEMORAL LOCKING IMPACTOR HEADS (CATALOG #: 02-029-19-1000). OF THESE COMPLAINT REPORTS, 9 WERE DUE TO DISASSEMBLY OF THE CAP SCREW AND WASHER. INVESTIGATIONS REVEALED THAT THE FAILURE OF THE CAP SCREW WAS DIRECTLY RELATED TO INADEQUATE WELDING BETWEEN THE CAP SCREW AND THE WASHER SUBCOMPONENT APPLIED DURING MANUFACTURING ASSEMBLY. THERE WAS ONLY A MINIMAL SPOT WELD APPLIED WHICH WAS VULNERABLE DURING IMPACTION USE. A REVIEW OF THE PRINT SHOWED THE WELD REQUIREMENT WAS CALLED OUT AS A ¿TACK WELD¿, WHICH WAS NOT SPECIFIC ENOUGH AND LEFT ROOM FOR INTERPRETATION. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: THE ROOT CAUSE OF THE DISASSEMBLED SCREW CAP AND WASHER, AS DETERMINED BY CAPA(B)(4), WAS LIKELY THE RESULT OF THE WELDING REQUIREMENT NOT BEING ADEQUATELY SPECIFIED. BASED ON SIMILAR PREVIOUS COMPLAINT INVESTIGATIONS, THE FRACTURED IMPACTOR FACES REPORTED ARE LIKELY THE RESULT OF MATERIAL DEGRADATION ASSOCIATED WITH NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. THERE IS NO REPORTED PATIENT IMPACT OR HOW THE DEVICES WERE USED. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00453 | 2022-04-25 | 2019-03-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRUL FEM IMPACTOR HEAD LOCKING | EXACTECH, INC. |  | 02-029-19-1000 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE INSTRUMENT THAT FRACTURED DURING USE. THIS DID NOT CAUSE A DELAY IN CASE, NOR ANY COMPLICATIONS. PATIENT LEFT THE OPERATING ROOM IN STABLE CONDITION. THE DEVICE WAS RECEIVED FOR ANALYSIS. THE REPRESENTATIVE WAS PRESENT AT THE TIME OF SURGERY. INACTIVE AGENCY WITH NO HOSPITAL/OTHER CONTACT INFORMATION FOR THE EVENT. NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. UNKNOWN DATE OF DEVICE RETURN. FINDINGS: THE REPORTED FRACTURED IMPACTOR FACE APPEARS CONSISTENT WITH MATERIAL DEGRADATION DUE TO NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. MOST LIKELY UNDERLYING CAUSE: BASED ON SIMILAR PREVIOUS COMPLAINT INVESTIGATIONS, THE FRACTURED IMPACTOR FACE REPORTED WAS LIKELY THE RESULT OF MATERIAL DEGRADATION ASSOCIATED WITH YEARS OF NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. RMR REVIEW: THE RMR FOR THIS TRULIANT FEMORAL LOCKING IMPACTOR HEAD, 750-2014-006-RMR-INSTRUMENTS REV C, WAS REVIEWED. THE RISK IS CAPTURED. IFU 700-096-181: INSTRUMENT INSPECTION. VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00452 | 2022-04-25 | 2018-12-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRUL FEM IMPACTOR HEAD LOCKING | EXACTECH, INC. |  | 02-029-19-1000 |  | IT WAS REPORTED FROM THE US THAT THE SURGEON WAS IMPACTING WITH THE DEVICE WHEN THE IMPACTOR BROKE AT THE BOLT THAT HOLDS THE CLAMP ARM ON. ALL PIECES WERE ACCOUNT FOR AND THERE WAS NOTHING LEFT IN WOUND SITE. THE AGENT WAS PRESENT AT SURGERY WITH NO ADVERSE EVENT TO THE PATIENT. THE AGENCY ASSOCIATED WITH THIS EVENT IS INACTIVE. NO FURTHER INFORMATION IS AVAILABLE. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE WAS RETURNED, UNABLE TO LOCATE FOR EVALUATION. DESIGN-RELATED ISSUES: THE DESIGN OF THE TRULIANT FEMORAL IMPACTOR LOCKING HEAD HAS BEEN IN THE FIELD SINCE 2017. EXACTECH IS AWARE OF 6 SIMILAR COMPLAINT REPORTS INVOLVING TRULIANT FEMORAL IMPACTOR LOCKING HEADS DUE TO A FRACTURED THREAD SHAFT AND AN INADEQUATE WELDING BETWEEN THE SCREW AND WASHER SUBCOMPONENT SINCE 2017. BECAUSE THE TRULIANT FEMORAL IMPACTOR LOCKING HEAD IS USED IN TOTAL KNEE SURGERIES, SALES DATA FOR TRULIANT FEMORAL COMPONENTS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4). THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THE INADEQUATE WELD FAILURE WAS PREVIOUSLY INVESTIGATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF (B)(4) UNITS, WHICH HAVE BEEN IN THE FIELD SINCE 2017. THIS ISSUE REGARDING THE INADEQUATE WELD WAS NOT FOUND TO BE MANUFACTURING RELATED. (B)(4) FOR ADDITIONAL INFORMATION. A REVIEW OF THE RISK MANAGEMENT REPORT(RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS " VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. INADEQUATE INFORMATION WAS PROVIDED TO DETERMINE THE CAUSE OF THE BROKEN DEVICE REPORTED. HOWEVER; IT MAY HAVE BEEN THE RESULT OF EITHER APPLYING A BENDING MOMENT TO THE LOCKING ARM, WHICH LED TO PLASTIC DEFORMATION AND ULTIMATE FRACTURE OF THE SCREW THREADS OR INADEQUATE WELDING BETWEEN THE SCREW AND THE WASHER SUBCOMPONENT APPLIED DURING MANUFACTURING ASSEMBLY. IFU 700-096-181: INSTRUMENT INSPECTION VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. THERE WAS NO REPORTED PATIENT ADVERSE EVENT. FAILURES CAN BE CONFIRMED BECAUSE THE DEVICE WAS NOT AVAILABLE FOR EVALUATION AND NO FURTHER INFORMATION WAS PROVIDED. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00451 | 2022-04-25 | 2019-04-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRUL FEM IMPACTOR HEAD LOCKING | EXACTECH, INC. |  | 02-029-19-1000 |  | IT WAS REPORTED FROM THE US THAT ¿UPON IMPACT THE PLASTIC SHIM BROKE¿. NO PATIENT OR EVENT INFORMATION WAS PROVIDED. THE DEVICE WAS RETURNED. NO INFORMATION PROVIDED DURING INITIAL INVESTIGATION. INACTIVE AGENCY WITH NO HOSPITAL/OTHER CONTACT INFORMATION FOR THE EVENT. NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH (B)(4) HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PATIENT ID, AGE, GENDER, WEIGHT, ETHNICITY AND RACE. FINDINGS: REVIEW OF THE COMPLAINT SHOWED CAP SCREW WAS LOOSENED DURING SURGERY. INVESTIGATION REVEALED THAT THE FAILURE OF THE CAP SCREW WAS DIRECTLY RELATED TO INADEQUATE WELDING BETWEEN THE CAP SCREW AND THE WASHER SUBCOMPONENT APPLIED DURING MANUFACTURING ASSEMBLY. THERE WAS ONLY A MINIMAL SPOT WELD APPLIED WHICH WAS VULNERABLE DURING IMPACTION USE. A REVIEW OF PRINT SHOWED THE WELD REQUIREMENT WAS CALLED OUT AS ¿TACK WELD¿, WHICH WAS NOT SPECIFIC ENOUGH AND LEFT ROOM FOR INTERPRETATION. SINCE THE INTRODUCTION OF THE TRULIANT FEMORAL LOCKING IMPACTOR HEAD IN 2017, THERE HAVE BEEN (B)(4) COMPLAINT REPORTS INVOLVING TRULIANT FEMORAL LOCKING IMPACTOR HEAD (CATALOG #: 02-029-19-1000). OF THESE COMPLAINT REPORTS, (B)(4) WERE DUE TO DISASSEMBLY OF THE CAP SCREW AND WASHER. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: THE ROOT CAUSE OF THE DISASSEMBLED SCREW CAP AND WASHER, AS DETERMINED BY CAPA (B)(4), WAS LIKELY THE RESULT OF THE WELDING REQUIREMENT NOT BEING ADEQUATELY SPECIFIED. CORRECTIVE ACTION TAKEN: THE CURRENT INVESTIGATION AND CAPA ARE STILL ONGOING AT THE TIME OF THIS EVALUATION, AND CORRECTIVE ACTIONS ARE SUBJECT TO CHANGE. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO DESIGN ISSUES. THERE IS NO PATIENT OR EVENT INFORMATION. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00450 | 2022-04-25 | 2019-07-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | FLEX SHAFT - MFR 2 TELEFLEX | EXACTECH, INC. |  | 181-00-02 |  | IT WAS REPORTED BY EXACTECH INTERNAL THAT THE DEVICE WAS FOUND ON (B)(6) 2019 IN 16ADBULK WITH A NOTE READING ¿BROKEN FLEXI DRILL GUIDE.¿ THERE IS NO REPORTED PATIENT ADVERSE EVENT DUE TO THE FRACTURED TIP FROM A FACILITY OR AGENT. NO OTHER INFORMATION ABOUT THIS DEVICE IS AVAILABLE. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT DUE TO THE FRACTURED TIP FROM A FACILITY OR AGENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. DESIGN-RELATED ISSUES: THE DESIGN OF THE FLEXIBLE DRILL SHAFT HAS BEEN IN THE FIELD SINCE 2013. (B)(4). THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF (B)(4) UNITS, WHICH HAVE BEEN IN THE FIELD SINCE 2016. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR. THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF EXCESSIVE BENDING FORCES APPLIED TO THE DEVICE, CAUSING CRACK INITIATION, PROPAGATION, AND ULTIMATE BRITTLE FRACTURE OF THE TIP. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT DUE TO THE FRACTURED TIP FROM A FACILITY OR AGENT. AN INVESTIGATION WAS CONDUCTED: THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF EXCESSIVE BENDING FORCES APPLIED TO THE DEVICE, CAUSING CRACK INITIATION, PROPAGATION, AND ULTIMATE BRITTLE FRACTURE OF THE TIP. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00454 | 2022-04-25 | 2019-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRUL FEM IMPACTOR HEAD LOCKING | EXACTECH, INC. |  | 02-029-19-1000 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE INSTRUMENT THAT FRACTURED DURING USE. THIS DID NOT CAUSE A DELAY IN CASE, NOR ANY COMPLICATIONS. PATIENT LEFT THE OPERATING ROOM IN STABLE CONDITION. THE DEVICE WAS RECEIVED FOR ANALYSIS. THE REPRESENTATIVE WAS PRESENT AT THE TIME OF SURGERY. INACTIVE AGENCY WITH NO HOSPITAL/OTHER CONTACT INFORMATION FOR THE EVENT. NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE REPORTED FRACTURED IMPACTOR FACE APPEARS CONSISTENT WITH MATERIAL DEGRADATION DUE TO NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. MOST LIKELY UNDERLYING CAUSE: BASED ON SIMILAR PREVIOUS COMPLAINT INVESTIGATIONS, THE FRACTURED IMPACTOR FACE REPORTED WAS LIKELY THE RESULT OF MATERIAL DEGRADATION ASSOCIATED WITH YEARS OF NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. RMR REVIEW: THE RMR FOR THIS TRULIANT FEMORAL LOCKING IMPACTOR HEAD, 750-2014-006-RMR-INSTRUMENTS REV C, WAS REVIEWED. THE RISK IS CAPTURED. IFU (B)(4): INSTRUMENT INSPECTION VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00466 | 2022-04-26 | 2021-08-21 | I, n, j, u, r, y | Manufacturer report | LOGIC CR TIB INSERT STD, SZ 2.5, 9 MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC CR TIB INSERT STD, SZ 2.5, 9 MM | 02-012-47-2509 |  | AS REPORTED BY THE LEGAL TEAM, APPROXIMATELY THREE YEARS POST INITIAL TKA, THIS (B)(6) MALE PATIENT UNDERWENT A REVISION FOR AN UNKNOW REASON. NO ADDITIONAL INFORMATION AVAILABLE. | CONCOMITANT MEDICAL PRODUCTS: LOGIC CR FEMORAL CEM, LEFT SZ 2.5 (CAT# 02-010-03-0225 / SERIAL# (B)(4)); LGC TIBIAL FIT TRAY CEM SZ 2.5F / 3.5T (CAT# 02-012-45-2535 / SERIAL# (B)(4)) ; THREE PEG PATELLA 32MM (CAT# 200-02-32 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10088 | 2022-04-26 | 2017-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | STEM TRIAL SIZE 3 RIGHT | EXACTECH, INC. |  | SK-015132 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGICAL PROCEDURE THAT AFTER THE TRIAL WAS USED, THE SCRUB TECH NOTICED THE TRIAL WAS BROKEN WHEN IT WAS ON THE BACK TABLE. THERE WERE NO PARTS OR PIECES THAT FELL INTO THE SURGICAL WOUND, THERE WAS NO DELAY IN SURGERY. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS ((B)(4)). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE COMPLAINT DEVICE WAS RECEIVED FOR ANALYSIS. THE COMPLAINT CONDITION WAS CONFIRMED. VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR 10X OPTICAL) T MISSING SCREW FROM STEM TRIAL ADAPTOR- THE STEM TRIAL ADAPTOR FEATURES TWO SCREW HOLES THAT UTILIZE A RETAINING RING TO HOLD A WASHER AND SCREW IN PLACE. THE MISSING SCREW APPEARS CONSISTENT WITH THE DISLOCATION OF THE RETAINING RING LIKELY FROM APPLYING A FORCE TO THE RING WHILE FASTENING THE ADAPTOR TO A FEMORAL TRIAL. AS THE ADAPTOR IS BEING FASTENED, THE RETAINING RING CAN BECOME COMPRESSED AGAINST THE SCREW HEAD. IF THE RING IS NOT SEATED INTO THE GROOVE PROPERLY, THE SCREW HAS THE POTENTIAL TO OVERCOME THE FORCE NEEDED TO KEEP THE RING IN ITS GROOVE. DEFORMED RETAINING RING- THE DEFORMATION APPEARS CONSISTENT WITH ATTEMPTING TO REMOVE AND/OR REASSEMBLE THE RING. IT CAN ALSO BE SEEN THAT THE RETAINING RING IN PLACE, THE HOSPITAL LIKELY REASSEMBLED THE DEVICE PRIOR TO RETURNING. THE IMAGE PROVIDED IN THE EXPERIENCE REPORT SHOWS THE RING, SCREW, AND WASHER HAD DISLOCATED AND IT WAS REPORTED THAT THE BROKEN INSTRUMENT HAD BEEN DISCOVERED AFTER USE ON THE BACK TABLE. HOWEVER, IF THE DEVICE WERE TO DISASSEMBLE INSIDE THE PATIENT, THE COMPONENTS ARE LARGE ENOUGH TO BE EASILY REMOVED FROM THE JOINT SPACE. THE MATERIALS FOR THE WASHER AND THE RETAINING RING, 18-8 STAINLESS STEEL, AND OF THE SCREW, 17-4 STAINLESS STEEL, ARE INCLUDED IN (B)(4) - MASTER BIOCOMPATIBILITY MEMO FOR INSTRUMENTS ASSOCIATED WITH PROJECT NUMBER (B)(4). ALL OTHER ASPECTS OF THE DEVICES APPEAR UNREMARKABLE. THIS IS THE ONLY PART FROM THIS LOT, WHICH HAS BEEN IN THE FIELD SINCE 2016. THE FREQUENCY OF OCCURRENCE RANKING SCALE IS VERY LOW; THEREFORE, THIS DOES NOT APPEAR TO BE DESIGN RELATED. THIS IS THE ONLY PART FROM THIS LOT, WHICH HAS BEEN IN THE FIELD SINCE 2016. IN ORDER TO CREATE THIS SPECIAL INSTRUMENT, THE STANDARD LOGIC STEM TRAIL ADAPTOR WAS FIRST DISASSEMBLED BY REMOVING THE RETAINING RING AND SCREW. THE DEVICE WAS THEN REASSEMBLED WITH THE ADDITION OF A WAVE WASHER. DURING THIS TIME, THE RETAINING RING WAS LIKELY NOT SEATED IN ITS GROOVE PROPERLY, ALLOWING THE DEVICE TO DISASSEMBLE DURING SURGERY. THE DISASSEMBLY REPORTED WAS LIKELY THE RESULT OF APPLYING A FORCE TO A RETAINING RING THAT WAS NOT PROPERLY SEATED IN ITS GROOVE, ALLOWING THE RING TO DISLOCATE. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY DESIGN ISSUES, NOR DID IT LEAD TO PATIENT IMPACT. THE DEVICE WAS DISCOVERED BROKEN ON THE BACK TABLE AFTER IT WAS USED IN THE SURGICAL PROCEDURE. AN INVESTIGATION WAS CONDUCTED; THE DISASSEMBLY REPORTED WAS LIKELY THE RESULT OF APPLYING A FORCE TO A RETAINING RING THAT WAS NOT PROPERLY SEATED IN ITS GROOVE, ALLOWING THE RING TO DISLOCATE. INFORMATION WAS NOT PROVIDED FOR: AGE OR DATE OF BIRTH, SEX, WEIGHT, ETHNICITY, RACE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10085 | 2022-04-26 | 2019-09-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | GLENOID SKID | EXACTECH, INC. |  | 315-03-01 |  | EXACTECH COMPLAINTS RECEIVED DEVICES VIA AN RG TRANSFER REPORT FROM AND AGENCY INDICATING THE DEVICES WERE DEFECTIVE/COMPLAINT ITEMS FROM RG479524. NOT ALL PIECES OF THE BROKEN DEVICES WERE RETURNED. NO CUSTOMER OR PATIENT IMPACT INFORMATION WAS PROVIDED. ATTEMPTS WERE MADE TO REQUEST ADDITIONAL INFORMATION FROM THE AGENCY AND A RESPONSE WAS RECEIVED TO ASK WHO REPORTED THE EVENT. THERE WAS NO REPORTING REPRESENTATIVE INDICATED ON THE RG. NO ADDITIONAL INFORMATION WAS PROVIDED ABOUT THESE DEVICES. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICES WERE AVAILABLE FOR ANALYSIS. ENGINEERING ANALYSIS COMPLETED BY (B)(6) ON 12/2/2019. FINDINGS: THE BROKEN GLENOID SKIDS APPEAR CONSISTENT WITH APPLYING A BENDING MOMENT DURING USE. THE OTHER PIECES OF THE GLENOID SKIDS WERE NOT RETURNED. MOST LIKELY UNDERLYING CAUSE: BASED ON PREVIOUS SIMILAR INVESTIGATIONS, THE BROKEN GLENOID SKIDS REPORTED IN (B)(4) WERE LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO ONE END OF THE INSTRUMENT, WHICH LED TO FRACTURE OF THE DEVICES. HOWEVER, NO INFORMATION REGARDING THEIR USE DURING SURGERY WAS PROVIDED. RMR REVIEW: THE RMR AND RACR FOR THIS GLENOID SKID, RMR-00006 REV - AND RACR-00009 REV -, WERE REVIEWED. THE RISK FOR THE REPORTED FAILURE IS NOT CAPTURED FOR THIS DEVICE. THE RACR WILL BE UPDATED UNDER DCRSP-19-541 TO CAPTURE THE APPROPRIATE RISK FOR THIS DEVICE. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICES USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT FROM THE AGENCY. NOT ALL PIECES OF THE DEVICES WERE RETURNED. AN INVESTIGATION WAS CONDUCTED: BASED ON PREVIOUS SIMILAR INVESTIGATIONS, THE BROKEN GLENOID SKIDS REPORTED WERE LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO ONE END OF THE INSTRUMENT, WHICH LED TO FRACTURE OF THE DEVICES. HOWEVER, NO INFORMATION REGARDING THEIR USE DURING SURGERY WAS PROVIDED. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10072 | 2022-04-26 | 2017-06-27 | I, n, j, u, r, y | Manufacturer report | EXACTECH | DRIVER, MODULAR CANNULATED TRI-DRIVE | EXACTECH, INC. |  | 315-25-00 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE INSTRUMENTATION. THE TIP APPEARS TO HAVE DISASSEMBLED FROM THE TRI-DRIVER. THE RETAINING RING IS MEANT TO HOLD THE SLEEVE IN PLACE BUT APPEARS TO HAVE DISLODGED. THE SLEEVE, SPRING, AND BALL BEARING WERE NOT RETURNED. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT FOLLOWING MULTIPLE ATTEMPTS. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE AVAILABLE FOR ANALYSIS. ANALYSIS COMPLETED ON 17 SEP 2018 BY MB. VISUAL EVALUATION CONDITIONS/FINDINGS: THE TIP APPEARS TO HAVE DISASSEMBLED. THE RETAINING RING IS MEANT TO HOLD THE SLEEVE IN PLACE BUT APPEARS TO HAVE DISLODGED. THE SLEEVE, SPRING, AND BALL BEARING WERE NOT RETURNED. THESE TRI-DRIVE HANDLES WERE MATED WITH A CANNULATED REAMER HEAD, BUT THE SLEEVES WOULD NOT FULLY RETRACT OR ATTACH TO THE MATING FEATURE OF THE REAMER HEAD. DESIGN RELATED ISSUES: THE DESIGN OF THE MODULAR CANNULATED TRI-DRIVE HAS BEEN IN THE FIELD SINCE 2009. EXACTECH HAS RECEIVED 38 SIMILAR COMPLAINT REPO1TS INVOLVING 54 DEVICES SINCE 2009. BECAUSE THIS DEVICE IS USED IN TOTAL SHOULDER SURGERIES, SALES DATA FOR HUMERAL STEMS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: DHRS WERE NOT REVIEWED BECAUSE EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS ASSOCIATED WITH THE LOT. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THESE TRI-DRIVE MODULAR REAMER DRIVERS, 750-2007-026-RMR-INSTRUMENTS REV H, WAS REVIEWED. THE RISK IS NOT CAPTURED. THIS DOCUMENT WILL BE UPDATED TO INCLUDE THIS FAILURE. REFER TO DCRSP-18-340 FOR DETAILS. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATES ARE " VERY LOW", AND THE RMR WILL BE UPDATED TO CAPTURE THESE RISKS. MOST LIKELY CAUSE: THE DISASSEMBLED DRIVER REPORTED IN EXPERIENCE C2017-365 WAS LIKELY THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY AND ALLOWED THE TIP TO DISASSEMBLE. THE JAM WAS LIKELY THE RESULT OF A BONE FRAGMENT OR OTHER FOREIGN OBJECT BECOMING LODGED IN THE MECHANISM PREVENTING THE SLEEVE FROM FULLY RETRACTING. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES. THE DISASSEMBLY OF THE DRIVER DID NOT LEAD TO PATIENT IMPACT. THE RETAINING RING IS MEANT TO HOLD THE SLEEVE IN PLACE BUT APPEARS TO HAVE DISLODGED. AN INVESTIGATION WAS CONDUCTED; THE DISASSEMBLED DRIVER REPORTED WAS LIKELY THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY AND ALLOWED THE TIP TO DISASSEMBLE. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10046 | 2022-04-26 | 2017-06-06 | I, n, j, u, r, y | Manufacturer report | EXACTECH | LOGIC FEMORAL, NOTCH CUTTING GUIDE, SIZE 5 | EXACTECH, INC. |  | 02-019-10-0150 |  | IT WAS REPORTED FROM THE US THAT DURING A SURGICAL PROCEDURE THE LOCKING RING POPPED OUT OF A FEMORAL NOTCH CUTTING GUIDE. ALSO, A LOCKING TIBIAL TRAY IMPACTOR WOULD NOT LOCK ONTO THE IMPLANT. INSTRUMENTS WERE TAKEN FROM OTHER SET TO COMPLETE THE SURGICAL CASE. THERE WAS NO NEGATIVE AFFECT AS REPORTED ABOUT THE PATIENT. NO PIECE WAS LEFT IN THE PATIENT AND THERE WAS NO SURGICAL DELAY. THE AGENT/AGENCY IS NO LONGER REPRESENTING THE COMPANY, NO ADDITIONAL INFORMATION IS AVAILABLE ABOUT THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE COMPLAINT PRODUCTS WERE RECEIVED FOR ANALYSIS. THE LOGIC FEMORAL, NOTCH CUTTING GUIDE WAS FOUND TO HAVE A MISSING RING, COMPLAINT CONFIRMED. THE LOCKING TIBIAL TRAY IMPACTOR WAS FOUND TO OPERATE AS INTENDED, COMPLAINT NOT CONFIRMED. VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR 10X OPTICAL)- A PORTION OF THE FEMORAL NOTCH CUTTING GUIDE WAS LIKELY IMPACTED TO SEAT THE GUIDE FLUSH AGAINST THE RESECTION FACE OF THE FEMORAL BONE. THIS APPEARS TO HAVE DEFORMED THE GUIDE ALLOWING FOR THE LOCKING RING TO DISASSOCIATE. THE FREQUENCY OF OCCURRENCE RANKING SCALE IS VERY LOW; THEREFORE, THIS DOES NOT APPEAR TO BE A DESIGN ISSUE. FEMORAL NOTCH CUTTING GUIDE, EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINT INVOLVING PARTS FROM THIS MANUFACTURING LOT, WHICH HAS BEEN IN THE FIELD SINCE 2012. A REVIEW OF THE DEVICE HISTORY RECORD ALSO SHOWED THAT THE NAMED DEVICE WAS ACCEPTED WITH CONFORMANCE TO THE DEVICE REQUIREMENTS. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING-RELATED. THERE ARE NO USER-RELATED ISSUES REPORTED THAT APPEAR TO HAVE CONTRIBUTED TO THE REPORTED EVENT. THE DESIGN OF THE LOGIC FEMORAL NOTCH CUTTING GUIDE HAS BEEN IN THE FIELD SINCE 2011. EXACTECH IS AWARE OF ONE COMPLAINT INVOLVING LOGIC FEMORAL NOTCH CUTTING GUIDES DUE TO LOCKING RING DISASSOCIATION SINCE 2014. SINCE THIS DEVICE IS USED TO PREPARE FOR LOGIC FEMORAL IMPLANTS, THE SALES DATA FOR LOGIC FEMORAL IMPLANTS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF LESS THAN (B)(4). THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR 10X OPTICAL)- THE LOCKING TIBIAL TRAY IMPACTOR HAD A FUNCTIONAL CHECK WHICH CONFIRMED THE INSTRUMENT SUCCESSFULLY MATED AND LOCKED WITH THE TIBIAL TRAY, PN: 02-012-41-5050 (B)(4).THEREFORE, THIS FAILURE COULD NOT BE RECREATED. THE FREQUENCY OF OCCURRENCE RANKING SCALE IS VERY LOW; THEREFORE, THIS DOES NOT APPEAR TO BE A DESIGN ISSUE. LOCKING TIBIAL TRAY IMPACTOR, EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT, WHICH HAS BEEN IN THE FIELD SINCE 2010. A REVIEW OF THE DEVICE HISTORY RECORD ALSO SHOWED THAT THE NAMED DEVICE WAS ACCEPTED WITH CONFORMANCE TO THE DEVICE REQUIREMENTS. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING-RELATED. THE USER APPEARS TO HAVE LOCKED THE IMPACTOR INCORRECTLY TO THE TIBIAL TRAY AND ASSUMED THE DEVICE WAS NO LONGER FUNCTIONAL. HOWEVER, THE FUNCTIONAL CHECK CONFIRMED THIS IS NOT THE CASE. THE DESIGN OF THE LOCKING TIBIAL TRAY IMPACTOR HAS BEEN IN THE FIELD SINCE 2003. EXACTECH IS AWARE OF ONE COMPLAINT INVOLVING LOCKING TIBIAL TRAY IMPACTORS INCORRECTLY MATING WITH TIBIAL TRAYS SINCE 2014. SINCE THIS DEVICE IS USED TO PREPARE FOR LOGIC FEMORAL IMPLANTS, THE SALES DATA FOR LOGIC FEMORAL IMPLANTS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF LESS THAN (B)(4). THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THE FEMORAL NOTCH CUTTING GUIDE ISSUE REPORTED IN (B)(6) LIKELY FAILED DUE TO IMPACTION OF THE GUIDE SURFACE, WHICH LED TO THE DEFORMATION OF THE GUIDE, AND DISASSOCIATION OF THE LOCKING RING. THE TIBIAL TRAY IMPACTOR ASSEMBLY ISSUE REPORTED IN EXPERIENCE (B)(6) COULD NOT BE REPRODUCED. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT IMPACT. AN INVESTIGATION WAS CONDUCTED; THE FEMORAL NOTCH CUTTING GUIDE LIKELY FAILED DUE TO IMPACTION OF THE GUIDE SURFACE, WHICH LED TO THE DEFORMATION OF THE GUIDE, AND DISASSOCIATION OF THE LOCKING RING AND THE TIBIAL TRAY IMPACTOR ISSUE COULD NOT BE RECREATED. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10086 | 2022-04-26 | 2017-01-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRIPAN REAMER HANDLE | EXACTECH, INC. |  | 315-25-00 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGICAL PROCEDURE THE UNIVERSAL TRIPAN HANDLE BEING USED FOR REAMING AND DRILLING A CENTRAL CAGE ON THE BASEPLATE DEVELOPED A FAULT. THE RETRACTABLE METAL SLEEVE AT THE DISTAL END SOMEHOW CAME LOOSE FROM ITS FIXED POSITION (IT REMAINED ON THE MAIN SHAFT BUT DISSOCIATED FROM THE SPRING/BALL BEARING MECHANISM THAT HOLDS IT IN PLACE) AND AS SUCH WOULD NOT HOLD THE REAMERS IN POSITION. THE SURGICAL TEAM HAD TO OPEN A REVERSE TRAY TO USE ANOTHER INSTRUMENT. IT WAS STATED THAT "NO PIECES/FRAGMENTS WERE LEFT IN THE PATIENT". THERE WAS NO DETRIMENT TO PATIENT OR FLOW OF SURGERY. EMAIL REQUEST WERE MADE FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RETURNED. THE REPORTED MALFUNCTION WAS CONFIRMED DURING ANALYSIS OF THE DEVICE. THE RETAINING RING THAT IS MEANT TO HOLD THE SLEEVE-SPRING MECHANISM, APPEARS TO HAVE DISLODGED CAUSING THE DEVICE TO DISASSEMBLE. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE AND CONSISTENT WITH NORMAL USE. THE DEVICE WAS MANUFACTURED IN 2014. THE DEVICE HISTORY RECORD WAS REVIEWED, AND ALL PARTS WERE ACCEPTED WITH CONFORMANCE TO THE DEVICE SPECIFICATIONS. THE DESIGN OF THE TRI-DRIVE MODULAR HANDLE BEEN IN THE FIELD SINCE 2009. 38 SIMILAR COMPLAINT REPORTS INVOLVING 54 DEVICES SINCE ITS INTRODUCTION. BECAUSE THIS DEVICE IS USED IN BOTH ANATOMIC AND REVERSE TOTAL SHOULDER SURGERIES, THE SALES DATA FOR HUMERAL STEMS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW". 2 OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT CONTAINING 30 UNITS, WHICH HAS BEEN IN THE FIELD SINCE 2014. THE CAUSE OF THE FAILURE CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY IS RELATED TO THE USER HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY AND ALLOWED THE TIP TO DISASSEMBLE. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT IMPACT. IT WAS REPORTED THAT "NO PIECES/FRAGMENTS WERE LEFT IN THE PATIENT." AN INVESTIGATION WAS CONDUCTED; THE CAUSE OF THE FAILURE CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY IS RELATED TO THE USER HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY AND ALLOWED THE TIP TO DISASSEMBLE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10076 | 2022-04-26 | 2018-07-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRI-DRIVE MODULAR HANDLE | EXACTECH, INC. |  | 315-25-00 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGICAL PROCEDURE THERE WAS AN ISSUE WITH THE MODULAR REVERSE PILOT TIP REAMER. THE MODULAR REVERSE PILOT TIP REAMER SLIPPED OUT AFTER REAMING WITH STARTER REAMER. THE PATIENT WAS FINE LEAVING THE OPERATING ROOM. THE REP WAS PRESENT AT THE SURGERY. RG301619 WAS ISSUED FOR THE RETURN OF THE DEVICE, BUT THE RG# WAS CANCELED ON (B)(6) 2018 WITH NO DEVICE RETURN. AN ATTEMPT WAS MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REP AND A RESPONSE WAS RECEIVED. THE REP STATED THAT NO PIECES OF THE DEVICE(S) FELL INTO THE SURGICAL SITE. NO OTHER ADDITIONAL INFORMATION WAS PROVIDED. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS NOT AVAILABLE FOR ANALYSIS. ENGINEERING ANALYSIS COMPLETED BY MB ON 9/21/2018. DESIGN-RELATED ISSUES: THE DESIGN OF THE TRI-DRIVE MODULAR HANDLE HAS BEEN IN THE FIELD SINCE 2009. EXACTECH IS AWARE OF 38 SIMILAR COMPLAINT REPORTS INVOLVING 54 TRI-DRIVES SINCE 2009. BECAUSE THE TRI-DRIVE HANDLE IS USED IN ANATOMIC AND REVERSE TOTAL SHOULDER SURGERIES, SALES DATA FOR HUMERAL STEMS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINT REPORTS INVOLVING PARTS FORM THIS MANUFACTURING LOT. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS TRI-DRIVE MODULAR REAMER DRIVER, 750-2007-026-RMR-INSTRUMENT REV J, WAS REVIEWED. THE RISK IS NOT CAPTURED. THIS DOCUMENT WILL BE UPDATED TO INCLUDE THIS FAILURE. REFER TO DCRSP-18-340 FOR DETAILS. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RMR WILL BE UPDATED TO CAPTURE THIS RISK. THE DISASSEMBLED DRIVER REPORTED IN EXPERIENCE C2018-586 WAS LIKELY THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY AND ALLOWED THE TIP TO DISASSEMBLE. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. AN INVESTIGATION WAS CONDUCTED; THE DISASSEMBLED DRIVER REPORTED WAS LIKELY THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FORM THE ASSEMBLY AND ALLOWED THE TIP TO DISASSEMBLE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10055 | 2022-04-26 |  | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRULIANT TIBIAL TAMP GUIDE | EXACTECH, INC. |  | 02-029-22-2000 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE TRULIANT TIBIAL TAMP GUIDE BROKE. THERE WERE NO PIECES THAT FELL INTO THE SURGICAL SITE. THERE WAS A 5-MINUTE DELAY TO SURGERY WITH NO ADVERSE EVENT TO THE PATIENT. THE AGENT WAS PRESENT AT THE TIME OF SURGERY. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. MFG DATE: 11/14/2017. ENG EVAL COMPLETED ON 11/15/2018 BY (B)(4). PER INVESTIGATION CONDUCTED UNDER CAPA(B)(4), (B)(4) COMPLAINTS INCLUDING (B)(4) DEVICES WERE EVALUATED AND THE REPORTED CONDITION WAS CONFIRMED. THE ROOT CAUSE WAS DETERMINED TO BE THAT THE DESIGN OF THE SNAP RING THAT HOLDS THE BUTTON IN PLACE WAS NOT ROBUST. THE AMOUNT OF THE OVERLAP BETWEEN THE SNAP RING AND THE LIP OF THE HOUSING WAS ON THE LOW SIDE, WHICH INCREASES THE RISK OF DISENGAGEMENT. INVESTIGATION CONFIRMED THAT ALL PARTS WERE MANUFACTURED TO SPECIFICATION. A REVIEW OF THE SPECIFICATIONS ON THE ENGINEERING PRINT REVEALED THAT THE OVERLAP BETWEEN THE SNAP RING'S OD (.720+.000/-.004) AND THE ID OF THE HOUSING LIP (.705+/-.003) IS ON THE LOW SIDE. FROM THE HEALTH RISK INDEX (HRI) CONDUCTED FOR (B)(4): THE BUTTON COMPONENT FELL OFF OF THE TRULIANT TIBIAL TAMP GUIDE DURING VARIOUS OCCASIONS (USAGE, CLEANING, TRANSPORTATION, ETC.). (B)(4). THERE IS NO ANTICIPATED HARM TO THE PATIENT DUE TO THE DELAY IN SURGERY. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. PER THE TRILLIANT PRIMARY KNEE SYSTEMS OPERATIVE TECHNIQUE (LIT 712-35-31, REV A), THE TIBIAL TAMP GUIDE IS USED FOR TAMPING OF THE TIBIAL BASEPLATE TRIAL DURING KNEE IMPLANT SURGERY VIA IMPACTION AND RETRO-IMPACTION. THE DEVICE IFU (700-096-0160) WARNS THAT ONLY QUALIFIED SURGEONS KNOWLEDGEABLE IN ANATOMY, BIOMECHANICS, AND RECONSTRUCTIVE SURGERY SHOULD USE THESE DEVICES. THE SURGEON MUST BE FULLY KNOWLEDGEABLE ABOUT ALL ASPECTS OF THE TRULIANT FEMORAL COMPONENTS, TIBIAL INSERTS AND TIBIAL TRAYS SURGICAL TECHNIQUE AND USE THESE IMPLANTS IN ACCORDANCE WITH THE RESPECTIVE INDICATIONS AND CONTRAINDICATIONS SUMMARIZED IN THIS DOCUMENT. IN ADDITION, THE SURGEON SHOULD BE FULLY KNOWLEDGEABLE ABOUT THE COMPATIBILITY BETWEEN SYSTEM COMPONENTS AND USE ACCORDINGLY. FINALLY, THE SURGEON MUST BE TRAINED ON THE PROPER OPERATIVE TECHNIQUE INCLUDING THE PROPER USE OF SYSTEM INSTRUMENTATION. IT IS ESSENTIAL THAT THE PROPER SURGICAL TECHNIQUE BE USED WHEN SECURING LOCKING COMPONENTS. AGGRESSIVE IMPACTION TECHNIQUES AND FORCEFUL ASSEMBLY METHODS MAY DAMAGE COMPONENTS AND LEAD TO DEVICE FAILURE. REPEATED ASSEMBLY AND DISASSEMBLY OF MODULAR COMPONENTS COULD COMPROMISE THE STRENGTH OF LOCKING MECHANISMS. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INST RUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. THERE WAS NO PATIENT ADVERSE EVENT. PER INVESTIGATION CONDUCTED UNDER CAPA(B)(4), THE ROOT CAUSE WAS DETERMINED TO BE THAT THE DESIGN OF THE SNAP RING THAT HOLDS THE BUTTON IN PLACE WAS NOT ROBUST. THE AMOUNT OF THE OVERLAP BETWEEN THE SNAP RING AND THE LIP OF THE HOUSING WAS ON THE LOW SIDE, WHICH INCREASES THE RISK OF DISENGAGEMENT. INFORMATION WAS NOT PROVIDED FOR: PT. INFO.; DATE OF EVENT; INITIAL REPORTER NAME AND ADDRESS: | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10071 | 2022-04-26 | 2017-06-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | DRIVER, MODULAR CANNULATED, TRI-DRIVE | EXACTECH, INC. |  | 315-25-00 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE TRI-DRIVE. THE TIP DISASSEMBLED ON THE DRIVER. ANOTHER SET WAS OPENED AND USED IN THE PROCEDURE IT WAS CONFIRMED THAT THERE WERE NO PIECES LEFT IN THE PATIENT AND THE SURGICAL OUTCOME WAS GOOD. INACTIVE AGENCY WITH NO HOSPITAL/OTHER CONTACT INFORMATION FOR THE EVENT. NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4) HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS AVAILABLE FOR ANALYSIS. CRC2019-04-05 - MODULAR CANNULATED TRI-DRIVE DISASSEMBLY THE DESIGN OF THE MODULAR CANNULATED TRIDRIVE HAS BEEN IN THE FIELD SINCE 2009. EXACTECH HAS RECEIVED 38 GLOBAL EXPERIENCE REPORTS DOCUMENTING DISASSEMBLY IN 40 MODULAR CANNULATED TRIDRIVES. THESE TRIDRIVES ARE USED IN BOTH ANATOMIC AND REVERSE TOTAL SHOULDER ARTHROPLASTY, SO SALES DATA FOR HUMERAL STEMS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF 0.11%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THE DISASSEMBLED MODULAR CANNULATED TRIDRIVE MAY RESULT IN A DELAYED SURGERY TO FIND A NEW ONE. TYPICALLY, IN ANATOMIC SHOULDER CASES BOTH ANATOMIC AND REVERSE SETS ARE IN THE OR. IF THE MODULAR CANNULATED TRIDRIVE WERE TO BREAK, A SECOND WOULD BE IN THE REVERSE SET. THESE BROKEN TRI-DRIVES WERE LIKELY THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY, ALLOWING THE TIP TO DISASSEMBLE. THE CRC DETERMINED THAT NO FURTHER ACTION WAS REQUIRED AT THIS TIME AS THIS ISSUE WILL BE ADDRESSED VIA THE EQUINOXE ERGO INSTRUMENTATION REDESIGN PROJECT. IN THE NEW EQUINOXE ERGO INSTRUMENTATION, THE TRI-DRIVE HANDLES WILL NO LONGER BE USED. SINCE THE DESIGN OF THE PILOT TIP REAMERS ARE CHANGING (ALL THE REAMERS ARE CANNULATED) THERE WILL NOW BE TWO DIFFERENT HANDLES; A PILOT TIP HANDLE AND A CANNULATED HANDLE, REPLACING THE MODULAR TRI-DRIVE. THIS NEW ERGO INSTRUMENTATION DESIGN WILL BE IMPLEMENTED AS PART OF A PHASED ROLL-OUT, FIRST IN THE US AND THEN TO OTHER MARKETS, BEGINNING IN MID-2019. EXACTECH IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. NO REPORTED PATIENT ADVERSE EVENT. THE TIP DISASSEMBLED ON THE DRIVER; IT WAS CONFIRMED THAT THERE WERE NO PIECES LEFT IN THE PATIENT. AN INVESTIGATION WAS CONDUCTED: THESE BROKEN TRI-DRIVES WERE LIKELY THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY, ALLOWING THE TIP TO DISASSEMBLE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00465 | 2022-04-26 | 2021-12-15 | I, n, j, u, r, y | Manufacturer report | NV GXL LINER NEUTRAL, 32MM ID GROUP 1 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LINER NEUTRAL, 32MM ID GROUP 1 CUPS | 130-32-51 |  | AS REPORTED BY THE LEGAL TEAM, APPROXIMATELY SIX YEARS POST INITIAL THA, THIS (B)(6) FEMALE PATIENT UNDERWENT A REVISION DUE TO PREMATURE POLYETHYLENE FAILURE. NO ADDITIONAL INFORMATION AVAILABLE. | CONCOMITANT MEDICAL PRODUCTS :12/14 ZIRCONIA HEAD 32MM -3.5MM NECK (CAT# 148-32-93 / SERIAL# (B)(4)) ; NOVATION ELEMENT RO S/O SZ 8 (CAT# 164-11-08 / SERIAL# (B)(4)) ; ALTEON 6.5MM SCREW, 30MM (CAT# 180-65-30 / SERIAL# (B)(4)) ; INTEGRIP CC, CLUSTER 48MM, G1 (CAT# 186-01-48 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00467 | 2022-04-26 | 2020-12-16 | I, n, j, u, r, y | Manufacturer report | LOGIC CR TIB INSERT SLOPE +, SZ 3.5, 11 MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC CR TIB INSERT SLOPE +, SZ 3.5, 11 MM | 02-012-48-3511 |  | AS REPORTED BY THE LEGAL TEAM, APPROXIMATELY THREE YEARS POST INITIAL TKA, THIS (B)(6) MALE PATIENT UNDERWENT A REVISION FOR AN UNKNOWN REASON. NO ADDITIONAL INFORMATION AVAILABLE. | CONCOMITANT MEDICAL PRODUCTS: LOGIC CR FEMORAL CEM, LEFT, SZ 3.5 (CAT# 02-010-03-0235 / SERIAL# (B)(4)); LGC TIBIAL FIT TRAY CEM SZ 3.5F / 4.5T (CAT# 02-012-45-3545 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10095 | 2022-04-26 | 2016-11-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | HUMERAL HEAD TRIAL | EXACTECH, INC. |  |  |  | IT WAS REPORTED FROM THE US THAT DURING AN INITIAL ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE INSTRUMENTATION. THE SURGEON HAD DIFFICULTY SEATING SIZE 13 HUMERAL STEM (CAPTURED IN ANOTHER COMPLAINT). THE SURGEON BROKE 2 HUMERAL HEAD TRIALS TRYING TO SEAT STEM. THE SURGEON ENDED UP IMPLANTING A SIZE 9 STEM. THERE WAS MINIMUM DELAY, AND THE PATIENT WAS NOT ADVERSELY AFFECTED. THE PATIENT WAS DOING WELL AT THE END OF SURGERY. THE REPRESENTATIVE WAS UNABLE TO PROVIDE ANY CATALOG AND LOT/SERIAL NUMBERS OF THE HUMERAL HEAD TRIALS. THE DEVICES WERE NOT RETURNED FOR ANALYSIS. ATTEMPTS WERE MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REPRESENTATIVE, BUT NO RESPONSES WERE RECEIVED. THERE HAS BEEN NO OTHER INFORMATION PROVIDED BY THE CONTACTS TO THIS EVENT. THIS IS REPORT 2 OF 2 FOR THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DESIGN RELATED ISSUES: THE PART NUMBERS FOR THE BROKEN INSTRUMENTS WERE NOT PROVIDED. HOWEVER; THE DESIGN OF THE HUMERAL HEAD TRIAL (ALL SIZES) HAS BEEN IN THE FIELD SINCE 2004. EXACTECH HAS NOT RECEIVED ANY OTHER COMPLAINT REPORTS INVOLVING BREAKAGE OF A HUMERAL HEAD TRIAL. THE SALES DATA FOR ALL HUMERAL HEADS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4). THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS DOES NOT APPEAR TO BE A DESIGN RELATED ISSUE. MFG-RELATED ISSUES: MFG-RELATED ISSUES: THE MANUFACTURING LOT NUMBER WAS NOT PROVIDED. THEREFORE, THE MANUFACTURING DATA CANNOT BE REVIEWED. THIS ISSUE IS UNLIKELY RELATED TO THE MANUFACTURING. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THESE HUMERAL HEAD TRIALS, (B)(4), WAS REVIEWED. THE RISK IS CAPTURED. MOST LIKELY CAUSE: THE BROKEN HUMERAL HEAD TRIALS REPORTED WERE LIKELY THE RESULT OF EXCESSIVE IMPACTION DUE TO DIFFICULTY SEATING THE HUMERAL STEM.THIS CANNOT BE CONFIRMED BECAUSE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES. NO PATIENT ADVERSE EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00469 | 2022-04-26 | 2021-10-27 | I, n, j, u, r, y | Manufacturer report | LOGIC CR TIB INSERT SLOPE+, SZ 3, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC CR TIB INSERT SLOPE+, SZ 3, 9MM | 02-012-48-3009 |  | AS REPORTED BY THE LEGAL TEAM, APPROXIMATELY THREE YEARS POST INITIAL TKA, THIS (B)(6) MALE PATIENT UNDERWENT REVISION SURGERY DUE TO POLYTHYLENE FAILURE. DURING REVISION SURGERY, IT WAS NOTED THAT THE POLYETHYLENE FROM THE TIBIAL TRAY AND IT WAS FRACTURED POSTEROMEDIALLY AND HAD SOME PITTING THROUGHOUT THE MEDIAL AND LATERAL ASPECT AND SOME WEARING DOWN OF THE MEDIAL PLATEAU. DESPITE UNDERGOING REVISION SURGERY, PATIENT EXPERIENCES DAILY PAIN AND DISCOMFORT IN RIGHT KNEE WHICH LIMITS PATIENT¿S ACTIVITIES OF DAILY LIVING AND IMPACTS PATIENT QUALITY OF LIFE. | SECTION D10: CONCOMITANT MEDICAL PRODUCTS: LOGIC CR FEMORAL CEM, RIGHT, SZ 3 (CAT# 02-010-03-0330 / SERIAL# (B)(4)). LGC TIBIAL FIT TRAY CEM SZ 3F / 3T (CAT# 02-012-45-3030 / SERIAL# (B)(4)). THREE PEG PATELLA 29MM (CAT# 200-02-29 / SERIAL# (B)(4)). HOLDING PIN MINI SHARP POINT 4 PK (CAT# 201-78-15 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00470 | 2022-04-26 | 2022-04-06 | I, n, j, u, r, y | Manufacturer report | NV GXL LINR, NTRL, 36MM ID, GROUP 3 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LINR, NTRL, 36MM ID, GROUP 3 CUPS | 130-36-53 |  | AS REPORTED BY THE LEGAL TEAM, APPROXIMATELY FOUR YEARS POST INITIAL THA, THIS (B)(6) MALE PATIENT UNDERWENT A REVISION DUE TO PREMATURE POLYETHYLENE FAILURE. NO ADDITIONAL INFORMATION AVAILABLE. | CONCOMITANT MEDICAL PRODUCTS: COCR FEM HEAD 36MM +3.5 OFFSET 12/14 (CAT# 142-36-03 / SERIAL# (B)(4)) INTEGRIP CC, CLUSTER 56MM,G3 (CAT# 186-01-56 / SERIAL# (B)(4)) WEDGE PLASMA X/O SZ 10 (CAT# 188-01-10 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00471 | 2022-04-26 | 2022-03-14 | I, n, j, u, r, y | Manufacturer report | LOGIC CR TIB INSERT SLOPE+, SZ 5, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC CR TIB INSERT SLOPE+, SZ 5, 9MM | 02-012-48-5009 |  | AS REPORTED BY THE LEGAL TEAM, APPROXIMATELY THREE YEARS POST INITIAL TKA, THIS (B)(6)MALE PATIENT INDICATES THAT PATIENT WAS DIAGNOSED WITH MECHANICAL FAILURE. PATIENT UNDERWENT REVISION SURGERY DUE TO SIGNIFICANT POLYETHYLENE WEAR AND CYSTIC FORMATION WITHIN THE PROXIMAL TIBIA AND THE DISTAL FEMUR. FOLLOWING THE REVISION SURGERY, PATIENT CONTINUES TO BE LIMITED IN ACTIVITIES OF DAILY LIVING, PATIENT HAS DIFFICULTY WITH DAILY ACTIVITIES SUCH AS GETTING DRESSED, CLIMBING STAIRS, AND BATHING. FURTHERMORE, DESPITE UNDERGOING THE REVISION SURGERY, PATIENT EXPERIENCES DAILY PAIN AND DISCOMFORT IN LEFT KNEE WHICH LIMITS PATIENT¿S ACTIVITIES OF DAILY LIVING AND IMPACTS PATIENT QUALITY OF LIFE. | CONCOMITANT MEDICAL PRODUCTS: LOGIC CR FEMORAL CEM, LEFT, SZ 5 (CAT# 02-010-03-0250 / SERIAL# (B)(4)); LGC TIBIAL FIT TRAY CEM SZ 5F / 4T (CAT# 02-012-45-5040 / SERIAL# (B)(4)); THREE PEG PATELLA 35MM (CAT# 200-02-35 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00472 | 2022-04-26 | 2022-04-27 | I, n, j, u, r, y | Manufacturer report | LOGIC CR TIB INSERT SLOPE++, SZ 4, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC CR TIB INSERT SLOPE++, SZ 4, 9MM | 02-012-49-4009 |  | AS REPORTED BY THE LEGAL TEAM, APPROXIMATELY SIX YEARS POST INITIAL TKA, THIS (B)(6) MALE PATIENT UNDERWENT A REVISION FOR AN UNKNOW REASON. NO ADDITIONAL INFORMATION AVAILABLE. | CONCOMITANT MEDICAL PRODUCTS: LOGIC CR FEMORAL CEM, RIGHT, SZ 4 (CAT# 02-010-03-0340 / SERIAL# (B)(4)). LGC TIBIAL FIT TRAY CEM SZ 4F / 4T (CAT# 02-012-45-4040 / SERIAL# (B)(4)). THREE PEG PATELLA 32MM (CAT# 200-02-32 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00473 | 2022-04-26 | 2022-04-06 | I, n, j, u, r, y | Manufacturer report | LOGIC CR TIB INSERT SLOPE +, SZ 1.5, 9 MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC CR TIB INSERT SLOPE +, SZ 1.5, 9 MM | 02-012-48-1509 |  | AS REPORTED BY THE LEGAL TEAM, APPROXIMATELY SIX YEARS POST INITIAL TKA, THIS (B)(6) Y/O MALE PATIENT UNDERWENT A REVISION FOR AN UNKNOW REASON. NO ADDITIONAL INFORMATION AVAILABLE. | CONCOMITANT MEDICAL PRODUCTS: LOGIC CR FEMORAL CEM, LEFT SZ 1.5 (CAT# 02-010-03-0215 / SERIAL# (B)(4)). LGC TIBIAL FIT TRAY CEM SZ 1.5F / 1.5T (CAT3 02-012-45-1515 / SERIAL# (B)(4)). THREE PEG PATELLA 29MM (CAT# 200-02-29 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00474 | 2022-04-26 | 2022-03-07 | I, n, j, u, r, y | Manufacturer report | LOGIC TIB INSERT IMPL CRC, SZ 5, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIB INSERT IMPL CRC, SZ 5, 9MM | 02-012-51-5009 |  | AS REPORTED BY LEGAL TEAM, APPROXIMATELY FIVE YEARS POST INITIAL TKA, THIS MALE PATIENT HAD A REVISION FOR AN UNKNOWN REASON. NO ADDITIONAL INFORMATION AVAILABLE. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10098 | 2022-04-26 | 2018-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | FLEXIBLE DRILL SHAFT | EXACTECH, INC. |  | 181-00-01 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGICAL PROCEDURE A FLEXIBLE DRILL SHAFT BROKE. BOTH PIECES WERE NOTED TO BE ON THE BACK TABLE, BY THE AGENT, DURING AND AFTER THE PROCEDURE. NO PIECES WERE BELIEVED TO BE IN THE PATIENT AND/OR LEFT IN THE PATIENT. THERE WAS NO DELAY TO SURGERY AND NO ADVERSE EVENT TO THE PATIENT. THE PATIENT'S HEALTH WAS STABLE LEAVING THE OR. MULTIPLE EMAIL REQUESTS HAVE BEEN MADE FOR ADDITIONAL INFORMATION; NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DESIGN-RELATED ISSUES: THE DESIGN OF THE FLEXIBLE DRILL SHAFT HAS BEEN IN THE FIELD SINCE 2013. EXACTECH HAS RECEIVED 7 SIMILAR COMPLAINT REPORTS INVOLVING 9 DEVICES SINCE 2013. BECAUSE THE FLEXIBLE DRILL SHAFT IS USED IN TOTAL HIP SURGERIES, SALES DATA FOR ACETABULAR CUPS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MANUFACTURING: NO INFORMATION REGARDING THE SERIAL NUMBER OR MANUFACTURING LOT WAS PROVIDED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE MOST LIKELY UNDERLYING CAUSE OF THE BROKEN FLEXIBLE DRILL SHAFT COULD NOT BE DETERMINED BECAUSE THE DEVICE WAS NOT AVAILABLE FOR EVALUATION, AND NO FURTHER DETAILS ABOUT THE BREAKAGE HAVE BEEN PROVIDED. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10096 | 2022-04-26 | 2017-06-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | RESURFACING REAMER | EXACTECH, INC. |  | 313-25-50 |  | IT WAS REPORTED FROM THE US BY THE AGENT THAT DURING AN INSPECTION OF INSTRUMENTS IN THEIR WAREHOUSE 2 RESURFACING REAMERS WERE FOUND TO BE BROKEN. THERE IS NO INFORMATION REGARDING ANY DELAY OF A SURGERY, PATIENT INJURY, ADVERSE EVENT, OR CLINICAL CONSEQUENCE TO ANY PATIENT. THE AGENT WHO ORIGINALLY REPORTED THIS, IS NO LONGER WITH THE AGENCY. NO ADDITIONAL INFORMATION IS AVAILABLE TO FURTHER INVESTIGATE THIS ISSUE. THIS IS 1 MDR OF 2 FOR THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR PATIENT. DESIGN: THIS DESIGN HAS BEEN IN THE FIELD SINCE 2013. EXACTECH HAS NOT RECEIVED ANY OTHER COMPLAINT REPORTS INVOLVING BREAKAGE OF THIS INSTRUMENT. MANUFACTURING: EXACTECH HAS NOT RECEIVED ANY OTHER COMPLAINTS REPORTS INVOLVING PARTS WITH THIS CATALOG NUMBER OF FROM THIS MANUFACTURING LOT NUMBER. MOST LIKELY CAUSE: THE USER MOST LIKELY APPLIED OFF-AXIS FORCE WHILE REAMING THE WHICH LED TO FRACTURING OF THE INSTRUMENTS. IFU 700-096-181: INSTRUMENT INSPECTION VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU ALSO STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED INSTRUMENT MALFUNCTION IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO ANY REPORTED ADVERSE PATIENT OUTCOME. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10099 | 2022-04-26 | 2017-12-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRULIANT FEMORAL IMPACTOR HEAD | EXACTECH, INC. |  | 02-029-19-1000 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC KNEE REVISION SURGERY OF COMPETITOR DEVICES THE SURGEON EXPERIENCED AN ISSUE WITH AN INSTRUMENT THAT BROKE IN SURGERY. THE LITTLE SCREW FELL OUT OF THE FEMORAL IMPACTOR HEAD AND THE SCRUB TECH KEPT IT ASIDE. AFTER THE INSTRUMENT WAS WASHED, WHILE TRYING TO DETERMINE HOW THE SCREW CAME OUT, THE ENTIRE INSTRUMENT FELL APART IN THE REPRESENTATIVE'S HAND. THERE WAS NO DELAY. THE DEVICE WAS RETURNED. THE REP WAS PRESENT AT THE TIME OF SURGERY. INACTIVE AGENCY WITH NO HOSPITAL/OTHER CONTACT INFORMATION FOR THE EVENT. NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. ENGINEERING EVAL COMPLETED ON 30 MAY 2018. "BASED ON EXAMINATION OF THE RETURNED DEVICE, IT WAS CONCLUDED THAT THE FAILURE MODE IS THE SAME INVESTIGATED IN CONJUNCTION WITH CAPA(B)(4) [WHICH IS CURRENTLY OPEN]. THIS COMPLAINT IS ALREADY INCLUDED IN THAT CAPA AND THE COMPLAINT WILL BE CLOSED WITHOUT FURTHER INVESTIGATION." (B)(4) AND CAPA(B)(4). (B)(4) EVALUATION SUMMARY: THE COMMITTEE HAS DETERMINED THAT THE OPERATIVE TECHNIQUE WILL BE UPDATED, AND A FIELD NOTICE WILL BE SENT TO THE AGENTS AND REPRESENTATIVES NOTIFYING THEM OF THE UPDATE. A STAGED MARKET WITHDRAWAL WILL BE PERFORMED TO REMOVE THE AFFECTED DEVICES FROM THE FIELD. HEALTH RISK- LOW: UP TO (B)(6) 2018, THE TRULIANT INSTRUMENTS HAVE BEEN USED IN OVER (B)(6)SURGERIES,(B)(4). IN THESE COMPLAINTS, THE BROKEN SCREW WAS COLLECTED AND NOT LEFT IN PATIENT'S WOUND. THE SCREW IS MADE OF STAINLESS STEEL AND IS RADIOPAQUE AND IS LARGE ENOUGH TO BE DETECTED. CAPA(B)(4) INVESTIGATION: REVIEW OF THE COMPLAINT SHOWED CAP SCREW WAS LOOSENED DURING SURGERY. INVESTIGATION REVEALED THAT THE FAILURE OF THE CAP SCREW WAS DIRECTLY RELATED TO INADEQUATE WELDING BETWEEN THE CAP SCREW AND THE WASHER SUBCOMPONENT APPLIED DURING MANUFACTURING ASSEMBLY. THERE WAS ONLY A MINIMAL SPOT WELD APPLIED WHICH WAS VULNERABLE DURING IMPACTION USE. A REVIEW OF PRINT SHOWED THE WELD REQUIREMENT WAS CALLED OUT AS "TACK WELD", WHICH WAS NOT SPECIFIC ENOUGH AND LEFT ROOM FOR INTERPRETATION. ROOT CAUSE: WELDING REQUIREMENT WAS NOT ADEQUATELY SPECIFIED. "BASED ON EXAMINATION OF THE RETURNED DEVICE, IT WAS CONCLUDED THAT THE FAILURE MODE IS THE SAME INVESTIGATED IN CONJUNCTION WITH CAPA(B)(4) [WHICH IS CURRENTLY OPEN]. THIS COMPLAINT IS ALREADY INCLUDED IN THAT CAPA AND THE COMPLAINT WILL BE CLOSED WITHOUT FURTHER INVESTIGATION." (B)(4) AND CAPA(B)(4). (B)(4) EVALUATION SUMMARY: THE COMMITTEE HAS DETERMINED THAT THE OPERATIVE TECHNIQUE WILL BE UPDATED, AND A FIELD NOTICE WILL BE SENT TO THE AGENTS AND REPRESENTATIVES NOTIFYING THEM OF THE UPDATE. A STAGED MARKET WITHDRAWAL WILL BE PERFORMED TO REMOVE THE AFFECTED DEVICES FROM THE FIELD. HEALTH RISK- LOW: UP TO (B)(6) 2018, THE TRULIANT INSTRUMENTS HAVE BEEN USED IN OVER (B)(6) SURGERIES, (B)(4). IN THESE COMPLAINTS, THE BROKEN SCREW WAS COLLECTED AND NOT LEFT IN PATIENT'S WOUND. THE SCREW IS MADE OF STAINLESS STEEL AND IS RADIOPAQUE AND IS LARGE ENOUGH TO BE DETECTED. CAPA(B)(4) INVESTIGATION: REVIEW OF THE COMPLAINT SHOWED CAP SCREW WAS LOOSENED DURING SURGERY. INVESTIGATION REVEALED THAT THE FAILURE OF THE CAP SCREW WAS DIRECTLY RELATED TO INADEQUATE WELDING BETWEEN THE CAP SCREW AND THE WASHER SUBCOMPONENT APPLIED DURING MANUFACTURING ASSEMBLY. THERE WAS ONLY A MINIMAL SPOT WELD APPLIED WHICH WAS VULNERABLE DURING IMPACTION USE. A REVIEW OF PRINT SHOWED THE WELD REQUIREMENT WAS CALLED OUT AS "TACK WELD", WHICH WAS NOT SPECIFIC ENOUGH AND LEFT ROOM FOR INTERPRETATION. ROOT CAUSE: WELDING REQUIREMENT WAS NOT ADEQUATELY SPECIFIED. EXACTECH IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT LED TO ANY PATIENT IMPACT. THERE IS NO INDICATION THAT THE SCREW FELL INTO THE SURGICAL SITE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10081 | 2022-04-26 | 2019-08-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | HEAD TRIAL, 12/14 36 +0 | EXACTECH, INC. |  | 143-36-00 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE TRIAL HEAD CRACKED IN HALF WHEN REMOVING FROM THE HIP. THE DEVICE WAS NOT RETURNED AS IT WAS DISPOSED OF AT THE HOSPITAL. THEREFORE; NOT AVAILABLE FOR EVALUATION. NO PIECES FELL INTO THE SURGICAL SITE AND THERE WAS NO PATIENT INJURY, ADVERSE EVENT, OR CLINICAL CONSEQUENCE TO THE PATIENT. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE DESTROYED BY FACILITY. ANALYSIS COMPLETED ON 02/03/2020 BY (B)(4). MOST LIKELY UNDERLYING CAUSE / ADDITIONAL COMMENTS: ACCORDING TO THE INFORMATION PROVIDED, "THE TRIAL HEAD (36MM, +0 TRIAL HEAD) CRACKED IN HALF WHEN REMOVING THE HIP." THE DEVICE WAS NOT RETURNED AS IT WAS DISPOSED OF AT THE HOSPITAL; THEREFORE, NOT AVAILABLE FOR EVALUATION. DESIGN-RELATED ISSUES: THE DESIGN OF THE 12/14 FEMORAL HEAD TRIAL, O-RING HAS BEEN IN THE FIELD SINCE 2007. (B)(4). THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: MANUFACTURING DATA COULD NOT BE REVIEWED BECAUSE THE SERIAL/LOT INFORMATION WAS NOT PROVIDED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) AND RISK ASSESSMENT AND CONTROLS REPORT (RACR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR AND RACR FOR THIS FEMORAL HEAD TRIAL, RMR-00027 REV - AND RACR-00030 REV -, WERE REVIEWED. THE RISK IS CAPTURED, HOWEVER THE RACR IS BEING UPDATED UNDER (B)(4) TO CAPTURE AN APPROPRIATE HAZARDOUS SITUATION FOR THE REPORTED EVENT. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", THE RISK IS CAPTURED IN THE RMR AND RACR, AND THE HAZARDOUS SITUATION WILL BE UPDATED. THE BROKEN FEMORAL HEAD TRIAL REPORTED IN (B)(4) WAS LIKELY THE RESULT OF AN EXPOSURE TO IMPACTION FORCES AND SUBSEQUENT STERILIZATION CYCLES WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE WHEN REMOVING THE HEAD TRIAL. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE WAS NO PATIENT ADVERSE EVENT, NONE OF THE DEVICE FRAGMENTS FELL INTO THE SURGICAL SITE. AN INVESTIGATION WAS CONDUCTED; THE BROKEN FEMORAL HEAD TRIAL REPORTED WAS LIKELY THE RESULT OF AN EXPOSURE TO IMPACTION FORCES AND SUBSEQUENT STERILIZATION CYCLES WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE WHEN REMOVING THE HEAD TRIAL. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10097 | 2022-04-26 | 2017-06-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | RESURFACING REAMER | EXACTECH, INC. |  | 313-25-50 |  | IT WAS REPORTED FROM THE US BY THE AGENT THAT DURING AN INSPECTION OF INSTRUMENTS IN THEIR WAREHOUSE 2 RESURFACING REAMERS WERE FOUND TO BE BROKEN. THERE IS NO INFORMATION REGARDING ANY DELAY OF A SURGERY, PATIENT INJURY, ADVERSE EVENT, OR CLINICAL CONSEQUENCE TO ANY PATIENT. THE AGENT WHO ORIGINALLY REPORTED THIS, IS NO LONGER WITH THE AGENCY. NO ADDITIONAL INFORMATION IS AVAILABLE TO FURTHER INVESTIGATE THIS ISSUE. THIS IS 2 MDR OF 2 FOR THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: AGE, SEX, WEIGHT, ETHNICITY, RACE. NO OTHER INFORMATION HAS BEEN PROVIDED ABOUT THE EVENT OR PATIENT. DESIGN: THIS DESIGN HAS BEEN IN THE FIELD SINCE 2013. EXACTECH HAS NOT RECEIVED ANY OTHER COMPLAINT REPORTS INVOLVING BREAKAGE OF THIS INSTRUMENT. MANUFACTURING: EXACTECH HAS NOT RECEIVED ANY OTHER COMPLAINTS REPORTS INVOLVING PARTS WITH THIS CATALOG NUMBER OF FROM THIS MANUFACTURING LOT NUMBER. MOST LIKELY CAUSE: THE USER MOST LIKELY APPLIED OFF-AXIS FORCE WHILE REAMING THE WHICH LED TO FRACTURING OF THE INSTRUMENTS. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU ALSO STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED INSTRUMENT MALFUNCTION IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO ANY REPORTED ADVERSE PATIENT OUTCOME. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00468 | 2022-04-26 | 2021-10-20 | I, n, j, u, r, y | Manufacturer report | NV GXL LINER NEUTRAL, 36MM ID, GROUP 2 CUPS | SEE H10 | EXACTECH, INC. | NV GXL LINER NEUTRAL, 36MM ID, GROUP 2 CUPS | 130-36-52 |  | AS REPORTED BY THE LEGAL TEAM, APPROXIMATELY THREE YEARS POST INITIAL THA, THIS (B)(6) Y/O MALE PATIENT UNDERWENT A REVISION DUE TO PREMATURE POLYETHYLENE FAILURE. NO ADDITIONAL INFORMATION AVAILABLE. | CONCOMITANT MEDICAL PRODUCTS : COCR FEM HEAD 36MM -3.5 OFFSET 12/14 (CAT# 142-36-93 / SERIAL# (B)(4)). INTEGRIP CC, CLUSTER 52MM, G2 (CAT# 186-01-52 / SERIAL# (B)(4)). WEDGE PLASMA X/O SZ 5 (CAT# 188-01-05 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10093 | 2022-04-26 | 2017-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | SHORT STEM PROTECTOR | EXACTECH, INC. |  | 301-07-20 |  | IT WAS REPORTED FROM OUS THAT, "THE INSTRUMENT COULD HAVE BEEN BROKEN DURING THE SURGERY. HOWEVER, THE BREAKAGE HAS BEEN NOTICED AT THE RECOVERY OF THE INSTRUMENTATION-SET BY OUR LOGISTICS DEPT., AFTER THE CLEANING REALIZED BY THE STERILIZATION TEAM OF THE HOSPITAL." THE SALES REPRESENTATIVE WAS NOT PRESENT AT THE SURGERY. THERE WAS NO REPORTED DELAY OF SURGERY AND THERE WAS NO REPORTED PATIENT INJURY, ADVERSE EVENT, OR CLINICAL CONSEQUENCE TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE REPORTER RELATED TO THIS EVENT OR PATIENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. PATIENT INFORMATION WAS NOT PROVIDED. DESIGN-RELATED ISSUES: THE DESIGN OF THE SHORT STEM PROTECTOR HAS BEEN IN THE FIELD SINCE 2004. EXACTECH HAS RECEIVED (B)(4) SIMILAR COMPLAINT REPORTS SINCE 2008. BECAUSE THE SHORT STEM PROTECTOR IS USED IN ANATOMIC AND REVERSE TOTAL SHOULDER SURGERIES, SALES DATA FOR HUMERAL STEMS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4). THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE DETAILED BELOW. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. EXACTECH HAS NOT RECEIVED ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT, WHICH HAS BEEN IN THE FIELD SINCE 2014. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING-RELATED. RISK MANAGEMENT: A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS SHORT STEM PROTECTOR, 750-2002-039-RMR REV F, WAS REVIEWED. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR MOST LIKELY CAUSE: THE BROKEN SHORT STEM PROTECTOR REPORTED WAS LIKELY THE RESULT OF DISASSEMBLING THE DEVICE FOR CLEANING AND NOT REASSEMBLING IT. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED INSTRUMENT MALFUNCTION IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO ANY PATIENT IMPACT. AN INVESTIGATION WAS CONDUCTED THAT THE MOST LIKELY ROOT CAUSE OF THE DEVICE BREAKAGE REPORTED IN THIS EVENT WAS LIKELY THE RESULT OF DISASSEMBLING THE DEVICE FOR CLEANING AND NOT REASSEMBLING IT. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10077 | 2022-04-26 | 2018-10-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | HUMERAL HEAD IMPACTOR TIP | EXACTECH, INC. |  | 311-07-07 |  | IT WAS REPORTED FROM THE US THAT DURING AND ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED DIFFICULTY WITH THE INSTRUMENTS. THE TRI-DRIVE BROKE WHILE DRILLING THE CENTER GLENOID POST (LARGE MALE WITH TIGHT EXPOSURE AND VERY HARD BONE), AND THE SURGEON NOTICED THE SHAFT RUBBING ON THE FUKUDA RETRACTOR. IN THE SAME CASE, THE IMPACTOR TIP HAD A SMALL PORTION OF THE BLACK PLASTIC BREAK AND FALL TO THE FLOOR WHEN THE SURGEON USED THE INSTRUMENT TO IMPACT THE LARGEST HUMERAL HEAD COMPONENT (53MM). THERE WAS NO COMPROMISE TO THE PATIENT OR SURGERY, AS THE SURGEON WAS DONE WITH THE INSTRUMENTS. NO DELAY TO SURGERY. THE PATIENT'S HEALTH WAS STABLE LEAVING THE OR. NO PIECES WERE LEFT IN THE PATIENT, AS PROVEN BY THE STANDARD POST OP X-RAY. AGENT WAS PRESENT AT TIME OF SURGERY. THE AGENCY IS INACTIVE, NO FURTHER INFORMATION IS AVAILABLE. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICES WERE RECEIVED FOR ANALYSIS. ENG EVAL COMPLETED ON 2019.08.06 BY (B)(4). TRI-DRIVE MODULAR HANDLE. (B)(4) AND CAPA(B)(4). THE DESIGN OF THE TRI-DRIVE MODULAR HANDLE HAS BEEN IN THE FIELD SINCE 2009. EXACTECH HAS RECEIVED (B)(4) SIMILAR COMPLAINT REPORTS INVOLVING (B)(4) TRI-DRIVE MODULAR HANDLE DUE TO SLEEVE DISASSEMBLY SINCE 2009. BECAUSE THE TRI-DRIVE IS USED IN ANATOMIC AND REVERSE TOTAL SHOULDER SURGERIES, (B)(4). THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE IN THE RMR. THE DISASSEMBLED TRI-DRIVE MODULAR REAMER DRIVER MAY RESULT IN A DELAY IN SURGERY TO FIND A NEW ONE. TYPICALLY, IN ANATOMIC SHOULDER CASES BOTH THE ANATOMIC AND THE REVERSE SETS ARE IN THE OR. IF THE TRI-DRIVE WERE TO BREAK, A SECOND TRI-DRIVE WOULD BE IN THE REVERSE SHOULDER SET. BASED ON THE INVESTIGATIONS CONDUCTED IT APPEARS THAT THIS ISSUE IS RELATED TO A USER ISSUE WHEREIN THE USER HOLDS ONTO THE COLOR OF THE TRI-DRIVE HANDLE WHEN THE DEVICE IS UNDER POWERED ROTATION. THIS CAUSES FORCE TO BE EXERTED ON THE INTERNAL PARTS WHICH RESULTS IN DISASSEMBLY. MOST LIKELY ROOT CAUSE-THE BROKEN TRI-DRIVE MODULAR HANDLE REPORTED IN EXPERIENCE (B)(4) WAS LIKELY THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY, ALLOWING THE TIP TO DISASSEMBLE. HUMERAL HEAD IMPACTOR VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR 10X OPTICAL UNLESS OTHERWISE SPECIFIED)- THE BROKEN HUMERAL HEAD IMPACTOR TIP APPEARS CONSISTENT WITH CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE RADEL IMPACTION SURFACES. THIS BREAKAGE WAS LIKELY THE RESULT OF YEARS OF NORMAL USE AND SUBSEQUENT STERILIZATIONS. HEAD IMPACTOR TIP DUE TO A FRACTURED RADEL IMPACTION SURFACE SINCE 2008. THE FREQUENCY OF OCCURRENCE RANKING SCALE IS "VERY LOW", THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN- RELATED. (B)(4), WHICH HAS BEEN IN THE FIELD SINCE 2017. THE DEVICE HISTORY RECORD FOR THE HUMERAL HEAD IMPACTOR TIP WAS REVIEWED, AND ALL PARTS ACCEPTED WITH CONFORMANCE TO THE DEVICE SPECIFICATIONS. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMRS. MOST LIKELY ROOT CAUSE-THE BROKEN HUMERAL HEAD IMPACTOR REPORTED IN EXPERIENCE (B)(4) WAS LIKELY THE RESULT OF CRACK INITIATION, PROPAGATION, AND ULTIMATE FAILURE OF THE RADEL IMPACTION SURFACE RELATED TO NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE WAS NO PATIENT ADVERSE EVENT, NO BROKEN PIECES OF THE DEVICE WERE LEFT IN THE PATIENT, THIS WAS CONFIRMED WITH THE ROUTINE POST-OP X-RAY. AN INVESTIGATION WAS CONDUCTED FOR THE TRI-DRIVE MODULAR HANDLE UNDER (B)(4) AND CAPA(B)(4) THE BROKEN TRI-DRIVE MODULAR HANDLE REPORTED WAS LIKELY THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY, ALLOWING THE TIP TO DISASSEMBLE. AN INVESTIGATION WAS CONDUCTED FOR THE HUMERAL HEAD IMPACTOR TIP MANUFACTURING; THE BROKEN HUMERAL HEAD IMPACTOR REPORTED WAS LIKELY THE RESULT OF CRACK INITIATION, PROPAGATION, AND ULTIMATE FAILURE OF THE RADEL IMPACTION SURFACE RELATED TO NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10083 | 2022-04-26 | 2019-11-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TIBIAL PLATE FB SZ 3 LT | EXACTECH, INC. |  | 350-11-03 |  | IT WAS REPORTED FROM THE VANTAGE TOTAL ANKLE STUDY, THAT A PATIENT EXPERIENCED POSSIBLE LOOSENING NOTED AT 2-YEAR FOLLOW UP X-RAY. ACTION TAKEN: MEDICATION. THE CASE REPORT FORM INDICATES THIS EVENT IS POSSIBLY RELATED TO THE DEVICES AND RELATED TO THE PROCEDURE. THE PATIENT IS ALSO SUFFERING PAIN ALONG POSTERIOR TIBIAL TENDON. NO ADDITIONAL INFORMATION IS AVAILABLE. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE REMAINS IMPLANTED. ENG EVAL COMPLETED ON 2020.03.24 BY (B)(4). DESIGN-RELATED ISSUES: THE DESIGN OF THE VANTAGE ANKLE HAS BEEN IN THE FIELD SINCE 2015. SINCE 2015, EXACTECH HAS RECEIVED (B)(4) COMPLAINT REPORTS ANKLE COMPONENTS DUE TO LOOSENING. SALES DATA FOR ALL TIBIAL ANKLE COMPONENTS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE DETAILED BELOW. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. (B)(4). EXACTECH IS NOT AWARE OF ANY COMPLAINTS INVOLVING PARTS FROM THESE MANUFACTURING LOTS, WHICH HAVE BEEN IN THE FIELD SINCE 2017. THE MANUFACTURING RECORDS FOR THIS DEVICE ONLY EXIST IN PAPER FORM AND ARE STORED IN CARDBOARD BOXES AT (B)(4). DUE TO RISKS ASSOCIATED WITH COVID-19 TRANSMISSION, THE PHYSICAL MANUFACTURING RECORDS FOR THIS COMPLAINT WILL NOT BE REVIEWED. EXACTECH WILL CONTINUE TO MONITOR THIS COMPLAINT TYPE AND, IF WARRANTED, WILL ESCALATE ANY POSITIVE TRENDS TO THE CORRECTIONS AND REMOVAL COMMITTEE (CRC) FOR ADDITIONAL ACTIONS. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE LOOSENING IS LISTED IN THE PRODUCT LABELING, THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. THE LOOSENING REPORTED IN EXPERIENCE (B)(4) WAS LIKELY THE RESULT OF AN INSUFFICIENT BOND BETWEEN THE IMPLANT AND THE BONE, WHICH LED TO POSSIBLE ASEPTIC (NON-INFECTED) LOOSENING. HOWEVER, THIS CANNOT BE CONFIRMED AS THE PATIENT HAS NOT BEEN REVISED. PROPERLY FUNCTIONING IMPLANTS DEPEND ON THEIR APPROPRIATE FIXATION TO THE BONE; FIXATION IS USUALLY ACHIEVED BY CEMENTING THE IMPLANT ONTO THE BONE. SOME SURGEONS PREFER TO USE BIOLOGIC (NONCEMENTED) FIXATION. ALTHOUGH IMPLANTS ARE FIRMLY FIXED AT THE INITIAL ANKLE REPLACEMENT SURGERY, THEY MAY BECOME LOOSE OVER TIME. THE CAUSE CAN BE RESIDUAL "BONEY" MALALIGNMENT OF THE FOOT OR POOR FUNCTIONAL BALANCE AND ROM OF THE JOINT SECONDARY TO MUSCULO-TENDON CONTRACTURES OR IMBALANCES BETWEEN ANTAGONISTIC MUSCLE GROUPS. EITHER OF THESE CAN INCREASE CONTACT PRESSURES AT THE ARTICULAR SURFACES OF THE COMPONENTS. ADDITIONALLY, FRICTION CAUSED BY THE JOINT SURFACES RUBBING AGAINST EACH OTHER WEARS AWAY THE SURFACES OF THE IMPLANT, CREATING TINY PARTICLES THAT ACCUMULATE AROUND THE JOINT. IN A PROCESS CALLED ASEPTIC (NONINFECTED) LOOSENING, THE BOND OF THE IMPLANT TO THE BONE IS DESTROYED BY THE BODY'S ATTEMPT TO DIGEST THE WEAR PARTICLES. DURING THIS PROCESS, NORMAL BONE IS ALSO DIGESTED (A CONDITION CALLED OSTEOLYSIS), WHICH CAN WEAKEN OR EVEN FRACTURE THE BONE. WHEN THE PROSTHESIS BECOMES LOOSE, THE PATIENT MAY EXPERIENCE PAIN, CHANGE IN ALIGNMENT, SUBSIDENCE, OR INSTABILITY. COMPONENT LOOSENING IS THE MOST COMMON MODE OF FAILURE OF ANKLE IMPLANTS(1,2). 1. ESCUDERO MI, LE V, BARAHONA M, ET AL. 2019. TOTAL ANKLE ARTHROPLASTY SURVIVAL AND RISK FACTORS FOR FAILURE. FOOT ANKLE INT. 40(9):997-1006. 2. PENNER M, DAVIS WH, WING K, ET AL. 2018. THE INFINITY TOTAL ANKLE SYSTEM: EARLY CLINICAL RESULTS WITH 2- TO 4- YEAR FOLLOW-UP. FOOT ANKLE SPEC :1938640018777601. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THE STUDY PATIENT IS REPORTED TO BE SUFFERING PAIN ALONG THE POSTERIOR TIBIAL TENDON AND POSSIBLE LOOSENING NOTED AT 2-YEAR FOLLOW UP X-RAY. LOOSENING IS A KNOWN CAUSE FOR ORTHOPEDIC REVISION SURGERY. AN INVESTIGATION WAS CONDUCTED; THE LOOSENING REPORTED WAS LIKELY THE RESULT OF AN INSUFFICIENT BOND BETWEEN THE IMPLANT AND THE BONE, WHICH LED TO POSSIBLE ASEPTIC (NON-INFECTED) LOOSENING. HOWEVER, THIS CANNOT BE CONFIRMED AS THE PATIENT HAS NOT BEEN REVISED. INFORMATION WAS NOT PROVIDED FOR: PT. INFO.. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10092 | 2022-04-26 | 2018-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | 2.0 NITINOL K-WIRE | EXACTECH, INC. |  | 315-35-00 |  | IT WAS REPORTED FROM THE US THAT PRIOR TO GOING INTO SURGERY, A PUNCTURE HOLE WAS NOTICED IN THE PACKAGING. THE AGENT DID NOT TAKE INTO THE OPERATING ROOM. THE DEVICE WAS RECEIVED FOR ANALYSIS ON 12/27/2018. THE K-WIRE PUNCTURED THROUGH THE CLEAR SIDE OF THE PEEL POUCH DURING HANDLING OR TRANSPORTATION, COMPROMISING THE STERILITY OF THE INSTRUMENT. THERE IS NO PATIENT INVOLVEMENT. THE STERILITY OF THE DEVICE IS COMPROMISED, THIS PRESENTS THE POSSIBILITY OF PATIENT INFECTION. INACTIVE AGENCY WITH NO HOSPITAL/OTHER CONTACT INFORMATION FOR THE EVENT. NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS AVAILABLE FOR ANALYSIS. ENGINEERING ANALYSIS COMPLETED BY AD ON 4/10/2019. DESIGN-RELATED ISSUES: EXACTECH HAS RECEIVED 9 COMPLAINT REPORTS INVOLVING HOLES IN STERILE INSTRUMENT/IMPLANT PEEL POUCHES (NO OTHER COMPLAINTS INVOLVING 315-35-00) SINCE 2008. THE SALES DATA FOR ALL GLENOID K-WIRES (315-35-00) WAS USED TO CALCULATE A COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF RECEIVING ANY OTHER COMPLAINT REPORTS INVOLVING A PART FROM THIS PACKAGING LOT OF 35 PIECES. THE PACKAGING IS VISUALLY INSPECTED PER WI 701-105-090. THEREFORE, THIS DOES NOT APPEAR TO BE RELATED TO THE MANUFACTURING OF THE PACKAGING MATERIALS. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS STERILE INSTRUMENT, 750-2007-026- RMR-INSTRUMENT REV K, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 12. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE IS "VERY LOW" AND THE RISK IS CAPTURED IN THE RMR. THE OUTER PACKAGING DAMAGE REPORTED IN EXPERIENCE C2018-532 WAS LIKELY THE RESULT OF THE END OF THE K-WIRE PUNCTURING THROUGH THE CLEAR SIDE OF THE PEEL POUCH DURING HANDLING OR TRANSPORTATION, COMPROMISING THE STERILITY OF THE INSTRUMENT. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO PATIENT INVOLVEMENT. THE STERILITY OF THE DEVICE IS COMPROMISED, THIS PRESENTS THE POSSIBILITY OF PATIENT INFECTION. AN INVESTIGATION WAS CONDUCTED: THE OUTER PACKAGING DAMAGE REPORTED WAS LIKELY THE RESULT OF THE END OF THE K-WIRE PUNCTURING THROUGH THE CLEAR SIDE OF THE PEEL POUCH DURING HANDLING OR TRANSPORTATION, COMPROMISING THE STERILITY OF THE INSTRUMENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00459 | 2022-04-26 | 2022-11-10 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 4 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 4 13MM | 02-012-35-4013 |  | AS REPORTED BY LEGAL TEAM, THIS (B)(6) MALE PATIENT HAS BEEN EXPERIENCING PAIN AND TIGHTNESS IN BOTH KNEES RIGHT KNEE TO BE REVISED IN THE FUTURE. LEFT KNEE WAS REVISED FOR POTENTIAL WEAR. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00460 | 2022-04-26 | 2022-03-31 | I, n, j, u, r, y | Manufacturer report | EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0 | PROSTHESIS, SHOULDER, NON-CONSTRAINED, METAL/POLYMER CEMENTED | EXACTECH, INC. | 320-38-00 | 320-10-00 |  | AS REPORTED, APPROXIMATELY 6 MONTHS POST OP THE INITIAL LEFT TSA, THIS (B)(6) MALE PATIENT WAS REVISED DUE TO DISLOCATION. PON OPENING THE JOINT AND ASSESSING THE IMPLANTS THE SURGEON FOUND THE HUMERAL PLATE WAS ROTATING. THE HUMERAL LINER (38 MM +0) WAS REMOVED FOLLOWED BY THE TORQUE SCREW AND THE HUMERAL ADAPTER TRAY (+0). THE SURGEON REPLACED THE FOREMENTIONED IMPLANTS WITH THE SAME SIZE COMPONENTS AND DETERMINED THE PATIENT TO BE STABLE. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. | DEVICE EVALUATED BY MFR: PENDING EVALUATION CONCOMITANT DEVICE(S): 320-20-00, 6954073 - EQ REVERSE TORQUE DEFINING SCREW KIT. 320-10-00, 6954073 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00461 | 2022-04-26 | 2022-04-04 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | CAGE GLENOID SMALL, ALPHA | EXACTECH, INC. | 314-13-02 | 314-13-02 |  | AS REPORTED, THIS (B)(6) FEMALE PATIENT THAT IS A FACTORY WORKER, RECEIVED A LEFT ANATOMIC TSA ON (B)(6) 2020. THIS IS HER NON-DOMINANT ARM. APPROXIMATELY 1.5 YRS LATER, SHE RETURNED TO THE SURGEON COMPLAINING OF PAIN & WAS OBSERVED TO HAVE A DISLODGED GLENOID COMPONENT ON X-RAY/CT AND WAS REVISED. NO TRAUMATIC EVENT WAS REPORTED. IT IS NOTED DURING THE REVISION PROCEDURE THAT THE GLENOID COMPONENT, INCLUDING ALL THREE PERIPHERAL PEGS, WERE FOUND TO BE DISLODGED FROM THE GLENOID. THE CENTRAL CAGE REMAINED WITHIN THE GLENOID BONE. THE HUMERAL HEAD WAS REMOVED, AS WELL AS THE TORQUE SCREW & REPLICATOR PLATE. THE CENTRAL CAGE WAS THEN REMOVED. THE HUMERAL STEM WAS ALSO FOUND TO BE LOOSE & WAS REMOVED WITH RELATIVE EASE. THE SEVERITY OF HUMERAL BONE LOSS WAS FOUND TO BE TOO SIGNIFICANT TO SUPPORT A NEW STEM. APPROXIMATELY 2CM OF STEM COULD BE SEEN DUE TO THE BONE LOSS. CANCELLOUS CHIPS & OPTECURE +CCC WERE USED TO FILL THE DEFECTS & A 7MM INTERSPACE ANTIBIOTIC SPACER WAS USED IN THE HUMERUS. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE PATIENT WILL LIKELY BE REFERRED TO A SURGEON FOR EVALUATION & POSSIBLE HUMERAL RECONSTRUCTION PROSTHESIS. | DEVICE EVALUATED BY MFR: PENDING EVALUATION. CONCOMITANT DEVICE(S): 300-01-11, 6476735 - EQUINOXE, HUMERAL STEM PRIMARY, PRESS FIT 11MM. 300-10-45, 6553666 - EQUINOXE REPLICATOR PLATE 4.5MM O/S. 300-20-02, 6536115 - EQUINOX SQUARE TORQUE DEFINE SCREW DRIVE KIT. 310-01-38, 6418258 - EQUINOXE, HUMERAL HEAD SHORT, 38MM (ALPHA). | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00462 | 2022-04-26 | 2022-04-05 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | SMALL SUPERIOR AUGMENT GLENOID PLATE | EXACTECH, INC. | UNK | 320-35-02 |  | DURING THE INITIAL RIGHT TSA FOR THIS (B)(6) FEMALE PATIENT, THERE WAS A CRACK IN THE CORACOID NOTED DUE TO LOOSENING OF THE CORACOID GPS TRACKER. AT THIS POINT, THE LAST TWO SCREWS WERE PLACED WITHOUT COMPUTER NAVIGATION. THERE WAS NO ACTION TAKEN AND THE OUTCOME IS NOTED AS: CONTINUE TO UPDATE THIS ADVERSE EVENT UNTIL IT HAS BEEN RESOLVED. CURRENTLY, THERE IS NO UPDATED INFORMATION IN THE CLINICAL STUDY DATABASE. THE CASE REPORT FORM INDICATES THIS EVENT IS UNLIKELY RELATED TO DEVICES AND DEFINITELY RELATED TO PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | PENDING EVALUATION: CONCOMITANT DEVICE(S): 320-36-00 - 36MM HUMERAL LINER +0 UNCONSTRAINED, 320-10-00 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0, 320-31-36 - GLENOSPHERE, 36MM, 300-30-06 - EQUINOXE PRESERVE STEM 6MM. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00463 | 2022-04-26 | 2022-02-10 | I, n, j, u, r, y | Manufacturer report | VANTAGE | TALAR IMPLANT SZ 4 LT | EXACTECH, INC. | 350-01-04 | 350-01-04 |  | AS REPORTED, APPROXIMATELY 9 MONTHS POST OP THE INITIAL LEFT ANKLE REPLACEMENT, THIS (B)(6) MALE PATIENT'S WAS EXPERIENCING PAIN IN ANKLE, CONSTANT, LOCALIZED TO MEDIAL AND ANTERIOR ANKLE. HE ALSO HAD INCREASED SWELLING AND REDNESS BUT DENIES ANY NEW INJURIES, FALLS, INFECTION. ORDERED LABS TO RULE OUT INFECTION. SYMPTOMS SLIGHTLY IMPROVED. FOLLOW-UP WILL BE BASED ON LABS. THE OUTCOME IS CONTINUING. THE CASE REPORT FORM INDICATES THIS EVENT IS DEFINITELY NOT RELATED TO DEVICES AND UNLIKELY RELATED TO PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | PENDING EVALUATION: CONCOMITANT DEVICE(S): 350-21-24, TIBIAL INSERT FB SZ 4 LT 8MM, 350-10-04, ANKLE SZ 4 LOCKING CLIP. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00464 | 2022-04-26 | 2022-04-06 | I, n, j, u, r, y | Manufacturer report | TRULIANT PS CEM FEM PS CEM RIGHT SZ 4 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT PS CEM FEM PS CEM RIGHT SZ 4 | 02-020-11-0340 |  | AS REPORTED, APPROXIMATELY THREE YEARS POST INITIAL RIGHT TKA, THIS (B)(6) MALE PATIENT HAD A REVISION DUE TO LOOSE FEMUR, PATIENT WAS COMPLAINING OF PAIN. THIS IMPLANT COULD BE PART OF RECALL. THE SURGEON USED ANOTHER COMPANIES IMPLANTS FOR THE REVISION. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. SALES REP WAS UNABLE TO OBTAIN PHOTOS/X-RAYS, THE DEVICE WILL BE RETURN. | CONCOMITANT MEDICAL PRODUCTS: TRULIANT TIB IMP PS INSERT SZ 4 9MM (CAT# 02-022-35-4009 / SERIAL# (B)(4)); TRULIANT TIB FIT TRAY CEM SZ 4F / 4T (CAT# 02-022-45-4040 / SERIAL# (B)(4)); ADVANCED PATELLA 32MM 3 PEG IMPLANT (CAT# 200-07-32 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10078 | 2022-04-26 |  | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRULIANT TIBIAL TAMP GUIDE | EXACTECH, INC. |  | 02-029-22-2000 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE TRULIANT TIBIAL TAMP GUIDE BROKE. THERE WERE NO PIECES THAT FELL INTO THE SURGICAL SITE. THERE WAS A 5-MINUTE DELAY TO SURGERY WITH NO ADVERSE EVENT TO THE PATIENT. THE AGENT WAS PRESENT AT THE TIME OF SURGERY. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. MFG DATE: 11/14/2017. ENG EVAL COMPLETED ON 11/15/2018 BY CC. PER INVESTIGATION CONDUCTED UNDER CAPA (B)(4), 9 COMPLAINTS INCLUDING 10 DEVICES WERE EVALUATED AND THE REPORTED CONDITION WAS CONFIRMED. THE ROOT CAUSE WAS DETERMINED TO BE THAT THE DESIGN OF THE SNAP RING THAT HOLDS THE BUTTON IN PLACE WAS NOT ROBUST. THE AMOUNT OF THE OVERLAP BETWEEN THE SNAP RING AND THE LIP OF THE HOUSING WAS ON THE LOW SIDE, WHICH INCREASES THE RISK OF DISENGAGEMENT. INVESTIGATION CONFIRMED THAT ALL PARTS WERE MANUFACTURED TO SPECIFICATION. A REVIEW OF THE SPECIFICATIONS ON THE ENGINEERING PRINT REVEALED THAT THE OVERLAP BETWEEN THE SNAP RING'S OD (.720+.000/-.004) AND THE ID OF THE HOUSING LIP (.705+/-.003) IS ON THE LOW SIDE. FROM THE HEALTH RISK INDEX (HRI) CONDUCTED FOR CRC-2018-08-31-02: THE BUTTON COMPONENT FELL OFF OF THE TRULIANT TIBIAL TAMP GUIDE DURING VARIOUS OCCASIONS (USAGE, CLEANING, TRANSPORTATION, ETC.). EXACTECH HAS RECEIVED 9 INCIDENT REPORTS AFTER MORE THAN 3000 CASES (OCCURRENCE RATE ~3%). THERE IS NO ANTICIPATED HARM TO THE PATIENT DUE TO THE DELAY IN SURGERY. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. PER THE TRILLIANT PRIMARY KNEE SYSTEMS OPERATIVE TECHNIQUE (LIT 712-35-31, REV A), THE TIBIAL TAMP GUIDE IS USED FOR TAMPING OF THE TIBIAL BASEPLATE TRIAL DURING KNEE IMPLANT SURGERY VIA IMPACTION AND RETRO-IMPACTION. THE DEVICE IFU (700-096-0160) WARNS THAT ONLY QUALIFIED SURGEONS KNOWLEDGEABLE IN ANATOMY, BIOMECHANICS, AND RECONSTRUCTIVE SURGERY SHOULD USE THESE DEVICES. THE SURGEON MUST BE FULLY KNOWLEDGEABLE ABOUT ALL ASPECTS OF THE TRULIANT FEMORAL COMPONENTS, TIBIAL INSERTS AND TIBIAL TRAYS SURGICAL TECHNIQUE AND USE THESE IMPLANTS IN ACCORDANCE WITH THE RESPECTIVE INDICATIONS AND CONTRAINDICATIONS SUMMARIZED IN THIS DOCUMENT. IN ADDITION, THE SURGEON SHOULD BE FULLY KNOWLEDGEABLE ABOUT THE COMPATIBILITY BETWEEN SYSTEM COMPONENTS AND USE ACCORDINGLY. FINALLY, THE SURGEON MUST BE TRAINED ON THE PROPER OPERATIVE TECHNIQUE INCLUDING THE PROPER USE OF SYSTEM INSTRUMENTATION. IT IS ESSENTIAL THAT THE PROPER SURGICAL TECHNIQUE BE USED WHEN SECURING LOCKING COMPONENTS. AGGRESSIVE IMPACTION TECHNIQUES AND FORCEFUL ASSEMBLY METHODS MAY DAMAGE COMPONENTS AND LEAD TO DEVICE FAILURE. REPEATED ASSEMBLY AND DISASSEMBLY OF MODULAR COMPONENTS COULD COMPROMISE THE STRENGTH OF LOCKING MECHANISMS. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INST RUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. THERE WAS NO PATIENT ADVERSE EVENT. PER INVESTIGATION CONDUCTED UNDER CAPA (B)(4), THE ROOT CAUSE WAS DETERMINED TO BE THAT THE DESIGN OF THE SNAP RING THAT HOLDS THE BUTTON IN PLACE WAS NOT ROBUST. THE AMOUNT OF THE OVERLAP BETWEEN THE SNAP RING AND THE LIP OF THE HOUSING WAS ON THE LOW SIDE, WHICH INCREASES THE RISK OF DISENGAGEMENT. INFORMATION WAS NOT PROVIDED FOR: AGE, SEX, WEIGHT, ETHNICITY, RACE; DATE OF EVENT; INITIAL REPORTER NAME AND ADDRESS. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10089 | 2022-04-26 | 2017-08-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | OPTETRAK LOGIC FIT TIBIAL TRAY, CEMENTED, SIZE 3.5F/2.5T | EXACTECH, INC. |  | 02-012-45-3525 |  | IT WAS REPORTED BY A SURGEON FROM AUSTRALIA THAT THE TIBIAL BASEPLATE (IMPLANT) WAS SHIFTED ALONG THE MEDIAL/LATERAL AS SEEN ON X-RAY, COMPARED TO THE BASEPLATE TRIAL, WHEREAS THE TRIAL WAS WELL ALIGNED. THE PATIENT RESULTS WERE REPORTED AS "GOOD" AT THE FIRST POST-OP VISIT. MULTIPLE ATTEMPTS WERE MADE FOR INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE COMPLAINT DEVICES WERE NOT RETURNED FOR ANALYSIS. A SIBLING DEVICE OF THE TIBIAL TRAY WAS RETRIEVED FOR EVALUATION. THE FREQUENCY OF OCCURRENCE RANKING IS VERY LOW; THEREFORE, THIS DOES NOT APPEAR TO BE A DESIGN ISSUE. A SIBLING DEVICE OF THE TIBIAL TRAY WAS RETRIEVED FOR EVALUATION, THE DEVICE MET REQUIREMENTS, THEREFORE, THIS ISSUED DOES NOT APPEAR TO BE MANUFACTURING-RELATED. THERE IS NO USER REPORTED ISSUES. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. THE COMPLAINT REPORTED IN EXPERIENCE (B)(4) WAS LIKELY THE RESULT OF ASEPTIC LOOSENING OF THE TIBIAL IMPLANT, WHICH DAMAGED THE BOND OF THE IMPLANTS TO THE BONE. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE COMPONENT WAS NOT RETURNED FOR EVALUATION. IN A REVIEW OF THE LABELING AND IFU 700-096-004 REV. N, THERE IS LISTED CONTRAINDICATION FOR USE AS PATIENT'S AGE, WEIGHT, OR ACTIVITY LEVEL THAT WOULD CAUSE THE SURGEON TO EXPECT EARLY FAILURE OF THE SYSTEM. ONLY QUALIFIED SURGEONS ARE TO USE THESE PRODUCTS WHO ARE FULLY KNOWLEDGEABLE ABOUT ALL ASPECTS OF THE SURGICAL TECHNIQUE AND USE OF THESE IMPLANTS, HAVE FULL KNOWLEDGE ABOUT THE SYSTEM COMPATIBILITY'S, AND MUST BE FULLY TRAINED TO PROPERLY USE THE SYSTEM AND INSTRUMENTATION. ALL PATIENTS SHOULD BE INSTRUCTED ON THE LIMITATIONS OF THE PROSTHESIS, THE PATIENT SHOULD BE CAUTIONED TO MONITOR ACTIVITIES AND PROTECT THE REPLACED JOINT FROM UNREASONABLE STRESSES AND FOLLOW THE WRITTEN INSTRUCTIONS OF THE PHYSICIAN WITH RESPECT TO FOLLOW-UP CARE AND TREATMENT. DEVICE SPECIFIC RISKS - FRACTURE, MIGRATION, LOOSENING, SUBLUXATION, OR DISLOCATION OF THE PROSTHESIS OR ANY OF ITS COMPONENTS, ANY OF WHICH MAY REQUIRE A SECOND SURGICAL INTERVENTION OR REVISION. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE HAS NOT BEEN A REVISION REPORTED FOR THIS SUBSIDENCE/MIGRATION OF AN IMPLANT. AN INVESTIGATION WAS CONDUCTED; THE EVENT WAS LIKELY THE RESULT OF ASEPTIC LOOSENING OF THE TIBIAL IMPLANT, WHICH DAMAGED THE BOND OF THE IMPLANTS TO THE BONE. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE COMPONENT WAS NOT RETURNED FOR EVALUATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10068 | 2022-04-26 | 2018-10-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | HUMERAL HEAD IMPACTOR TIP | EXACTECH, INC. |  | 311-07-07 |  | IT WAS REPORTED FROM THE US THAT DURING AN INITIAL ORTHOPEDIC SURGICAL PROCEDURE THE SURGEON EXPERIENCED EVENTS WITH THE INSTRUMENTATION. ONE TRI-DRIVE CANULATED SHAFT BROKE IN THE HANDS OF THE SCRUB TECH AS HE WAS LOADING A DRILL TIP PRIOR TO SURGICAL USE. THE SECOND TRI-DRIVE CANULATED SHAFT BROKE AT THE BACK TABLE AFTER BEING USED FOR GLENOID DRILL HOLE. THE REPORTED DEVICE MALFUNCTION OF A DISASSEMBLED TRI-DRIVE MODULAR REAMER HANDLE DID NOT CAUSE OR CONTRIBUTE TO SERIOUS INJURY OR DEATH AND IS NOT LIKELY TO CAUSE OR CONTRIBUTE TO SERIOUS INJURY OR DEATH IF THE MALFUNCTION WERE TO RECUR. ANOTHER INSTRUMENT IS READILY AVAILABLE. IN THIS SAME CASE WHEN THE SURGEON WAS IMPACTING THE HUMERAL HEAD IMPLANT A SMALL PORTION OF THE IMPACTOR TIP (BLACK PLASTIC PORTION) BROKE OFF. IT DID NOT ENTER THE PATIENT AND WAS DISCARDED BY THE SCRUB TECH. THE PATIENT WAS STABLE LEAVING THE OR. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. MFG-RELATED ISSUES: EXACTECH IS AWARE OF 2 OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT, WHICH HAS BEEN IN THE FIELD SINCE 2014. THE DEVICE HISTORY RECORD WAS REVIEWED, AND ALL PARTS WERE ACCEPTED WITH CONFORMANCE TO THE DEVICE SPECIFICATIONS. THEREFORE; THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. THE CAUSE OF THE BROKEN TRI-DRIVE MODULAR REAMER HANDLE REPORTED IN EXPERIENCE C2018-660 COULD NOT BE DETERMINED BECAUSE ADEQUATE INFORMATION REGARDING THE FAILURE OF THE DEVICE WAS NOT PROVIDED, AND THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. 311-07-07 DESIGN-RELATED ISSUES: THE DESIGN OF THE HUMERAL HEAD IMPACTOR TIP HAS BEEN IN THE FIELD SINCE 2004. EXACTECH IS AWARE OF 6 COMPLAINT REPORTS INVOLVING BROKEN HUMERAL HEAD IMPACTOR TIPS SINCE 2008. THE FREQUENCY OF OCCURRENCE RANKING SCALE IS VERY LOW. THEREFORE; THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT, WHICH HAS BEEN IN THE FIELD SINCE 2012. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. THE BROKEN HUMERAL HEAD IMPACTOR TIP WAS LIKELY THE RESULT OF EXCESSIVE IMPACTION FORCES, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. DEVICES WERE RECEIVED FOR ANALYSIS. ENG EVAL COMPLETED ON 2019.06.18 BY MB. 315-25-00 DESIGN-RELATED ISSUES: THE DESIGN OF THE TRI-DRIVE MODULAR REAMER HANDLE HAS BEEN IN THE FIELD SINCE 2009. EXACTECH IS AWARE OF 39 COMPLAINT REPORTS INVOLVING 47 BROKEN TRI-DRIVE MODULAR REAMER HANDLES SINCE THEIR INTRODUCTION. THE FREQUENCY OF OCCURRENCE RANKING SCALE IS VERY LOW; THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THE REPORTED DEVICE MALFUNCTION OF A DISASSEMBLED TRI-DRIVE MODULAR REAMER HANDLE DID NOT CAUSE OR CONTRIBUTE TO SERIOUS INJURY OR DEATH AND IS NOT LIKELY TO CAUSE OR CONTRIBUTE TO SERIOUS INJURY OR DEATH IF THE MALFUNCTION WERE TO RECUR. ANOTHER INSTRUMENT IS READILY AVAILABLE. THERE WAS NO PATIENT ADVERSE EVENT, THE BROKEN TIP DID NOT FALL INTO THE SURGICAL SITE. AN INVESTIGATION WAS CONDUCTED; THE BROKEN HUMERAL HEAD IMPACTOR TIP WAS LIKELY THE RESULT OF EXCESSIVE IMPACTION FORCES, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE. HOWEVER; THIS CANNOT BE CONFIRMED BECAUSE THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. THE CAUSE OF THE BROKEN TRI-DRIVE MODULAR REAMER HANDLE REPORTED COULD NOT BE DETERMINED BECAUSE ADEQUATE INFORMATION REGARDING THE FAILURE OF THE DEVICE WAS NOT PROVIDED, AND THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. CATALOG # 315-25-00. SERIAL/LOT # (B)(4). DESCRIPTION: TRI-DRIVE MODULAR REAMER HANDLE X2. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10074 | 2022-04-26 | 2018-03-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRULIANT LOCKING FEMORAL IMPACTOR | EXACTECH, INC. |  | 02-029-19-1000 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE IMPACTOR. A SCREW AND WASHER WERE FOUND AT THE END OF THE CASE. ALL PIECES WERE ACCOUNTED FOR. NO PIECES WERE LEFT IN THE PATIENT. THE PATIENT'S HEALTH WAS STABLE LEAVING THE OPERATING ROOM THERE WAS A 5-MINUTE DELAY, BUT THE PATIENT WAS NOT ADVERSELY AFFECTED. THE DEVICE WAS NOT RETURNED FOR ANALYSIS, THE REPRESENTATIVE STATED IT HAD BEEN RETURNED. PHOTOS WERE PROVIDED. THE REP WAS PRESENT AT THE TIME OF SURGERY. AN ATTEMPT WAS MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REPRESENTATIVE AND A RESPONSE WAS RECEIVED. PER THE REPRESENTATIVE, THERE IS NO MORE INFORMATION AVAILABLE. THE DOCTOR HAS NO IDEA OR REMEMBER THE INCIDENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS NOT AVAILABLE FOR ANALYSIS. INVESTIGATION # AND DESCRIPTION: CAPA (B)(4): INVESTIGATION REVEALED THAT THE FAILURE OF THE CAP SCREW WAS DIRECTLY RELATED TO INADEQUATE WELDING BETWEEN THE CAP SCREW AND THE WASHER SUBCOMPONENT APPLIED DURING MANUFACTURING ASSEMBLY. THERE WAS ONLY A MINIMAL SPOT WELD APPLIED WHICH WAS VULNERABLE DURING IMPACTION USE. A REVIEW OF PRINT SHOWED THE WELD REQUIREMENT WAS CALLED OUT AS "TACK WELD", WHICH WAS NOT SPECIFIC ENOUGH AND LEFT ROOM FOR INTERPRETATION. FINDINGS: A SCREW AND WASHER WERE FOUND AT THE END OF THE CASE. THE DISASSEMBLY OF THE DEVICE WAS LIKELY DUE TO INADEQUATE WELDING OF THE SCREW AS COVERED IN CAPA (B)(4). THE DEVICE AND DISASSEMBLED PARTS WERE NOT RETURNED TO EXACTECH AT THE TIME OF THIS EVALUATION. THE COMPLAINT WILL BE REOPENED IF THE DEVICES ARE RETURNED. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: THE REPORTED COMPLAINT IN (B)(4) WAS LIKELY THE RESULT OF THE CAP SCREW NOT BEING PROPERLY WELDED TO THE BODY OF THE DEVICE RESULTING IN THE SCREW LOOSENING AND FALLING OFF WITH THE ASSOCIATED WASHER. CORRECTIVE ACTION TAKEN: THE COMMITTEE HAS DETERMINED THAT THE OPERATIVE TECHNIQUE WILL BE UPDATED, AND A FIELD NOTICE WILL BE SENT TO THE AGENTS AND REPRESENTATIVES NOTIFYING THEM OF THE UPDATE. ONCE THE DESIGN CHANGE IS VALIDATED AND REPLACEMENT DEVICES ARE AVAILABLE, THE DEVICES WILL BE REMOVED FROM THE FIELD. THE CURRENT INVESTIGATION AND CAPA ARE STILL ONGOING AT THE TIME OF THIS EVALUATION AND CORRECTIVE ACTIONS ARE SUBJECT TO CHANGE. REFER TO CAPA (B)(4) FOR MOST UP TO DATE INFORMATION. EXACTECH IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED INSTRUMENT MALFUNCTION IS RELATED TO ANY MANUFACTURING ISSUES NOR DID IT LEAD TO ANY PATIENT IMPACT. AN INVESTIGATION WAS CONDUCTED: "THE REPORTED COMPLAINT WAS LIKELY THE RESULT OF THE CAP SCREW NOT BEING PROPERLY WELDED TO THE BODY OF THE DEVICE RESULTING IN THE SCREW LOOSENING AND FALLING OFF WITH THE ASSOCIATED WASHER." | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10091 | 2022-04-26 | 2018-02-19 | I, n, j, u, r, y | Manufacturer report | EXACTECH | NOVATION CROWN CUP GXL, +5MM LINER | EXACTECH, INC. |  | 136-40-53 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE LINER. THE IMPLANT WAS MISSING A TAB; THE STAFF IS UNSURE IF TAB WAS MISSING UPON OPENING OF IMPLANT OR IF IT BROKE OFF AFTER IMPLANTATION. IMPLANT WAS REMOVED FROM PATIENT, BACKUP IMPLANT WAS OPENED AND IMPLANTED, AND CASE PROCEEDED WITHOUT FURTHER INCIDENT. THE DEFORMED AND BROKEN TABS ON THE CROWN CUP LINER APPEAR TO BE CONSISTENT WITH ATTEMPTING TO REMOVE THE LINER WHICH LIKELY DAMAGED THE LOCKING TABS. PATIENT LEFT OPERATING ROOM IN STABLE CONDITION AND REP WAS PRESENT FOR SURGERY. THE DEVICE WAS RETURNED. THE REP WAS PRESENT AT THE TIME OF SURGERY. PATIENT DEMOGRAPHIC INFORMATION WAS NOT PROVIDED. INACTIVE AGENT/AGENCY, NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4) HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS AVAILABLE FOR ANALYSIS. ENGINEERING ANALYSIS COMPLETED BY MB. VISUAL EXAMINATION: THE DEFORMED AND BROKEN TABS ON THE CROWN CUP LINER APPEAR TO BE CONSISTENT WITH ATTEMPTING TO REMOVE THE LINER WHICH LIKELY DAMAGED THE LOCKING TABS. THE DESIGN OF THE NOVATION CROWN CUP LINER HAS BEEN IN THE FIELD SINCE 2011. EXACTECH HAS RECEIVED 2 SIMILAR COMPLAINT REPORTS INVOLVING NOVATION CROWN CUP LINER IMPLANTS SINCE 2011. BECAUSE THE NOVATION CROWN CUP LINER IS USED IN TOTAL HIP SURGERIES, SALES DATA FOR ACETABULAR SHELLS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4). THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH HAS NOT RECEIVED ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 7 UNITS, WHICH HAS BEEN IN THE FIELD SINCE 2017. ACCORDING TO THE SALES DATA, 4 PIECES OF THIS 7-PIECE LOT HAVE NOT BEEN SOLD. THIS IMPLIES THAT THE OTHER 2 PIECES HAVE BEEN SUCCESSFULLY IMPLANTED. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS NOVATION CROWN CUP, 750-2011-012-RMR-IMPLANTS REV D, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 28. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. THE BROKEN TABS REPORTED IN EXPERIENCE (B)(6) WERE LIKELY THE RESULT OF ATTEMPTING TO REMOVE THE LINER AT AN ANGLE, CAUSING THREE OF THE TABS TO DEFORM. THE DEFORMED LOCKING BUTTON AND SCRATCHES WERE LIKELY CAUSED DURING THE REMOVAL OF THE LINER. EXACTECH IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO ANY PATIENT IMPACT. THE DEFORMED AND BROKEN TABS ON THE CROWN CUP LINER APPEAR TO BE CONSISTENT WITH ATTEMPTING TO REMOVE THE LINER WHICH LIKELY DAMAGED THE LOCKING TABS. AN INVESTIGATION WAS CONDUCTED: THE BROKEN TABS REPORTED WERE LIKELY THE RESULT OF ATTEMPTING TO REMOVE THE LINER AT AN ANGLE, CAUSING THREE OF THE TABS TO DEFORM. THE DEFORMED LOCKING BUTTON AND SCRATCHES WERE LIKELY CAUSED DURING THE REMOVAL OF THE LINER. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10073 | 2022-04-26 | 2017-10-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRI-DRIVE, LONG COLLAR | EXACTECH, INC. |  | BSK-9533 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE TRI-DRIVE. WHILE ATTEMPTING TO DRILL GLENOID PEG HOLES, THE INSTRUMENT CAME APART. IT APPEARED THAT A METAL PIN WAS MISSING, BUT IT COULD NOT BE LOCATED. NO POST-OP X-RAYS WERE TAKEN TO CONFIRM THE PIN WASN'T LEFT IN THE PATIENT. THE SURGEON DOES NOT BELIEVE THE PIN WAS LEFT IN THE PATIENT AS THE PIN IS ON THE PROXIMAL END AND WASN'T OVER THE WOUND. NOTHING DETRIMENTAL HAPPENED TO THE PATIENT. THERE WAS A DELAY, BUT THE PATIENT WAS NOT ADVERSELY AFFECTED. NO PATIENT DEMOGRAPHIC INFORMATION WAS PROVIDED. THE REP WAS PRESENT AT THE TIME OF SURGERY. THE DEVICE WAS RETURNED ON 10/30/2017 ON RG406687. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT FOLLOWING MULTIPLE ATTEMPTS. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS AVAILABLE FOR ANALYSIS. ENGINEERING ANALYSIS COMPLETED BY (B)(6) ON 8/10/2018. DESIGN-RELATED ISSUES: THE DESIGN OF THE MODULAR TRI-DRIVE HANDLE HAS BEEN IN THE FIELD SINCE 2009. EXACTECH HAS RECEIVED 8 COMPLAINT REPORTS INVOLVING 9 TRI-DRIVES DUE TO HOLDING PIN DISASSOCIATION SINCE 2009. THE FEATURE INVOLVED IN THIS COMPLAINT DEVICE IS THE SAME AS THE FEATURE ON THE MODULAR TRI-DRIVE HANDLE (315-25-00); THEREFORE, THE SAME SALES DATA WILL BE PULLED FOR THIS COMPLAINT. BECAUSE TRI-DRIVE HANDLES ARE USED IN TOTAL SHOULDER SURGERIES, SALES DATA FOR HUMERAL STEMS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4) THIS IS CONSIDERED " VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS AWARE OF 1 OTHER COMPLAINT INVOLVING A PART FROM THIS MANUFACTURING LOT, WHICH HAS BEEN IN THE FIELD SINCE 2011. HOWEVER, THE OTHER COMPLAINT IS DUE TO A DIFFERENT FAILURE MODE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED FOR THE MODULAR TRI- DRIVE HANDLE TOENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THIS RMR IS APPROPRIATE BECAUSE THE MODE OF FAILURE IS NOT ASSOCIATED WITH THE CHANGE MADE TO THE ORIGINAL DESIGN TO PRODUCE THIS SPECIAL LONG COLLAR TRI- DRIVE . THE RMR, 750-2007-026-RMR-INSTRUMENT REV H, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 20. NO CORRECTIVE ACTION IS REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. THE BROKEN INSTRUMENT REPORTED IN EXPERIENCE (B)(4) WAS LIKELY THE RESULT OF THE DOWEL PIN DISASSOCIATING FROM THE ASSEMBLY, WHICH ALLOWED FOR THE REAR CONNECTION COMPONENT TO DETACH. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. EXACTECH IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED INSTRUMENT MALFUNCTION IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. WHILE ATTEMPTING TO DRILL GLENOID PEG HOLES, THE INSTRUMENT CAME APART. IT APPEARED THAT A METAL PIN WAS MISSING, BUT IT COULD NOT BE LOCATED. ALL INSTRUMENTS SHOULD BE INSPECTED BEFORE USE. AN INVESTIGATION WAS CONDUCTED: THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF THE DOWEL PIN DISASSOCIATING FROM THE ASSEMBLY, WHICH ALLOWED FOR THE REAR CONNECTION COMPONENT TO DETACH. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10084 | 2022-04-26 | 2019-09-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | GLENOID SKID | EXACTECH, INC. |  | 315-03-01 |  | EXACTECH COMPLAINTS RECEIVED DEVICES VIA AN RG TRANSFER REPORT FROM AND AGENCY INDICATING THE DEVICES WERE DEFECTIVE/COMPLAINT ITEMS FROM (B)(4). NOT ALL PIECES OF THE BROKEN DEVICES WERE RETURNED. NO CUSTOMER OR PATIENT IMPACT INFORMATION WAS PROVIDED. ATTEMPTS WERE MADE TO REQUEST ADDITIONAL INFORMATION FROM THE AGENCY AND A RESPONSE WAS RECEIVED TO ASK WHO REPORTED THE EVENT. THERE WAS NO REPORTING REPRESENTATIVE INDICATED ON THE RG. NO ADDITIONAL INFORMATION WAS PROVIDED ABOUT THESE DEVICES. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICES WERE AVAILABLE FOR ANALYSIS. ENGINEERING ANALYSIS COMPLETED BY (B)(4) ON 12/2/2019. FINDINGS: THE BROKEN GLENOID SKIDS APPEAR CONSISTENT WITH APPLYING A BENDING MOMENT DURING USE. THE OTHER PIECES OF THE GLENOID SKIDS WERE NOT RETURNED. MOST LIKELY UNDERLYING CAUSE: BASED ON PREVIOUS SIMILAR INVESTIGATIONS, THE BROKEN GLENOID SKIDS REPORTED IN (B)(4) WERE LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO ONE END OF THE INSTRUMENT, WHICH LED TO FRACTURE OF THE DEVICES. HOWEVER, NO INFORMATION REGARDING THEIR USE DURING SURGERY WAS PROVIDED. RMR REVIEW: THE RMR AND RACR FOR THIS GLENOID SKID, RMR-00006 REV - AND RACR-00009 REV -, WERE REVIEWED. THE RISK FOR THE REPORTED FAILURE IS NOT CAPTURED FOR THIS DEVICE. THE RACR WILL BE UPDATED UNDER (B)(4) TO CAPTURE THE APPROPRIATE RISK FOR THIS DEVICE. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICES USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT FROM THE AGENCY. NOT ALL PIECES OF THE DEVICES WERE RETURNED. AN INVESTIGATION WAS CONDUCTED: BASED ON PREVIOUS SIMILAR INVESTIGATIONS, THE BROKEN GLENOID SKIDS REPORTED WERE LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO ONE END OF THE INSTRUMENT, WHICH LED TO FRACTURE OF THE DEVICES. HOWEVER, NO INFORMATION REGARDING THEIR USE DURING SURGERY WAS PROVIDED. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10069 | 2022-04-26 | 2018-10-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | : DRIVER, MODULAR CANNULATED TRI-DRIVE | EXACTECH, INC. |  | 315-25-00 |  | IT WAS REPORTED FROM THE US THAT DURING AND ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED DIFFICULTY WITH THE INSTRUMENTS. THE TRI-DRIVE BROKE WHILE DRILLING THE CENTER GLENOID POST (LARGE MALE WITH TIGHT EXPOSURE AND VERY HARD BONE), AND THE SURGEON NOTICED THE SHAFT RUBBING ON THE FUKUDA RETRACTOR. IN THE SAME CASE, THE IMPACTOR TIP HAD A SMALL PORTION OF THE BLACK PLASTIC BREAK AND FALL TO THE FLOOR WHEN THE SURGEON USED THE INSTRUMENT TO IMPACT THE LARGEST HUMERAL HEAD COMPONENT (53MM). THERE WAS NO COMPROMISE TO THE PATIENT OR SURGERY, AS THE SURGEON WAS DONE WITH THE INSTRUMENTS. NO DELAY TO SURGERY. THE PATIENT'S HEALTH WAS STABLE LEAVING THE OPERATING ROOM. NO PIECES WERE LEFT IN THE PATIENT, AS PROVEN BY THE STANDARD POST OP X-RAY. AGENT WAS PRESENT AT TIME OF SURGERY. THE AGENCY IS INACTIVE, NO FURTHER INFORMATION IS AVAILABLE. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICES WERE RECEIVED FOR ANALYSIS. ENG EVAL COMPLETED ON 2019.08.06 BY NF. TRI-DRIVE MODULAR HANDLE: (B)(4) AND CAPA (B)(4). THE DESIGN OF THE TRI-DRIVE MODULAR HANDLE HAS BEEN IN THE FIELD SINCE 2009. EXACTECH HAS RECEIVED 45 SIMILAR COMPLAINT REPORTS INVOLVING 47 TRI-DRIVE MODULAR HANDLE DUE TO SLEEVE DISASSEMBLY SINCE 2009. BECAUSE THE TRI-DRIVE IS USED IN ANATOMIC AND REVERSE TOTAL SHOULDER SURGERIES, SALES DATA FOR ANATOMIC GLENOID COMPONENTS AND GLENOSPHERE COMPONENTS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4). THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE IN THE RMR. THE DISASSEMBLED TRI-DRIVE MODULAR REAMER DRIVER MAY RESULT IN A DELAY IN SURGERY TO FIND A NEW ONE. TYPICALLY, IN ANATOMIC SHOULDER CASES BOTH THE ANATOMIC AND THE REVERSE SETS ARE IN THE OPERATING ROOM. IF THE TRI-DRIVE WERE TO BREAK, A SECOND TRI-DRIVE WOULD BE IN THE REVERSE SHOULDER SET. BASED ON THE INVESTIGATIONS CONDUCTED IT APPEARS THAT THIS ISSUE IS RELATED TO A USER ISSUE WHEREIN THE USER HOLDS ONTO THE COLOR OF THE TRI-DRIVE HANDLE WHEN THE DEVICE IS UNDER POWERED ROTATION. THIS CAUSES FORCE TO BE EXERTED ON THE INTERNAL PARTS WHICH RESULTS IN DISASSEMBLY. MOST LIKELY ROOT CAUSE-THE BROKEN TRI-DRIVE MODULAR HANDLE REPORTED IN EXPERIENCE C2018-663 WAS LIKELY THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY, ALLOWING THE TIP TO DISASSEMBLE. HUMERAL HEAD IMPACTOR: VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR 10X OPTICAL UNLESS OTHERWISE SPECIFIED)- THE BROKEN HUMERAL HEAD IMPACTOR TIP APPEARS CONSISTENT WITH CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE RADEL IMPACTION SURFACES. THIS BREAKAGE WAS LIKELY THE RESULT OF YEARS OF NORMAL USE AND SUBSEQUENT STERILIZATIONS. HEAD IMPACTOR TIP DUE TO A FRACTURED RADEL IMPACTION SURFACE SINCE 2008. THE FREQUENCY OF OCCURRENCE RANKING SCALE IS "VERY LOW", THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN- RELATED. EXACTECH IS AWARE OF ONE OTHER COMPLAINT INVOLVING PARTS FROM THIS HUMERAL HEAD IMPACTOR TIP MANUFACTURING LOT OF (B)(4) UNITS, WHICH HAS BEEN IN THE FIELD SINCE 2017. THE DEVICE HISTORY RECORD FOR THE HUMERAL HEAD IMPACTOR TIP WAS REVIEWED, AND ALL PARTS ACCEPTED WITH CONFORMANCE TO THE DEVICE SPECIFICATIONS. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMRS. MOST LIKELY ROOT CAUSE-THE BROKEN HUMERAL HEAD IMPACTOR REPORTED IN EXPERIENCE C2018-663 WAS LIKELY THE RESULT OF CRACK INITIATION, PROPAGATION, AND ULTIMATE FAILURE OF THE RADEL IMPACTION SURFACE RELATED TO NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE WAS NO PATIENT ADVERSE EVENT, NO BROKEN PIECES OF THE DEVICE WERE LEFT IN THE PATIENT, THIS WAS CONFIRMED WITH THE ROUTINE POST-OP X-RAY. AN INVESTIGATION WAS CONDUCTED FOR THE TRI-DRIVE MODULAR HANDLE UNDER (B)(4) AND CAPA (B)(4) THE BROKEN TRI-DRIVE MODULAR HANDLE REPORTED WAS LIKELY THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY, ALLOWING THE TIP TO DISASSEMBLE. AN INVESTIGATION WAS CONDUCTED FOR THE HUMERAL HEAD IMPACTOR TIP MANUFACTURING; THE BROKEN HUMERAL HEAD IMPACTOR REPORTED WAS LIKELY THE RESULT OF CRACK INITIATION, PROPAGATION, AND ULTIMATE FAILURE OF THE RADEL IMPACTION SURFACE RELATED TO NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10062 | 2022-04-26 | 2018-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRULIANT TIBIAL TAMP GUIDE | EXACTECH, INC. |  | 02-029-22-2000 |  | IT WAS REPORTED FROM THE US THAT DURING A BILATERAL TOTAL KNEE SURGERY THE SURGEON UTILIZED TRULIANT KNEE INSTRUMENTATION TO PREPARE THE KNEES TO ACCEPT TRULIANT POSTERIOR STABILIZED IMPLANTS. THE RIGHT KNEE WAS PREPPED FIRST. DURING PREPARATION OF THE RIGHT KNEE, WHEN THE TIBIAL TAMP HANDLE WAS HANDED TO THE SURGEON BY THE SURGICAL TECH, THE MODULAR TAMP HEAD FELL OFF THE HANDLE. THEY PUT IT BACK ON AND PROCEEDED WITH TAMPING THE TIBIA WITHOUT INCIDENT. THE LEFT KNEE WAS PREPPED AND THE TIBIAL TAMP USED WITHOUT INCIDENT. WHEN THE INSTRUMENTS WERE BROUGHT TO THE DECONTAMINATION ROOM FOR CLEANING, THE CLEANING PERSON NOTICED THAT THE TIBIAL TAMP HANDLE BUTTON WAS OUT OF POSITION, AND OUT OF THE HANDLE. THE SPRING THAT SITS BEHIND THE BUTTON IN THE HANDLE WAS MISSING. A CURSORY SEARCH FOR THE SPRING WAS MADE IN THE DECONTAMINATION ROOM. A MORE THOROUGH SEARCH WAS PLANNED FOR ON THE CLEAN SIDE (IN THE CLEAN ROOM). AT THIS POINT, IT WAS DETERMINED TO CHECK IF THE PATIENT WAS INVOLVED, BECAUSE THE SPRING WAS UNACCOUNTED FOR. THE SURGEON HAD LEFT THE HOSPITAL, THE PA WAS PRESENT AND CALLED FOR X-RAY REVIEW. THE SURGEON ALREADY HAD ORDERED POST-OPERATIVE X-RAYS. THE X-RAY REVIEW REVEALED NO EVIDENCE OF THE SPRING IN THE PATIENT. UPON FURTHER SEARCH FOR THE SPRING IN THE CLEAN ROOM THE SPRING WAS LOCATED. INACTIVE AGENCY, NO FURTHER INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE WAS RECEIVED FOR ANALYSIS. ENG EVAL COMPLETED ON 10/10/2018 BY AD. CAPA2018-35 ENGINEERING PRINTS OF 02-029-122-2000 WERE REVIEWED. THE OVERLAP BETWEEN THE SNAP RING'S OD AND THE ID OF THE HOUSING LIP IS .008 FOR THE WORST-CASE SCENARIO. ALTHOUGH THE OVERLAP IS POSITIVE, THE SMALL AMOUNT COULD CAUSE DISENGAGEMENT ROOT CAUSE PER CAPA2018-35, THE ROOT CAUSE FOR THE BUTTON DISENGAGEMENT WAS LIKELY THE RESULT OF INSUFFICIENT OVERLAP BETWEEN THE SNAP RING AND THE LIP OF THE HOUSING TO PREVENT THE OCCURRENCE OF DISENGAGEMENT. CORRECTIVE ACTION TAKEN: PER CRC2018-08-31-02, THE COMMITTEE HAS DETERMINED THAT THE DESIGN WILL BE UPDATED, TESTED, AND IMPLEMENTED, AFTER WHICH A STAGED MARKET WITHDRAWAL WAS CONDUCTED. PER CAPA(B)(4): A STOCK AUDIT WAS PERFORMED FOR THE DEVICES LOCATED IN 01 (COC). SEE NPR-40446 AND NPR-45222 FOR DETAILS. DEFECTIVE DEVICES WERE DISPOSITIONED FOR REWORK. THE DESIGN WAS UPDATED TO INCREASE THE OVERLAP OF THE SNAP RING AND THE HOUSING LIP COMPONENTS. SEE DCRTC-18-0401 AND DCRTC-19-0446 FOR DETAILS. THE DEFECTED DEVICES RETURNED FROM THE FIELD ARE TO BE REWORKED. SEE STAGED MARKET WITHDRAWAL OF TRULIANT TIBIAL TAMP GUIDE DOCUMENTED IN THE CAPA FOR DETAILS. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. THE PATIENT WAS X-RAYED, AND IT WAS VERIFIED THAT THE MISSING SPRING WAS NOT RETAINED BY THE PATIENT. AN INVESTIGATION WAS CONDUCTED UNDER CAPA(B)(4); THE ROOT CAUSE FOR THE BUTTON DISENGAGEMENT WAS LIKELY THE RESULT OF INSUFFICIENT OVERLAP BETWEEN THE SNAP RING AND THE LIP OF THE HOUSING TO PREVENT THE OCCURRENCE OF DISENGAGEMENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10080 | 2022-04-26 | 2018-05-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | PATELLA CLAMP | EXACTECH, INC. |  | 201-62-00 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE CLAMP. THE PATELLA CLAMP BROKE WHILE CEMENTING THE IMPLANT. PULSLAVAGE BEFORE CLOSING, AND ANTIBIOTHERAPY FOR THE REST. "NOTHING SPECIAL." THE PATIENT'S HEALTH WAS STABLE LEAVING THE OPERATING ROOM. THE DEVICE WAS RETURNED ON 3/4/2019 ON RG301536. THE REP WAS PRESENT AT THE TIME OF SURGERY. NEW INFORMATION 1 JUL 2021: "THE INSTRUMENT COULD HAVE BEEN BROKEN DURING THE SURGERY. HOWEVER, THE BREAKAGE HAS BEEN NOTICED AT THE RECOVERY OF THE INSTRUMENTATION-SET BY OUR LOGISTICS DEPT., AFTER THE CLEANING REALIZED BY THE STERILIZATION TEAM OF THE HOSPITAL. NO INFORMATION HAS BEEN RECEIVED FROM THE HOSPITAL REGARDING THIS ISSUE AND [THE AGENCY] DIDN'T RECEIVE ANY REQUEST FOR COMPLEMENTARY INFORMATION REGARDING THIS ISSUE AS WELL." NO ADDITIONAL INFORMATION IS AVAILABLE. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS AVAILABLE FOR ANALYSIS. ENGINEERING ANALYSIS COMPLETED BY KN ON 3/15/2019. FINDINGS: THE PATELLAR CUP FRACTURED FROM NORMAL USE. MOST LIKELY UNDERLYING CAUSE: THE BROKEN PIECE WAS LIKELY THE RESULT OF MATERIAL DEGRADATION ASSOCIATED WITH YEARS OF NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. RMR REVIEW: 750-1992-001-RMR-INSTRUMENTS REV -, LINE 12. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. THE PATELLA CLAMP BROKE AND THE ACTIONS TAKEN WERE REPORTED AS "PULSLAVAGE BEFORE CLOSING, AND ANTIBIOTHERAPIE". AN INVESTIGATION WAS CONDUCTED: THE BROKEN PIECE WAS LIKELY THE RESULT OF MATERIAL DEGRADATION ASSOCIATED WITH YEARS OF NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10090 | 2022-04-26 | 2017-12-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | DRIVER, MODULAR CANNULATED TRI-DRIVE | EXACTECH, INC. |  | 315-25-00 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGICAL PROCEDURE THE SPRING-LOADED COLLET PORTION OR MODULAR TRI-DRIVE SHAFT DISSOCIATED DURING GLENOID PREPARATION. THE ISSUE WAS RESOLVED BY USING ANOTHER MODULAR CANNULATED TRI-DRIVE WHICH WAS PRESENT IN THE SET OF INSTRUMENTS. REPRESENTATIVE WAS PRESENT AT THE TIME OF SURGERY. THIS ISSUE HAD NO EFFECT ON THE PATIENT'S OUTCOME. THERE WAS NO DELAY TO SURGERY. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICES WAS RECEIVED FOR ANALYSIS. ENG ANALYSIS COMPLETED 09/19/2018 BY MB THE TIP APPEARS TO HAVE DISASSEMBLED. THE RETAINING RING IS MEANT TO HOLD THE SLEEVE AND SPRING COMPONENTS IN PLACE BUT APPEARS TO HAVE DISLODGED. THE USER LIKELY HELD THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY AND ALLOWED THE TIP TO DISASSEMBLE. THE 0.125 BALL BEARING LIKELY DETACHED WHEN THE TIP DISASSEMBLED. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE AND CONSISTENT WITH YEARS OF USE. DESIGN: THIS DESIGN HAS BEEN IN THE FIELD SINCE 2009. THERE HAVE BEEN NO OTHER COMPLAINT REPORTS INVOLVING THIS DEVICE FOR ANY REASON SINCE 2008. THE OCCURRENCE RATE IS LESS THAN (B)(4), "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THIS DOES NOT APPEAR TO BE DESIGN RELATED. MANUFACTURING: THERE HAVE BEEN 5 OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF (B)(4) UNITS, WHICH HAS BEEN IN THE FIELD SINCE 2012. THE DESIGN HISTORY RECORD WAS REVIEWED DURING A PREVIOUS COMPLAINT INVESTIGATION, ALL PIECES WERE ACCEPTED WITH CONFORMANCE TO THE DEVICE REQUIREMENTS. THIS DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE DISASSEMBLED DRIVER REPORTED WAS LIKELY THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY AND ALLOWED THE TIP TO DISASSEMBLE. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THE RETAINING RING BROKE LOOSE FROM THE ASSEMBLY AND ALLOWED THE TIP TO DISASSEMBLE. THERE WAS NO ADVERSE EVENT TO THE PATIENT. AN INVESTIGATION WAS CONDUCTED; THE DISASSEMBLED DRIVER REPORTED WAS LIKELY THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY AND ALLOWED THE TIP TO DISASSEMBLE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10082 | 2022-04-26 | 2019-10-31 | I, n, j, u, r, y | Manufacturer report | EXACTECH | NARROW REV 38MM REAMER MODULAR, CANNULATED | EXACTECH, INC. |  | 321-55-38 |  | IT WAS REPORTED FROM THE US THAT THE "SURGEON PUT BOW TIE REAMER ON MODULAR HANDLE (ON DRILL NOT REAM). PUT IT OVER A K WIRE AND RAN IT DOWN TO THE GLENOID. STARTED TO REAM THE GLENOID AND THE PIECE THAT ATTACHES TO THE HANDLE SHEARED OFF. IT DID END UP NOTCHING THE GLENOID, BUT WE WERE ABLE TO FACILITATE THE BASEPLATE ON WITHOUT COMPLICATIONS." MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. NO DEVICE RECEIVED FOR ANALYSIS. ANALYSIS COMPLETED ON 04/10/2020 BY AH. DESIGN-RELATED ISSUES: EXACTECH IS AWARE OF 5 COMPLAINT REPORTS INVOLVING A FRACTURED REAMER SHAFT SINCE 2008. BECAUSE THIS DEVICE IS USED IN TOTAL REVERSE SHOULDER SURGERIES, SALES DATA FOR GLENOSPHERES WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. MFG-RELATED ISSUES: MANUFACTURING-RELATED INFORMATION COULD NOT BE REVIEWED BECAUSE NO INFORMATION REGARDING MANUFACTURING LOTS WAS PROVIDED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) AND THE RISK ASSESSMENT AND CONTROLS REPORT (RACR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR EQUINOXE NARROW REAMERS, RMR-00007 REV B AND RACR-00010 REV B, WERE REVIEWED. THE RISK IS CAPTURED IN RISK ID 5.3. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR AND RACR. THE BROKEN REAMER REPORTED IN (B)(4) WAS LIKELY THE RESULT OF APPLYING A FORCE GREATER THAN YIELD STRENGTH OF THE MATERIAL AND SUBSEQUENT STERILIZATION CYCLES WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FAILURE. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE COMPONENT WAS NOT AVAILABLE FOR EVALUATION. ADDITIONAL COMMENTS: ACCORDING TO THE EXPERIENCE REPORT, "SURGEON PUT BOW TIE REAMER ON MODULAR HANDLE (ON DRILL NOT REAM). PUT IT OVER A K WIRE AND RAN IT DOWN TO THE GLENOID. STARTED TO REAM THE GLENOID AND THE PIECE THAT ATTACHES TO THE HANDLE SHEARED OFF." IT WAS REPORTED THAT THIS HAS OCCURRED ON 3 SEPARATE OCCASIONS. THE OTHER TWO EVENTS ARE CAPTURED IN (B)(4) AND (B)(4). IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY DESIGN ISSUES. THE GLENOID WAS REPORTED TO BE NOTCHED DUE TO THIS EVENT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN REAMER REPORTED WAS LIKELY THE RESULT OF APPLYING A FORCE GREATER THAN YIELD STRENGTH OF THE MATERIAL AND SUBSEQUENT STERILIZATION CYCLES WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FAILURE. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE COMPONENT WAS NOT AVAILABLE FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10070 | 2022-04-26 | 2017-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | HANDLE, MODULAR BROACH | EXACTECH, INC. |  | 301-03-01 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE BROACH HANDLE. THE COLLAR OF THE BROACH HANDLE CAME OFF DURING SURGERY WHILE THE BROACH HANDLE WAS BEING USED DURING BROACHING. NO PIECES ENDED UP IN THE WOUND SITE. THERE WAS A SLIGHT DELAY. MORE OF AN INCONVENIENCE FOR THE SURGEON. PATIENT WAS NOT AFFECTED AT ALL. GOOD SURGERY OUTCOME. SURGEON WAS ABLE TO SLIDE COLLAR BACK ON AND FINISH SURGERY. REPLACED HANDLE WITH NEW BROACH HANDLE. THE REP WAS PRESENT AT THE TIME OF SURGERY. NO PATIENT DEMOGRAPHIC INFORMATION PROVIDED. THE DEVICE WAS RETURNED ON (B)(6) 2017. INACTIVE AGENCY WITH NO HOSPITAL/OTHER CONTACT INFORMATION FOR THE EVENT. NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS AVAILABLE FOR ANALYSIS. ENGINEERING ANALYSIS COMPLETED BY MB ON 7/30/2018. DESIGN-RELATED ISSUES: THIS DESIGN HAS BEEN IN THE FIELD SINCE 2004. EXACTECH HAS RECEIVED 6 SIMILAR COMPLAINT REPORTS INVOLVING THIS DEVICE SINCE 2008. BECAUSE THIS DEVICE IS USED IN BOTH PRIMARY AND REVERSE SHOULDER REPLACEMENTS, SALES DATA FOR HUMERAL STEMS WERE USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH HAS NOT RECEIVED ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT, WHICH HAS BEEN IN THE FIELD SINCE 2014. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS MODULAR BROACH HANDLE, 750-2002-039-RMR-INSTRUMENTS ADDITIONS REV F, WAS REVIEWED. THE RISK IS CAPTURED IN ROW 17. THE BROKEN INSTRUMENT REPORTED IN EXPERIENCE C2017-192 WAS LIKELY THE RESULT OF THE DOWEL PINS BACKING OUT, ALLOWING FOR THE FOOT COMPONENT TO DETACH FROM THE HANDLE. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. NO PATIENT ADVERSE EVENT, THE SURGICAL OUTCOME WAS "GOOD". NO PIECES FROM THE DEVICE FELL IN THE WOUND SITE. AN INVESTIGATION WAS CONDUCTED: THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF THE DOWEL PINS BACKING OUT, ALLOWING FOR THE FOOT COMPONENT TO DETACH FROM THE HANDLE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10075 | 2022-04-26 | 2018-05-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | DRIVER, MODULAR CANNULATED TRI-DRIVE | EXACTECH, INC. |  | 315-25-00 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE TRI-DRIVE. THE LOCKING SYSTEM OF THE DEVICE HAS UNLOCKED DURING THE REAMING OF THE GLENOID SURFACE. AFTER PASSING THE STARTER REAMER, THE SURGEON POSITIONED HIMSELF TO USE THE REAMER SIZE 38MM. AT THE BEGINNING OF THE REAM, THE LOCKING SYSTEM OF THE DEVICE HAS UNLOCKED WITHOUT RUBBING ON THE RETRACTORS OR AN ESPECIALLY DENSE BONE. THE RESULT OF THE SURGERY WAS GOOD. PATIENT'S HEALTH WAS STABLE LEAVING THE OR. THE SECOND DEVICE PRESENT IN THE TOOLBOX WAS USED. THE BROKEN ONE HAS BEEN REPORTED TO THE LOGISTIC SERVICE. THE REP WAS PRESENT AT THE TIME OF SURGERY. NEW INFORMATION ON 1 JUL 2021: THE INSTRUMENT COULD HAVE BEEN BROKEN DURING THE SURGERY. HOWEVER; THE BREAKAGE HAS BEEN NOTICED AT THE RECOVERY OF THE INSTRUMENTATION-SET BY OUR LOGISTICS DEPT AFTER THE CLEANING REALIZED BY THE STERILIZATION TEAM OF THE HOSPITAL. NO INFORMATION HAS BEEN RECEIVED FROM THE HOSPITAL REGARDING THIS ISSUE AND THE AGENCY DIDN'T RECEIVE ANY REQUEST FOR COMPLEMENTARY INFORMATION REGARDING THIS ISSUE AS WELL. THE DEVICE WAS RECEIVED ON 3/4/2019. NO ADDITIONAL INFORMATION IS AVAILABLE. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS AVAILABLE FOR ANALYSIS. ENGINEERING ANALYSIS COMPLETED BY (B)(4). 82402015 MFG DATE: 04/03/2017. DESIGN-RELATED ISSUES: THE DESIGN OF THE TRI-DRIVE MODULAR REAMER HANDLE HAS BEEN IN THE FIELD SINCE 2009. (B)(4). THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS TRI-DRIVE MODULAR REAMER HANDLE, 750-2007-026-RMR-INSTRUMENTS REV K, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 8. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. THE BROKEN DEVICE REPORTED IN EXPERIENCE (B)(4) WAS LIKELY THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY, WHICH ALLOWED THE TIP TO DISASSEMBLE. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY, WHICH ALLOWED THE TIP TO DISASSEMBLE. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10079 | 2022-04-26 | 2017-01-24 | I, n, j, u, r, y | Manufacturer report | EXACTECH | EQUINOXE TORQUE DEFINING SCREW SQUARE DRIVE KIT | EXACTECH, INC. |  | 300-20-02 |  | IT WAS REPORTED FROM THE US THAT DURING A TSA ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE SCREW. THE HEAD OF THE SCREW WAS NOT RETAINED IN THE DRIVER. THERE WAS A SLIGHT SURGICAL DELAY AS X-RAY CONFIRMED THAT THE PART WAS NOT LEFT IN THE PATIENT. THE OUTCOME OF THE PATIENT WAS A TSA AS PLANNED. NO ADVERSE EVENT REPORTED. THE REP WAS PRESENT AT THE TIME OF SURGERY. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PATIENT DEMOGRAPHICS NOT AVAILABLE. IT WAS NOT REPORTED IF THE DEVICE FRAGMENT WAS REMOVED BY THE SURGEON. NO OTHER INFORMATION WAS PROVIDED ABOUT THE EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4) HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DESIGN-RELATED ISSUES: THE DESIGN OF THE TORQUE DEFINING SCREW KIT HAS BEEN IN THE FIELD SINCE 2005. EXACTECH HAS RECEIVED 12 SIMILAR COMPLAINT REPORTS SINCE 2008. SALES DATA FOR TORQUE DEFINING SCREW KITS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH HAS NOT RECEIVED ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT. ACCORDING TO SALES DATA, 1 PIECE OF THE 50-PIECE LOT HAS NOT BEEN SOLD. THIS IMPLIES THAT THE OTHER 48 PIECES HAVE BEEN USED TO IMPLANT THE TORQUE DEFINING SCREW SUCCESSFULLY. A SIBLING WAS NOT AVAILABLE FOR EVALUATION. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. THE DESIGN OF THE TORQUE DEFINING SCREW INCORPORATES A SMALL PIECE OF POLYETHYLENE IN THE END OF THE SCREW THAT BECOMES WEDGED IN THE DRIVER WHEN FULLY INSERTED. IF THIS POLYETHYLENE PIECE IS UNDERSIZED OR NOT FULLY ENGAGED, THE SCREW WILL NOT BE RETAINED BY THE DRIVER. THE TORQUE DEFINING SCREW KIT ISSUE REPORTED IN EXPERIENCE (B)(6) WAS LIKELY THE RESULT OF EITHER THE SCREW NOT BEING FULLY INSERTED AND SECURED TO THE DRIVER OR THE POLYETHYLENE PLUG BEING UNDERSIZED, WHICH ALLOWED THE SCREW HEAD TO SLIP OUT OF THE DRIVER WITHOUT INTERFERENCE. UNFORTUNATELY, NEITHER THE KIT USED, NOR A SIBLING OF THE KIT WERE AVAILABLE FOR EVALUATION. THEREFORE, THE EXACT CAUSE OF THE REPORTED EVENT CANNOT BE DETERMINED. EXACTECH IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED INSTRUMENT MALFUNCTION IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE WAS NO PATIENT ADVERSE EVENT. THE HEAD OF THE SCREW WAS NOT RETAINED IN THE DRIVER. X-RAY CONFIRMED THAT THE PART WAS NOT LEFT IN THE PATIENT. AN INVESTIGATION WAS CONDUCTED: THE TORQUE DEFINING SCREW KIT ISSUE REPORTED WAS LIKELY THE RESULT OF EITHER THE SCREW NOT BEING FULLY INSERTED AND SECURED TO THE DRIVER OR THE POLYETHYLENE PLUG BEING UNDERSIZED, WHICH ALLOWED THE SCREW HEAD TO SLIP OUT OF THE DRIVER WITHOUT INTERFERENCE. UNFORTUNATELY, NEITHER THE KIT USED, NOR A SIBLING OF THE KIT WERE AVAILABLE FOR EVALUATION. THEREFORE, THE EXACT CAUSE OF THE REPORTED EVENT CANNOT BE DETERMINED. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10094 | 2022-04-26 | 2016-11-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | HUMERAL HEAD TRIAL | EXACTECH, INC. |  |  |  | IT WAS REPORTED FROM THE US THAT DURING AN INITIAL ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE INSTRUMENTATION. THE SURGEON HAD DIFFICULTY SEATING SIZE 13 HUMERAL STEM (CAPTURED IN ANOTHER COMPLAINT). THE SURGEON BROKE 2 HUMERAL HEAD TRIALS TRYING TO SEAT STEM. THE SURGEON ENDED UP IMPLANTING A SIZE 9 STEM. THERE WAS MINIMUM DELAY, AND THE PATIENT WAS NOT ADVERSELY AFFECTED. THE PATIENT WAS DOING WELL AT THE END OF SURGERY. THE REPRESENTATIVE WAS UNABLE TO PROVIDE ANY CATALOG AND LOT/SERIAL NUMBERS OF THE HUMERAL HEAD TRIALS. THE DEVICES WERE NOT RETURNED FOR ANALYSIS. ATTEMPTS WERE MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REPRESENTATIVE, BUT NO RESPONSES WERE RECEIVED. THERE HAS BEEN NO OTHER INFORMATION PROVIDED BY THE CONTACTS TO THIS EVENT. THIS IS 1 REPORT OF 2 FOR THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DESIGN RELATED ISSUES: THE PART NUMBERS FOR THE BROKEN INSTRUMENTS WERE NOT PROVIDED. HOWEVER, THE DESIGN OF THE HUMERAL HEAD TRIAL (ALL SIZES) HAS BEEN IN THE FIELD SINCE 2004. EXACTECH HAS NOT RECEIVED ANY OTHER COMPLAINT REPORTS INVOLVING BREAKAGE OF A HUMERAL HEAD TRIAL. THE SALES DATA FOR ALL HUMERAL HEADS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS DOES NOT APPEAR TO BE A DESIGN RELATED ISSUE. MFG-RELATED ISSUES: MFG-RELATED ISSUES: THE MANUFACTURING LOT NUMBER WAS NOT PROVIDED. THEREFORE, THE MANUFACTURING DATA CANNOT BE REVIEWED. THIS ISSUE IS UNLIKELY RELATED TO THE MANUFACTURING. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THESE HUMERAL HEAD TRIALS, 750-2002-039-RMR­ INSTRUMENT ADDITIONS REV F, WAS REVIEWED. THE RISK IS CAPTURED. MOST LIKELY CAUSE: THE BROKEN HUMERAL HEAD TRIALS REPORTED WERE LIKELY THE RESULT OF EXCESSIVE IMPACTION DUE TO DIFFICULTY SEATING THE HUMERAL STEM. THIS CANNOT BE CONFIRMED BECAUSE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES. NO PATIENT ADVERSE EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-10087 | 2022-04-26 | 2017-02-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | DRIVER MODULAR CANNULATED TRI-DRIVE | EXACTECH, INC. |  | 315-25-00 |  | IT WAS REPORTED FROM (B)(6) THAT DURING SURGERY THE INSTRUMENT BROKE DURING SURGERY. NO ADDITIONAL INFORMATION REGARDING THE NATURE OF THE DEVICE BREAKAGE WAS AVAILABLE. THERE WAS NO SURGICAL DELAY AND NO PATIENT INJURY WAS REPORTED. NO PATIENT OR SURGEON INFORMATION WAS PROVIDED DUE TO CONFIDENTIALITY. THE REPRESENTATIVE WAS NOT PRESENT DURING THE SURGERY. ALL BROKEN PIECES OF THE DEVICE WERE ACCOUNTED FOR. AFTER ADDITIONAL INVESTIGATION THE AGENT REPORTED, THE INSTRUMENT COULD HAVE BEEN BROKEN DURING THE SURGERY. HOWEVER, THE BREAKAGE HAS BEEN NOTICED AT THE RECOVERY OF THE INSTRUMENTATION-SET BY OUR LOGISTICS DEPT., AFTER THE CLEANING REALIZED BY THE STERILIZATION TEAM OF THE HOSPITAL. NO INFORMATION HAS BEEN RECEIVED FROM THE HOSPITAL REGARDING THIS ISSUE AND THE AGENCY DIDN'T RECEIVE ANY REQUEST FOR COMPLEMENTARY INFORMATION REGARDING THIS ISSUE AS WELL. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN IFU (B)(6) 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED FOR ANALYSIS. THE DEVICE WAS RETURNED IN FOUR PIECES. THE RETAINING RING IS MEANT TO HOLD THE SLEEVE-SPRING MECHANISM IN PLACE BUT APPEARS TO HAVE DISLODGED. THE BALL BEARINGS WERE NOT RETURNED WITH THE DEVICE. ALL OTHER ASPECTS OF THE DEVICE APPEAR NORMAL. THIS DEVICE LOT WAS MANUFACTURED IN 2016. THE DEVICE HISTORY RECORD WAS REVIEWED, AND ALL PARTS WERE ACCEPTED WITH CONFORMANCE TO THE DEVICE SPECIFICATIONS. THERE IS NO EVIDENCE TO SUGGEST A DESIGNOR MANUFACTURING ISSUE RELATED TO THE DEVICE AND THERE ARE NO KNOWN PATIENT FACTORS THAT MAY HAVE CONTRIBUTED. THE MOST LIKELY CAUSE OF THE REPORTED EVENT A USER-RELATED ISSUE, SPECIFICALLY HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY AND ALLOWED THE TIP TO DISASSEMBLE. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICALINTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLEWITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION,THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT IMPACT. AN INVESTIGATION WAS CONDUCTED; THE MOST LIKELY CAUSE OF THE REPORTED EVENT A USER-RELATED ISSUE, SPECIFICALLY HOLDING THE SLEEVE STATIONARY WHILE THE INSTRUMENT WAS SPINNING, WHICH LED TO THE RETAINING RING BREAKING LOOSE FROM THE ASSEMBLY AND ALLOWED THE TIP TO DISASSEMBLE. INFORMATION WAS NOT PROVIDED. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00475 | 2022-04-27 | 2022-04-07 | I, n, j, u, r, y | Manufacturer report | TRULIANT PS CEM FEM PS CEM RIGHT SZ 6 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT PS CEM FEM PS CEM RIGHT SZ 6 | 02-020-11-0360 |  | AS REPORTED, APPROXIMATELY ONE YEAR POST INITIAL RIGHT TKA, THIS (B)(6) MALE PATIENT HAD REVISION DUE TO LOOSE FEMUR AND COMPLAINED OF PAIN. PATIENT WAS REVISED WITH COMPETITOR'S IMPLANTS. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. SALES REP WAS UNABLE TO OBTAIN PHOTOS/X-RAYS. THE DEVICE WILL BE RETURN. | CONCOMITANT MEDICAL PRODUCTS TRULIANT TIB IMP PS INSERT SZ 6 9MM (CAT# 02-022-35-6009 / SERIAL(B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00476 | 2022-04-27 | 2018-07-18 | I, n, j, u, r, y | Manufacturer report | LOGIC TIB INSERT IMPL CRC, SZ 2, 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIB INSERT IMPL CRC, SZ 2, 13MM | 02-010-03-0220 |  | AS REPORTED BY LEGAL TEAM, APPROXIMATELY FIVE YEARS POST INITIAL LEFT TKA, THIS FEMALE PATIENT HAD A REVISION DUE TO IMPLANT BROKE AND CAUSED HER FEMUR TO BREAK AS WELL. THEN, TWO YEARS POST LEFT TKA REVISION, PATIENT RETURNED TO SURGEON BECAUSE HER KNEE DIDN¿T FEEL RIGHT, AND SHE HAD SEVERE PAIN, PROBLEMS WITH HER MOBILITY, AND HAD TO BE IN A WHEELCHAIR SINCE 2013. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00478 | 2022-04-27 | 2022-04-07 | I, n, j, u, r, y | Manufacturer report | NV GXL LNR, +5LAT, 32MM G1-48/50MM CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LNR, +5LAT, 32MM G1-48/50MM CUPS | 136-32-51 |  | AS REPORTED, APPROXIMATELY SEVEN YEARS POST INITIAL LEFT THA, THIS (B)(6) FEMALE PATIENT WAS BROUGHT BACK TO SURGERY FOR PAINFUL LEFT HIP. THE HIP WAS DISLOCATED, FEMORAL HEAD REMOVED, STEM CHECKED AND FOUND TO BE WELL FIXED. ATTENTION TURNED TO THE CUP. THE LINER WAS REMOVED WITH THE HELP OF AN OSTEOTOME. THE CUP WAS CHECKED AND FOUND TO BE WELL FIXED. A LARGE CYST WAS CLEANED OUT POSTERIOR SUPERIOR OF THE CUP. BONE GRAFT WAS INJECTED BEHIND TO FILL THE VOID. A NEW XLE 10 DEGREE LINER WAS IMPLANTED, AS WELL AS +7 BIOLOX HEAD. THERE WAS NO BREAKAGE OF DEVICES OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. X-RAY RECEIVED. THE DEVICE WILL BE RETURN. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00479 | 2022-04-27 | 2022-04-07 | I, n, j, u, r, y | Manufacturer report | BIOLOX DELTA FEMORAL HEAD 36MM OD, +3.5MM | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | BIOLOX DELTA FEMORAL HEAD 36MM OD, +3.5MM | 170-36-03 |  | AS REPORTED, APPROXIMATELY EIGHTEEN DAYS POST INITIAL LEFT THA, THIS FEMALE PATIENT WAS BROUGHT BACK FOR LEFT HIP TRAUMA, LESSER TROCHANTERIC FRACTURE, THE HIP WAS DISLOCATED, FEMORAL HEAD REMOVED AND STEM REMOVED. THE FRACTURE WAS REDUCED USING CABLE SYSTEM, A SIZE 6 HPS STEM WAS CEMENTED WITH A 36+3.5 BIOLOX HEAD. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. X-RAY RECEIVED. THE DEVICES WILL BE RETURN. | CONCOMITANT MEDICAL PRODUCTS: WEDGE PLASMA S/O SZ 5 (CAT# 188-00-05 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00500 | 2022-04-28 | 2019-12-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LOCKING FEMORAL IMPACTOR 0-5 | EXACTECH, INC. |  | 213-64-01 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE IMPACTOR BROKE. WHEN IMPACTING THE FINAL FEMORAL COMPONENT, THE BACK PLASTIC PIECE CHIPPED OFF FROM THE IMPACTOR. THE INSTRUMENT CAN NO LONGER HOLD A FEMUR. THIS DID NOT LEAD TO INJURY OF A PATIENT OR USER, NOR WERE ANY PARTS OF A BROKEN DEVICE/DEVICE FRAGMENTS LEFT IN THE PATIENT. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. ACCORDING TO THE INFORMATION PROVIDED, THE REPORTED DEVICE IS EXPECTED TO RETURN. THE DEVICE WAS NOT AVAILABLE AT THE TIME OF THIS INVESTIGATION; HOWEVER, AN IMAGE WAS PROVIDED THAT CONFIRMED THE FAILURE REPORTED. THEREFORE, THE PHYSICAL DEVICE IS NOT NEEDED TO PERFORM THE ENGINEERING ANALYSIS ON THIS COMPLAINT. IF THE DEVICE WERE TO RETURN, THE COMPLAINT FILE WILL BE UPDATED TO LOG THE DATE THE DEVICE HAS BEEN RETURNED, AND THE PART WILL AWAIT ASSESSMENT FOR POSSIBLE REWORK/REPAIR/SCRAP. THE INVESTIGATION TASK WILL NOT NEED TO BE REOPENED. PREVIOUS INVESTIGATION COMPLAINT NUMBERS: PREVIOUS INVESTIGATION. FINDINGS: THE REPORTED FRACTURED IMPACTOR FACE APPEARS CONSISTENT WITH MATERIAL DEGRADATION DUE TO NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. RMR REVIEW: THE RMR AND RACR FOR THIS INSTRUMENT, RMR-00039 REV ¿ AND RACR-00051 REV -, WERE REVIEWED. THIS RISK IS CAPTURED. HOWEVER, THE RACR IS BEING UPDATED UNDER DCRSP-19-527 TO CAPTURE AN APPROPRIATE HARM FOR THE HAZARDOUS SITUATION. MOST LIKELY UNDERLYING CAUSE: BASED ON SIMILAR PREVIOUS COMPLAINT INVESTIGATIONS, THE FRACTURED IMPACTOR FACE REPORTED WAS LIKELY THE RESULT OF MATERIAL DEGRADATION ASSOCIATED WITH YEARS OF NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. IFU 700-096-181: INSTRUMENT INSPECTION, VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT FROM THE BROKEN DEVICE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00489 | 2022-04-28 | 2022-03-15 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 4 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 4 11MM | 02-012-35-4011 |  | AS REPORTED, THIS (B)(6) MALE PATIENT HAD A REVISION DUE TO WEAR OF SIZE 4 PS LOGIC INSERT. REVISED INSERT BECAME DISLODGED FROM THE TIBIA TRAY WHEN THE KNEE WAS PUT IN FLEXION. PATIENT WAS REVISED TO TRULIANT SIZE 4 PSC 13MM INSERT. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. IMAGES AND X-RAYS RECEIVED. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO HSS KEEPS ALL RETRIEVALS. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00497 | 2022-04-28 | 2019-02-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | DRILL PINS | EXACTECH, INC. |  |  |  | IT WAS REPORTED FROM OUS THAT A SURGEON SNAPPED A DRILL IN DISTAL FEMORAL NAVIGATION CUTTING BLOCK, TWICE. ONE SNAPPED IN ONE BLOCK, THEN ANOTHER SNAPPED IN THE NEXT BLOCK. THE PATIENT HAD VERY HARD BONE. AS THE SURGEON WAS DRIVING YOU COULD SEE THE PINS FLEX WITH THE AMOUNT OF FORCE REQUIRED TO BREAK THE CORTICAL BONE. ALL PIECES WERE RETRIEVED. NO PATIENT IMPACT WAS REPORTED. PRODUCT RETURN CONTAINED THE 2 FEMORAL CUTTING BLOCKS ONLY, NOT THE DRILL PINS. THE DRILL PINS WERE CONFIRMED TO BE EXACTECH DEVICES. THERE WERE NO PARTS, OR PIECES THAT FELL INTO THE SURGICAL SITE. THE PATIENT WAS NOTED TO BE FINE, NO COMPLICATIONS AT DISCHARGE AND FOLLOW UP WAS ALL ¿NORMAL¿. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. THIS IS 1 OF 2 MDRS FOR THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: NO CATALOG OR SERIAL/LOT NUMBER CONCOMITANT MEDICAL PRODUCTS: CAT #521-11-01; SERIAL/LOT#(B)(4); KNEE GPS CUTTING BLOCK / CAT#521-40-12. SERIAL/LOT# (B)(4); FEMORAL GPS DISTAL BLOCK CAPTURED LOW PROFILE. DESIGN-RELATED ISSUES: (B)(4). THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE, THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN-RELATED. MFG-RELATED ISSUES: MANUFACTURING-RELATED DATA COULD NOT BE REVIEWED BECAUSE NO INFORMATION REGARDING THE MANUFACTURING LOT NUMBERS WAS PROVIDED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) COULD NOT BE CONDUCTED BECAUSE ADEQUATE INFORMATION REGARDING THE CATALOG NUMBERS OF THE DEVICES WAS NOT PROVIDED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿. MOST LIKELY CAUSE: THE CAUSE OF THE BROKEN DRILL PINS REPORTED COULD NOT BE DETERMINED BECAUSE THE DEVICES WERE NOT AVAILABLE FOR EVALUATION, AND NO FURTHER INFORMATION WAS PROVIDED REGARDING THE INCIDENT. IFU 700-096-181: INSTRUMENT INSPECTION VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY DESIGN ISSUES. THIS EVENT DID NOT LEAD TO PATIENT ADVERSE EVENT, NO PIECES OF THE PINS FELL INTO THE SURGICAL SITE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00498 | 2022-04-28 | 2019-02-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | DRILL PINS | EXACTECH, INC. |  |  |  | IT WAS REPORTED FROM OUS THAT A SURGEON SNAPPED A DRILL IN DISTAL FEMORAL NAVIGATION CUTTING BLOCK, TWICE. ONE SNAPPED IN ONE BLOCK, THEN ANOTHER SNAPPED IN THE NEXT BLOCK. THE PATIENT HAD VERY HARD BONE. AS THE SURGEON WAS DRIVING YOU COULD SEE THE PINS FLEX WITH THE AMOUNT OF FORCE REQUIRED TO BREAK THE CORTICAL BONE. ALL PIECES WERE RETRIEVED. NO PATIENT IMPACT WAS REPORTED. PRODUCT RETURN CONTAINED THE 2 FEMORAL CUTTING BLOCKS ONLY BUT NOT THE DRILL PINS. THE DRILL PINS WERE CONFIRMED TO BE EXACTECH DEVICES. THERE WERE NO PARTS, OR PIECES THAT FELL INTO THE SURGICAL SITE. THE PATIENT WAS NOTED TO BE FINE, NO COMPLICATIONS AT DISCHARGE AND FOLLOW UP WAS ALL ¿NORMAL¿. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. THIS IS 2 OF 2 MDRS FOR THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4) HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. H.10 INFORMATION WAS NOT PROVIDED FOR: D4- NO CATALOG OR SERIAL/LOT NUMBER CAT #521-11-01; SERIAL/LOT#37190009; KNEE GPS CUTTING BLOCK / CAT#521-40-12 ; SERIAL/LOT# 63552008; FEMORAL GPS DISTAL BLOCK CAPTURED LOW PROFILE DESIGN-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER SIMILAR COMPLAINT REPORTS INVOLVING GPS FEMORAL CUT BLOCKS AND BROKEN DRILL BITS SINCE 2009. BECAUSE THE POSSIBLE FEMORAL BLOCK AND POSSIBLE DRILL PINS ARE USED IN APPROXIMATELY 30% OF EXACTECH KNEE GPS PROCEDURES, DATA FROM THE GPS WEB SERVICE WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE, THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN-RELATED. MFG-RELATED ISSUES: MANUFACTURING-RELATED DATA COULD NOT BE REVIEWED BECAUSE NO INFORMATION REGARDING THE MANUFACTURING LOT NUMBERS WAS PROVIDED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) COULD NOT BE CONDUCTED BECAUSE ADEQUATE INFORMATION REGARDING THE CATALOG NUMBERS OF THE DEVICES WAS NOT PROVIDED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿. MOST LIKELY CAUSE: THE CAUSE OF THE BROKEN DRILL PINS REPORTED COULD NOT BE DETERMINED BECAUSE THE DEVICES WERE NOT AVAILABLE FOR EVALUATION, AND NO FURTHER INFORMATION WAS PROVIDED REGARDING THE INCIDENT. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. ¿ CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY DESIGN ISSUES. THIS EVENT DID NOT LEAD TO PATIENT ADVERSE EVENT, NO PIECES OF THE PINS FELL INTO THE SURGICAL SITE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00493 | 2022-04-28 | 2021-01-01 | I, n, j, u, r, y | Manufacturer report | PS TIBIAL INSERTS SZ 4, 18MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | PS TIBIAL INSERTS SZ 4, 18MM | 204-24-18 |  | AS REPORTED BY LEGAL TEAM, PATIENT ORIGINAL SURGERY WAS PERFORMED BY DR. (B)(6) AT (B)(6) HOSPITAL IN (B)(6) IN 2014. PATIENT MOVED FROM (B)(6) TO (B)(6) AFTER THE PRIMARY SURGERY. PATIENT HAS HAD TWO (2) REVISIONS IN (B)(6) AT (B)(6) MEDICAL CENTER WITH A DR. (B)(6) SAYS HAD PAIN AND ISSUES FOLLOWING THE PRIMARY SURGERY. WHEN PATIENT HAD FIRST REVISION IN 2017, THE SURGEON CORRECTED A ¿PATELLAR ISSUE¿ (THE PATELLAR BUTTON WAS LOOSE HE THOUGHT), AND ¿TIGHTENED THE KNEE¿. HOWEVER, ONLY THE INSERT WAS EXCHANGED. THE PATIENT CONTINUING PAIN AND MOBILITY ISSUES, THEN IN (B)(6) 2021 HAD A SECOND REVISION. PATIENT BELIEVES ALL THE COMPONENTS WERE REMOVED IN 2021, AND SURGEON SAID HE HAD SIGNIFICANT LYSIS. PATIENT HAS HAD TO WORK FROM HOME SINCE THE SURGERY, HAS LEG LENGTH ISSUES, CONTINUING PAIN, ETC. PATIENT IS MALE, (B)(6) AND HAS VERY LIMITED MOBILITY. | CONCOMITANT PRODUCTS: TRAPEZOID TIBIAL TRAY SZ 4F/5T (CAT# 204-04-45 / SERIAL# (B)(4)) THREE PEG PATELLA 41MM (CAT# 200-02-41 / SERIAL# (B)(4)) OPTETRAK ASY,PS CEMENTED FEMORAL, SZ 4, (CAT# 234-03-04 / SERIAL# (B)(4)) ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00492 | 2022-04-28 | 2017-12-12 | I, n, j, u, r, y | Manufacturer report | PS TIBIAL INSERTS SZ 4, 15MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | PS TIBIAL INSERTS SZ 4, 15MM | 204-24-15 |  | AS REPORTED BY LEGAL TEAM, PATIENT ORIGINAL SURGERY WAS PERFORMED IN 2014. PATIENT MOVED TO A DIFFERENT CITY AFTER THE PRIMARY SURGERY. PATIENT HAS HAD TWO (2) REVISIONS, AND PATIENT HAD PAIN AND ISSUES FOLLOWING THE PRIMARY SURGERY. WHEN PATIENT HAD FIRST REVISION IN 2017, THE SURGEON CORRECTED A ¿PATELLAR ISSUE¿ (THE PATELLAR BUTTON WAS LOOSE HE THOUGHT), AND ¿TIGHTENED THE KNEE¿. HOWEVER, ONLY THE INSERT WAS EXCHANGED. THE PATIENT CONTINUING PAIN AND MOBILITY ISSUES, THEN IN (B)(6) 2021 HAD A SECOND REVISION. PATIENT BELIEVES ALL THE COMPONENTS WERE REMOVED IN 2021, AND SURGEON SAID HE HAD SIGNIFICANT LYSIS. PATIENT HAS HAD TO WORK FROM HOME SINCE THE SURGERY, HAS LEG LENGTH ISSUES, CONTINUING PAIN, ETC. PATIENT IS MALE, (B)(6) AND HAS VERY LIMITED MOBILITY. | "D10: 200-02-41 - THREE PEG PATELLA 41MM. 204-04-45 - TRAPEZOID TIBIAL TRAY SZ 4F/5T. 234-03-04 - OPTETRAK ASY, PS CEMENTED FEMORAL, SZ 4. MULTIPLE MDR REPORTS WERE FILED FOR THIS EVENT, PLEASE SEE ASSOCIATED REPORTS: 1038671-2025-00671. THIS FOLLOW-UP REPORT IS BEING SUBMITTED TO RELAY ADDITIONAL AND/OR CORRECTED INFORMATION. THE REASON FOR THE REVISION REPORTED CANNOT BE CONFIRMED FROM THE INFORMATION PROVIDED BUT MAY BE THE RESULT OF PROSTHESIS WEAR, INSTABILITY AND PATELLAR LOOSENING OR DUE TO INCLUSION OF THE POLYETHYLENE IN THE PACKAGING RECALL. POTENTIAL CONTRIBUTIONS OF USER AND PATIENT-RELATED CONSIDERATIONS TO THE EVENT COULD NOT BE ASSESSED AS THE DEVICES WERE NOT AVAILABLE FOR EVALUATION AND IMAGES, RADIOGRAPHS, AND RELEVANT CLINICAL INFORMATION WERE NOT PROVIDED. THE PRODUCT ASSOCIATED WITH THE REPORTED EVENT IS WITHIN THE SCOPE OF RECALL Z-0019-2022; HOWEVER, THERE IS INSUFFICIENT INFORMATION TO EVALUATE WHETHER THE SUBJECT ISSUE OF THE RECALL WAS A CAUSE OR CONTRIBUTOR TO THE REPORTED EVENT. IF ANY FURTHER INFORMATION IS OBTAINED THAT WOULD CHANGE OR ALTER ANY INFORMATION PROVIDED, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY." | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00490 | 2022-04-28 | 2022-04-07 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 3 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 3 11MM | 02-012-35-3011 |  | AS REPORTED BY LEGAL TEAM, APPROXIMATELY NINE YEARS POST INITIAL TKA, THIS MALE PATIENT HAD A REVISION FOR AN UNKNOWN REASON. NO ADDITIONAL INFORMATION AVAILABLE. | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION RECEIVED THE FOLLOWING SECTIONS G3, G6, H2, H3, H6 AND H9 HAVE BEEN UPDATED ACCORDINGLY. (H3) AS REPORTED BY LEGAL TEAM, APPROXIMATELY NINE YEARS POST INITIAL TKA, THIS MALE PATIENT HAD A REVISION FOR AN UNKNOWN REASON. NO ADDITIONAL INFORMATION AVAILABLE. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN ISSUES. THE CAUSE OF THE SUBSEQUENT REVISION CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION. THESE DEVICES ARE USED FOR TREATMENT NOT DIAGNOSIS. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00483 | 2022-04-28 | 2022-04-04 | I, n, j, u, r, y | Manufacturer report | NV GXL LINER NEUTRAL, 36MM ID, GROUP 2 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LINER NEUTRAL, 36MM ID, GROUP 2 CUPS | 130-36-52 |  | AS REPORTED, APPROXIMATELY SIX YEARS POST INITIAL THA, THIS MALE PATIENT HAD A REVISION DUE TO WEAR OF NOVATION OF G2 36MM GXL LINER. PATIENT WAS REVISED TO NOVATION G2 36MM GXE INSERT. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PHOTOS RECEIVED. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO HSS KEEPS ALL RETRIEVALS. | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION. (H3) THE PATIENT INVOLVED WAS REPORTED FOR REVISION DUE TO EARLY PROSTHESIS WEAR APPROXIMATELY 6 YEARS AFTER THE INDEX SURGERY. THE PROSTHESIS WEAR WAS ABLE TO BE CONFIRMED FROM THE PROVIDED RADIOGRAPH AND PHOTO OF THE EXPLANTED DEVICES; HOWEVER, THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. BASED ON THE AVAILABLE INFORMATION, THE PATIENT INVOLVED MEETS THE FOLLOWING RISK CRITERIA FOR EARLY PROSTHESIS WEAR/OSTEOLYSIS COMBINATION OF LARGEST AVAILABLE FEMORAL HEAD AND/OR THINNEST AVAILABLE ACETABULAR LINER WAS USED. THE MOST LIKELY CAUSE FOR THE REVISION REPORTED DUE TO EARLY PROSTHESIS WEAR IS A COMBINATION OF THE RISK FACTORS INCLUDING, USE ERROR, IMPLANT POSITIONING, IMPLANT SIZE SELECTION, AND PATIENT FACTORS (FITNESS FOR SURGERY, BIOMECHANICS, ACTIVITY LEVEL AND LOCAL TISSUE OXIDATION POTENTIAL COULD HAVE ALL CONTRIBUTED TO THE INCREASED TREND IN WEAR/OSTEOLYSIS RELATED COMPLAINTS). A RISK ASSESSMENT HAS BEEN INITIATED TO ESCALATE THIS ISSUE. | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00486 | 2022-04-28 | 2018-10-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | DRILL GUIDE | EXACTECH, INC. |  | 181-00-05 |  | IT WAS REPORTED FROM (B)(6) THAT THE DRILL GUIDE SLEEVE FELL OFF DURING THE ORTHOPEDIC SURGERY. ¿NOT ENOUGH STRENGTH (WELDING) AGAINST THE STRESS OF DRILLING.¿ THE TIP WAS BROKEN OFF THE SHAFT. THE PATIENT¿S HEALTH WAS STABLE LEAVING THE OPERATING ROOM. THERE WAS NO DELAY TO SURGERY. THE REPRESENTATIVE WAS NOT PRESENT AT TIME OF SURGERY. PATIENT DEMOGRAPHIC INFORMATION WAS NOT PROVIDED BY THE CONTACT. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR L0X OPTICAL UNLESS OTHERWISE SPECIFIED)- THE SHEARED SURFACES APPEAR TO BE CONSISTENT WITH APPLYING A BENDING MOMENT TO THE DRILL GUIDE WHILE USING A DRILL, EVENTUALLY BREAKING THE WELD BETWEEN THE DRILL GUIDE TIP AND SHAFT. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE. DESIGN-RELATED ISSUES: THE DESIGN OF THE DRILL GUIDE HAS BEEN IN THE FIELD SINCE 2013. EXACTECH HAS RECEIVED 6 SIMILAR COMPLAINT REPORTS INVOLVING BROKEN DRILL GUIDES SINCE 2013. THE FREQUENCY OF OCCURRENCE RANKING SCALE IS VERY LOW; THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF (B)(4) UNITS, WHICH HAS BEEN IN THE FIELD SINCE 2017. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE DRILL GUIDE WHILE USING A DRILL, EVENTUALLY BREAKING THE WELD BETWEEN THE DRILL GUIDE TIP AND THE SHAFT. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT IMPACT. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00485 | 2022-04-28 | 2018-10-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | M-SERIES 12/14 FEMORAL HEAD IMPACTOR | EXACTECH, INC. |  | 153-00-02 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE IMPACTOR HEAD BROKE. UPON IMPACTING FEMORAL HEAD, THE WHITE IMPACTOR HEAD SPLIT. THE WHITE PART WAS THROWN AWAY BY HOSPITAL STAFF. ALL PIECES OF THE IMPACTOR HEAD WERE ACCOUNTED FOR, NOTHING FELL INTO THE SURGICAL SITE. THERE WAS NO DELAY TO THE CASE AND NOT ADVERSE EFFECTS TO THE PATIENT. THE AGENT WAS PRESENT AT THE TIME OF SURGERY. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR 10X OPTICAL UNLESS OTHERWISE SPECIFIED) MISSING IMPACTION TIP. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE AND CONSISTENT WITH YEARS OF USE. DESIGN-RELATED ISSUES: THE DESIGN OF THE M-SERIES 12/14 FEMORAL HEAD IMPACTOR HAS BEEN IN THE FIELD SINCE 2000. EXACTECH IS AWARE OF (B)(4) SIMILAR COMPLAINT REPORTS INVOLVING THIS DEVICE SINCE 2008. THE FREQUENCY OF OCCURRENCE RANKING SCALE IS ¿VERY LOW¿. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT, WHICH HAVE BEEN IN THE FIELD SINCE 2013. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING-RELATED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RISK MANAGEMENT REPORT. THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF YEARS OF NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE RADEL TIP. HOWEVER, THIS CANNOT BE CONFIRMED AS THE FRACTURED IMPACTOR TIP WAS NOT AVAILABLE FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THE BROKEN PIECES DID NOT FALL INTO THE SURGICAL SITE. THERE WAS NO PATIENT ADVERSE EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00484 | 2022-04-28 | 2018-10-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | MDLR BROACH HNDL | EXACTECH, INC. |  | 301-03-01 |  | IT WAS REPORTED FROM THE OUS THAT DURING AN ORTHOPEDIC SURGERY THE INSTRUMENT BROKE. WHEN THE SURGEON WAS ¿RAISING THE SIZE IN TURN, TARGETING #15, RASP HANDLE BROKE AROUND #11¿. THE TIP FELL OFF, BUT RASPING WAS CONTINUOUSLY MANAGED BY THE SURGEON, AND FINALLY TARGETED RASPING #15 WAS SUCCESSFULLY FINISHED. NO PIECES WERE LEFT IN THE PATIENT. THE PATIENT WAS STABLE LEAVING THE O.R. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: AGE, SEX, WEIGHT, ETHNICITY, RACE. VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR L0X OPTICAL UNLESS OTHERWISE SPECIFIED) THE DETACHED FOOT APPEARS TO HAVE DISASSOCIATED FROM THE HANDLE ASSEMBLY. THE FOOT IS ATTACHED TO THE REST OF THE DEVICE BY TWO 316 STAINLESS STEEL DOWEL PINS. THE PINS APPEAR TO HAVE BACKED OUT, ALLOWING THE FOOT TO DETACH FROM THE ASSEMBLY. THE SALES REPRESENTATIVE CONFIRMED THAT NO PIECES WERE LEFT IN THE PATIENT. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE. DESIGN-RELATED ISSUES: THE DESIGN OF THE EQUINOXE SHOULDER MODULAR BROACH HANDLE HAS BEEN IN THE FIELD SINCE 2004. EXACTECH IS AWARE OF 4 OTHER SIMILAR COMPLAINT REPORTS INVOLVING THE EQUINOXE MODULAR BROACH HANDLE SINCE 2008. BECAUSE THE EQUINOXE SHOULDER MODULAR BROACH HANDLE IS USED IN BOTH PRIMARY AND REVERSE SHOULDER SURGERIES, SALES DATA FOR HUMERAL STEMS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 60 UNITS, WHICH HAS BEEN IN THE FIELD SINCE 2017. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS " VERY LOW", AND THE RISK IS CAPTURED IN THE RISK MANAGEMENT REPORT. THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF THE DOWEL PINS BACKING OUT, ALLOWING THE FOOT COMPONENT TO DETACH FROM THE HANDLE. IFU 700-096-181: INSTRUMENT INSPECTION. VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT IMPACT. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00482 | 2022-04-28 | 2022-04-01 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 5 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 5 9MM | 02-012-35-5009 |  | AS REPORTED, APPROXIMATELY SEVEN YEARS POST INITIAL BILATERAL TKA, THIS MALE PATIENT HAD A REVISION DUE TO WEAR OF PS AND PSC INSERTS. PATIENT HAS BILATERAL LOGIC KNEES. THERE WAS SOME LYSIS SEEN MEDIALLY ON THE RIGHT TIBIA DEFECT FILLED WITH A BIOLOGIC. COMPONENTS OF BOTH KNEES WERE WELL FIXED EXCEPT THE 35MM PATELLA ON THE RIGHT KNEE. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PHOTOS RECEIVED. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO HSS KEEPS ALL RETRIEVALS. | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION RECEIVED THE FOLLOWING SECTIONS G3, G6, H2, H3, H6, H7 AND H9 HAVE BEEN UPDATED ACCORDINGLY. (H3) THE REVISION REPORTED MAY HAVE BEEN THE RESULT OF MALALIGNMENT BETWEEN THE FEMORAL AND TIBIAL COMPONENTS, THIRD BODY WEAR, PATIENT-RELATED CONDITIONS, OR ANY COMBINATION OF THESE POSSIBILITIES. AN ADDITIONAL CONTRIBUTING FACTOR TO THE REVISION MAY HAVE BEEN INCLUSION OF THE IMPLANTED POLYETHYLENE COMPONENT IN THE PACKAGING RECALL. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION AND PRE-REVISION X-RAYS WERE NOT PROVIDED. THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE REPORTED EVENT OF "PROSTHESIS WEAR¿ IS ASSOCIATED WITH MATERIAL DAMAGE TO A SURFACE, USUALLY INVOLVING PROGRESSIVE LOSS OR DISPLACEMENT OF MATERIAL, DUE TO RELATIVE MOTION BETWEEN THAT SURFACE AND A CONTACTING SUBSTANCE OR SUBSTANCES. | Y | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00480 | 2022-04-28 | 2022-03-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EQUINOXE | MDLR CNLTD TRIDRV DRVR | EXACTECH, INC. | 315-25-00 | 315-25-00 |  | IT WAS REPORTED THAT DURING SURGICAL USE, THE REAMER CAME APART, SPRING CAME LOOSE. THE TRI-DRIVE CAME LOOSE DURING REAMING SO SURGEON STOPPED USING IT AND USED ANOTHER ONE. NO PARTS OR PIECES FELL INTO THE (B)(6) PATIENT'S WOUND SITE. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. | SECTION H10: (H3) THE BROKEN TRI-DRIVER REPORTED WAS LIKELY THE RESULT OF THE RETAINING RING RESPONSIBLE FOR MAINTAINING ASSEMBLY OF THE SLEEVE ON THE TRI-DRIVER SHAFT BECOMING DEFORMED AS A RESULT OF THE SLEEVE BEING HELD STATIONARY WHILE REAMING OR THE SLEEVE BEING OBSTRUCTED BY SOFT TISSUE OR BONE, WHICH LED TO THE DISASSEMBLY OF THE SPRING WHEN ATTEMPTING TO LOAD A DRILL ON THE DEVICE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00481 | 2022-04-28 | 2022-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EQUINOXE | HUM LNR IMPACT TIP 42M | EXACTECH, INC. | 321-07-42 | 321-07-42 |  | IT WAS REPORTED THAT SURGEON IMPACTED HUMERAL LINER, DEVICE DEVELOPED A CRACK. PART DIDN'T ACTUALLY BREAK, RATHER CRACKED, SO NO PIECES CAME LOOSE. NO PARTS FELL INTO THE(B)(6) PATIENT'S WOUND. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. | SECTION H10: (H3) BASED ON HISTORICAL COMPLAINT INVESTIGATIONS AND THE RETURNED DEVICE, THE CRACKED HUMERAL LINER IMPACTOR TIP REPORTED WAS LIKELY THE RESULT OF A STRESS RISER INTRODUCED DURING IMPACTION, WHICH LED TO CRACK INITIATION AND PROPAGATION. ADDITIONALLY, THE DEVICE MAY HAVE EXPERIENCED MATERIAL DEGRADATION IF STERILIZED BEYOND THE RECOMMENDED PARAMETERS LISTED IN THE REPROCESSING INSTRUCTION WHICH COULD HAVE LED TO THE OBSERVED FAILURE. SECTION H11: THE FOLLOWING SECTIONS HAVE CORRECTED INFORMATION: (H6) COMPONENT CODE: 4755, PART/COMPONENT/SUB-ASSEMBLY TERM NOT APPLICABLE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00488 | 2022-04-28 | 2018-07-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | HUM LNR IMPACT TIP 38M | EXACTECH, INC. |  | 321-07-38 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE IMPACTING PROCESS THE TIP FRACTURED AND SEPARATED FROM THE IMPACTOR HEAD. THE SURGEON TEMPLATED FOR A 36MM GLENOSPHERE/LINER COMBINATION. WHEN IMPACTING THE 36MM HUMERAL LINER, THE CLOSEST FITTING LINER WAS THE 38MM LINER IMPACTOR TIP. THE SURGERY WAS COMPLETED SUCCESSFULLY, WITH NO PATIENT ADVERSE EVENT OR DELAY TO SURGERY. NO PIECES FELL INTO THE SURGICAL SITE. ALL PIECES WERE RETURNED TO THE HEAD OFFICE AFTER BEING STERILIZED AND ARE BEING REPLACED. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT OR PATIENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR L0X OPTICAL UNLESS OTHERWISE SPECIFIED)-THE BROKEN DEVICE APPEARS CONSISTENT WITH CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE CAUSED BY A STRESS RISER INTRODUCED DURING IMPACTION. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE AND CONSISTENT WITH YEARS OF USE. DESIGN-RELATED ISSUES: THE DESIGN OF THE HUMERAL LINER IMPACTOR TIP HAS BEEN IN THE FIELD SINCE 2007. EXACTECH IS AWARE OF 19 SIMILAR COMPLAINT REPORTS INVOLVING 24 HUMERAL LINER IMPACTOR TIP SINCE 2008. THE FREQUENCY OF OCCURRENCE RANKING SCALE IS VERY LOW. THE DESIGN OF THIS DEVICE HAS SINCE BEEN UPDATED TO INCREASE THE THICKNESS AT THE TIP AND MODIFY THE CONNECTION FEATURES DIMENSIONS IN DCRTC- 14-0272. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT, WHICH HAS BEEN IN THE FIELD SINCE 2013. THIS DEVICE WAS MANUFACTURED PRIOR TO THE DESIGN UPDATE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE DESIGN OF THIS DEVICE HAS SINCE BEEN UPDATED, THE OCCURRENCE RATE IS " VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF A STRESS RISER INTRODUCED DURING IMPACTION, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE. THE DESIGN HAS SINCE BEEN UPDATED. IFU 700-096-181: INSTRUMENT INSPECTION VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY DESIGN ISSUES OR MANUFACTURING ISSUES. THERE WAS NO REPORTED PATIENT ADVERSE EVENT DUE TO THE IMPACTOR TIP FRAGMENTS. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
MW5109455 | 2022-04-29 | 2022-03-23 | I, n, j, u, r, y | Voluntary report | EXACTECH CONNEXION GXL ACETABULAR POLYETHYLENE LINERS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER, CEMENTED OR NON-POROUS | EXACTECH INC. | 132-36-53 | 132-36-53 |  | PATIENT HAD A TOTAL HIP IMPLANT FROM EXACTECH PLACED (B)(6) 2019. COMPONENTS OF THIS HIP REPLACEMENT SYSTEM HAVE A RECALL NOTICE ADVISING PHYSICIANS OF EARLY WEAR AND TO EVALUATE PATIENTS WITH X-RAYS TO DETERMINE IF EARLY WEAR IS AFFECTING THE PATIENTS IMPLANT. IF IT IS, A TOTAL HIP REVISION IS TO BE CONSIDERED. THIS PATIENT HAD A TOTAL HIP REVISION DUE TO THE COMPONENTS OF THE HIP IMPLANTS PLACED INITIALLY IN 2019. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00502 | 2022-04-29 | 2022-04-08 | I, n, j, u, r, y | Manufacturer report | LOGIC CR TIB INSERT STD, SZ 4, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC CR TIB INSERT STD, SZ 4, 9MM | 02-012-47-4009 |  | AS REPORTED, APPROXIMATELY THREE YEARS POST INITIAL RIGHT TKA, THIS MAL PATIENT HAD A REVISION DUE TO DEEP INFECTION. PATIENT REVISED WITH A NEW INSERT. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. SALES REP WAS UNABLE TO OBTAIN PHOTOS/X-RAYS. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICE WILL BE RETURN. | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION RECEIVED THE FOLLOWING SECTIONS G3, G6, H2, H3, H6 AND H9 HAVE BEEN UPDATED ACCORDINGLY. (H3) AS REPORTED, APPROXIMATELY 9 MONTHS POST REVISED RIGHT TKA, THIS MALE PATIENT HAD A REVISION DUE TO DEEP INFECTION. PATIENT REVISED WITH A NEW INSERT. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. SALES REP WAS UNABLE TO OBTAIN PHOTOS/X-RAYS. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. NO ADDITIONAL INFORMATION AVAILABLE. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN OR MANUFACTURING ISSUES. THE CAUSE OF SURGICAL REVISION FOR INFECTION CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION AND PLACING AN IMPLANT IN AN INFECTED JOINT. THESE DEVICES ARE USED FOR TREATMENT NOT DIAGNOSIS. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00491 | 2022-04-29 | 2018-12-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | HUM HD IMPCTR TIP | EXACTECH, INC. |  | 311-07-07 |  | IT WAS REPORTED FROM THE US THAT UPON IMPACTING THE HUMERAL ANATOMIC HEAD ON REPLICATOR PLATE, THE TIP OF THE IMPACTOR BROKE INTO TWO PIECES. THE AGENT WAS PRESENT AT SURGERY. THERE WAS NO DELAY TO THE SURGERY AND NO CLINICAL CONSEQUENCE TO THE PATIENT. PATIENT¿S HEALTH WAS STABLE LEAVING THE OR. NO PIECES WERE LEFT IN THE PATIENT. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT OR PATIENT. | H10. AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI (B)(6), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. VISUAL EVALUATION CONDITION/FINDINGS THE BROKEN IMPACTOR TIP APPEARS CONSISTENT WITH CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE. DESIGN-RELATED ISSUES: THE DESIGN OF THE HUMERAL HEAD IMPACTOR TIP HAS BEEN IN THE FIELD SINCE 2004. EXACTECH IS AWARE OF 9 SIMILAR COMPLAINT REPORTS INVOLVING THIS DEVICE SINCE 2008. BECAUSE THE HUMERAL HEAD IMPACTOR TIP IS USED IN ANATOMIC TOTAL SHOULDER SURGERIES, SALES DATA FOR HUMERAL HEADS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT, WHICH HAS BEEN IN THE FIELD SINCE 2015. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING-RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS HUMERAL HEAD IMPACTOR TIP, 750-2002-039-RMR-INSTRUMENT ADDITIONS (B)(4), WAS REVIEWED. THE RISK IS CAPTURED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS " VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE RADEL IMPACTION TIP. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES. THERE WAS NO REPORTED PATIENT ADVERSE EVENT, NO PIECES WERE LEFT IN THE PATIENT. THERE IS NO INDICATION OR REPORT THAT PIECES FELL IN THE WOUND SITE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00499 | 2022-04-29 | 2019-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | GLENOID IMPACTOR TIP | EXACTECH, INC. |  | 315-07-06 |  | IT WAS REPORTED FROM OUS THAT THE GLENOID IMPACTOR FRACTURED AND SNAPPED IN HALF. THE INSTRUMENT WAS DISCARDED. IT WAS REPORTED THAT NO DEVICE FRAGMENTS FELL INTO THE SURGICAL SITE. THERE WAS NO DELAY OF SURGERY AND THERE WAS NO PATIENT INJURY, ADVERSE EVENT, OR CLINICAL CONSEQUENCE TO THE PATIENT. THE PATIENT WAS DISCHARGED HOME WITH NO COMPLICATIONS. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. FINDINGS: THE REPORTED FRACTURED GLENOID IMPACTOR TIP WAS LIKELY THE RESULT OF CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE WHICH LIKELY RESULTED FROM ABNORMAL OR AGGRESSIVE USE. RMR REVIEW: THE RMR FOR THIS GLENOID IMPACTOR TIP, 750-2007-026-RMR-INSTRUMENTS REV L, WAS REVIEWED. THE RISK IS CAPTURED. MOST LIKELY CAUSE: BASED ON SIMILAR PREVIOUS COMPLAINT INVESTIGATIONS, THE REPORTED FRACTURED DEVICE REPORTED WAS LIKELY THE RESULT OF CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE WHICH WAS LIKELY CAUSED BY ABNORMAL OR AGGRESSIVE USE INCONSISTENT WITH THE INTENDED USE OF THIS DEVICE. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT, NO PIECES FELL INTO THE SURGICAL SITE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00496 | 2022-04-29 | 2018-12-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | 3.2MM DRILL GUIDE INSERT | EXACTECH, INC. |  | 181-00-06 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE SLEEVE OF THE DRILL GUIDE HAD FALLEN OFF FROM THE BODY OF THE DEVICE. THE DOCTOR USED A SIMILAR DRILL GUIDE FROM ANOTHER COMPANY INSTEAD OF THIS BROKEN DRILL GUIDE. THIS CAUSE A DELAY OF ABOUT 10 MINS. THIS DID NOT ADVERSELY AFFECT THE PATIENT AND THE PATIENT WAS STABLE LEAVING THE OPERATING ROOM. THE DEVICE WAS RETURNED FOR ANALYSIS. MULTIPLE ATTEMPTS HAVE BEEN MADE FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT OR PATIENT. THIS IS 2 OF 2 MDRS FOR THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. FINDINGS: THE SHEARED SURFACES APPEAR CONSISTENT WITH APPLYING A BENDING MOMENT TO THE DRILL GUIDE WHILE USING A DRILL, CAUSING THE WELD BETWEEN THE DRILL GUIDE TIP AND THE DRILL GUIDE SHAFT TO BREAK. MOST LIKELY UNDERLYING CAUSE: THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE DRILL GUIDE WHILE USING A DRILL, WHICH LED TO FRACTURE OF THE WELD BETWEEN THE DRILL GUIDE TIP AND THE SHAFT. RMR REVIEW: THE RMR FOR THIS DRILL GUIDE, 750-2005-028-RMR-INSTRUMENTS REV U, WAS REVIEWED. THE RISK IS CAPTURED. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. ¿ CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THESE DEVICES ARE USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT IMPACT. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00501 | 2022-04-29 | 2022-03-30 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REV LOCKING SCREW | EXACTECH, INC. | 320-15-05 | 320-15-05 |  | AS REPORTED, 8 DAYS POST OP THE INITIAL RIGHT TSA, THIS (B)(6) FEMALE PATIENT WAS REVISED BECAUSE THE ORIGINAL RTSA FAILED TO SEAT THE GLENOSPHERE CORRECTLY ON THE 8 DEGREE POSTERIOR AUGMENT. SURGEON TOOK SCREW OUT, SEATED GLENOSPHERE AND USED NEW LOCKING SCREW. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICE WILL NOT BE RETURNING, DISPOSED BY HOSPITAL. | SECTION H10: (H3) THE REVISION REPORTED WAS LIKELY THE RESULT OF EITHER OFF-AXIS SEATING OF THE GLENOSPHERE ON THE BASEPLATE OR INCOMPLETE SEATING OF THE GLENOSPHERE LOCKING SCREW AT THE TIME OF IMPLANTATION. INCOMPLETE SEATING OF THE SCREW MAY HAVE BEEN DUE TO BONE GRAFT OR ALLOGRAFT BEING RETAINED ON THE SCREW-HOLE THREADS OF THE BASEPLATE, THE GLENOSPHERE LOCKING SCREW BEING CROSS-THREADED INTO THE BASEPLATE, OR A COMBINATION OF THE TWO, WHICH LED TO THE INABILITY TO FULLY SEAT THE GLENOSPHERE ONTO THE BASEPLATE. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT RETURNED FOR EVALUATION AND PRE-REVISION X-RAYS WERE NOT PROVIDED. | Y | N | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00494 | 2022-04-29 | 2018-12-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | DRILL GUIDE | EXACTECH, INC. |  | 181-00-05 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE SLEEVE OF THE DRILL GUIDE HAD FALLEN OFF FROM THE BODY OF THE DEVICE. THE DOCTOR USED A SIMILAR DRILL GUIDE FROM ANOTHER COMPANY INSTEAD OF THIS BROKEN DRILL GUIDE. THIS CAUSE A DELAY OF ABOUT 10 MINS. THIS DID NOT ADVERSELY AFFECT THE PATIENT AND THE PATIENT WAS STABLE LEAVING THE OPERATING ROOM. THE DEVICE WAS RETURNED FOR ANALYSIS. MULTIPLE ATTEMPTS HAVE BEEN MADE FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. THIS IS 1 OF 2 MDRS FOR THIS EVENT. | H10. AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. FINDINGS: THE SHEARED SURFACES APPEAR CONSISTENT WITH APPLYING A BENDING MOMENT TO THE DRILL GUIDE WHILE USING A DRILL, CAUSING THE WELD BETWEEN THE DRILL GUIDE TIP AND THE DRILL GUIDE SHAFT TO BREAK. MOST LIKELY UNDERLYING CAUSE: THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT TO THE DRILL GUIDE WHILE USING A DRILL, WHICH LED TO FRACTURE OF THE WELD BETWEEN THE DRILL GUIDE TIP AND THE SHAFT. RMR REVIEW: THE RMR FOR THIS DRILL GUIDE, 750-2005-028-RMR-INSTRUMENTS (B)(4), WAS REVIEWED. THE RISK IS CAPTURED. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THESE DEVICES ARE USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT IMPACT. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00487 | 2022-04-29 | 2018-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | GLENOID IMPACTOR TIP | EXACTECH, INC. |  | 315-07-06 |  | IT WAS REPORTED FROM OUS THAT DURING THE IMPACTION PROCESS THE TIP FRACTURED AND SEPARATED FROM THE IMPACTOR HEAD DURING AN ORTHOPEDIC SURGERY. SURGERY WAS COMPLETED SUCCESSFULLY, AND THE EVENT DID NOT INTERFERE WITH SURGERY. ALL PIECES WERE RETURNED TO THE HEAD OFFICE AFTER BEING STERILIZED AND WERE REPLACED. NO PIECES FELL INTO THE SURGICAL SITE. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT OR PATIENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. PATIENT INFORMATION WAS NOT PROVIDED. VISUAL EVALUATION (CONDUCTED WITH 7X OR 10X OPTICAL) THE FRACTURED GLENOID IMPACTOR TIP APPEARS CONSISTENT WITH ABNORMAL OR AGGRESSIVE USE. THE DAMAGE TO THE IMPACTION SURFACE APPEARS CONSISTENT WITH CONTACTING ANOTHER SURGICAL INSTRUMENT DURING USE. THE DAMAGE TO THE BASE OF THE MATING FEATURE APPEARS CONSISTENT WITH A MATING DEVICE (E.G. AN IMPACTION HANDLE) CONTACTING THIS SURFACE DURING IMPACTION. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE AND CONSISTENT WITH YEARS OF USE. DESIGN-RELATED ISSUES: THE DESIGN OF THE GLENOID IMPACTOR TIP HAS BEEN IN THE FIELD SINCE 2013. EXACTECH IS NOT AWARE OF ANY OTHER SIMILAR COMPLAINT REPORTS INVOLVING THIS DEVICE SINCE 2013. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT, WHICH HAS BEEN IN THE FIELD SINCE 2015. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING-RELATED. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS " VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF ABNORMAL OR AGGRESSIVE USE INCONSISTENT WITH THE INTENDED USE OF THIS DEVICE. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. NO PIECES FELL INTO THE SURGICAL SITE. | N | Y | HEALTH PROFESSIONAL |  | 2022-04
1038671-2022-00503 | 2022-05-03 | 2022-04-06 | I, n, j, u, r, y | Manufacturer report | PS TIBIAL INSERTS SZ 4, 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | PS TIBIAL INSERTS SZ 4, 11MM | 204-24-11 |  | AS REPORTED, APPROXIMATELY EIGHT YEARS POST INITIAL LEFT TKA, THIS (B)(6) MALE PATIENT UNDERWENT REVISION OF THE TOTAL TKR OPTETRAK DUE TO POLY WEAR AND DEBRIS IN THE KNEE WERE INDICATED. THERE WAS NO BREAKAGE OF A DEVICE. THERE WAS NO SURGICAL DELAYS REPORTED. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. X-RAYS WERE NOT ABLE TO BE OBTAINED BY SALES REP. DEVICE IMAGES RECEIVED. THE DEVICE WILL BE RETURN. | CONCOMITANT PRODUCTS: OPTETRAK 3 PEG PATELLA (CAT# 200-02-32 / SERIAL# (B)(4)) OPTETRAK TIBIAL TRAY (CAT# 204-04-44 / SERIAL# (B)(4)) OPTETRAK ASYMMETRIC FEMORAL POSTERIAR STABILIZER (CAT# 234-02-04 / SERIAL# (B)(4)) ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00504 | 2022-05-03 | 2022-04-01 | I, n, j, u, r, y | Manufacturer report | NV GXL LINER NEUTRAL, 36MM ID, GROUP 2 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LINER NEUTRAL, 36MM ID, GROUP 2 CUPS | 130-36-52 |  | AS REPORTED, APPROXIMATELY NINE YEARS POST INITIAL RIGHT THA, THIS 66 Y/O MALE PATIENT HAD A REVISION DUE TO SIGNIFICANT WEAR OF G2 36MM GXL NOVATION LINER. PATIENT WAS REVISED TO 36MM OPTION HEAD/+0 SLEEVE AND G2 36MM GXE NOVATION LINER. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO HSS KEEPS ALL RETRIEVALS. | CONCOMITANT PRODUCTS: BIOLOX DELTA FEMORAL HEAD 36MM OD, +0MM (CAT# 170-36-00 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00506 | 2022-05-03 | 2022-03-30 | I, n, j, u, r, y | Manufacturer report | NV GXL LNR, +5LAT, 36MM G2-52/54MM CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LNR, +5LAT, 36MM G2-52/54MM CUPS | 136-36-52 |  | AS REPORTED, APPROXIMATELY EIGHT YEARS POST INITIAL RIGHT THA, THIS (B)(6) MALE PATIENT WAS INFORMED THAT HIP ACETAB LINER WAS WORN-OUT AND NEED REPLACEMENT. PATIENT WAS REVISED WITH A NV EHXL NTRL LNR G2 36MM. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO HOSPITAL WILL NOT ALLOW REMOVAL OF IMPLANTS. | H3: BASED ON THE AVAILABLE INFORMATION, THE PATIENT INVOLVED MEETS THE FOLLOWING RISK CRITERIA FOR EARLY PROSTHESIS WEAR/OSTEOLYSIS IMPLANTED WITH A LATERALIZED LINER. THE MOST LIKELY CAUSE FOR THE REVISION REPORTED DUE TO EARLY PROSTHESIS WEAR IS A COMBINATION OF THE RISK FACTORS INCLUDING, USE ERROR, IMPLANT POSITIONING, IMPLANT SIZE SELECTION, AND PATIENT FACTORS MAY HAVE ALSO BEEN A CONTRIBUTING FACTOR TO THE EARLY PROSTHESIS WEAR. HOWEVER, THE PROSTHESIS WEAR CANNOT BE CONFIRMED AS THE DEVICES WERE NOT AVAILABLE FOR EVALUATION AND LIMITED CLINICAL INFORMATION WAS AVAILABLE AT THE TIME OF EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00505 | 2022-05-03 | 2022-04-12 | I, n, j, u, r, y | Manufacturer report | TRULIANT PS CEM FEM PS CEM RIGHT SZ 5 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT PS CEM FEM PS CEM RIGHT SZ 5 | 02-020-11-0350 |  | AS REPORTED, APPROXIMATELY 1.5 YEARS POST INITIAL RIGHT TKA, THIS (B)(6) MALE PATIENT COMPLAINED OF PAIN AND LOSE FEMUR. THERE WAS NO BREAKAGE OF DEVICES OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PHOTOS RECEIVED. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO CHAIN OF CUSTODY BETWEEN HOSPITAL AND PATIENT; THEREFORE. SALES REP WAS UNABLE TO RETRIEVE IT. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00518 | 2022-05-04 | 2022-04-11 | I, n, j, u, r, y | Manufacturer report | NOVATION STEM | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NOVATION STEM | UNKNOWN |  | AS REPORTED, THE FEMORAL STEM, WHICH IS A NOVATION ELEMENT 12/14 STEM APPEARS TO HAVE SUBSIDED IN THE FEMUR. THIS (B)(6) FEMALE PATIENT UNDERWENT A REVISION TOTAL HIP SURGERY TO TRY TO REMOVE THE STEM. THE STEM WOULD NOT MOVE DURING REMOVAL IMPACTION. THE SURGEON DID NOT REVISE THE CUP OR LINER. LINER OPTIONS WERE NOT REQUESTED TO BE BROUGHT TO THE REVISION SURGERY. ONLY FEMORAL HEADS WERE REQUESTED TO BE BROUGHT. A NEW HEAD WITH AN ADAPTOR +7MM AND 36MM WAS REPLACED DURING THIS REVISION. THERE WAS NO BREAKAGE OF DEVICES OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE IMPLANT INFORMATION IS UNKNOWN. X-RAYS RECEIVED. THE HEAD IS AVAILABLE FOR RETURN TO BE EXAMINED. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00523 | 2022-05-04 | 2019-01-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LOGIC CR INSERT TRIAL SIZE 2.5, NEUTRAL, 9 MM | EXACTECH, INC. |  | 02-013-57-2509 |  | IT WAS REPORTED FROM OUS THAT UPON RETURN OF THE LOAN KIT THE DEVICE WAS DISCOVERED TO HAVE FRACTURED ON THE FACE OF THE DEVICE. THIS WAS DISCOVERED IN THE WAREHOUSE, NO COMPLAINTS FROM SURGERY. NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DESIGN-RELATED ISSUES: THE DESIGN OF THE LOGIC CR TIBIAL INSERT TRIAL HAS BEEN IN THE FIELD SINCE 2012. EXACTECH IS AWARE OF 48 SIMILAR COMPLAINT REPORTS INVOLVING 68 CRACKED/FRACTURED TIBIAL INSERT TRIALS (17 CRACKED; 51 FRACTURED) SINCE 2012. BECAUSE THE LOGIC CR TIBIAL INSERT TRIAL IS USED IN TOTAL KNEE SURGERIES, SALES DATA FOR TIBIAL INSERTS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4). THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF (B)(4) UNITS, WHICH HAVE BEEN IN THE FIELD SINCE 2014. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS LOGIC CR TIBIAL INSERT TRIAL, (B)(4), WAS REVIEWED. THE RISK IS CAPTURED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR. THE DAMAGED LOGIC CR TIBIAL INSERT TRIAL REPORTED WAS LIKELY THE RESULT OF EITHER YEARS OF NORMAL USE AND STERILIZATION CYCLES, THE DEVICE COMING IN CONTACT WITH SHARP SURGICAL TOOLS DURING TRIALING, OR A COMBINATION OF THE TWO. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE WAS NO REPORT FROM AN AGENT OR FACILITY ABOUT AN ISSUE IN SURGERY DUE TO THE TRIAL FRAGMENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00522 | 2022-05-04 | 2019-01-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | HUM LNR IMPACT TIP 38M | EXACTECH, INC. |  | 321-07-38 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE IMPACTOR TIP. THE DEVICE WAS BROKEN BY MALLETING WHEN THE IMPLANTS WERE INSERTED. THE PATIENT¿S HEALTH WAS STABLE LEAVING THE OR. THERE WAS NO DELAY, THE DEVICE WAS REPLACED BY A 42MM IMPACTOR. THE AGENT STATED THE INSTRUMENT¿S DEBRIS WAS COMPLETELY REMOVED FROM THE PATIENT¿S BODY. NEW INFORMATION FROM 27 JUL 2021: NO PART/PIECES FELL INTO THE SURGICAL SITE. THE DEVICE WAS NOT RECEIVED FOR ANALYSIS. NO PATIENT INFORMATION WAS PROVIDED. ATTEMPTS WERE MADE TO REQUEST ADDITIONAL INFORMATION FROM THE AGENT AND A RESPONSE WAS RECEIVED. PER THE AGENT, NO PROBLEMS WERE NOTED AT THE POST-OP FOLLOW UP. NO REPRESENTATIVE PRESENT DURING SURGERY. NO ADDITIONAL INFORMATION IS AVAILABLE. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: COUNTRY-(B)(4) THE DEVICE WAS NOT AVAILABLE FOR ANALYSIS. DESIGN-RELATED ISSUES: THE DESIGN OF THE HUMERAL LINER IMPACTOR TIP HAS BEEN IN THE FIELD SINCE 2007. EXACTECH IS AWARE OF 28 SIMILAR COMPLAINT REPORTS INVOLVING 34 BROKEN HUMERAL LINER IMPACTOR TIPS SINCE 2008. THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 30 UNITS, WHICH HAS BEEN IN THE FIELD SINCE 2015. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS HUMERAL LINER IMPACTOR TIP, 750-2005-021-RMR-INSTRUMENTS REV N, WAS REVIEWED. THE RISK IS CAPTURED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS VERY LOW, AND THE RISK IS CAPTURED IN THE RMR. THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF A STRESS RISER INTRODUCED DURING IMPACTION, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. THE IMPACTOR TIP BROKE, NO PARTS/PIECES FELL INTO THE SURGICAL SITE. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00521 | 2022-05-04 | 2019-02-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | GLENOID IMPACTOR TIP | EXACTECH, INC. |  | 315-07-06 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE IMPACTOR TIP. THE TIP BROKE DURING IMPACTING. THE PATIENT¿S HEALTH WAS STABLE LEAVING THE OPERATING ROOM. NO PIECES WERE LEFT IN THE PATIENT. THERE WAS NO SURGICAL DELAY, AND THE PATIENT WAS NOT ADVERSELY AFFECTED. THE REP WAS PRESENT AT THE TIME OF SURGERY. NEW INFORMATION RECEIVED ON 18 JUL 21: NO PARTS/PIECES OF THE DEVICE FELL INTO THE SURGICAL SITE AND THE PATIENT WAS KNOWN TO BE STABLE. NO PATIENT INFORMATION WAS PROVIDED. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DESIGN-RELATED ISSUES: THE DESIGN OF THE GLENOID IMPACTOR TIP HAS BEEN IN THE FIELD SINCE 2013. EXACTECH IS AWARE OF 8 COMPLAINT REPORTS INVOLVING 9 BROKEN GLENOID IMPACTOR TIPS SINCE 2013. THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 100 UNITS, WHICH HAS BEEN IN THE FIELD SINCE 2017. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING-RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS GLENOID IMPACTOR TIP, (B)(4)-INSTRUMENT REV L, WAS REVIEWED. THE RISK IS CAPTURED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR. THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF A COMBINATION OF ABNORMAL AND AGGRESSIVE USE AND AN IMPACTOR HANDLE TIP CONTACTING THE INNER SURFACE OF THE IMPACTOR TIP DURING IMPACTION. IFU 700-096-181: INSTRUMENT INSPECTION : VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THE TIP BROKE DURING IMPACTING, IT WAS REPORTED THAT THERE WERE NO PARTS OR PIECES THAT FELL INTO THE SURGICAL SITE. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00520 | 2022-05-04 | 2019-03-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | M-SERIES 12/14 FEMORAL HEAD IMPACTOR | EXACTECH, INC. |  | 153-00-02 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE INSTRUMENT. UPON IMPACTING FEMORAL HEAD, THE WHITE IMPACTOR HEAD SPLIT. THE WHITE PART WAS THROWN AWAY BY HOSPITAL STAFF. ALL PIECES WERE ACCOUNTED FOR. THERE IS NO REPORT OR INDICATION THAT DEVICE PIECES FELL INTO THE SURGICAL SITE. THERE WAS NO DELAY TO THE CASE AND NO ADVERSE EFFECTS TO THE PATIENT. REP UNSURE THE EXACT DATE OF SURGERY JUST REMEMBERS IT WAS ON A THURSDAY. THE DEVICE WAS RETURNED. THE REP WAS PRESENT AT THE TIME OF SURGERY. ADDITIONAL ATTEMPTS WERE MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REP, BUT NO RESPONSES WERE RECEIVED. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DESIGN-RELATED ISSUES: THE DESIGN OF THE M-SERIES 12/14 FEMORAL HEAD IMPACTOR HAS BEEN IN THE FIELD SINCE 2000. EXACTECH IS AWARE OF 3 SIMILAR COMPLAINT REPORTS INVOLVING THIS DEVICE SINCE 2008. BECAUSE THE M-SRIES 12/14 FEMORAL HEAD IMPACTOR IS USED IN TOTAL HIP SURGERIES, SALES DATA FOR 12/14 FEMORAL HEADS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT, WHICH HAS BEEN IN THE FIELD SINCE 2009. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS M-SERIES 12/14 FEMORAL HEAD IMPACTOR, 750-1997-032-RMR-INSTRUMENTS REV A, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 11. HOWEVER, THE MITIGATION IN THIS LINE IS NOT APPROPRIATE FOR THIS REPORTED FAILURE. THEREFORE, THIS RMR WILL BE UPDATED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR. THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF YEARS OF NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE RADEL TIP. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00519 | 2022-05-04 | 2019-03-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LINER DRIVER HEAD 32MM | EXACTECH, INC. |  | 141-01-32 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE INSTRUMENT. THE 32MM LINER DRIVER HEAD FRACTURED DURING IMPACTION OF THE LINER. SURGERY COMPLETED AS PLANNED, PATIENT DOING FINE. NO DELAY TO SURGERY. THE DEVICE WAS NOT RECEIVED FOR ANALYSIS. THE REPRESENTATIVE WAS PRESENT AT THE TIME OF SURGERY. NO PATIENT DEMOGRAPHIC INFORMATION WAS PROVIDED. ATTEMPTS WERE MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REPRESENTATIVE, BUT NO RESPONSE WAS RECEIVED. NO ADDITIONAL INFORMATION IS AVAILABLE. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS NOT AVAILABLE FOR ANALYSIS. DESIGN-RELATED ISSUES: THE DESIGN OF THE LINER DRIVER HEAD HAS BEEN IN THE FIELD SINCE 2005. EXACTECH IS AWARE OF 4 SIMILAR COMPLAINT REPORTS INVOLVING BROKEN LINER DRIVER HEADS SINCE 2008. BECAUSE THE LINER DRIVER HEAD IS USED IN TOTAL HIP SURGERIES, SALES DATA FOR ACETABULAR LINERS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4). THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: MANUFACTURING-RELATED DATA COULD NOT BE REVIEWED BECAUSE NO INFORMATION REGARDING THE MANUFACTURING LOT NUMBER WAS PROVIDED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS LINER DRIVER HEAD, 750-2003-022-RMR-INSTRUMENTS REV F, WAS REVIEWED. THE RISK IS CAPTURED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR. THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF A CRACK INITIATING WHERE THE TIP OF THE LINER DRIVER HANDLE MEETS THE LINER DRIVER HEAD, WHICH PROPAGATED OVER TIME RESULTING IN ULTIMATE FRACTURE OF THE LINER DRIVER HEAD. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. ¿ CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00524 | 2022-05-04 | 2019-01-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | HUM LNR IMPACT TIP 42M | EXACTECH, INC. |  | 321-07-42 |  | IT WAS REPORTED BY THE OUS THAT THE IMPACTOR TIP WAS BROKEN BY MALLETING WHEN THE IMPLANTS WERE INSERTED. NO PIECES FELL INTO THE SURGICAL SITE, NO PIECES WERE LEFT IN THE PATIENT. THE SURGEON REPLACED THE 38MM IMPACTOR HEAD. THERE WAS NO DELAY AND THE PATIENT WAS NOT ADVERSELY AFFECTED. THE PATIENT¿S HEALTH WAS STABLE LEAVING THE OPERATING ROOM. THE REP WAS PRESENT AT THE TIME OF SURGERY. TWO ATTEMPTS WERE MADE TO REQUEST ADDITIONAL INFORMATION DURING REMEDIATION AND A RESPONSE WAS RECEIVED. PER THE REP, THERE WERE NO PROBLEMS NOTED AT THE POST-OP FOLLOW-UP. THE DEVICE WAS NOT RETURNED. PATIENT DEMOGRAPHIC INFORMATION WILL NOT BE PROVIDED BY CONTACT. NO ADDITIONAL INFORMATION ABOUT THE EVENT HAS BEEN PROVIDED. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH (B)(4) HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PATIENT ID, AGE, GENDER, WEIGHT, ETHNICITY AND RACE. INITIAL REPORTER COUNTRY: (B)(6). THE DEVICE WAS NOT AVAILABLE FOR ANALYSIS. DESIGN-RELATED ISSUES: THE DESIGN OF THE HUMERAL LINER IMPACTOR TIP HAS BEEN IN THE FIELD SINCE 2007. EXACTECH IS AWARE OF (B)(4) SIMILAR COMPLAINT REPORTS INVOLVING (B)(4) BROKEN HUMERAL LINER IMPACTOR TIPS SINCE 2008. BECAUSE THE HUMERAL LINER IMPACTOR TIP IS USED IN REVERSE TOTAL SHOULDER SURGERIES, SALES DATA FOR HUMERAL LINERS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4). THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF (B)(4) UNITS, WHICH HAS BEEN IN THE FIELD SINCE 2015. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS HUMERAL LINER IMPACTOR TIP, 750-2005-021-RMR INSTRUMENTS REV N, WAS REVIEWED. THE RISK IS CAPTURED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR. THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF A STRESS RISER INTRODUCED DURING IMPACTION, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. THE TIP WAS BROKEN; NO PIECES FELL INTO THE SURGICAL SITE, NO PIECES WERE LEFT IN THE PATIENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00516 | 2022-05-04 | 2022-04-13 | I, n, j, u, r, y | Manufacturer report | TRULIANT TIB IMP CRC INSERT SZ 1.5, 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT TIB IMP CRC INSERT SZ 1.5, 11MM | 02-022-51-1511 |  | AS REPORTED, APPROXIMATELY 2.5 YEARS POST INITIAL RIGHT TKA, THIS (B)(6) MALE PATIENT'S SURGEON BOOKED A REVISION TKA WITH PATELLECTOMY. THE PATIENT HAD A FALL ON THEIR TOTAL KNEE AND POST THE FALL THE PATIENT WAS HAVING PAIN AND INSTABILITY. THE POLY IMPLANT THAT WAS IN THE PATIENT AND REMOVED DURING THE REVISION SURGERY WAS AFFECTED BY THE EXACTECH MANUFACTURER KNEE POLY RECALL. THE SURGEON WENT FROM WHAT WAS IN AN 11 MM CRC TO TRAILING AND IMPLANTING A 15 MM CRC POLY. THE SURGEON REMOVED BONE FROM THE PATELLA AND CLOSED THE KNEE. THE PATIENT LEFT THE OR STABLE. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PHOTOS & X-RAYS RECEIVED. THE DEVICES ARE NOT AVAILABLE FOR EVALUATION AS THE DEVICES WERE DISCARDED OR LOST IN TRANSIT FROM THE SX ROOM TO DECONTAMINATION AND SPD. | (H3) AS REPORTED, APPROXIMATELY 3.5 YEARS POST INITIAL RIGHT TKA, THIS 75 Y/O MALE PATIENT'S SURGEON BOOKED A REVISION TKA WITH PATELLECTOMY. THE PATIENT HAD A FALL ON THEIR TOTAL KNEE AND POST THE FALL THE PATIENT WAS HAVING PAIN AND INSTABILITY. THE POLY IMPLANT THAT WAS IN THE PATIENT AND REMOVED DURING THE REVISION SURGERY WAS AFFECTED BY THE EXACTECH MANUFACTURER KNEE POLY RECALL. THE SURGEON WENT FROM WHAT WAS IN AN 11 MM CRC TO TRAILING AND IMPLANTING A 15 MM CRC POLY. THE SURGEON REMOVED BONE FROM THE PATELLA AND CLOSED THE KNEE. THE PATIENT LEFT THE OR STABLE. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PHOTOS & X-RAYS RECEIVED. THE DEVICES ARE NOT AVAILABLE FOR EVALUATION AS THE DEVICES WERE DISCARDED OR LOST IN TRANSIT FROM THE OPERATING ROOM TO DECONTAMINATION AND SPD. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN ISSUES. THE CAUSE OF THE SUBSEQUENT REVISION CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION AND A TRAUMA. THESE DEVICES ARE USED FOR TREATMENT NOT DIAGNOSIS. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00515 | 2022-05-04 | 2022-04-13 | I, n, j, u, r, y | Manufacturer report | LOGIC CC FEMORAL SIZE 2, RIGHT | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC CC FEMORAL SIZE 2, RIGHT | 02-010-06-0320 |  | AS REPORTED, A REPRESENTATIVE INFORMED THAT ON (B)(6) 2022 THAT A SURGEONS WILL BE REPLACING A LOOSE FEMORAL STEM EXTENSION SCREW IN ONE OF HIS PATIENT'S RIGHT KNEE. THE REVISION TO OCCUR ON (B)(6) 2022. THE X-RAY IMAGES PROVIDED INDICATE THAT THIS IS A LOGIC CC FEMORAL COMPONENT. THE PLAN IS TO RETAIN THE FEMORAL COMPONENT, REMOVE THE LOOSE SCREW, AND REPLACE WITH A NEW SCREW THAT THEY WILL TIGHTEN DOWN INTRAOPERATIVELY. NO ADDITIONAL INFORMATION AVAILABLE. | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION RECEIVED THE FOLLOWING SECTIONS D6B, G3, G6, H2, H3 AND H6 HAVE BEEN UPDATED ACCORDINGLY. (H3) THE REVISION REPORTED MAY HAVE BEEN THE RESULT OF NOT FULLY TIGHTENING THE SCREW INTO THE FEMORAL COMPONENT USING THE TORQUE-LIMITING DRIVER HANDLE, WHICH IS PRESET TO RELEASE WHEN THE APPROPRIATE TORQUE HAS BEEN ACHIEVED, ALLOWING FOR THE SCREW TO BACK OUT. HOWEVER, THIS CANNOT BE CONFIRMED AS THE SCREW WAS NOT RETURNED FOR EVALUATION. THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE REPORTED EVENT OF "DISASSEMBLED/MISASSEMBLED¿ IS ASSOCIATED WITH THE UNWANTED DISASSEMBLY OF THE DEVICE, OR INCORRECT ASSEMBLY OF THE DEVICE (ATTACHING MATING INSTRUMENTATION INCORRECTLY, ASSEMBLING AN IMPLANT CONSTRUCT OFF-AXIS, ETC.). SECTION D1: BRAND NAME SECTION D4: MODEL NUMBER, CATALOG NUMBER, SERIAL NUMBER, EXPIRATION DATE, UNIQUE IDENTIFIER (UDI) # SECTION D9: DEVICE AVAILABLE FOR EVALUATION SECTION D10: CONCOMITANTS PRODUCTS - TRULIANT TIB IMP PSC INSERT SZ 2, 13MM (CAT# 02-022-44-2013 / SERIAL# (B)(4)) - TRULIANT TIB IMP PSC INSERT SZ 2, 15MM (CAT# 02-022-44-2015 / SERIAL# (B)(4)) | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00513 | 2022-05-04 | 2019-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRULIANT PIN PULLER SYRINGE STYLE | EXACTECH, INC. |  | 02-029-90-4100 |  | IT WAS REPORTED BY EXACTECH INTERNAL THAT A BROKEN DEVICE WAS FOUND DURING 16ADBULK REVIEW. THERE HAS BEEN NO REPORT ABOUT THESE DEVICES FROM AN AGENT OR FACILITY. NO ADDITIONAL INFORMATION IS AVAILABLE. THIS IS 2 OF 2 MDRS FOR THIS CASE. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. FINDINGS: ONE OR MORE OF THE METAL PRONGS USED TO CLAMP A PIN HEAD ARE MISSING FROM THE RECEIVED TRULIANT SYRINGE PIN PULLERS. WHEN ANALYZING THE FRACTURE SURFACE AS PART OF ANOTHER COMPLAINT, THE FAILURE APPEARS CONSISTENT WITH BRITTLE FRACTURE. THIS TYPE OF FAILURE IS MOST LIKELY THE RESULT OF THE PRONGS EXPERIENCING EXCESSIVE FORCES. FURTHERMORE, IF THE PIN PULLER WAS USED OFF-AXIS RELATIVE TO THE BONE PIN, AN UNEVEN FORCE DISTRIBUTION MAY HAVE BEEN APPLIED TO THE METAL PRONGS AND COULD EXPLAIN WHY THEY FRACTURED FROM ONLY ONE SIDE. OVER TIME, THESE EXCESSIVE OFF-AXIS FORCES COULD POTENTIALLY LEAD TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE MATERIAL. RMR REVIEW: A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THESE TRULIANT SYRINGE PIN PULLERS, 750-2014-006-RMR-INSTRUMENTS REV C, WAS REVIEWED. THE RISK IS CAPTURED. MOST LIKELY UNDERLYING CAUSE: THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF THE PRONGS OF THE PIN PULLER EXPERIENCING EXCESSIVE FORCES DURING THE EXTRACTION OF BONE PINS. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT FROM AN AGENT OR FACILITY. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00514 | 2022-05-04 | 2019-04-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | DRL GD HNDL | EXACTECH, INC. |  | 315-07-20 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE INSTRUMENTATION. THE DRILL GUIDE HANDLE BROKE DURING INSERTION. NO PARTS/PIECES FELL INTO THE SURGICAL SITE. THERE WAS NO DELAY AND NO INJURY/CLINICAL CONSEQUENCE/ADVERSE EVENT. MAGES WERE PROVIDED. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. REQUEST FOR ADDITIONAL INFORMATION WAS SENT TO THE REP AND A RESPONSE WAS RECEIVED. THERE IS NO ADDITIONAL INFORMATION AVAILABLE FROM THE CONTACTS ABOUT THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. SEOUL, KOREA. DESIGN-RELATED ISSUES: THE DESIGN OF THE DRILL GUIDE HANDLE HAS BEEN IN THE FIELD SINCE 2004. EXACTECH IS AWARE OF 5 SIMILAR COMPLAINT REPORTS INVOLVING BROKEN DRILL GUIDE HANDLES DUE TO FRACTURED THREADS SINCE 2008. BECAUSE THE DRILL GUIDE HANDLE IS USED IN BOTH ANATOMIC AND REVERSE SHOULDER SURGERIES, SALES DATA FOR HUMERAL STEMS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 30 UNITS, WHICH HAS BEEN IN THE FIELD SINCE 2015. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) AND RISK ASSESSMENT AND CONTROLS REPORT (RACR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR AND RACR FOR THIS DRILL GUIDE HANDLE, RMR-00006 REV ¿ AND RACR-00009 REV, WERE REVIEWED. THE RISK IS CAPTURED, HOWEVER THE RACR IS BEING UPDATED UNDER DCRSP-19-541 TO CAPTURE AN APPROPRIATE HARM FOR THIS HAZARDOUS SITUATION. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, THE RISK IS CAPTURED IN THE RMR AND RACR, AND THE HARM WILL BE UPDATED. THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF APPLYING AN OFF-AXIS TORQUE TO THE HANDLE DURING USE WHICH ALLOWED THE THREADS TO FRACTURE. HOWEVER; THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT DUE TO THE BROKEN DEVICE. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00512 | 2022-05-04 | 2019-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRULIANT PIN PULLER SYRINGE STYLE | EXACTECH, INC. |  | 02-029-90-4100 |  | IT WAS REPORTED BY EXACTECH INTERNAL THAT A BROKEN DEVICE WAS FOUND DURING 16ADBULK REVIEW. THERE HAS BEEN NO REPORT ABOUT THESE DEVICES FROM AN AGENT OR FACILITY. NO ADDITIONAL INFORMATION IS AVAILABLE. THIS IS 1 OF 2 MDRS FOR THIS CASE. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. FINDINGS: ONE OR MORE OF THE METAL PRONGS USED TO CLAMP A PIN HEAD ARE MISSING FROM THE RECEIVED TRULIANT SYRINGE PIN PULLERS. WHEN ANALYZING THE FRACTURE SURFACE AS PART OF ANOTHER COMPLAINT, THE FAILURE APPEARS CONSISTENT WITH BRITTLE FRACTURE. THIS TYPE OF FAILURE IS MOST LIKELY THE RESULT OF THE PRONGS EXPERIENCING EXCESSIVE FORCES. FURTHERMORE, IF THE PIN PULLER WAS USED OFF-AXIS RELATIVE TO THE BONE PIN, AN UNEVEN FORCE DISTRIBUTION MAY HAVE BEEN APPLIED TO THE METAL PRONGS AND COULD EXPLAIN WHY THEY FRACTURED FROM ONLY ONE SIDE. OVER TIME, THESE EXCESSIVE OFF-AXIS FORCES COULD POTENTIALLY LEAD TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE MATERIAL. RMR REVIEW: A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THESE TRULIANT SYRINGE PIN PULLERS, 750-2014-006-RMR-INSTRUMENTS REV C, WAS REVIEWED. THE RISK IS CAPTURED. MOST LIKELY UNDERLYING CAUSE: THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF THE PRONGS OF THE PIN PULLER EXPERIENCING EXCESSIVE FORCES DURING THE EXTRACTION OF BONE PINS. (B)(4): INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. ¿ CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT FROM AN AGENT OR FACILITY. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00507 | 2022-05-04 | 2022-03-29 | I, n, j, u, r, y | Manufacturer report | TRULIANT TIB IMP CR INS STD SZ 1.5, 9 MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT TIB IMP CR INS STD SZ 1.5, 9 MM | 02-012-47-1509 |  | AS REPORTED, APPROXIMATELY THREE YEARS POST INITIAL LEFT TKA, SURGEON REMOVED THE EXISTING INSERT OF THIS 70 Y/O FEMALE PATIENT AND REPLACED WITH A NEW INSERT. IMO THE FEMUR APPEARED TO BE LOOSE. SURGEON TOOK SAMPLES FOR CULTURE AND ELECTED TO CHANGE ONLY THE TIBIAL INSERT AT THIS TIME. PATIENT WAS REVISED TO A TRULIANT TIB IMP CR INS STD SZ 1.5, 9 MM. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PHOTOS AND X-RAY RECEIVED. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO FACILITY WILL NOT RELEASE TO REP. AS OF NOTE" IT APPEARED THAT PATIENT WAS EXPERIENCING CARDIAC ISSUES THAT ACCORDING TO STAFF IN ROOM HAD NOT APPEARED ON THE CARDIAC CLEARANCE". | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00517 | 2022-05-04 | 2019-03-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRULIANT PIN PULLER SYRINGE STYLE | EXACTECH, INC. |  | 02-029-90-4100 |  | IT WAS REPORTED BY EXACTECH INTERNAL THAT THE INSTRUMENT WAS FOUND IN 16 ADBULK ON (B)(6) 2019 NOT WORKING PROPERLY. ONE OR MORE OF THE METAL PRONGS USED TO CLAMP A PIN HEAD ARE MISSING FROM THE RECEIVED TRULIANT SYRINGE PIN PULLERS. THERE WAS NO COMPLAINT FROM A FACILITY OR AGENT ABOUT THIS DEVICE. NO ADDITIONAL INFORMATION IS AVAILABLE. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. FINDINGS: ONE OR MORE OF THE METAL PRONGS USED TO CLAMP A PIN HEAD ARE MISSING FROM THE RECEIVED TRULIANT SYRINGE PIN PULLERS. WHEN ANALYZING THE FRACTURE SURFACE AS PART OF ANOTHER COMPLAINT, THE FAILURE APPEARS CONSISTENT WITH BRITTLE FRACTURE. THIS TYPE OF FAILURE IS MOST LIKELY THE RESULT OF THE PRONGS EXPERIENCING EXCESSIVE FORCES. FURTHERMORE, IF THE PIN PULLER WAS USED OFF-AXIS RELATIVE TO THE BONE PIN, AN UNEVEN FORCE DISTRIBUTION MAY HAVE BEEN APPLIED TO THE METAL PRONGS AND COULD EXPLAIN WHY THEY FRACTURED FROM ONLY ONE SIDE. OVER TIME, THESE EXCESSIVE OFF-AXIS FORCES COULD POTENTIALLY LEAD TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE MATERIAL. FURTHER INVESTIGATION WILL OCCUR AS PART OF (B)(4). MOST LIKELY UNDERLYING CAUSE: THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF THE PRONGS OF THE PIN PULLER EXPERIENCING EXCESSIVE FORCES DURING THE EXTRACTION OF BONE PINS. RMR REVIEW: A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THESE TRULIANT SYRINGE PIN PULLERS, 750-2014-006-RMR-INSTRUMENTS REV C, WAS REVIEWED. THE RISK IS CAPTURED. IFU 700-096-181: INSTRUMENT INSPECTION . VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00528 | 2022-05-05 | 2022-04-05 | I, n, j, u, r, y | Manufacturer report | NV GXL LNR, LIPPED 32MM ID GROUP 2 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LNR, LIPPED 32MM ID GROUP 2 CUPS | 132-32-52 |  | AS REPORTED, APPROXIMATELY NINE YEARS POST INITIAL LEFT THA, THIS MALE PATIENT HAD A REVISION DUE TO OSTEOLYSIS CAUSED BY WEAR OF NOVATION CROWN CUP WITH GXL POLY. PATIENT WAS REVISED TO A Z/B CUSTOM TRI-FLANGE CUP. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT LEFT OPERATING ROOM (OR) IN STABLE CONDITION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PHOTOS AND X-RAYS PROVIDED. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO HSS KEEPS ALL RETRIEVALS. | CONCOMITANT MEDICAL PRODUCTS: NV GXL LNR, LIPPED 32MM ID GROUP 2 CUPS (CAT# 132-32-52 / (B)(4)); 12/14 BIOLOX FEMORAL HD 32MM +0 (CAT# 140-32-00 / (B)(4)); NV SPLINED HA RDD S/O SZ 13 (CAT# 160-32-13 / (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10110 | 2022-05-05 | 2018-07-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LOCKING FEMORAL IMPACTOR 0-5 | EXACTECH, INC. |  | 213-64-01 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE IMPACTOR HANDLE BROKE. WHILE THE SURGEON WAS IMPACTING THE FEMORAL TRIAL, THE PLASTIC PIECE OF THE IMPACTOR BROKE OFF THE FEMORAL IMPACTOR. THE SURGERY WENT FINE THE IMPACTORS WERE SWITCHED OUT. THE PATIENT'S HEALTH WAS STABLE LEAVING THE OR. AGENT WAS PRESENT AT TIME OF SURGERY. THERE WAS A BRIEF DELAY TO SURGERY WHILE THE IMPACTOR WAS REPLACED, NO ADVERSE EVENT TO PATIENT. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED FOR ANALYSIS. ENG EVAL COMPLETED ON 22 MAR 2019 BY MB. THE BROKEN IMPACTION PLATE APPEARS CONSISTENT WITH CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE IMPACTION PLATE. THE BROKEN WELD APPEARS CONSISTENT WITH FORCIBLY DETACHING THE HANDLE FROM THE IMPACTOR BODY, LIKELY DURING CLEANING, WHICH BROKE THE WELD AT THE JUNCTION BETWEEN THE TWO COMPONENTS. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE AND CONSISTENT WITH YEARS OF USE. DESIGN-RELATED ISSUES: THE DESIGN OF THE LPI LOCKING FEMORAL IMPACTOR HAS BEEN IN THE FIELD SINCE 2005. EXACTECH IS AWARE OF 16 SIMILAR COMPLAINT REPORTS INVOLVING 17 FEMORAL IMPACTOR HANDLES SINCE 2008. THE FREQUENCY OF OCCURRENCE RANKING SCALE IS "VERY LOW". THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS AWARE OF 2 OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 52 UNITS, WHICH HAS BEEN IN THE FIELD SINCE 2013. THE DEVICE HISTORY RECORD FOR THIS LOCKING FEMORAL IMPACTOR WAS REVIEWED. ONE PART WAS REJECTED FROM THIS LOT FOR AN UNRELATED ISSUE. ALL OTHER PARTS WERE ACCEPTED WITH CONFORMANCE TO THE DEVICE SPECIFICATIONS, AND THE MATERIAL CERTIFICATION FOR THE RADEL IMPACTION PLATE WAS REVIEWED. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING-RELATED. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. THE BROKEN IMPACTOR FACE REPORTED IN EXPERIENCE C2018-546 WAS LIKELY THE RESULT OF MATERIAL DEGRADATION ASSOCIATED WITH YEARS OF NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. THE BROKEN WELD WAS LIKELY THE RESULT OF UNAUTHORIZED DISASSEMBLY OF THE DEVICE DURING CLEANING. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. THERE WAS A DELAY TO SURGERY WITH NO ADVERSE EVENT TO THE PATIENT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN IMPACTOR FACE REPORTED WAS LIKELY THE RESULT OF MATERIAL DEGRADATION ASSOCIATED WITH YEARS OF NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. THE BROKEN WELD WAS LIKELY THE RESULT OF UNAUTHORIZED DISASSEMBLY OF THE DEVICE DURING CLEANING. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10101 | 2022-05-05 | 2016-12-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | ADJUSTABLE PATELLA CUTTING GUIDE, SLOTTED | EXACTECH, INC. |  | 213-68-00 |  | IT WAS REPORTED FROM OUS THAT AFTER AN ORTHOPEDIC SURGICAL PROCEDURE THE LOCKING NUT OF PATELLA CUTTING HANDLE WAS NOTICED TO BE LOOSE. THE AGENT PRESENT AT THE TIME OF SURGERY. THE SURGEON USED ANOTHER PATELLA CUTTING GUIDE. IT WAS CONFIRMED THAT THE DEVICE DID NOT BREAK DURING SURGERY AND NO PIECE WAS LEFT IN THE PATIENT DUE TO THE DEVICE HAVING A LOOSE LOCKING NUT. THERE WAS NO INJURY/CLINICAL CONSEQUENCE/ADVERSE EVENT TO THE PATIENT BECAUSE OF THIS EVENT. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED FOR ANALYSIS. ONE OF THE SCREWS APPEARS TO HAVE DISASSEMBLED FROM THE INSTRUMENT. ACCORDING TO THE REPORT THE SCREW WAS LOOSE. HOWEVER, IT HAD COMPLETELY DISASSEMBLED AND WAS NOT RETURNED WITH THE INSTRUMENT WHEN IT WAS PROVIDED FOR EVALUATION. THE SCREW WAS LIKELY MISPLACED DURING TRANSPORTATION/HANDLING. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE. DESIGN: THIS DESIGN HAS BEEN IN THE FIELD SINCE 2011. 3 COMPLAINT REPORTS INVOLVING THIS DEVICE SINCE ITS INTRODUCTION. THE OCCURRENCE RATE OF LESS THAN 0.5% IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MANUFACTURING: THERE HAVE NOT BEEN ANY OTHER COMPLAINTS FROM THIS LOT NUMBER, THEREFORE, IT DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. MOST LIKELY CAUSE: THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF THE LOCKING NUT BECOMING LOOSE AND DISENGAGING FROM THE ASSEMBLY, MAKING THE INSTRUMENT UNUSABLE. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT IMPACT. IT WAS CONFIRMED THAT THE DEVICE DID NOT BREAK DURING SURGERY AND NO PIECE WAS LEFT IN THE PATIENT DUE TO THE DEVICE HAVING A LOOSE LOCKING NUT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF THE LOCKING NUT BECOMING LOOSE AND DISENGAGING FROM THE ASSEMBLY, MAKING THE INSTRUMENT UNUSABLE. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10107 | 2022-05-05 | 2017-09-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | HUMERAL LINER IMPACTOR TIP, 38MM | EXACTECH, INC. |  | 321-07-38 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE HUMERAL LINER IMPACTOR TIPS BROKE. THE REPRESENTATIVE WAS PRESENT AT THE TIME OF SURGERY. IT WAS STATED THAT THERE WERE NO PIECES WERE LEFT IN THE PATIENT. NO DELAY TO SURGERY AND NO ADVERSE EVENT TO PATIENT. THE ISSUE WAS RESOLVED BY USING ANOTHER IMPACTOR FOR THE SURGICAL PROCEDURE. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICES WERE NOT RECEIVED FOR ANALYSIS. ENG ANALYSIS COMPLETED 05/01/2018 BY MB DESIGN: THIS DESIGN HAS BEEN IN THE FIELD SINCE 2006. THERE HAVE BEEN 18 COMPLAINT REPORTS INVOLVING 26 BROKEN HUMERAL LINER IMPACTOR TIPS SINCE 2008.THE OCCURRENCE RATE IS LESS THAN 0.5%, "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THIS DOES NOT APPEAR TO BE DESIGN RELATED. MANUFACTURING: THERE HAVE NOT BEEN ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT, WHICH HAS BEEN IN THE FIELD SINCE 2014. THIS DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE BROKEN IMPACTOR TIPS WERE LIKELY THE RESULT OF A STRESS RISER INTRODUCED DURING IMPACTION, CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICES WERE NOT AVAILABLE FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. IT WAS REPORTED THERE WAS NO SURGICAL DELAY, NO PIECES LEFT IN THE PATIENT AND NO PATIENT ADVERSE EVENT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN IMPACTOR TIPS WERE LIKELY THE RESULT OF A STRESS RISER INTRODUCED DURING IMPACTION, CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICES WERE NOT AVAILABLE FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10103 | 2022-05-05 | 2017-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | OPTETRAK LOCKING FEMORAL IMPACTOR | EXACTECH, INC. |  | 213-64-01 |  | IT WAS REPORTED BY ORTHO-TEAM S.A. THAT THE INSTRUMENT COULD HAVE BEEN BROKEN DURING THE SURGERY. HOWEVER, THE BREAKAGE HAS BEEN NOTICED AT THE RECOVERY OF THE INSTRUMENTATION SET BY OUR LOGISTICS DEPARTMENT AFTER THE CLEANING REALIZED BY THE STERILIZATION TEAM OF THE HOSPITAL. IT IS NOT CONFIRMED THAT THE SALES REPRESENTATIVE WAS PRESENT AT THE SURGERY. THERE WAS NO REPORTED DELAY OF SURGERY AND THERE WAS NO REPORTED PATIENT INJURY, ADVERSE EVENT, OR CLINICAL CONSEQUENCE TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE REPORTER RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS NOT RETUNED TO EXACTECH FOR ANALYSIS. DESIGN-RELATED ISSUES: THE DESIGN OF THE OPTETRAK LOCKING FEMORAL IMPACTOR HAS BEEN IN THE FIELD SINCE 2005. EXACTECH HAS RECEIVED 20 COMPLAINTS INVOLVING THIS DEVICE DUE TO BEING BROKEN SINCE 2008. BECAUSE THIS DEVICE IS USED IN TOTAL KNEE SURGERIES, SALES DATA FOR FEMORAL COMPONENTS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH HAS NOT RECEIVED ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT, WHICH HAS BEEN IN THE FIELD SINCE 2011. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING-RELATED. RISK MANAGEMENT: A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS FEMORAL IMPACTOR, (B)(4) REV-, WAS REVIEWED. THE RISK IS CAPTURED IN ROW 18. NO CORRECTIVE ACTION IS REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE BROKEN FEMORAL IMPACTOR REPORTED WAS LIKELY THE RESULT OF YEARS OF NORMAL USE WHICH POSSIBLY LED TO FRACTURE OF THE IMPACTION SURFACE. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. EXACTECH IFU 700-096-004 STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED INSTRUMENT MALFUNCTION IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO ANY PATIENT IMPACT. AN INVESTIGATION WAS CONDUCTED THAT THE MOST LIKELY ROOT CAUSE OF THE DEVICE BREAKAGE REPORTED IN THIS EVENT WAS LIKELY DUE TO YEARS OF NORMAL USE WHICH POSSIBLY LED TO FRACTURE OF THE IMPACTION SURFACE. ATTEMPTS HAVE BEEN MADE FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. NO ADDITIONAL ATTEMPTS WILL BE MADE FOR INFORMATION REGARDING THIS EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10100 | 2022-05-05 | 2016-11-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LINER DRIVER HEAD, 36MM | EXACTECH, INC. |  | 141-01-36 |  | IT WAS REPORTED FROM THE US THAT DURING A SURGICAL PROCEDURE THE LINER DRIVER BROKE IN HALF. ALL PIECES WERE ACCOUNTED FOR. THERE WAS NO DELAY TO SURGERY AND PATIENT WAS DOING WELL AFTER SURGERY. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED FOR ANALYSIS. VISUAL EVALUATION CONDUCTED WITH A 7X OR 10X OPTICAL-THE DEVICE WAS RETURNED IN TWO PIECES, THE FRACTURE SURFACES INDICATE A BRITTLE FRACTURE THAT LIKELY INITIATED FROM THE CENTER OF THE DEVICE. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE AND CONSISTENT WITH NORMAL USE. MOST LIKELY CAUSE: THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF A BRITTLE FRACTURE INITIATING FROM THE CENTER OF THE DEVICE WHERE THE TIP OF THE LINER DRIVER HANDLE MEETS THE LINER DRIVER HEAD. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST ; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT IMPACT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF A BRITTLE FRACTURE INITIATING FROM THE CENTER OF THE DEVICE WHERE THE TIP OF THE LINER DRIVER HANDLE MEETS THE LINER DRIVER HEAD. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10111 | 2022-05-05 | 2018-08-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LOCKING FEMORAL IMPACTOR 0-5 | EXACTECH, INC. |  | 213-64-01 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE INSTRUMENT BROKE DURING IMPACTION. THERE WAS NO DELAY IN SURGERY, NO PIECES WERE LEFT IN THE PATIENT. THE PATIENT WAS STABLE LEAVING THE OR. NO PATIENT DEMOGRAPHIC INFORMATION WILL BE PROVIDED. NEW INFORMATION RECEIVED 25 MAR 2021; THE INSTRUMENT COULD HAVE BEEN BROKEN DURING THE SURGERY. HOWEVER, THE BREAKAGE HAS BEEN NOTICED AT THE RECOVERY OF THE INSTRUMENTATION-SET BY OUR LOGISTICS DEPT., AFTER THE CLEANING REALIZED BY THE STERILIZATION TEAM OF THE HOSPITAL. NO INFORMATION HAS BEEN RECEIVED FROM THE HOSPITAL REGARDING THIS ISSUE AND [THE AGENCY] DIDN'T RECEIVE ANY REQUEST FOR COMPLEMENTARY INFORMATION REGARDING THIS ISSUE AS WELL. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED. ENG EVAL COMPLETED ON 11/29/2018 BY (B)(4). THE SCREW HOLE SHOWS THE RETAINED END OF THE HEX SCREW AND THE PIN THAT IS INTENDED TO HOLD THE SCREWS -IN PLACE. THE BROKEN HEX SCREW APPEARS TO HAVE FAILED AT THE PIN-SCREW INTERSECTION. THE LOCATION OF THE PIN LIKELY IMPACTED THE MECHANICAL INTEGRITY OF THE SCREW. THE DETACHED HEAD OF THE HEX SCREW WAS NOT RETURNED FOR EVALUATION. THE DEFORMATION ON THE HEAD OF THE HEX SCREW APPEARS CONSISTENT WITH ATTEMPTING TO REMOVE THE SCREW FROM THE IMPACTOR BODY. THE SCREWS ARE NOT INTENDED TO BE REMOVED FROM THE IMPACTOR BODY. THE MISSING IMPACTOR FACE APPEARS LIKELY TO HAVE FRACTURED DURING IMPACTION, HOWEVER, WAS NOT RETURNED FOR EVALUATION. THE SALES REPRESENTATIVE REPORTED THE INSTRUMENT BROKEN DURING IMPACTION AND NO PIECES WERE LEFT IN THE PATIENT. THE BROKEN WELD APPEARS CONSISTENT WITH UNAUTHORIZED DI ASSEMBLY OF THE DEVICE DURING CLEANING. THE COMPRESSION SPRING REMAINED INTACT AND WAS UNAFFECTED BY THE BREAKAGE. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE AND CONSISTENT WITH YEARS OF USE. DESIGN-RELATED ISSUES: THE DESIGN OF THE LPI LOCKING FEMORAL IMPACTOR HAS BEEN IN THE FIELD SINCE 2005. EXACTECH IS AWARE OF (B)(4) OTHER SIMILAR COMPLAINT REPORTS INVOLVING THE LPI LOCKING FEMORAL IMPACTOR SINCE 2008. BECAUSE THE LPI LOCKING FEMORAL IMPACTOR IS USED IN ALL TOTAL KNEE REPLACEMENT SURGERIES, (B)(4). THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THE DESIGN OF THE DEVICE HAS SINCE BEEN UPDATED AND THE LOCATION OF THE PIN HAS BEEN MOVED TO PREVENT THE SCREW BREAKAGE SHOWN ABOVE. REFER TO (B)(4), AND CAPA(B)(4) FOR FURTHER INFORMATION. MFG-RELATED ISSUES: THIS ISSUE HAS BEEN ADDRESSED IN CAPA(B)(4) AND THE DESIGN HAS BEEN UPDATED. THIS DOES NOT APPEAR TO BE MANUFACTURING-RELATED. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RISK MANAGEMENT REPORT. THE FRACTURED HEX SCREW REPORTED WAS LIKELY THE RESULT OF THE LOCATION OF THE PIN-SCREW INTERSECTION, WHICH IMPACTED THE MECHANICAL INTEGRITY OF THE SCREW, ALLOWING FOR THE SCREW BREAKAGE. THE DESIGN OF THE DEVICE HAS SINCE BEEN UPDATED. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. IT DID NOT LEAD TO PATIENT ADVERSE EVENT. THE MISSING IMPACTOR FACE APPEARS LIKELY TO HAVE FRACTURED DURING IMPACTION. AN INVESTIGATION WAS CONDUCTED; THE FRACTURED HEX SCREW REPORTED WAS LIKELY THE RESULT OF THE LOCATION OF THE PIN-SCREW INTERSECTION, WHICH IMPACTED THE MECHANICAL INTEGRITY OF THE SCREW, ALLOWING FOR THE SCREW BREAKAGE. THE DESIGN OF THE DEVICE HAS SINCE BEEN UPDATED. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10105 | 2022-05-05 | 2017-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | COMBO TRAY ADJUSTABLE TIBIAL TAMP GUIDE | EXACTECH, INC. |  | 213-75-00 |  | IT WAS REPORTED FROM OUS THAT DURING TRANSPORTATION OF THE INSTRUMENT FROM THE HOSPITAL IT WAS BROKEN, NO PATIENT INVOLVEMENT. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED FOR ANALYSIS. COMPLETED ON 07/06/2018 BY (B)(4). THE GUIDE BUTTON AND SPRING APPEAR TO BE MISSING FROM THE RETURNED ASSEMBLY. THE BUTTON AND SPRING WERE NOT RETURNED. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE AND CONSISTENT WITH YEARS OF USE. DESIGN-RELATED ISSUES: THE DESIGN OF THE TIBIAL TAMP GUIDE HAS BEEN IN THE FIELD SINCE 2010. EXACTECH HAS RECEIVED 3 SIMILAR COMPLAINT REPORTS SINCE 2010. BECAUSE THE TIBIAL TAMP GUIDE IS USED IN TOTAL KNEE SURGERIES, SALES DATA FOR TIBIAL TRAYS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4). THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH HAS NOT RECEIVED ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT, WHICH HAS BEEN IN THE FIELD SINCE 2014. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RISK IS CAPTURED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE BROKEN INSTRUMENT REPORTED IN EXPERIENCE (B)(6) WAS LIKELY THE RESULT OF CRACK INITIATION AND PROPAGATION ON THE GUIDE BUTTON WHICH LED TO ITS ULTIMATE FRACTURE AND DISASSOCIATION (ALONG WITH THE COMPRESSION SPRING) FROM THE ASSEMBLY. HOWEVER, THIS CANNOT BE CONFIRMED AS THE BUTTON AND SPRING WERE NOT AVAILABLE FOR EVALUATION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. NO PATIENT INVOLVEMENT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF CRACK INITIATION AND PROPAGATION ON THE GUIDE BUTTON WHICH LED TO ITS ULTIMATE FRACTURE AND DISASSOCIATION (ALONG WITH THE COMPRESSION SPRING) FROM THE ASSEMBLY. HOWEVER, THIS CANNOT BE CONFIRMED AS THE BUTTON AND SPRING WERE NOT AVAILABLE FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10109 | 2022-05-05 |  | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | WEDGE BROACH/RASP, SZ 8 | EXACTECH, INC. |  | 189-02-08 |  | IT WAS REPORTED FROM THE US THAT THE WEDGE BROACH/RASP, SZ 8 WAS BROKEN. MULTIPLE ATTEMPTS HAVE BEEN MADE FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED FOR ANALYSIS. COMPLETED ON (B)(6) 2018. BY SH. VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR 10X OPTICAL UNLESS OTHERWISE SPECIFIED BM1 MARKINGS WERE FOUND THAT APPEAR CONSISTENT WITH HITTING THE BROACH HANDLE IMPACT PLATE FROM BELOW TO AID IN REMOVING THE BROACH. THE IMPACT MARKS ON THE STEM OF THE HANDLE INDICATE SOME OF THE FORCE FROM THE MALLET IMPACT MAY HAVE BEEN DIRECTED OFF AXIS. ALL OTHER ASPECTS OF THE BROACH HANDLE APPEAR UNREMARKABLE AND CONSISTENT WITH YEARS OF USE. TOP OF THE BROACH SHOWING SCRATCHES AND MATERIAL DEFORMATION MOST LIKELY CREATED WHILE ATTEMPTING TO REMOVE THE DEVICE FROM THE FEMORAL CANAL (LEFT) AND THE FRACTURE SURFACE (RIGHT). THE BROACH WAS RETURNED IN 2 PIECES. SCRATCHES AND MATERIAL DEFORMATION ON THE DEVICE APPEAR CONSISTENT WITH TOOL MARKS MADE WHILE ATTEMPTING TO REMOVE THE BROACH. ADDITIONALLY, THE FRACTURE SURFACE WHERE THE BROACH POST ATTACHES TO THE BROACH. THE LACK OF PLASTIC DEFORMATION INDICATES A BRITTLE ULTIMATE FRACTURE. MATERIAL DEFORMATION ON SIDE OF BROACH MOST LIKELY CREATED WHILE ATTEMPTING TO REMOVE THE DEVICE FROM THE FEMORAL CANAL. THE MATERIAL DEFORMATION AND SCRATCHES ON THE SIDE OF THE BROACH APPEAR CONSISTENT WITH TOOL MARKINGS CREATED WHILE ATTEMPTING TO REMOVE THE DEVICE FROM THE FEMORAL CANAL. [BM2] [BM3] ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE AND CONSISTENT WITH YEARS OF USE. DESIGN-RELATED ISSUES: THE DESIGN OF THE TAPERED WEDGE BROACH, SIZE 8 HAS BEEN IN THE FIELD SINCE 2014. EXACTECH IS NOT AWARE OF ANY SIMILAR COMPLAINT REPORTS INVOLVING THE TAPERED WEDGE BROACH SINCE 2014. THE TAPERED WEDGE BROACH IS USED IN TOTAL HIP SURGERIES WHERE TAPERED WEDGE FEMORAL STEMS ARE IMPLANTED. THEREFORE, SALES DATA FOR THE TAPERED WEDGE FEMORAL STEM WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 10 UNITS, WHICH HAVE BEEN IN THE FIELD SINCE 2014. THEREFORE; THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS DEVICE, 750-2011-010-RMR-INSTRUMENTS REV B, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 12. CORRECTIVE ACTION IS NOT REQUIRED, THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE BROKEN DEVICE REPORTED IN EXPERIENCE C2018-365 WAS LIKELY THE RESULT OF EXCESSIVE FORCES APPLIED OFF AXIS TO THE BROACH POST, CAUSING CRACK INITIATION AND PROPAGATION AT THE BASE OF THE BROACH POST LEADING TO ULTIMATE FRACTURE. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES. THERE WAS ON REPORTED PATIENT ADVERSE EVENT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF EXCESSIVE FORCES APPLIED OFF AXIS TO THE BROACH POST, CAUSING CRACK INITIATION AND PROPAGATION AT THE BASE OF THE BROACH POST LEADING TO ULTIMATE FRACTURE. EVENT DATE NOT REPORTED. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10108 | 2022-05-05 | 2017-11-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LPI LOCKING FEMORAL IMPACTOR | EXACTECH, INC. |  | 213-64-01 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGICAL PROCEDURE A BROKEN INSTRUMENT WAS NOTICED. UPON INITIAL USE OF AN IMPACTOR, IT WAS NOTED THAT THE BLACK PAD WAS CRACKED OFF. ALL PIECES ARE ACCOUNTED FOR, NOTHING WAS LEFT IN THE PATIENT. THERE WAS NO ADVERSE EFFECT TO THE PATIENT AND NO SURGICAL DELAY. THE REPRESENTATIVE HAD A SPARE INSTRUMENT SET ON HAND AND USED AN INSTRUMENT FROM THAT SET. THE PATIENT WAS REPORTED AS STABLE UPON LEAVING THE OPERATING ROOM. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. MFR DATE: 08/08/14. THE LPI LOCKING FEMORAL IMPACTOR WAS RECEIVED FOR ANALYSIS. THE REPORTED MALFUNCTION (BLACK PAD IS CRACKED OFF) WAS CONFIRMED DURING ANALYSIS. VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR 10X OPTICAL UNLESS OTHERWISE SPECIFIED) REVEALED A FRACTURED IMPACTOR FACE. THE BROKEN IMPACT FACE APPEARS CONSISTENT WITH MATERIAL DEGRADATION ASSOCIATED WITH YEARS OF NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE AND CONSISTENT WITH YEARS OF USE. DESIGN: THE DESIGN OF THE LPI LOCKING FEMORAL IMPACTOR HAS BEEN IN THE FIELD SINCE 2005. EXACTECH HAS RECEIVED 18 SIMILAR COMPLAINT REPORTS INVOLVING LPI LOCKING FEMORAL IMPACTOR DEVICES SINCE 2014. BECAUSE THE LPI LOCKING FEMORAL IMPACTOR IS USED IN ALL TOTAL KNEE REPLACEMENT SURGERIES, SALES DATA FROM 2014 TO 2018 FOR FEMORAL COMPONENTS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MANUFACTURING: EXACTECH IS NOT AWARE OF RECEIVING ANY COMPLAINT REPORTS INVOLVING ANOTHER PART FROM THIS MANUFACTURING LOT OF 28 PIECES WHICH HAVE BEEN IN THE FIELD SINCE 2014. EXACTECH COMPLAINT DATA FROM 2014 THROUGH 2018 INVOLVING THE REPORTED FAILURE FOR THIS FAMILY OF DEVICES WAS REVIEWED. THE INVESTIGATIONS FOR THOSE INCIDENCES HAVE NOT IDENTIFIED ANY EVIDENCE THAT A MANUFACTURING ISSUE CAUSED OR CONTRIBUTED TO THIS REPORTED ISSUE. THEREFORE, THIS DOES NOT APPEAR TO BE MANUFACTURING-RELATED. BASED ON THE ANALYSIS, THE BROKEN IMPACTOR FACE REPORTED IN EXPERIENCE C2018-253 WAS LIKELY THE RESULT OF MATERIAL DEGRADATION ASSOCIATED WITH YEARS OF NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. AFTER THE DEVICE BROKE, ALL PIECES WERE ACCOUNTED FOR. AN INVESTIGATION WAS CONDUCTED; THE BROKEN IMPACTOR FACE REPORTED WAS LIKELY THE RESULT OF MATERIAL DEGRADATION ASSOCIATED WITH YEARS OF NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10102 | 2022-05-05 | 2017-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | OPTETRAK LOCKING FEMORAL IMPACTOR | EXACTECH, INC. |  | 213-64-01 |  | IT WAS REPORTED BY ORTHO-TEAM S.A. THAT THE INSTRUMENT COULD HAVE BEEN BROKEN DURING THE SURGERY. HOWEVER, THE BREAKAGE HAS BEEN NOTICED AT THE RECOVERY OF THE INSTRUMENTATION SET BY OUR LOGISTICS DEPARTMENT AFTER THE CLEANING REALIZED BY THE STERILIZATION TEAM OF THE HOSPITAL. IT IS NOT CONFIRMED THAT THE SALES REPRESENTATIVE WAS PRESENT AT THE SURGERY. THERE WAS NO REPORTED DELAY OF SURGERY AND THERE WAS NO REPORTED PATIENT INJURY, ADVERSE EVENT, OR CLINICAL CONSEQUENCE TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE REPORTER RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS NOT RETUNED TO EXACTECH FOR ANALYSIS. DESIGN-RELATED ISSUES: THE DESIGN OF THE OPTETRAK LOCKING FEMORAL IMPACTOR HAS BEEN IN THE FIELD SINCE 2005. EXACTECH HAS RECEIVED 20 COMPLAINTS INVOLVING THIS DEVICE DUE TO BEING BROKEN SINCE 2008. BECAUSE THIS DEVICE IS USED IN TOTAL KNEE SURGERIES, SALES DATA FOR FEMORAL COMPONENTS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH HAS NOT RECEIVED ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT, WHICH HAS BEEN IN THE FIELD SINCE 2011. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING-RELATED. RISK MANAGEMENT: A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS FEMORAL IMPACTOR, (B)(4) REV-, WAS REVIEWED. THE RISK IS CAPTURED IN ROW 18. NO CORRECTIVE ACTION IS REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE BROKEN FEMORAL IMPACTOR REPORTED WAS LIKELY THE RESULT OF YEARS OF NORMAL USE WHICH POSSIBLY LED TO FRACTURE OF THE IMPACTION SURFACE. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. EXACTECH IFU 700-096-004 STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR TOTAL KNEE ARTHROPLASTY SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL EXACTECH INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED INSTRUMENT MALFUNCTION IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO ANY PATIENT IMPACT. AN INVESTIGATION WAS CONDUCTED THAT THE MOST LIKELY ROOT CAUSE OF THE DEVICE BREAKAGE REPORTED IN THIS EVENT WAS LIKELY DUE TO YEARS OF NORMAL USE WHICH POSSIBLY LED TO FRACTURE OF THE IMPACTION SURFACE. ATTEMPTS HAVE BEEN MADE FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. NO ADDITIONAL ATTEMPTS WILL BE MADE FOR INFORMATION REGARDING THIS EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10106 | 2022-05-05 | 2017-09-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | HUMERAL LINER IMPACTOR TIP, 42MM | EXACTECH, INC. |  | 321-07-42 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE HUMERAL LINER IMPACTOR TIPS BROKE. THE REPRESENTATIVE WAS PRESENT AT THE TIME OF SURGERY. IT WAS STATED THAT THERE WERE NO PIECES WERE LEFT IN THE PATIENT. NO DELAY TO SURGERY AND NO ADVERSE EVENT TO PATIENT. THE ISSUE WAS RESOLVED BY USING ANOTHER IMPACTOR FOR THE SURGICAL PROCEDURE. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICES WERE NOT RECEIVED FOR ANALYSIS. ENG ANALYSIS COMPLETED 05/01/2018 BY MB DESIGN: THIS DESIGN HAS BEEN IN THE FIELD SINCE 2006. THERE HAVE BEEN 18 COMPLAINT REPORTS INVOLVING 26 BROKEN HUMERAL LINER IMPACTOR TIPS SINCE 2008. THE OCCURRENCE RATE IS LESS THAN 0.5%, "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THIS DOES NOT APPEAR TO BE DESIGN RELATED. MANUFACTURING: THERE HAVE NOT BEEN ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT, WHICH HAS BEEN IN THE FIELD SINCE 2014. THIS DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE BROKEN IMPACTOR TIPS WERE LIKELY THE RESULT OF A STRESS RISER INTRODUCED DURING IMPACTION, CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICES WERE NOT AVAILABLE FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. IT WAS REPORTED THERE WAS NO SURGICAL DELAY, NO PIECES LEFT IN THE PATIENT AND NO PATIENT ADVERSE EVENT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN IMPACTOR TIPS WERE LIKELY THE RESULT OF A STRESS RISER INTRODUCED DURING IMPACTION, CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICES WERE NOT AVAILABLE FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10112 | 2022-05-05 | 2019-11-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | 38MM BLUE IMPACTOR TIP | EXACTECH, INC. |  | 321-07-38 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE IMPACTOR TIP BROKE. WHILE USING THE 38MM IMPACTOR TIP TO IMPACT THE 38MM LINER, THE TIP BROKE IN HALF. THE SURGEON DECIDED TO GO TO THE 42MM OPTION AND OPERATION PROCEEDED WITHOUT FURTHER INCIDENT. THERE WERE NO PARTS/PIECES OF THE BROKEN DEVICE THAT FELL INTO THE PATIENT WOUND. THERE WAS NO SURGICAL DELAY OR ANY ADVERSE EVENT TO THE PATIENT. THERE ARE NO REPORTED RESULTS ABOUT THE PATIENT'S OVERALL OUTCOME. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. ENG EVAL ATTACHED BY MV TO SS ON (B)(6) 2020. ALTHOUGH STILL UNDER THE THRESHOLD OF <0.5% FOR VERY LOW' OCCURRENCE RATE, DUE TO A TREND OF THE RECENT COMPLAINTS HAVING APPROXIMATELY A 5X HIGHER OCCURRENCE RATE IN THE KOREAN MARKET, (B)(4) AND CAPA (B)(4) WERE ESTABLISHED TO INVESTIGATE ADDITIONAL POTENTIAL CAUSES OF THE HIGHER OCCURRENCE RATE SPECIFIC TO THE KOREAN MARKET. PREVIOUSLY, CAPA(B)(4) IDENTIFIED IMPROPER REPROCESSING PROCEDURES AS A ROOT CAUSE FOR A HIGHER INCIDENCE OF FRACTURED RADEL TIBIAL INSERT TRIALS IN KOREA AND SINGAPORE. BECAUSE THE DEVICES INVESTIGATED IN CAPA (B)(4) ARE MADE FROM THE SAME TYPE OF MATERIAL AS THE HUMERAL LINER IMPACTOR TIPS AND THE GEOGRAPHIC LOCATION OF THE FAILURES IS THE SAME, CAPA (B)(4) IS INVESTIGATING THE POSSIBILITY OF A SIMILAR CORRELATION BETWEEN THE HIGHER INCIDENCE OF FRACTURES AND THE REPROCESSING PROCEDURE FOR THE IMPACTOR TIPS. IF THE BROKEN PIECE IS LEFT IN THE PATIENT'S WOUND WITHOUT BEING NOTICED, IT COULD CAUSE INJURY TO THE PATIENT AND WOULD REQUIRE PROFESSIONAL MEDICAL INTERVENTION. THE MATERIAL OF THESE INSTRUMENTS, RADEL PER ETS-0205, IS INCLUDED IN THE BIOCOMPATIBILITY ASSESSMENT TECHNICAL MEMO, TM-2013-465. RADEL PASSES THE BIOCOMPATIBILITY REQUIREMENTS FOR SURGICAL USE, AND THERE ARE NO ANTICIPATED ADVERSE HEALTH CONSEQUENCES. THE NEED FOR ANY CORRECTION ACTIONS WILL BE DETERMINED BY CAPA(B)(4). MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: BASED ON INFORMATION AVAILABLE AT THIS TIME, NO CONCLUSION CAN BE MADE REGARDING THE HIGHER MARKET SPECIFIC OCCURRENCE. ADDITIONAL INVESTIGATION FOR A MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE RELATED TO MARKET SPECIFIC OCCURRENCE IS STILL ONGOING. REFER TO CAPA(B)(4) FOR ADDITIONAL INFORMATION. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES. NONE OF THE BROKEN PARTS/PIECES FELL INTO THE SURGICAL SITE AND THERE WAS NO ADVERSE EVENT TO THE PATIENT. AN INVESTIGATION WAS CONDUCTED UNDER CAPA2020-01; BASED ON INFORMATION AVAILABLE AT THIS TIME, NO CONCLUSION CAN BE MADE REGARDING THE HIGHER MARKET SPECIFIC OCCURRENCE. ADDITIONAL INVESTIGATION FOR A MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE RELATED TO MARKET SPECIFIC OCCURRENCE IS STILL ONGOING. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00527 | 2022-05-05 | 2022-01-21 | I, n, j, u, r, y | Manufacturer report | THREE PEG PATELLA 32MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED, | EXACTECH, INC. | THREE PEG PATELLA 32MM | 200-02-32 |  | IT WAS REPORTED FROM AN EXACTECH EGPS LEVEL 1 CLINICAL STUDY, THAT A MALE PATIENT,(B)(6) IN WEIGHT, WITH A CONFIRMED DIAGNOSIS OF PRIMARY OSTEOARTHRITIS, WHO UNDERWENT AN INITIAL KNEE PROCEDURE ON (B)(6) 2021, REPORTED INSTABILITY ON THEIR 6 MONTH FOLLOW UP VISIT. LATERAL INSTABILITY OF THE PATELLA DUE TO ARTHROSCOPY FAILURE WAS DETERMINED REQUIRING ADDITIONAL SURGERY. THE STUDY INDICATES IT IS POSSIBLY RELATED TO THE PROCEDURE AND POSSIBLY RELATED TO THE DEVICES. THE STUDY INDICATES THE OUTCOME AS RESOLVED. | H6: BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN OR MANUFACTURING ISSUES. THE CAUSE OF INSTABILITY AND SUBSEQUENT REVISION CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION AS ASSOCIATED WITH THE INTERACTION BETWEEN THE IMPLANTED DEVICE AND THE PATIENT DUE TO PATIENT ILLNESS, UNIQUE ANATOMY, OR OTHER CONDITION THAT IMPACTS THE PERFORMANCE OF THE DEVICE. THESE DEVICES ARE USED FOR TREATMENT NOT DIAGNOSIS. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00509 | 2022-05-05 | 2019-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRULIANT TIB TAMP GUIDE | EXACTECH, INC. |  | 02-029-22-2000 |  | IT WAS REPORTED FROM THE US THAT THE TIBIA PUNCH TOWER 02-029-22-2000 FUNCTIONED DURING SURGERY BUT AFTER SPD CLEANING THE RELEASE BUTTON WAS FOUND MINUS THE SPRING. THE BUTTON CAN COME LOOSE AND POP OUT DURING USE. THE DEVICE WAS RECEIVED FOR ANALYSIS. THE REPRESENTATIVE WAS PRESENT AT THE TIME OF SURGERY. THERE IS NO ADDITIONAL INFORMATION FOR THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. FINDINGS: THE BUTTON COMPONENT (02-029-222-2005) OF THE TIBIAL TAMP GUIDE DISENGAGED FROM THE INSTRUMENT (02-029-22-2000). ENGINEERING PRINTS OF 02-029-122-2000 WERE REVIEWED. THE OVERLAP BETWEEN THE SNAP RING¿S OD AND THE ID OF THE HOUSING LIP IS .008 FOR THE WORST-CASE SCENARIO. ALTHOUGH THE OVERLAP IS POSITIVE, THE SMALL AMOUNT COULD CAUSE DISENGAGEMENT. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: THE ROOT CAUSE FOR THE BUTTON DISENGAGEMENT, AS DETERMINED BY CAPA(B)(4), WAS LIKELY THE RESULT OF INSUFFICIENT OVERLAP BETWEEN THE SNAP RING AND THE LIP OF THE HOUSING TO PREVENT THE OCCURRENCE OF DISENGAGEMENT. CORRECTIVE ACTION TAKEN: PER (B)(4), THE COMMITTEE HAS DETERMINED THAT THE DESIGN WILL BE UPDATED, TESTED, AND IMPLEMENTED, AFTER WHICH A STAGED MARKET WITHDRAWAL WILL BE CONDUCTED. IN THE MEANTIME, THE EXACTECH TEAM WILL CONTINUE TO MONITOR THIS INSTRUMENT FOR THE ISSUE. CAPA(B)(4) IS STILL ONGOING AT THE TIME OF THIS EVALUATION, AND CORRECTIVE ACTIONS ARE SUBJECT TO CHANGE. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. THERE IS NO REPORTED PATIENT INVOLVEMENT. THE BUTTON CAN DISENGAGE DURING USE. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00526 | 2022-05-05 | 2022-04-13 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 42MM HUMERAL LINER +0 | EXACTECH, INC. | 320-42-00 | 320-42-00 |  | AS REPORTED, APPROXIMATELY 2 WEEKS POST OP THE INITIAL LEFT TSA, THIS (B)(6) MALE PATIENT WAS REVISED DUE TO DISLOCATION. PATIENT CLAIMS THEY WERE GOING PT AND "SOMETHING DIDN'T FEEL RIGHT". PATIENT DISLOCATED AT AND UNKNOWN TIME, POSSIBLY DURING PT. UPON XRAY PATIENT WAS NOTED TO HAVE LOOSENED HUMERAL COMPONENTS. HUMERAL COMPONENTS WERE REMOVED AND REVISED TO IMPLANTS LISTED BELOW. PATIENT DEEMED STABLE. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETURNING DUE TO HOSPITAL POLICY. | PENDING EVALUATION. CONCOMITANT DEVICE(S): 300-30-06, 7075177 - EQ PRESERVE STEM, 320-10-00, 7265032 - EQ REV ADAPTER PLATE TRAY +0, 320-20-00, 7274300 - EQ REV TORQUE SCREW KIT. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00510 | 2022-05-05 | 2019-05-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | A+ STERN INSERTER | EXACTECH, INC. |  | 167-00-00 |  | IT WAS REPORTED FROM THE US THAT THE TIP OF THE STEM INSERTER BROKE OFF WHILE IMPACTING THE STEM. THERE IS NO INDICATION THAT PIECES FELL INTO THE SURGICAL SITE. ¿NO PIECES WERE LEFT IN THE PATIENT [;] HOWEVER [,] WE BELIEVE THE TIP WAS THROWN AWAY BY CENTRAL STERILE STAFF¿. NO DELAY TO SURGERY, NO ADVERSE EFFECTS TO THE PATIENT. PATIENT LEFT THE OR IN STABLE CONDITION. THE DEVICE WAS RECEIVED ON RG466324. THE REPRESENTATIVE WAS NOT PRESENT AT THE TIME OF SURGERY. ATTEMPTS WERE MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REPRESENTATIVE, BUT NO RESPONSES WERE RECEIVED. THERE HAS BEEN NO OTHER INFORMATION PROVIDED BY THE CONTACTS ABOUT THE EVENT OR PATIENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PATIENT INFORMATION. DESIGN-RELATED ISSUES: THE DESIGN OF THE NOVATION A+ STEM INSERTER HAS BEEN IN THE FIELD SINCE 2009. EXACTECH IS AWARE OF ONE SIMILAR COMPLAINT REPORT INVOLVING NOVATION A+ STEM INSERTERS DUE TO FRACTURED TIPS SINCE 2009. BECAUSE THE NOVATION A+ STEM INSERTER IS USED IN TOTAL HIP SURGERIES, SALES DATA FOR FEMORAL STEMS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 15 UNITS, WHICH HAVE BEEN IN THE FIELD SINCE 2010. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS NOVATION A+ STEM INSERTER, 750-2008-007-RMR REV C, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 9. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR. THE FRACTURED STEM INSERTER REPORTED WAS LIKELY THE RESULT OF LONG-TERM, RIGOROUS USE, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE STEM INSERTER TIP. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. ¿ CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. THERE IS NO INDICATION THAT PIECES FELL INTO THE SURGICAL SITE. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00511 | 2022-05-05 | 2019-05-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | MDLR PRM RMR LG PLT TIP | EXACTECH, INC. |  | 315-25-14 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE INSTRUMENT. THE ANATOMICAL (PURPLE) REAMER WAS BROKEN WHILE REAMING AT A 4 DEG POSTERIOR GPS CASE. THE PATIENT HAD VERY ARTHRITIC HARD BONE AND THE REAMER BROKE. THERE WERE NO PARTS/PIECES THAT FELL INTO THE SURGICAL SITE. THE REPRESENTATIVE WAS PRESENT DURING THE CASE. THERE WAS NO DELAY OR PATIENT ADVERSE EVENT. THE OVERALL PATIENT OUTCOME WAS NOTED AS GOOD RESULTS. THERE WAS NO PATIENT DEMOGRAPHIC INFORMATION PROVIDED. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED BY THE CONTACTS FOR THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. VISUAL EVALUATION CONDITIONS/FINDINGS: THE FRACTURED BLADES APPEAR CONSISTENT WITH BRITTLE FRACTURE LIKELY CAUSE BY REAMING OFF-AXIS INTO ¿VERY ARTHRITIC HARD BONE.¿ ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE AND CONSISTENT WITH YEARS OF USE DESIGN-RELATED ISSUES: THE DESIGN OF THE LARGE PRIMARY PILOT TIP REAMER HAS BEEN IN THE FIELD SINCE 2009. EXACTECH IS AWARE OF 11 SIMILAR COMPLAINT REPORTS INVOLVING PRIMARY PILOT TIP REAMERS DUE TO FRACTURED BLADES SINCE 2009. BECAUSE THE PRIMARY PILOT TIP REAMER IS USED IN ANATOMIC TOTAL SHOULDER SURGERIES, SALES DATA FOR ANATOMIC GLENOID COMPONENTS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4). THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 60 UNITS, WHICH HAVE BEEN IN THE FIELD SINCE 2017. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS LARGE PRIMARY PILOT TIP REAMER, 750-2007-026-RMR-INSTRUMENTS REV L, WAS REVIEWED. THE RISK IS CAPTURED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE BROKEN REAMER REPORTED WAS LIKELY THE RESULT OF REAMING OFF-AXIS UNDER HIGH LOADS INTO ¿VERY ARTHRITIC HARD BONE,¿ LEADING TO FRACTURE OF THE BLADES. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES. NO PATIENT ADVERSE EVENT REPORTED, NO PARTS/PIECES FELL INTO THE SURGICAL SITE | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00525 | 2022-05-05 | 2022-04-13 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 42MM HUMERAL LINER +0 | EXACTECH, INC. | 320-42-00 | 320-42-00 |  | AS REPORTED, APPROXIMATELY 1 MONTH POST OP THE INITIAL RIGHT TSA, THIS (B)(6) MALE PATIENT WAS REVISED DUE INFECTION. PATIENT DEVELOPED INFECTION AT WOUND SITE. I&D WAS PERFORMED AND POLY WAS UPSIZED TO 42MM +2.5 CONSTRAINED. LOCKING TORQUE SCREW WAS ALSO REPLACED. PATIENT DEEMED STABLE. THERE AS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETUNING DUE TO HOSPITAL POLICY. | (H3) BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN OR MANUFACTURING ISSUES. THE CAUSE OF SURGICAL REVISION FOR INFECTION CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10104 | 2022-05-05 | 2017-09-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | FEMORAL HEAD IMPACTOR | EXACTECH, INC. |  | 01-001-03-0001 |  | IT WAS REPORTED FROM (B)(6) THAT DURING SURGERY, THE IMPACTORS HEAD WAS BROKEN WHEN THE DOCTOR IMPACTED THE IMPLANT. NO DAMAGE TO THE PATIENT WAS CAUSED BY THIS INCIDENT. THE DOCTOR IMPACTED FIVE TIMES. THE BREAKAGE HAPPENED AT THE LAST IMPACT. THE AGENT WAS PRESENT AT THE SURGERY AND CONFIRMED NO BROKEN PIECES WERE LEFT IN THE PATIENT. THERE WAS NO DELAY TO SURGERY. ATTEMPTS HAVE BEEN MADE FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. NO ADDITIONAL ATTEMPTS WILL BE MADE FOR INFORMATION REGARDING THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED FOR ANALYSIS. ENG EVAL COMPLETED ON (B)(6) 2018 BY AD. DESIGN-RELATED ISSUES: THE DESIGN OF THE FEMORAL HEAD IMPACTOR HAS BEEN IN THE FIELD SINCE 2013. EXACTECH HAS RECEIVED 2 OTHER COMPLAINT REPORTS INVOLVING BREAKAGE OF THIS INSTRUMENT. BECAUSE THE FEMORAL HEAD IMPACTOR IS USED IN TOTAL HIP SURGERIES, SALES DATA FOR FEMORAL HEADS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH HAS NOT RECEIVED ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 25 UNITS, WHICH HAS BEEN IN THE FIELD SINCE 2017. THE VENDOR HAS BEEN NOTIFIED OF THIS ISOLATED EVENT. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS FEMORAL HEAD IMPACTOR, (B)(4) REV G, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 11. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THIS IS AN ISOLATED EVENT, THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. THE VENDOR HAS BEEN NOTIFIED OF THIS ISOLATED EVENT. MOST LIKELY CAUSE: THE BROKEN FEMORAL HEAD IMPACTOR REPORTED IN EXPERIENCE C2017-501 WAS LIKELY THE RESULT OF A MATERIAL DEFECT, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE. THE VENDOR HAS BEEN NOTIFIED OF THIS ISOLATED EVENT. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN FEMORAL HEAD IMPACTOR REPORTED WAS LIKELY THE RESULT OF A MATERIAL DEFECT, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE. THE VENDOR HAS BEEN NOTIFIED OF THIS ISOLATED EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00532 | 2022-05-06 | 2022-04-13 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 2, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 2, 9MM | 02-012-44-2009 |  | AS REPORTED, APPROXIMATELY EIGHT YEARS POST INITIAL LEFT TKA, THIS (B)(6) FEMALE PATIENT HAD A REVISION FOR AN INSERT EXCHANGE OF A SIZE #2 PSC 9MM LINER DUE TO POLY WEAR. PATIENT WAS REVISED TO A SIZE #2 PSC 13MM INSERT. THERE WAS NO BREAKAGE OF A DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PHOTOS AND X-RAYS RECEIVED. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO HSS KEEPS ALL RETRIEVALS. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00347 | 2022-05-06 | 2022-03-14 | I, n, j, u, r, y | Manufacturer report | NV GXL LINER NEUTRAL, 32MM ID GROUP 1 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LINER NEUTRAL, 32MM ID GROUP 1 CUPS | 130-32-51 |  | AS REPORTED, APPROXIMATELY SIX YEARS POST INITIAL RIGHT THA, THIS FEMALE PATIENT WAS BROUGHT BACK FOR RIGHT HIP PAIN. X-RAYS SHOWED SIGNS OF OSTEOLYSIS. THE HIP WAS DISLOCATED, THE FEMORAL HEAD WAS REMOVED, THE STEM WAS CHECKED AND FOUND TO BE WELL FIXED. ATTENTION TURNED TO CUP, THE LINER WAS REMOVED WITH THE HELP OF A CHISEL, THE CUP WAS CHECKED AND FOUND TO BE WELL FIXED. BONE GRAFT WAS PLACED BEHIND THE CUP. A NEW XLE G1 LIPPED 32MM POLY WAS INSERTED ALONG WITH A 32 +0 BIOLOX HEAD. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. X-RAYS AND PHOTOS RECEIVED. THE DEVICE IS NOT AVAILABLE FOR EVALUATION AS THE HOSPITAL KEPT THE DEVICES. | CONCOMITANT MEDICAL PRODUCTS: COCR FEM HEAD 32MM -3.5 OFFSET 12/14 (CAT# 142-32-93 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00348 | 2022-05-06 | 2022-03-15 | I, n, j, u, r, y | Manufacturer report | LOGIC CR TIB INSERT SLOPE++, SZ 5, 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC CR TIB INSERT SLOPE++, SZ 5, 11MM | 02-012-49-5011 |  | AS REPORTED, APPROXIMATELY THIRTEEN YEARS POST INITIAL RIGHT TKA, THIS (B)(6) MALE PATIENT HAD A REVISION DUE TO INCREASE WEAR OF THE POLYETHYLENE INSERT. THERE WAS NO BREAKAGE OF THE DEVICE AND NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. SALES REP WAS UNABLE TO OBTAIN PHOTOS/X-RAYS. THE DEVICE WILL BE RETURN. | CONCOMITANT MEDICAL PRODUCTS THREE PEG PATELLA 41MM (CAT# 200-02-41 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00495 | 2022-05-06 | 2022-02-14 | I, n, j, u, r, y | Manufacturer report | 2.5, 13MM SLOPE ++ CR INSERT | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | 2.5, 13MM SLOPE ++ CR INSERT | UNKNOWN |  | AS REPORTED, A CR INSERT THAT IS DELAMINATED FROM A REVISION CASE THE WEEK OF (B)(6) 2022. DEVICE ENGRAVINGS INDICATE IT IS A SIZE 2.5, 13MM SLOPE++ CR INSERT. NO FURTHER INFORMATION IS KNOWN AT THIS TIME. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00508 | 2022-05-06 | 2022-04-12 | I, n, j, u, r, y | Manufacturer report | TRULIANT CR CEM FEM CR CEM RIGHT SZ 3 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT CR CEM FEM CR CEM RIGHT SZ 3 | 02-020-13-0330 |  | AS REPORTED, APPROXIMATELY 2.5 YEARS POST INITIAL TKA, THIS FEMALE PATIENT WAS REVISED DUE TO PAIN. A MRI AND CT SCAN DID NOT SHOW ANYTHING SIGNIFICANT PER SURGEON BUT SUSPECTED POLY WEAR AND POTENTIALLY AN ISSUE WITH THE TIBIA. HOWEVER DURING SURGERY NO SIGNIFICANT POLY WEAR WAS NOTED AND THE TIBIA WAS NOT LOOSE. THE FEMUR HOWEVER WAS FOUND TO BE DE-BONDED FROM CEMENT MANTLE. IT WAS STILL IN ORIGINAL POSITION BUT WAS EASILY WORKED LOOSE WITH MINIMAL STRIKES WITH A FLEXIBLE OSTEOTOME. THE SURGEON FELT THE NEED TO SWITCH TO COMPETITOR'S DEVICE FOR THE REVISION IMPLANTS. PATIENT LAST KNOWN TO BE IN STABLE CONDITION. NO FURTHER INFORMATION AVAILABLE. IMAGES AND PHOTOS RECEIVED. EXPLANTS NOT AVAILABLE TO THE REP. | CONCOMITANT MEDICAL PRODUCTS: TRULIANT TIB FIT TRAY CEM SZ 3F / 2T (CAT# 02-022-45-3020 / SERIAL# (B)(4)). TRULIANT TIB IMP CRC INSERT SZ 3, 11MM (CAT# 02-022-51-3011 / SERIAL# (B)(4)). THREE PEG PATELLA 32MM (CAT# 200-02-32 / SERIAL# (B)(4)). HOLDING PIN MINI SHARP POINT 4 PK (CAT# 201-78-15 / SERIAL# (B)(4)). THREADED PIN SIZE 2.3 COLLARED 2PK (CAT# 521-78-23 / SERIAL# (B)(4)). THREADED PIN SIZE 3.0 COLLARLESS 2PK (CAT# 521-78-32 / SERIAL# (B)(4)). THREADED PIN SIZE 3.0 COLLARLESS 2PK (CAT# 521-78-32 / SERIAL# (B)(4)). THREADED PIN SIZE 5.1 COLLARLESS 2PK (CAT# 521-78-36 / SERIAL# (B)(4)). GPS IMPLANT KIT V2 (CAT# A10012 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00531 | 2022-05-06 | 2022-04-16 | I, n, j, u, r, y | Manufacturer report | VANTAGE | TIBIAL INSERT FB SZ 3 LT 7MM | EXACTECH, INC. | 350-21-13 | 350-21-13 |  | AS REPORTED, APPROXIMATELY 1.5 YEARS POST OP THE INITIAL LEFT ANKLE REPLACEMENT WITH A CUSTOM-MADE TALUS REPLACEMENT, THIS (B)(6) PATIENT WAS REVISED BECAUSE THE TOTAL TALUS COMPONENT HAS DRIFTED INTO VALGUS. SINCE THE VANTAGE TIBIAL BEARING WAS PART OF THE AFFECTED RECALL, THE SURGEON WILL BE REPLACING THE TIBIAL BEARING. THERE WAS NO SURGICAL DELAY/PROLONGATION. DEVICE WILL BE RETURNING, SURGEON REQUESTED ENGINEERING TO EVALUATE THE DEVICE. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. OF NOTE: THE INITIAL LEFT ANKLE REPLACEMENT FOR THIS PATIENT IS DUE TO PREVIOUS TRAUMA. | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00533 | 2022-05-06 | 2019-04-30 | I, n, j, u, r, y | Manufacturer report | NOVATION GXL LNR +5 LAT 40MM GROUP 3 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NOVATION GXL LNR +5 LAT 40MM GROUP 3 CUPS | 136-40-53 |  | IT WAS REPORTED VIA LEGAL DOCUMENTATION THAT APPROXIMATELY 36 MONTHS AFTER A LEFT TOTAL HIP REPLACEMENT, THE PATIENT UNDERWENT LEFT HIP REVISION TO ADDRESS PROSTHESIS WEAR. A POST OPERATIVE OP REPORT WAS PROVIDED. POST OP DIAGNOSIS NOTED CHRONICALLY FAILED TOTAL HIP ARTHROPLASTY DUE TO POLYETHYLENE WEAR. INTRAOPERATIVELY, THE SURGEON OBSERVED METALLOSIS AND SYNOVITIS AND FRAGMENTED POLYETHYLENE LINER PIECES IN THE JOINT CAUSING THE METALLOSIS. ADDITIONALLY, THE SURGEON OBSERVED A LINEAR SCRATCH ON THE SURFACE OF THE FEMORAL HEAD. BOTH LINER AND FEMORAL HEAD COMPONENTS WERE REPLACED. IT WAS NOTED THAT THE PATIENT'S WEIGHT CONTRIBUTED TO A MORE COMPLEX APPROACH AND POSITIONING DIFFICULTIES, INTRAOPERATIVELY. NO SURGICAL OR MEDICAL INTERVENTIONS WERE REPORTED INFORMATION AND THE DEVICES WERE NOT AVAILABLE FOR EVALUATION. | THIS FOLLOW-UP REPORT IS BEING SUBMITTED TO RELAY ADDITIONAL AND/OR CORRECTED INFORMATION. THIS FOLLOW-UP REPORT IS BEING SUBMITTED TO RELAY ADDITIONAL AND/OR CORRECTED INFORMATION. THE MOST LIKELY UNDERLYING CAUSE FOR THE REVISION REPORTED IS PROSTHESIS WEAR AND FRACTURE AND A COMBINATION OF RISK FACTORS: SUCH AS USE ERROR, IMPLANT POSITIONING, IMPLANT SIZE SELECTION, AND PATIENT FACTORS (FITNESS FOR SURGERY, BIOMECHANICS, ACTIVITY LEVEL AND LOCAL TISSUE OXIDATION POTENTIAL). HOWEVER, THIS CANNOT BE CONFIRMED FROM THE REPORTED INFORMATION AND THE DEVICES WERE NOT AVAILABLE FOR EVALUATION. IF ANY FURTHER INFORMATION IS OBTAINED THAT WOULD CHANGE OR ALTER ANY INFORMATION PROVIDED, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00477 | 2022-05-06 | 2020-01-01 | I, n, j, u, r, y | Manufacturer report | CC TIBIAL INSERT SZ 2, 15MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | CC TIBIAL INSERT SZ 2, 15MM | 208-22-15 |  | AS REPORTED BY LEGAL TEAM, APPROXIMATELY TWO YEARS POST LEFT TKA REVISION, PATIENT RETURNED TO SURGEON BECAUSE HER KNEE DIDN¿T FEEL RIGHT, AND SHE HAD SEVERE PAIN, PROBLEMS WITH HER MOBILITY, AND HAD TO BE IN A WHEELCHAIR SINCE 2013. | H3: AS REPORTED BY LEGAL TEAM, APPROXIMATELY TWO YEARS POST LEFT TKA REVISION, PATIENT RETURNED TO SURGEON BECAUSE HER KNEE DIDN¿T FEEL RIGHT, AND SHE HAD SEVERE PAIN, PROBLEMS WITH HER MOBILITY, AND HAD TO BE IN A WHEELCHAIR SINCE 2013. NO ADDITIONAL INFORMATION. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN ISSUES. THE CAUSE OF THE SUBSEQUENT PAIN CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION. THESE DEVICES ARE USED FOR TREATMENT NOT DIAGNOSIS. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00534 | 2022-05-06 | 2022-04-15 | I, n, j, u, r, y | Manufacturer report | LOGIC FEMORAL PS CEM LEFT SZ 2.5 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC FEMORAL PS CEM LEFT SZ 2.5 | 02-010-01-0225 |  | AS REPORTED, APPROXIMATELY SEVEN YEARS AFTER A LEFT TOTAL KNEE REPLACEMENT, THE PATIENT EXPERIENCED LOOSENING AND UNDERWENT A REVISION PROCEDURE. A REVISION OPERATIVE REPORT WAS PROVIDED. POST OPERATIVE DIAGNOSIS WAS NOTED AS POLYETHYLENE WEAR AND A LOOSE FEMORAL COMPONENT. INTRAOPERATIVELY THE SURGEON OBSERVED OSTEOLYSIS BEHIND THE FEMORAL COMPONENT. THE TIBIAL AND PATELLAR COMPONENTS WERE NOTED AS WELL FIXED. IT WAS NOTED THE PATIENT WAS REVISED TO A COMPETITOR'S OXINIUM CONSTRUCT AS THE SURGEON FEARED THE PATIENT HAD A NICKEL ALLERGY. NO MEDICAL OR SURGICAL INTERVENTIONS WERE REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. | THIS FOLLOW-UP REPORT IS BEING SUBMITTED TO RELAY ADDITIONAL AND/OR CORRECTED INFORMATION. THE REVISION REPORTED WAS LIKELY THE RESULT OF FEMORAL LOOSENING. ADDITIONAL REASONS FOR THE REPORTED REVISION MAY BE PROSTHESIS WEAR, INSTABILITY AND INCLUSION OF THE POLYETHYLENE IN THE PACKAGING RECALL. HOWEVER, THE REPORTED PROSTHESIS WEAR AND INSTABILITY COULD NOT BE CONFIRMED FROM THE PROVIDED INFORMATION. POTENTIAL CONTRIBUTIONS OF USER AND PATIENT-RELATED CONSIDERATIONS TO THE EVENT COULD NOT BE ASSESSED AS THE DEVICES WERE NOT AVAILABLE FOR EVALUATION AND IMAGES WERE NOT PROVIDED. IF ANY FURTHER INFORMATION IS OBTAINED THAT WOULD CHANGE OR ALTER ANY INFORMATION PROVIDED, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00535 | 2022-05-09 | 2019-02-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | NV GXL LINR, NTRL, 36MM ID, GROUP 3 CUPS | EXACTECH, INC. |  | 130-36-53 |  | IT WAS REPORTED FROM THE EXACTECH HIP REPLACEMENT STUDY, THE PATIENT RECEIVED A RIGHT TOTAL HIP REPLACEMENT ON (B)(6) 2010. APPROXIMATELY 8 YEARS LATER, THE PATIENT EXPERIENCED PAIN AFTER A FALL. AN X-RAY REPORTEDLY CONFIRMED THAT THE ¿HARDWARE [WAS] IN GOOD ALIGNMENT.¿ THE PATIENT HAD AN MRI APPROXIMATELY 2 MONTHS LATER, ON (B)(6) 2019, THAT REPORTEDLY IDENTIFIED POLYETHYLENE WEAR. THE CASE REPORT FORM INDICATES EVENT IS RELATED TO THE DEVICE BUT NOT RELATED TO THE PROCEDURE. NO REVISION HAS TAKEN PLACE, THIS STUDY IS CONTINUING. ACCORDING TO THE CLINICAL FORM-NO PREOP DEMOGRAPHIC DATA, THE COMORBIDITIES WERE NOT AVAILABLE. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED ABOUT THIS EVENT OR PATIENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. CONCOMITANT MEDICAL PRODUCTS: ACETABULAR SHELL (CAT# 180-01-56, SN (B)(4)). FEMORAL STEM (CAT# 160-01-16, SN (B)(4)). FEMORAL HEAD (CAT# 142-36-03, SN (B)(4)). BONE SCREW 6.5MM DIA X 30MM LONG (CAT# 120-65-30, SN (B)(4)). BONE SCREW 6.5MM DIA X 35MM LONG (CAT#120-65-35, SN (B)(4)). INVESTIGATION COMPLETED (B)(6) 2020. DESIGN-RELATED ISSUES: EXACTECH IS AWARE OF 80 COMPLAINT REPORTS INVOLVING GXL ACETABULAR LINERS DUE TO PROSTHESIS WEAR AND/OR OSTEOLYSIS SINCE 2008. SALES DATA FOR ALL GXL LINERS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE DETAILED BELOW. A CORRECTION AND REMOVAL COMMITTEE MEETING WAS HELD TO ADDRESS AN INQUIRY FROM A SURGEON REGARDING REVISIONS DUE TO GXL POLYETHYLENE WEAR. A HEALTH RISK INDEX WAS COMPLETED. BASED ON THE ¿VERY LOW¿ OCCURRENCE RATE, THE COMMITTEE DETERMINED THAT NO ACTION WAS NECESSARY. REFER TO (B)(4) FOR ADDITIONAL INFORMATION. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 65 UNITS. THE MANUFACTURING RECORDS FOR THIS DEVICE ONLY EXIST IN PAPER FORM AND ARE STORED IN CARDBOARD BOXES AT IRON MOUNTAIN. DUE TO RISKS ASSOCIATED WITH COVID-19 TRANSMISSION, THE PHYSICAL MANUFACTURING RECORDS FOR THIS COMPLAINT WILL NOT BE REVIEWED. EXACTECH WILL CONTINUE TO MONITOR THIS COMPLAINT TYPE AND, IF WARRANTED, WILL ESCALATE ANY POSITIVE TRENDS TO THE CORRECTIONS AND REMOVAL COMMITTEE (CRC) FOR ADDITIONAL ACTIONS. THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS ACUMATCH GXL LINER, 750-2005-028-RMR-IMPLANTS REV F, WAS REVIEWED. THE RISK IS CAPTURED. THE PROSTHESIS WEAR REPORTED MAY HAVE BEEN THE RESULT OF MISALIGNMENT OF THE ACETABULAR CUP, EDGE LOADING/IMPINGEMENT, OR A COMBINATION OF THE TWO, AS A RESULT OF THE PATIENT¿S FALL, WHICH LED TO POLYETHYLENE WEAR. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICES ARE STILL IMPLANTED; THEREFORE, THE ACETABULAR LINER WAS NOT AVAILABLE FOR EVALUATION. IT IS NOTED IN THE IFU; PHYSICIANS SHOULD CONSIDER FEMORAL AND ACETABULAR COMPONENT PLACEMENT AND MALPOSITION, FEMORAL HEAD COMPONENT DIAMETER AND THE EFFECT OF FEMORAL HEAD COMPONENT ¿SLEEVES AND SKIRTS¿ ON RANGE OF MOTION. A MALPOSITIONED COMPONENT MAY INCREASE THE PROBABILITY FOR IMPINGEMENT, DISASSOCIATION, DISLOCATION, WEAR, AND REVISION SURGERY. IT IS A KNOWN RISK THAT THERE MAY BE SUBSEQUENT REVISIONS OR SURGICAL INTERVENTIONS OF JOINT REPLACEMENT DEVICES. ALSO, THAT WEIGHT AND ACTIVITY COULD INFLUENCE THE LIFETIME OF AN IMPLANT. IT IS KNOWN THAT EXTREME CARE MUST BE USED WITH PATIENTS AFTER JOINT REPLACEMENT SURGERY. ALSO, THAT EXCESSIVE ACTIVITY AND TRAUMA AFFECTING JOINT REPLACEMENTS HAVE BEEN ASSOCIATED WITH PREMATURE FAILURE. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THE REPORTED PATIENT EVENT IS PAIN POLY WEAR WITHOUT A REVISION SURGERY. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00530 | 2022-05-09 | 2022-04-12 | I, n, j, u, r, y | Manufacturer report | PS TIBIAL INSERTS SZ 3, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | PS TIBIAL INSERTS SZ 3, 9MM | 204-23-09 |  | AS REPORTED, APPROXIMATELY TWELVE YEARS POST LEFT TKA, THE (B)(6) FEMALE PATIENT'S POLY FAILED OVER TIME. DURING SURGERY, A PIECE OF THE TRIAL BROKE OFF BUT WAS RETRIEVED. THERE WERE NO PARTS OR PIECES THAT FELL INTO WOUND SITE. IMPLANT WAS REMOVED AND REPLACED WITH 15MM IMPLANT, BUT WAS DEEMED TOO TIGHT SO IT WAS REMOVED AND DOWNSIZED TO A 13MM PS POLY IMPLANT. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT DEEMED STABLE. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PROVIDED IMAGES SHOW THE TIBIAL INSERT HAS FRACTURED. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. THIS IS ONE OF TWO PRODUCTS INVOLVED WITH THE REPORTED EVENT AND THE ASSOCIATED MANUFACTURER REPORT NUMBERS ARE 1038671-2022-00529. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00529 | 2022-05-09 | 2022-04-12 | I, n, j, u, r, y | Manufacturer report | TIBIAL INSERT TRIAL SZ 3 13 MM | ORTHOPEDIC MANUAL SURGICAL INSTRUMENT | EXACTECH, INC. | TIBIAL INSERT TRIAL SZ 3 13 MM | 201-23-13 |  | AS REPORTED, APPROXIMATELY TWELVE YEARS POST LEFT TKA, THE (B)(6) FEMALE PATIENT'S POLY FAILED OVER TIME. DURING SURGERY, A PIECE OF THE TRIAL BROKE OFF BUT WAS RETRIEVED. THERE WERE NO PARTS OR PIECES THAT FELL INTO WOUND SITE. IMPLANT WAS REMOVED AND REPLACED WITH 15MM IMPLANT, BUT WAS DEEMED TOO TIGHT SO IT WAS REMOVED AND DOWNSIZED TO A 13MM PS POLY IMPLANT. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT DEEMED STABLE. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PROVIDED IMAGES SHOW THE TIBIAL INSERT HAS FRACTURED. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. THIS IS ONE OF TWO PRODUCTS INVOLVED WITH THE REPORTED EVENT AND THE ASSOCIATED MANUFACTURER REPORT NUMBERS ARE 1038671-2022-00530. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00536 | 2022-05-09 | 2022-04-08 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 2 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 2 9MM | 02-012-35-2009 |  | AS REPORTED BY THE EXACTECH KNEE REPLACEMENT STUDY; APPROXIMATELY FIVE YEARS POST INITIAL TKA, THIS (B)(6) FEMALE PATIENT EXPERIENCED OSTEOLYSIS OF BONE IN MEDIAL TIBIAL CONDYLE OF RIGHT KNEE. EVENT IS ONGOING. PER CLINICAL STUDY CASE REPORT FORM INDICATES THIS EVENT IS POSSIBLY RELATED TO DEVICES AND DEFINITELY NOT RELATED TO PROCEDURE. | CONCOMITANTS PRODUCTS: LOGIC FEMORAL PS CEM RIGHT SZ 2 (CAT# 02-010-01-0320); LGC TIBIAL FIT TRAY CEM SZ 2F /2T (CAT# 02-012-45-2020); THREE PEG PATELLA 32MM (CAT# 200-02-32). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00538 | 2022-05-09 | 2022-04-18 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 5 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 5 11MM | 02-012-35-5011 |  | AS REPORTED, APPROXIMATELY 7.5 YEARS POST INITIAL RIGHT TKA, THE (B)(6) MALE PATIENT HAD A REVISION DUE TO KNEE FAILURE OF AN EXACTECH RECALLED POLY. SURGEON FELT IT WAS NECESSARY TO PERFORM A TOTAL KNEE REVISION ON THIS PATIENT, BECAUSE THE PATIENT WAS EXPERIENCING PAIN AND MOBILITY ISSUES DUE TO EXTREME WEAR ON BOTH THE TIBIAL POLY INSERT AND THE POLY PATELLA. SURGEON REMOVED ENTIRE EXACTECH KNEE AND NOTICED THAT THE POLY WAS IN ONE PIECE; HOWEVER, POLY HAD EXTREME WEAR. THERE WERE NO PARTS OR PIECES FELL INTO THE PATIENT WOUND SITE. PATIENT WAS REVISED TO A COMPETITOR'S DEVICES. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PHOTOS/X-RAYS RECEIVED. THE DEVICES ARE NOT AVAILABLE FOR EVALUATION. | CONCOMITANTS PRODUCTS: LOGIC FEMORAL PS CEM RIGHT SZ 5 (CAT# 02-010-01-0350 / SERIAL# (B)(4)); LGC TIBIAL FIT TRAY CEM SZ 5F / 6T (CAT# 02-012-45-5060 / SERIAL# (B)(4)); THREE PEG PATELLA 41MM (CAT# 200-02-41 / SERIAL# (B)(4)); TIBIAL STEM EXT. SCREW (CAT# 204-70-00 / SERIAL# (B)(4)); FLUTED STEM EXTENSION 40L X 14 MM (CAT# 204-34-04 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00539 | 2022-05-09 | 2022-04-19 | I, n, j, u, r, y | Manufacturer report | ALT XLE LNR 10° FC G4 36 | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | ALT XLE LNR 10° FC G4 36 | 01-030-48-0436 |  | AS REPORTED, APPROXIMATELY FIFTEEN DAYS POST INITIAL LEFT SIDE THA, THIS (B)(6) FEMALE PATIENT WAS DISLOCATING SO SURGEON HAD TO GO BACK IN AND REVISE TO A DIFFERENT LINER. THE PATIENT HAD A 48 ALTEON CLUSTER CUP IN AND REP GAVE SURGEON A G4 36 10 FACE CHANGING LINER LATERALIZED TO HELP WITH STABILITY. SALES REP GAVE SURGEON THE WRONG LINER. THEREFORE, PATIENT HAD TO BE REVISED DUE TO HAVING THE WRONG LINER. SURGEON TOOK OUT OUR LINER WHICH WAS THE WRONG ONE AND OUT IN A COMPETITOR'S CUP. THERE WAS NO BREAKAGE OF DEVICE. THERE WAS A 30-45 MINS SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. SALES REP WAS UNABLE TO OBTAIN X-RAYS. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO THE HOSPITAL HAS IT IN THEIR POSSESSION. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00540 | 2022-05-09 | 2022-04-20 | I, n, j, u, r, y | Manufacturer report | FEMORAL COMPONENTS | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | FEMORAL COMPONENTS | UNKNOWN |  | AS REPORTED, SURGEON INDICATED THAT THE FEMORAL COMPONENT WAS DEBONDED. PHOTO RECEIVED. NO OTHER INFORMATION REPORTED. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00541 | 2022-05-09 | 2022-04-19 | I, n, j, u, r, y | Manufacturer report | NV GXL LNR, +5LAT, 36MM G4-60/62MM CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LNR, +5LAT, 36MM G4-60/62MM CUPS | 136-36-54 |  | AS REPORTED, SURGEON PERFORMED A REVISION DUE TO A FAILED GXL LINER. X-RAYS RECEIVED. NO ADDITIONAL INFORMATION RECEIVED. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00543 | 2022-05-10 | 2021-10-20 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | SMALL GLENOID PLATE | EXACTECH, INC. | 320-35-01 | 320-35-01 |  | AS REPORTED, APPROXIMATELY 1 YEAR POST OP THE INITIAL TSA, THIS (B)(6) MALE PATIENT'S WAS REVISED TO A REVERSE DUE TO INFECTION. PATIENT EXPERIENCED DEEP PAIN AFTER LIFTING A HEAVY SUITCASE IN 2021. NO IMPROVEMENT IN SYMPTOMS, LABS CONFIRMED INFECTION. THE CASE REPORT FORM INDICATES THIS EVENT IS DEFINITELY NOT RELATED TO DEVICES AND DEFINITELY RELATED TO PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | CONCOMITANT DEVICE(S): 300-01-11 - EQUINOXE, HUMERAL STEM PRIMARY, PRESS FIT 11MM; 320-36-00 - 36MM HUMERAL LINER +0 UNCONSTRAINED; 320-10-00 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0; 320-35-01 - SMALL GLENOID PLATE. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00542 | 2022-05-10 | 2022-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | VANTAGE | LP TALAR IMPACTOR INSERT | EXACTECH, INC. | 351-07-04 | 351-07-04 |  | IT WAS REPORTED THAT DURING IMPACTION OF THE TALAR IMPLANT, THE TALAR IMPLANT INSERTER BROKE. NO PARTS OR PIECES FELL INTO THE PATIENT WOUND SITE. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICE WILL BE RETUNING. | PENDING EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00544 | 2022-05-11 | 2022-04-20 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | SMALL LEFT SUPERIOR POSTERIOR BASEPLATE | EXACTECH, INC. |  | UNK |  | AS REPORTED, APPROXIMATELY 2 YEARS POST OP THE INITIAL LEFT TSA, THIS MALE PATIENT'S WAS REVISED DUE TO INFECTION. STEM AND BASEPLATE BOTH CAME OUT EASILY. AND THE SCREWS WERE BROKEN-MIDSHAFT, WITH NO REPORT OF FALL OR TRAUMA. SURGEON HAD WANTED TO UPSIZE STEM AND DO STANDARD BASEPLATE BUT AFTER BOTH CAME OUT SO EASILY DECIDED SPACER WAS BEST OPTION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES ARE NOT RETURNING DUE TO BEING DISPOSED OF AFTER REMOVAL. | SECTION H10: (H3) THE REVISION REPORTED WAS LIKELY THE RESULT OF INFECTION WHICH CAUSED EROSION AROUND THE CAGE AND ALLOWED THE CENTER CAGE TO LOOSEN IN GLENOID BONE. THE GLENOID LOOSENING LIKELY LED TO CRACK INITIATION, PROPAGATION AND ULTIMATE FRACTURE OF THE COMPRESSION SCREWS. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICES WERE NOT RETURNED FOR EVALUATION. SECTION H11: *THE FOLLOWING SECTIONS HAVE CORRECTED INFORMATION: (H6) COMPONENT CODE: 568, SCREW. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00545 | 2022-05-11 | 2022-04-21 | I, n, j, u, r, y | Manufacturer report | PS TIBIAL INSERTS SZ 4, 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | PS TIBIAL INSERTS SZ 4, 13MM | 204-24-13 |  | AS REPORTED, APPROXIMATELY FIFTEEN YEARS POST INITIAL LEFT TKA, THE (B)(6) MALE PATIENT HAD A FULL REVISION GOING TO A CONSTRAINED KNEE. PATIENT'S REVISION WAS DUE TO INSTABILITY OF FEMORAL COMPONENT. THERE WAS POLY WEAR AND METALLOSIS DUE TO THE 8MM TIBIAL SPACER AND OSTEOLYSIS WAS ALSO PRESENT. PATIENT WAS REVISED TO A SIZE 4 LEFT CONSTRAINED FEMUR WITH SPLINED STEMS IN THE FEMUR AND TIBIA AS WELL AS METAPHYSEAL CONES AND A 25 MM CONSTRAINED LINER. THERE WAS NO DEVICE BREAKAGE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. X-RAYS AND PHOTOS RECEIVED. THE DEVICES WILL BE RETURNED. AS OF NOTE: "PATIENT WAS REVISED IN 2014 WITH LOOSENING OF THE TIBIA AND THE SIZE 4/13MM POSTERS STABILIZED TIBIAL INSERT REFERENCE NUMBER 204¿24¿13 WITH SERIAL NUMBER (B)(4). IS A NON-CONFORMED INSERT THAT IT IS ON THE LIST OF POLYS". | CONCOMITANT MEDICAL PRODUCTS: FLUTED STEM EXTENSION 120L X 14 MM (CAT# 204-34-12 / SERIAL# (B)(4)). TIBIAL STEM EXT. SCREW (CAT# 204-70-00 / SERIAL# (B)(4)). LOGIC PTS, SIZE 4, 8MM (CAT# 02-012-42-4008 / SERIAL# (B)(4)). LGC TIBIAL FIT TRAY CEM SZ 4F / 3T (CAT# 02-012-45-4030 / SERIAL3 (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00546 | 2022-05-11 | 2019-09-30 | I, n, j, u, r, y | Manufacturer report | TRULIANT TIB IMP PSC INSERT SZ 4, 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT TIB IMP PSC INSERT SZ 4, 13MM | 02-022-44-4013 |  | AS REPORTED BY THE EXACTECH KNEE REPLACEMENT STUDY, APPROXIMATELY FOUR MONTHS POST INITIAL TKA, THIS PATIENT PRESENTED TO ED ON (B)(6) 2019 WITH RIGHT INFECTED KNEE. PATIENT TAKEN TO OR ON (B)(6) 2019 FOR EXPLANT OF TKA AND PLACEMENT OF ANTIBIOTIC SPACER. EVENT IS ONGOING. PER CLINICAL STUDY CASE REPORT FORM INDICATES THIS EVENT IS UNLIKELY RELATED TO DEVICES AND DEFINITELY RELATED TO THE PROCEDURE. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00547 | 2022-05-12 | 2022-04-14 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | KEELED GLENOID | EXACTECH, INC. | UNK | UNK |  | AS REPORTED, APPROXIMATELY 10 YEARS POST OP THE INITIAL RIGHT TSA, THIS (B)(6) FEMALE PATIENT WAS REVISED DUE TO CUFF FAILURE AND ALL POLY KEELED GLENOID LOOSENING. CONVERTED TO A REVERSE, ORIGINAL STEM WAS WELL FIXED. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETUNING DUE TO HOSPITAL POLICY. | PENDING EVALUATION: CONCOMITANT DEVICE(S): REPLICATOR PLATE AND HUMERAL HEAD. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00548 | 2022-05-12 | 2022-04-20 | I, n, j, u, r, y | Manufacturer report | NOVATION GXL LNR +5 LAT 40MM ID GROUP 4 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NOVATION GXL LNR +5 LAT 40MM ID GROUP 4 CUPS | 136-40-54 |  | AS REPORTED, THIS (B)(6) PATIENT PRESENTED TO THE EMERGENCY ROOM WITH SEVERE RIGHT HIP PAIN. CT AND X-RAY IMAGING SHOW EVIDENCE OF SIGNIFICANT ACETABULAR AND FEMORAL OSTEOLYSIS AND ECCENTRICITY OF THE FEMORAL HEAD IN THE ACETABULAR LINER. X-RAY RECEIVED. NO ADDITIONAL INFORMATION AVAILABLE. | CONCOMITANTS MEDICAL PRODUCTS: ELEMENT-STEM, COLLARED W/HA, STD OFFSET, SZ 11 (CAT# 164-03-11), NV CROWN CUP CLSTR HOLE 60MM GROUP 4 (CAT# 180-01-60), COCR FEMORAL HEAD 40MM, +0 OFFSET 12/14 (CAT# 142-40-00), 3.2MM DRILL BIT 20MM 1PK (CAT# 101-05-20), ALTEON 6.5MM SCREW, 15MM (CAT# 180-65-15). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00549 | 2022-05-12 | 2015-01-01 | I, n, j, u, r, y | Manufacturer report | FEMORAL COMPONENTS | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | FEMORAL COMPONENTS | UNKNOWN |  | AS REPORTED BY LEGAL TEAM, THIS FEMALE PATIENT HAD BILATERAL TKA. LEFT KNEE WAS DONE ON (B)(6) 2010 AND RIGHT KNEE ON (B)(6) 2011. PATIENT HAD A LEFT KNEE REVISION ON 2015, AND SINCE THEN HAS NOT SPENT A DAY WITHOUT PAIN, PATIENT HAS BEEN IN PHYSICAL THERAPY (PT) CONTINUOUSLY. NO OTHER INFORMATION PROVIDED. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00537 | 2022-05-12 | 2022-04-08 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 2 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED, | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 2 9MM | 02-012-35-2009 |  | AS REPORTED BY THE EXACTECH KNEE REPLACEMENT STUDY, APPROXIMATELY TEN YEARS POST INITIAL TKA, THIS 77 Y/O FEMALE PATIENT EXPERIENCED WEAR OF POLYETHYLENE INSERT LEFT KNEE WITHOUT EVIDENCE OF OSTEOLYSIS. THE EVENT IS STILL ONGOING. PER CLINICAL STUDY CASE REPORT FORM INDICATES THIS EVENT IS POSSIBLY RELATED TO DEVICES AND PROCEDURE. | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION RECEIVED THE FOLLOWING SECTIONS G3, G6, H2, H3, H6 AND H9 HAVE BEEN UPDATED ACCORDINGLY. (H3) THE PROSTHESIS WEAR REPORTED MAY HAVE BEEN THE RESULT OF MALALIGNMENT BETWEEN THE FEMORAL AND TIBIAL COMPONENTS, THIRD BODY WEAR, PATIENT-RELATED CONDITIONS, OR ANY COMBINATION OF THESE POSSIBILITIES. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION AND PRE-REVISION X-RAYS WERE NOT PROVIDED. THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE REPORTED EVENT OF "PROSTHESIS WEAR¿ IS ASSOCIATED WITH MATERIAL DAMAGE TO A SURFACE, USUALLY INVOLVING PROGRESSIVE LOSS OR DISPLACEMENT OF MATERIAL, DUE TO RELATIVE MOTION BETWEEN THAT SURFACE AND A CONTACTING SUBSTANCE OR SUBSTANCES. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00550 | 2022-05-15 | 2022-04-25 | I, n, j, u, r, y | Manufacturer report | LOGIC CR TIB INSERT SLOPE+, SZ 3, 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC CR TIB INSERT SLOPE+, SZ 3, 11MM | 02-012-48-3011 |  | AS REPORTED, APPROXIMATELY FIVE YEARS POST INITIAL TKA, THE FEMALE PATIENT RETURNED TO THE CLINIC COMPLAINING OF PAIN. A MRI SHOWED SYNOVITIS AND X-RAY SHOWED ASYMMETRIC WEAR. A REVISION WAS PERFORMED AND THE COMPONENTS WERE SECURE WITH NO SIGNS OF LOOSENING. THE KNEE WAS WASHED OUT AND A NEW 12 MM CRC INSERT WAS IMPLANTED. THE KNEE WAS STABLE WITH GOOD RANGE OF MOTION (ROM). SURGEON STATED THAT THE ORIGINAL INSERT WAS BAD WITH SIGNIFICANT DAMAGE TO ARTICULATE SURFACE. PATIENT LAST KNOWN TO BE IN STABLE CONDITION. PHOTO RECEIVED. THE EXPLANT HAS NOT BEEN RELEASED FOR RETURN. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00552 | 2022-05-15 | 2022-03-22 | I, n, j, u, r, y | Manufacturer report | TRULIANT TIB IMP PS INSERT SZ 3 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT TIB IMP PS INSERT SZ 3 9MM | 02-022-35-3009 |  | AS REPORTED BY THE EXACTECH TRULIANT KNEE CLINICAL STUDY, APPROXIMATELY TWO MONTHS POST OP INITIAL TKA, THIS 65 Y/O FEMALE PATIENT VISITED CLINIC WITH LIMITED RANGE OF MOTION (ROM) AND COMPLAINED OF STIFFNESS AND PAIN. PATIENT WAS PRESCRIBED A LEFT KNEE MANIPULATION UNDER ANESTHESIA. MANIPULATION SCHEDULED ON 28-APR-2022. | CONCOMITANT MEDICAL PRODUCTS: TRULIANT PS CEM FEM PS CEM LEFT SZ 3 (CAT# 02-020-11-0230), TRULIANT TRAY, CEM SZ 3F/2.5T (CAT# 02-022-45-3025), ADVANCED PATELLA 29M 3 PEG IMPLANT (CAT# 200-07-29). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00551 | 2022-05-15 | 2022-04-25 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 5 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 5 9MM | 02-012-35-5009 |  | AS REPORTED, APPROXIMATELY SIX YEARS POST OP INITIAL RIGHT TKA, THE 73 Y/O MALE PATIENT COMPLAINED OF "CLUNKING" ISSUES WITH THE PATELLA. SURGEON DID AN INSERT EXCHANGE. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PHOTO RECEIVED. SALES REP WAS UNABLE TO OBTAIN X-RAYS. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO HOSPITAL IS KEEPING THE IMPLANT. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00558 | 2022-05-17 | 2022-04-01 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 2, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 2, 9MM | 02-012-44-2009 |  | AS REPORTED, FEMALE PATIENT HAD A WASHOUT PERFORMED DUE TO A WOUND. PATIENT'S INSERT WAS REPLACED. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. SALES REP WAS UNABLE TO OBTAIN X-RAYS/PHOTO. THE DEVICE IS NOT AVAILABLE FOR EVALUATION AS IT WAS DISPOSED BY HOSPITAL. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00559 | 2022-05-17 | 2022-04-21 | I, n, j, u, r, y | Manufacturer report | LOGIC CR TIB INSERT SLOPE++, SZ 2, 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC CR TIB INSERT SLOPE++, SZ 2, 13MM | 02-012-49-2013 |  | AS REPORTED, APPROXIMATELY SIX YEARS POST INITIAL LEFT TKA, THE (B)(6) FEMALE PATIENT HAD A REVISION DUE TO PHYSICIAN REQUESTED IMPLANTS FOR REVISION OF WORN INSERT. PATIENT HAD SYNOVITIS REMOVED AND PATIENT IS EXPECTING BETTER MOTION AND RELIEF OF PAIN. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO HOSPITAL WOULD NOT ALLOW REMOVAL OF THE EXPLANT. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00557 | 2022-05-17 | 2021-12-23 | I, n, j, u, r, y | Manufacturer report | LOGIC CR TIB INSERT SLOPE+, SZ 3, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | UNK | 02-012-48-3009 |  | AS REPORTED BY THE LEGAL TEAM, APPROXIMATELY THREE YEARS POST INITIAL TKA, THE (B)(6)FEMALE PATIENT WAS DIAGNOSED WITH FAILED RIGHT TOTAL KNEE ARTHORPLASTY AND UNDERWENT REVISION SURGERY. DURING THE REVISION SURGERY, PATIENT'S SURGEON STATED, "PATIENT HAD A PREVIOUS RIGHT TOTAL KNEE ARTHROPLASTY. THIS HAD FAILED SECONDARY TO CATASTROPHIC POLYETHYLENE WEAR. THIS CREATED LARGE CYSTS WITHIN BOTH THE MEDIAL AND LATERAL POSTERIOR CONDYLES, AS WELL AS THE MEDIAL ASPECT OF THE PROXIMAL TIBIA. SURGEON REMOVED THE POLYETHYLENE, WHICH WAS NOTED LO HAVE SEVERE POLYETHYLENE WEAR." FOLLOWING THE REVISION SURGERY, PATIENT CONTINUES TO BE LIMITED IN PATIENT'S ACTIVITIES OF DAILY LIVING. PATIENT HAS DIFFICULTY WITH DAILY ACTIVITIES SUCH AS GETTING DRESSED, CLIMBING STAIRS, AND BATHING. AS A DIRECT RESULT, PATIENT HAS SUFFERED AND CONTINUES TO SUFFER PERMANENT AND DEBILITATING INJURES AND DAMAGES, INCLUDING BUT NOT LIMITED TO, SIGNIFICANT PAIN AND DISCOMFORT, GAIT IMPAIRMENT, POOR BALANCE, DIFFICULTY WALKING, COMPONENT PART LOOSENING, SOFT TISSUE DAMAGE, BONE LOSS, AND OTHER INJURIES PRESENTLY UNDIAGNOSED, WHICH ALL REQUIRE ONGOING MEDICAL CARE. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00555 | 2022-05-17 | 2022-04-26 | I, n, j, u, r, y | Manufacturer report | FEMORAL LINER | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | FEMORAL LINER | UNKNOWN |  | AS REPORTED, APPROXIMATELY SIX YEARS POST OP INITIAL LEFT TKA, PATIENT PRESENTED CLEAR EVIDENCE OF MEDIAL INSERT POLYETHYLENE WEAR WITH MEDIAL TIBIAL OSTEOLYSIS. THERE IS ALSO EVIDENCE OF MEDIAL FEMORAL OSTEOLYSIS AT THE MEDIAL FEMORAL CONDYLE. X-RAYS RECEIVED. NO OTHER INFORMATION PROVIDED. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00554 | 2022-05-17 | 2022-04-26 | I, n, j, u, r, y | Manufacturer report | NV GXL LINER NEUTRAL, 32MM ID GROUP 1 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LINER NEUTRAL, 32MM ID GROUP 1 CUPS | 130-32-51 |  | AS REPORTED, APPROXIMATELY FIVE YEARS POST OP INITIAL LEFT THA, THE 71 Y/O FEMALE PATIENT HAD A REVISION AS THERE WAS CLEAR EVIDENCE OF ASYMMETRIC POLYETHYLENE WEAR. THERE ALSO APPEARED TO BE MINIMAL PERIACETABULAR OSTEOLYSIS. X-RAYS RECEIVED. NO OTHER INFORMATION REPORTED. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00553 | 2022-05-17 | 2022-04-26 | I, n, j, u, r, y | Manufacturer report | NV GXL LINR, NTRL, 36MM ID, GROUP 4 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LINR, NTRL, 36MM ID, GROUP 4 CUPS | 130-36-54 |  | AS REPORTED, PATIENT HAD A REVISION DUE TO PREMATURE POLYETHYELENE WEAR, GXL LINER, WITH PERIACETABULAR OSTEOLYSIS. THERE WAS NO DELAY, PATIENT IS STABLE CONDITION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. IMAGES RECEIVED. THE DEVICE WILL BE RETURN. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00556 | 2022-05-17 | 2022-04-25 | I, n, j, u, r, y | Manufacturer report | CEMENTED FINNED TIB. TRA SZ 2F/2T | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | CEMENTED FINNED TIB. TRA SZ 2F/2T | 200-04-22 |  | AS REPORTED, APPROXIMATELY TWELVE YEARS POST INITIAL LEFT TKA, THE FEMALE PATIENT PRESENTED TO THE SURGEON'S OFFICE WITH PAIN IN HER LEFT KNEE. THE X-RAYS SHOWED POSSIBLE WEAR OR VARUS STRESS. THE TIBIA BONE LOOKS SUSPECT ON X-RAY. THE PATIENT WAS SCHEDULED FOR A REVISION TKA POSSIBLE POLY SWAP. UPON ENTERING THE KNEE JOINT, REMOVING THE POLY, AND TESTING/CHECKING THE FEMORAL AND TIBIAL COMPONENTS, IT WAS FOUND THAT THE TIBIAL COMPONENT WAS LOOSE ON THE BONE. THE SURGEON WAS ABLE TO LIFT IT UP OFF THE BONE EASILY WITH AN OSTEOTOME. THE SURGEON REMOVED ALL OF THE EXACETCH IMPLANTS AND COMPLETED A TOTAL KNEE REVISION USING COMPETITOR'S REVISION KNEE SYSTEM. THERE WAS NO BREAKAGE OF DEVICES. THERE WAS A 30-45 MINUTE SURGICAL DELAY/PROLONGATION DUE TO THE LOOSE IMPLANTS CAUSING A LONGER PROCEDURE THAN ANTICIPATED OF A POLY SWAP BUT THERE WAS NO ADVERSE EVENT AS A RESULT. THE PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE REVISION PROCEDURE. THE DEVICES ARE NOT AVAILABLE FOR EVALUATION DUE TO IMPLANTS WERE DISPOSED OF BY THE FACILITY. | CONCOMITANT MEDICAL PRODUCTS: PS TIBIAL INSERTS SZ 2, 13MM (CA# 204-22-13 / (B)(4)); OPTETRAK ASY,PS CEMENTED FEMORAL, SZ 2 (CAT# 234-02-02 / (B)(4)); THREE PEG PATELLA 29MM (CAT# 200-02-29 / (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00560 | 2022-05-18 | 2022-04-27 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 4 15MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 4 15MM | 02-012-35-4015 |  | AS REPORTED, APPROXIMATELY FIVE YEARS POST INITIAL RIGHT TKA, THIS PATIENT IS PRESENTING WITH OSTEOLYSIS ON THE MEDIAL FEMORAL EPICONDYLE WHICH IS CURRENTLY ASYMPTOMATIC BUT SURGEON BELIEVES IT¿S DUE TO POLY WEAR. AT THE MOMENT THE PATIENT IS CONSIDERING THEIR OPTIONS BUT SURGEON ANTICIPATES THE POLY EXCHANGE WILL OCCUR WITHIN THE NEXT THREE MONTHS. X-RAY RECEIVED. NO OTHER INFORMATION PROVIDED. | CONCOMITANT MEDICAL PRODUCTS: LOGIC FEMORAL PS CEM RIGHT SZ 4 (CAT# 02-010-01-0340 / SERIAL# (B)(4)). LGC TIBIAL FIT TRAY CEM SZ 4F / 4T (CAT# 02-012-45-4040 / SERIAL# (B)(4)). THREE PEG PATELLA 41MM (CAT# 200-02-41 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00562 | 2022-05-18 | 2022-04-25 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 42MM HUMERAL LINER +2.5 | EXACTECH, INC. | 320-42-03 | 320-42-03 |  | AS REPORTED, 2 DAYS POST OP THE INITIAL LEFT TSA, THIS (B)(6) MALE PATIENT WAS REVISED DUE TO DISLOCATION. PATIENT DISLOCATED ANTERIORLY AFTER FRACTURE REVERSE REPLACEMENT ON APRIL 23RD. REMOVED HUMERAL TRAY AND LINER AND REPLACED WITH A +5MM TRAY AND A 2.5MM CONSTRAINED HUMERAL LINER. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETUNING, DISPOSED OF AT HOSPITAL. | PENDING EVALUATION. CONCOMITANT DEVICE(S): 320-20-00, 7277597 - EQ REVERSE TORQUE DEFINING SCREW KIT. 320-10-00, 6692270 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00572 | 2022-05-19 | 2019-01-22 | I, n, j, u, r, y | Manufacturer report | EXACTECH | HUM LNR IMPACT TIP 42M | EXACTECH, INC. |  | 321-07-42 |  | IT WAS REPORTED FROM OUS THAT THE IMPACTOR TIP WAS BROKEN BY MALLETING WHEN THE IMPLANTS WERE INSERTED. NO PIECES WERE LEFT IN THE PATIENT. THERE WAS NO DELAY, AND THE PATIENT WAS NOT ADVERSELY AFFECTED. THE SURGEON RESOLVED THE ISSUE BY REPLACING THE 38MM IMPACTOR HEAD. THE PATIENT¿S HEALTH WAS STABLE LEAVING THE OR. THE REP WAS PRESENT AT THE TIME OF SURGERY. ATTEMPTS WERE MADE TO REQUEST ADDITIONAL INFORMATION DURING REMEDIATION AND NO RESPONSE WAS RECEIVED. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INITIAL REPORTER NAME AND ADDRESS. KOREA. DESIGN-RELATED ISSUES: THE DESIGN OF THE HUMERAL LINER IMPACTOR TIP HAS BEEN IN THE FIELD SINCE 2007. EXACTECH IS AWARE OF 28 SIMILAR COMPLAINT REPORTS INVOLVING 34 BROKEN HUMERAL LINER IMPACTOR TIPS SINCE 2008. BECAUSE THE HUMERAL LINER IMPACTOR TIP IS USED IN REVERSE TOTAL SHOULDER SURGERIES, SALES DATA FOR HUMERAL LINERS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF (B)(4) UNITS, WHICH HAS BEEN IN THE FIELD SINCE 2015. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS HUMERAL LINER IMPACTOR TIP, 750-2005-021-RMR-INSTRUMENTS REV N, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 15. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR. THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF A STRESS RISER INTRODUCED DURING IMPACTION, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. ¿ CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00579 | 2022-05-19 | 2019-01-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | SPECIFIC DEVICE NOT REPORTED | EXACTECH, INC. |  |  |  | IT WAS REPORTED FROM OUS THAT UPON RETURN OF THE LOAN KIT THE IMPACTOR TIP FRACTURE AND CRACKED IN HALF. ON (B)(6) 2019 THE PRODUCT LOCATION HAS NOT BEEN DETERMINED. IT IS EITHER MISPLACED OR WAS THROWN AWAY INADVERTENTLY. THE DEVICE CATALOG AND SERIAL/LOT NUMBER IS NOT AVAILABLE. THE LAST HOSPITAL AND SURGEON TO USE THE DEVICE IS NOT AVAILABLE. IT WAS CONFIRMED THAT THERE WAS NO COMPLAINT FROM ANY HOSPITAL OR USER THAT HAD THE DEVICE THAT WAS BROKEN. NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. ENGINEERING EVALUATION: DESIGN-RELATED ISSUES: EXACTECH IS AWARE OF 8 SIMILAR COMPLAINT REPORTS INVOLVING FEMORAL HEAD IMPACTORS DUE TO DAMAGED IMPACTOR TIPS SINCE 2008. BECAUSE THE FEMORAL HEAD IMPACTOR IS USED IN TOTAL HIP SURGERIES, SALES DATA FOR FEMORAL HEADS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4). THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: MANUFACTURING DATA COULD NOT BE REVIEWED BECAUSE THE SERIAL/LOT INFORMATION WAS NOT PROVIDED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW.¿ THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FAILURE OF THE RADEL TIP. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT FROM A FACILITY OR USER. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00578 | 2022-05-19 | 2019-09-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | SPEC DUAL STRIKE PLATE BROACH HANDLE | EXACTECH, INC. |  | SK-025375 |  | IT WAS REPORTED FROM THE US THE BROACH HANDLE BROKE. IT IS NOT INDICATED IF THIS OCCURRED DURING SURGERY. THE AGENT WAS ABLE TO FIX DEVICE. BASED ON THE INFORMATION AVAILABLE, IT IS UNCLEAR WHEN THE DEVICE DISASSEMBLED DURING SURGERY. THE ENGINEERING TEAM SPOKE ON THE PHONE WITH THE SURGEON ON 10/10/2019 AND DISCUSSED THE BROACH HANDLES. THE SURGEON HAS NOT NOTICED ANYTHING WITH THE HANDLES INTRA-OPERATIVELY AND UNDERSTOOD HIS SALES AGENTS WERE WORKING HARD TO MAKE SURE HIS CASES WENT FLAWLESSLY. THE AGENT WAS ABLE TO GET THE DEVICE WORKING AGAIN. THE SURGEON IS HAPPY WITH THE PERFORMANCE OF THE BROACH HANDLES AND HAS THE SPECIAL HANDLES AND STANDARD DUAL OFFSET AVAILABLE AT EVERY CASE. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. INITIAL REPORTER ALSO SENT REPORT TO FDA? UNKNOWN. VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR 10X OPTICAL UNLESS OTHERWISE SPECIFIED) THE DEFORMATION MARKS APPEAR CONSISTENT WITH BEING IMPACTED DURING SURGERY. THE LABEL WAS LIKELY ADDED PRIOR TO RETURNING TO EXACTECH BY THE HOSPITAL. THE DEFORMATION APPEARS CONSISTENT WITH A WELD FAILURE. THIS WAS LIKELY DUE TO IMPACTION FORCES AND SUBSEQUENT STERILIZATION CYCLES THAT LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE. BASED ON THE PART DRAWING, THERE IS NOT A WELD CALLOUT FOR THIS COMPONENT. THE ADDED WELD WAS LIKELY DUE TO THE DOWEL PIN NEEDING TO BE SECURED TO THE ACTUATOR SINCE IT WAS NOT BEING FULLY RETAINED WITHIN THE BROACH HANDLE BODY, AS IN THE PRODUCTION DEVICE (01-003-02-0001/3). BECAUSE THE BODY OF THE BROACH HANDLE DOES NOT FULLY ENCAPSULATE THE DOWEL PIN, ONCE FRACTURED, THE PIN WAS ABLE TO DISASSEMBLE FROM THE HANDLE, ALLOWING FOR THE SUBSEQUENT DISASSEMBLY OF THE ACTUATOR FROM THE LEVER. THE DEVICE WAS REPORTEDLY RETURNED TO EXACTECH WITH THE MISSING DOWEL PIN, BUT WAS LOST PRIOR TO BEING DELIVERED TO COMPLAINTS. WHEN EXAMINING THE BODY OF THE BROACH HANDLE, THERE DOES NOT APPEAR TO BE EVIDENCE OF A WELD. THIS WELD WOULD HAVE PREVENTED THE RETAINING PIN FROM BEING ABLE TO DISASSEMBLE FROM THE BODY. THE DISASSEMBLY OF BOTH THE DOWEL PIN AND THE RETAINING PIN WOULD HAVE ALLOWED FOR THE SUBSEQUENT DISASSEMBLY OF THE ACTUATOR FROM THE CAVITY OF THE BROACH HANDLE BODY. THE DESIGN OF THIS SPECIAL INSTRUMENT WAS RELEASED ON TM-2019-0629. ACCORDING TO THE DOCUMENT, 6 DUAL STRIKE PLATE BROACH HANDLE BODIES WERE TO BE MADE, HOWEVER 3 WOULD BE USED TO CREATE SK-025375 REV 3 AND 3 WOULD BE USED TO CREATE SK-028526 REV 2. FOR SK-025375 REV 3 (DSR219060), THE BODY, ALL EXTERNAL COMPONENTS, AND SOME INTERNAL COMPONENTS (BELL CRANK, LEVER ACTUATOR, AND PIVOT RIVET) WERE TO BE MANUFACTURED. THE REMAINING INTERNAL COMPONENTS WOULD BE TAKEN FROM A DISASSEMBLED STANDARD BROACH HANDLE (CN: 01-003-02-0001/3). ACCORDING TO THE EXPERIENCE REPORT, THE SPECIAL BROACH HANDLE BROKE DURING A CASE AND ALL PIECES WERE LOCATED OUTSIDE OF THE PATIENT. BASED ON THE INFORMATION AVAILABLE, IT IS UNCLEAR WHEN THE DEVICE DISASSEMBLED DURING SURGERY. A CORRECTION AND REMOVAL COMMITTEE (CRC) MEETING WAS HELD TO DETERMINE THE NEED FOR FIELD ACTION. DESIGN-RELATED ISSUES: THE DESIGN OF THIS SPECIAL BROACH HANDLE HAS BEEN IN THE FIELD SINCE MARCH 2019. EXACTECH IS AWARE OF 2 COMPLAINT REPORTS INVOLVING THE DISASSEMBLY OF THESE SPECIAL BROACH HANDLE SINCE MARCH 2019. SINCE THESE DEVICES ARE USED WITH ALTEON HA STEMS AND BY SPECIFIED SURGEONS, SALES DATA FOR HA FEMORAL STEMS IMPLANTED BY THE SPECIFIED SURGEONS DURING THE POSSESSION OF THE SPECIAL BROACH HANDLES WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE (B)(4). THIS IS CONSIDERED ¿MODERATE¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. ADDITIONALLY, THE DOWEL PIN THAT SECURES THE LEVER ACTUATOR TO THE LEVER WAS NOT RETAINED WITHIN THE BODY OF THE BROACH HANDLE, UNLIKE THE DESIGN OF THE PRODUCTION PART. THIS LIKELY ALLOWED FOR THE DISASSEMBLY OF THE ACTUATOR. MFG-RELATED ISSUES: EXACTECH IS AWARE OF (B)(4) COMPLAINT REPORTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF (B)(4) UNITS, WHICH HAS BEEN IN THE FIELD SINCE JUNE 2019. AS DOCUMENTED ABOVE, THE LEAF SPRING, LEVER ACTUATOR, AND BODY CAVITY ALL APPEAR TO BE OUT OF SPECIFICATION. THE LATTER TWO SUBCOMPONENTS WERE MANUFACTURED FOR THIS SPECIAL INSTRUMENT. THE LEAF SPRING SHOULD HAVE BEEN TAKEN FROM A DISASSEMBLED STANDARD BROACH HANDLE; HOWEVER, IT DOES APPEAR TO HAVE BEEN MODIFIED. THESE OUT OF SPECIFICATION CONDITIONS LIKELY CONTRIBUTED TO THE DISASSEMBLY OF THE DEVICE. THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF A COMBINATION OF A WELD FAILURE AS A RESULT OF IMPACTION FORCES, THE DOWEL PIN NOT BEING RETAINED WITHIN THE BODY OF THE BROACH HANDLE, AND THE OUT OF SPECIFICATION CONDITIONS WHICH ALLOWED FOR THE SUBSEQUENT DISASSEMBLY OF THE DEVICE. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. THE DEVICE DISASSEMBLED, IT IS UNCLEAR IF THIS OCCURRED DURING A SURGICAL CASE, THERE WAS NO REPORTED PATIENT ADVERSE EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00577 | 2022-05-19 | 2019-12-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | HEAD PUSHER, MOD, LOWPRO | EXACTECH, INC. |  | 01-001-03-0001 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THERE WAS A BROKEN INSTRUMENT. IN THE LAST FEW MINUTES OF THE CASE, THE SURGEON WAS IMPACTING THE FEMORAL HEAD WITH OUR INSTRUMENT AND THE BLACK IMPACTOR HEAD CRACKED AND BROKE INTO PIECES. NO PIECES FELL INTO THE SURGICAL SITE AND NO PIECES WERE LEFT INSIDE OF PATIENT. THERE WAS A DELAY OF 1 OR 2 MINUTES WHILE ANOTHER TRAY WAS OPENED. THERE WAS NO ADVERSE EVENT TO THE PATIENT. THE DEVICE WAS NOT RETURNED FOR ANALYSIS, IT WAS DISPOSED OF AT THE HOSPITAL. THE PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION WITH NO ISSUES. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. ENGINEERING EVALUATION: THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF A STRESS RISER INTRODUCED DURING IMPACTION WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE COMPONENT WAS NOT RETURNED FOR EVALUATION. DESIGN-RELATED ISSUES: THE DESIGN OF THIS FEMORAL HEAD PUSHER, LOW PROFILE/MODIFIED FEMORAL HEAD IMPACTOR HAS BEEN IN THE FIELD SINCE 2013. EXACTECH IS AWARE OF 3 TOTAL COMPLAINTS INVOLVING FEMORAL HEAD PUSHERS SINCE 2013. BECAUSE THE FEMORAL HEAD PUSHER IS USED IN TOTAL HIP REPLACEMENTS, THE SALES DATA FOR NON-CERAMIC FEMORAL HEADS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THE ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: THE MANUFACTURING LOT INFORMATION WAS NOT PROVIDED. THEREFORE, NO INFORMATION REGARDING THE MANUFACTURING OF THE SPECIFIC DEVICE LOT CAN BE GATHERED AT THIS TIME. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) AND RISK ASSESSMENT AND CONTROLS REPORT (RACR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. RMR-00025 REV- AND RACR-00028 REV- WERE REVIEWED. THE RISK IS CAPTURED. HOWEVER, THE RACR WILL BE UPDATED UNDER (B)(4) TO CAPTURE AN APPROPRIATE HARM FOR THIS HAZARDOUS SITUATION. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW,¿ THE RISK IS CAPTURED IN THE RMR/RACR, AND THE RACR WILL BE UPDATED TO CAPTURE AN APPROPRIATE HARM FOR THIS HAZARDOUS SITUATION. IFU 700-096-181: INSTRUMENT INSPECTION VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY DESIGN ISSUES. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00576 | 2022-05-19 | 2018-12-10 | I, n, j, u, r, y | Manufacturer report | EXACTECH | HUM HD IMPCTR TIP | EXACTECH, INC. |  | 311-07-07 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE HUMERAL HEAD IMPACTOR TIP THAT BROKE DURING IMPACTION. NO INCIDENT TO PATIENT. NO DELAY TO SURGERY. NEW INFORMATION RECEIVED ON 23 JUL 2021: SUMMARY OF REPORTED INCIDENT: ¿THE INSTRUMENT COULD HAVE BEEN BROKEN DURING THE SURGERY. HOWEVER, THE BREAKAGE HAS BEEN NOTICED AT THE RECOVERY OF THE INSTRUMENTATION-SET BY OUR LOGISTICS DEPT., AFTER THE CLEANING REALIZED BY THE STERILIZATION TEAM OF THE HOSPITAL. NO INFORMATION HAS BEEN RECEIVED FROM THE HOSPITAL REGARDING THIS ISSUE AND [THE AGENCY] DIDN¿T RECEIVE ANY REQUEST FOR COMPLEMENTARY INFORMATION REGARDING THIS ISSUE AS WELL.¿ MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INITIAL REPORTER ALSO SENT REPORT TO FDA IS UNKNOWN. DESIGN-RELATED ISSUES: THE DESIGN OF THE HUMERAL HEAD IMPACTOR TIP HAS BEEN IN THE FIELD SINCE 2004. EXACTECH IS AWARE OF 7 SIMILAR COMPLAINT REPORTS INVOLVING BROKEN HUMERAL HEAD IMPACTOR TIPS SINCE 2008. BECAUSE THE HUMERAL HEAD IMPACTOR TIP IS USED IN ANATOMIC TOTAL SHOULDER SURGERIES, SALES DATA FOR HUMERAL HEADS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF (B)(4) UNITS, WHICH HAS BEEN IN THE FIELD SINCE MARCH 2018. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, THE RISK IS CAPTURED IN THE RMR AND RACR, AND THE HARM WILL BE UPDATED. THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE RADEL IMPACTION SURFACE RELATED TO NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. ¿ CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00575 | 2022-05-19 | 2019-04-22 | I, n, j, u, r, y | Manufacturer report | EXACTECH | M-SERIES 12/14 FEMORAL HEAD IMPACTOR | EXACTECH, INC. |  | 153-00-02 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE INSTRUMENTATION. THE FEMORAL HEAD IMPACTOR BROKE WHEN IMPACTING THE FINAL FEMORAL HEAD WITH A MALLET. THERE WAS NO REPORTED DELAY TO SURGERY OR PATIENT ADVERSE EVENT. THIS DEVICE IS IN THE SURGICAL SITE AND THERE IS A POSSIBILITY OF FRAGMENTS TO BE REMOVED FROM THE SITE. THE DEVICE WAS RECEIVED. THE REPRESENTATIVE WAS PRESENT AT THE TIME OF SURGERY. TWO ATTEMPTS WERE MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REPRESENTATIVE, BUT NO RESPONSES WERE RECEIVED. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED BY THE CONTACTS FOR THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. FINDINGS: THE FRACTURED IMPACTION TIP APPEARS CONSISTENT WITH CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE. THE DISCOLORATION OF THE FRACTURE SURFACE APPEARS CONSISTENT WITH HAVING BEEN CRACKED FOR A SIGNIFICANT PERIOD OF TIME BEFORE ULTIMATE FRACTURE, WHICH ALLOWED FOR BIOLOGICAL MATERIAL TO BE RETAINED WITHIN THE CRACK AND STAIN THE RADEL. MOST LIKELY UNDERLYING CAUSE: BASED ON PREVIOUS SIMILAR INVESTIGATIONS, THE BROKEN M-SERIES 12/14 FEMORAL HEAD IMPACTOR REPORTED WAS LIKELY THE RESULT OF NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES, WHICH LEAD TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE RADEL TIP. RMR REVIEW: A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00574 | 2022-05-19 | 2019-03-05 | I, n, j, u, r, y | Manufacturer report | EXACTECH | TRULIANT CC FEM TRIAL SIZE 3 | EXACTECH, INC. |  | 02-011-06-1130 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE FEMORAL TRIAL. THE FEMORAL TRIAL WITH THE TRIAL STEM WERE IMPACTED ONTO THE FEMUR AT THAT POINT BOTH TRIAL CONDYLES BROKE OFF. ALL PIECES WERE ACCOUNTED FOR. ACCORDING TO THE REPRESENTATIVE, AT THIS TIME THE SURGEON HAD ALREADY CONFIRMED THE SIZE AND INSERT THICKNESS AND THE SURGEON COMPLETED THE SURGERY WITHOUT DELAY OR ANY COMPROMISE TO THE PATIENT. IMAGES WERE PROVIDED. THE DEVICE WAS NOT RECEIVED FOR ANALYSIS. ATTEMPTS WERE MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REP, BUT NO RESPONSES WERE RECEIVED. NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE EVALUATED BY MFR: NOT RETURNED, ENGINEERING EVALUATION CONDUCTED. DESIGN-RELATED ISSUES: THE DESIGN OF THIS TRULIANT FEMORAL TRIAL HAS BEEN IN THE FIELD SINCE 2017. EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINT REPORTS INVOLVING A BROKEN FEMORAL TRIAL SINCE 2017. (B)(4).MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF RECEIVING ANY OTHER COMPLAINT REPORTS INVOLVING A PART FROM THIS MANUFACTURING LOT OF (B)(4) PIECES, WHICH HAVE BEEN IN THE FIELD SINCE 2018. THEREFORE, THIS DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS TRULIANT FEMORAL TRIAL, 750-2011-032-RMR-INSTRUMENTS REV E, WAS REVIEWED. THE RISK IS CAPTURED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR. THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF APPLYING EXCESS FORCE TO THE FEMORAL TRIAL DURING IMPACTION. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE COMPONENT WAS NOT RETURNED FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION. VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED INSTRUMENT MALFUNCTION IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO ANY REPORTED PATIENT IMPACT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00573 | 2022-05-19 | 2018-10-19 | I, n, j, u, r, y | Manufacturer report | EXACTECH, | TRU TIBIAL ADAPTOR PLATE TRIAL/IMPLANT, SIZE 3.5 | EXACTECH, INC. |  | 02-029-25-4035 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON HAD AN ISSUE WITH THE INSTRUMENTS. WHILE TRIALING THE TIBIAL INSERT ON REAL IMPLANT THE SURGEON WAS HAVING TROUBLE GETTING THE TRIAL OUT AND BENT THE TRIAL HANDLE AND BROKE OFF THE TAP ON THE SIZE 3.5 TRIAL SHIM. ¿NO PIECES WERE LEFT IN THE PATIENT¿. THE PATIENT LEFT THE OPERATING ROOM IN STABLE CONDITION. THERE WAS NO DELAY TO SURGERY AND NO PATIENT ADVERSE EVENT. THE BENT HANDLE CAUSED NO HARM TO THE PATIENT. THE AGENT WAS PRESENT AT TIME OF SURGERY. INACTIVE AGENCY WITH NO HOSPITAL/OTHER CONTACT INFORMATION FOR THE EVENT. NO ADDITIONAL INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. CONCOMITANT MEDICAL PRODUCTS: CAT #02-029-29-1000; LOT;S/N 90106060; DESCRIPTION TRULIANT TIBIAL TRIAL HANDLE. CAPAS: THE ROOT CAUSE ASSOCIATED RESIDES IN THE MISUSE/ MISTREATMENT OF THE ADAPTOR PLATES OR CANNOT BE IDENTIFIED DUE TO INSUBSTANTIAL EVIDENCE . THE CAPA TEAM HAS DETERMINED, HOWEVER, THAT USE CONSIDERATIONS INDICATE THAT THE FLEXURE DESIGN IS SUSCEPTIBLE TO DAMAGE FROM SAID MISUSE/MISTREATMENT. DESIGN: A REVIEW OF THE TOLERANCE STACK-UP PERFORMED IN TM-2017-0417 WAS CONDUCTED. AS PART OF THIS INVESTIGATION, THE TOLERANCE STACK-UP BETWEEN THE TRULIANT TIBIAL TRIAL ADAPTOR PLATES AND THE CORRESPONDING TRULIANT AND LOGIC TIBIAL TRAYS WAS REVISITED TO CONFIRM IT HAD BEEN ADEQUATELY PERFORMED. THE REVIEW WAS FOUND TO BE ACCURATE. THERE IS NO NEW INFORMATION THAT HAS BEEN FOUND THAT WOULD ALTER THE RESULTS OR METHODS OF SAID ANALYSIS. IN ADDITION, ATTEMPTS WERE MADE IN THE LABORATORY TO REPLICATE FAILURE MODES WITHOUT SUCCESS. THE DESIGN FUNCTIONS AS INTENDED UNDER WORST CASE CONDITIONS. MANUFACTURING: REVIEW OF THE DEVICE DHR CONCLUDED THAT ALL PARTS WERE MADE WITHIN SPECIFICATION. USE CONSIDERATIONS: THE DESIGN DOES LEND TO BEING SUSCEPTIBLE TO BREAKAGE IF USED INCORRECTLY DUE TO THE USE OF ULTEM MATERIAL WITH A THINNER ASPECT UNDER LOAD. THIS WAS OBSERVED IN A COMPLAINT WHEN A INSTRUMENT TRAY WAS DROPPED ON THE INSTRUMENT. IN ADDITION, IN OTHER CASES, THE DEVICES BROKE AT SOME POINT DURING REPROCESSING AT ONE PARTICULAR INSTITUTE. A REVIEW WAS CONDUCTED AT THIS INSTITUTE¿S REPROCESSING PROCEDURES AND RESULTS INDICATED THAT CORRECT REPROCESSING INSTRUCTIONS PROVIDED BY EXACTECH WAS FOLLOWED. A REVIEW OF THE INSTRUMENT TRAY CASE INDICATED THAT IT IS POSSIBLE THAT THE ADAPTOR PLATES MAY NOT BE FULLY CONSTRAINED DURING REPROCESSING. DUE TO THE NATURE OF THE REPROCESSING PROCESS IT IS POSSIBLE THE PLATES BECAME DISLODGED AND DAMAGED ELSEWHERE IN THE MACHINE. HUMAN ERROR: HUMAN ERROR WAS CONSIDERED AS PART OF THE INVESTIGATION AND FOUND TO BE THE ROOT CAUSE OF ONE CASE. THE ADAPTOR PLATE THAT WAS USED IN SAID COMPLAINT WAS THE INCORRECT SIZE RELATIVE TO THE TIBIAL TRAY THAT HAD BEEN IMPLANTED. THIS MAY HAVE CAUSED THE PLATE TO DISLODGE UNPREDICTABLY. IN ADDITION A HOHMANN RETRACTOR WAS USED TO REMOVE THE PLATE AND NOT THE TRULIANT DROP ROD HANDLE, AS STATED PER THE OPERATIVE TECHNIQUE, 712-35-31. THE RETRACTOR IS CURVED AND FEATURES A MUCH SMALLER NOSE WHICH COULD INADVERTENTLY FACILITATE CONTACT WITH THE FLEXURE DURING REMOVAL. THE GEOMETRY OF THE TRULIANT DROP ROD HANDLE, IN ADDITION TO AN APPROPRIATELY CONSTRAINED ADAPTOR PLATE, AS DESIGNED CREATES A SCENARIO IN WHICH CONTACT WITH THE FLEXURE IS EXTREMELY DIFFICULT. IFU 700-096-181: INSTRUMENT INSPECTION. VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00570 | 2022-05-19 | 2019-12-11 | I, n, j, u, r, y | Manufacturer report | EXACTECH | HUM LNR IMPACT TIP 38M | EXACTECH, INC. |  | 321-07-38 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE IMPACTOR TIP BROKE. THE IMPACTOR TIP WAS USED TO IMPACT THE LINER. THE HANDLE BROKE THROUGH THE PLASTIC AND IT WAS IN 2 PIECES NOTHING WAS LEFT IN PATIENT AND THERE WAS NO NEGATIVE IMPACT TO THE PATIENT. THE FACILITY CENTRAL STERILE CENTER CALLED THE AGENT TO SAY THEY THREW THE BROKEN DEVICE AWAY. THERE WAS NO INJURY OF THE PATIENT OR USER, AND NO PARTS OF THE BROKEN DEVICE/DEVICE FRAGMENTS LEFT IN THE PATIENT. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: AGE, SEX, WEIGHT, ETHNICITY, RACE. LOT/SERIAL NUMBER UNKNOWN, WAS NOT REPORTED. INITIAL REPORTER ALSO SENT REPORT TO FDA: UNKNOWN. DUE TO A TREND OF THE RECENT COMPLAINTS HAVING A HIGHER OCCURRENCE RATE IN THE KOREAN MARKET, CRC2020-01-10-07 AND CAPA(B)(4) WERE ESTABLISHED TO INVESTIGATE ANY ADDITIONAL POTENTIAL CAUSE OF THE HIGHER OCCURRENCE RATE IN THIS SPECIFIC MARKET. PER CRC2020-01-10-07: HRI SCORE OF ¿1¿ ¿ NO OR LOW RISK TO HEALTH. EXACTECH HAS RECEIVED 46 GLOBAL EXPERIENCE REPORTS DOCUMENTING BREAKAGE OF THE HUMERAL LINER IMPACTOR TIPS SINCE 2011. THESE IMPTACOR TIPS ARE USED FOR IMPLANTATION OF THE REVERSE SHOULDER HUMERAL LINERS, WHICH HAVE SALES OF 49,608 TO DATE SINCE 2011. THIS CORRESPONDS TO A 0.09% OCCURRENCE RATE. THIS WON¿T CAUSE INJURY TO THE PATIENT. OF NOTE, THERE HAS BEEN AN INCREASED NUMBER OF COMPLAINTS COMING FROM KOREA. THERE HAVE BEEN 8 COMPLAINTS SINCE 2011 FROM KOREA. IN KOREA, THE HUMERAL LINERS HAVE SALES OF 1,860 TO DATE SINCE 2011. THIS CORRESPONDS TO A 0.43% OCCURRENCE RATE WHICH IS STILL BELOW OUR THRESHOLD OF 0.5% HRI SCORE OF ¿3¿ ¿ POTENTIAL RISK TO HEALTH. EXACTECH HAS NOT RECEIVED ANY COMPLAINTS DETAILING PIECES OF THE HUMERAL LINER IMPACTION TIP BEING LEFT IN THE PATIENT. IF THE BREAKAGE HAPPENS, IT¿S HIGHLY NOTICEABLE AND THE HEALTHCARE PROFESSIONALS ARE UNLIKELY TO OMIT IT. THE BREAKAGE PIECES ARE USUALLY LARGE AND NOT HARD TO LOCATE. IT CAN BE CONFIRMED IF ALL BROKEN PIECES ARE COLLECTED BY CHECKING THE COMPLETENESS OF THE ASSEMBLED PART. IF THE BROKEN PIECE IS LEFT IN THE PATIENT¿S WOUND WITHOUT BEING NOTICED, IT COULD CAUSE INJURY TO THE PATIENT AND WOULD REQUIRE PROFESSIONAL MEDICAL INTERVENTION. THE MATERIAL OF THESE INSTRUMENTS, RADEL PER ETS-0205, IS INCLUDED IN THE BIOCOMPATIBILITY ASSESSMENT TECHNICAL MEMO, TM-2013-465. RADEL PASSES THE BIOCOMPATIBILITY REQUIREMENTS FOR SURGICAL USE, AND THERE ARE NO ANTICIPATED ADVERSE HEALTH CONSEQUENCES. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: AS OF THIS TIME, INVESTIGATION FOR AN ADDITIONAL MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE RELATED TO MARKET SPECIFIC OCCURRENCE IS STILL ONGOING. REFER TO CAPA(B)(4) FOR ADDITIONAL INFORMATION. RMR-00004 REV A AND RACR-00007 REV A WERE REVIEWED. THE RISK IS CAPTURED. MOST LIKELY UNDERLYING CAUSE: BASED ON PREVIOUS SIMILAR INVESTIGATIONS, THE FRACTURED HUMERAL LINER IMPACTOR TIP REPORTED WAS LIKELY THE RESULT OF A STRESS RISER INTRODUCED DURING IMPACTION WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THE BROKEN PIECES OF INSTRUMENT WERE REPORTED NOT LEFT IN THE PATIENT AND THERE WAS NO PATIENT ADVERSE EVENT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00569 | 2022-05-19 | 2019-12-03 | I, n, j, u, r, y | Manufacturer report | EXACTECH | HUM LNR IMPACT TIP 38M | EXACTECH, INC. |  | 321-07-38 |  | IT WAS REPORTED FROM OUS THAT DURING A REVERSE SHOULDER PROCEDURE THE INSTRUMENT BROKE. THE IMPACTOR TIP BROKE IN HALF WHEN IMPACTING THE DEFINITIVE HUMERAL LINER. THE ISSUE DID NOT NEED TO BE RESOLVED, AS THE IMPLANTS WERE IMPACTED FAR ENOUGH. IT WAS CONFIRMED THAT NO PIECES WERE LEFT IN THE PATIENT. ¿NO ISSUE TO PATIENT¿ NO DELAY TO THE PROCEDURE AND NO ADVERSE EVENT TO THE PATIENT. PRODUCT WAS NOT AVAILABLE FOR RETURN. THE AGENT WAS PRESENT AT SURGERY. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE EVALUATED BY MFR: NOT RETURNED FINDINGS: THE FRACTURED HUMERAL LINER IMPACTOR TIP IN THE PHOTOGRAPH APPEARS CONSISTENT WITH A STRESS RISER INTRODUCED DURING IMPACTION, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE. AN ACCURATE LOT NUMBER WAS NOT PROVIDED FOR THIS EXPERIENCE, SO IT CANNOT BE DETERMINED IF IT WAS MANUFACTURED PRIOR TO DEVICE DESIGN CHANGES TO INCREASE THICKNESS AND MODIFY THE CONNECTION FEATURE DIMENSIONS. MOST LIKELY UNDERLYING CAUSE: BASED ON PREVIOUS SIMILAR INVESTIGATIONS, THE FRACTURED HUMERAL LINER IMPACTOR TIP REPORTED WAS LIKELY THE RESULT OF A STRESS RISER INTRODUCED DURING IMPACTION WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE. RMR REVIEW: RMR-00004 REV- AND RACR-00007 REV- WERE REVIEWED. THE RISK IS CAPTURED IN RISK ID 9.7 OF THE RACR. HOWEVER, IT IS BEING UPDATED UNDER (B)(4) TO INCLUDE A HARM OF ¿NONE¿. IFU 700-096-181: INSTRUMENT INSPECTION VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES. THERE WAS NO REPORTED PATIENT ADVERSE EVENT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00568 | 2022-05-19 | 2019-11-26 | I, n, j, u, r, y | Manufacturer report | EXACTECH | TRULIANT PIN PULLER SYRINGE STYLE | EXACTECH, INC. |  | 02-029-90-4100 |  | IT WAS REPORTED FROM THE US THAT DURING A SURGICAL PROCEDURE THE PIN PUSHER CRACKED AND BROKE REMOVING A PIN FROM THE TIBIAL TRAY. THERE WERE NO PARTS OF THE BROKEN INSTRUMENT LEFT IN THE PATIENT. THERE WAS NO REPORTED PATIENT ADVERSE EVENT. THE AGENCY IS INACTIVE, NO FURTHER INFORMATION. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. (B)(4) AND CAPA (B)(4). FINDINGS: ONE OR MORE OF THE METAL PRONGS USED TO CLAMP A PIN HEAD ARE MISSING FROM THE REPORTED TRULIANT SYRINGE PIN PULLERS. WHEN ANALYZING THE FRACTURE SURFACE AS A PART OF [A COMPLAINT], THE FAILURE APPEARS CONSISTENT WITH BRITTLE FRACTURE. THIS TYPE OF FAILURE IS MOST LIKELY THE RESULT OF THE PRONGS EXPERIENCING EXCESSIVE FORCES. FURTHERMORE, IF THE PIN PULLER WAS USED OFF-AXIS RELATIVE TO THE BONE PIN, AN UNEVEN FORCE DISTRIBUTION MAY HAVE BEEN APPLIED TO THE METAL PRONGS AND COULD EXPLAIN WHY THEY FRACTURED FROM ONLY ONE SIDE. OVER TIME, THESE EXCESSIVE OFF-AXIS FORCES COULD POTENTIALLY LEAD TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE MATERIAL. MOST LIKELY CAUSE: THE ROOT CAUSES OF THE PIN PULLER FAILURES AS IDENTIFIED IN CAPA (B)(4)., ARE: 1. THE BREAKAGE AREA OF THE COMPONENT IS RELATIVELY THIN WHICH DOES NOT OFFER GREAT SAFETY MARGIN WHEN THE INSTRUMENT IS SUBJECT TO OFF-AXIS LOADING. 2. OFF-AXIS LOADING WHEN THE USER APPLIES A BENDING MOMENT TO THE DEVICE DURING PIN DRIVING/PULLING CORRECTIVE ACTION TAKEN: DESIGN UPDATES WERE IMPLEMENTED TO IMPROVE THE STRENGTH OF THE BREAKAGE AREA AND THE OP TECH WAS UPDATED TO INSTRUCT USERS TO AVOID USING THE INSTRUMENT OFF-AXIS. IFU 700-096-181: INSTRUMENT INSPECTION, VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. THE PIN PUSHER CRACKED AND BROKE REMOVING A PIN FROM THE TIBIAL TRAY. THERE WERE NO PARTS OF THE BROKEN INSTRUMENT LEFT IN THE PATIENT. THERE WAS NO REPORTED PATIENT ADVERSE EVENT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00567 | 2022-05-19 | 2019-09-16 | I, n, j, u, r, y | Manufacturer report | EXACTECH | PRM STM INSRTR EXTRCT | EXACTECH, INC. |  | 301-07-10 |  | IT WAS ORIGINALLY REPORTED FROM OUS THAT THE INSTRUMENT WAS DAMAGED DURING SURGERY. NO COMPLICATIONS TO PATIENT. IT WAS CONFIRMED THAT ALL PIECES WERE RETRIEVED FROM THE PATIENT. THE AGENCY WILL NOT PROVIDE SURGEON/HOSPITAL INFORMATION. THE PATIENT¿S HEATH WAS STABLE LEAVING THE OR. NO PIECES WERE LEFT IN THE PATIENT. NO DELAY TO SURGERY. AGENT WAS PRESENT AT TIME OF SURGERY. NEW INFORMATION RECEIVED ON 1 APRIL 2021 STATES: ¿THE INSTRUMENT COULD HAVE BEEN BROKEN DURING THE SURGERY. HOWEVER, THE BREAKAGE HAS BEEN NOTICED AT THE RECOVERY OF THE INSTRUMENTATION-SET BY OUR LOGISTICS DEPT., AFTER THE CLEANING REALIZED BY THE STERILIZATION TEAM OF THE HOSPITAL. NO INFORMATION HAS BEEN RECEIVED FROM THE HOSPITAL REGARDING THIS ISSUE AND [THE AGENCY] DIDN¿T RECEIVE ANY REQUEST FOR COMPLEMENTARY INFORMATION REGARDING THIS ISSUE AS WELL.¿ MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR 10X OPTICAL UNLESS OTHERWISE SPECIFIED) - THE MARKINGS ON THE SURFACE OF THE STRIKE PLATE APPEAR CONSISTENT WITH THE DEVICE BEING IMPACTED DURING SURGICAL USE. THE THREAD DAMAGE APPEARS CONSISTENT WITH REPEATED USE AND CROSS THREADING THE DEVICE WITH A MATING STEM. THE DEFORMATION APPEARS CONSISTENT WITH TOOL MARKS AND/OR CONTACTING A MATING HUMERAL STEM. THE HANDLE OF THE STEM INSERTER/EXTRACTOR HAS FRACTURED. THIS APPEARS CONSISTENT WITH EXPOSURE TO IMPACTION FORCES FROM REPEATED USE AND SUBSEQUENT STERILIZATION CYCLES WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FAILURE. ALL OTHER ASPECTS OF THE DEVICES APPEAR UNREMARKABLE. DESIGN-RELATED ISSUES: EXACTECH IS AWARE OF 31 COMPLAINT REPORTS INVOLVING A BROKEN OR DEFORMED STEM INSERTER/EXTRACTOR SINCE 2008. MFG-RELATED ISSUES: EXACTECH IS AWARE OF 1 OTHER COMPLAINT INVOLVING PARTS FROM THIS MANUFACTURING LOT OF (B)(4) UNITS, HOWEVER THE FAILURE MODES DO NOT APPEAR RELATED. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. USER-RELATED ISSUES: THERE ARE NO USER-RELATED ISSUES REPORTED THAT CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR AND RACR. THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF REPEATED USE AND CROSS-THREADING THE DEVICE WITH A MATING STEM WHICH LED TO THREAD DAMAGE. THE FRACTURED HANDLE WAS LIKELY THE RESULT OF REPEAT EXPOSURE TO IMPACTION FORCES AND SUBSEQUENT STERILIZATION CYCLES WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FAILURE. IFU 700-096-181: INSTRUMENT INSPECTION. VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THIS WAS REPORTED TO HAVE OCCURRED DURING SURGERY, THE HANDLE IS BROKEN, AND THE THREADS ARE DAMAGED. IT IS STATED THAT NO PIECES WERE LEFT IN THE PATIENT, NO PATIENT ADVERSE EVENT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00566 | 2022-05-19 | 2019-03-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTEC | MIDDLE SEGMENT TRIAL 75 | EXACTECH, INC. |  | 309-10-75 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE TAB BROKE OFF DURING TRIAL ASSEMBLY. NO PIECES FELL INTO THE SURGICAL SITE, IT WAS STILL ATTACHED TO THE MAIN PART OF THE TRIAL. POST OP THE PATIENT WAS FINE. THE ISSUE WAS RESOLVED BY USING 2ND TRIAL SEGMENT IN THE STERILE SET. THE AGENT WAS PRESENT AT SURGERY WITH NO DELAY AND NO ADVERSE EVENT TO THE PATIENT. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: SEX, WEIGHT, ETHNICITY, RACE. VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR L0X OPTICAL UNLESS OTHERWISE SPECIFIED) THE BROKEN MATING FEATURE APPEARS CONSISTENT WITH NORMAL USE DURING TRIALING WITH OTHER MATING DEVICE (E.G. PROXIMAL BODY TRIALS). ACCORDING TO THE INFORMATION PROVIDED, NO PIECES WERE LEFT IN THE PATIENT. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE AND CONSISTENT WITH YEARS OF USE. THE OTHER THREE DEVICES WERE FOUND TO BE BROKEN DURING AN INSTRUMENT SERVICING REVIEW. BECAUSE THE MIDDLE SEGMENT TRIAL IS USED IN TOTAL HUMERAL RECONSTRUCTIVE SURGERIES, SALES DATA FOR MIDDLE SEGMENT IMPLANTS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE BETWEEN 1% AND 5%. THIS IS CONSIDERED "MODERATE" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE DETAILED BELOW. A CORRECTION AND REMOVAL COMMITTEE (CRC) MEETING WAS HELD TO ADDRESS THESE BROKEN INSTRUMENTS FOUND IN SERVICING. BASED ON THE EVIDENCE OF TWO REPORTED BREAKAGES DURING USE IN SURGERY AND NO COMPLAINTS OF RESULTING INJURY OR NEED FOR SUBSEQUENT SURGICAL INTERVENTION, IT IS CONCLUDED THAT POTENTIAL RISK TO HEALTH IS LOW AND CONTINUED USE OF THE DEVICE IS ACCEPTABLE. DESIGN-RELATED ISSUES: SINCE INTRODUCTION OF THE EQUINOXE HUMERAL RECONSTRUCTION PROSTHESIS SYSTEM IN 2015, THERE HAVE BEEN THREE LOGGED COMPLAINT REPORTS INVOLVING 5 MIDDLE SEGMENT TRIALS, TWO OF WHICH DOCUMENTED BREAKAGE DURING USE IN SURGERY. A REVIEW OF THE RMR AND RACR WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR AND RACR FOR THIS MIDDLE SEGMENT TRIAL, RMR-00014 REV - AND RACR-00017 REV -, WERE REVIEWED. THE RISK IS CAPTURED IN LINE 9.13. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF THE TRIALING PROCESS. IFU 700-096-181: INSTRUMENT INSPECTION VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE WAS NO PATIENT ADVERSE EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00564 | 2022-05-19 | 2018-10-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | 3.5MM HEX SCREWDRIVER | EXACTECH, INC. |  | 521-11-04 |  | IT WAS REPORTED FROM THE US THAT DURING A FUNCTIONAL CHECK BEFORE AN ORTHOPEDIC SURGERY THE HEX SCREWDRIVER TIP WAS FOUND TO BE DAMAGED/BROKEN. THE INSTRUMENT WAS REPLACED. THERE WAS NO PATIENT INVOLVEMENT. THERE WAS NO DELAY TO SURGERY OR PATIENT ADVERSE EVENT. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR L0X OPTICAL UNLESS OTHERWISE SPECIFIED). THE FRACTURE SURFACE APPEARS CONSISTENT WITH APPLYING A BENDING MOMENT TO THE TIP OF THE SCREWDRIVER CAUSING BRITTLE FRACTURE OF THE HEX TIP. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE. DESIGN-RELATED -THE DESIGN OF THE GPS 3.5MM HEX SCREWDRIVER HAS BEEN IN THE FIELD SINCE 2010. EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINT REPORTS INVOLVING THIS DEVICE DUE TO TIP FRACTURE SINCE THE DESIGN WAS UPDATED IN 2013. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RISK MANAGEMENT REPORT. THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF APPLYING A BENDING MOMENT WITH THE TIP OF THE DEVICE ENGAGED, WHICH LED TO FRACTURE INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE HEX TIP. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT IMPACT. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00565 | 2022-05-19 | 2018-10-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRU TIBIAL ADAPTOR PLATE TRIAL/IMPLANT, SIZE 3.5 | EXACTECH, INC. |  | 02-029-25-4035 |  | IT WAS REPORTED FROM THE US THAT AT SOME POINT IN THE ORTHOPEDIC SURGICAL CASE THE MIDDLE BROKE IN A TIBIAL ADAPTOR PLATE TRIAL/IMPLANT. IT IS UNKNOWN WHEN THE DEVICE BROKE, LIKELY ON BACK TABLE. THERE WERE NO PIECES THAT FELL INTO THE SURGICAL SITE. THEREFORE; NO PIECES WERE LEFT IN THE PATIENT. THERE WAS NO DELAY TO SURGERY AND NO ADVERSE EVENT TO THE PATIENT. THE AGENT WAS PRESENT AT THE TIME OF SURGERY. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. HE ROOT CAUSE ASSOCIATED WITH THE CAPA RESIDES IN THE MISUSE/ MISTREATMENT OF THE ADAPTOR PLATES OR CANNOT BE IDENTIFIED DUE TO INSUBSTANTIAL EVIDENCE . THE CAPA TEAM HAS DETERMINED. HOWEVER; THAT USE CONSIDERATIONS INDICATE THAT THE FLEXURE DESIGN IS SUSCEPTIBLE TO DAMAGE FROM SAID MISUSE/MISTREATMENT. DESIGN A REVIEW OF THE TOLERANCE STACK-UP PERFORMED IN TM-2017-0417 WAS CONDUCTED. AS PART OF THIS INVESTIGATION, THE TOLERANCE STACK-UP BETWEEN THE TRULIANT TIBIAL TRIAL ADAPTOR PLATES AND THE CORRESPONDING TRULIANT AND LOGIC TIBIAL TRAYS WAS REVISITED TO CONFIRM IT HAD BEEN ADEQUATELY PERFORMED. THE REVIEW WAS FOUND TO BE ACCURATE. THERE IS NO NEW INFORMATION THAT HAS BEEN FOUND THAT WOULD ALTER THE RESULTS OR METHODS OF SAID ANALYSIS. IN ADDITION, ATTEMPTS WERE MADE IN THE LABORATORY TO REPLICATE FAILURE MODES WITHOUT SUCCESS. THE DESIGN FUNCTIONS AS INTENDED UNDER WORST CASE CONDITIONS. MANUFACTURING REVIEW OF THE DEVICE DHR CONCLUDED THAT ALL PARTS WERE MADE WITHIN SPECIFICATION. USE CONSIDERATIONS THE DESIGN DOES LEND TO BEING SUSCEPTIBLE TO BREAKAGE IF USED INCORRECTLY DUE TO THE USE OF ULTEM MATERIAL WITH A THINNER ASPECT UNDER LOAD. THIS WAS OBSERVED IN ANOTHER COMPLAINT WHEN A INSTRUMENT TRAY WAS DROPPED ON THE INSTRUMENT. IN ADDITION THE DEVICES BROKE AT SOME POINT DURING REPROCESSING AT ONE INSTITUTE. A REVIEW WAS CONDUCTED AT THIS INSTITUTE¿S REPROCESSING PROCEDURES AND RESULTS INDICATED THAT CORRECT REPROCESSING INSTRUCTIONS PROVIDED BY EXACTECH WAS FOLLOWED. A REVIEW OF THE INSTRUMENT TRAY CASE INDICATED THAT IT IS POSSIBLE THAT THE ADAPTOR PLATES MAY NOT BE FULLY CONSTRAINED DURING REPROCESSING. DUE TO THE NATURE OF THE REPROCESSING PROCESS IT IS POSSIBLE THE PLATES BECAME DISLODGED AND DAMAGED ELSEWHERE IN THE MACHINE. HUMAN ERROR HUMAN ERROR WAS CONSIDERED AS PART OF THE INVESTIGATION AND FOUND TO BE THE ROOT CAUSE OF ON COMPLAINT. THE ADAPTOR PLATE THAT WAS USED IN SAID COMPLAINT WAS THE INCORRECT SIZE RELATIVE TO THE TIBIAL TRAY THAT HAD BEEN IMPLANTED. THIS MAY HAVE CAUSED THE PLATE TO DISLODGE UNPREDICTABLY. IN ADDITION, A HOHMANN RETRACTOR WAS USED TO REMOVE THE PLATE AND NOT THE TRULIANT DROP ROD HANDLE, AS STATED PER THE OPERATIVE TECHNIQUE, 712-35-31. THE RETRACTOR IS CURVED AND FEATURES A MUCH SMALLER NOSE WHICH COULD INADVERTENTLY FACILITATE CONTACT WITH THE FLEXURE DURING REMOVAL. THE GEOMETRY OF THE TRULIANT DROP ROD HANDLE, IN ADDITION TO AN APPROPRIATELY CONSTRAINED ADAPTOR PLATE, AS DESIGNED CREATES A SCENARIO IN WHICH CONTACT WITH THE FLEXURE IS EXTREMELY DIFFICULT THE ROOT CAUSE ASSOCIATED WITH THE CAPA RESIDES IN THE MISUSE/MISTREATMENT OF THE ADAPTOR PLATES OR CANNOT BE IDENTIFIED DUE TO INSUBSTANTIAL EVIDENCE. THE CAPA TEAM HAS DETERMINED, HOWEVER, THAT USE CONSIDERATIONS INDICATE THAT THE FLEXURE DESIGN IS SUSCEPTIBLE TO DAMAGE FROM SAID MISUSE/MISTREATMENT. IFU 700-096-181: INSTRUMENT INSPECTION VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. NO ADVERSE EVENT TO THE PATIENT, THE TRIAL WAS BROKEN. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00563 | 2022-05-19 | 2022-04-20 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 38MM GLENOSPHERE | EXACTECH, INC. | 320-01-38 | 320-01-38 |  | AS REPORTED, THIS (B)(6) PATIENT EXPERIENCED A MIDSHAFT HUMERAL FRACTURE. FISSURE OF THE HUMERAL SHAFT DURING IMPACTION OF THE HUMERAL STEM. INSTALLATION OF 3 ORTHOCORD CERCLAGES TO STRENGTHEN THE HUMERUS WEAKENED BY THE SEQUELAE OF OLD FRACTURE. THE CASE REPORT FORM INDICATES THIS EVENT IS DEFINITELY NOT RELATED TO DEVICES AND UNLIKELY RELATED TO PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | CONCOMITANT DEVICE(S): 300-01-13 - EQUINOXE, HUMERAL STEM PRIMARY, PRESS FIT 13MM; 320-38-00 - EQUINOXE REVERSE 38MM HUMERAL LINER +0; 320-10-00 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0; 320-15-04 - RS GLENOID PLATE R POST AUG, 8 DEG, RIGHT. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00571 | 2022-05-19 | 2020-02-05 | I, n, j, u, r, y | Manufacturer report | EXACTECH | GLENOID IMPACTOR TIP | EXACTECH, INC. |  | 315-07-06 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGICAL PROCEDURE THE GLENOID IMPACTOR TIP BROKE IN HALF WHILE IMPACTING THE GLENOID. THE EVENT DID NOT LEAD TO SERIOUS INJURY OF PATIENT OR USER AND NO PARTS OF THE BROKEN DEVICE/DEVICE FRAGMENTS WERE LEFT IN THE PATIENT. THE DEVICE WAS THROWN AWAY IN SPD. THE AGENCY IS INACTIVE, NO FURTHER INFORMATION IS AVAILABLE. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INITIAL REPORTER ALSO SENT REPORT TO FDA IS UNKNOWN. DEVICE EVALUATED BY MFR: NOT RETURNED. FINDINGS: THE REPORTED FRACTURED GLENOID IMPACTOR TIP WAS LIKELY THE RESULT OF CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE WHICH LIKELY RESULTED FROM ABNORMAL OR AGGRESSIVE USE. THE RMR AND RACR FOR THIS GLENOID IMPACTOR TIP, RMR-00007 REV B AND RACR-00010 REV B, WERE REVIEWED. THE RISK IS CAPTURED. BASED ON SIMILAR PREVIOUS COMPLAINT INVESTIGATIONS, THE REPORTED FRACTURED DEVICE REPORTED WAS LIKELY THE RESULT OF CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE WHICH WAS LIKELY CAUSED BY ABNORMAL OR AGGRESSIVE USE INCONSISTENT WITH THE INTENDED USE OF THIS DEVICE. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THE EVENT DID NOT LEAD TO SERIOUS INJURY OF PATIENT OR USER AND NO PARTS OF THE BROKEN DEVICE/DEVICE FRAGMENTS WERE LEFT IN THE PATIENT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00591 | 2022-05-20 | 2022-04-29 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 9MM | 02-012-44-4009 |  | AS REPORTED BY THE LEGAL TEAM, APPROXIMATELY SIX YEARS POST INITIAL RIGHT TKA, MALE PATIENT IS EXPERIENCING CONSTANT PAIN, SWELLING AND INSTABILITY. PATIENT NEEDS A REVISION; HE HAS THE SAME FINDINGS OF PREMATURE POLY WEAR WITH OSTEOLYSIS AND POSSIBLE FEMUR LOOSENING. A YEAR AFTER PATIENT'S RTKR, HE REQUIRED AN ARTHROSCOPY WITH EXTENSIVE SYNOVECTOMY AND LYSIS OF ADHESIONS. INFECTION HAS BEEN RULED OUT. NO ADDITIONAL INFORMATION AVAILABLE. THE DEVICE IS NOT AVAILABLE FOR EVALUATION,. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00594 | 2022-05-20 | 2022-05-02 | I, n, j, u, r, y | Manufacturer report | OPTETRACK PS TIBIAL INSERT SIZE 3 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | OPTETRACK PS TIBIAL INSERT SIZE 3 13MM | UNKNOWN |  | AS REPORTED, PATIENT WAS SCHEDULED FOR A REVISION DUE TO A POSSIBLE INFECTION. THE SURGEON REQUESTED PS REVISION COMPONENTS THOUGH THE PRIMARY KNEE WAS A CC KNEE. SURGEON REQUESTED A 13MM SIZE 3 PS INSERT FOR A POLY SWAP FOR A PATIENT WITH A POSSIBLE INFECTION. X-RAYS WERE RECEIVED AND IMPLANTS WERE IDENTIFIED AT OPTETRAK CC IMPLANTS. SALES REP WAS UNAWARE OF SURGEON¿S REQUEST FOR PS IMPLANTS, SALES REP WAS AT THE FACILITY WITH ONLY CC POLY INSERTS. SALES REP NOTIFIED SURGEON THAT WE DID NOT HAVE PS IMPLANTS IMMEDIATELY AFTER HE FINISHED HIS PREVIOUS SURGERY. HE DECIDED TO GO AHEAD WITH CC POLY. IN SURGERY, IT WAS REVEALED THE PATIENT HAD A PROXIMAL TIBIA SPACER ON HIS TIBIAL TRAY. THIS MADE IT IMPOSSIBLE TO IMPLANT THE CC POLYETHYLENE WITH THE STIFFENING SCREW. SURGEON WAS NOT HAPPY WITH THIS DEVELOPMENT AND OPTED TO RESCHEDULE THE PATIENT FOR ANOTHER POLY SWAP TO THE PS POLYETHYLENE. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO TIBIAL INSERT DISCARDED AFTER SURGERY. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00582 | 2022-05-20 | 2019-03-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | HUM LNR IMPACT TIP 38M | EXACTECH, INC. |  | 321-07-38 |  | IT WAS REPORTED BY EXACTECH KOREA THAT THE HUMERAL LINER IMPACTOR TIP WAS BROKEN WHEN THE SURGEON IMPACTED THE PRE-ASSEMBLED IMPLANT (STEM,TRAY,LINER) DURING AN ORTHOPEDIC SURGERY. THE SURGEON HIT 2-3 TIMES FOR INSERTING IMPLANT. 4-5 MINUTES DELAY EACH TIME. THE PATIENT'S HEALTH WAS STABLE LEAVING THE OR. ATTEMPTS TO REQUEST ADDITIONAL INFORMATION WERE SENT TO THE REPRESENTATIVE, AND A RESPONSE WAS RECEIVED. NO PARTS/PIECES OF THE DEVICE FELL INTO THE SURGICAL SITE. THERE WERE NO PROBLEMS NOTED AT DISCHARGE. THERE IS NO PATIENT DEMOGRAPHIC INFORMATION PROVIDED. MULTIPLE ATTEMPTS HAVE BEEN MADE FOR INFORMATION; NO ADDITIONAL INFORMATION IS AVAILABLE FROM THE CONTACTS ABOUT THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. ADDITIONAL MFR NARRATIVE INFORMATION WAS NOT PROVIDED FOR: PT. INFO. INITIAL REPORTER: COUNTRY -KOREA. INITIAL REPORTER ALSO SENT REPORT TO FDA? UNKNOWN. THE DEVICE WAS NOT AVAILABLE FOR ANALYSIS. ENGINEERING EVALUATION: DESIGN-RELATED ISSUES: THE DESIGN OF THE HUMERAL LINER IMPACTOR TIP HAS BEEN IN THE FIELD SINCE 2006. EXACTECH IS AWARE OF 6 SIMILAR COMPLAINT REPORTS INVOLVING 7 BROKEN HUMERAL LINER IMPACTOR TIPS SINCE 2008. BECAUSE THE HUMERAL LINER IMPACTOR TIP IS USED IN REVERSE TOTAL SHOULDER SURGERIES, SALES DATA FOR HUMERAL LINERS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINT INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 30 UNITS, WHICH HAS BEEN IN THE FIELD SINCE 2017. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS HUMERAL LINER IMPACTOR TIP, 750-2005-021-RMR INSTRUMENTS REV P, WAS REVIEWED. THE RISK IS CAPTURED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR. THE BROKEN HUMERAL LINER IMPACTOR TIP REPORTED WAS LIKELY THE RESULT OF A STRESS RISER INTRODUCED DURING IMPACTION, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. ¿ CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00588 | 2022-05-20 | 2019-08-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | HUM LNR IMPACT TIP 38M | EXACTECH, INC. |  | 321-07-38 |  | IT WAS REPORTED FROM OUS THAT THE HUMERAL LINER IMPACT TIP BROKE. THERE WAS A 4-5 MINUTES DELAYED. THERE WAS NO REPORTED PATIENT INJURY. THE REP WAS PRESENT AT THE TIME OF THE SURGERY. INITIALLY, HUMERAL LINER IMPACTOR TIP WAS BROKEN WHEN SURGEON IMPACTED PRE-ASSEMBLED IMPLANT (STEM,TRAY,LINER), AS THE SURGEON HIT 2-3 TIMES FOR INSERTING IMPLANT. AFTER BROKEN THESE IMPACTOR TIPS, THE REP PREPARED AN IMPACTOR TIP INSTEAD OF 38 IMPACTOR TIPS. PATIENT WAS DISCHARGED WITHOUT COMPLICATIONS. NO PIECES WERE LEFT IN THE PATIENT. THE DEVICE WAS NOT RECEIVED FOR EVALUATION. ATTEMPT WAS MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REP AND A RESPONSE WAS RECEIVED. PER THE REP, NO PARTS/PIECES OF THE DEVICE FELL INTO THE SURGICAL SITE, GOT LEFT IN THE SURGICAL SITE OR WERE REMOVED FROM THE SURGICAL SITE. THE EVENT DID NOT RESULT IN A REVISION. NO DEMOGRAPHIC INFORMATION IS AVAILABLE. NO OTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS ABOUT THE PATIENT/DEVICE/EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. COUNTRY: KOREA. DEVICE NOT RETURNED, CAPA INVESTIGATION : INVESTIGATION # AND DESCRIPTION: ALTHOUGH STILL UNDER THE THRESHOLD OF <0.5% FOR ¿VERY LOW¿ OCCURRENCE RATE, DUE TO A TREND OF THE RECENT COMPLAINTS HAVING APPROXIMATELY A 5X HIGHER OCCURRENCE RATE IN THE KOREAN MARKET [ CAPA ] ESTABLISHED TO INVESTIGATE ADDITIONAL POTENTIAL CAUSES OF THE HIGHER OCCURRENCE RATE SPECIFIC TO THE KOREAN MARKET. PREVIOUSLY, CAPA IDENTIFIED IMPROPER REPROCESSING PROCEDURES AS A ROOT CAUSE FOR A HIGHER INCIDENCE OF FRACTURED RADEL TIBIAL INSERT TRIALS IN KOREA AND SINGAPORE. BECAUSE THE DEVICES INVESTIGATED IN CAPA ARE MADE FROM THE SAME TYPE OF MATERIAL AS THE HUMERAL LINER IMPACTOR TIPS AND THE GEOGRAPHIC LOCATION OF THE FAILURES IS THE SAME, CAPA IS INVESTIGATING THE POSSIBILITY OF A SIMILAR CORRELATION BETWEEN THE HIGHER INCIDENCE OF FRACTURES AND THE REPROCESSING PROCEDURE FOR THE IMPACTOR TIPS. IN KOREA, THE HUMERAL LINERS HAVE SALES OF (B)(4) TO DATE SINCE 2011. THIS CORRESPONDS TO A (B)(4) OCCURRENCE RATE WHICH IS STILL BELOW OUR THRESHOLD OF 0.5%. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: BASED ON INFORMATION AVAILABLE AT THIS TIME, NO CONCLUSION CAN BE MADE REGARDING THE HIGHER MARKET SPECIFIC OCCURRENCE. ADDITIONAL INVESTIGATION FOR A MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE RELATED TO MARKET SPECIFIC OCCURRENCE IS STILL ONGOING {AT THE TIME OF THE EVALUATION}. IFU 700-096-181: INSTRUMENT INSPECTION VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. FOLLOW REPROCESSING INSTRUCTIONS. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00596 | 2022-05-20 | 2022-04-01 | I, n, j, u, r, y | Manufacturer report | EQUINOXE CAGE GLENOID MEDIUM, ALPHA | PROSTHESIS, SHOULDER, SEMI-CONSTRAINED, METAL/POLYMER CEMENTED | EXACTECH, INC. |  | 314-13-03 |  | AS REPORTED, APPROXIMATELY 10 YEARS POST OP THE INITIAL TSA, THIS (B)(6) YEAR OLDFEMALE PATIENT WAS REVISED DUE ASEPTIC HUMERAL LOOSENING. PATIENT HAD PROGRESSIVE HUMERAL LUCENCY SINCE 2016 ON XRAY - NOW WITH GLENOID LOOSENING NOTED IN 2/2022. NO SPECIFIC MOI NOTED. SURGEON REMOVED THE STANDARD HUMERAL STEM, REPLICATOR PLATE, TORQUE SCREW, HUMERAL HEAD, GLENOID. THE CASE REPORT FORM INDICATES THIS EVENT IS POSSIBLY RELATED TO DEVICES AND/OR PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | PENDING EVALUATION (CONCOMITANT DEVICE(S): REPLICATOR PLATE. TORQUE SCREW. HUMERAL HEAD. GLENOID. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00595 | 2022-05-20 | 2022-04-29 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 38MM GLENOSPHERE | EXACTECH, INC. | 320-01-38 | 320-01-38 |  | AS REPORTED, APPROXIMATELY 5 YEARS POST OP THE INITIAL RIGHT RTSA, THIS PATIENT WAS REVISED FOR PAIN. BASEPLATE SCREWS WERE REMOVED AND REPLACED. GLENOSPHERE, GLENOSPHERE LOCKING SCREW AND HUMERAL LINER WERE ALSO REPLACED. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICE IS NOT RETURNING, IT WAS DISCARDED. | CONCOMITANT DEVICE(S): 320-20-18, (B)(4)- COMPRESSION SCREW/LOCKING CAP KIT 4.5M X 18MM SCREW WHITE. 320-20-18, (B)(4)- COMPRESSION SCREW/LOCKING CAP KIT 4.5M X 18MM SCREW WHITE. 320-20-34, (B)(4)- EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 34MM. 320-20-30, (B)(4)- EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 30MM. 320-15-05, (B)(4)- EQ REV LOCKING SCREW. 320-38-00, (B)(4)- EQUINOXE REVERSE 38MM HUMERAL LINER +0. 320-15-01, (B)(4)- EQ REV GLENOID PLATE. 300-01-11, (B)(4)- EQUINOXE, HUMERAL STEM PRIMARY, PRESS FIT 11MM. 320-20-00, (B)(4)- EQ REVERSE TORQUE DEFINING SCREW KIT. 320-10-00, (B)(4)- EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00585 | 2022-05-20 | 2019-05-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | HUM LNR IMPACT TIP 42M | EXACTECH, INC. |  | 321-07-42 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH AN INSTRUMENT. THE SURGEON WAS ASSEMBLING THE HUMERAL IMPLANTS OUTSIDE OF THE PATIENT AS HE USUALLY DOES. UPON IMPACTION OF THE HUMERAL LINER THE PLASTIC LINER IMPACTOR BROKE INTO TWO PIECES. THE BROKEN PIECES WERE REMOVED FROM THE FIELD AND THE SURGERY PROGRESSED AS USUAL. THE DEVICE WAS NOT AVAILABLE FOR ANALYSIS. THE REPRESENTATIVE WAS PRESENT DURING THE SURGERY. THERE WAS NO DELAY. THE PATIENT WAS NOT NEGATIVELY AFFECTED. THERE IS NO ADDITIONAL INFORMATION PROVIDED BY THE CONTACTS FOR THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH (B)(4) HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR PATIENT. INITIAL REPORTER: UNKNOWN. ENGINEERING EVALUATION PROVIDED. DESIGN-RELATED ISSUES: THE DESIGN OF THE HUMERAL LINER IMPACTOR TIP HAS BEEN IN THE FIELD SINCE 2007. EXACTECH IS AWARE OF 34 COMPLAINT REPORTS INVOLVING 42 BROKEN HUMERAL LINER IMPACTOR TIPS SINCE 2008. SALES DATA FOR ALL HUMERAL LINERS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4). THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) AND RISK ASSESSMENT AND CONTROLS REPORT (RACR) WERE CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR AND RACR FOR THIS HUMERAL LINER IMPACTOR TIP, RMR-00004 REV - AND RACR-00007 REV -, WERE REVIEWED. THE RISK IS CAPTURED, HOWEVER THE RACR IS BEING UPDATED UNDER (B)(4) TO CAPTURE AN APPROPRIATE HARM FOR THIS HAZARDOUS SITUATION. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, THE RISK IS CAPTURED IN THE RMR AND RACR, AND THE HARM WILL BE UPDATED. THE BROKEN IMPACTOR TIP REPORTED WAS LIKELY THE RESULT OF A STRESS RISER INTRODUCED DURING IMPACTION, WHICH LED TO CRACK INITIATION, PROPAGATION AND ULTIMATE FRACTURE OF THE DEVICE. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00593 | 2022-05-20 | 2022-04-21 | I, n, j, u, r, y | Manufacturer report | LOGIC CR TIB INSERT SLOPE+, SZ 5, 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC CR TIB INSERT SLOPE+, SZ 5, 13MM | 02-012-48-5013 |  | AS REPORTED, APPROXIMATELY 3 YEARS POST INITIAL TKA, THE 79 Y/O MALE PATIENT HAD A REVISION DUE TO POLY WEAR. SURGEON REQUESTED INSET AND IMP FOR POLY EXCHANGE. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PHOTOS RECEIVED. SALES REP WAS UNABLE TO OBTAIN X-RAYS. THE DEVICES ARE NOT AVAILABLE FOR EVALUATION AS HOSPITAL WOULD NOT ALLOW REMOVAL OF EXPLANTS. | CONCOMITANT MEDICAL PRODUCTS - THREE PEG PATELLA 38MM (CAT# 200-02-38 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00581 | 2022-05-20 | 2019-03-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECJ | HUM LNR IMPACT TIP 38M | EXACTECH, INC. |  | 321-07-38 |  | IT WAS REPORTED BY EXACTECH KOREA THAT THE HUMERAL LINER IMPACTOR TIP WAS BROKEN WHEN THE SURGEON IMPACTED THE PRE-ASSEMBLED IMPLANT (STEM,TRAY,LINER) DURING AN ORTHOPEDIC SURGERY. THE SURGEON HIT 2-3 TIMES FOR INSERTING IMPLANT. 4-5 MINUTES DELAY EACH TIME. THE PATIENT'S HEALTH WAS STABLE LEAVING THE OR. ATTEMPTS TO REQUEST ADDITIONAL INFORMATION WERE SENT TO THE REPRESENTATIVE, AND A RESPONSE WAS RECEIVED. NO PARTS/PIECES OF THE DEVICE FELL INTO THE SURGICAL SITE. THERE WERE NO PROBLEMS NOTED AT DISCHARGE. THERE IS NO PATIENT DEMOGRAPHIC INFORMATION PROVIDED. MULTIPLE ATTEMPTS WERE MADE FOR INFORMATION, NO ADDITIONAL INFORMATION IS AVAILABLE FROM THE CONTACTS ABOUT THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. (B)(6). FINDINGS: THE FRACTURED HUMERAL LINER IMPACTOR TIP APPEARS CONSISTENT WITH A STRESS RISER INTRODUCED DURING IMPACTION, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE. MOST LIKELY UNDERLYING CAUSE: BASED ON PREVIOUS SIMILAR INVESTIGATIONS, THE FRACTURED HUMERAL LINER IMPACTOR TIP REPORTED WAS LIKELY THE RESULT OF A STRESS RISER INTRODUCED DURING IMPACTION, WHICH LED TO CRACK INITIATION PROPAGATION, AND ULTIMATE FRACTURE. RMR REVIEW: THE RMR FOR THIS HUMERAL LINER IMPACTOR TIP, 750-2005-021-RMR-INSTRUMENTS REV P, WAS REVIEWED. THE RISK IS CAPTURED. THE DEVICE WAS NOT AVAILABLE FOR ANALYSIS. IFU 700-096-181: INSTRUMENT INSPECTION. VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. THERE WERE NO PARTS/PIECES THAT FELL INTO THE SURGICAL SITE. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00589 | 2022-05-20 | 2019-08-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | A+ BROACH HANDLE RIGHT | EXACTECH, INC. |  | 167-02-00 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE INSTRUMENTATION. THE LOCKING MECHANISM ON THE A+ BROACH HANDLE WAS BROKEN AFTER THE CASE; THE STAFF BELIEVE IT HAPPENED SOMETIME DURING THE WASHING PROCESS. THERE WAS A DELAY TO SURGERY, BUT THE PATIENT WAS NOT ADVERSELY AFFECTED. THE DEVICE WAS RECEIVED FOR ANALYSIS, IT WAS DISASSEMBLED. THERE IS NO OTHER INFORMATION ABOUT THE EVENT/DEVICE PROVIDED BY THE CONTACTS.. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. UNKNOWN. DESIGN-RELATED ISSUES: THE DESIGN OF THE NOVATION A+ BROACH HANDLE HAS BEEN IN THE FIELD SINCE 2009. (B)(4). THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THE NOVATION A+ BROACH HANDLE IN THIS COMPLAINT WAS MANUFACTURED IN 2010. THE DESIGN OF THE DEVICE HAS SINCE BEEN UPDATED ((B)(4)) TO REPLACE THE PIVOT PIN THAT IS RETAINED WITH A SPIRAL RETAINING RING WITH A PIVOT RIVET THAT IS MECHANICALLY DEFORMED. ENDURANCE TESTING WAS PERFORMED ON THIS MECHANICALLY DEFORMED (SWAGED) DESIGNED AND IS DOCUMENTED IN TM-2012-477. ADDITIONALLY, ENDURANCE TESTING WAS PERFORMED ON A SPECIAL A+ STRAIGHT BROACH HANDLE WITH A PROTOTYPE DESIGN CHANGE TO THE RELEASE LATCH AND LATCH PIN. RESULTS OF THIS TEST WERE USED TO DETERMINE THE NEED FOR AN UPDATED ROLL PIN, RELEASE LATCH, AND LATCH PIN. SEE TM-2013-264 AND DCRTC-14-0171. MFG-RELATED ISSUES: EXACTECH IS AWARE OF 3 OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT, WHICH HAVE BEEN IN THE FIELD SINCE 2010. THIS DEVICE WAS MANUFACTURED BEFORE THE DESIGN WAS UPDATED. THE DEVICE HISTORY RECORD WAS REVIEWED AS PART OF [ANOTHER COMPLAINTE] FOR THIS DEVICE, AND ALL OF THE PARTS WERE ACCEPTED WITH CONFORMANCE TO THE DEVICE SPECIFICATIONS. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) AND THE RISK ASSESSMENT AND CONTROLS REPORT WERE CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR AND RACR FOR THIS NOVATION ANTERIOR BROACH HANDLE, RMR-00034 REV- AND RACR-00037 REV-, WERE REVIEWED. THE RISK IS CAPTURED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR AND RACR. THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF WEAR DUE TO YEARS OF NORMAL USE, WHICH LED TO THE DISASSEMBLY OF THE LEVER PIVOT PIN, STAINLESS-STEEL RETAINING RING, COMPRESSION SPRING AND THE BROACH LOCKING PIN. THE DESIGN HAS BEEN UPDATED TO REPLACE THE LEVER PIVOT PIN AND STAINLESS-STEEL RETAINING RING WITH A LEVER PIVOT RIVET TO IMPROVE THE ROBUSTNESS OF THE DESIGN. IFU 700-096-181: INSTRUMENT INSPECTION VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. ¿ CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. FOLLOW REPROCESSING PROCEDURES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00587 | 2022-05-20 | 2019-08-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | HUM LNR IMPACT TIP 42M | EXACTECH, INC. |  | 321-07-42 |  | IT WAS REPORTED FROM OUS THAT THE HUMERAL LINER IMPACT TIP BROKE. THERE WAS A 4-5 MINUTES DELAYED. THERE WAS NO REPORTED PATIENT INJURY. THE REP WAS PRESENT AT THE TIME OF THE SURGERY. INITIALLY, HUMERAL LINER IMPACTOR TIP WAS BROKEN WHEN SURGEON IMPACTED PRE-ASSEMBLED IMPLANT (STEM,TRAY,LINER), AS THE SURGEON HIT 2-3 TIMES FOR INSERTING IMPLANT. AFTER BROKEN THESE IMPACTOR TIPS, THE REP PREPARED AN IMPACTOR TIP INSTEAD OF 38 IMPACTOR TIPS. PATIENT WAS DISCHARGED WITHOUT COMPLICATIONS. NO PIECES WERE LEFT IN THE PATIENT. THE DEVICE WAS NOT RECEIVED FOR EVALUATION. AN ATTEMPT WAS MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REP AND A RESPONSE WAS RECEIVED. PER THE REP, NO PARTS/PIECES OF THE DEVICE FELL INTO THE SURGICAL SITE, GOT LEFT IN THE SURGICAL SITE OR WERE REMOVED FROM THE SURGICAL SITE. THE EVENT DID NOT RESULT IN A REVISION. NO DEMOGRAPHIC INFORMATION IS AVAILABLE. NO OTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS ABOUT THE PATIENT/DEVICE/EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INITIAL REPORTER COUNTRY -(B)(6). ENGINEERING ANALYSIS: THE DEVICE WAS NOT AVAILABLE FOR ANALYSIS. INVESTIGATION # AND DESCRIPTION: ALTHOUGH STILL UNDER THE THRESHOLD OF <0.5% FOR ¿VERY LOW¿ OCCURRENCE RATE, DUE TO A TREND OF THE RECENT COMPLAINTS HAVING APPROXIMATELY A 5X HIGHER OCCURRENCE RATE IN THE (B)(6) MARKET, [CAPA] WERE] ESTABLISHED TO INVESTIGATE ADDITIONAL POTENTIAL CAUSES OF THE HIGHER OCCURRENCE RATE SPECIFIC TO THE KOREAN MARKET. PREVIOUSLY, A CAPA IDENTIFIED IMPROPER REPROCESSING PROCEDURES AS A ROOT CAUSE FOR A HIGHER INCIDENCE OF FRACTURED RADEL TIBIAL INSERT TRIALS IN KOREA AND SINGAPORE. BECAUSE THE DEVICES INVESTIGATED IN CAPA (B)(4) ARE MADE FROM THE SAME TYPE OF MATERIAL AS THE HUMERAL LINER IMPACTOR TIPS AND THE GEOGRAPHIC LOCATION OF THE FAILURES IS THE SAME, CAPA (B)(4) IS INVESTIGATING THE POSSIBILITY OF A SIMILAR CORRELATION BETWEEN THE HIGHER INCIDENCE OF FRACTURES AND THE REPROCESSING PROCEDURE FOR THE IMPACTOR TIPS. FINDINGS: FOR THIS COMPLAINT -THE FRACTURED HUMERAL LINER IMPACTOR TIP WAS REPORTED TO HAVE FRACTURED WHILE BEING IMPACTED DURING SURGERY. NO OR LOW RISK TO HEALTH. EXACTECH HAS RECEIVED 46 GLOBAL EXPERIENCE REPORTS DOCUMENTING BREAKAGE OF THE HUMERAL LINER IMPACTOR TIPS SINCE 2011. THESE IMPTACOR TIPS ARE USED FOR IMPLANTATION OF THE REVERSE SHOULDER HUMERAL LINERS, WHICH HAVE SALES OF 49,608 TO DATE SINCE 2011. THIS CORRESPONDS TO A 0.09% OCCURRENCE RATE. THIS WON¿T CAUSE INJURY TO THE PATIENT. OF NOTE, THERE HAS BEEN AN INCREASED NUMBER OF COMPLAINTS COMING FROM KOREA. THERE HAVE BEEN 8 COMPLAINTS SINCE 2011 FROM (B)(6). IN (B)(6), THE HUMERAL LINERS HAVE SALES OF 1,860 TO DATE SINCE 2011. THIS CORRESPONDS TO A 0.43% OCCURRENCE RATE WHICH IS STILL BELOW OUR THRESHOLD OF 0.5% HRI SCORE OF ¿3¿ ¿ POTENTIAL RISK TO HEALTH. EXACTECH HAS NOT RECEIVED ANY COMPLAINTS DETAILING PIECES OF THE HUMERAL LINER IMPACTION TIP BEING LEFT IN THE PATIENT. IF THE BREAKAGE HAPPENS, IT¿S HIGHLY NOTICEABLE AND THE HEALTHCARE PROFESSIONALS ARE UNLIKELY TO OMIT IT. THE BREAKAGE PIECES ARE USUALLY LARGE AND NOT HARD TO LOCATE. IT CAN BE CONFIRMED IF ALL BROKEN PIECES ARE COLLECTED BY CHECKING THE COMPLETENESS OF THE ASSEMBLED PART. IF THE BROKEN PIECE IS LEFT IN THE PATIENT¿S WOUND WITHOUT BEING NOTICED, IT COULD CAUSE INJURY TO THE PATIENT AND WOULD REQUIRE PROFESSIONAL MEDICAL INTERVENTION. THE MATERIAL OF THESE INSTRUMENTS, RADEL PER ETS-0205, IS INCLUDED IN THE BIOCOMPATIBILITY ASSESSMENT TECHNICAL MEMO, TM-2013-465. RADEL PASSES THE BIOCOMPATIBILITY REQUIREMENTS FOR SURGICAL USE, AND THERE ARE NO ANTICIPATED ADVERSE HEALTH CONSEQUENCES. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: BASED ON INFORMATION AVAILABLE AT THIS TIME, NO CONCLUSION CAN BE MADE REGARDING THE HIGHER MARKET SPECIFIC OCCURRENCE. ADDITIONAL INVESTIGATION FOR A MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE RELATED TO MARKET SPECIFIC OCCURRENCE IS STILL ONGOING. IFU 700-096-181: INSTRUMENT INSPECTION ¿ VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. ¿ CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. ¿ IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. FOLLOW REPROCESSING INSTRUCTIONS. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00586 | 2022-05-20 | 2019-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | HUM LNR IMPACT TIP 38M | EXACTECH, INC. |  | 321-07-38 |  | IT WAS REPORTED FROM OUS THAT THE HUMERAL LINER IMPACTOR TIP WAS BROKEN WHEN THE SURGEON IMPACTED PRE-ASSEMBLED IMPLANT (STEM, TRAY, LINER). THE SURGEON HIT 3-4 TIMES FOR INSERTING IMPLANT. SURGERY WAS ¿GOOD¿. THERE WAS A 5-MINUTE DELAY WITH NO PATIENT ADVERSE EVENT. TWO ATTEMPTS WERE MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REPRESENTATIVE AND A RESPONSE WAS RECEIVED. PER THE REPRESENTATIVE, NO PARTS/PIECES FELL INTO THE SURGICAL SITE. ¿THE OUTCOME WAS NOTHING SPECIAL¿. THERE IS NO ADDITIONAL INFORMATION PROVIDED BY THE CONTACTS FOR THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. COUNTRY - (B)(6). ENGINEERING ANALYSIS: THE DEVICE WAS NOT AVAILABLE FOR ANALYSIS. DESIGN-RELATED ISSUES: THE DESIGN OF THE HUMERAL LINER IMPACTOR TIP HAS BEEN IN THE FIELD SINCE 2006. EXACTECH IS AWARE OF 6 SIMILAR COMPLAINT REPORTS INVOLVING 7 BROKEN HUMERAL LINER IMPACTOR TIPS SINCE 2008. BECAUSE THE HUMERAL LINER IMPACTOR TIP IS USED IN REVERSE TOTAL SHOULDER SURGERIES, SALES DATA FOR HUMERAL LINERS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4). THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINT INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 30 UNITS, WHICH HAS BEEN IN THE FIELD SINCE 2015. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS HUMERAL LINER IMPACTOR TIP, 750-2005-021-RMR-INSTRUMENTS REV P, WAS REVIEWED. THE RISK IS CAPTURED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR. THE BROKEN HUMERAL LINER IMPACTOR TIP REPORTED WAS LIKELY THE RESULT OF A STRESS RISER INTRODUCED DURING IMPACTION, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00584 | 2022-05-20 | 2019-03-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | OFFSET CUP IMPACTOR | EXACTECH, INC. |  | 121-01-02 |  | THE DEVICE WAS RETURNED TO THE COMPANY BROKEN FROM THE US. NO FURTHER INFORMATION WAS PROVIDED. MULTIPLE ATTEMPTS WERE MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REPRESENTATIVE, BUT NO RESPONSES WERE RECEIVED. THERE IS NO INFORMATION ABOUT THE PATIENT OR EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DESIGN-RELATED ISSUES: THE DESIGN OF THE OFFSET CUP IMPACTOR HAS BEEN IN THE FIELD SINCE 2006. EXACTECH IS AWARE OF 9 SIMILAR COMPLAINT REPORTS INVOLVING CUP IMPACTORS DUE TO DISASSEMBLY OR BROKEN TIPS SINCE 2008. BECAUSE THE OFFSET CUP IMPACTOR IS USED IN TOTAL HIP SURGERIES, SALES DATA FOR ACETABULAR SHELLS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF (B)(4) UNITS, WHICH HAVE BEEN IN THE FIELD SINCE 2009. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS OFFSET CUP IMPACTOR, 750-2005-028-RMR-INSTRUMENTS REV U, WAS REVIEWED. THE RISK IS CAPTURED. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW¿, AND THE RISK IS CAPTURED IN THE RMR. THE FRACTURED OFFSET CUP IMPACTOR TIP REPORTED WAS LIKELY THE RESULT OF FATIGUE FRACTURE INITIATED OVER TIME WHICH LED TO ULTIMATE FRACTURE OF THE FLEXIBLE LOCKING TAB. THE DISASSEMBLY OF THE CUP IMPACTOR COMPONENTS WAS LIKELY THE RESULT OF YEARS OF NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES, WHICH LED TO DEGRADATION OF THE THREADLOCKER USED TO SECURE THE IMPACTOR STRIKE HEAD ON THE OFFSET CUP IMPACTOR HANDLE. IFU 700-096-181: INSTRUMENT INSPECTION VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO PATIENT OR EVENT INFORMATION FROM THE CONTACTS. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00583 | 2022-05-20 | 2019-03-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | GLENOID IMPACTOR TIP | EXACTECH, INC. |  | 315-07-06 |  | IT WAS REPORTED FROM THE US THAT THE INSTRUMENT WAS RETURNED TO THE COMPANY BROKEN. THERE WAS NO EVENT REPORT PROVIDED. THE COMPANY IS UNABLE TO OBTAIN ANY INFORMATION FROM THE AGENCY/REPRESENTATIVE AFTER MULTIPLE ATTEMPTS. NO ADDITIONAL INFORMATION IS AVAILABLE FROM THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH (B)(4) HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR PATIENT. EVENT DATE: USED 1ST DAY OF THE MONTH. INITIAL REPORTER: UNKNOWN. EVALUATION FINDINGS: THE FRACTURED GLENOID IMPACTOR TIP APPEARS CONSISTENT WITH ABNORMAL OR AGGRESSIVE USE. MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE: BASED ON SIMILAR PREVIOUS COMPLAINT INVESTIGATIONS, THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF ABNORMAL OR AGGRESSIVE USE INCONSISTENT WITH THE INTENDED USE OF THIS DEVICE. RMR REVIEW: THE RMR FOR THIS GLENOID IMPACTOR TIP, 750-2007-026-RMR-INSTRUMENTS REV K, WAS REVIEWED. THE RISK IS CAPTURED. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00580 | 2022-05-20 | 2019-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | STEM INSERTER | EXACTECH, INC. |  |  |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY A HIP INSTRUMENTS BROKE DURING THE CASE. CONTACT WAS MADE VIA TELEPHONE AND THE REPRESENTATIVE DOES NOT RECALL ANY INFORMATION ON THIS CASE. THE DEVICE WAS RETURNED, BUT WAS NOT RETURNED TO COMPLAINTS, AND WAS UNABLE TO BE LOCATED FOR ENGINEERING ANALYSIS. NO ADDITIONAL INFORMATION IS AVAILABLE FROM THE CONTACTS FOR THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE WAS RETURNED TO COMPANY, NOT AVAILABLE FOR ANALYSIS. ENGINEERING EVALUATION COMPLETED. DESIGN-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY COMPLAINT REPORTS INVOLVING BROKEN STEM INSERTERS SINCE 2008. BECAUSE THE STEM INSERTER IS USED IN TOTAL HIP SURGERIES, SALES DATA FOR FEMORAL STEMS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4). THIS IS CONSIDERED ¿VERY LOW¿ ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: MANUFACTURING DATA COULD NOT BE REVIEWED BECAUSE THE SERIAL/LOT INFORMATION WAS NOT PROVIDED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) COULD NOT BE CONDUCTED BECAUSE ADEQUATE INFORMATION WAS NOT PROVIDED TO CONDUCT A REVIEW. CORRECTIVE ACTIONS ARE NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS ¿VERY LOW.¿ THE BROKEN STEM INSERTER REPORTED WAS LIKELY THE RESULT OF MISALIGNMENT OF THE STEM INSERTER WHEN IMPACTING THE FEMORAL STEM, WHICH ALLOWED FOR THE TIP TO BREAK. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR INFORMATION AND NO FURTHER INFORMATION WAS PROVIDED. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00592 | 2022-05-20 | 2022-05-02 | I, n, j, u, r, y | Manufacturer report | TRULIANT TIB IMP CRC INSERT SZ 2.5,10MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED, | EXACTECH, INC. | TRULIANT TIB IMP CRC INSERT SZ 2.5,10MM | 02-022-51-2510 |  | AS REPORTED, APPROXIMATELY THREE YEARS POST TKA, PATIENT PRESENTED TO CLINIC WITH PAIN. A MAGNETIC RESONANCE IMAGING (MRI) SHOWED SYNOVITIS AND A REVISION WAS PERFORMED. SURGEONS COMMENTED THAT THE FEMUR AND TIBIA WERE BOTH WELL ALIGNED AND NOT MECHANICALLY LOOSE. THE PATELLA HAD SOME OSSIFICATIONS THAT THEY REMOVED. SURGEONS ALSO NOTED SYNOVITIS TISSUE AND REMOVED IT. THEY SAW SOME POSTERIOR WEAR ON THE INSERT WITH SOME DAMAGE TO THE POSTERIOR LEADING EDGE WITH SUSPECTED DAMAGE UPON ORIGINAL INSERTION. A NEW INSERT WAS PLACED OF SAME THICKNESS (10 CRC). PATIENT WAS LAST KNOW TO BE IN STABLE CONDITION. THE INSERT IS NOT AVAILABLE FOR RETURN. NO FURTHER INFORMATION AVAILABLE. X-RAYS AND PHOTOS RECEIVED. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00597 | 2022-05-21 | 2022-04-28 | I, n, j, u, r, y | Manufacturer report | NV GXL LNR, +5LAT, 36MM G2-52/54MM CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LNR, +5LAT, 36MM G2-52/54MM CUPS | 136-36-52 |  | AS REPORTED, APPROXIMATELY SIX YEARS POST INITIAL LEFT THA, THE 40 Y/O MALE PATIENT VISITED SURGEON AND LEFT HIP HAS CLEAR PREMATURE WEAR. THE LEFT HIP HAS ASYMMETRIC POLY WEAR, BUT LUCKILY NO LYSIS. PATIENT WAS TOLD TO COME BACK IN 6 MONTHS. NO OTHER INFORMATION PROVIDED | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION RECEIVED THE FOLLOWING SECTIONS D6B, G3, G6, H2 AND H3 HAVE BEEN UPDATED ACCORDINGLY. SECTION D6B: EXPLANTED, GIVE DATE (H3) THE REVISION REPORTED MAY HAVE BEEN THE RESULT OF INCREASED BIOMECHANICAL STRESSORS ON THE JOINTS SECONDARY TO PATIENT-RELATED CONDITIONS AND A COMBINATION OF THE RISK FACTORS INCLUDING, USE ERROR, IMPLANT POSITIONING, IMPLANT SIZE SELECTION, AND PATIENT FACTORS MAY HAVE ALSO BEEN A CONTRIBUTING FACTOR TO THE EARLY PROSTHESIS WEAR. MOREOVER, INCLUDING BUT NOT LIMITED TO SIGNIFICANT EDGE LOADING OF THE FEMORAL HEAD ON THE ACETABULAR LINER, BEING IMPLANTED WITH A LATERALIZED LINER, AND COMBINATION OF LARGEST AVAILABLE FEMORAL HEAD AND/OR THINNEST AVAILABLE ACETABULAR LINER. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICES WERE NOT AVAILABLE FOR EVALUATION. THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE REPORTED EVENT OF "PROSTHESIS WEAR¿ IS ASSOCIATED WITH MATERIAL DAMAGE TO A SURFACE, USUALLY INVOLVING PROGRESSIVE LOSS OR DISPLACEMENT OF MATERIAL, DUE TO RELATIVE MOTION BETWEEN THAT SURFACE AND A CONTACTING SUBSTANCE OR SUBSTANCES. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10145 | 2022-05-24 |  | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | SMALL STEM PROTECTOR | EXACTECH, INC. |  | 301-07-60 |  | IT WAS REPORTED FROM OUS THAT DURING AN INITIAL ORTHOPEDIC SURGERY THE SMALL STEM PROTECTOR BROKE. NO DELAY TO SURGERY AND NO ADVERSE EVENT TO THE PATIENT. THE AGENT WAS PRESENT AT TIME OF SURGERY. THE PATIENT'S HEALTH WAS STABLE LEAVING THE OR. NO PATIENT DEMOGRAPHIC INFORMATION WAS PROVIDED. NEW INFORMATION PROVIDED ON (B)(6) 2021: "THE INSTRUMENT COULD HAVE BEEN BROKEN DURING THE SURGERY. HOWEVER, THE BREAKAGE HAS BEEN NOTICED AT THE RECOVERY OF THE INSTRUMENTATION-SET BY OUR LOGISTICS DEPT., AFTER THE CLEANING REALIZED BY THE STERILIZATION TEAM OF THE HOSPITAL. NO INFORMATION HAS BEEN RECEIVED FROM THE HOSPITAL REGARDING THIS ISSUE AND [THE AGENCY] DIDN'T RECEIVE ANY REQUEST FOR COMPLEMENTARY INFORMATION REGARDING THIS ISSUE AS WELL." MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: DESIGN-RELATED ISSUES: THE DESIGN OF THE SMALL STEM PROTECTOR HAS BEEN IN THE FIELD SINCE 2004. EXACTECH IS AWARE OF 6 OTHER SIMILAR COMPLAINT REPORTS INVOLVING STEM PROTECTOR DEVICES SINCE 2008. BECAUSE THE SMALL STEM PROTECTOR IS USED IN ANATOMIC AND REVERSE TOTAL SHOULDER SURGERIES, SALES DATA FOR HUMERAL STEMS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THESE MANUFACTURING LOTS. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING-RELATED. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RISK MANAGEMENT REPORT. THE BROKEN INSTRUMENTS REPORTED WERE LIKELY THE RESULT OF IMPACTION OF THE CAP COMPONENT DURING SURGERY OR REMOVAL OF THE CAP COMPONENT DURING REPROCESSING AND NOT REPLACING IT. HOWEVER, FURTHER DETAILS REGARDING HOW THE DEVICES BECAME BROKEN WERE NOT PROVIDED. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. THIS IS MDR 4 OF 4 FROM THE SAME CUSTOMER REPORT NOT THE SAME PATIENT OR EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10137 | 2022-05-24 |  | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | DRILL GUIDE | EXACTECH, INC. |  | T6163 |  | IT WAS REPORTED FROM THE US THAT THE PINS COME OUT OF THE DRILL GUIDES. THE EVENT DID NOT OCCUR DURING SURGERY. THE ISSUE WAS NOTICED BY THE AGENT AS HE WAS PUTTING THE INSTRUMENTS BACK TOGETHER. PHOTO WAS PROVIDED. THE AGENT STATED: "FROM WHAT I CAN REMEMBER, THE PIN THAT ALLOWS THE BLUE ADJUSTABLE DRILL GUIDE TO ADJUST ANGLES CAME OUT. THIS CAUSED THE DRILL GUIDE TO BE UNABLE TO LOCK IN ON AN ANGLE. I NOTICED IT AS I WAS PUTTING INSTRUMENTS BACK IN TRAYS. TO MY KNOWLEDGE, IT MUST HAVE HAPPENED AS INSTRUMENTS WERE BEING CLEANED. THE GUIDE WAS NOT DROPPED, AND THE SURGEON DID NOT NOTIFY ME OF THE PROBLEM DURING SURGERY, AND HE WOULD HAVE IF IT WAS NOT WORKING PROPERLY." NO ADDITIONAL INFORMATION WAS PROVIDED. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE(S) NOT AVAILABLE FOR EVALUATION. EVALUATION COMPLETED BY CC ON 12/11/2018. THE ENGINEERING MEMO STATES THAT THE DEVICE WAS SENT TO SERVICING. MOST LIKELY UNDERLYING CAUSE / ADDITIONAL COMMENTS: THE DEVICE WAS NOT AVAILABLE FOR EVALUATION AT THE TIME OF THIS INVESTIGATION. IF THE DEVICE IS RECEIVED IN THE FUTURE, THE FILE WILL BE REOPENED. A PHOTO OF THE DISASSEMBLED ADJUSTABLE DRILL GUIDE WAS PROVIDED. ACCORDING TO THE EXPERIENCE REPORT THE PIN DISENGAGED FROM THE DRILL GUIDE. THE EVENT DID NOT HAPPEN DURING SURGERY. THE SALES REPRESENTATIVE NOTICED THE ISSUE WHILE GATHERING INSTRUMENTS. DESIGN-RELATED ISSUES: THE ADJUSTABLE DRILL GUIDE HAS BEEN IN THE FIELD SINCE 2003. EXACTECH IS AWARE OF NO OTHER COMPLAINT REPORTS INVOLVING ADJUSTABLE DRILL GUIDES REPORTED AS DISASSEMBLED/MISASSEMBLED SINCE 2014. THE ADJUSTABLE DRILL GUIDE IS USED IN TOTAL HIP ARTHROPLASTY SURGERIES IN THE NOVATION SYSTEM, SO SALES DATA FOR NOVATION HIP ACETABULAR IMPLANTS WERE USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: NOT APPLICABLE. NO INFORMATION WAS PROVIDED REGARDING THE SERIAL OR LOT INFORMATION SO MANUFACTURING INFORMATION COULD NOT BE REVIEWED. MOST LIKELY CAUSE: THE DISASSEMBLED ADJUSTABLE DRILL GUIDE REPORTED WAS LIKELY THE RESULT OF A STUCK LEVER WITHIN THE MECHANISM THAT ALLOWS THE ADJUSTABLE DRILL GUIDE TO CHANGE ANGLES, CAUSING THE PINS TO POP OUT. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE COMPONENT WAS NOT RETURNED FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES. NO PATIENT INVOLVEMENT. THE PINS CAME OUT OF THE DEVICE. AN INVESTIGATION WAS CONDUCTED: THE DISASSEMBLED ADJUSTABLE DRILL GUIDE REPORTED WAS LIKELY THE RESULT OF A STUCK LEVER WITHIN THE MECHANISM THAT ALLOWS THE ADJUSTABLE DRILL GUIDE TO CHANGE ANGLES, CAUSING THE PINS TO POP OUT. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE COMPONENT WAS NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10129 | 2022-05-24 | 2017-12-28 | I, n, j, u, r, y | Manufacturer report | EXACTECH | STEM PROTECTOR, SMALL | EXACTECH, INC. |  | 301-07-60 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGICAL PROCEDURE THE INSTRUMENT BROKE. THE AGENT WAS PRESENT AT THE TIME OF SURGERY. THE PATIENT'S HEALTH WAS STABLE LEAVING THE OR. ALL PIECES WERE RETRIEVED FROM THE PATIENT. NEW INFORMATION RECEIVED ON 15 MAR 2021: "THE INSTRUMENT COULD HAVE BEEN BROKEN DURING THE SURGERY. HOWEVER, THE BREAKAGE HAS BEEN NOTICED AT THE RECOVERY OF THE INSTRUMENTATION-SET BY OUR LOGISTICS DEPT., AFTER THE CLEANING REALIZED BY THE STERILIZATION TEAM OF THE HOSPITAL. NO INFORMATION HAS BEEN RECEIVED FROM THE HOSPITAL REGARDING THIS ISSUE AND [THE AGENCY] DIDN'T RECEIVE ANY REQUEST FOR COMPLEMENTARY INFORMATION REGARDING THIS ISSUE AS WELL." MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED FOR ANALYSIS. ENG ANALYSIS COMPLETED ON 08/14/2018 BY AD DESIGN: THIS DESIGN HAS BEEN IN THE FIELD SINCE 2005. THE OCCURRENCE RATE IS (B)(4), "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THIS DOES NOT APPEAR TO BE DESIGN RELATED. MANUFACTURING: THERE HAVE BEEN NO OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF (B)(4) UNITS, WHICH HAS BEEN IN THE FIELD SINCE 2014. THIS DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF BEING INADVERTENTLY IMPACTED, LEADING THE CAP COMPONENT TO BREAK. HOWEVER, THIS CANNOT BE CONFIRMED AS THE CAP COMPONENT WAS NOT AVAILABLE FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. IT WAS REPORTED THAT ALL PIECES WERE RETRIEVED. AN INVESTIGATION WAS CONDUCTED; THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF BEING INADVERTENTLY IMPACTED, LEADING THE CAP COMPONENT TO BREAK. HOWEVER, THIS CANNOT BE CONFIRMED AS THE CAP COMPONENT WAS NOT AVAILABLE FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10133 | 2022-05-24 | 2018-08-10 | I, n, j, u, r, y | Manufacturer report | EXACTECH | ANGLE DRIVER | EXACTECH, INC. |  | 315-15-10 |  | IT WAS REPORTED FROM THE US THAT DURING A PRIMARY SHOULDER ARTHROPLASTY THE ANGLE DRIVER BROKE. THE ISSUE WAS RESOLVED WITH AN ALTERNATE INSTRUMENT. IT WAS CONFIRMED THERE WERE NO PIECES LEFT IN THE PATIENT AND NO DELAY TO SURGERY. THE AGENT WAS PRESENT AT THE TIME OF SURGERY. THE PATIENT LEFT THE ROOM IN STABLE CONDITION. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS ((B)(4)). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED FOR ANALYSIS. ENG EVAL COMPLETED ON 10/18/2018 BY (B)(6). THE FRACTURED SURFACES APPEAR TO BE CONSISTENT WITH CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE INNER SHAFT. THE DEFORMATION AND SCRATCHES APPEAR TO BE CONSISTENT WITH NORMAL USE. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE AND CONSISTENT WITH YEARS OF USE. DESIGN-RELATED ISSUES: THE DESIGN OF THE ANGLE DRIVER HAS BEEN IN THE FIELD SINCE 2007. EXACTECH IS AWARE OF ONE OTHER SIMILAR COMPLAINT REPORT INVOLVING ANGLE DRIVERS SINCE 2007. BECAUSE THE ANGLE DRIVER IS USED IN TOTAL SHOULDER SURGERIES, SALES DATA FOR HUMERAL STEMS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED " VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 14 UNITS, WHICH HAS BEEN IN THE FIELD SINCE 2014. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS " VERY LOW", AND THE RISK IS CAPTURED IN THE RISK MANAGEMENT REPORT. THE BROKEN INSTRUMENT REPORTED IN EXPERIENCE (B)(4) IS LIKELY THE RESULT OF CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE INNER SHAFT. IFU (B)(4): INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE WAS NO REPORTED PATIENT ADVERSE EVENT, IT WAS CONFIRMED THAT NO PIECES WERE LEFT IN THE PATIENT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN INSTRUMENT REPORTED IS LIKELY THE RESULT OF CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE INNER SHAFT. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00601 | 2022-05-24 | 2022-04-01 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 5, 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 5, 11MM | 02-012-44-5011 |  | AS REPORTED, APPROXIMATELY TWO YEARS POST INITIAL BILATERAL TKA, THIS MALE PATIENT HAD A REVISION DUE TO WEAR OF PS AND PSC INSERTS. PATIENT HAS BILATERAL LOGIC KNEES. THERE WAS SOME LYSIS SEEN MEDIALLY ON THE RIGHT TIBIA DEFECT FILLED WITH A BIOLOGIC. COMPONENTS OF BOTH KNEES WERE WELL FIXED EXCEPT THE 35MM PATELLA ON THE RIGHT KNEE. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PHOTOS RECEIVED. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO HSS KEEPS ALL RETRIEVALS. | CONCOMITANT MEDICAL PRODUCTS: THREE PEG PATELLA 35MM (CAT# 200-02-35 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10118 | 2022-05-24 | 2017-09-29 | I, n, j, u, r, y | Manufacturer report | EXACTECH | LPI IMPACTOR HANDLE | EXACTECH, INC. |  | 213-65-00 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGICAL PROCEDURE THE LPI HANDLE BROKE. THE PLASTIC GRIP BROKE INTO 3 PIECES WHEN THE SURGEON IMPACTED. NO BROKEN PIECES WERE LEFT IN THE PATIENT. THE SURGEON USED ANOTHER INSTRUMENT TO COMPLETE THE PROCEDURE. THERE WAS NO SURGICAL DELAY OR PATIENT ADVERSE EVENT. SURGERY WAS SUCCESSFUL. THE REPRESENTATIVE WAS PRESENT AT THE TIME OF SURGERY. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED FOR ANALYSIS. ENG EVAL COMPLETED ON 11/15/2018 BY CC. AN INVESTIGATION OF COMPLAINTS RELATED TO LPI HANDLE - BROKEN NON-FUNCTIONAL, SPECIFICALLY RELATED TO CRACKS/BREAKAGE FOR THE PLASTIC HANDLE WAS CONDUCTED UNDER TM 2018-XX. CRC-2018-03-02-02 WAS INITIATED ON 03/02/2018 TO EVALUATE THIS ISSUE. THE LPI HANDLE CONSISTS OF A METAL STRIKE PLATE, METAL END PIECE WHICH ENGAGES WITH MATING INSTRUMENTS AND IMPLANTS, AND A PLASTIC CENTRAL HANDLE. THE METAL END PIECES ARE ASSEMBLED TO THE PLASTIC HANDLE VIA THERMAL ASSEMBLY WHICH INVOLVES HEATING THE METAL AND PRESSING THE HOT METAL PIECES INTO THE HANDLE. THE METAL PIECES HAVE CYLINDRICAL BOSSES WHICH FEATURE A PRESS-FIT WITH A FEMALE HOLE ON THE PLASTIC HANDLE. THE LPI HANDLE ASSEMBLES WITH TWO END PIECES FOR USE THREE TIMES DURING THE SURGICAL PROCEDURE: IMPACTION OF THE FEMORAL TRIAL, IMPACTION OF THE TIBIAL TRAY IMPLANT, AND IMPACTION OF THE FEMORAL IMPLANT. AVAILABLE DEVICES FROM PREVIOUSLY REPORTED COMPLAINTS WERE EVALUATED. COMPLAINT DEVICES APPEARED TO HAVE SIMILAR FAILURE MODES, AND EITHER WERE CRACKED OR FRACTURED. CRACKS WERE CONSISTENT WITH STARTING INSIDE THE PLASTIC HANDLE AT THE BASE OF ONE OF THE TWO METAL ATTACHMENT FEATURES AND SPREADING OUTWARD. CRACKS WERE EITHER LINEAR, ALONG THE HANDLE AXIS OR HORIZONTAL, NEARLY IN LINE WITH THE ORIGIN. THE REVIEW OF THE LPI IMPACTOR HANDLES (213-65-00) CITED IN COMPLAINTS FOR BROKEN OR CRACKED HANDLES SHOWED A CONSISTENT FAILURE MODE. A REVIEW OF THE DEVICE ORDER HISTORY DID NOT SHOW A TREND IN THOSE LOTS WITH COMPLAINT DEVICES, NOR DID A REVIEW OF DESIGN CHANGES SHOW CHANGES MADE TO THE DESIGN WHICH COULD CONTRIBUTE TO THE COMPLAINT. PER 750-2009-023-RMR-INSTRUMENT, REV B - DEVICE FRACTURE/BREAKAGE (LINE ITEM 18) CAN LEAD TO A SURGICAL DELAY. THE PREDICTED SEVERITY OF THIS HARM IS MODERATE (3), DEFINED AS "ANY INSTRUMENT-RELATED HAZARD THAT COULD RESULT IN MARKED INTRA-OPERATIVE INCONVENIENCE TO THE SURGEON AND STAFF", AND THE PREDICTED OCCURRENCE IS LOW (2), DEFINED AS 1%. BASED ON THE AVAILABLE DATA, THE DEVICE IS PERFORMING WITHIN ITS PREDICTED RISK. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE WAS NO PATIENT ADVERSE EVENT. THERE ARE NO PIECES LEFT IN THE PATIENT. AN INVESTIGATION WAS CONDUCTED UNDER CRC-2018-03-02-02; THE MOST LIKELY CAUSE OF THE REPORTED EVENTS IS RELATED TO HIGH IMPACTION FORCES REQUIRED DURING IMPLANT SURGERY. INFORMATION WAS NOT PROVIDED FOR: AGE, SEX, WEIGHT, ETHNICITY, RACE. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10135 | 2022-05-24 |  | I, n, j, u, r, y | Manufacturer report | EXACTECH | SPECIAL 14MM DRILL GUIDE BUSHING | EXACTECH, INC. |  | SK-016000 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE DRILL GUIDE BUSHING. THE DRILL STOP DOES NOT HOLD POSITION WHEN REAMING DOWN THROUGH THE DRILL/PUNCH GUIDE. THE PART WAS USED DURING SURGERY, IT JUST PUNCHED DOWN A LITTLE FURTHER THAN THE SETTING. NO DELAYS AND NO ADVERSE EFFECTS TO PATIENTS. AGENT WAS PRESENT AT THE SURGERY. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT FOLLOWING MULTIPLE ATTEMPTS. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS ((B)(4)). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS UNABLE TO BE LOCATED AT THE TIME OF THIS EVALUATION. ANALYSIS COMPLETED ON 09/24/2018 BY (B)(6). THIS COMPLAINT HAS BEEN ADDED TO CAPA (B)(4). DESIGN-RELATED ISSUES: THE DESIGN OF THIS SPECIAL TRULIANT DRILL STOP HAS BEEN IN THE FIELD SINCE 2017. EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINT REPORTS REGARDING THE STANDARD 14MM TRULIANT DRILL STOPS (CATALOG #: 02-029-22-1200 OR 02-029-22-0310). THE FAILURE OF THIS INSTRUMENT APPEARS TO BE DIRECTLY RELATED TO THE FEATURE THAT MADE THIS A SPECIAL INSTRUMENT AND DIFFERENT FROM THE PARENT/PREDICATE INSTRUMENT. CAPA (B)(4) WAS INITIATED TO ENSURE THAT SPECIAL INSTRUMENTS ARE DESIGNED AND MANUFACTURED AFTER SUFFICIENT DESIGN CONTROL ACTIVITIES ARE COMPLETED AND DOCUMENTED. THE SPECIAL INSTRUMENT SOP IS BEING UPDATED AS PART OF THE CAPA TO IMPROVE THE CONNECTION BETWEEN INPUTS, PREDICATE DEVICES, RISK ASSESSMENT, AND POST-MARKET FEEDBACK. MFG-RELATED ISSUES: THIS BREAKAGE APPEARS TO BE RELATED TO THE DESIGN OF THE SPECIAL INSTRUMENT. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THE STANDARD TRULIANT PILOT DRILL GUIDES, 750-2014-006-RMR-INSTRUMENTS REV C, WAS REVIEWED. THE RISK IS CAPTURED IN LINE 17. CORRECTIVE ACTIONS WILL BE DETERMINED BASED ON THE OUTCOME OF FUTURE INVESTIGATIONS AS A RESULT OF THE CAPA. THE BROKEN INSTRUMENT REPORTED IN (B)(4) WAS LIKELY THE RESULT OF THE DESIGN OF THE INSTRUMENT. CAPA (B)(4) HAS BEEN INITIATED TO ENSURE THAT SPECIAL INSTRUMENTS ARE DESIGNED AND MANUFACTURED AFTER SUFFICIENT DESIGN CONTROL ACTIVITIES ARE COMPLETED AND DOCUMENTED. IFU (B)(4): INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES. THE PART WAS USED DURING SURGERY, "JUST PUNCHED DOWN A LITTLE FURTHER THAN THE SETTING". THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF THE DESIGN OF THE INSTRUMENT. CAPA (B)(4) HAS BEEN INITIATED TO ENSURE THAT SPECIAL INSTRUMENTS ARE DESIGNED AND MANUFACTURED AFTER SUFFICIENT DESIGN CONTROL ACTIVITIES ARE COMPLETED AND DOCUMENTED. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10140 | 2022-05-24 | 2019-09-11 | I, n, j, u, r, y | Manufacturer report | EXACTECH | HUM LINER IMPACT TIP 42M | EXACTECH, INC. |  | 321-07-42 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE HUMERAL LINER IMPACT TIP BROKE INTO SMALL PIECES DURING IMPACTION. THE EVENT OCCURRED WHEN IMPACTING THE DEFINITIVE CONSTRUCT INTO THE HUMERUS. THE HUMERAL LINER IMPACTOR TIP (WHICH IS A PLASTIC PART) BROKE IN SEVERAL PIECES WHEN THE SURGEON IMPACTED THE PROTHESIS. ALL PLASTIC PART HAVE BEEN REMOVED FROM THE PATIENT & THE SURGEON REPLACED THE BROKEN TIP BY THE 38 MM IMPACTOR TIP TO FINISH THE IMPACTION. THE ORIGINAL REPORT STATED: THERE WAS A MISSING INSTRUMENT INSIDE ONE OF THE KITS. NO INDICATION PROVIDED BY THE CUSTOMER WHEN THE KITS OF INSTRUMENTS WERE RETURNED. THE PIECE DID NOT BREAK/FALL INTO THE PATIENT. THE AGENT WAS PRESENT DURING SURGERY AND STATED THAT THE ITEM BROKE INTO SMALL PIECES DURING IMPACTION & THAT THE MEDICAL STAFF TOLD HER TO THROW IT AWAY & THEY WILL INFORM US AFTERWARDS. (WHICH THEY DID NOT). MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: AGE, SEX, WEIGHT, ETHNICITY, RACE. ENGINEERING ANALYSIS WITH NO DEVICE RETURN: FINDINGS: BASED ON PREVIOUS SIMILAR INVESTIGATIONS, THE FRACTURED HUMERAL LINER IMPACTOR TIP REPORTED IS CONSISTENT WITH A STRESS RISER INTRODUCED DURING IMPACTION, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE. NO LOT NUMBER WAS PROVIDED IN THE EXPERIENCE REPORT AND BECAUSE THE DEVICE WAS NOT RETURNED, IT CANNOT BE DETERMINED IF THIS DEVICE WAS MANUFACTURED PRIOR TO DESIGN CHANGES TO INCREASE THICKNESS AND MODIFY THE CONNECTION FEATURE DIMENSIONS. RMR REVIEW: RMR-00004 REV A AND RACR-00007 REV A WERE REVIEWED. THE RISK IS CAPTURED. MOST LIKELY UNDERLYING CAUSE: BASED ON PREVIOUS SIMILAR INVESTIGATIONS, THE FRACTURED HUMERAL LINER IMPACTOR TIP REPORTED WAS LIKELY THE RESULT OF A STRESS RISER INTRODUCED DURING IMPACTION WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE. AS OF THIS TIME ((B)(6) 2019), INVESTIGATION FOR AN ADDITIONAL MOST LIKELY UNDERLYING CAUSE/ROOT CAUSE RELATED TO MARKET SPECIFIC OCCURRENCE IS STILL ONGOING. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10123 | 2022-05-24 | 2016-09-14 | I, n, j, u, r, y | Manufacturer report | EXACTECH | LOGIC PS TIBIA, MODULAR INSERT TRIAL, SIZE 1, 9MM | EXACTECH, INC. |  | 02-013-35-1009 |  | IT WAS REPORTED FROM ITALY BY THE HOSPITAL THAT THE LOGIC PS TIBIA, MODULAR INSERT WAS BROKEN. THE AGENT WAS PRESENT AT SURGERY. THERE WAS NO DELAY TO THE PROCEDURE AND NO HARM TO THE PATIENT. IT WAS REPORTED BY THE REPRESENTATIVE THAT NO PIECES WERE LEFT IN THE PATIENT. THE AGENT WAS PRESENT AT SURGERY WITH NO DELAY. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS THE ISSUE WAS RESOLVED WITH THE USE OF SECONDARY INSTRUMENTATION. MULTIPLE ATTEMPTS HAVE BEEN MADE FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS NOT RECEIVED FOR ANALYSIS. THE ANALYSIS WAS COMPLETED ON 15 MAY 2018 BY (B)(4). DESIGN-RELATED ISSUES: THE DESIGN OF THE LOGIC PS TIBIAL INSERT TRIAL HAS BEEN IN THE FIELD SINCE 2008. EXACTECH HAS RECEIVED (B)(4) COMPLAINT REPORTS INVOLVING (B)(4) TRIALS DUE TO BEING BROKEN SINCE 2008. BECAUSE TIBIAL INSERT TRIALS ARE USED IN TOTAL KNEE SURGERIES, SALES DATA FOR TIBIAL INSERTS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4). THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE; THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH HAS NOT RECEIVED ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT, WHICH HAS BEEN IN THE FIELD SINCE 2012. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. NO CORRECTIVE ACTION IS REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE BROKEN TIBIAL INSERT TRIAL REPORTED IN EXPERIENCE (B)(6) WAS LIKELY THE RESULT OF IMPACTING OR DROPPING THE DEVICE ONTO A HARD SURFACE, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE TRIAL. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION. VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. IT WAS REPORTED BY THE REPRESENTATIVE THAT NO PIECES WERE LEFT IN THE PATIENT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN TIBIAL INSERT TRIAL REPORTED WAS LIKELY THE RESULT OF IMPACTING OR DROPPING THE DEVICE ONTO A HARD SURFACE, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE TRIAL. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10121 | 2022-05-24 | 2017-11-06 | I, n, j, u, r, y | Manufacturer report | EXACTECH | LPI IMPACTOR HANDLE | EXACTECH, INC. |  | 213-65-00 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGICAL PROCEDURE THE LPI HANDLE BROKE. THE REPRESENTATIVE WAS PRESENT AT THE TIME OF SURGERY. THE ISSUE WAS RESOLVED BY USING ANOTHER IMPACTOR. THERE WAS NO DELAY IN SURGERY AND NO ADVERSE EVENT TO THE PATIENT. IT WAS REPORTED THAT NO PIECES WERE LEFT IN THE PATIENT. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI (B)(6), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. ENG EVAL 11/15/2018 BY CC AN INVESTIGATION OF COMPLAINTS RELATED TO LPI HANDLE - BROKEN NON-FUNCTIONAL, SPECIFICALLY RELATED TO CRACKS/BREAKAGE FOR THE PLASTIC HANDLE WAS CONDUCTED UNDER TM 2018-XX. THE LPI HANDLE CONSISTS OF A METAL STRIKE PLATE, METAL END PIECE WHICH ENGAGES WITH MATING INSTRUMENTS AND IMPLANTS, AND A PLASTIC CENTRAL HANDLE. THE METAL END PIECES ARE ASSEMBLED TO THE PLASTIC HANDLE VIA THERMAL ASSEMBLY WHICH INVOLVES HEATING THE METAL AND PRESSING THE HOT METAL PIECES INTO THE HANDLE. THE METAL PIECES HAVE CYLINDRICAL BOSSES WHICH FEATURE A PRESS -FIT WITH A FEMALE HOLE ON THE PLASTIC HANDLE. THE LPI HANDLE ASSEMBLES WITH TWO END PIECES FOR USE THREE TIMES DURING THE SURGICAL PROCEDURE: IMPACTION OF THE FEMORAL TRIAL. IMPACTION OF THE TIBIAL TRAY IMPLANT. IMPACTION OF THE FEMORAL IMPLANT. AVAILABLE DEVICES FROM PREVIOUSLY REPORTED COMPLAINTS WERE EVALUATED. COMPLAINT DEVICES APPEARED TO HAVE SIMILAR FAILURE MODES, AND EITHER WERE CRACKED OR FRACTURED. CRACKS WERE CONSISTENT WITH STARTING INSIDE THE PLASTIC HANDLE AT THE BASE OF ONE OF THE TWO METAL ATTACHMENT FEATURES AND SPREADING OUTWARD. CRACKS WERE EITHER LINEAR, ALONG THE HANDLE AXIS OR HORIZONTAL, NEARLY IN LINE WITH THE ORIGIN. THE REVIEW OF THE LPI IMPACTOR HANDLES (213-65-00) CITED IN COMPLAINTS FOR BROKEN OR CRACKED HANDLES SHOWED A CONSISTENT FAILURE MODE. A REVIEW OF THE DEVICE ORDER HISTORY DID NOT SHOW A TREND IN THOSE LOTS WITH COMPLAINT DEVICES, NOR DID A REVIEW OF DESIGN CHANGES SHOW CHANGES MADE TO THE DESIGN WHICH COULD CONTRIBUTE TO THE COMPLAINT. PER 750-2009-023-RMR-INSTRUMENT, REV B - DEVICE FRACTURE/BREAKAGE (LINE ITEM 18) CAN LEAD TO A SURGICAL DELAY. THE PREDICTED SEVERITY OF THIS HARM IS MODERATE (3), DEFINED AS "ANY INSTRUMENT - RELATED HAZARD THAT COULD RESULT IN MARKED INTRA-OPERATIVE INCONVENIENCE TO THE SURGEON AND STAFF", AND THE PREDICTED OCCURRENCE IS LOW (2), DEFINED AS LESS THAN OR EQUAL TO 1%. CRC-2018-03-02-02 WAS INITIATED ON 03/02/2018 TO EVALUATE THIS ISSUE. BASED ON THE AVAILABLE DATA, THE DEVICE IS PERFORMING WITHIN ITS PREDICTED RISK. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE WAS NO DELAY IN SURGERY AND NO ADVERSE EVENT TO THE PATIENT. IT WAS REPORTED THAT NO PIECES WERE LEFT IN THE PATIENT. AN INVESTIGATION WAS CONDUCTED; DEVICES APPEARED TO HAVE SIMILAR FAILURE MODES, AND EITHER WERE CRACKED OR FRACTURED. CRACKS WERE CONSISTENT WITH STARTING INSIDE THE PLASTIC HANDLE AT THE BASE OF ONE OF THE TWO METAL ATTACHMENT FEATURES AND SPREADING OUTWARD. CRACKS WERE EITHER LINEAR, ALONG THE HANDLE AXIS OR HORIZONTAL, NEARLY IN LINE WITH THE ORIGIN. BASED ON THE AVAILABLE DATA, THE DEVICE IS PERFORMING WITHIN ITS PREDICTED RISK. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. INFORMATION WAS NOT PROVIDED FOR: AGE, SEX, WEIGHT, ETHNICITY, RACE. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10113 | 2022-05-24 | 2016-12-20 | I, n, j, u, r, y | Manufacturer report | EXACTECH | LPI IMPACTOR HANDLE | EXACTECH, INC. |  | 213-65-00 |  | IT WAS REPORTED FROM OUS THAT THE INSTRUMENT BROKE DURING AN ORTHOPEDIC SURGICAL PROCEDURE. ALL PIECES WERE RETRIEVED FROM THE PATIENT. THE SURGERY WAS SUCCESSFUL WITH NO DELAY. ADDITIONAL INFORMATION RECEIVED ON 9 MAR 2021- "THE INSTRUMENT COULD HAVE BEEN BROKEN DURING THE SURGERY. HOWEVER, THE BREAKAGE HAS BEEN NOTICED AT THE RECOVERY OF THE INSTRUMENTATION-SET BY OUR LOGISTICS DEPT., AFTER THE CLEANING REALIZED BY THE STERILIZATION TEAM OF THE HOSPITAL. NO INFORMATION HAS BEEN RECEIVED FROM THE HOSPITAL REGARDING THIS ISSUE AND [THE AGENCY] DIDN'T RECEIVE ANY REQUEST FOR COMPLEMENTARY INFORMATION REGARDING THIS ISSUE AS WELL." MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DESIGN OF THE LPI IMPACTOR HANDLE HAS BEEN IN THE FIELD SINCE 2009. 6 COMPLAINT REPORTS INVOLVING THE HANDLE SINCE ITS INTRODUCTION. THE APPROXIMATE COMPLAINT OCCURRENCE RATE IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THERE HAVE NOT BEEN ANY OTHER COMPLAINTS FROM MANUFACTURING LOT 49039003. THE INSTRUMENT EVENT REPORTED WAS LIKELY THE RESULT OF WEAR AND DEGRADATION DUE TO MANY USES AND SUBSEQUENT STERILIZATION CYCLES. THE ASSEMBLY ISSUE INVOLVING THE STEM TRIAL WAS LIKELY THE RESULT OF THREAD DEFORMATION ON THE TRIAL OR THE MATING INSTRUMENT. HOWEVER, THE MOST LIKELY UNDERLYING CAUSES OF THE REPORTED ISSUES CANNOT BE CONFIRMED BECAUSE THE INSTRUMENT WAS NOT AVAILABLE FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10134 | 2022-05-24 |  | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | SMALL STEM PROTECTOR | EXACTECH, INC. |  | 301-07-60 |  | IT WAS REPORTED FROM OUS THAT DURING AN INITIAL ORTHOPEDIC SURGERY THE SMALL STEM PROTECTOR BROKE. NO DELAY TO SURGERY AND NO ADVERSE EVENT TO THE PATIENT. THE AGENT WAS PRESENT AT TIME OF SURGERY. THE PATIENT'S HEALTH WAS STABLE LEAVING THE OR. NO PATIENT DEMOGRAPHIC INFORMATION WAS PROVIDED. NEW INFORMATION PROVIDED ON (B)(6)2021: "THE INSTRUMENT COULD HAVE BEEN BROKEN DURING THE SURGERY. HOWEVER, THE BREAKAGE HAS BEEN NOTICED AT THE RECOVERY OF THE INSTRUMENTATION-SET BY OUR LOGISTICS DEPT., AFTER THE CLEANING REALIZED BY THE STERILIZATION TEAM OF THE HOSPITAL. NO INFORMATION HAS BEEN RECEIVED FROM THE HOSPITAL REGARDING THIS ISSUE AND [THE AGENCY] DIDN'T RECEIVE ANY REQUEST FOR COMPLEMENTARY INFORMATION REGARDING THIS ISSUE AS WELL." MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INFORMATION WAS NOT PROVIDED FOR: AGE OR DATE OF BIRTH, SEX, WEIGHT, ETHNICITY, RACE; DATE OF EVENT. DESIGN-RELATED ISSUES: THE DESIGN OF THE SMALL STEM PROTECTOR HAS BEEN IN THE FIELD SINCE 2004. EXACTECH IS AWARE OF 6 OTHER SIMILAR COMPLAINT REPORTS INVOLVING STEM PROTECTOR DEVICES SINCE 2008. BECAUSE THE SMALL STEM PROTECTOR IS USED IN ANATOMIC AND REVERSE TOTAL SHOULDER SURGERIES, SALES DATA FOR HUMERAL STEMS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4). THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THESE MANUFACTURING LOTS. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING-RELATED. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RISK MANAGEMENT REPORT. THE BROKEN INSTRUMENTS REPORTED WERE LIKELY THE RESULT OF IMPACTION OF THE CAP COMPONENT DURING SURGERY OR REMOVAL OF THE CAP COMPONENT DURING REPROCESSING AND NOT REPLACING IT. HOWEVER, FURTHER DETAILS REGARDING HOW THE DEVICES BECAME BROKEN WERE NOT PROVIDED. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. THIS IS MDR 1OF 4 FROM THE SAME CUSTOMER REPORT NOT THE SAME PATIENT OR EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10141 | 2022-05-24 | 2019-09-07 | I, n, j, u, r, y | Manufacturer report | EXACTECH | HUMERAL LINER IMPACTOR TIP, 38MM | EXACTECH, INC. |  | 321-07-38 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE SURGEON EXPERIENCED AN ISSUE WITH THE INSTRUMENTATION. THE HUMERAL LINER IMPACTOR TIP 38MM WAS BROKEN DURING IMPACTION. "THERE WERE NO PIECES LEFT IN THE PATIENT" THERE WAS NO DELAY TO THE SURGERY. THE SURGERY WAS COMPLETED WITH NO CONSEQUENCES TO THE PATIENT. THE PATIENT IS FINE. THE REPRESENTATIVE WAS PRESENT AT THE TIME OF THE SURGERY. ATTEMPTS WERE MADE TO REQUEST ADDITIONAL INFORMATION FROM THE REP AND A RESPONSE WAS RECEIVED WITH NO ADDITIONAL INFORMATION. THERE IS NO OTHER INFORMATION ABOUT THE EVENT/PATIENT/DEVICE FROM THE CONTACT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. INITIAL REPORTER INFORMATION: UNKNOWN. ENGINEERING EVALUATION: THE FRACTURED HUMERAL LINER IMPACTOR TIP IN THE PHOTOGRAPH APPEARS CONSISTENT WITH A STRESS RISER INTRODUCED DURING IMPACTION, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE. BASED UPON THE REPORTED LOT NUMBER, THIS DEVICE WAS MANUFACTURED AFTER DESIGN CHANGES TO INCREASE THICKNESS AND MODIFY THE CONNECTION FEATURE DIMENSIONS. MOST LIKELY UNDERLYING CAUSE: BASED ON PREVIOUS SIMILAR INVESTIGATIONS, THE FRACTURED HUMERAL LINER IMPACTOR TIP REPORTED WAS LIKELY THE RESULT OF A STRESS RISER INTRODUCED DURING IMPACTION WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE DEVICE. DUE TO A TREND OF THE RECENT COMPLAINTS HAVING A HIGHER OCCURRENCE RATE IN THE KOREAN MARKET, [CAPA] WERE ESTABLISHED TO INVESTIGATE ANY ADDITIONAL POTENTIAL CAUSE OF THE HIGHER OCCURRENCE RATE IN THIS SPECIFIC MARKET.ROOT CAUSE: THE STERILIZATION OF THESE INSTRUMENTS EXCEEDS THE RECOMMENDED PARAMETERS LISTER IN THE REPROCESSING INSTRUCTIONS. THE MATERIAL DEGRADATION FROM STERILIZATION OF THESE INSTRUMENTS BEYOND OUR RECOMMENDED PROCEDURE HAS LIKELY LED TO THE OBSERVED FAILURES. SALES REPRESENTATIVES FROM KOREA WERE NOTIFIED TO FOLLOW REPROCESSING INSTRUCTIONS. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. FOLLOW DEVICE REPROCESSING INSTRUCTIONS. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00561 | 2022-05-24 | 2022-04-20 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | HUMERAL STEM PRIMARY, PRESS FIT 13MM | EXACTECH, INC. | NA | 300-01-13 |  | AS REPORTED, APPROXIMATELY 1 YEAR POST OP THE INITIAL RIGHT TSA, THIS MALE PATIENT WAS REVISED DUE TO AN INFECTION. PATIENT HAD AN INFECTED SHOULDER AND THE SURGEON OPTED TO REMOVE ALL IMPLANTS, WASH OUT AND REPLACE WITH A CEMENT SPACER. THIS IS A SECOND STAGE REVISION AFTER AND INITIAL WASH OUT AND HUMERAL LINER AND GLENOSPHERE EXCHANGE. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. | PLEASE DISREGARD SUBMISSION. EVENT WAS REPORTED IN ERROR. THIS INFORMATION WILL BE CAPTURED IN (B)(4)/ MDR# 1038671-2021-00419. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00598 | 2022-05-24 | 2021-12-19 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | HUMERAL STEM PRIMARY, PRESS FIT 7MM | EXACTECH, INC. | 300-01-07 | 300-01-07 |  | AS REPORTED, 2 DAYS POST OP THE INITIAL TSA, THIS 80 Y/O MALE PATIENT WAS REVISED DUE TO A MIDSHAFT HUMERAL FRACTURE. UNDIAGNOSED INTRA-OPERATIVE HUMERAL FRACTURE CAUSING LOOSENING OF THE HUMERAL STEM AND DISLOCATION. THE CASE REPORT FORM INDICATES THIS EVENT IS POSSIBLY RELATED TO THE DEVICE AND DEFINITELY RELATED TO THE PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | SECTION H10: (H3) PLEASE DISREGARD THE PREVIOUS REPORT. THIS EVENT HAS BEEN DETERMINED TO BE A DUPLICATE AND ALL FURTHER INFORMATION WILL BE CAPTURED IN 1038671-2022-00599 (B)(4). | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00600 | 2022-05-24 | 2022-04-27 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 42MM HUMERAL CONST LINER +2.5 | EXACTECH, INC. | 320-42-13 | 320-42-13 |  | BETWEEN REVISION ON 4/13 AND 4/26, PATIENT HAD REVERSE SHOULDER SURGERY ON (B)(6) 2022. ON (B)(6) 2022, PATIENT HAD REVISION DUE TO AN INFECTION AT THE WOUND SITE. HUMERAL LINER WAS UPSIZED TO A 42MM +2.5 CONSTRAINED. AN EXPERIENCE REPORT WAS SUBMITTED ON (B)(6) 2022. ON (B)(6) 2022 PATIENT HAD ANOTHER REVISION DUE TO SUSPECTED INFECTION. SURGEON EXCHANGED STEM, GLENOSPHERE, ADAPTER PLATE TRAY AND HUMERAL LINER. FIRST REVISION IS CAPTURED IN (B)(6) / MDR# 1038671-2022-00525. | PENDING EVALUATION. CONCOMITANT DEVICE(S): 320-20-00, 7272969 - EQ REVERSE TORQUE DEFINING SCREW KIT. 300-30-07, 7216415 - EQUINOXE PRESERVE STEM 7MM. 320-10-00, 7066657 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0. 320-15-05, 7246544 - EQ REV LOCKING SCREW. 320-06-42, 7219545 - GLENOSPHERE 42MM. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10116 | 2022-05-24 |  | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | BROACH HANDLE | EXACTECH, INC. |  | 161-04-00 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGICAL PROCEDURE A BROACH HANDLE DID NOT FUNCTION. THERE WAS NO PART OF THE BROACH HANDLE THAT FRAGMENTED OFF INTO THE PATIENT. THE AGENT STATED THAT THE BROACH WAS BROKEN AS IN IT DOES NOT WORK. NO DELAY TO THE SURGICAL PROCEDURE, THE SURGEON ALWAYS HAS OTHER BROACH HANDLES AVAILABLE. THE ENGINEERING ANALYSIS FOUND THE TIP OF THE BROACH LEVER THAT ATTACHES TO THE BROACH LOCKING PIN FRACTURED. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE COMPLAINT DEVICE WAS RECEIVED FOR ANALYSIS. THE CONDITION OF BROKEN WAS CONFIRMED. VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR 10X OPTICAL) THE TIP OF THE BROACH LEVER THAT ATTACHES TO THE BROACH LOCKING PIN WAS FRACTURED OFF AND NOT RETURNED. THIS FRACTURE APPEARS CONSISTENT WITH BRITTLE FAILURE DUE TO THE APPLICATION OF EXCESSIVE FORCE. BONE FRAGMENTS OR TISSUE BECOMING LODGED BETWEEN THE INTERFACING SURFACES OF THE BROACH AND BROACH HANDLE MAY NOT ALLOW FOR PROPER ENGAGEMENT OF THE MECHANISM. THIS MAY LEAD TO THE APPLICATION OF EXCESSIVE FORCES WHILE ATTEMPTING TO CLOSE THE BROACH LEVER. THE TIP THAT FRACTURED OFF WAS NOT RETURNED WITH THE DEVICE. THE DISASSEMBLED LATCH RELEASE APPEARS TO BE CONSISTENT WITH NOT STAYING IN PLACE DUE TO FAILURE OF THE ROLL PIN THAT FIXES THE LATCH RELEASE TO THE LOCKING PIN. THE ROLL PIN APPEARS TO HAVE SHEARED OFF. FROM HOW THE MATERIAL DEFORMATION OF THE ROLL PIN IS FOLDED OVER IN ONE DIRECTION, IT APPEARS THE SHEARING OCCURRED. THIS IS THE SAME DIRECTION AS IF ONE WERE TO APPLY A FORCE TO RELEASE THE MECHANISM. NO OTHER MARKINGS ARE APPARENT ON THE DEVICE TO INDICATE THE USE OF EXCESSIVE FORCES IN ATTEMPTING TO RELEASE THE LATCH. BASED ON THE LASER MARKINGS, THIS DEVICE HAS BEEN REWORKED TO AT LEAST REV B OF THE DESIGN. THE RELEASE LATCH (167-201-06) FROM THE FIRST DESIGN ITERATION OF THE DEVICE. (B)(4) CALLS FOR THIS COMPONENT TO BE REPLACED BY 167-201-09 WITH A SMALLER 0.039IN DIAMETER FEATURE THAN THE LARGER ONE WHEN THE DEVICE IS REWORKED TO REV B. THUS, THE ROLL PIN WAS NOT ABLE TO FIXATE THE RELEASE LATCH AS IT WAS DESIGNED TO. IT APPEARS THIS ALLOWED FOR THE INERTIAL FORCES OF THE RELEASE LATCH TO FATIGUE THE ROLL PIN DURING NORMAL USE OF THE DEVICE. THIS RESULTED IN THE DUCTILE FAILURE OF THE ROLL PIN. THIS FAILURE DOES NOT COMPROMISE THE USE OF THE DEVICE. ONE CAN STILL HOLD THE LOCKING LEVER IN PLACE BY HAND WHEN IMPACTING THE BROACH HANDLE. A CORRECTION AND REMOVAL COMMITTEE MEETING WILL BE HELD TO ADDRESS THIS MANUFACTURING ISSUE. REFER TO (B)(4) FOR DETAILS. THE FREQUENCY OF OCCURRENCE RANKING SCALE IS VERY LOW; THEREFORE, THIS DOES NOT APPEAR TO BE DESIGN-RELATED. THE FAILURE OF THE ROLE PIN AND RELEASE LATCH MECHANISM WAS FOUND TO BE MANUFACTURING RELATED. IN THE REWORK PROCESS THE OLD RELEASE LATCH (167-201-06) WAS NOT REPLACED WITH THE NEW RELEASE LATCH (167-201-09) WHICH LEAD TO THE FAILURE OF THE ROLL PIN. THIS WILL BE ADDRESSED IN (B)(4). EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF (B)(4) UNITS, WHICH HAS BEEN IN THE FIELD SINCE 2014. THEREFORE, THE FAILURE OF THE BROACH LEVER TIP DOES NOT APPEAR TO BE MANUFACTURING-RELATED. THE BROKEN DEVICE REPORTED IN EXPERIENCE (B)(4) WAS MOST LIKELY A RESULT OF BONE OR TISSUE BECOMING LODGED BETWEEN THE MATING SURFACES OF THE BROACH AND BROACH HANDLE. THIS MAY HAVE PREVENTED PROPER ENGAGEMENT OF THE BROACH LOCKING MECHANISM AND THE USE OF EXCESSIVE FORCE ON THE BROACH LEVER LEADING TO ULTIMATE FAILURE OF THE BROACH LEVER TIP. THE FAILED ROLL PIN WAS MOST LIKELY THE RESULT OF A MANUFACTURING ERROR WHERE THE INCORRECT PIECE WAS INSTALLED. A CORRECTION AND REMOVAL COMMITTEE MEETING WILL BE HELD TO ADDRESS THIS MANUFACTURING ISSUE. REFER TO (B)(4) FOR DETAILS. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT IMPACT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN DEVICE REPORTED WAS MOST LIKELY A RESULT OF BONE OR TISSUE BECOMING LODGED BETWEEN THE MATING SURFACES OF THE BROACH AND BROACH HANDLE. THIS MAY HAVE PREVENTED PROPER ENGAGEMENT OF THE BROACH LOCKING MECHANISM AND THE USE OF EXCESSIVE FORCE ON THE BROACH LEVER LEADING TO ULTIMATE FAILURE OF THE BROACH LEVER TIP. THE FAILED ROLL PIN WAS MOST LIKELY THE RESULT OF A MANUFACTURING ERROR WHERE THE INCORRECT PIECE WAS INSTALLED. INFORMATION WAS NOT PROVIDED FOR: PT. INFO., INITIAL REPORTER NAME AND ADDRESS. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10142 | 2022-05-24 | 2017-08-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRULIANT TIBIAL INSERT TRIAL/TRAY ADAPTER | EXACTECH, INC. |  | 02-029-25-4050 |  | IT WAS DISCOVERED IN THE COMPANY RECEIVING AREA THAT THE TIBIAL INSERT TRIAL/TRAY ADAPTER WAS BROKEN. THE PLATES' CENTRAL FLEXURE DISASSOCIATED FROM THE PLATES. IT IS NOT CLEAR WHEN AND WHERE THE DISASSOCIATION TOOK PLACE. THERE IS NO ADDITIONAL INFORMATION AVAILABLE ABOUT THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE ROOT CAUSE ASSOCIATED WITH CAPA(B)(4) RESIDES IN THE MISUSE/MISTREATMENT OF THE ADAPTOR PLATES OR CANNOT BE IDENTIFIED DUE TO INSUBSTANTIAL EVIDENCE. THE CAPA TEAM HAS DETERMINED, HOWEVER, THAT USE CONSIDERATIONS INDICATE THAT THE FLEXURE DESIGN IS SUSCEPTIBLE TO DAMAGE FROM SAID MISUSE/MISTREATMENT. DESIGN -A REVIEW OF THE TOLERANCE STACK-UP PERFORMED IN TM-2017-0417 WAS CONDUCTED. AS PART OF THIS INVESTIGATION, THE TOLERANCE STACK-UP BETWEEN THE TRULIANT TIBIAL TRIAL ADAPTOR PLATES AND THE CORRESPONDING TRULIANT AND LOGIC TIBIAL TRAYS WAS REVISITED TO CONFIRM IT HAD BEEN ADEQUATELY PERFORMED. THE REVIEW WAS FOUND TO BE ACCURATE. THERE IS NO NEW INFORMATION THAT HAS BEEN FOUND THAT WOULD ALTER THE RESULTS OR METHODS OF SAID ANALYSIS. IN ADDITION, ATTEMPTS WERE MADE IN THE LABORATORY TO REPLICATE FAILURE MODES WITHOUT SUCCESS. THE DESIGN FUNCTIONS AS INTENDED UNDER WORST CASE CONDITIONS. MANUFACTURING - REVIEW OF THE DEVICE DHR CONCLUDED THAT ALL PARTS WERE MADE WITHIN SPECIFICATION. USE CONSIDERATIONS - THE DESIGN DOES LEND TO BEING SUSCEPTIBLE TO BREAKAGE IF USED INCORRECTLY DUE TO THE USE OF ULTEM MATERIAL WITH A THINNER ASPECT UNDER LOAD. IN ADDITION, THE DEVICES BROKE AT SOME POINT DURING REPROCESSING AT ONE PARTICULAR INSTITUTE. A REVIEW WAS CONDUCTED AT THIS INSTITUTE'S REPROCESSING PROCEDURES AND RESULTS INDICATED THAT CORRECT REPROCESSING INSTRUCTIONS PROVIDED BY EXACTECH WAS FOLLOWED. A REVIEW OF THE INSTRUMENT TRAY CASE INDICATED THAT IT IS POSSIBLE THAT THE ADAPTOR PLATES MAY NOT BE FULLY CONSTRAINED DURING REPROCESSING. DUE TO THE NATURE OF THE REPROCESSING PROCESS IT IS POSSIBLE THE PLATES BECAME DISLODGED AND DAMAGED ELSEWHERE IN THE MACHINE. HUMAN ERROR - HUMAN ERROR WAS CONSIDERED AS PART OF THE INVESTIGATION AND FOUND TO BE THE ROOT CAUSE OF ONE COMPLAINT. THE ADAPTOR PLATE THAT WAS USED IN SAID COMPLAINT WAS THE INCORRECT SIZE RELATIVE TO THE TIBIAL TRAY THAT HAD BEEN IMPLANTED. THIS MAY HAVE CAUSED THE PLATE TO DISLODGE UNPREDICTABLY. IN ADDITION, A HOHMANN RETRACTOR WAS USED TO REMOVE THE PLATE AND NOT THE TRULIANT DROP ROD HANDLE, AS STATED PER THE OPERATIVE TECHNIQUE, 712-35-31. THE RETRACTOR IS CURVED AND FEATURES A MUCH SMALLER NOSE WHICH COULD INADVERTENTLY FACILITATE CONTACT WITH THE FLEXURE DURING REMOVAL. THE GEOMETRY OF THE TRULIANT DROP ROD HANDLE, IN ADDITION TO AN APPROPRIATELY CONSTRAINED ADAPTOR PLATE, AS DESIGNED CREATES A SCENARIO IN WHICH CONTACT WITH THE FLEXURE IS EXTREMELY DIFFICULT. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10127 | 2022-05-24 | 2017-10-19 | I, n, j, u, r, y | Manufacturer report | EXACTECH | LPI IMPACTOR HANDLE | EXACTECH, INC. |  | 213-65-00 |  | IT WAS ORIGINALLY REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGICAL PROCEDURE THE LPI HANDLE BROKE. ALL PIECES WERE RETRIEVED FROM THE PATIENT. THE SURGERY WAS SUCCESSFUL. THERE WAS NO DELAY TO SURGERY AND NO ADVERSE EFFECT TO THE PATIENT. NO MORE DETAILS AVAILABLE. ON 9 MAR 2021, ADDITIONAL INFORMATION WAS RECEIVED: "THE INSTRUMENT COULD HAVE BEEN BROKEN DURING THE SURGERY. HOWEVER, THE BREAKAGE HAS BEEN NOTICED AT THE RECOVERY OF THE INSTRUMENTATION-SET BY OUR LOGISTICS DEPARTMENT, AFTER THE CLEANING REALIZED BY THE STERILIZATION TEAM OF THE HOSPITAL. NO INFORMATION HAS BEEN RECEIVED FROM THE HOSPITAL REGARDING THIS ISSUE AND [THE AGENCY] DIDN'T RECEIVE ANY REQUEST FOR COMPLEMENTARY INFORMATION REGARDING THIS ISSUE AS WELL. PATIENT INFORMATION: SURGERIES WERE ALL SUCCESSFUL AND NO DELAYS TO ANY PATIENTS." MULTIPLE ATTEMPTS WERE MADE FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED. EVAL COMPLETED (B)(6) 2018 BY CC. AN INVESTIGATION OF COMPLAINTS RELATED TO LPI HANDLE - BROKEN NON-FUNCTIONAL, SPECIFICALLY RELATED TO CRACKS/BREAKAGE FOR THE PLASTIC HANDLE WAS CONDUCTED UNDER TM 2018-XX. THE LPI HANDLE CONSISTS OF A METAL STRIKE PLATE, METAL END PIECE WHICH ENGAGES WITH MATING INSTRUMENTS AND IMPLANTS, AND A PLASTIC CENTRAL HANDLE. THE METAL END PIECES ARE ASSEMBLED TO THE PLASTIC HANDLE VIA THERMAL ASSEMBLY WHICH INVOLVES HEATING THE METAL AND PRESSING THE HOT METAL PIECES INTO THE HANDLE. THE METAL PIECES HAVE CYLINDRICAL BOSSES WHICH FEATURE A PRESS -FIT WITH A FEMALE HOLE ON THE PLASTIC HANDLE. THE LPI HANDLE ASSEMBLES WITH TWO END PIECES FOR USE THREE TIMES DURING THE SURGICAL PROCEDURE: IMPACTION OF THE FEMORAL TRIAL, IMPACTION OF THE TIBIAL TRAY IMPLANT AND IMPACTION OF THE FEMORAL IMPLANT. AVAILABLE DEVICES FROM PREVIOUSLY REPORTED COMPLAINTS WERE EVALUATED. COMPLAINT DEVICES APPEARED TO HAVE SIMILAR FAILURE MODES, AND EITHER WERE CRACKED OR FRACTURED. CRACKS WERE CONSISTENT WITH STARTING INSIDE THE PLASTIC HANDLE AT THE BASE OF ONE OF THE TWO METAL ATTACHMENT FEATURES AND SPREADING OUTWARD. CRACKS WERE EITHER LINEAR, ALONG THE HANDLE AXIS OR HORIZONTAL, NEARLY IN LINE WITH THE ORIGIN. THE REVIEW OF THE LPI IMPACTOR HANDLES (213-65-00) CITED IN COMPLAINTS FOR BROKEN OR CRACKED HANDLES SHOWED A CONSISTENT FAILURE MODE. A REVIEW OF THE DEVICE ORDER HISTORY DID NOT SHOW A TREND IN THOSE LOTS WITH COMPLAINT DEVICES, NOR DID A REVIEW OF DESIGN CHANGES SHOW CHANGES MADE TO THE DESIGN WHICH COULD CONTRIBUTE TO THE COMPLAINT. PER 750-2009-023-RMR-INSTRUMENT, REV B - DEVICE FRACTURE/BREAKAGE (LINE ITEM 18) CAN LEAD TO A SURGICAL DELAY. THE PREDICTED SEVERITY OF THIS HARM IS MODERATE (3), DEFINED AS "ANY INSTRUMENT - RELATED HAZARD THAT COULD RESULT IN MARKED INTRA-OPERATIVE INCONVENIENCE TO THE SURGEON AND STAFF", AND THE PREDICTED OCCURRENCE IS LOW (2), DEFINED AS LESS THAN OR EQUAL TO 1%. (B)(6) WAS INITIATED ON (B)(6) 2018 TO EVALUATE THIS ISSUE. BASED ON THE AVAILABLE DATA, THE DEVICE IS PERFORMING WITHIN ITS PREDICTED RISK. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. IT WAS REPORTED THAT THE HANDLE BROKE AND ALL PIECES WERE RETRIEVED FROM THE PATIENT. AN INVESTIGATION WAS CONDUCTED; THE MOST LIKELY CAUSE OF THE REPORTED EVENTS IS RELATED TO HIGH IMPACTION FORCES REQUIRED DURING IMPLANT SUGERY. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10131 | 2022-05-24 | 2018-03-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | CC SPINE TRIAL SZ 4 | EXACTECH, INC. |  | 209-52-04 |  | IT WAS REPORTED FROM AUSTRALIA THAT THERE WAS A BROKEN TRIALS SPINE. THE BROKEN INSTRUMENTS WERE FOUND WHEN THE LOANER KIT WAS RECEIVED IN THE WAREHOUSE. THE PATIENT WAS STABLE AFTER SURGERY. THERE WAS NO REPORTED ADVERSE EVENT FOR THE PATIENT. THERE WERE NO AGENTS PRESENT AT THE TIME OF SURGERY. NO ADDITIONAL INFORMATION IS AVAILABLE. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. 5955-002 - MFG DATE: 03/30/01 34235006 - MFG DATE: 03/19/10. THE COMPLAINT PRODUCT WAS RECEIVED FOR ANALYSIS. THE REPORTED MALFUNCTION FOR BROKEN TRIAL SPINE WAS CONFIRMED. VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR 10X OPTICAL) FRACTURED SURFACE OF THE SIZE 3 SPINE TRIAL , APPEARS CONSISTENT WITH THE SPINE TRIAL EXPERIENCING A HIGH STRESS CONCENTRATION THAT WAS LIKELY CREATED WHEN REMOVING THE SPINE TRIAL FROM THE TIBIAL INSERT TRIAL. MISSING COLOR DOT ON THE SIZE 3 SPINE TRIAL WAS LIKELY THE RESULT OF YEARS OF USE AND THE STERILIZATION AND REPROCESSING PROCEDURES PERFORMED BY THE HOSPITAL.. FRACTURED SURFACE OF THE SIZE 4 SPINE TRIAL APPEARS CONSISTENT WITH THE SPINE TRIAL EXPERIENCING A HIGH STRESS CONCENTRATION THAT WAS LIKELY CREATED WHILE REMOVING THE SPINE TRIAL FROM THE TIBIAL INSERT TRIAL. EXACTECH IS NOT AWARE OF RECEIVING ANY OTHER COMPLAINT REPORTS INVOLVING A PART FROM THE MANUFACTURING LOT 34235006 OF 50 PIECES THAT HAS BEEN IN THE FIELD SINCE 2010. EXACTECH IS NOT AWARE OF RECEIVING ANY OTHER COMPLAINT REPORTS INVOLVING A PART FROM THE MANUFACTURING LOT 5955-002 OF 35 PIECES THAT HAS BEEN IN THE FIELD SINCE 2001. THEREFORE, THIS DOES NOT APPEAR TO BE MANUFACTURING-RELATED. THE COMPLAINT REPORTED IN EXPERIENCE (B)(4) WAS LIKELY THE RESULT OF EXCESSIVE FORCES APPLIED DURING THE REMOVAL OF THE DEVICE. MATERIAL DEGRADATION FROM YEARS OF USE AND CLEANING CYCLES MAY HAVE ALSO CONTRIBUTED TO THE FAILURE MODE. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. AN INVESTIGATION WAS CONDUCTED; THE EVENT REPORTED WAS LIKELY THE RESULT OF EXCESSIVE FORCES APPLIED DURING THE REMOVAL OF THE DEVICE. MATERIAL DEGRADATION FROM YEARS OF USE AND CLEANING CYCLES MAY HAVE ALSO CONTRIBUTED TO THE FAILURE MODE. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10130 | 2022-05-24 | 2018-03-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | CC SPINE TRIAL SZ 4 | EXACTECH, INC. |  | 209-52-04 |  | IT WAS REPORTED FROM AUSTRALIA THAT THERE WAS A BROKEN TRIALS SPINE. THE BROKEN INSTRUMENTS WERE FOUND WHEN THE LOANER KIT WAS RECEIVED IN THE WAREHOUSE. THE PATIENT WAS STABLE AFTER SURGERY. THERE WAS NO REPORTED ADVERSE EVENT FOR THE PATIENT. THERE WERE NO AGENTS PRESENT AT THE TIME OF SURGERY. NO ADDITIONAL INFORMATION IS AVAILABLE. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. 5955-002 - MFG DATE: 03/30/2001 34235006 - MFG DATE: 03/19/2010. THE COMPLAINT PRODUCT WAS RECEIVED FOR ANALYSIS. THE REPORTED MALFUNCTION FOR BROKEN TRIAL SPINE WAS CONFIRMED. VISUAL EVALUATION CONDITION/FINDINGS (CONDUCTED WITH A 7X OR 10X OPTICAL) FRACTURED SURFACE OF THE SIZE 3 SPINE TRIAL , APPEARS CONSISTENT WITH THE SPINE TRIAL EXPERIENCING A HIGH STRESS CONCENTRATION THAT WAS LIKELY CREATED WHEN REMOVING THE SPINE TRIAL FROM THE TIBIAL INSERT TRIAL. MISSING COLOR DOT ON THE SIZE 3 SPINE TRIAL WAS LIKELY THE RESULT OF YEARS OF USE AND THE STERILIZATION AND REPROCESSING PROCEDURES PERFORMED BY THE HOSPITAL. FRACTURED SURFACE OF THE SIZE 4 SPINE TRIAL APPEARS CONSISTENT WITH THE SPINE TRIAL EXPERIENCING A HIGH STRESS CONCENTRATION THAT WAS LIKELY CREATED WHILE REMOVING THE SPINE TRIAL FROM THE TIBIAL INSERT TRIAL. EXACTECH IS NOT AWARE OF RECEIVING ANY OTHER COMPLAINT REPORTS INVOLVING A PART FROM THE MANUFACTURING LOT 34235006 OF (B)(4) PIECES THAT HAS BEEN IN THE FIELD SINCE 2010. EXACTECH IS NOT AWARE OF RECEIVING ANY OTHER COMPLAINT REPORTS INVOLVING A PART FROM THE MANUFACTURING LOT 5955-002 OF (B)(4) PIECES THAT HAS BEEN IN THE FIELD SINCE 2001. THEREFORE, THIS DOES NOT APPEAR TO BE MANUFACTURING-RELATED. THE COMPLAINT REPORTED IN EXPERIENCE (B)(4) WAS LIKELY THE RESULT OF EXCESSIVE FORCES APPLIED DURING THE REMOVAL OF THE DEVICE. MATERIAL DEGRADATION FROM YEARS OF USE AND CLEANING CYCLES MAY HAVE ALSO CONTRIBUTED TO THE FAILURE MODE. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. AN INVESTIGATION WAS CONDUCTED; THE EVENT REPORTED WAS LIKELY THE RESULT OF EXCESSIVE FORCES APPLIED DURING THE REMOVAL OF THE DEVICE. MATERIAL DEGRADATION FROM YEARS OF USE AND CLEANING CYCLES MAY HAVE ALSO CONTRIBUTED TO THE FAILURE MODE. INFORMATION WAS NOT PROVIDED FOR: PT. INFO. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10124 | 2022-05-24 | 2017-11-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LPI LOCKING FEMORAL IMPACTOR | EXACTECH, INC. |  | 213-64-01 |  | IT WAS REPORTED FROM OUS THAT WHEN THE FEMORAL COMPONENT WAS IMPACTED WITH THE LOCKING IMPACTOR DURING THE PROCEDURE, THE PLASTIC WEDGE BROKE. THE PATIENT'S HEALTH WAS STABLE LEAVING THE OR. THE AGENT WAS PRESENT AT THE SURGERY AND REPORTED NO PIECES WERE LEFT IN THE PATIENT. NO PIECES FELL INTO THE SURGICAL SITE. THE ISSUE WAS RESOLVED BY USING A SPARE SET. THERE WAS NO DELAY TO THE SURGERY AND NO CLINICAL CONSEQUENCE TO THE PATIENT. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE WAS NOT RETURNED BECAUSE IT WAS LOST. ENG EVAL COMPLETED ON 02/27/2019 BY SH. DESIGN-RELATED ISSUES: THE DESIGN OF THE LPI LOCKING FEMORAL IMPACTOR HAS BEEN IN THE FIELD SINCE 2005. EXACTECH HAS RECEIVED 18 SIMILAR COMPLAINT REPORTS INVOLVING LPI LOCKING FEMORAL IMPACTOR DEVICES SINCE 2014. THE FREQUENCY OF OCCURRENCE RANKING SCALE IS VERY LOW; THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF RECEIVING ANY COMPLAINT REPORTS INVOLVING ANOTHER PART FROM THIS MANUFACTURING LOT OF 30 PIECES WHICH HAS BEEN IN THE FIELD SINCE 2007. EXACTECH COMPLAINT DATA FROM 2014 THROUGH 2018 INVOLVING THE REPORTED FAILURE FOR THIS FAMILY OF DEVICES WAS REVIEWED. THE INVESTIGATIONS FOR THOSE INCIDENCES HAVE NOT IDENTIFIED ANY EVIDENCE THAT A MANUFACTURING ISSUE CAUSED OR CONTRIBUTED TO THIS REPORTED ISSUE. THEREFORE, THIS DOES NOT APPEAR TO BE MANUFACTURING-RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE BROKEN FEMORAL IMPACTOR REPORTED IN C2017-642 WAS LIKELY THE RESULT OF YEARS OF NORMAL USE WHICH POSSIBLY LED TO FRACTURE OF THE IMPACTION SURFACE. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES. THERE WAS NO PATIENT ADVERSE EVENT. NO PIECES FELL INTO THE SURGICAL SITE. AN INVESTIGATION WAS CONDUCTED; THE BROKEN FEMORAL IMPACTOR REPORTED WAS LIKELY THE RESULT OF YEARS OF NORMAL USE WHICH POSSIBLY LED TO FRACTURE OF THE IMPACTION SURFACE. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10126 | 2022-05-24 | 2017-10-19 | I, n, j, u, r, y | Manufacturer report | EXACTECH | OPTETRAK LOCKING FEMORAL IMPACTOR | EXACTECH, INC. |  | 213-64-01 |  | IT WAS ORIGINALLY REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGICAL PROCEDURE THE INSTRUMENT BROKE DURING IMPACTION. IT WAS STATED THAT THERE WAS NO DELAY TO SURGERY AND NO ADVERSE EVENT TO PATIENT. "ALL PIECES WERE RETRIEVED FROM THE PATIENT. THE SURGERY WAS SUCCESSFUL." ADDITIONAL INFORMATION RECEIVED ON 9 MAR 2021 STATES THE FOLLOWING: "THE INSTRUMENT COULD HAVE BEEN BROKEN DURING THE SURGERY. HOWEVER, THE BREAKAGE HAS BEEN NOTICED AT THE RECOVERY OF THE INSTRUMENTATION-SET BY OUR LOGISTICS DEPT., AFTER THE CLEANING REALIZED BY THE STERILIZATION TEAM OF THE HOSPITAL. NO INFORMATION HAS BEEN RECEIVED FROM THE HOSPITAL REGARDING THIS ISSUE AND [THE AGENCY] DIDN'T RECEIVE ANY REQUEST FOR COMPLEMENTARY INFORMATION REGARDING THIS ISSUE AS WELL. PATIENT INFORMATION: SURGERIES WERE ALL SUCCESSFUL AND NO DELAYS TO ANY PATIENTS. " NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DEVICE WAS RECEIVED FOR ANALYSIS. ENG ANALYSIS COMPLETED (B)(6) 2018 BY AD THE BROKEN IMPACTOR FACE APPEARS CONSISTENT WITH MATERIAL DEGRADATION ASSOCIATED WITH YEARS OF NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. THE OTHER HALF OF THE BROKEN FACE WAS NOT RETURNED. BROKEN WELD ON IMPACTOR BODY AND IMPACTOR HANDLE. THE BROKEN WELD APPEARS CONSISTENT WITH UNAUTHORIZED DISASSEMBLY OF THE DEVICE DURING CLEANING. THE COMPRESSION SPRING REMAINED INTACT AND WAS UNAFFECTED BY THE BREAKAGE. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE AND CONSISTENT WITH YEARS OF USE. DESIGN: THE OCCURRENCE RATE IS LESS THAN 0.5%, "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THIS DOES NOT APPEAR TO BE DESIGN RELATED. MANUFACTURING: THERE HAVE NOT BEEN ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 70 UNITS, WHICH HAS BEEN IN THE FIELD SINCE 2010. THIS DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE BROKEN IMPACTOR FACE REPORTED WAS MOST LIKELY THE RESULT OF MATERIAL DEGRADATION ASSOCIATED WITH YEARS OF NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. THE BROKEN WELD WAS LIKELY THE RESULT OF UNAUTHORIZED DISASSEMBLY OF THE DEVICE DURING CLEANING. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. IT WAS STATED THAT "ALL PIECES WERE RETRIEVED FROM THE PATIENT". AN INVESTIGATION WAS CONDUCTED; THE BROKEN IMPACTOR FACE REPORTED WAS MOST LIKELY THE RESULT OF MATERIAL DEGRADATION ASSOCIATED WITH YEARS OF NORMAL USE AND SUBSEQUENT STERILIZATION CYCLES. THE BROKEN WELD WAS LIKELY THE RESULT OF UNAUTHORIZED DISASSEMBLY OF THE DEVICE DURING CLEANING. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10128 | 2022-05-24 | 2017-12-29 | I, n, j, u, r, y | Manufacturer report | EXACTECH | SHORT STEM PROTECTOR | EXACTECH, INC. |  | 301-07-20 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGICAL PROCEDURE AN INSTRUMENT BROKE. THE REPRESENTATIVE WAS PRESENT AT THE SURGERY. THE PATIENT'S HEALTH WAS STABLE LEAVING THE OR. ALL PIECES WERE RETRIEVED FROM THE PATIENT. NO DELAY TO SURGERY. NEW INFORMATION WAS RECEIVED 15 MAR 2021, "THE INSTRUMENT COULD HAVE BEEN BROKEN DURING THE SURGERY. HOWEVER, THE BREAKAGE HAS BEEN NOTICED AT THE RECOVERY OF THE INSTRUMENTATION-SET BY OUR LOGISTICS DEPT., AFTER THE CLEANING REALIZED BY THE STERILIZATION TEAM OF THE HOSPITAL. NO INFORMATION HAS BEEN RECEIVED FROM THE HOSPITAL REGARDING THIS ISSUE AND [THE AGENCY] DIDN'T RECEIVE ANY REQUEST FOR COMPLEMENTARY INFORMATION REGARDING THIS ISSUE AS WELL. SURGERY." MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED FOR ANALYSIS. ENG ANALYSIS COMPLETED ON (B)(6) 2018 BY AD THE CAP FOR THE COMPONENT WAS NOT RETURNED. NO FAILURE OF THE RETURNED PLATE COMPONENT WAS OBSERVED. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE. DESIGN: THIS DESIGN HAS BEEN IN THE FIELD SINCE 2004. THE OCCURRENCE RATE IS LESS THAN 0.5%, "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THIS DOES NOT APPEAR TO BE DESIGN RELATED. MANUFACTURING: THERE HAVE BEEN NO OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 30 UNITS, WHICH HAS BEEN IN THE FIELD SINCE 2016. THIS DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF BEING INADVERTENTLY IMPACTED, LEADING TO THE CAP COMPONENT TO BREAK. HOWEVER, THIS CANNOT BE CONFIRMED AS THE CAP COMPONENT WAS NOT AVAILABLE FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. IT WAS STATED THAT ALL PIECES WERE RETRIEVED FROM THE PATIENT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF BEING INADVERTENTLY IMPACTED, LEADING TO THE CAP COMPONENT TO BREAK. HOWEVER, THIS CANNOT BE CONFIRMED AS THE CAP COMPONENT WAS NOT AVAILABLE FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10132 | 2022-05-24 | 2018-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | MCS LINER TRIAL | EXACTECH, INC. |  |  |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC REVISION SURGERY THE MCS LINER TRIAL BROKE DURING TRIAL REDUCTION. THE LINER IMPACTOR HANDLE HAS A DAMAGED TIP AND IT WILL NOT ACCEPT THE LINER IMPACTOR TIPS, THE RUBBER RING CAME OFF. THERE WAS NO DELAY TO SURGERY OR PATIENT ADVERSE EVENT. THE PATIENT'S HEALTH WAS STABLE LEAVING THE OR. THE AGENT WAS PRESENT AT THE TIME OF SURGERY. THE AGENCY IS INACTIVE AND NO FURTHER INFORMATION IS AVAILABLE. THE REVISION SURGERY WAS CAPTURED IN 1038671-2018-00587 (FEMORAL HEAD) & 1038671-2018-00569 (MCS LINER). | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. LINER IMPACTOR HANDLE; CATALOG AND LOT NUMBER NOT REPORTED. DEVICES NOT AVAILABLE FOR ANALYSIS. DESIGN-RELATED ISSUES: NOT ENOUGH INFORMATION REGARDING THE CATALOG NUMBERS WAS PROVIDED TO DETERMINE IF THERE WERE ANY DESIGN-RELATED ISSUES. MFG-RELATED ISSUES: MANUFACTURING-RELATED ISSUES COULD NOT BE DETERMINED BECAUSE NO INFORMATION REGARDING THE SERIAL NUMBER OR MANUFACTURING LOT WAS PROVIDED. THE BROKEN DEVICES REPORTED COULD NOT BE CONFIRMED AS THE DEVICES WERE NOT AVAILABLE FOR EVALUATION AND NOT ENOUGH IDENTIFYING INFORMATION WAS PROVIDED. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT IMPACT. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT IMPACT. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10139 | 2022-05-24 | 2016-12-22 | I, n, j, u, r, y | Manufacturer report | EXACTECH | DRIVER, MODULAR CANNULATED TRI-DRIVE | EXACTECH, INC. |  | 315-25-00 |  | IT WAS REPORTED FROM OUS THAT THE INSTRUMENT BROKE DURING AN ORTHOPEDIC SURGICAL PROCEDURE. ALL PIECES WERE RETRIEVED FROM THE PATIENT. THE SURGERY WAS SUCCESSFUL WITH NO DELAY. ADDITIONAL INFORMATION RECEIVED ON 9 MAR 2021- "THE INSTRUMENT COULD HAVE BEEN BROKEN DURING THE SURGERY. HOWEVER, THE BREAKAGE HAS BEEN NOTICED AT THE RECOVERY OF THE INSTRUMENTATION-SET BY OUR LOGISTICS DEPT., AFTER THE CLEANING REALIZED BY THE STERILIZATION TEAM OF THE HOSPITAL. NO INFORMATION HAS BEEN RECEIVED FROM THE HOSPITAL REGARDING THIS ISSUE AND [THE AGENCY] DIDN'T RECEIVE ANY REQUEST FOR COMPLEMENTARY INFORMATION REGARDING THIS ISSUE AS WELL." MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. ENGINEERING EVALUATION: THE DESIGNS OF THE CANNULATED TRI-DRIVE HANDLE HAVE BEEN IN THE FIELD SINCE 2009. EXACTECH HAS RECEIVED 36 COMPLAINT REPORTS INVOLVING THE TRI-DRIVE HANDLE SINCE ITS INTRODUCTION. THIS IS RATED "LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. MANUFACTURING: ONE OTHER COMPLAINT REPORT HAS BEEN RECEIVED INVOLVING A PART FROM LOT 68106010. THE INSTRUMENT EVENT REPORTED WAS LIKELY THE RESULT OF WEAR AND DEGRADATION DUE TO MANY USES AND SUBSEQUENT STERILIZATION CYCLES. THE ASSEMBLY ISSUE INVOLVING THE STEM TRIAL WAS LIKELY THE RESULT OF THREAD DEFORMATION ON THE TRIAL OR THE MATING INSTRUMENT. HOWEVER, THE MOST LIKELY UNDERLYING CAUSES OF THE REPORTED ISSUES CANNOT BE CONFIRMED BECAUSE THE INSTRUMENT WAS NOT AVAILABLE FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10122 | 2022-05-24 | 2017-09-08 | I, n, j, u, r, y | Manufacturer report | EXACTECH | LPI IMPACTOR HANDLE | EXACTECH, INC. |  | 213-65-00 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC SURGERY THE LPI HANDLE BROKE. IT WAS STATED THAT THERE WERE NO PIECES WERE LEFT IN THE PATIENT. NO DELAY TO SURGERY AND NO ADVERSE EVENT TO PATIENT. THE ISSUE WAS RESOLVED BY USING ANOTHER IMPACTOR FOR THE SURGICAL PROCEDURE. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED. EVAL COMPLETED 11/15/2018 BY (B)(6). AN INVESTIGATION OF COMPLAINTS RELATED TO LPI HANDLE - BROKEN NON-FUNCTIONAL, SPECIFICALLY RELATED TO CRACKS/BREAKAGE FOR THE PLASTIC HANDLE WAS CONDUCTED UNDER TM 2018-XX. THE LPI HANDLE CONSISTS OF A METAL STRIKE PLATE, METAL END PIECE WHICH ENGAGES WITH MATING INSTRUMENTS AND IMPLANTS, AND A PLASTIC CENTRAL HANDLE. THE METAL END PIECES ARE ASSEMBLED TO THE PLASTIC HANDLE VIA THERMAL ASSEMBLY WHICH INVOLVES HEATING THE METAL AND PRESSING THE HOT METAL PIECES INTO THE HANDLE. THE METAL PIECES HAVE CYLINDRICAL BOSSES WHICH FEATURE A PRESS -FIT WITH A FEMALE HOLE ON THE PLASTIC HANDLE. THE LPI HANDLE ASSEMBLES WITH TWO END PIECES FOR USE THREE TIMES DURING THE SURGICAL PROCEDURE: 1) IMPACTION OF THE FEMORAL TRIAL 2) IMPACTION OF THE TIBIAL TRAY IMPLANT AND 3) IMPACTION OF THE FEMORAL IMPLANT. AVAILABLE DEVICES FROM PREVIOUSLY REPORTED COMPLAINTS WERE EVALUATED. COMPLAINT DEVICES APPEARED TO HAVE SIMILAR FAILURE MODES, AND EITHER WERE CRACKED OR FRACTURED. CRACKS WERE CONSISTENT WITH STARTING INSIDE THE PLASTIC HANDLE AT THE BASE OF ONE OF THE TWO METAL ATTACHMENT FEATURES AND SPREADING OUTWARD. CRACKS WERE EITHER LINEAR, ALONG THE HANDLE AXIS OR HORIZONTAL, NEARLY IN LINE WITH THE ORIGIN. THE REVIEW OF THE LPI IMPACTOR HANDLES (213-65-00) CITED IN COMPLAINTS FOR BROKEN OR CRACKED HANDLES SHOWED A CONSISTENT FAILURE MODE. A REVIEW OF THE DEVICE ORDER HISTORY DID NOT SHOW A TREND IN THOSE LOTS WITH COMPLAINT DEVICES, NOR DID A REVIEW OF DESIGN CHANGES SHOW CHANGES MADE TO THE DESIGN WHICH COULD CONTRIBUTE TO THE COMPLAINT. PER 750-2009-023-RMR-INSTRUMENT, REV B - DEVICE FRACTURE/BREAKAGE (LINE ITEM 18) CAN LEAD TO A SURGICAL DELAY. THE PREDICTED SEVERITY OF THIS HARM IS MODERATE (3), DEFINED AS "ANY INSTRUMENTRELATED HAZARD THAT COULD RESULT IN MARKED INTRA-OPERATIVE INCONVENIENCE TO THE SURGEON AND STAFF", AND THE PREDICTED OCCURRENCE IS LOW (2), DEFINED AS LESS THAN OR EQUAL TO 1%. CRC-2018-03-02-02 WAS INITIATED ON 03/02/2018 TO EVALUATE THIS ISSUE. BASED ON THE AVAILABLE DATA, THE DEVICE IS PERFORMING WITHIN ITS PREDICTED RISK. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. IT WAS REPORTED THAT THE HANDLE BROKE AND NO PIECES WERE LEFT IN THE PATIENT. AN INVESTIGATION WAS CONDUCTED; THE MOST LIKELY CAUSE OF THE REPORTED EVENTS IS RELATED TO HIGH IMPACTION FORCES REQUIRED DURING IMPLANT SURGERY. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10138 | 2022-05-24 | 2018-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | EQUINOXE REVERSE SHOULDER COMPRESSION SCREW/LOCKING CAP KIT | EXACTECH, INC. |  | 320-20-30 |  | IT WAS REPORTED BY EXACTECH UK THAT WHILE FIXING THE GLENOID PLATE, THE SURGEON HAD DIFFICULTY INSERTING THE SCREW. HE STATED THAT IT WASN'T THREADING AS IT SHOULD. HE DECIDED TO REMOVE THE SCREW AND IMPLANT ANOTHER SCREW. UPON INSPECTION OF THE SCREW THERE APPEARED TO BE A WIRE LIKE PIECE OF METAL WRAPPING AROUND THE SCREW. A PHOTO WAS PROVIDED. POST-SURGERY X-RAYS LOOK GOOD. THE SURGEON SAID HE WAS HAPPY WITH THE POST-OP RESULT. PATIENT'S HEALTH WAS STABLE LEAVING THE OR. ISSUE WAS RESOLVED BY USING ANOTHER SCREW. THE REP WAS PRESENT AT THE TIME OF SURGERY. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PHOTO WAS PROVIDED. ATTEMPTS WERE MADE TO REQUEST ADDITIONAL INFORMATION AND A RESPONSE WAS RECEIVED: NO PARTS/PIECES OF THE DEVICE FELL INTO THE SURGICAL SITE. THERE WAS A MINOR DELAY OF 10 MINUTES. NO ADDITIONAL INFORMATION IS AVAILABLE FOR THE EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. ENGINEERING ANALYSIS COMPLETED BY MB ON 8/3/2018. DESIGN-RELATED ISSUES: THE DESIGN OF THE EQUINOXE REVERSE SHOULDER COMPRESSION SCREW HAS BEEN IN THE FIELD SINCE 2007. EXACTECH HAS RECEIVED 3 SIMILAR COMPLAINT REPORTS INVOLVING EQUINOXE REVERSE SHOULDER COMPRESSION SCREWS SINCE 2008. SALES DATA FOR COMPRESSION SCREWS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH HAS NOT RECEIVED ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT OF 50 UNITS, WHICH HAS BEEN IN THE FIELD SINCE 2017. ACCORDING TO THE SALES DATA, 2 PIECES OF THIS 50-PIECE LOT HAVE NOT BEEN SOLD. THIS IMPLIES THAT THE OTHER 47 PIECES HAVE BEEN SUCCESSFULLY IMPLANTED. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING RELATED. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS EQUINOXE REVERSE SHOULDER COMPRESSION SCREW, 750-2005-021-RMR-IMPLANT REV K, WAS REVIEWED. THE RISK IS CAPTURED IN 49. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. THE DAMAGED DEVICE REPORTED IN EXPERIENCE (B)(4) WAS LIKELY THE RESULT OF CROSS-THREADING THAT LED TO THE REPORTED THREAD DAMAGE. HOWEVER; THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES. THERE IS NO REPORTED PATIENT ADVERSE EVENT. THE SCREW APPEARED TO HAVE A WIRE LIKE PIECE OF METAL WRAPPING AROUND IT, NO PARTS OR PIECES FELL INTO THE PATIENT. POST-OP X-RAYS WERE "GOOD". AN INVESTIGATION WAS CONDUCTED: THE DAMAGED DEVICE REPORTED WAS LIKELY THE RESULT OF CROSS-THREADING THAT LED TO THE REPORTED THREAD DAMAGE. HOWEVER; THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10114 | 2022-05-24 | 2017-04-19 | I, n, j, u, r, y | Manufacturer report | EXACTECH | STEM PROTECTOR, SMALL | EXACTECH, INC. |  | 301-07-60 |  | IT WAS REPORTED THAT DURING AN ORTHOPEDIC KNEE "THE SMALL BLACK PLASTIC PART CAME OFF AFTER SEVERAL USES. THE ISSUE WAS RESOLVED, BUT NO DETAILS ARE AVAILABLE." THE SALES REPRESENTATIVE WAS NOT PRESENT WHEN THE DEVICE MALFUNCTION WAS DISCOVERED. THERE WAS NO REPORTED DELAY OF SURGERY AND THERE WAS NO REPORTED PATIENT INJURY, ADVERSE EVENT, OR CLINICAL CONSEQUENCE TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE REPORTER RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS NOT RETURNED TO EXACTECH FOR ANALYSIS. DESIGN-RELATED ISSUES: THE DESIGN OF THE SMALL STEM PROTECTOR HAS BEEN IN THE FIELD SINCE 2005. EXACTECH HAS RECEIVED 5 SIMILAR COMPLAINT REPORTS SINCE 2008. BECAUSE THE SMALL STEM PROTECTOR IS USED IN ANATOMIC AND REVERSE TOTAL SHOULDER SURGERIES, SALES DATA FOR HUMERAL STEMS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED '"VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE DETAILED BELOW. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH HAS NOT RECEIVED ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT, WHICH HAS BEEN IN THE FIELD SINCE 2014. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING-RELATED. RISK MANAGEMENT: A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. THE RMR FOR THIS SMALL STEM PROTECTOR, 750-2002-039-RMR REV F, WAS REVIEWED. THE RISK IS CAPTURED IN ROW 17. NO CORRECTIVE ACTION IS REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE BROKEN SMALL STEM PROTECTOR REPORTED IN EXPERIENCE C2017-219 WAS LIKELY THE RESULT OF DISASSEMBLING THE DEVICE FOR CLEANING AND NOT REASSEMBLING IT. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON A REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED INSTRUMENT MALFUNCTION IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO ANY PATIENT IMPACT. AN INVESTIGATION WAS CONDUCTED THAT THE MOST LIKELY ROOT CAUSE OF THE DEVICE BREAKAGE REPORTED IN THIS EVENT WAS LIKELY THE RESULT OF DISASSEMBLING THE DEVICE FOR CLEANING AND NOT REASSEMBLING IT CORRECTLY. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. ATTEMPTS HAVE BEEN MADE FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. NO ADDITIONAL ATTEMPTS WILL BE MADE FOR INFORMATION REGARDING THIS EVENT. INFORMATION WAS NOT PROVIDED FOR: AGE, SEX, WEIGHT, ETHNICITY, RACE. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10115 | 2022-05-24 | 2017-07-24 | I, n, j, u, r, y | Manufacturer report | EXACTECH | HUMERAL HEAD IMPACTOR TIP | EXACTECH, INC. |  | 311-07-07 |  | IT WAS REPORTED FROM THE US THAT DURING AN ORTHOPEDIC SURGICAL PROCEDURE WHILE THE SURGEON WAS IMPACTING THE HUMERAL HEAD TO THE STEM, THE HUMERAL HEAD IMPACTOR WAS USED, WITH THE FIRST STRIKE THE TIP SPLIT INTO TWO PIECES. IT WAS CONFIRMED BY THE REPRESENTATIVE THAT THERE WERE NO PIECES LEFT IN THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED FOR ANALYSIS. THE DEVICE WAS RETURNED IN 2 PIECES. THE DESIGN OF THE HUMERAL HEAD IMPACTOR TIP HAS BEEN IN THE FIELD SINCE 2004. THERE HAVE BEEN 4 SIMILAR COMPLAINT REPORTS INVOLVING THIS DEVICE SINCE 2008. THE FREQUENCY OF OCCURRENCE RANKING SCALE IS "VERY LOW". THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. THERE HAVE NOT BEEN ANY OTHER COMPLAINTS INVOLVING PARTS FROM THE MANUFACTURING LOT, WHICH HAS BEEN IN THE FIELD SINCE 2006. THEREFORE, THIS DOES NOT APPEAR TO BE A MANUFACTURING RELATED ISSUE. A REVIEW OF THE RISK MANAGEMENT REPORT (RMR) AND RISK ASSESSMENT AND CONTROLS REPORT (RACR) WAS CONDUCTED TO ENSURE THE RISK WAS INCLUDED AND THE OCCURRENCE IS BELOW THE THRESHOLD. MOST LIKELY CAUSE: THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF IMPACTING THE DEVICE AT AN ANGLE OR APPLYING A BENDING MOMENT DURING IMPACTION WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE INSTRUMENT. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT IMPACT. IT WAS CONFIRMED BY THE REPRESENTATIVE THAT THERE WERE NO PIECES LEFT IN THE PATIENT. AN INVESTIGATION WAS CONDUCTED; THE BROKEN INSTRUMENT REPORTED WAS LIKELY THE RESULT OF IMPACTING THE DEVICE AT AN ANGLE OR APPLYING A BENDING MOMENT DURING IMPACTION WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE INSTRUMENT. INFORMATION WAS NOT PROVIDED FOR: AGE, SEX, ETHNICITY, RACE. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10119 | 2022-05-24 | 2017-10-18 | I, n, j, u, r, y | Manufacturer report | EXACTECH | LPI IMPACTOR HANDLE | EXACTECH, INC. |  | 213-65-00 |  | IT WAS REPORTED FROM OUS THAT DURING AN ORTHOPEDIC THE PLASTIC HANDLE WAS "DESTROYED" WHEN A MALLET WAS USED TO IMPACT THE REAR SIDE. INSTRUMENT WAS REPLACED WITH A SECOND SET. THERE WAS A DELAY TO SURGERY WITH NO ADVERSE EVENT TO THE PATIENT. THERE WERE NO FRAGMENTS OF THE INSTRUMENT LEFT IN THE PATIENT. THE AGENT WAS PRESENT DURING THE TIME OF SURGERY. "THIS SEEMS TO BE CAUSED BY FREQUENT USE OR FREQUENT STERILIZATION (AVERAGE OF 27 OPERATIONS PER MONTH FOR 3 YEARS)." THE PATIENT WAS DISCHARGED WITHOUT ANY PROBLEMS. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN (B)(6) 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RETURNED FOR ANALYSIS. MFG DATE: 01/29/2013. AN INVESTIGATION OF COMPLAINTS RELATED TO LPI HANDLE - BROKEN NON-FUNCTIONAL, SPECIFICALLY RELATED TO CRACKS/BREAKAGE FOR THE PLASTIC HANDLE WAS CONDUCTED UNDER TM 2018-XX. THE LPI HANDLE CONSISTS OF A METAL STRIKE PLATE, METAL END PIECE WHICH ENGAGES WITH MATING INSTRUMENTS AND IMPLANTS, AND A PLASTIC CENTRAL HANDLE. THE METAL END PIECES ARE ASSEMBLED TO THE PLASTIC HANDLE VIA THERMAL ASSEMBLY WHICH INVOLVES HEATING THE METAL AND PRESSING THE HOT METAL PIECES INTO THE HANDLE. THE METAL PIECES HAVE CYLINDRICAL BOSSES WHICH FEATURE A PRESS-FIT WITH A FEMALE HOLE ON THE PLASTIC HANDLE. THE LPI HANDLE ASSEMBLES WITH TWO END PIECES FOR USE THREE TIMES DURING THE SURGICAL PROCEDURE: IMPACTION OF THE FEMORAL TRIAL, IMPACTION OF THE TIBIAL TRAY IMPLANT AND IMPACTION OF THE FEMORAL IMPLANT. AVAILABLE DEVICES FROM PREVIOUSLY REPORTED COMPLAINTS WERE EVALUATED. COMPLAINT DEVICES APPEARED TO HAVE SIMILAR FAILURE MODES, AND EITHER WERE CRACKED OR FRACTURED. CRACKS WERE CONSISTENT WITH STARTING INSIDE THE PLASTIC HANDLE AT THE BASE OF ONE OF THE TWO METAL ATTACHMENT FEATURES AND SPREADING OUTWARD. CRACKS WERE EITHER LINEAR, ALONG THE HANDLE AXIS OR HORIZONTAL, NEARLY IN LINE WITH THE ORIGIN. THE REVIEW OF THE LPI IMPACTOR HANDLES (213-65-00) CITED IN COMPLAINTS FOR BROKEN OR CRACKED HANDLES SHOWED A CONSISTENT FAILURE MODE. A REVIEW OF THE DEVICE ORDER HISTORY DID NOT SHOW A TREND IN THOSE LOTS WITH COMPLAINT DEVICES, NOR DID A REVIEW OF DESIGN CHANGES SHOW CHANGES MADE TO THE DESIGN WHICH COULD CONTRIBUTE TO THE COMPLAINT. PER 750-2009-023-RMR-INSTRUMENT, REV B - DEVICE FRACTURE/BREAKAGE (LINE ITEM 18) CAN LEAD TO A SURGICAL DELAY. THE PREDICTED SEVERITY OF THIS HARM IS MODERATE (3), DEFINED AS "ANY INSTRUMENT-RELATED HAZARD THAT COULD RESULT IN MARKED INTRA-OPERATIVE INCONVENIENCE TO THE SURGEON AND STAFF", AND THE PREDICTED OCCURRENCE IS LOW (2), DEFINED AS 1%. CRC-2018-03-02-02 WAS INITIATED ON 03/02/2018 TO EVALUATE THIS ISSUE. BASED ON THE AVAILABLE DATA, THE DEVICE IS PERFORMING WITHIN ITS PREDICTED RISK. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. IT WAS REPORTED THAT THE HANDLE BROKE AND ALL PIECES WERE RETRIEVED FROM THE PATIENT. AN INVESTIGATION WAS CONDUCTED UNDER CRC-2018-03-02-02; THE MOST LIKELY CAUSE OF THE REPORTED EVENTS IS RELATED TO HIGH IMPACTION FORCES REQUIRED DURING IMPLANT SURGERY. INFORMATION WAS NOT PROVIDED FOR: AGE, SEX, WEIGHT, ETHNICITY, RACE. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10136 | 2022-05-24 |  | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | DRILL GUIDE | EXACTECH, INC. |  | T6163 |  | IT WAS REPORTED FROM THE US THAT THE PINS COME OUT OF THE DRILL GUIDES. THE EVENT DID NOT OCCUR DURING SURGERY. THE ISSUE WAS NOTICED BY THE AGENT AS HE WAS PUTTING THE INSTRUMENTS BACK TOGETHER. PHOTO WAS PROVIDED. THE AGENT STATED: "FROM WHAT I CAN REMEMBER, THE PIN THAT ALLOWS THE BLUE ADJUSTABLE DRILL GUIDE TO ADJUST ANGLES CAME OUT. THIS CAUSED THE DRILL GUIDE TO BE UNABLE TO LOCK IN ON AN ANGLE. I NOTICED IT AS I WAS PUTTING INSTRUMENTS BACK IN TRAYS. TO MY KNOWLEDGE, IT MUST HAVE HAPPENED AS INSTRUMENTS WERE BEING CLEANED. THE GUIDE WAS NOT DROPPED, AND THE SURGEON DID NOT NOTIFY ME OF THE PROBLEM DURING SURGERY, AND HE WOULD HAVE IF IT WAS NOT WORKING PROPERLY." NO ADDITIONAL INFORMATION WAS PROVIDED. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH (B)(4) HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE(S) NOT AVAILABLE FOR EVALUATION. EVALUATION COMPLETED BY (B)(4) ON 12/11/2018. THE ENGINEERING MEMO STATES THAT THE DEVICE WAS SENT TO SERVICING. MOST LIKELY UNDERLYING CAUSE / ADDITIONAL COMMENTS: THE DEVICE WAS NOT AVAILABLE FOR EVALUATION AT THE TIME OF THIS INVESTIGATION. IF THE DEVICE IS RECEIVED IN THE FUTURE, THE FILE WILL BE REOPENED. A PHOTO OF THE DISASSEMBLED ADJUSTABLE DRILL GUIDE WAS PROVIDED. ACCORDING TO THE EXPERIENCE REPORT THE PIN DISENGAGED FROM THE DRILL GUIDE. THE EVENT DID NOT HAPPEN DURING SURGERY. THE SALES REPRESENTATIVE NOTICED THE ISSUE WHILE GATHERING INSTRUMENTS. DESIGN-RELATED ISSUES: THE ADJUSTABLE DRILL GUIDE HAS BEEN IN THE FIELD SINCE 2003. EXACTECH IS AWARE OF NO OTHER COMPLAINT REPORTS INVOLVING ADJUSTABLE DRILL GUIDES REPORTED AS DISASSEMBLED/MISASSEMBLED SINCE 2014. THE ADJUSTABLE DRILL GUIDE IS USED IN TOTAL HIP ARTHROPLASTY SURGERIES IN THE NOVATION SYSTEM, SO SALES DATA FOR NOVATION HIP ACETABULAR IMPLANTS WERE USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF (B)(4). THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: NOT APPLICABLE. NO INFORMATION WAS PROVIDED REGARDING THE SERIAL OR LOT INFORMATION SO MANUFACTURING INFORMATION COULD NOT BE REVIEWED. MOST LIKELY CAUSE: THE DISASSEMBLED ADJUSTABLE DRILL GUIDE REPORTED WAS LIKELY THE RESULT OF A STUCK LEVER WITHIN THE MECHANISM THAT ALLOWS THE ADJUSTABLE DRILL GUIDE TO CHANGE ANGLES, CAUSING THE PINS TO POP OUT. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE COMPONENT WAS NOT RETURNED FOR EVALUATION. IFU 700-096-181: INSTRUMENT INSPECTION: VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. DEVICE(S) USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES. NO PATIENT INVOLVEMENT. THE PINS CAME OUT OF THE DEVICE. AN INVESTIGATION WAS CONDUCTED: THE DISASSEMBLED ADJUSTABLE DRILL GUIDE REPORTED WAS LIKELY THE RESULT OF A STUCK LEVER WITHIN THE MECHANISM THAT ALLOWS THE ADJUSTABLE DRILL GUIDE TO CHANGE ANGLES, CAUSING THE PINS TO POP OUT. HOWEVER, THIS CANNOT BE CONFIRMED BECAUSE THE COMPONENT WAS NOT RETURNED FOR EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10144 | 2022-05-24 |  | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | SMALL STEM PROTECTOR | EXACTECH, INC. |  | 301-07-60 |  | IT WAS REPORTED FROM OUS THAT DURING AN INITIAL ORTHOPEDIC SURGERY THE SMALL STEM PROTECTOR BROKE. NO DELAY TO SURGERY AND NO ADVERSE EVENT TO THE PATIENT. THE AGENT WAS PRESENT AT TIME OF SURGERY. THE PATIENT'S HEALTH WAS STABLE LEAVING THE OR. NO PATIENT DEMOGRAPHIC INFORMATION WAS PROVIDED. NEW INFORMATION PROVIDED ON (B)(6)2021: "THE INSTRUMENT COULD HAVE BEEN BROKEN DURING THE SURGERY. HOWEVER, THE BREAKAGE HAS BEEN NOTICED AT THE RECOVERY OF THE INSTRUMENTATION-SET BY OUR LOGISTICS DEPT., AFTER THE CLEANING REALIZED BY THE STERILIZATION TEAM OF THE HOSPITAL. NO INFORMATION HAS BEEN RECEIVED FROM THE HOSPITAL REGARDING THIS ISSUE AND [THE AGENCY] DIDN'T RECEIVE ANY REQUEST FOR COMPLEMENTARY INFORMATION REGARDING THIS ISSUE AS WELL." MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DESIGN-RELATED ISSUES: THE DESIGN OF THE SMALL STEM PROTECTOR HAS BEEN IN THE FIELD SINCE 2004. EXACTECH IS AWARE OF 6 OTHER SIMILAR COMPLAINT REPORTS INVOLVING STEM PROTECTOR DEVICES SINCE 2008. BECAUSE THE SMALL STEM PROTECTOR IS USED IN ANATOMIC AND REVERSE TOTAL SHOULDER SURGERIES, SALES DATA FOR HUMERAL STEMS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THESE MANUFACTURING LOTS. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING-RELATED. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RISK MANAGEMENT REPORT. THE BROKEN INSTRUMENTS REPORTED WERE LIKELY THE RESULT OF IMPACTION OF THE CAP COMPONENT DURING SURGERY OR REMOVAL OF THE CAP COMPONENT DURING REPROCESSING AND NOT REPLACING IT. HOWEVER, FURTHER DETAILS REGARDING HOW THE DEVICES BECAME BROKEN WERE NOT PROVIDED. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. THIS IS MDR 3 OF 4 FROM THE SAME CUSTOMER REPORT NOT THE SAME PATIENT OR EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10117 | 2022-05-24 | 2017-08-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | TRULIANT TIBIAL INSERT TRIAL/TRAY ADAPTER | EXACTECH, INC. |  | 02-029-25-4025 |  | IT WAS DISCOVERED IN THE COMPANY RECEIVING AREA THAT THE TIBIAL INSERT TRIAL/TRAY ADAPTER WAS BROKEN. THE PLATES' CENTRAL FLEXURE DISASSOCIATED FROM THE PLATES. IT IS NOT CLEAR WHEN AND WHERE THE DISASSOCIATION TOOK PLACE. THERE IS NO ADDITIONAL INFORMATION AVAILABLE ABOUT THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI 1038671, HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE ROOT CAUSE ASSOCIATED WITH CAPA2017-53 RESIDES IN THE MISUSE/MISTREATMENT OF THE ADAPTOR PLATES OR CANNOT BE IDENTIFIED DUE TO INSUBSTANTIAL EVIDENCE. THE CAPA TEAM HAS DETERMINED, HOWEVER, THAT USE CONSIDERATIONS INDICATE THAT THE FLEXURE DESIGN IS SUSCEPTIBLE TO DAMAGE FROM SAID MISUSE/MISTREATMENT. DESIGN -A REVIEW OF THE TOLERANCE STACK-UP PERFORMED IN TM-2017-0417 WAS CONDUCTED. AS PART OF THIS INVESTIGATION, THE TOLERANCE STACK-UP BETWEEN THE TRULIANT TIBIAL TRIAL ADPTOR PLATES AND THE CORRESPONDING TRULIANT AND LOGIC TIBIAL TRAYS WAS REVISITED TO CONFIRM IT HAD BEEN ADEQUATELY PERFORMED. THE REVIEW WAS FOUND TO BE ACCURATE. THERE IS NO NEW INFORMATION THAT HAS BEEN FOUND THAT WOULD ALTER THE RESULTS OR METHODS OF SAID ANALYSIS. IN ADDITION, ATTEMPTS WERE MADE IN THE LABORATORY TO REPLICATE FAILURE MODES WITHOUT SUCCESS. THE DESIGN FUNCTIONS AS INTENDED UNDER WORST CASE CONDITIONS. MANUFACTURING - REVIEW OF THE DEVICE DHR CONCLUDED THAT ALL PARTS WERE MADE WITHIN SPECIFICATION. USE CONSIDERATIONS - THE DESIGN DOES LEND TO BEING SUSCEPTIBLE TO BREAKAGE IF USED INCORRECTLY DUE TO THE USE OF ULTEM MATERIAL WITH A THINNER ASPECT UNDER LOAD. IN ADDITION, THE DEVICES BROKE AT SOME POINT DURING REPROCESSING AT ONE PARTICULAR INSTITUTE. A REVIEW WAS CONDUCTED AT THIS INSTITUTE'S REPROCESSING PROCEDURES AND RESULTS INDICATED THAT CORRECT REPROCESSING INSTRUCTIONS PROVIDED BY EXACTECH WAS FOLLOWED. A REVIEW OF THE INSTRUMENT TRAY CASE INDICATED THAT IT IS POSSIBLE THAT THE ADAPTOR PLATES MAY NOT BE FULLY CONSTRAINED DURING REPROCESSING. DUE TO THE NATURE OF THE REPROCESSING PROCESS IT IS POSSIBLE THE PLATES BECAME DISLODGED AND DAMAGED ELSEWHERE IN THE MACHINE. HUMAN ERROR - HUMAN ERROR WAS CONSIDERED AS PART OF THE INVESTIGATION AND FOUND TO BE THE ROOT CAUSE OF ONE COMPLAINT. THE ADAPTOR PLATE THAT WAS USED IN SAID COMPLAINT WAS THE INCORRECT SIZE RELATIVE TO THE TIBIAL TRAY THAT HAD BEEN IMPLANTED. THIS MAY HAVE CAUSED THE PLATE TO DISLODGE UNPREDICTABLY. IN ADDITION, A HOHMANN RETRACTOR WAS USED TO REMOVE THE PLATE AND NOT THE TRULIANT DROP ROD HANDLE, AS STATED PER THE OPERATIVE TECHNIQUE, 712-35-31. THE RETRACTOR IS CURVED AND FEATURES A MUCH SMALLER NOSE WHICH COULD INADVERTENTLY FACILITATE CONTACT WITH THE FLEXURE DURING REMOVAL. THE GEOMETRY OF THE TRULIANT DROP ROD HANDLE, IN ADDITION TO AN APPROPRIATELY CONSTRAINED ADAPTOR PLATE, AS DESIGNED CREATES A SCENARIO IN WHICH CONTACT WITH THE FLEXURE IS EXTREMELY DIFFICULT. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10120 | 2022-05-24 | 2017-05-29 | I, n, j, u, r, y | Manufacturer report | EXACTECH | LOGIC PS TIBIA, MODULAR INSERT TRIAL, SIZE 2, 11MM | EXACTECH, INC. |  | 02-013-35-2011 |  | IT WAS REPORTED FROM ITALY THAT DURING AN ORTHOPEDIC SURGERY, THE INSERT TRAIL BROKE. NO PIECES WERE LEFT IN THE PATIENT, THERE IS NO REPORT OF HARM OR ADVERSE EVENT TO THE PATIENT. THE PATIENT WAS STABLE LEAVING THE OR. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. CAPA (B)(4) INVESTIGATION STATES: THE CAUSE OF THIS FAILURE COULD BE REPROCESSING OF THESE PARTS. TO UNDERSTAND WHY THESE DEVICES ARE FAILING IN SINGAPORE AND KOREA, A QUESTIONNAIRE WAS SENT TO THE SALES FORCE IN KOREA AND SINGAPORE TO INQUIRE ABOUT TKA TRIALING, GENERAL USE, AND STERILIZATION/REPROCESSING TECHNIQUES. ACCORDING TO AN EMAIL FROM THE SALES MANAGER AT EXACTECH KOREA, THESE INSTRUMENTS ARE BEING STERILIZED FOR 40-50 MINUTES AT 134 DEGREES (CELSIUS). THESE REPROCESSING METHODS ARE OUTSIDE OF EXACTECH'S RECOMMENDED STERILIZATION INSTRUCTIONS PER (B)(4) - REPROCESSING INSTRUCTIONS FOR REUSABLE SURGICAL INSTRUMENTS, WHICH STATES PARAMETERS FOR STERILIZATION. (B)(4) IS THE INITIAL SCOPE ASSESSMENT. ROOT CAUSE: THE REPROCESSING OF THESE INSTRUMENTS EXCEEDS THE RECOMMENDED PARAMETERS LISTED IN THE REPROCESSING INSTRUCTIONS. THE MATERIAL DEGRADATION FROM EXCESSIVE STERILIZATION OF THESE INSTRUMENTS HAS LIKELY LED TO THE OBSERVED FAILURE. NO ADDITIONAL ACTION REQUIRED. IFU (B)(4): INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: (1) APPROPRIATE READING OF THE LITERATURE, AND (2) TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND (3) REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI- OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT IMPACT. NO PIECES WERE LEFT IN THE PATIENT. IT WAS REPORTED THAT NO PIECES WERE LEFT IN THE PATIENT. AN INVESTIGATION WAS CONDUCTED UNDER CAPA (B)(4); THE ROOT CAUSE, THE REPROCESSING OF THESE INSTRUMENTS EXCEEDS THE RECOMMENDED PARAMETERS LISTED IN THE REPROCESSING INSTRUCTIONS. THE MATERIAL DEGRADATION FROM EXCESSIVE STERILIZATION OF THESE INSTRUMENTS HAS LIKELY LED TO THE OBSERVED FAILURE. NO ADDITIONAL ACTION REQUIRED. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10125 | 2022-05-24 | 2017-11-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | LOGIC PS TIBIA, MODULAR INSERT TRIAL, SIZE 1.5, 13MM | EXACTECH, INC. |  | 02-013-35-1513 |  | IT WAS REPORTED FROM OUS THAT THE INSERT TRIAL WAS DAMAGED DURING SURGERY. THE TRIAL WAS BROKEN, NO PIECES FELL INTO THE SURGICAL SITE. PATIENT'S HEALTH WAS STABLE LEAVING THE OR. THE SURGERY PROCEEDED WITHOUT USE OF THE SPECIAL INSTRUMENT. THE AGENT WAS PRESENT AT SURGERY. THERE WAS NO DELAY TO THE SURGERY AND NO CLINICAL CONSEQUENCE TO THE PATIENT. MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. THE DEVICE WAS RECEIVED FOR ANALYSIS. COMPLETED ON 31 JUL 2018 BY MB. VISUAL EVALUATION CONDITION/FINDINGS: (CONDUCTED WITH A 7X OR LOX OPTICAL UNLESS OTHERWISE SPECIFIED) THE FRACTURE SURFACE APPEARS CONSISTENT WITH CRACK PROPAGATION, AND ULTIMATE FRACTURE. ALL OTHER ASPECTS OF THE DEVICE APPEAR UNREMARKABLE AND CONSISTENT WITH YEARS OF USE. DESIGN-RELATED ISSUES: THE DESIGN OF THE LOGIC TIBIAL TRIAL HAS BEEN IN THE FIELD SINCE 2013. EXACTECH HAS RECEIVED 37 SIMILAR COMPLAINT REPORTS INVOLVING 48 TIBIAL TRIALS SINCE 2013. THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH HAS NOT RECEIVED ANY OTHER COMPLAINTS INVOLVING PARTS FROM THIS MANUFACTURING LOT, WHICH HAS BEEN IN THE FIELD SINCE 2014. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING-RELATED. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RMR. MOST LIKELY CAUSE: THE BROKEN DEVICE REPORTED IN EXPERIENCE C2017-643 WAS LIKELY THE RESULT OF IMPACTING OR DROPPING THE DEVICE ONTO A HARD SURFACE, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE TRIAL. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES AND/OR DESIGN ISSUES. THERE WAS NO PATIENT ADVERSE EVENT, THE DEVICE PIECES DID NOT FALL INTO THE SURGICAL SITE. AN INVESTIGATION WAS CONDUCTED; THE BROKEN DEVICE REPORTED WAS LIKELY THE RESULT OF IMPACTING OR RM NUMBER (B)(4) REV: - EFFECTIVE DATE: 10/15/18. CONTROLLED DOCUMENT COMPANY CONFIDENTIAL PAGE 3 OF 3 DROPPING THE DEVICE ONTO A HARD SURFACE, WHICH LED TO CRACK INITIATION, PROPAGATION, AND ULTIMATE FRACTURE OF THE TRIAL. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-10143 | 2022-05-24 |  | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EXACTECH | SMALL STEM PROTECTOR | EXACTECH, INC. |  | 301-07-60 |  | IT WAS REPORTED FROM OUS THAT DURING AN INITIAL ORTHOPEDIC SURGERY THE SMALL STEM PROTECTOR BROKE. NO DELAY TO SURGERY AND NO ADVERSE EVENT TO THE PATIENT. THE AGENT WAS PRESENT AT TIME OF SURGERY. THE PATIENT'S HEALTH WAS STABLE LEAVING THE OR. NO PATIENT DEMOGRAPHIC INFORMATION WAS PROVIDED. NEW INFORMATION PROVIDED ON 25 MAR 2021: "THE INSTRUMENT COULD HAVE BEEN BROKEN DURING THE SURGERY. HOWEVER, THE BREAKAGE HAS BEEN NOTICED AT THE RECOVERY OF THE INSTRUMENTATION-SET BY OUR LOGISTICS DEPT., AFTER THE CLEANING REALIZED BY THE STERILIZATION TEAM OF THE HOSPITAL. NO INFORMATION HAS BEEN RECEIVED FROM THE HOSPITAL REGARDING THIS ISSUE AND [THE AGENCY] DIDN'T RECEIVE ANY REQUEST FOR COMPLEMENTARY INFORMATION REGARDING THIS ISSUE AS WELL." MULTIPLE EMAIL REQUESTS WERE SENT TO THE CONTACTS FOR ADDITIONAL INFORMATION. NO FURTHER INFORMATION HAS BEEN PROVIDED BY THE CONTACTS RELATED TO THIS EVENT. | AS A RESULT OF AN FDA INSPECTION CONDUCTED IN JAN 2020, EXACTECH, FEI (B)(4), HAS COMMITTED TO REMEDIATING 3 YEARS OF COMPLAINTS (2017-2020). THIS MDR IS BEING SUBMITTED AS PART OF THAT REMEDIATION. DESIGN-RELATED ISSUES: THE DESIGN OF THE SMALL STEM PROTECTOR HAS BEEN IN THE FIELD SINCE 2004. EXACTECH IS AWARE OF 6 OTHER SIMILAR COMPLAINT REPORTS INVOLVING STEM PROTECTOR DEVICES SINCE 2008. BECAUSE THE SMALL STEM PROTECTOR IS USED IN ANATOMIC AND REVERSE TOTAL SHOULDER SURGERIES, SALES DATA FOR HUMERAL STEMS WAS USED TO CALCULATE AN APPROXIMATE COMPLAINT OCCURRENCE RATE OF <0.5%. THIS IS CONSIDERED "VERY LOW" ACCORDING TO THE FREQUENCY OF OCCURRENCE RANKING SCALE. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE DESIGN RELATED. MFG-RELATED ISSUES: EXACTECH IS NOT AWARE OF ANY OTHER COMPLAINTS INVOLVING PARTS FROM THESE MANUFACTURING LOTS. THEREFORE, THIS ISSUE DOES NOT APPEAR TO BE MANUFACTURING-RELATED. CORRECTIVE ACTION IS NOT REQUIRED BECAUSE THE OCCURRENCE RATE IS "VERY LOW", AND THE RISK IS CAPTURED IN THE RISK MANAGEMENT REPORT. THE BROKEN INSTRUMENTS REPORTED WERE LIKELY THE RESULT OF IMPACTION OF THE CAP COMPONENT DURING SURGERY OR REMOVAL OF THE CAP COMPONENT DURING REPROCESSING AND NOT REPLACING IT. HOWEVER, FURTHER DETAILS REGARDING HOW THE DEVICES BECAME BROKEN WERE NOT PROVIDED. IFU 700-096-181: INSTRUMENT INSPECTION · VISUALLY INSPECT THE INSTRUMENTS FOR DAMAGE SUCH AS FRACTURES; CRACKS; GOUGES; DEFORMATION; BURRS; DISCOLORATION, CORROSION, OR RUST; EXCESSIVE COMPONENT WEAR; NICKS ON CUTTING SURFACES, MISSING OR LOOSE COMPONENTS; BLOCKAGES IN CANNULAE, CLEANING HOLES OR OTHER CAVITIES THAT CANNOT BE REMOVED VIA STANDARD CLEANING; WORN OR DIFFICULT TO READ MARKINGS/ENGRAVINGS; OR OTHER APPARENT DAMAGE. · CHECK THE FUNCTION OF MECHANISMS BY ACTUATING ANY LEVERS, KNOBS, SWITCHES, CONNECTORS, SLIDING FEATURES, HINGES, OR OTHER MECHANICAL INTERFACE FEATURES. ENSURE SMOOTH OPERATION OF THESE FEATURES OVER THEIR FUNCTIONAL RANGE OF MOTION. · IF DAMAGE, WEAR, OR NON-FUNCTIONING/POORLY FUNCTIONING MECHANISMS ARE FOUND, DO NOT USE THE INSTRUMENT, AND CONTACT THE SALES REPRESENTATIVE OR CUSTOMER SERVICE FOR DISPOSITION. IFU STATES: THE SURGEON SHALL BECOME THOROUGHLY FAMILIAR WITH THE TECHNIQUE OF IMPLANTATION OF THE PROSTHESES BY: APPROPRIATE READING OF THE LITERATURE, AND TRAINING IN THE OPERATIVE SKILLS AND TECHNIQUES REQUIRED FOR SURGERY, AND REVIEWING INFORMATION REGARDING USE OF INSTRUMENTATION. IT IS A CLINICAL STANDARD OF PRACTICE IN THE OPERATING ROOM THAT ALL INSTRUMENTS SHOULD BE VISUALLY AND FUNCTIONALLY INSPECTED BEFORE USE, THESE GUIDELINES ARE FROM STANDARD ASSOCIATION OF PERI-OPERATIVE REGISTERED NURSES (AORN) GUIDELINES. SURGEONS ARE TO BE FAMILIAR AND KNOWLEDGEABLE WITH ALL INSTRUMENTATION, DEVICES AND PROFICIENT WITH SURGICAL TECHNIQUES. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST THE REPORTED EVENT IS RELATED TO ANY MANUFACTURING ISSUES OR DESIGN ISSUES, NOR DID IT LEAD TO PATIENT ADVERSE EVENT. THIS IS MDR 2 OF 4 FROM THE SAME CUSTOMER REPORT NOT THE SAME PATIENT OR EVENT. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00607 | 2022-05-25 | 2022-05-03 | I, n, j, u, r, y | Manufacturer report | NV GXL LNR, NEUTRAL, 32MM ID, GROUP 3 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LNR, NEUTRAL, 32MM ID, GROUP 3 CUPS | 130-32-53 |  | AS REPORTED, THE (B)(6) PATIENT REQUIRES A RIGHT HIP GXL REVISION DUE TO POLYWEAR. REVISION HAS NOT BEEN SCHEDULED AS OF YET. NO ADDITIONAL INFORMATION AVAILABLE. | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION RECEIVED THE FOLLOWING SECTIONS B5, G3, G6, H2, H3, H6, H7 AND H9 HAVE BEEN UPDATED ACCORDINGLY. SECTION B5: ADDITIONAL INFORMATION RECEIVED INDICATES THAT THE PATIENT REFUSED YEARLY CONTROL. NO ADDITIONAL INFORMATION FORTHCOMING. (H3) THE MOST LIKELY CAUSE FOR THE MONITORING REPORTED DUE TO EARLY PROSTHESIS WEAR MAY HAVE BEEN DUE TO A COMBINATION OF THE RISK FACTORS INCLUDING BUT NOT LIMITED TO USE ERROR, IMPLANT POSITIONING, AND PATIENT FACTORS. HOWEVER, THIS CANNOT BE CONFIRMED FROM THE REPORTED INFORMATION AND THE DEVICES WERE NOT AVAILABLE FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00606 | 2022-05-25 | 2022-05-03 | I, n, j, u, r, y | Manufacturer report | NV GXL LINR, NTRL, 32MM ID, GROUP 2 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LINR, NTRL, 32MM ID, GROUP 2 CUPS | 130-32-52 |  | AS REPORTED, THE (B)(6) MALE PATIENT REQUIRES A LEFT HIP GXL REVISION DUE TO POLYWEAR. REVISION HAS NOT BEEN SCHEDULED AS OF YET. NO ADDITIONAL INFORMATION AVAILABLE. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00604 | 2022-05-25 | 2022-05-05 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | MDLR RV RMR 38MM PLT TIP | EXACTECH, INC. | 321-25-38 | 321-25-38 |  | IT WAS REPORTED THAT DURING SURGICAL USE, THE SURGEON WAS REAMING THE GLENOID AND SUDDENLY THE REAMER BEGAN TO SLIP. UPON INVESTIGATION THE TIP WAS BROKEN. THE SURGEON REMOVED THE TIP AND CONTINUED TO REAM WITH THE 42MM REAMER. PART OF THE DEVICE WAS REMOVED FROM THE (B)(6) MALE PATIENT. THERE WAS NO DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICE WILL BE RETURNING. | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00602 | 2022-05-25 | 2021-08-18 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 5 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 5 9MM | 02-012-35-5009 |  | AS REPORTED, APPROXIMATELY FIVE YEARS POST LTKA, THE (B)(6) MALE PATIENT VISITED THE SURGEON FOR REASON NOT REPORTED. | SECTION H10: (H3) THE POTENTIAL PROSTHESIS WEAR REPORTED IN (B)(4) CANNOT BE CONFIRMED FROM THE PROVIDED INFORMATION AND THE DEVICES ARE NOT AVAILABLE FOR EVALUATION AS THEY REMAIN IMPLANTED. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00605 | 2022-05-25 | 2022-05-05 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | HUM LNR IMPACT TIP 38M | EXACTECH, INC. | 321-07-38 | 321-07-38 |  | IT WAS REPORTED THAT THE 38MM HUMERAL LINER IMPACTOR TIP SPLIT DURING IMPACTION OF THE LINER ONTO THE 0MM HUMERAL ADAPTER TRAY. NO DAMAGE TO 38MM +0MM HUMERAL LINER WAS NOTED AND IMPLANTATION WAS NOT AFFECTED. NO PARTS OR PIECES FELL INTO THE FEMALE PATIENT'S WOUND SITE. THERE WAS A SMALL SURGICAL DELAY, LESS THAN 5 MINUTES. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL BE RETURNING. | PENDING EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00599 | 2022-05-25 | 2022-05-04 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | HUMERAL STEM PRIMARY, PRESS FIT 7MM | EXACTECH, INC. | 300-01-07 | 300-01-07 |  | AS REPORTED BY THE EQUINOXE SHOULDER STUDY, 2 DAYS POST OP THE INITIAL TSA, THIS (B)(6) MALE PATIENT WAS REVISED DUE TO AN UNDIAGNOSED INTRA-OPERATIVE MIDSHAFT HUMERAL FRACTURE CAUSING LOOSENING OF THE HUMERAL STEM AND DISLOCATION. THE CASE REPORT FORM INDICATES THIS EVENT IS NOT RELATED TO DEVICES OR PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | SECTION H10: (H3) BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN OR MANUFACTURING ISSUES. THE CAUSE OF THE SURGICAL EVENT CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION OF MULTIPLE SURGERIES AND COMORBIDITIES. THESE DEVICES ARE USED FOR TREATMENT NOT DIAGNOSIS. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00603 | 2022-05-25 | 2022-04-13 | I, n, j, u, r, y | Manufacturer report | SMALL SUPERIOR/POSTERIOR AUG GLENOID PLATE,LEFT | PROSTHESIS, SHOULDER, NON-CONSTRAINED, METAL/POLYMER CEMENTED | EXACTECH, INC. | 320-35-07 | 320-35-07 |  | AS REPORTED, APPROXIMATELY 2 YEARS POST OP THE TSA, THIS (B)(6) FEMALE PATIENT WAS REVISED TO A STANDARD REVERSE DUE TO ASEPTIC GLENOID LOOSENING. BASEPLATE LOOSENED WHICH WAS FOLLOWED BY STEM LOOSENING. THE CASE REPORT FORM INDICATES THIS EVENT IS POSSIBLY RELATED TO DEVICES AND/OR PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | SECTION H10: (H3) THE REVISION REPORTED MAY HAVE BEEN THE RESULT OF AN INSUFFICIENT BOND BETWEEN THE IMPLANTS AND THE BONE, PATIENT-RELATED CONDITIONS, OR A COMBINATION OF THESE POSSIBILITIES, WHICH LED TO ASEPTIC (NON-INFECTED) GLENOID AND HUMERAL LOOSENING. HOWEVER, THIS CANNOT BE CONFIRMED AS THE DEVICES WERE NOT RETURNED FOR EVALUATION AND PRE-REVISION X-RAYS WERE UNABLE TO BE OBTAINED. SECTION H11: *THE FOLLOWING SECTIONS HAVE CORRECTED INFORMATION: (D1) BRAND NAME: SMALL SUPERIOR/POSTERIOR AUG GLENOID PLATE,LEFT (D2B) COMMON DEVICE NAME: PROSTHESIS, SHOULDER, NON-CONSTRAINED, METAL/POLYMER CEMENTED (D4) CATALOG NUMBER: 320-35-07, SERIAL NUMBER: (B)(6), EXPIRATION DATE: 29-SEP-2029, UNIQUE IDENTIFIER (UDI) #: 10885862534910 (D10) CONCOMITANT DEVICE(S): 320-20-18, (B)(6) - EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 18MM 320-20-30, (B)(6) - EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 30MM 320-20-34, (B)(6)- EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 34MM 320-20-34, (B)(6)- EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 34MM. (H4) DEVICE MANUFACTURE DATE: 01-OCT-2019 (H6) COMPONENT CODE: 4755, PART/COMPONENT/SUB-ASSEMBLY TERM NOT APPLICABLE. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00614 | 2022-05-26 | 2022-04-22 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 5 15MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 5 15MM | 02-012-35-5015 |  | AS REPORTED BY THE EXACTECH KNEE REPLACEMENT STUDY, APPROXIMATELY TEN YEARS POST INITIAL TKA, THE (B)(6) MALE PATIENT HAS VERY MINIMAL BONE LYSIS UNDER THE MEDIAL TIBIAL BASE PLATE. EVENT IS ONGOING. PER CLINICAL STUDY CASE REPORT FORM INDICATES THIS EVENT IS POSSIBLY RELATED TO DEVICES AND/OR PROCEDURE. | CONCOMITANT MEDICAL PRODUCTS: LOGIC FEMORAL PS CEM RIGHT SZ 5 (CAT# 02-010-01-0350). LGC TIBIAL FIT TRAY CEM SZ 5F / 5T (CAT# 02-012-45-5050). THREE PEG PATELLA 38MM (CAT# 200-02-38). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00617 | 2022-05-26 | 2021-07-01 | I, n, j, u, r, y | Manufacturer report | TRULIANT TIB IMP CR INS STD SZ 3, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT TIB IMP CR INS STD SZ 3, 9MM | 02-022-47-3009 |  | AS REPORTED BY THE EXACTECH KNEE REPLACEMENT STUDY, APPROXIMATELY FIVE MONTHS POST INITIAL TKA, THE 54 Y/O FEMALE PATIENT COMPLAINT OF STIFFNESS TO RIGHT KNEE. NO IMPROVEMENT WITH PHYSICAL THERAPY (PT). THE EVENT IS STILL ONGOING. PER CLINICAL STUDY CASE REPORT FORM INDICATES THIS EVENT IS UNLIKELY RELATED TO DEVICES AND POSSIBLY RELATED TO PROCEDURE. | (H3) AS REPORTED BY THE EXACTECH KNEE REPLACEMENT STUDY, APPROXIMATELY FIVE MONTHS POST INITIAL TKA FOR OSTEOARTHRITIS, THE 54 Y/O OBESE FEMALE PATIENT COMPLAINT OF STIFFNESS TO RIGHT KNEE. NO IMPROVEMENT WITH PHYSICAL THERAPY (PT). OTHER ACTION TAKEN: PRE-OP COUNSELING FOR LYSIS OF ADHESIONS. RISKS AND BENEFITS GIVEN. THE EVENT IS STILL ONGOING. PER CLINICAL STUDY CASE REPORT FORM INDICATES THIS EVENT IS UNLIKELY RELATED TO DEVICES AND POSSIBLY RELATED TO PROCEDURE. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN OR MANUFACTURING ISSUES. THE CAUSE OF THE SUBSEQUENT PAIN/STIFFNESS CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION. THESE DEVICES ARE USED FOR TREATMENT NOT DIAGNOSIS. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00616 | 2022-05-26 | 2022-05-02 | I, n, j, u, r, y | Manufacturer report | ALT HA S CLR EXT SZ 7 | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | ALT HA S CLR EXT SZ 7 | 190-31-07 |  | AS REPORTED, THE (B)(6) MALE PATIENT HAD A REVISION DUE TO STEM SUBSIDENCE. PATIENT WAS REVISED WITH A ZB STEM, CUP WAS RETAINED AND A NEW LINER WAS IMPLANTED. THERE WAS NO BREAKAGE OF A DEVICE. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PHOTOS AND X-RAYS RECEIVED. PRODUCT WILL NOT BE RETURNED. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00615 | 2022-05-26 | 2022-05-05 | I, n, j, u, r, y | Manufacturer report | TRULIANT TIB IMP PS INSERT SZ 2 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT TIB IMP PS INSERT SZ 2 9MM | 02-022-35-2009 |  | AS REPORTED, SURGEON WENT TO RESURFACE PATELLA AND JUST DECIDED TO EXCHANGE POLY WHILE HE WAS DOING THE PROCEDURE. PATIENT WAS REVISED TO A NEW POLY. THERE WAS NO BREAKAGE IN THE DEVICE AND NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. SALES REP WAS UNABLE TO OBTAIN PHOTOS OR X-RAYS. | AS REPORTED, SURGEON WENT TO RESURFACE PATELLA AND JUST DECIDED TO EXCHANGE POLY WHILE HE WAS DOING THE PROCEDURE. PATIENT WAS REVISED TO A NEW POLY. THERE WAS NO BREAKAGE IN THE DEVICE AND NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. SALES REP WAS UNABLE TO OBTAIN PHOTOS OR X-RAYS. UPON REVIEW, THERE IS NO ALLEGATION AGAINST THE DEVICE AS THE REASON FOR REVISION WAS NOT REPORTED. THE MOST LIKELY CAUSE OF THE REPORTED EVENT IS PATIENT CONDITIONS. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00613 | 2022-05-26 | 2018-01-01 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 4 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 4 11MM | 02-012-35-4011 |  | AS REPORTED BY THE EXACTECH KNEE REPLACEMENT STUDY, APPROXIMATELY SIX YEARS POST INITIAL TKA, THE 67 Y/O MALE PATIENT EXPERIENCED SWOLLEN KNEE (LEFT) DUE TO POLYETHYLENE WEAR WITH EVIDENCE OF OSTEOLYSIS SEEN ON X-RAY. THE EVENT IS STILL ONGOING. PER CLINICAL STUDY CASE REPORT FORM INDICATES THIS EVENT IS POSSIBLY RELATED TO DEVICES AND DEFINITELY RELATED TO PROCEDURE. | CONCOMITANT MEDICAL PRODUCTS: LOGIC FEMORAL PS CEM LEFT SZ 4 (CAT# 02-010-01-0240). LGC TIBIAL FIT TRAY CEM SZ 4F / 4T (CAT# 02-012-45-4040). THREE PEG PATELLA 35MM (CAT# 200-02-35). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00609 | 2022-05-26 | 2022-05-03 | I, n, j, u, r, y | Manufacturer report | NV GXL LNR, +5 LAT 32MM ID, GROUP 3 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LNR, +5 LAT 32MM ID, GROUP 3 CUPS | 136-32-53 |  | AS REPORTED, A (B)(6) MALE PATIENT HAD A SCHEDULE REVISION TO RIGHT HIP ON (B)(6) 2022 DUE TO GXL POLY WEAR. NO OTHER PATIENT INFORMATION AVAILABLE. | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION RECEIVED THE FOLLOWING SECTIONS D6B, G3, G6, H2, H3, H6, H7 AND H9 HAVE BEEN UPDATED ACCORDINGLY. (H3) BASED ON THE AVAILABLE INFORMATION, THE PATIENT INVOLVED MEETS THE FOLLOWING RISK CRITERIA FOR EARLY PROSTHESIS WEAR IMPLANTED WITH A LATERALIZED LINER. THE MOST LIKELY CAUSE FOR THE REVISION REPORTED DUE TO EARLY PROSTHESIS WEAR IS A COMBINATION OF THE RISK FACTORS INCLUDING, USE ERROR, IMPLANT POSITIONING, IMPLANT SIZE SELECTION, AND PATIENT FACTORS (FITNESS FOR SURGERY, BIOMECHANICS, ACTIVITY LEVEL AND LOCAL TISSUE OXIDATION POTENTIAL COULD HAVE ALL CONTRIBUTED TO THE INCREASED TREND IN WEAR/OSTEOLYSIS RELATED COMPLAINTS). HOWEVER, THIS CANNOT BE CONFIRMED FROM THE REPORTED INFORMATION AND THE DEVICES WERE NOT AVAILABLE FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00611 | 2022-05-26 | 2022-05-03 | I, n, j, u, r, y | Manufacturer report | NV GXL LNR, +5LAT, 32MM G1-48/50MM CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LNR, +5LAT, 32MM G1-48/50MM CUPS | 136-32-51 |  | AS REPORTED, THE (B)(6) FEMALE PATIENT HAD A REVISION DUE TO POLY WEAR. PATIENT REVISED TO A NV EHXL NTRL LNR G1 32MM. NO OTHER PATIENT INFORMATION AVAILABLE. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00610 | 2022-05-26 | 2021-01-05 | I, n, j, u, r, y | Manufacturer report | TRULIANT TIB IMP CR STD, SIZE 3, 10MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT TIB IMP CR STD, SIZE 3, 10MM | 02-022-47-3010 |  | AS REPORTED BY THE EXACTECH KNEE REPLACEMENT STUDY, APPROXIMATELY ONE YEAR POST INITIAL TKA, THE (B)(6) MALE PATIENT COMPLAINS OF DISCOMFORT AND SOME LOSS OF MOTION. THE EVENT IS STILL ONGOING. PER CLINICAL STUDY CASE REPORT FORM INDICATES THIS EVENT IS UNLIKELY RELATED TO DEVICES AND POSSIBLY RELATED TO PROCEDURE. | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION RECEIVED THE FOLLOWING SECTIONS G3, G6, H2, H3 AND H6 HAVE BEEN UPDATED ACCORDINGLY. (H3) AS REPORTED BY THE EXACTECH KNEE REPLACEMENT STUDY, APPROXIMATELY ONE YEAR POST INITIAL TKA FOR OSTEOARTHRITIS, THE 58 Y/O OBESE FEMALE PATIENT COMPLAINS OF DISCOMFORT AND SOME LOSS OF MOTION DUE TO ARTHROFIBROSIS. ACTION TAKEN: PATIENT IS SCHEDULED FOR PRE-OP VISIT. THE EVENT IS STILL ONGOING. PER CLINICAL STUDY CASE REPORT FORM INDICATES THIS EVENT IS UNLIKELY RELATED TO DEVICES AND POSSIBLY RELATED TO PROCEDURE. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN OR MANUFACTURING ISSUES. THE CAUSE OF THE SUBSEQUENT KNEE DISCOMFORT/PAIN CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION. THESE DEVICES ARE USED FOR TREATMENT NOT DIAGNOSIS. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00608 | 2022-05-26 | 2022-05-03 | I, n, j, u, r, y | Manufacturer report | NV GXL LINER NEUTRAL, 32MM ID GROUP 1 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LINER NEUTRAL, 32MM ID GROUP 1 CUPS | 130-32-51 |  | AS REPORTED, THE (B)(6) MALE PATIENT HAD A SCHEDULE REVISION ON (B)(6) 2022 DUE TO GXL POLY WEAR. NO OTHER PATIENT INFORMATION AVAILABLE. | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION RECEIVED THE FOLLOWING SECTIONS G3, G6, H2, H3, H6, H7 AND H9 HAVE BEEN UPDATED ACCORDINGLY. (H3) BASED ON THE AVAILABLE INFORMATION, THE PATIENT INVOLVED MEETS THE FOLLOWING RISK CRITERIA FOR EARLY PROSTHESIS WEAR COMBINATION OF LARGEST AVAILABLE FEMORAL HEAD AND/OR THINNEST AVAILABLE ACETABULAR LINER WAS USED. THE MOST LIKELY CAUSE FOR THE PLANNED REVISION REPORTED DUE TO EARLY PROSTHESIS WEAR IS A COMBINATION OF THE RISK FACTORS INCLUDING, USE ERROR, IMPLANT POSITIONING, IMPLANT SIZE SELECTION, AND PATIENT FACTORS MAY HAVE ALSO BEEN A CONTRIBUTING FACTOR TO THE EARLY PROSTHESIS WEAR. HOWEVER, THIS CANNOT BE CONFIRMED FROM THE REPORTED INFORMATION AND THE DEVICES WERE NOT AVAILABLE FOR EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00612 | 2022-05-26 | 2022-03-21 | I, n, j, u, r, y | Manufacturer report | NV GXL LINER NEUTRAL, 32MM ID GROUP 1 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LINER NEUTRAL, 32MM ID GROUP 1 CUPS | 130-32-51 |  | AS REPORTED, THE 59 Y/O FEMALE PATIENT HAD A GXL REVISION DUE TO POLY WEAR. PATIENT REVISED TO A NV EHXL NTRL LNR G1 32MM. NO OTHER PATIENT INFORMATION AVAILABLE. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00618 | 2022-05-27 | 2022-04-19 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | 8.5MM PLATFORM FX STEM LEFT | EXACTECH, INC. | 304-21-09 | 304-21-09 |  | AS REPORTED, APPROXIMATELY 1 YEAR POST OP THE INITIAL LEFT TSA, THIS 67 Y/O MALE PATIENT WAS REVISED FOR LOOSENING OF THE STEM DUE TO INFECTION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICE NOT RETURNED, DISPOSED OF BY HOSPITAL. NO ADDITIONAL INFORMATION. | SECTION H10: (H3) BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN OR MANUFACTURING ISSUES. THE CAUSE OF SURGICAL REVISION FOR LOOSENING DUE TO INFECTION CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION. THESE DEVICES ARE USED FOR TREATMENT NOT DIAGNOSIS. SECTION H11: *THE FOLLOWING SECTIONS HAVE CORRECTED INFORMATION: (B5) DESCRIBE EVENT OR PROBLEM: AS REPORTED, APPROXIMATELY 1 YEAR POST OP THE INITIAL LEFT TSA, THIS 67 Y/O MALE PATIENT WAS REVISED FOR LOOSENING OF THE STEM DUE TO INFECTION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICE NOT RETURNED, DISPOSED OF BY HOSPITAL. NO ADDITIONAL INFORMATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00620 | 2022-05-27 | 2022-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EQUINOXE | MDLR CNLTD TRIDRV DRVR | EXACTECH, INC. | 315-25-00 | 315-25-00 |  | IT WAS REPORTED THAT THE INSTRUMENT CAME PART DURING REAMING. NO PARTS OR PIECES FELL INTO THE (B)(6) MALE PATIENT WOUND SITE, ALL PIECES ACCOUNTED FOR. PATIENT OUTCOME UNAFFECTED. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. | PENDING EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00621 | 2022-05-27 | 2022-05-02 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | LINER | EXACTECH, INC. |  | UNK |  | AS REPORTED, APPROXIMATELY 8 MONTHS POST OP THE INITIAL LEFT TSA, THIS (B)(6) MALE PATIENT WAS REVISED DUE TO INFECTION. SURGEON USED 2 SCREWS LESS THAN WERE ORIGINALLY USED- SO 6 REMOVED AND 4 NEW REIMPLANTED. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICE WILL NOT BE RETUNING DUE TO HOSPITAL POLICY. | SECTION H10: (H3) BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN OR MANUFACTURING ISSUES. THE CAUSE OF SURGICAL REVISION FOR INFECTION CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION. THESE DEVICES ARE USED FOR TREATMENT NOT DIAGNOSIS. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00622 | 2022-05-27 | 2022-05-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | VANTAGE | LP TALAR IMPACTOR INSERT | EXACTECH, INC. | 351-07-04 | 351-07-04 |  | IT WAS REPORTED THAT DURING IMPACTION OF THEN TALAR IMPLANT, THE TALAR IMPLANT INSERTER BROKE. NO PARTS OR PIECES FELL INTO THE PATIENTS WOUND SITE. THERE WAS NO SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICE IS RETURNING. | PENDING EVALUATION. | N | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00623 | 2022-05-27 | 2022-04-27 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | PRESERVE STEM 6MM | EXACTECH, INC. | 300-30-06 | 300-30-06 |  | AS REPORTED BY THE EQUINOXE SHOULDER STUDY, THIS 77 Y/O FEMALE PATIENT AT 250 LBS., EXPERIENCED A MID-SHAFT HUMERAL FRACTURE INTRAOPERATIVELY, THE SURGEON MANAGED THIS CORTICAL SPLIT USING CABLES. DURING THE PROCEDURE, THE PATIENT HAD AN AUTOGRAFT FROM THE HUMERAL HEAD; THE BONE GRAFT COMMENTS STATE THAT THERE WAS ¿A SPLIT ABOUT THE ANTERIOR MEDIAL CORTEX WAS GRAFTED¿. THE PATIENT ALSO HAD A BICEP TENODESIS DURING THIS PROCEDURE. THE PREOPERATIVE DIAGNOSIS FOR THIS PATIENT WAS A PROXIMAL HUMERUS FRACTURE. THE CASE REPORT FORM INDICATES THIS EVENT IS DEFINITELY RELATED TO DEVICES AND PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. NO ADDITIONAL INFORMATION. | (H3) BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN OR MANUFACTURING ISSUES. THE CAUSE OF THE SURGICAL EVENT CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION OF A PRE-SURGICAL DIAGNOSIS OF HUMERAL FRACTURE AND COMORBIDITIES. SECTION H11: THE FOLLOWING SECTIONS HAVE CORRECTED INFORMATION: (B5) DESCRIBE EVENT OR PROBLEM: AS REPORTED BY THE EQUINOXE SHOULDER STUDY, THIS 77 Y/O FEMALE PATIENT AT 250 LBS., EXPERIENCED A MID-SHAFT HUMERAL FRACTURE INTRAOPERATIVELY, THE SURGEON MANAGED THIS CORTICAL SPLIT USING CABLES. DURING THE PROCEDURE, THE PATIENT HAD AN AUTOGRAFT FROM THE HUMERAL HEAD; THE BONE GRAFT COMMENTS STATE THAT THERE WAS ¿A SPLIT ABOUT THE ANTERIOR MEDIAL CORTEX WAS GRAFTED¿. THE PATIENT ALSO HAD A BICEP TENODESIS DURING THIS PROCEDURE. THE PREOPERATIVE DIAGNOSIS FOR THIS PATIENT WAS A PROXIMAL HUMERUS FRACTURE. THE CASE REPORT FORM INDICATES THIS EVENT IS DEFINITELY RELATED TO DEVICES AND PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. NO ADDITIONAL INFORMATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00625 | 2022-05-27 | 2022-05-06 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 4 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 4 11MM | 02-012-35-4011 |  | AS REPORTED, APPROXIMATELY 8.5 YEARS POST INITIAL TKA, THE (B)(6) MALE PATIENT HAD A REVISION DUE TO PAIN AND DISCOMFORT. PATIENT'S WAS REVISED TO A COMPETITOR'S KNEE IMPLANTS. THERE WAS NO BREAKAGE OF A DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. X-RAYS RECEIVED. SALES REP WAS UNABLE TO OBTAIN PHOTOS. THE DEVICES ARE NOT AVAILABLE FOR EVALUATION. | CONCOMITANT MEDICAL PRODUCTS. LGC TIBIAL FIT TRAY CEM SZ 4F / 4T CAT# 02-012-45-4040 / SERIAL# (B)(4). LOGIC FEMORAL PS CEM RIGHT SZ 4 CAT# 02-010-01-0340 / SERIAL# (B)(4). THREE PEG PATELLA 35MM CAT# 200-02-35 / SERIAL# (B)(4). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00624 | 2022-05-27 | 2022-05-05 | I, n, j, u, r, y | Manufacturer report | NV CROWN CUP CLSTR HOLE 56MM GROUP 3 | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV CROWN CUP CLSTR HOLE 56MM GROUP 3 | 180-01-56 |  | AS REPORTED, APPROXIMATELY 10.5 YEARS POST INITIAL RIGHT THA, THE (B)(6) YEAR OLD MALE PATIENT WAS BROUGHT BACK FOR OSTEOLYSIS OF THE RIGHT HIP. THE HIP WAS DISLOCATED, FEMORAL HEAD REMOVED, STEM FOUND TO BE WELL FIXED. ATTENTION TURNED TO CUP. THE LINER WAS REMOVED WITH THE USE OF A CHISEL, THE BONE SCREW WAS REMOVED. THE CUP FOUND TO BE LOOSE AND WAS REMOVED. THE PATIENT RECEIVED 40CC OF OPTECURE MIXED WITH A CRUSHED FEMORAL HEAD ALLOGRAFT. A BIOMET G7 MULTIHOLE CUP WITH DUAL MOBILITY WAS IMPLANTED. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PHOTOS AND X-RAYS RECEIVED. THE DEVICES WILL BE RETURN. | CONCOMITANT MEDICAL PRODUCTS: BONE SCREW 6.5MM DIA X 25MM LONG (CAT# 120-65-25 / SERIAL# (B)(4) ). NV GXL LNR, LIPPED, 36MM ID, GROUP 3 CUPS (CAT# 132-36-53 / SERIAL# (B)(4) ). COCR FEM HEAD 36MM +0 OFFSET 12/14 (CAT# 142-36-00 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00627 | 2022-05-30 | 2022-05-09 | I, n, j, u, r, y | Manufacturer report | ELEMENT-STEM, COLLARED W/HA, STD OFFSET, SZ 8 | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | ELEMENT-STEM, COLLARED W/HA, STD OFFSET, SZ 8 | 164-03-08 |  | AS REPORTED, APPROXIMATELY 7.5 YEARS POST INITIAL RIGHT THA, THE FEMALE PATIENT WAS BROUGHT BACK FOR RIGHT HIP PAIN. SIGNS OF OSTEOLYSIS AND A LOOSE STEM. THE HIP WAS DISLOCATED, THE FEMORAL HEAD REMOVED, STEM CHECKED FOUND LOOSE AND REMOVED. ATTENTION TURNED TO CUP. LINER WAS REMOVED WITH A CHISEL, THE BONE SCREW WAS REMOVED. THE CUP WAS CHECKED AND FOUND TO BE WELL FIXED. A NEW XLE LINER WAS IMPLANTED ALONG WITH A HPS STEM. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PHOTOS AND X-RAYS RECEIVED. THE DEVICE IS NOT AVAILABLE FOR EVALUATION AS PATIENT ASKED TO KEEP IMPLANTS. | CONCOMITANT MEDICAL PRODUCTS: BONE SCREW 6.5MM DIA X 20MM LONG (CAT# 120-65-20 / SERIAL# (B)(4)). COCR FEM HEAD 32MM +3.5 OFFSET 12/14 (CAT# 142-32-03 / SERIAL# (B)(4)). NV GXL LINER NEUTRAL, 32MM ID GROUP 1 CUPS (CAT# 130-32-51 / SERIAL# (B)(4). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00626 | 2022-05-30 | 2022-05-10 | I, n, j, u, r, y | Manufacturer report | TRULIANT CR CEM FEM CR CEM RIGHT SZ 4.5 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT CR CEM FEM CR CEM RIGHT SZ 4.5 | 02-020-13-0345 |  | AS REPORTED, APPROXIMATELY THREE YEARS POST INITIAL RIGHT TKA, THE 73 Y/O MALE PATIENT COMPLAINED OF PAIN AND HAD A REVISION DUE TO LOOSE FEMUR. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PHOTO RECEIVED. SALES REP WAS UNABLE TO OBTAIN X-RAYS. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO HOSPITAL HAS DECIDED TO KEEP THE IMPLANTS. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-05
1038671-2022-00619 | 2022-06-01 | 2022-04-28 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 38MM HUMERAL LINER +0 | EXACTECH, INC. | 320-38-00 | 320-38-00 |  | AS REPORTED, APPROXIMATELY 1 YEAR POST OP THE INITIAL LEFT TSA, THIS (B)(6) MALE PATIENT'S WAS REVISED DUE TO THE PATIENT COMPLAINING OF PAIN. THERE WAS SOME NOTCHING NOTED. PT. GLENOSPHERE WAS REPLACE WITH +4MM EXPANDED. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES WILL NOT BE RETUNING DUE TO HOSPITAL POLICY. | PENDING EVALUATION: CONCOMITANT DEVICE(S): 320-10-00, 6728535 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0, 320-15-05, 6629035 - EQ REV LOCKING SCREW, 320-20-00, 6750662 - EQ REVERSE TORQUE DEFINING SCREW KIT, 320-06-38, 6581846 - GLENOSPHERE 38MM. WEIGHT,ETHINICITY,REVELANT TESTS, LABORATORYDATA, INCLUDING DATES, OTHER REVELANT HISTORY, INCLUDING PREEXISTING MEDICAL CONDITIONS, SUSPECTS PRODUCTS,WAS THE DEVICE SERVICED BYA THIRD PARTY? ADVERESE EVENT TERMS, IF REMEDIAL ACTION INITATED, CHECK TYPE, AND IF ACTION REPORTED TO FDA UNDER 21 USC 360 I(F), LIST CORRECTION/REMOVAL REPORTING NUMBER. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00628 | 2022-06-01 | 2022-02-09 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 42MM HUMERAL LINER +2.5 | EXACTECH, INC. | 320-42-03 | 320-42-03 |  | AS REPORTED BY AN EQUINOXE SHOULDER STUDY, APPROXIMATELY 3 MONTHS POST OP THE INITIAL TSA, THIS (B)(6) Y/O MALE PATIENT WAS REVISED DUE TO DEEP INFECTION. PATIENT HAD SCAR DISCHARGE AND SEVERE PAIN SINCE FEBRUARY 1ST. BIOLOGICAL ANALYSIS CONFIRMED INFECTION. THE CASE REPORT FORM INDICATES THIS EVENT IS DEFINITELY NOT RELATED TO DEVICES AND UNLIKELY RELATED TO PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | PENDING EVALUATION CONCOMITANT DEVICE(S): 320-01-42 - EQUINOXE REVERSE 42MM GLENOSPHERE. 300-01-13 - EQUINOXE, HUMERAL STEM PRIMARY, PRESS FIT 13MM. 320-10-00 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0. 320-15-02 - RS GLENOID PLATE SUP AUG, 10 DEG. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00629 | 2022-06-01 | 2022-02-19 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | RS GLENOID PLATE L POST AUG, 8 DEG, LEFT | EXACTECH, INC. | 320-15-03 | 320-15-03 |  | AS REPORTED BY AN EQUINOXE SHOULDER STUDY, APPROXIMATELY 3 YEARS POST OP THE INITIAL TSA, THIS (B)(6) Y/O MALE PATIENT WAS EXPERIENCING ASEPTIC GLENOID LOOSENING. PATIENT WAS PUSHING A SHOPPING CART, FELT POP, AND HAD INCREASED PAIN. SURGEON ORDERED A CT SCAN AND THE OUTCOME IS STILL CONTINUING. THE CASE REPORT FORM INDICATES THIS EVENT IS UNLIKELY RELATED TO DEVICES AND DEFINITELY NOT RELATED TO PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | CONCOMITANT DEVICE(S): 300-30-06, 5639974 - EQUINOXE PRESERVE STEM 6MM. 320-01-46, 5767789 - EQUINOXE REVERSE 46MM GLENOSPHERE. 320-10-00, 5888995 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0. 320-46-03, 5285948 - EQUINOXE REVERSE 46MM HUMERAL LINER +2.5. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00630 | 2022-06-01 | 2022-01-01 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | SHOULDER COMPONENTS | EXACTECH, INC. |  | UNK |  | IT WAS REPORTED, THIS (B)(6) Y/O PATIENT EXPERIENCED PAIN WITHOUT INJURY FOR 2-3 MONTHS. PATIENT IMMOBILIZED FOR 5 WEEKS. OUTCOME INDICATED RESOLVED ON (B)(6) 2022. ONSET DATE OF REPORTED PAIN IS UNKNOWN. SPECIFIC DEVICES WERE NOT REPORTED. THE CASE REPORT FORM INDICATES THIS EVENT IS UNLIKELY RELATED TO THE DEVICES AND IS DEFINITELY NOT RELATED TO THE PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00590 | 2022-06-02 | 2022-04-01 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 4 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 4 11MM | 02-012-35-4011 |  | AS REPORTED, APPROXIMATELY EIGHT YEARS POST INITIAL LEFT TKA, THE MALE PATIENT HAD A REVISION DUE TO POLY WEAR AND OSTEOLYSIS OF LOGIC PS INSERT. PATIENT WAS REVISED TO A LOGIC PSC SIZE 4 13MM INSERT. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS STABLE UPON LEAVING THE OR. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. X-RAYS AND PHOTOS RECEIVED. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO HSS KEEPS ALL RETRIEVALS. | CONCOMITANT MEDICAL PRODUCTS LOGIC FEMORAL PS CEM LEFT SZ 4 (CAT# 02-010-01-0240 / SERIAL# (B)(4)) LOGIC TIBIA TRAPTRAY CEM SZ 4F/4T (CAT# 02-012-41-4040 / SERIAL# (B)(4)) THREE PEG PATELLA 38MM (CAT#200-02-38 / SERIAL# (B)(4)) ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00631 | 2022-06-02 | 2022-05-12 | I, n, j, u, r, y | Manufacturer report | OPTETRAK SIZE 4, 11MM POLY | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | OPTETRAK SIZE 4, 11MM POLY | UNKNOWN |  | AAS REPORTED, SURGEON HAD TO REVISE A 17 YEAR OLD OPTETRAK SIZE 4, 11MM POLY AND PATELLA. PATIENT WAS REVISED TO A NEW PATELLA AND SIZE 4, 9MM TIBIAL POLY. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. SALES REP WAS UNABLE TO OBTAIN PHOTOS OR X-RAYS. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICE IS NOT AVAILABLE FOR EVALUATION AS DOCTOR KEPT THE IMPLANTS. | CONCOMITANT MEDICAL PRODUCTS: PATELLA. ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00632 | 2022-06-05 | 2022-05-11 | I, n, j, u, r, y | Manufacturer report | TRULIANT PS CEM FEM PS CEM LEFT SZ 5 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT PS CEM FEM PS CEM LEFT SZ 5 | 02-020-11-0250 |  | AS REPORTED, APPROXIMATELY TWO YEARS POST INITIAL LEFT TKA, THE (B)(6) MALE PATIENT COMPLAINED OF PAIN AND HAD A REVISION DUE TO LOOSE FEMUR. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PHOTO RECEIVED. SALES REP WAS UNABLE TO OBTAIN X-RAYS. THE DEVICE IS NOT AVAILABLE FOR EVALUATION BECAUSE HOSPITAL IS NOT RELEASING THE IMPLANT. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00640 | 2022-06-07 | 2022-05-18 | I, n, j, u, r, y | Manufacturer report | NV GXL LINER NEUTRAL, 40MM ID, GROUP 3 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LINER NEUTRAL, 40MM ID, GROUP 3 CUPS | 130-40-53 |  | AS REPORTED, APPROXIMATELY SIX YEARS POST INITIAL LEFT THA, THE (B)(6) FEMALE PATIENT HAD A REVISION DUE G2/36MM NEUTRAL GXL LINER HAS EARLY ECCENTRIC WEAR. SURGEON SWAPPED THE WORN POLY FOR A G2/36MM NEUTRAL GXLE AND A CERAMIC 36MM HEAD WITH +3.5 SLEEVE OPTION. PATIENT'S DIAGNOSIS IS POLYETHYLENE WEAR AND OSTEOLYSIS. X-RAYS RECEIVED. THE DEVICE WILL BE RETURN. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00639 | 2022-06-07 | 2022-05-18 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 4 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 4 13MM | 02-012-35-4013 |  | AS REPORTED, APPROXIMATELY FIVE YEARS POST INITIAL RIGHT TKA, THE (B)(6) MALE PATIENT HAD A REVISION DUE TO LOOSE FEMUR AND WORN POLY. PATIENT WAS DIAGNOSED WITH POLYETHYLENE WEAR, OSTEOLYSIS, POLYETHYLENE SYNOVITIS AND COMPONENT LOOSENING. PATIENT WAS REVISED TO TRULIANT CC COMPONENTS. X-RAYS RECEIVED. THE DEVICE WILL BE RETURN. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00638 | 2022-06-07 | 2022-05-18 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 5 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 5 9MM | 02-012-35-5009 |  | AS REPORTED, APPROXIMATELY TWO YEARS POST INITIAL RIGHT TKA, THE MALE PATIENT HAD A REVISION DUE TO INFECTION. PATIENT HAD A LINER REPLACED. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. SALES REP WAS UNABLE TO OBTAIN PHOTOS/X-RAYS. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO DEVICES WAS DISPOSED BY HOSPITAL. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00633 | 2022-06-07 | 2020-06-01 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 46MM HUMERAL LINER +0 | EXACTECH, INC. | 320-46-00 | 320-46-00 |  | IT WAS REPORTED THAT THIS PATIENT HAD A REVISION WHICH WAS NOT PREVIOUSLY REPORTED. REASON HAS NOT BEEN REPORTED. THIS WAS FOUND DURING THE INVESTIGATION OF (B)(4). (MDR# 1038671-2022-00072). | CONCOMITANT DEVICE(S): 320-02-46. 6480696 - 46X21MM GLENOSPHERE HIGH MOMENT ARM. 320-10-00 , 6487055- EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0. 320-15-05 , 6362580 - EQ REV LOCKING SCREW. 320-20-00, 6464845 - EQ REVERSE TORQUE DEFINING SCREW KIT. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00636 | 2022-06-07 | 2022-04-23 | I, n, j, u, r, y | Manufacturer report | LOGIC CC TIB INSERT SIZE 3, 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC CC TIB INSERT SIZE 3, 13MM | 02-012-65-3013 |  | AS REPORTED, APPROXIMATELY 5.5 YEARS POST INITIAL RIGHT TKA, SURGEON PREFORMED A POLY SWAP TO THE (B)(6) FEMALE PATIENT DUE TO INSTABILITY. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. SALES REP WAS UNABLE TO OBTAIN PHOTOS/X-RAYS. THE PRODUCT IS NOT AVAILABLE FOR RETURN. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00637 | 2022-06-07 | 2022-05-17 | I, n, j, u, r, y | Manufacturer report | LOGIC CR TIB INSERT STD, SZ 2.5, 9 MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC CR TIB INSERT STD, SZ 2.5, 9 MM | 02-012-47-2509 |  | AS REPORTED, APPROXIMATELY THREE YEARS POST INITIAL RIGHT TKA, THE FEMALE PATIENT HAD A REVISION DUE TO EFFECTED POLY CAUSED EXCESSIVE WEAR WHICH HAS LED TO LYSIS AND PAIN. PATIENT WAS REVISED TO A 11MM INSERT. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PHOTOS RECEIVED. SALES REP WAS UNABLE TO OBTAIN X-RAYS. THE DEVICE WILL BE RETURN. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00635 | 2022-06-07 | 2022-05-13 | I, n, j, u, r, y | Manufacturer report | TRULIANT TIB IMP PS INSERT SZ 3 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT TIB IMP PS INSERT SZ 3 13MM | 02-022-35-3013 |  | AS REPORTED, APPROXIMATELY TWO YEARS POST INITIAL RIGHT TKA, SURGEON PREFORMED A POLY SWAP OF A (B)(6) Y/O MALE PATIENT DUE TO INSTABILITY. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. SALES REP WAS UNABLE TO OBTAIN PHOTOS/X-RAYS. THE PRODUCT IS NOT AVAILABLE FOR RETURN. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00650 | 2022-06-08 | 2022-01-20 | I, n, j, u, r, y | Manufacturer report | EUINOXE | RS GLENOID PLATE SUP AUG, 10 DEG | EXACTECH, INC. | 320-15-02 | 320-15-02 |  | AS REPORTED BY THE EQUINOXE SHOULDER STUDY, APPROXIMATELY 1.5 YEARS POST OP THE INITIAL TSA, THIS (B)(6) YEAR OLD MALE PATIENT WAS REVISED DUE TO SEPTIC LOOSENING. AT THE (B)(6) 2022 VISIT, THERE WAS SMALL SUSPICION FOR POSSIBLE INFECTION, SUBJECT HAD LABS DONE AND MET WITH ID IN GREENVILLE WHO PRESCRIBED ANTIBIOTICS INDEFINITELY. AT THE (B)(6) 2021 VISIT, XRAYS SUGGESTED LOOSENING OF THE GLENOID COMPONENT. AT (B)(6) 2021 VISIT, LOOSENING HAD PROGESSED. SCHEDULED FOR REVISION (B)(6) 2022. THE CASE REPORT FORM INDICATES THIS EVENT IS DEFINITELY NOT RELATED TO DEVICES AND POSSIBLY RELATED TO PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | CONCOMITANT DEVICE(S): 300-01-11 - EQUINOXE, HUMERAL STEM PRIMARY, PRESS FIT 11MM. 320-38-00 - EQUINOXE REVERSE 38MM HUMERAL LINER +0. 320-10-00 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0. 320-02-38 - RS EXPANDED GLENOSPHERE 38MM, +4MM OFFSET. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00655 | 2022-06-08 | 2022-05-19 | I, n, j, u, r, y | Manufacturer report | ADVANCED PATELLA 32MM 3 PEG IMPLANT | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | ADVANCED PATELLA 32MM 3 PEG IMPLANT | 200-07-32 |  | AS REPORTED BY THE EXACTECH TRULIANT KNEE CLINICAL STUDY, APPROXIMATELY ONE YEAR POST INITIAL TKA, THE (B)(6) Y/O FEMALE PATIENT EXPERIENCED PAIN IN RIGHT KNEE PATELLAR REGION AND CLICKING WHEN ASCENDING STAIRS AND WHEN BENDING RIGHT KNEE. PATIENT IS SCHEDULED FOR A RIGHT KNEE ARTHROSCOPIC WITH SYNOVECTOMY AT SUBJECT'S CONVENIENCE. | CONCOMITANTS MEDICAL PRODUCTS: TRULIANT TIB FIT TRAY CEM SZ 3.5F / 2.5T (CAT# 02-022-45-3525). TRULIANT PS CEM FEM PS CEM RIGHT SZ 3.5 (CAT# 02-020-11-0335) . TRULIANT TIB FIT TRAY CEM SZ 3.5F / 2.5T (CAT# 02-022-45-3525). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00654 | 2022-06-08 | 2021-03-09 | I, n, j, u, r, y | Manufacturer report | TRULIANT TIB IMP CR INS STD SZ 3, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT TIB IMP CR INS STD SZ 3, 9MM | 02-022-47-3009 |  | AS REPORTED BY THE EXACTECH KNEE REPLACEMENT STUDY, APPROXIMATELY FIVE MONTHS POST INITIAL TKA, THE 73 Y/O FEMALE PATIENT COMPLAINTS OF STIFFNESS RIGHT KNEE AND PAIN, LIMITED RANG OF MOTION (ROM) AND MINIMAL DISCOMFORT. THE EVENT IS STILL ONGOING, PATIENT IS TAKING MEDICATION AND ORDERED PHYSICAL THERAPY (PT). PER CLINICAL STUDY CASE REPORT FORM INDICATES THIS EVENT IS UNLIKELY RELATED TO DEVICES AND POSSIBLY RELATED TO PROCEDURE. | CONCOMITANTS MEDICAL PRODUCTS: TRULIANT CR CEM FEM CR CEM RIGHT SZ 3 (CAT# 02-020-13-0330), TRULIANT TIB FIT TRAY CEM SZ 3F / 2T (CAT# 02-022-45-3020), ADVANCED PATELLA 32MM 3 PEG IMPLANT (CAT# 200-07-32). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00652 | 2022-06-08 | 2019-12-16 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | HUMERAL LINER, 40MM, +2.5 | EXACTECH, INC. | 320-40-03 | 320-40-03 |  | AS REPORTED BY EQUINOXE SHOULDER STUDY, APPROXIMATELY 2 MONTHS POST OP THE INITIAL TSA, THIS (B)(6) MALE PATIENT EXPERIENCED REDNESS AND SWELLING TO HIS PROXIMAL INFECTION. HE ALSO HAD SURGICAL REMOVAL OF A HAEMATOMA. THE CASE REPORT FORM INDICATES THIS EVENT IS DEFINITELY NOT RELATED TO DEVICES AND DEFINITELY RELATED TO PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | DEVICE EVALUATED BY MFR: PENDING EVALUATION. CONCOMITANT MEDICAL PRODUCTS: 320-32-40 - EXPANDED GLENOSPHERE, 40MM, FOR SMALL REVERSE. 300-30-12 - EQUINOXE PRESERVE STEM 12MM. 320-10-10 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +10 REVERSE GLENOSPHERE BASEPLATE. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00653 | 2022-06-08 | 2020-01-23 | I, n, j, u, r, y | Manufacturer report | TRULIANT TIB FIT TRAY CEM SZ 3F / 3T | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT TIB FIT TRAY CEM SZ 3F / 3T | 02-022-45-3030 |  | AS REPORTED BY THE EXACTECH KNEE REPLACEMENT STUDY, APPROXIMATELY THREE MONTHS POST INITIAL TKA, THE (B)(6) FEMALE PATIENT COMPLAINS THAT BECAUSE OF POST -OP INFECTIONS AND LOSS OF RANGE OF MOTION. PATIENT DOCUMENTS THAT SHE IS MORE CRIPPLED THAN BEFORE SURGERY. THE EVEN IS STILL ONGOING. PER CLINICAL STUDY CASE REPORT THIS EVENT IS POSSIBLY RELATED TO DEVICES AND TO PROCEDURE. | CONCOMITANTS MEDICAL PRODUCT: FEMORAL COMPONENT (CAT# 02-020-13-3011), TIBIAL INSERT (CAT# 02-022-47-3030), THREE PEG PATELLA 32MM (CAT# 200-02-32), OPTETRAK (4) HOLDING PINS (CAT# 201-78-15). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00651 | 2022-06-08 | 2022-05-17 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 1, 13MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED, | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 1, 13MM | 02-012-44-1013 |  | AS REPORTED, APPROXIMATELY TWO YEARS POST INITIAL LEFT TKA, THE 66 Y/O FEMALE PATIENT HAD A REVISION. PATIENT WAS COMPLAINING SINCE 6 MONTHS OF ORIGINAL TKA OF PAIN, SWELLING AND DISCOMFORT. PATIENT HAD LITTLE INFLAMMATORY SYNDROME BUT PRESENCE OF GRANULUM. THE PE IS HEAVILY WORN BEHIND THE POSTERIOR CAM AND IN THE ANGLE OF THE EXTERNAL FEMORAL CONDYLE. ONLY PATIENT'S INSERT HAS BEEN REPLACED. X-RAYS AND PHOTOS RECEIVED. THE DEVICE WILL BE RETURN. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00648 | 2022-06-08 | 2022-04-15 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE TRAY ADAPTER PLATE TRAY +0 | EXACTECH, INC. | 320-10-00 | 320-10-00 |  | AS REPORTED, APPROXIMATELY 7 MONTHS POST OP THE INITIAL RIGHT TSA, THIS 68 Y/O FEMALE PATIENT'S WAS REVISED DUE TO DISLOCATION. THE PATIENT CURRENTLY HAS AN EXACTECH STEM FROM A PREVIOUS SURGERY. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICE IS NOT RETURNING. | PENDING EVALUATION. CONCOMITANT DEVICE(S): 320-06-42, 6783069: GLENOSPHERE 42MM, 320-20-00, 6959084: EQ REVERSE TORQUE DEFINING SCREW KIT, 320-15-05, 6945027: EQ REV LOCKING SCREW, 320-42-00, S219490: EQUINOXE REVERSE 42MM HUMERAL LINER +0. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00649 | 2022-06-08 | 2021-11-24 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | HUMERAL STEM PRIMARY PRESS FIT 17MM | EXACTECH, INC. | 300-01-17 | 300-01-17 |  | AS REPORTED BY THE EQUINOXE SHOULDER STUDY, APPROXIMATELY 7 YEARS POST OP THE INITIAL TSA, THIS 51 Y/O FEMALE PATIENT'S WAS REVISED TO A STANDARD REVERSE DUE TO ASEPTIC HUMERAL LOOSENING. PATIENT HAD PERIPROSTHETIC PAIN, NEGATIVE FOR INFECTION, AND SURGEON PLANNED AB SPACER PLACEMENT. THE CASE REPORT FORM INDICATES THIS EVENT IS POSSIBLY RELATED TO DEVICES AND/OR PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | PENDING EVALUATION. CONCOMITANT DEVICE(S): 320-01-38 - EQUINOXE REVERSE 38MM GLENOSPHERE; 320-38-00 - EQUINOXE REVERSE 38MM HUMERAL LINER +0; 320-10-00 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0; 320-15-01 - EQ REV GLENOID PLATE. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00646 | 2022-06-08 | 2022-03-28 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | PRESERVE STEM 10MM | EXACTECH, INC. | 300-30-10 | 300-30-10 |  | IT WAS REPORTED THAT THIS (B)(6) MALE PATIENT'S RIGHT SHOULDER WAS REVISED DUE TO LOOSENING. THE PATIENT'S PRESERVE STEM APPEARED TO BE LOOSE AT TIME OF REVISION. SURGEON REMOVED EXISTING EXACTECH IMPLANTS AND REPLACED THEM WITH A 42+4 GLENOSPHERE. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES ARE NOT AVAILABLE FOR RETURN. | PENDING EVALUATION. CONCOMITANT DEVICE(S): 320-06-38, 7063923 - GLENOSPHERE 38MM. 320-10-00, 7052245 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0. 320-38-00, UNKNOWN - EQUINOXE REVERSE 38MM HUMERAL LINER +0. 320-06-38, 7063923 - GLENOSPHERE 38MM. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00645 | 2022-06-08 | 2022-05-11 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | LINER | EXACTECH, INC. | UNK | UNK |  | IT WAS REPORTED THAT THIS (B)(6) MALE PATIENT'S RIGHT SHOULDER WAS REVISED DUE TO INFECTION. SURGEON PERFORMED A WASH OUT AND REPLACED THE GLENOSPHERE, LINER, TRAY, GLENOSPHERE LOCKING SCREW AND TORQUE SCREW. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. THE DEVICES WILL NOT BE RETURNING DUE TO HOSPITAL POLICY. | PENDING EVALUATION. CONCOMITANT DEVICE(S): TRAY, GLENOSPHERE, TORQUE SCREW, GLENO LOCKING SCEW. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00644 | 2022-06-08 | 2022-05-04 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | SHOULDER COMPONENT | EXACTECH, INC. | UNK | UNK |  | IT WAS REPORTED THAT A (B)(6) FEMALE PATIENT UNDERWENT A RIGHT SHOULDER REVISION. A LOOSE BASEPLATE WAS REPORTED. THE PATIENT WAS REVISED WITH A EQUINOX CTA HEAD. THE PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. NO SPECIFIC DEVICE REPORTED. NO PRODUCT RETURNING DUE TO HOSPITAL POLICY. | PENDING EVALUATION. CONCOMITANT MEDICAL PRODUCTS: TRAY, LINER, BASEPLATE, GLENOSPHERE. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00643 | 2022-06-08 | 2022-05-02 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | PRESERVE STEM 6MM | EXACTECH, INC. | 300-30-06 | 300-30-06 |  | IT WAS REPORTED THAT A (B)(6) Y/O MALE PATIENT, UNDERWENT A REVISION OF EXACTECH IMPLANTS ON HIS RIGHT SHOULDER. THE HUMERAL COMPONENTS WERE REPLACED WITH ARTHREX COMPONENTS. THE PATIENT HAD BEEN DISLOCATING HIS SHOULDER. THE LINER APPEARED TO STILL BE LOCKED INTO THE TRAY AT TIME OF REVISION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. UNABLE TO OBTAIN PHOTOS OR X-RAYS. THE PRODUCT IS NOT AVAILABLE FOR RETURN. | (DEVICE EVALUATED BY MFR) PENDING EVALUATION. CONCOMITANT DEVICE(S): 320-10-00, 49041934 - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0 UNKNOWN, UNKNOWN - 42 LINER. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00642 | 2022-06-08 | 2022-03-28 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 38MM GLENOSPHERE | EXACTECH, INC. | 320-01-38 | 320-01-38 |  | IT WAS REPORTED THAT A (B)(6) FEMALE PATIENT UNDERWENT A LEFT SIDE REVISION OF EXACTECH IMPLANTS, IN WHICH HER REVERSE TOTAL SHOULDER WAS CONVERTED TO A HEMI SHOULDER. THE GLENOID SIDE WAS REMOVED, AND 47 TALL CTA HEAD AND 0 REPLICATOR PLATE WAS IMPLANTED. NO PHOTOS OR X-RAYS OBTAINED. THE PRODUCT IS NOT AVAILABLE FOR RETURN. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. | PENDING EVALUATION. CONCOMITANT MEDICAL DEVICE(S): 320-10-05, 57850001: EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +5. UNKNOWN, UNKNOWN: 38 NEUTRAL LINER. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00641 | 2022-06-08 | 2022-05-09 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | STEMLESS HUMERAL COMP LASER CAGE, SIZE 1 | EXACTECH, INC. | 300-60-01 | 300-60-01 |  | IT WAS REPORTED THAT THE SURGEON CONVERTED A (B)(6) YEAR OLD MALE PATIENT FROM A STEMLESS ANATOMIC SHOULDER TO A REVERSE TOTAL SHOULDER. PATIENT WAS REVISED TO AN EQUINOXE REVERSE SHOULDER GLENOID PLATE +10MM EXTENDED CAGE AND AN EQUINOXE REVERSE SHOULDER GLENOSPHERE & EXTENDED LOCKING CAPS 42MM, +4 OFFSET. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICE IS NOT RETURNING. | (CONCOMITANT DEVICE(S): 310-61-50, 153399003 - STEMLESS HUMERAL HEAD 50MM X 19MM X BETA UNKNOWN, UNKNOWN - CAGE GLENOID. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00647 | 2022-06-08 | 2022-05-10 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | CAGE GLENOID L, POST AUG, RIGHT | EXACTECH, INC. | 314-13-34 | 314-13-34 |  | SURGEON USED A CAGED AUGMENTED GLENOID IN HIS CASE. THEY FOLLOWED THE SURGICAL STEPS AND THEN TRIED TO IMPLANT THE GLENOID. THE (B)(6) PATIENT'S BONE WAS VERY HARD AND THEY WERE UNABLE TO SEAT THE IMPLANT. THE SURGEON THEN TRIED TO REMOVE THE IMPLANT BUT TWO OF THE METAL PIECES DISENGAGED AND STAYED IN THE GLENOID. IT WAS VERY DIFFICULT TO REMOVE THE TWO PARTS. THEY THEN IMPLANTED A ALL POLY ONE TO COMPLETE THE CASE. THERE WAS A 5-15 MINS SURGICAL DELAY/PROLONGATION, NO ADVERSE EVENTS AS A RESULT OF THE DELAY PROLONGATION. THE PATIENT HAD NO ADVERSE AFFECTS. | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00664 | 2022-06-09 | 2021-01-21 | I, n, j, u, r, y | Manufacturer report | FEMORAL COMPONENTS | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | FEMORAL COMPONENTS | UNKNOWN |  | AS REPORTED BY LEGAL BRIEF, ON (B)(6) 2011, PATIENT UNDERWENT BILATERAL KNEE REPLACEMENT SURGERY AND WAS IMPLANTED WITH OPTETRAK DEVICES. PATIENT DEVELOPED SEVERE PAIN IN HIS RIGHT KNEE AND WAS ADVICE HE WOULD REQUIRE REVISION SURGERY DUE TO EXTENSIVE OSTEOLYSIS OF THE FEMORAL COMPONENT AND ADDITIONAL OSTEOLYSIS NOTED IN THE TIBIA. ON (B)(6) 2021, PATIENT UNDERWENT REVISION SURGERY OF HIS RIGHT KNEE IN WHICH ¿MARKED SYNOVITIS WAS ENCOUNTERED WITH TISSUE CHANGES SUGGESTIVE OF OSTEOLYSIS¿; FURTHERMORE, INSPECTION OF THE POLYETHYLENE COMPONENT DEMONSTRATED PITTER WEAR, DELAMINATION AS WELL AS SOME BACKSIDE WEAR AND THERE WAS EXTENSIVE DAMAGE TO THE BONY ARCHITECTURE WAS NOTED SECONDARY TO OSTEOLYSIS UPON CEMENT REMOVAL. PATIENT ALSO WAS REQUIRED TO UNDERGO REVISION SURGERY ON HIS LET KNEE DUE TO PREMATURE FAILURE OF THE OPTETRAK DEVICE. PATIENT DEVELOPED SEVERE PAIN IN HIS LEFT KNEE AND WAS DIAGNOSED WITH EXTENSIVE OSTEOLYSIS. ON (B)(6) 2021, PATIENT UNDERWENT REVISION SURGERY OF HIS LEFT KNEE IN WHICH AGAIN MARKED SYNOVITIS WAS ENCOUNTERED WITH TISSUE CHANGES SUGGESTIVE OF OSTEOLYSIS. PATIENT¿S SURGEON REMOVED LARGE AMOUNTS OF OSTEOLYTIC TISSUE AND MEMBRANE AND DOCUMENTED EXTENSIVE BONE LOSS OF THE FEMUR. AS A DIRECT AND PROXIMATE RESULT, PATIENT HAS SUFFERED AND CONTINUE TO SUFFER PERMANENT AND DEBILITATING INJURIES AND DAMAGES, INCLUDING BUT NOT LIMITED TO, SIGNIFICANT PAIN AND DISCOMFORT; GAIT IMPAIRMENT; POOR BALANCE; DIFFICULTY WALKING; COMPONENT PART LOOSENING; SOFT TISSUE DAMAGE; BONE LOSS; AND OTHER INJURIES PRESENTLY UNDIAGNOSED, WHICH ALL REQUIRE ONGOING MEDICAL CARE. AS FURTHER DIRECT, PATIENT HAS SUSTAINED AND WILL SUSTAIN FUTURE DAMAGES, INCLUDING BUT NOT LIMITED TO COST OF MEDICAL CARE, REHABILITATION, HOME HEALTH CARE, LOSS OF EARNING CAPACITY, MENTAL AND EMOTIONAL DISTRESS AND PAIN AND SUFFERING. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00667 | 2022-06-09 | 2022-05-23 | I, n, j, u, r, y | Manufacturer report | LOGIC CR TIB INSERT SLOPE++, SZ 4, 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC CR TIB INSERT SLOPE++, SZ 4, 11MM | 02-012-49-4011 |  | AS REPORTED, APPROXIMATELY THREE YEARS POST INITIAL LEFT TKA, THE FEMALE PATIENT HAD A REVISION AS THERE WAS OBVIOUS WEAR ON THE SURFACE OF THE POLY LINER, ALSO, DURING THE CHECK-UP PRIOR TO REVISION SURGERY, THE PATIENT¿S BONE SCAN SHOWED RED SPOTS ON PATELLA. PATIENT HAD SWOLLEN KNEE AND ANTERIOR PAIN. PATIENT WAS REVISE TO A CRC LINER. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION REPORTED. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. X-RAYS AND PHOTOS RECEIVED. THE DEVICE WILL BE RETURN. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00666 | 2022-06-09 | 2022-05-20 | I, n, j, u, r, y | Manufacturer report | NV GXL LINER NEUTRAL, 32MM ID GROUP 1 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LINER NEUTRAL, 32MM ID GROUP 1 CUPS | 130-32-51 |  | AS REPORTED, THE 66 Y/O FEMALE PATIENT HAD A REVISION OF TOTAL HIP DUE TO WEAR. THE STEM WAS RETAINED AND A NEW EXATEXH HEAD WAS USED. THE CUP AND LINER WAS REPLACED WITH COMPETITORS DEVICE. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PHOTOS AND X-RAYS RECEIVED. THE DEVICE WILL BE RETURN. | CONCOMITANT MEDICAL PRODUCTS: NV CROWN CUP CLSTR HOLE 48MM GROUP 1 (CAT# 180-01-48 / SERIAL# (B)(4), COCR FEM HEAD 32MM +0 OFFSET 12/14 (CAT# 142-32-00 / SERIAL# (B)(4). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | Y | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00665 | 2022-06-09 | 2022-01-04 | I, n, j, u, r, y | Manufacturer report | PS TIBIAL INSERTS SZ 3, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | PS TIBIAL INSERTS SZ 3, 9MM | 204-23-09 |  | AS REPORTED BY LEGAL BRIEF, ON OR ABOUT (B)(6) 2014, PATIENT UNDERWENT A LEFT TOTAL KNEE REPLACEMENT SURGERY IN WHICH AN OPTETRAK SYSTEM WAS IMPLANTED TO TREAT LEFT KNEE DEGENERATIVE JOINT DISEASE. AFTER IMPLANTATION PATIENT EXPERIENCED PAIN AND OTHER SYMPTOMS AND WAS ULTIMATELY DIAGNOSED IN 2021 WITH FAILED LEFT TOTAL KNEE ARTHROPLASTY. ON (B)(6) 2022, AS A DIRECT AND PROXIMATE RESULT THE PATIENT REQUIRED LEFT TOTAL KNEE REVISION SURGERY TO REMOVE FAILED DEVICE AND REPLACE IT WITH A NEW TOTAL KNEE REPLACEMENT SYSTEM. | MEDICAL PRODUCTS: CONCOMITANT MEDICAL PRODUCTS: 3 PEG 29MM PATELLA, SIZE 3 TIBIAL BASEPLATE, SIZE 3 PS FEMORAL KNEE. ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00668 | 2022-06-09 | 2022-06-15 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT | SEE H10 | EXACTECH, INC. | UNK | 02-012-44-3009 |  | AS REPORTED BY LEGAL BRIEF, 0N (B)(6) 2015, PATIENT UNDERWENT A RIGHT KNEE REPLACEMENT SURGERY AND WAS IMPLANTED WITH AN OPTETRAK DEVICE. APPROXIMATELY SEVEN YEARS POST INITIAL TKA, PATIENT DEVELOPED SEVERE PAIN IN HIS RIGHT KNEE AND IS SCHEDULED TO UNDERGO REVISION SURGERY ON (B)(6) 2022. UPON INFORMATION AND BELIEF, PATIENT REQUIRES REVISION SURGERY FOR ISSUES INCLUDING BUT NOT LIMITED TO POLYETHYLENE WEAR, BONE LOSS, OSTEOLYSIS, AND/OR COMPONENT LOOSENING. PATIENT EXPERIENCES DAILY PAIN AND DISCOMFORT IN HIS KNEE WHICH LIMITS HIS ACTIVITIES OF DAILY LIVING AND IMPACTS HIS QUALITY OF LIFE. AS A DIRECT, PROXIMATE AND RESULT, PATIENT HAS SUFFERED AND CONTINUES TO SUFFER PERMANENT AND DEBILITATING INJURES AND DAMAGES, INCLUDING BUT NOT LIMITED TO, SIGNIFICANT PAIN AND DISCOMFORT; GAIT IMPAIRMENT; POOR BALANCE; DIFFICULTY WALKING; COMPONENT PART LOOSENING; SOFT TISSUE DAMAGE; BONE LOSS; AND OTHER INJURIES PRESENTLY UNDIAGNOSED, WHICH ALL REQUIRE ONGOING MEDICAL CARE. | PENDING EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00663 | 2022-06-09 | 2008-10-28 | I, n, j, u, r, y | Manufacturer report | FEMORAL COMPONENTS | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | FEMORAL COMPONENTS | UNKNOWN |  | AS REPORTED BY THE LEGAL BRIEF, IN 2006, PATIENT SOUGHT TREATMENT FOR KNEE PAIN. AT THE RECOMMENDATION OF THE ORTHOPEDIC SURGEON, SHE UNDERWENT KNEE REPLACEMENT SURGERY ON HER RIGHT KNEE WITH AN OPTETRAK KNEE REPLACEMENT SYSTEM. IN 2008, LESS THAN TWO YEARS AFTER HER KNEE REPLACEMENT SURGERY, PATIENT DEVELOPED LOOSENING OF THE DEVICE AND PAIN IN HER RIGHT KNEE. ON (B)(6) 2008, PATIENT UNDERWENT A SURGICAL REVISION OF HER RIGHT KNEE DUE TO "FAILED RIGHT KNEE WITH A LOOSE TIBIAL COMPONENT". AT THE TIME OF SURGERY, IT WAS DISCOVERED THAT PATIENT RIGHT KNEE WAS ALSO INFECTED AND, THEREFORE, THE OPTETRAK KNEE REPLACEMENT SYSTEM AND ITS COMPONENTS WERE REMOVED WHILE PATIENT UNDERWENT TREATMENT FOR THE INFECTION. PATIENT CONTINUED TREATMENT FOR THE INFECTION UNTIL (B)(6) 2009, WHEN ANOTHER OPTETRAK KNEE REPLACEMENT SYSTEM WAS IMPLANTED IN HER RIGHT KNEE ALONG WITH ANOTHER DEFECTIVE INSERT. AT THAT TIME, UNBEKNOWNST TO PATIENT HAD INTRODUCED A NEW FINNED TIBIAL TRAY, FIT, TO REPLACE THE DEFECTIVE TIBIAL TRAY DUE TO NUMEROUS REPORTS OF LOOSENING AND BONE LOSS. PATIENT CONTINUED TO HAVE PAIN AND INSTABILITY IN CONNECTION WITH HER OPTETRAK KNEE REPLACEMENT SYSTEM AND, IN (B)(6) 2012, IT WAS DETERMINED THAT THE DEVICE HAD AGAIN LOOSENED AND FAILED, AND THAT, THEREFORE, SHE WAS REQUIRED TO UNDERGO ANOTHER REVISION. ON (B)(6) 2012, SURGERY WAS ONCE AGAIN PERFORMED ON PATIENT'S RIGHT KNEE. AT THIS RE-REVISION, ANOTHER DEFECTIVE INSERT WAS IMPLANTED IN PATIENT'S KNEE. PATIENT CONTINUED TO HAVE DIFFICULTY WITH HER RIGHT KNEE, AND, IN (B)(6) 2021, SHE SWITCHED HER CARE. ON (B)(6) 2021, PATIENT UNDERWENT ANOTHER REVISION DUE TO LOOSENING AND BONE LOSS RELATED TO THE DEFECTIVE INSERT. AS A RESULT OF THIS REVISION SURGERY, PATIENT'S POPLITEAL ARTERY WAS SEVERED, AND PATIENT CONTINUES TO SUFFER FROM MEDICAL COMPLICATIONS AND RISKS. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00659 | 2022-06-09 | 2020-12-17 | I, n, j, u, r, y | Manufacturer report | LOGIC FEMORAL PS CEM RIGHT SZ 2 | KNEE FEMORAL COMPONENTS | EXACTECH, INC. | LOGIC FEMORAL PS CEM RIGHT SZ 2 | 02-010-01-0320 |  | AS REPORTED BY THE LEGAL BRIEF, ON (B)(6) 2016. PATIENT UNDERWENT TOTAL KNEE REPLACEMENT SURGERY ON HER RIGHT KNEE, IN WHICH AN OPTETRAK DEVICE WAS IMPLANTED. IN THE YEARS FOLLOWING THE SURGERIES, PATIENT EXPERIENCED PAIN, SWELLING, INSTABILITY AND BONE LOSS IN HER RIGHT KNEE CAUSED BY EARLY AND ACCELERATED POLYETHYLENE WEAR AND/OR COMPONENT LOOSENING. ULTIMATELY, ON (B)(6) 2021, PATIENT UNDERWENT AN EXTENSIVE REVISION SURGERY ON HER RIGHT KNEE. AN X-RAY IMAGE TAKEN ON (B)(6) 2020, DEMONSTRATED ¿OBVIOUS FEMORAL LOOSENING¿ OF HER EXACTECH DEVICE. DURING EXPLANATION OF THE OPTETRAK DEVICES, PATIENT¿S SURGEON OBSERVED SIGNIFICANT GRANULATION TISSUE IN HER SYNOVIUM, CONSISTENT WITH POLYETHYLENE WEAR. BOTH THE TIBIAL AND FEMORAL COMPONENTS WERE LOOSE AND COULD BE REMOVED WITHOUT SIGNIFICANT EFFORT. THERE WERE NO SIGNS OF INFECTION, AND IT WAS DETERMINED THAT THIS WAS A MECHANICAL/ASEPTIC LOOSENING OF THE HARDWARE. AS A DIRECT AND PROXIMATE RESULTS, PATIENT HAS SUFFERED AND CONTINUES TO SUFFER INJURY AND DAMAGES, INCLUDING HAVING TO UNDERGO A PAINFUL REVISION SURGERY; HAVING REQUIRED AND WILL CONTINUE TO REQUIRE ADDITIONAL MEDICAL CARE AND TREATMENT, HAVING REQUIRE AND WILL CONTINUE TO REQUIRE PHYSICAL THERAPY AND PAIN MANAGEMENT, AND HAVING EXPERIENCED AND WILL CONTINUE TO EXPERIENCE PROLONGED AND LASTING PAIN AND SUFFERING AND LOSS OF ENJOYMENT OF LIFE. | NO INFORMATION PROVIDED IN THE FOLLOWING SECTION(S): AGE OR DATE OF BIRTH, WEIGHT, ETHNICITY, DATE OF REPORT, RELEVANT TESTS/LABORATORY DATA, BRAND NAME, LOT #, DEVICE MANUFACTURE DATE, RECALL OR CORRECTION/REMOVAL NUMBER. | Y | Y | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00660 | 2022-06-09 | 2022-02-26 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | SQR DRV RTCHT WRNCH CXN | EXACTECH, INC. | 321-15-21 | 321-15-21 |  | IT WAS REPORTED THAT THE DEVICE HAS FRACTURED WHEN FORCE WAS APPLIED TO THE SNAP OFF SCREW. SURGEON REMOVED THE PARTS AND PIECES THAT FELL INTO THE PATIENT'S WOUND. THERE WAS NO SURGICAL DELAY/PROLONGATION. DEVICE IS RETURNING. | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00658 | 2022-06-09 | 2022-03-01 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA IMPLANT PSC INSERT, SZ 2.5, 15MM | KNEE FEMORAL COMPONENTS | EXACTECH, INC. | LOGIC TIBIA IMPLANT PSC INSERT, SZ 2.5, 15MM | 02-012-44-2515 |  | AS REPORTED BY THE LEGAL BRIEF, ON (B)(6) 2020, PATIENT UNDERWENT A TKA ON HER RIGHT KNEE AND WAS IMPLANTED WITH AN OPTETRAK TKA SYSTEM. PATIENT'S OPTETRAK TKA SYSTEM SUBSEQUENTLY FAILED IN (B)(6) 2022-LESS THAN TWO (2) YEARS AFTER IT WAS INITIALLY IMPLANTED. AS A RESULT, PATIENT REQUIRED AN ADDITIONAL SURGERY TO REMOVE AND REPLACE IT WITH ANOTHER JOINT REPLACEMENT SYSTEM. BECAUSE PATIENT REQUIRED AN ADDITIONAL REVISION SURGERY, PATIENT HAS SUFFERED SIGNIFICANT AND CONTINUING PERSONAL INJURIES, IS LIMITED IN HER ACTIVITIES OF DAILY LIVING, REQUIRES ADDITIONAL PHYSICAL THERAPY, AND HAS INCURRED SUBSTANTIAL MEDICAL BILLS AND EXPENSES. AS A DIRECT, PROXIMATE AND RESULT, PATIENT HAS SUFFERED AND CONTINUES TO SUFFER SIGNIFICANT, PERMANENT, AND CONTINUING PERSONAL INJURIES, AS DESCRIBED HEREIN. | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00657 | 2022-06-09 | 2022-05-23 | I, n, j, u, r, y | Manufacturer report | NOVATION GXL LNR +5 LAT 40MM GROUP 3 CUPS | HIP ACETABULAR COMPONENTS | EXACTECH, INC. | NOVATION GXL LNR +5 LAT 40MM GROUP 3 CUPS | 136-40-53 |  | AS REPORTED, PATIENT HAD A LEFT HIP REVISION DUE TO LOOSENING. ALL COMPONENTS WERE REMOVED AND REPLACED WITH COMPETITOR'S DEVICE. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. X-RAYS RECEIVED. THE DEVICE WILL BE RETURNED. | (H3) THE REVISION REPORTED WAS LIKELY THE RESULT OF INSUFFICIENT BOND BETWEEN THE FEMORAL STEM AND THE BONE AND PATIENT-RELATED CONDITIONS, WHICH LED TO ASEPTIC (NON-INFECTED) FEMORAL STEM LOOSENING. THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE REPORTED EVENT OF ¿LOOSENING¿ IS ASSOCIATED WITH WEAKENED INTEGRATION OF THE FEMORAL STEM AT THE BONE-IMPLANT INTERFACE DUE TO LOSS OF FIBROUS AND/OR BONY TISSUE AND LEADING TO COMPROMISED ANCHORAGE OF THE DEVICE. | Y | Y | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00656 | 2022-06-09 | 2022-05-19 | I, n, j, u, r, y | Manufacturer report | THREE PEG PATELLA 35MM | CR TIB INSERT SLOPE+, SZ 3, 9MM | EXACTECH, INC. | THREE PEG PATELLA 35MM | 200-02-35 |  | AS REPORTED. APPROXIMATELY 4.5 YEARS POST INITIAL RIGHT TKA, THE (B)(6) FEMALE PATIENT HAD A REVISION DUE TO DELAMINATING/PITTING AND WEAR THAT CAN BE SEEN. PATELLA WARN RIDING ON THE SHARP LATERAL EDGE OF TROC. PATIENT'S INSERT AND PATELLA REPLACED. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PHOTOS/X-RAYS RECEIVED. THE DEVICE WILL BE RETURN. | PENDING EVALUATION. NO INFORMATION PROVIDED IN THE FOLLOWING SECTION(S): WEIGHT, ETHNICITY, RELEVANT TESTS/LABORATORY DATA, OTHER RELEVANT HISTORY. | Y | Y | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00662 | 2022-06-09 | 2022-05-19 | I, n, j, u, r, y | Manufacturer report | TRULIANT POR TIB TRAY SIZE 5F/4T | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT POR TIB TRAY SIZE 5F/4T | 02-022-55-5040 |  | AS REPORTED, APPROXIMATELY 1.5 YEARS POST INITIAL LEFT TKA, THE 50 Y/O MALE PATIENT HAD A SUCCESSFULLY REVISION SURGERY DUE TO LOOSE POROUS TIBIA NEEDED TO BE REPLACED. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. NO ADDITIONAL INFORMATION AVAILABLE. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | CONCOMITANTS MEDICAL PRODUCTS TRULIANT TIB IMP CRC INSERT SZ 5, 9MM (CAT# 02-022-51-5009 / SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00669 | 2022-06-10 | 2021-11-16 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 3 9MM | KNEE FEMORAL COMPONENTS | EXACTECH, INC. | UNK | 02-012-35-3009 |  | AS REPORTED BY THE LEGAL BRIEF, APPROXIMATELY 10 YEARS POST INITIAL BILATERAL TKA, THE FEMALE PATIENT WAS DIAGNOSED WITH LOOSENING OF HER RIGHT TOTAL KNEE REPLACEMENT AND OSTEOLYSIS AND UNDERWENT REVISION SURGERY. DURING THE REVISION SURGERY, PATIENT'S SURGEON NOTED, "THE PATIENT WAS NOTED TO HAVE ASEPTIC LOOSENING OF THE FEMORAL COMPONENT AND MARKED OSTEOLYSIS," AND "THERE WAS NOTED TO BE WEAR OF THE LINER." AS A DIRECT, PROXIMATE RESULT, PATIENT HAS SUFFERED AND CONTINUES TO SUFFER PERMANENT AND DEBILITATING INJURES AND DAMAGES, INCLUDING BUT NOT LIMITED TO, SIGNIFICANT PAIN AND DISCOMFORT; GAIT IMPAIRMENT; POOR BALANCE; DIFFICULTY WALKING; COMPONENT PART LOOSENING; SOFT TISSUE DAMAGE; BONE LOSS; AND OTHER INJURIES PRESENTLY UNDIAGNOSED, WHICH ALL REQUIRE ONGOING MEDICAL CARE. AS OF NOTE: "LEFT KNEE REVISION PREVIOUSLY CAPTURED UNDER (B)(4)". | PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00670 | 2022-06-10 | 2021-01-19 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | CAGE GLENOID MEDIUM, ALPHA | EXACTECH, INC. | 314-13-03 | 314-13-03 |  | AS REPORTED BY THE EQUINOXE SHOULDER STUDY, APPROXIMATELY 3 YEARS POST OP THE INITIAL TSA, THIS 61 Y/O FEMALE PATIENT WAS REVISED DUE TO SEPTIC LOOSENING. PATIENT HAD CUFF FAILURE AND HUMERAL LOOSENING. THE CASE REPORT FORM INDICATES THIS EVENT IS UNLIKELY RELATED TO DEVICES AND DEFINITELY RELATED TO PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | DEVICE EVALUATED BY MFR: PENDING EVALUATION. CONCOMITANT MEDICAL DEVICE(S): 300-01-09 - EQUINOXE, HUMERAL STEM PRIMARY, PRESS FIT 9MM, 310-01-44 - EQUINOXE, HUMERAL HEAD SHORT, 44MM (ALPHA), 300-10-15 - EQUINOXE REPLICATOR PLATE 1.5MM O/S. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00671 | 2022-06-10 | 2020-08-27 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | SHOULDER COMPONENTS | EXACTECH, INC. | 314-13-13 | 314-13-13 |  | AS REPORTED BY AN EQUINOXE SHOULDER STUDY, APPROXIMATELY 8 YEARS POST OP THE INITIAL TSA, THIS 51 Y/O MALE PATIENT'S WAS REVISED DUE TO ASEPTIC GLENOID LOOSENING. PATIENT HAD A LOOSE CAGE GLENOID WITH SQUEAKING AND PAIN. THE CASE REPORT FORM INDICATES THIS EVENT IS POSSIBLY RELATED TO DEVICES AND DEFINITELY RELATED TO PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | DEVICE EVALUATED BY MFR: PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00672 | 2022-06-10 | 2022-05-23 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 42MM HUMERAL LINER +0 | EXACTECH, INC. | 320-42-00 | 320-42-00 |  | AS REPORTED, APPROXIMATELY 7 MONTHS POST OP THE INITIAL LEFT TSA, THIS 74 Y/O MALE PATIENT WAS REVISED DUE TO INFECTION. HE PRESENTED A MONTH AGO WITH AN INFECTED CYST ON THE INCISION LINE FROM SURGERY. THE SURGEON THOUGHT THAT MAYBE IT WAS FROM THE SUTURE. CULTURES WERE TAKEN AND THE INFECTED SITE TREATED WITH BETADINE. THE IMPLANTS WERE ALL SOLIDLY FIXED BUT THE SURGEON WANTED TO CHANGE THE LINER JUST TO BE SURE. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. DEVICES ARE NOT RETURNING. | DEVICE EVALUATED BY MFR: PENDING EVALUATION. CONCOMITANT DEVICE(S): 300-01-17, (B)(4) - EQUINOXE HUMERAL STEM PRIMARY PRESS FIT 17MM, 320-01-42, (B)(4) - EQUINOXE REVERSE 42MM GLENOSPHERE, 320-10-05, (B)(4) - EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +5, 320-15-03, (B)(4) - RS GLENOID PLATE POST AUG, 8 DEG, LEFT. 320-15-05, (B)(4) - EQ REV LOCKING SCREW. 320-20-00, (B)(4) - EQ REVERSE TORQUE DEFINING SCREW KIT. 320-20-30, (B)(4) - EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 30MM. 320-20-34, (B)(4) - EQ REV COMPRESS SCREW LCK CAP KIT, 4.5 X 34MM. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00673 | 2022-06-10 | 2020-08-26 | I, n, j, u, r, y | Manufacturer report | UNK | KNEE FEMORAL COMPONENTS | EXACTECH, INC. | UNK | UNK |  | AS REPORTED BY THE LEGAL DEPARTMENT, ON (B)(6) 2008, PATIENT UNDERWENT A TKA ON HER RIGHT KNEE AND WAS IMPLANTED WITH AN EXACTECH TKA SYSTEM. PATIENT¿S EXACTECH TKA SYSTEM SUBSEQUENTLY FAILED IN (B)(6) 2020. AS A RESULT, PATIENT REQUIRED AN ADDITIONAL SURGERY TO REMOVE THE IMPLANT AND REPLACE IT WITH ANOTHER JOINT REPLACEMENT SYSTEM. BECAUSE SHE REQUIRED AN ADDITIONAL REVISION SURGERY, PATIENT HAS SUFFERED SIGNIFICANT AND CONTINUING PERSONAL INJURIES, IS LIMITED IN HER ACTIVITIES OF DAILY LIVING, REQUIRES ADDITIONAL PHYSICAL THERAPY, AND HAS INCURRED SUBSTANTIAL MEDICAL BILLS AND EXPENSES. AS A DIRECT, PROXIMATE AND LEGAL CONSEQUENCE OF THE DEFECTIVE NATURE OF THE EXACTECH TKA SYSTEM, PATIENT HAS SUFFERED AND CONTINUES TO SUFFER SIGNIFICANT, PERMANENT, AND CONTINUING PERSONAL INJURIES, AS DESCRIBED HEREIN. | DEVICE EVALUATED BY MFR: PENDING EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00674 | 2022-06-10 | 2022-05-23 | I, n, j, u, r, y | Manufacturer report | LOGIC TIB INSERT IMPL CRC, SZ 3, 11MM | TIB INSERT IMPL CRC, SZ 3, 13MM | EXACTECH, INC. | LOGIC TIB INSERT IMPL CRC, SZ 3, 11MM | 02-012-51-3011 |  | AS REPORTED, THE 78 Y/O MALE PATIENT HAD A REVISION FOR PRE-MATURE WEAR ON THE POLY-ETHYLENE. REPLACED POLY WITH A NEW, PRE-PACKAGED POLY. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PHOTO RECEIVED. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | DEVICE EVALUATED BY MFR: PENDING EVALUATION. | Y | Y | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00678 | 2022-06-13 | 2021-03-15 | I, n, j, u, r, y | Manufacturer report | LOGIC FEMORAL PS | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | FEMORAL COMPONENTS | 02-010-01-0240 |  | AS REPORTED BY THE LEGAL BRIEF, ON (B)(6) 2011, PATIENT UNDERWENT BILATERAL KNEE REPLACEMENT SURGERY AND WAS IMPLANTED WITH OPTETRAK DEVICES. PATIENT DEVELOPED SEVERE PAIN IN HIS LEFT KNEE AND WAS ADVISED HE WOULD REQUIRE REVISION SURGERY DUE TO PREMATURE POLYETHYLENE WEAR OF THE TIBIAL INSERT. AS A RESULT, HE UNDERWENT REVISION SURGERY OF HIS LEFT KNEE ON (B)(6) 2021. FOLLOWING THE REVISION SURGERY OF HIS LEFT KNEE, PATIENT CONTINUES TO BE LIMITED IN HIS ACTIVITIES OF DAILY LIVING. DESPITE UNDERGOING REVISION SURGERY, PATIENT EXPERIENCES DAILY PAIN AND DISCOMFORT IN HIS LEFT KNEE WHICH LIMITS HIS ACTIVITIES OF DAILY LIVING AND IMPACTS HIS QUALITY OF LIFE. PATIENT ALSO CONTINUES TO BE LIMITED DUE TO ONGOING PAIN IN HIS RIGHT KNEE WHICH HAS NOT YET BEEN REVISED. PATIENT IS IN THE PROCESS OF SCHEDULING A REVISION SURGERY FOR HIS RIGHT KNEE. AS A DIRECT, PROXIMATE RESULTS, PATIENT HAS SUFFERED AND CONTINUES TO SUFFER PERMANENT AND DEBILITATING INJURES AND DAMAGES, INCLUDING BUT NOT LIMITED TO, SIGNIFICANT PAIN AND DISCOMFORT; GAIT IMPAIRMENT; POOR BALANCE; DIFFICULTY WALKING; COMPONENT PART LOOSENING; SOFT TISSUE DAMAGE; BONE LOSS; AND OTHER INJURIES PRESENTLY UNDIAGNOSED, WHICH ALL REQUIRE ONGOING MEDICAL CARE. AS A FURTHER DIRECT RESULT, PATIENT HAS SUSTAINED AND WILL SUSTAIN FUTURE DAMAGES, INCLUDING BUT NOT LIMITED TO COST OF MEDICAL CARE; REHABILITATION; HOME HEALTH CARE; LOSS OF EARNING CAPACITY; MENTAL AND EMOTIONAL DISTRESS AND PAIN AND SUFFERING. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00677 | 2022-06-13 | 2022-05-19 | I, n, j, u, r, y | Manufacturer report | TRULIANT PS POR FEM PS POR RIGHT SZ 3.5 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT PS POR FEM PS POR RIGHT SZ 3.5 | 02-020-12-0335 |  | AS REPORTED, APPROXIMATELY FIVE MONTHS POST INITIAL RIGHT TKA, THE 66 Y/O FEMALE PATIENT COMPLAINED OF PAIN AND HAD A REVISION DUE TO LOOSE FEMUR. SURGEON USED A DEVICE FROM ANOTHER COMPANY FOR REVISION. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. SALES REP WAS UNABLE TO OBTAIN PHOTOS/X-RAYS. THE DEVICE IS NOT AVAILABLE FOR EVALUATION AS HOSPITAL WILL NOT RELEASE THE IMPLANTS. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00661 | 2022-06-13 | 2022-03-23 | I, n, j, u, r, y | Manufacturer report | NV GXL LNR, LIPPED, 36MM ID, GROUP 3 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LNR, LIPPED, 36MM ID, GROUP 3 CUPS | 132-36-53 |  | AS REPORTED, PATIENT HAD A TOTAL HIP IMPLANT. COMPONENTS OF THIS HIP REPLACEMENT SYSTEM HAVE A RECALL NOTICE ADVISING PHYSICIANS OF EARLY WEAR AND TO EVALUATE PATIENTS WITH X-RAYS TO DETERMINE IF EARLY WEAR IS AFFECTING THE PATIENTS IMPLANT. IT IS, A TOTAL HIP REVISION IS TO BE CONSIDERED, THIS PATIENT HAD A TOTAL HIP REVISION DUE TO THE COMPONENTS OF THE HIP IMPLANTS PLACED INITIALLY IN 2019. THE EVENT WAS REPORTED DIRECTLY TO FDA UNDER MEDWATCH# MW5109455. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00675 | 2022-06-13 | 2022-03-09 | I, n, j, u, r, y | Manufacturer report | NV GXL LINER NEUTRAL, 32MM ID GROUP 1 CUPS | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | NV GXL LINER NEUTRAL, 32MM ID GROUP 1 CUPS | 130-32-51 |  | AS REPORTED, APPROXIMATELY SIX YEARS POST INITIAL LEFT THA, THE FEMALE PATIENT HAD A REVISION FOR AN UNKNOWN REASON. NO ADDITIONAL INFORMATION AVAILABLE. | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION RECEIVED. (H3) THE 72 YEAR OLD FEMALE PATIENT INVOLVED WAS REPORTED AS REVISION DUE TO EARLY PROSTHESIS WEAR AND OSTEOLYSIS APPROXIMATELY 4 YEARS AFTER THE INDEX SURGERY. THE PROSTHESIS WEAR AND OSTEOLYSIS WAS ABLE TO BE CONFIRMED FROM THE PROVIDED RADIOGRAPH. BASED ON THE AVAILABLE INFORMATION, THE PATIENT INVOLVED MEETS THE FOLLOWING RISK CRITERIA FOR EARLY PROSTHESIS WEAR/OSTEOLYSIS, COMBINATION OF LARGEST AVAILABLE FEMORAL HEAD AND/OR THINNEST AVAILABLE ACETABULAR LINER WAS USED. THE MOST LIKELY CAUSE FOR THE REVISION REPORTED DUE TO EARLY PROSTHESIS WEAR/OSTEOLYSIS IN IS A COMBINATION OF THE RISK FACTORS INCLUDING, USE ERROR, IMPLANT POSITIONING, IMPLANT SIZE SELECTION, AND PATIENT FACTORS (FITNESS FOR SURGERY, BIOMECHANICS, ACTIVITY LEVEL AND LOCAL TISSUE OXIDATION POTENTIAL COULD HAVE ALL CONTRIBUTED TO THE INCREASED TREND IN WEAR/OSTEOLYSIS RELATED COMPLAINTS). SECTION (H9): Z-2112-2021; Z-2113-2021; Z-2114-2021; Z-2115-2021; Z-2116-2021; Z-2117-2021; Z-2118-2021; Z-2119-2021; Z-2120-2021; Z-2121-2021; Z-2122-2021; Z-2123-2021; Z-2124-2021; Z-2125-2021; Z-2126-2021; Z-2127-2021; Z-2128-2021; Z-2129-2021; Z-2130 THRU 33-2021. | Y | Y | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00676 | 2022-06-13 | 2022-05-24 | I, n, j, u, r, y | Manufacturer report | TRULIANT PS POR FEM PS POR RIGHT SZ 4 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT PS POR FEM PS POR RIGHT SZ 4 | 02-020-12-0340 |  | AS REPORTED, APPROXIMATELY EIGHT MONTHS POST INITIAL RIGHT TKA, THE 38 Y/O FEMALE PATIENT COMPLAINED OF KNEE PAIN AND HAD A POLY EXCHANGE. PATIENT WAS REVISED WITH A DEVICE FROM ANOTHER COMPANY. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. X-RAY RECEIVED. THE DEVICE IS NOT AVAILABLE FOR EVALUATION DUE TO HOSPITAL WILL NOT RELEASE IMPLANT. | ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00684 | 2022-06-14 | 2017-09-09 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | REVERSE 38MM GLENOSPHERE | EXACTECH, INC. | 320-01-38 | 320-01-38 |  | AS REPORTED, APPROXIMATELY 3 MONTHS POST OP THE INITIAL TSA, THIS 66 Y/O MALE PATIENT WAS REVISED DUE TO DISLOCATION. THE PROSTHESIS DISLOCATED WHEN A WEIGHT WAS LIFTED. THE CASE REPORT FORM INDICATES THIS EVENT IS DEFINITELY NOT RELATED TO DEVICES OR PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | SECTION H10: (H3) BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN OR MANUFACTURING ISSUES. THE CAUSE OF THE SURGICAL REVISION FOR DISLOCATION CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION AND ACTIVITY. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00679 | 2022-06-14 | 2022-05-23 | I, n, j, u, r, y | Manufacturer report | BIOLOX DELTA FEMORAL HEAD 36MM OD, +7MM | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | EXACTECH, INC. | BIOLOX DELTA FEMORAL HEAD 36MM OD, +7MM | 170-36-07 |  | AS REPORTED, APPROXIMATELY TWENTY TWO DAYS AFTER INITIAL THA, THE FEMALE PATIENT WAS BROUGHT BACK FOR A LESSER TROCHANTERIC FRACTURE. THE HEAD WAS REMOVED, THE STEM REMOVED. THE FRACTURE WAS REDUCED AND CABLES. A SIZE 10 HPS STEM WAS CEMENTED ALONG WITH A NEW BIOLOX HEAD. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. X-RAY RECEIVED. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | CONCOMITANTS MEDICAL PRODUCTS: SIZE 7 STD TAPER WEDGE STEM (CAT# 188-07-01, (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00680 | 2022-06-14 | 2022-05-24 | I, n, j, u, r, y | Manufacturer report | LOGIC FEMORAL PS CEM LEFT SZ 2 | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC FEMORAL PS CEM LEFT SZ 2 | 02-010-01-0220 |  | AS REPORTED, APPROXIMATELY SEVEN YEARS POST INITIAL LEFT TKA, THE 65 Y/O FEMALE PATIENT STARTED HAVING PAIN. SURGEON HAD REMOVED A LOOSE FEMORAL IMPLANT THAT NEED REVISED TO A SIZE 2 CONDYLAR CONSTRAINT FEMUR, ALONG WITH REPLACING THE POLY TO A 17MM CONSTRAIN LINER. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. X-RAYS AND PHOTOS RECEIVED. THE DEVICES ARE AVAILABLE FOR EVALUATION. | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION RECEIVED THE FOLLOWING SECTIONS G3, G6, H2, H3 AND H6 HAVE BEEN UPDATED ACCORDINGLY. (H3) THE REVISION REPORTED WAS LIKELY THE RESULT OF BOTH PATIENT-RELATED CONDITIONS AND AN INSUFFICIENT BOND BETWEEN THE FEMORAL COMPONENT AND THE BONE, WHICH LED TO ASEPTIC (NON-INFECTED) FEMORAL LOOSENING. THE LOOSENING LIKELY LED TO GROSS INSTABILITY AND ALLOWED THE FEMORAL COMPONENT TO MIGRATE INTO FLEXION AND INTERNALLY ROTATE, WHICH CAUSED AN INCREASE IN CEMENT AND BONE PARTICLES IN THE JOINT SPACE AND RESULTED IN TIBIAL INSERT WEAR (ARTICULATING SURFACE AND BACKSIDE) AND TIBIAL INSERT SUBSIDENCE. THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE REPORTED EVENT OF ¿LOOSENING - FEMORAL¿ IS ASSOCIATED WITH WEAKENED INTEGRATION OF THE FEMORAL COMPONENT AT THE BONE-IMPLANT INTERFACE DUE TO LOSS OF FIBROUS AND/OR BONY TISSUE AND LEADING TO COMPROMISED ANCHORAGE OF THE DEVICE. FURTHERMORE, THE MOST PROBABLE ROOT CAUSE ASSOCIATED WITH THE EVENT OF ¿PROSTHESIS WEAR¿ IS ASSOCIATED WITH MATERIAL DAMAGE TO A SURFACE, USUALLY INVOLVING PROGRESSIVE LOSS OR DISPLACEMENT OF MATERIAL, DUE TO RELATIVE MOTION BETWEEN THAT SURFACE AND A CONTACTING SUBSTANCE OR SUBSTANCES. | Y | Y | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00681 | 2022-06-14 | 2022-06-25 | I, n, j, u, r, y | Manufacturer report | TRULIANT TIB IMP CRC INSERT SZ 3.5, 9MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | TRULIANT TIB IMP CRC INSERT SZ 3.5, 9MM | 02-022-51-3509 |  | AS REPORTED, APPROXIMATELY THREE MONTHS POST INITIAL LEFT TKA, THE 54 Y/O PATIENT HAD A REVISION. THE POLY WAS EXCHANGED FOR A NEW ONE OF THE SAME SIZE DURING A SYNOVECTOMY. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. SALES REP WAS UNABLE TO OBTAIN PHOTOS/X-RAYS. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION RECEIVED THE FOLLOWING SECTIONS G3, G6, H2, H3 AND H6 HAVE BEEN UPDATED ACCORDINGLY. (H3) AS REPORTED, APPROXIMATELY THREE MONTHS POST INITIAL LEFT TKA, THE 54 Y/O PATIENT HAD A REVISION. THE POLY WAS EXCHANGED FOR A NEW ONE OF THE SAME SIZE DURING A SYNOVECTOMY. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. SALES REP WAS UNABLE TO OBTAIN PHOTOS/X-RAYS. THE DEVICE IS NOT AVAILABLE FOR EVALUATION. NO ADDITIONAL INFORMATION AVAILABLE. BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN OR MANUFACTURING ISSUES. THE CAUSE OF THE SUBSEQUENT REVISION CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION. THESE DEVICES ARE USED FOR TREATMENT NOT DIAGNOSIS. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00682 | 2022-06-14 | 2022-05-25 | I, n, j, u, r, y | Manufacturer report | THREE PEG PATELLA 38MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | THREE PEG PATELLA 38MM | 200-02-38 |  | AS REPORTED, APPROXIMATELY SEVEN YEARS POST INITIAL LEFT TKA, THE 78 Y/O MALE PATIENT HAS A SWOLLEN KNEE AND DURING THE CHECK UP IT WAS FOUND THAT THE PATELLA WAS LOOSE. THE PATELLA AND LINER WERE REPLACED. THERE WAS NO BREAKAGE OF DEVICE OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. PHOTOS AND X-RAYS RECEIVED. THE DEVICE WILL BE RETURN. | CONCOMITANT MEDICAL PRODUCTS: LOGIC TIBIAL INSERT PS, SIZE 5, 11MM (CAT# 02-012-35-5011 /SERIAL# (B)(4)). ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. | Y | Y | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00634 | 2022-06-14 | 2022-05-16 | I, n, j, u, r, y | Manufacturer report | LOGIC TIBIA PS MOD INSRT SZ 3.5 11MM | PROSTHESIS, KNEE, PATELLOFEMOROTIBIAL, SEMI-CONSTRAINED, CEMENTED | EXACTECH, INC. | LOGIC TIBIA PS MOD INSRT SZ 3.5 11MM | 02-012-35-3511 |  | AS REPORTED, APPROXIMATELY THREE YEARS POST INITIAL RIGHT THA, THE 65 Y/O MALE PATIENT HAD A REVISION DUE TO POLY HAS SIGNIFICANT WEAR. PATIENT HAD A POLY EXCHANGE AND FEMUR REPLACEMENT. THERE WAS NO BREAKAGE OF DEVICES OR SURGICAL DELAY/PROLONGATION. PATIENT WAS LAST KNOWN TO BE IN STABLE CONDITION FOLLOWING THE EVENT. SALES REP WAS UNABLE TO OBTAIN PHOTOS/X-RAYS. THE DEVICE WILL BE RETURN. | AFTER FURTHER REVIEW OF ADDITIONAL INFORMATION RECEIVED THE FOLLOWING SECTIONS D1, D4, D9, G3, G4, G6, H2 AND H3 HAVE BEEN UPDATED ACCORDINGLY. ADDITIONAL INFORMATION, INCLUDING THE PRODUCT INVESTIGATION, WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT SECTION D1: PRODUCT CODE. | Y | Y | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00683 | 2022-06-14 | 2021-01-19 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | CAGE GLENOID MEDIUM, ALPHA | EXACTECH, INC. | 314-13-03 | 314-13-03 |  | AS REPORTED BY THE EQUINOXE SHOULDER STUDY, APPROXIMATELY 2.5 YEARS POST OP THE INITIAL TSA, THIS 61 Y/O FEMALE PATIENT WITH A BMI OF 38 WAS REVISED DUE TO SEPTIC LOOSENING. PATIENT HAD CUFF FAILURE AND HUMERAL LOOSENING. THE CASE REPORT FORM INDICATES THIS EVENT IS UNLIKELY RELATED TO DEVICES AND DEFINITELY RELATED TO PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | SECTION H10: (H3) BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN OR MANUFACTURING ISSUES. THE CAUSE OF SURGICAL REVISION FOR INFECTION/SEPTIC LOOSENING CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION. THESE DEVICES ARE USED FOR TREATMENT NOT DIAGNOSIS. SECTION H11: *THE FOLLOWING SECTIONS HAVE CORRECTED INFORMATION: (B5) DESCRIBE EVENT OR PROBLEM: AS REPORTED BY THE EQUINOXE SHOULDER STUDY, APPROXIMATELY 2.5 YEARS POST OP THE INITIAL TSA, THIS 61 Y/O FEMALE PATIENT WITH A BMI OF 38 WAS REVISED DUE TO SEPTIC LOOSENING. PATIENT HAD CUFF FAILURE AND HUMERAL LOOSENING. THE CASE REPORT FORM INDICATES THIS EVENT IS UNLIKELY RELATED TO DEVICES AND DEFINITELY RELATED TO PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. | Y | N | HEALTH PROFESSIONAL |  | 2022-06
1038671-2022-00685 | 2022-06-14 | 2022-05-11 | I, n, j, u, r, y | Manufacturer report | EQUINOXE | SHOULDER COMPONENTS | EXACTECH, INC. | UNK | UNK |  | AS REPORTED, APPROXIMATELY 7 YEARS POST OP THE INITIAL TSA, THIS 78 Y/O MALE PATIENT WAS REVISED DUE TO SEPTIC LOOSENING. SUBJECT STARTING HAVING SEVERE LEFT SHOULDER PAIN AFTER UNRELATED HOSPITALIZATION-EXACT MOI/DATE UNK. THE CASE REPORT FORM INDICATES THIS EVENT IS UNLIKELY RELATED TO DEVICES OR PROCEDURE. THIS EVENT REPORT WAS RECEIVED THROUGH CLINICAL DATA COLLECTION ACTIVITIES. OUTCOME IS CONTINUING. | SECTION H10: (H3) BASED ON REVIEW OF ALL AVAILABLE INFORMATION, THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED EVENT IS RELATED TO ANY DESIGN OR MANUFACTURING ISSUES. THE CAUSE OF SURGICAL REVISION FOR INFECTION CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS MOST LIKELY RELATED TO THE PATIENT¿S UNDERLYING CONDITION. | Y | N | HEALTH PROFESSIONAL |  | 2022-06

## Monthly Summary
Month | Count | Date
2021-12 | 53 | 2021-12-01
2022-01 | 95 | 2022-01-01
2022-02 | 146 | 2022-02-01
2022-03 | 144 | 2022-03-01
2022-04 | 225 | 2022-04-01
2022-05 | 173 | 2022-05-01
2022-06 | 60 | 2022-06-01

## Event Types
Event Type | Count
I, n, j, u, r, y | 634
M, a, l, f, u, n, c, t, i, o, n | 262

## Brand Names
Brand Name | Count
EXACTECH | 307
EQUINOXE | 144
OPTETRAK | 22
FEMORAL COMPONENTS | 21
VANTAGE | 17
LOGIC TIBIA PS MOD INSRT SZ 5 9MM | 10
NV GXL LINER NEUTRAL, 32MM ID GROUP 1 CUPS | 10
SPECIFIC DEVICE NOT REPORTED | 8
LOGIC TIBIA PS MOD INSRT SZ 4 11MM | 7
LOGIC TIBIA PS MOD INSRT SZ 3 9MM | 5
LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 9MM | 5
LOGIC TIBIA PS MOD INSRT SZ 4 9MM | 5
TRULIANT TIB IMP PS INSERT SZ 4 9MM | 4
NV GXL LINER NEUTRAL, 36MM ID, GROUP 2 CUPS | 4
PS TIBIAL INSERTS SZ 4, 11MM | 4
LOGIC TIBIA PS MOD INSRT SZ 3 11MM | 4
LOGIC TIBIA PS MOD INSRT SZ 4 13MM | 4
LOGIC FEMORAL PS CEM RIGHT SZ 2.5 | 4
LOGIC TIBIA IMPLANT PSC INSERT, SZ 2, 9MM | 4
LOGIC CR TIB INSERT STD, SZ 2.5, 9 MM | 4
LOGIC TIBIA IMPLANT PSC INSERT, SZ 2, 11MM | 4
PS TIBIAL INSERTS SZ 3, 9MM | 4
LOGIC TIBIA PS MOD INSRT SZ 2 9MM | 3
CC TIBIAL INSERT SZ 2, 13MM | 3
LOGIC TIBIA PS MOD INSRT SZ 2 11MM | 3
LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 13MM | 3
THREE PEG PATELLA 32MM | 3
NV GXL LNR, +5LAT, 32MM G1-48/50MM CUPS | 3
LOGIC TIBIA IMPLANT PSC INSERT, SZ 5, 11MM | 3
LOGIC TIBIA IMPLANT PSC INSERT, SZ 2.5, 15MM | 3
LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 11MM | 3
LOGIC | 3
PS TIBIAL INSERTS SZ 4, 9MM | 3
LOGIC TIBIA IMPLANT PSC INSERT, SZ 4, 15MM | 3
NV GXL LINR, NTRL, 36MM ID, GROUP 3 CUPS | 3
LOGIC FEMORAL PS CEM LEFT SZ 2 | 3
NV GXL LINR, NTRL, 32MM ID, GROUP 2 CUPS | 3
LOGIC CR TIB INSERT SLOPE++, SZ 2, 13MM | 3
LOGIC TIBIA PS MOD INSRT SZ 5 11MM | 3
LOGIC CR TIB INSERT SLOPE+, SZ 3, 9MM | 3
ACETABULAR COMPONENTS | 3
LOGIC FEMORAL PS CEM LEFT SZ 3 | 3
NV GXL LINR, NTRL, 36MM ID, GROUP 4 CUPS | 2
EQUINOX | 2
LOGIC TIBIA IMPLANT PSC INSERT, SZ 5, 9MM | 2
LOCKING FEMORAL IMPACTOR 0-5 | 2
THREE PEG PATELLA 35MM | 2
LOGIC FEMORAL PS CEM RIGHT SZ 3.5 | 2
NV GXL LNR, LIPPED, 36MM ID, GROUP 3 CUPS | 2
NV GXL LNR, +5LAT, 36MM G2-52/54MM CUPS | 2
TRULIANT TIB IMP CR INS STD SZ 3, 9MM | 2
LOGIC TIBIA IMPLANT PSC INSERT, SZ 3, 13MM | 2
NV GXL LINER LIPPED 32MM ID, GROUP 1 CUPS | 2
NV GXL LINER LIPPED 36MM ID GROUP 2 | 2
EQUINOXE REVERSE 46MM HUMERAL LINER +0 | 2
NV GXL LNR, LIPPED, 28MM ID, GROUP 1 CUPS | 2
LOGIC CR TIB INSERT STD, SZ 4, 11MM | 2
LOGIC TIB INSERT IMPL CRC, SZ 5, 9MM | 2
TRULIANT PS CEM FEM PS CEM LEFT SZ 4.5 | 2
LOGIC FEMORAL PS CEM LEFT SZ 2.5 | 2
TRULIANT PS CEM FEM PS CEM RIGHT SZ 5 | 2
THREE PEG PATELLA 38MM | 2
NOVATION GXL LNR +5 LAT 40MM GROUP 3 CUPS | 2
FEMORAL COMPONENT | 2
CC TIBIAL INSERT SZ 4, 13MM | 2
LOGIC CR TIB INSERT SLOPE++, SZ 3, 13MM | 2
OPTETRAK LOGIC | 2
TRULIANT PIN PULLER SYRINGE STYLE | 2
LOGIC TIBIA IMPLANT PSC INSERT, SZ 2, 13MM | 2
LOGIC TIBIA PS MOD INSRT SZ 3 15MM | 2
TRAPEZOID TIBIAL TRAY SZ 5F/4T | 2
TRULIANT TIB IMP CRC INSERT SZ 1.5, 11MM | 1
LOGIC CC FEMORAL SIZE 2, RIGHT | 1
NOVATION STEM | 1
NV GXL LNR, LIPPED 32MM ID GROUP 2 CUPS | 1
LOGIC CR TIB INSERT SLOPE++, SZ 5, 11MM | 1
ALT HA S CLR STD SZ 3 | 1
2.5, 13MM SLOPE ++ CR INSERT | 1
TRULIANT CR CEM FEM CR CEM RIGHT SZ 3 | 1
PS TIBIAL INSERTS SZ 2, 18MM | 1
LOGIC CR TIB INSERT STD, SZ 4, 9MM | 1
CC TIBIAL INSERT SZ 2, 15MM | 1
TRULIANT TIB IMP CR INS STD SZ 1.5, 9 MM | 1
LOGIC TIB INSERT IMPL CRC, SZ 2, 13MM | 1
EXACTECH CONNEXION GXL ACETABULAR POLYETHYLENE LINERS | 1
PS TIBIAL INSERTS SZ 4, 15MM | 1
TRULIANT TIB IMP PS INSERT SZ 6 12MM | 1
STEM EXTENSION 22MM X 80 L. | 1
LOGIC TIBIA IMPLANT PSC INSERT, SZ 1, 11MM | 1
LOGIC CR TIB INSERT SLOPE +, SZ 3.5, 11 MM | 1
LOGIC CR TIB INSERT SLOPE+, SZ 5, 9MM | 1
LOGIC CR TIB INSERT SLOPE++, SZ 4, 9MM | 1
BIOLOX DELTA FEMORAL HEAD 36MM OD, -3.5MM | 1
LOGIC TIBIA PS MOD INSRT SZ 3.5 9MM | 1
LOGIC CR TIB INSERT SLOPE +, SZ 1.5, 9 MM | 1
TIBIAL INSERT TRIAL SZ 3 13 MM | 1
EQUINOXE REVERSE TRAY ADAPTER PLATE TRAY +0 | 1
+5MM 36MM LATERALIZED GXL LINER | 1
ACUMATCH GXL 15DEG LINER 28MM SZ G | 1
TRULIANT PS CEM FEM PS CEM RIGHT SZ 4 | 1
TRULIANT PS CEM FEM PS CEM RIGHT SZ 6 | 1
LOGIC CR TIB INSERT STD, SZ 2, 9MM | 1
BIOLOX DELTA FEMORAL HEAD 36MM OD, +3.5MM | 1
TRULIANT TIB IMP PSC INSERT SZ 4, 9MM | 1
PS TIBIAL INSERTS SZ 4, 18MM | 1
RBK PS TIBIAL INSERT SZ 3, 09MM | 1
LOGIC FEMORAL PS CEM LEFT SZ 4 | 1
ALT XLE LNR 10° FC G4 36 | 1
NV CROWN CUP CLSTR HOLE 56MM GROUP 3 | 1
TRULIANT CR CEM FEM CR CEM RIGHT SZ 4.5 | 1
OPTETRAK SIZE 4, 11MM POLY | 1
TRULIANT PS CEM FEM PS CEM LEFT SZ 5 | 1
NV GXL LINER NEUTRAL, 40MM ID, GROUP 3 CUPS | 1
LOGIC CC TIB INSERT SIZE 3, 13MM | 1
TRULIANT TIB IMP PS INSERT SZ 3 13MM | 1
EUINOXE | 1
ADVANCED PATELLA 32MM 3 PEG IMPLANT | 1
TRULIANT TIB FIT TRAY CEM SZ 3F / 3T | 1
LOGIC TIBIA IMPLANT PSC INSERT, SZ 1, 13MM | 1
LOGIC CR TIB INSERT SLOPE++, SZ 4, 11MM | 1
LOGIC TIBIA IMPLANT PSC INSERT | 1
LOGIC FEMORAL PS CEM RIGHT SZ 2 | 1
TRULIANT POR TIB TRAY SIZE 5F/4T | 1
UNK | 1
LOGIC TIB INSERT IMPL CRC, SZ 3, 11MM | 1
LOGIC FEMORAL PS | 1
TRULIANT PS POR FEM PS POR RIGHT SZ 3.5 | 1
TRULIANT PS POR FEM PS POR RIGHT SZ 4 | 1
BIOLOX DELTA FEMORAL HEAD 36MM OD, +7MM | 1
TRULIANT TIB IMP CRC INSERT SZ 3.5, 9MM | 1
ELEMENT-STEM, COLLARED W/HA, STD OFFSET, SZ 8 | 1
TRULIANT TIB IMP CR STD, SIZE 3, 10MM | 1
NV GXL LNR, +5LAT, 36MM G4-60/62MM CUPS | 1
NV GXL LNR, +5 LAT 32MM ID, GROUP 3 CUPS | 1
PS TIBIAL INSERTS SZ 4, 13MM | 1
TRULIANT TIB IMP PSC INSERT SZ 4, 13MM | 1
NOVATION GXL LNR +5 LAT 40MM ID GROUP 4 CUPS | 1
LOGIC CR TIB INSERT SLOPE+, SZ 3, 11MM | 1
TRULIANT TIB IMP PS INSERT SZ 3 9MM | 1
FEMORAL LINER | 1
CEMENTED FINNED TIB. TRA SZ 2F/2T | 1
LOGIC TIBIA PS MOD INSRT SZ 4 15MM | 1
EXACTECH, | 1
EXACTEC | 1
OPTETRACK PS TIBIAL INSERT SIZE 3 13MM | 1
EQUINOXE CAGE GLENOID MEDIUM, ALPHA | 1
LOGIC CR TIB INSERT SLOPE+, SZ 5, 13MM | 1
EXACTECJ | 1
TRULIANT TIB IMP CRC INSERT SZ 2.5,10MM | 1
NV GXL LNR, NEUTRAL, 32MM ID, GROUP 3 CUPS | 1
CC TIBIAL INSERT SZ 4, 15MM | 1
SMALL SUPERIOR/POSTERIOR AUG GLENOID PLATE,LEFT | 1
LOGIC TIBIA PS MOD INSRT SZ 5 15MM | 1
ALT HA S CLR EXT SZ 7 | 1
TRULIANT TIB IMP PS INSERT SZ 2 9MM | 1
LOGIC CR TIB INSERT STD, SZ 5, 9MM | 1
TRULIANT TIB IMP PSC INSERT SZ 1.5, 15MM | 1
PATELLA | 1
LOGIC CR TIB INSERT SLOPE ++, SZ 3.5, 11 MM | 1
ALT XLE LNR EXTCOV G7 40 | 1
PS TIBIAL INSERTS SZ 2, 13MM | 1
LOGIC CC TIB INSERT SIZE 2, 13MM | 1
CEMENTED CR FEMORAL | 1
PS TIBIAL INSERTS SZ 6, 15MM | 1
LOGIC TIBIA IMPLANT PSC INSERT, SZ 3, 11MM | 1
CC TIBIAL INSERT SZ 2, 9MM | 1
LOGIC TIBIA IMPLANT PSC INSERT, SZ 3.5, 15MM | 1
TRULIANT TIB IMP PS INSERT SZ 5 11MM | 1
LOGIC TIBIA PS MOD INSRT SZ 5 13MM | 1
LOGIC TIBIA PS MOD INSRT SZ 3.5 13MM | 1
TRULIANT CR CEM FEM CR CEM RIGHT SZ 2.5 | 1
LGC TIBIAL FIT TRAY CEM SZ 4F / 4T | 1
CC TIBIAL INSERT SZ 4, 9MM | 1
CC TIBIAL INSERT SZ 5, 11MM | 1
TRULIANT TIB IMP PS INSERT SZ 4 11MM | 1
PS TIBIAL INSERTS SZ 6, 13MM | 1
OPTETRAK HI-FLEX TIBIAL INSERT SZ 5 9MM | 1
OPTERTRAK LOGIC TIBIA IMPLANT PSC INSERT, SZ 5, 11MM | 1
LOGIC TIBIA IMPLANT PSC INSERT, SZ 3, 15MM | 1
LOGIC TIBIA IMPLANT PSC INSERT, SZ 3, 9MM | 1
LOGIC TIBIA PS MOD INSRT SZ 3 13MM | 1
CC TIBIAL INSERT SZ 3, 11MM | 1
CR TIBIAL INSERT SZ 3, 9MM, SLOPE + | 1
LOGIC CR TIB INSERT SLOPE+, SZ 4, 9MM | 1
LOGIC CC FEMORAL SIZE 3, LEFT | 1
CC TIBIAL INSERT SZ 2, 18MM | 1
LOGIC FEMORAL PS CEM RIGHT SZ 1.5 | 1
TRULIANT TIB IMP PS INSERT SZ 3 10MM | 1
OPTETRAK ASY,PS CEMENTED FEMORAL, SZ 3, | 1
PS TIBIAL INSERTS SZ 2, 9MM | 1
TRULIANT TIB IMP PS INSERT SZ 4.5 11MM | 1
LOGIC FEMORAL PS CEM LEFT SZ 3.5 | 1
PS TIBIAL INSERTS SZ 3, 11MM | 1
PS TIBIAL INSERTS SZ 3, 13MM | 1
CR TIBIAL INSERT SZ4, 13MM | 1
LOGIC FEMORAL PS CEM RIGHT SZ 4 | 1
EQ REV LOCKING SCREW | 1
NOV 6.5MM HEX DRIVER ARTICULATING | 1
TRULIANT PS POR FEM PS POR LEFT SZ 2.5 | 1
OPTETRACK CR | 1
TRULIANT CR CEM FEM CR CEM LEFT SZ 4 | 1
SHLDR GPS IMPACT HANDLE | 1
A+ BROACH HANDLE RIGHT | 1
GXL POLYETHYLENE LINER 32MM DIAMETER HEAD | 1
TRULIANT TIB IMP CR INS STD SZ 5, 9MM | 1
LOGIC CR TIB INSERT SLOPE++, SZ 1, 13MM | 1
TRULIANT TIB IMP CRC INSERT SZ 3.5, 17MM | 1
R3 52 MM ACETABULAR COMPONENT | 1
M-SERIES STEM | 1
STEM EXTENSION W/SLOT 120L X18 MM | 1
CC FEMORAL SZ 4 | 1
LOGIC RBK INSERT SZ 3.5, 9MM | 1
ALT HA S CLR EXT SZ 8 | 1
MED, LEFT, 16 POST AUG UHMWPE GLENOID | 1
LOGIC CR TIB INSERT SLOPE++, SZ 3, 11MM | 1
2.5 TRULIANT CRC POLYETHYLEN | 1
CR TIBIAL INSERT SZ 2, 11MM, SLOPE ++ | 1
LOGIC TIB INSERT IMPL CRC, SZ 2.5, 9MM | 1
TRULIANT TIB IMP PS INSERT SZ 4.5 9MM | 1
TIBIAL INSERT SIZE 3 9++ | 1
TRULIANT PS CEM FEM PS CEM RIGHT SZ 3 | 1
TRULIANT PS CEM FEM PS CEM LEFT SZ 3.5 | 1
TRULIANT CR CEM FEM CR CEM LEFT SZ 4.5 | 1
TRULIANT TIB IMP PSC INSERT SZ 4, 12MM | 1
ACETABULAR COMPOENTS | 1
TRU DISTAL FEM CUT BLOCK | 1
NV CROWN CUP CLSTR HOLE 48MM GROUP 1 | 1
COCR FEM HEAD 32MM +7 OFFSET 12/14 | 1
TRULIANT PS CEM FEM PS CEM RIGHT SZ 4.5 | 1
LOGIC FEMORAL PS CEM LEFT SZ 5 | 1
LOGIC TIB INSERT IMPL CRC, SZ 4, 9MM | 1
NV GXL LNR, 10 DEG FACE, 32MM ID, GRP 1 CUP | 1
LOGIC TIBIA PS MOD INSRT SZ 3.5 15MM | 1
LOGIC TIBIA IMPLANT PSC INSERT, SZ 2.5, 11MM | 1
ACUMATCH GXL 15DEG LINER 32MM SZ H | 1
CR TIBIAL INSERT SZ 5, 13MM, SLOPE ++ | 1
NV GXL LNR, +5LAT, 36MM G3-56/58MM CUPS | 1
TRULIANT PS POR FEM PS POR LEFT SZ 5 | 1
GXL NUETRAL LINER, G0 28MM ID | 1
ACUMATCH 0 DEGREE LINER 28MM 54/56MM | 1
OPTETRAK ASY HI-FLEX PS CEM FEM, SZ 3, RIGHT | 1
TRULIANT TIBIAL TRAY | 1
BIOLOX DELTA FEMORAL HEAD 36MM OD, +0MM | 1
OPTETRAK ASY HI-FLEX PS CEM FEM, SZ 2, RIGHT | 1
TIBIAL INSERT TRIAL SZ 5 15 MM | 1
ERGO BROACH HANDLE | 1
NV GXL LNR, LIPPED, 36MM ID GROUP 4 CUPS | 1
LOGIC TIBIA PS MOD INSRT SZ 3.5 11MM | 1

## Case Metadata
Case Name | MDL Number | MDL Filing Date | Court | MAUDE Start Date | MAUDE End Date | Total Records | Date Range
Exactech Joint Implants | 3044 | 2022-06-14 | E.D. New York | 2021-12-14 | 2022-06-14 | 896 | 2021-12-14 00:00:00 to 2022-06-14 00:00:00