2518422-2021-04636 | 2021-09-29 | 2021-08-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04407 | 2021-09-29 | 2021-08-30 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SKIN CANCER. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5104293 | 2021-09-29 | 2021-06-06 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I HAD BEEN SUFFERING FROM SEVERE SHORTNESS OF BREATH, ASTHMA, FATIGUE, HEADACHES, CHEST PRESSURE FOR OVER A YEAR AND THEN ON (B)(6) 2021 DIAGNOSED WITH DVT. ON (B)(6) 2021 WAS HOSPITALIZED AND PULMONARY EMBOLISM FOUND ON CTA OF CHEST. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5104302 | 2021-09-29 | 2020-05-15 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS DREAMSTATION BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS INC |  |  |  | CALLER HAS BEEN EXPERIENCING PERSISTENT COUGH AND SORE THROAT FOR SOMETIME NOW DUE TO A FAULTY DEVICE. CALLER ADVISED THE HUMIDIFICATION NO LONGER WORKS AND THE MANUFACTURER HAS ALSO ISSUED A RECALL NOTICE. CALLER WANTS A REPLACEMENT AS HE NEEDS THE DEVICE TO SLEEP. CALLER ADVISED THE MANUFACTURER HAS NOT RESPONDED TO HIM AND WANTS THE FDA TO EVALUATE AND FIND A RESOLUTION FOR HIM AND THE OTHER USERS OF THE DEVICE WHO NEED THIS DEVICE TO SLEEP. |  | N | Y | I |  | 2021-09
2518422-2021-04614 | 2021-09-29 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO SERIES ASSY, INT | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | 751P | 751P |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04403 | 2021-09-29 | 2021-08-12 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/CELL, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | CAX400H12C | CAX400H12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP BRAIN CANCER AND A GROWTH IN THEIR KIDNEYS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04699 | 2021-09-29 | 2021-07-22 | I, n, j, u, r, y | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED RESPIRATORY DISTRESS AND CARBON DIOXIDE POISONING. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF HOSPITALIZATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE /RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM THAT BECAME DEGRADED AND CAUSED THE PATIENT TO HOSPITALIZE DUE TO CARBONDIOXIDE POISONING AND RESPIRATORY DISTRESS. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER INSPECTED THE EXTERNAL AND INTERNAL PARTS OF THE DEVICE AND FOUND PRESENCE OF SMALL AMOUNTS OF DIRT, DUST, GRIM, AND OTHER FOREIGN PARTICLES NORMALLY FOUND INSIDE AIR/FLOW PATH OVERTIME. CONTAMINATION CAN BE FOUND FROM THE INLET PORT, ON/INTO AIRPATH, THEN BLOWER MOTOR BELLOWS, BLOWER MOTOR IMPELLER BLADES, INTO THE BLOWER MOTOR, TRANSITION TUBING, AND INTO CLEAR FLOW PATH TUBING DURING NORMAL USAGE. ALL CONTAMINATION FOUND INSIDE THE DEVICE APPEARS TO BE OF FOREIGN MATTER EXTERNAL TO THE DEVICE: DIRT, DUST, AND GRIM THROUGHOUT FLOW PATH. THIS FOREIGN CONTAMINATION IS SEEN INSIDE THE TRANSITION TUBING AND CLEAR FLOW PATH TUBING ON THROUGH TO THE OUTLET PORT. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS OBSERVED IN THIS DEVICE. SECTION B2, H1, H4, H7, H9 AND SECTION H6 HAVE BEEN UPDATED/CORRECTED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04663 | 2021-09-29 | 2021-09-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2000X11B | DS2000X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TESTING. THERE WAS NO HARM OR INJURY REPORTED. THE MANUFACTURER IS CURRENTLY INVESTIGATING THIS ISSUE AND A FOLLOW-UP REPORT WILL BE FILED WHEN THE INVESTIGATION IS COMPLETE. | IT WAS INITIALLY REPORTED, THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TESTING. THERE WAS NO HARM OR INJURY REPORTED. THE REPORTER OF THE EVENT CONFIRMED THE DEVICE WAS RE-TESTED AND IT PASSED ALL STEPS. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04623 | 2021-09-29 | 2021-09-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04624 | 2021-09-29 | 2021-09-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE WITH THE VENTILATOR'S INTERNAL BATTERY. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04597 | 2021-09-29 | 2021-09-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, A VENTILATOR'S POWER BUTTON REACTED POORLY. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S POWER BUTTON WAS REPEATEDLY PRESSED AND THE BUTTON RESPONDED. THE DEVICE WAS SCRAPPED BY THE MANUFACTURER. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04695 | 2021-09-29 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04596 | 2021-09-29 | 2021-09-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, A VENTILATOR'S LCD SCREEN WAS BLANK DURING TESTING. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04578 | 2021-09-29 | 2021-09-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A SERVICE REQUIRED ALARM OCCURRED RELATED TO "CIRCUIT DISCONNECT". THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE FAILED A TEST STEP DURING TESTING AND THE DEVICE'S SENSOR BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A SERVICE REQUIRED ALARM OCCURRED RELATED TO "CIRCUIT DISCONNECT" AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE FAILED A TEST STEP DURING TESTING AND THE SENSOR BOARD NEEDED REPLACED TO ADDRESS THE ISSUE. IT WAS REPORTED A SERVICE REQUIRED CODE WAS OBSERVED AND THE OXYGEN BLENDING MODULE BOARD NEEDED REPLACED TO ADDRESS THE ISSUE. THE OXYGEN BLENDING MODULE BOARD NEEDS REPLACED TO DO A TEST STEP FAILURE, NOT A SERVICE REQUIRED CODE. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04581 | 2021-09-29 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A GROWTH ON A PATIENT'S LIVER. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF OXYGEN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON 11 NOV, 2024 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER PREVIOUSLY REPORTED RECEIVED INFORMATION ALLEGING A VENTILATOR'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A GROWTH ON A PATIENT'S LIVER, PARTICLES IN AIRWAY. EYES ARE SWOLLEN AND TAKES TIME TO SWELLING TO GO DOWN. PATIENT STATED BLACK PARTICLES WILL BE SEEN IN COUGHING. NIGHT, PATIENT WILL WAKING UP COUGHING. THE FILTERS ARE VERY DIRTY ALL TIME. PATIENT IS HAVING HEADACHES IN MORNING WHEN FIRST WAKING UP 4 DAYS OF 7 DAYS OF WEEK. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF OXYGEN. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECTS: CLINICAL CODES AND HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04464 | 2021-09-29 | 2021-08-04 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP NOSE BLEEDS. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING NOSE BLEED RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND DIRT/DUST CONTAMINATION ON BLOWER MOTOR CASING. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGREDATION. THE DEVICE DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERROR. THE MANUFACTURER CONCLUDES THE DEVICE HASTHE PRESENCE OF CONTAMINATION IN THE AIRPATH. PIL CAN CONFIRM WATER INGRESS INTO THE BLOWER BOX. THERE WAS NO EVIDANCE OF SOUND ABATEMENT FOAM DEGREDATION. SECTION D8, D9 AND H3, H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04321 | 2021-09-29 | 2021-08-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS EVALUATED BY A FIELD SERVICE ENGINEER. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE BY THE FIELD SERVICE ENGINEER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE MANUFACTURERS' EVALUATION IS STILL ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR WAS EVAUATED BY A FIELD SERVICE ENGINEER AND DURING THE EVALUATION OF THE DEVICE BY THE FIELD SERVICE ENGINEER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE VENTILATOR'S INTERNAL BATTERY WAS CYCLED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04600 | 2021-09-29 | 2021-09-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR HAD A BATTERY ISSUE. THERE WAS NO HARM OR INJURY REPORTED. THE REPORTING FACILITY REVIEWED THE DEVICE'S ERROR LOG AND IT WAS DETERMINED THE INTERNAL BATTERY NEEDS REPLACED TO ADDRESS THE ISSUE. THE INTERNAL BATTERY WILL BE REPLACED BY THE REPORTING FACILITY. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04595 | 2021-09-29 | 2021-09-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, A VENTILATOR FAILED A TEST STEP DURING TESTING. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE NEEDS TO BE RECALIBRATED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04621 | 2021-09-29 | 2021-09-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | FX2100X15B | FX2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S DISPLAY WAS BLANK. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD, DISPLAY BOARD AND DISPLAY SCREEN WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04638 | 2021-09-30 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGED DEVICE IS NOT BLOWING OUT LIKE IT USED TO, IT'S BLOWING OUT LESS AIR RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04724 | 2021-09-30 | 2021-07-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04709 | 2021-09-30 | 2021-09-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED ON THIS DEVICE IN MDR 2518422-2022-04459. THIS REPORT WAS SUBMITTED AS A DUPLICATE REPORT OF THE PREVIOUSLY SUBMITTED REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04528 | 2021-09-30 | 2021-09-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION B4 WHICH HAS BEEN INCORRECTLY ADDED IN THE PREVIOUS REPORT WAS CORRECTED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04646 | 2021-09-30 | 2021-08-03 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SHORTNESS OF BREATH. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING ITCHYNESS/IRRITATION IN SKIN AND EYES, SLEEP DYSFUNTION, NASAL BLOCKAGE, COTTON MOUTH.THROAT, SHORTNESS OF BREATH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04649 | 2021-09-30 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A COLLAPSED LUNG, BLOOD CLOTS, AND SHORTNESS OF BREATH. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO EXPERIENCE DEVELOP A COLLAPSED LUNG, BLOOD CLOTS, AND SHORTNESS OF BREATH. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENT OF DEVELOP A COLLAPSED LUNG, BLOOD CLOTS, AND SHORTNESS OF BREATH AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION AND AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTIONS B1, B2 HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. SECTION G1 CORRECTED AND H10 WRONGLY CHECKED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04648 | 2021-09-30 | 2021-05-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400T11 | DSX400T11 |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103584) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CRYPTOGENIC ORGANIZING PNEUMONIA. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION FROM VOLUNTARY MEDWATCH FORM ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED TO PATIENT TO DEVELOP CRYPTOGENIC ORGANIZING PNEUMONIA.THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THERE IS NO CUSTOMER INFORMATION AND WE CANNOT REACH OUT TO THE CUSTOMER, HENCE, ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION CAN NOT BE MADE AT THIS TIME. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS UPDATED REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04647 | 2021-09-30 | 2017-10-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION BIPAP PRO W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX600H11 | DSX600H11 |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103694) ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA AND ATELECTASIS. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04725 | 2021-09-30 | 2021-09-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04514 | 2021-09-30 | 2021-08-12 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A BACTERIAL INFECTION AND A SPOT ON THE THROAT. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN WHICH PRESCRIBED MEDICATION WAS GIVEN. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING BACTERIAL INFECTION AND A SPOT ON THE THROAT RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM.THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN WHICH PRESCRIBED MEDICATION WAS GIVEN. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE MANUFACTURER EVALUATED THE DEVICE EXTERNALLY /INTERNALLY AND OBSERVED UNKNOWN DUST/DIRT CONTAMINATION INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM OBSERVED THROUGHOUT DEVICE ENCLOSURE AND AIRPATH SUGGESTING A SOURCE EXTERNAL TO THE DEVICE. THE MANUFACTURER FOUND SOUND ABATEMENT FOAM DEGRADATION AND BREAKDOWN WAS VISIBLY OBSERVED IN THE BASE UNIT AND THROUGHOUT THE BLOWER BOX AND WHERE THE BLOWER RESTS IN THE BLOWER BOX,A BLACK SUBSTANCES WAS LOCATED ON THE BLOWER. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THEY CONFIRMED THE PRESENCE OF CONTAMINATION IN THE AIRPATH AND CONFIRMED SOUND ABATEMENT FOAM DEGRADATION. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04741 | 2021-09-30 | 2021-08-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11C | DSX400H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING BLACK PARTICLES IN THE AIRWAY OF THE DEVICE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04532 | 2021-09-30 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED ON THIS DEVICE IN MDR 2518422-2021-03257-2. THIS REPORT WAS SUBMITTED AS A DUPLICATE REPORT OF THE PREVIOUSLY SUBMITTED REPORT. SECTION D1, D4 AND H4 HAS BEEN CORRECTED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04726 | 2021-09-30 | 2021-09-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04727 | 2021-09-30 | 2021-09-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04728 | 2021-09-30 | 2021-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED HAVING INFLAMMATION IN HER NOSE AND DRY THROAT. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 HAS BEEN CORRECTED AND UPDATED IN THIS REPORT CLINICAL CODE WAS PREVIOUSLY NOT CAPTURED IN THE INITIAL. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04537 | 2021-09-30 | 2021-08-03 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11 | DSX700T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE DIFFICULTY IN BREATHING AND LUNG INFECTION. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE EVENT ARE CURRENTLY UNKNOWN. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY SUBMITTED MDR 2518422-2021-04537-1 WITH INCORRECT SECTIONS B1, B2, H1, H6. CORRECTIONS TO PREVIOUS MDR ARE MADE IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM (ONLY PRODUCT PROBLEM WAS CHECKED IN PREVIOUS MDR). SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS (PREVIOUSLY IT WAS BLANK). SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6- HEALTH IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04652 | 2021-09-30 | 2021-08-05 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION EXPERT W/HUM/HT, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | CAX501T12 | CAX501T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP NODULES ON LUNGS. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE EVENT IS CURRENTLY UNKNOWN. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED TO DEVELOP NODULES ON LUNGS. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. MEDICAL INTERVENTION IS NOT SPECIFIED BY THE PATIENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL AND INTERNAL VISUAL INSPECTION OF DEVICE WERE COMPLETED BY THE MANUFACTURER AND FOUND EVIDENCE OF RESIDUE ON THE INSIDE OF TOP PANEL AND INSIDE OF FRONT PANEL, DUST ON OUTSIDE OF PLASTIC BLOWER BOX COVER, INSIDE OF PLASTIC BLOWER BOX COVER, ON TOP OF BLOWER BOX, AND INSIDE OF BOTTOM PANEL. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERROR LOGGED. THE DEVICE WAS APPLIED POWER AND THE DEVICE OPERATED PROPERLY. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH RESIDUE ON THE INSIDE OF TOP PANEL AND INSIDE OF FRONT PANEL, DUST ON OUTSIDE OF PLASTIC BLOWER BOX COVER, INSIDE OF PLASTIC BLOWER BOX COVER, ON TOP OF BLOWER BOX, AND INSIDE OF BOTTOM PANEL. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04643 | 2021-09-30 | 2021-07-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP STAGE 2 KIDNEY DISEASE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED STAGE 2 KIDNEY DISEASE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE DEVICE WAS EVALUATED AND THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL AND INTERNAL SHOWED NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS SOUND ABATEMENT FOAM DEGRADATION. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04540 | 2021-09-30 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING THROAT HURT, CHEST PAINS, ASTHMA FLARE UPS MORE THAN OFTEN, DIZZINESS, LIGHT HEADEDNESS, HEADACHES ALSO HAS BEEN AN INCREASE IN HEART PROBLEMS AND NASAL DRIPPING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. CORRECTED INFORMATION WAS PROVIDED IN SECTIONS A1, E1, G1, H6. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04535 | 2021-09-30 | 2021-08-01 | D, e, a, t, h | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT HAS SINCE DIED. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4).. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVED AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND ALLEGEDLY CAUSED A PATIENT DEATH. DESPITE MULTIPLE ATTEMPTS, THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04541 | 2021-09-30 | 2021-09-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, JP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | RJPX500S16 | RJPX500S16 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04539 | 2021-09-30 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/HT/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400T11C | DSX400T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURE SUBMITTED MDR 2518422-2021-04539-1 INCORRECTLY WITH SECTIONS E1 AND H6 IT WAS CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04533 | 2021-09-30 | 2021-07-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE SINUS INFECTION. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTION FOR ANTIBIOTICS. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 01, 2023, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE A SORE THROAT, NASAL DRAINAGE AND SCARRING OF THE LUNGS. THE PATIENT WILL BE SEEING A SPECIALIST TO RECEIVE MEDICAL INTERVENTION.). THE PATIENT DID NOT REPORT TO RECEIVE MEDICAL INTERVENTION. SECTION B1 WAS CORRECTED FOR BOTH ADVERSE EVENT AND PRODUCT PROBLEM. ( PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION. TO SERIOUS INJURY. SECTION H6 WAS CORRECTED AND UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04538 | 2021-09-30 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGED HEADACHES, NASAL/THROAT IRRITATION, AND BLACK PARTICLES IN THE TUBE AND TANK RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE DEVICE WAS EVALUATED AND THERE WAS NO SIGNS OF CONTAMINATION ON THE EXTERNAL PART OF THE DEVICE. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND FOUND THAT THERE WERE 0 BLOWER HOURS AND 0 THERAPY HOURS WERE LOGGED, SUGGESTING THE DEVICE DATA WAS RESET PRIOR TO PIL RECEIPT. 3026.9 MACHINE HOURS WERE LOGGED. CARE ORCHESTRATOR DATA WAS UNAVAILABLE. DUST/DIRT CONTAMINATION INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM WAS OBSERVED THROUGHOUT DEVICE ENCLOSURE, AND AIRPATH SUGGESTING A SOURCE EXTERNAL TO THE DEVICE. EVIDENCE OF WATER INGRESS ON THE BLOWER AND BLOWER BOX. NO ERRORS WERE LOGGED. THE DEVICE RISK TAGS (ER (B)(4) V20) ASSOCIATED WITH THIS FAILURE MODE INCLUDE: IRRIT02, INFECT01. THE RESULTS OF THIS INVESTIGATION DO NOT IMPACT THE CALCULATED RISKS. THE MANUFACTURER CONCLUDES NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION AND BREAKDOWN IN THE DEVICE. SECTION H6 CORRECTED AND WERE UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04569 | 2021-09-30 | 2021-08-10 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION BIPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX600H11 | DSX600H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE CANCER. THE PATIENT DID NOT REPORT RECEIVING ANY MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING TO HAVE CANCER RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURER CONCLUDES NO FURTHER ACTION IS NECESSARY. THE PATIENT DID NOT REPORT RECEIVING ANY MEDICAL INTERVENTION. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04566 | 2021-09-30 | 2021-08-12 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO BE SEPTIC AND EXPERIENCE PNEUMONIA AND A SINUS INFECTION. THE PATIENT WAS HOSPITALIZED AND RECEIVED MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP,AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING CAUSED THE PATIENT TO BE SEPTIC AND EXPERIENCE PNEUMONIA AND A SINUS INFECTION RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT WAS HOSPITALIZED AND RECEIVED MEDICAL INTERVENTION. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04740 | 2021-09-30 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A (CPAP, BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS DRY MOUTH. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. SECTION H4 AND H6 WAS CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04645 | 2021-09-30 | 2021-07-26 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT ALSO ALLEGED OF HAVING CANCER, WHEEZING AND BREATHING DIFFICULTY. THE EVENT IS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE EVENT IS CURRENTLY UNKNOWN. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE. POTENTIAL DARK KERATIN PARTICLES WERE FOUND ON BLOWER BOX OUTLET PORT. THE FINS OF THE BLOWER MOTOR AND IT'S OUTLET PORT ARE WHITE. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION POTENTIAL (CLINICAL CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS) WERE UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04708 | 2021-09-30 | 2021-07-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP METASTATIC BREAST CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04656 | 2021-09-30 | 2021-09-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04570 | 2021-09-30 | 2021-08-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE BLOOD CLOTS. THE PATIENT DID NOT REPORT RECEIVING ANY MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO BLOOD CLOTS IN THE LUNGS, DEVICE HAS STRANGE ODOR, BURNING / SMOKE / ELECTRICAL ODOR. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENT OF BLOOD CLOTS IN THE LUNGS AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION AND AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTIONS B1, B2 HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT-CLINICAL CODE CORRECTED (REPORTED INCORRECTLY IN THE INITIAL REPORT). SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04700 | 2021-09-30 | 2021-09-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04702 | 2021-09-30 | 2021-09-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2200X11B | DS2200X11B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04701 | 2021-09-30 | 2021-09-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04714 | 2021-09-30 | 2021-09-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS EVALUATED BY A FIELD SERVICE ENGINEER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04716 | 2021-09-30 | 2021-09-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A MALFUNCTION OF A VENTILATOR'S OXYGEN SENSOR. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A MALFUNCTION OF A VENTILATOR'S OXYGEN SENSOR. THERE WAS NO HARM OR INJURY REPORTED. THE OXYGEN SENSOR IS AN FIO2 SENSOR ASSEMBLY AND IS AN ACCESSORY TO THE VENTILATOR. THIS ACCESSORY WOULD NOT AFFECT THE VENTILATOR TO DELIEVER THERAPY AS DESIGNED. THERE WAS NO MALFUNCTION RELATED TO THE VENTILATOR. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04715 | 2021-09-30 | 2021-09-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A MALFUNCTION OF A VENTILATOR'S OXYGEN SENSOR. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A MALFUNCTION OF A VENTILATOR'S OXYGEN SENSOR. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS EVALUATED BY A THIRD PARTY FIELD SERVICE ENGINEER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04717 | 2021-09-30 | 2021-09-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "HIGH PRIORITY ALARM" ON A VENTILATOR. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A "HIGH PRIORITY ALARM" ON A VENTILATOR. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS EVALUATED BY A THIRD PARTY FIELD SERVICE ENGINEER AND THE CUSTOMERS COMPLAINT WAS CONFIRMED. VENTILATOR INOPERATIVE CODES WERE OBSERVED IN THE EVENT LOG. THE DEVICE'S SYSTEM BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04718 | 2021-09-30 | 2021-09-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04719 | 2021-09-30 | 2021-09-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICES LCD DISPLAY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04721 | 2021-09-30 | 2021-09-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE RELATED TO THE INTERNAL BATTERY WAS OBSERVED IN THE DOWNLOADED EVENT LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04720 | 2021-09-30 | 2021-09-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04644 | 2021-09-30 | 2021-07-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP DIFFICULTY BREATHING. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED DIFFICULTY BREATHING AND PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THE PATIENT DID NOT REPORT TO RECEIVE MEDICAL INTERVENTION. THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04582 | 2021-09-30 | 2021-09-02 | I, n, j, u, r, y | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED CARBON DIOXIDE POISONING AND WORSENING KIDNEY AND CHRONIC OBSTRUCTIVE PULMONARY DISEASE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04696 | 2021-09-30 | 2021-04-21 | D, e, a, t, h | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CONTRIBUTED TO A PATIENT EXPIRING. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR'S SOUND ABATEMENT FOAM BECAME DEGRADED. THE PATIENT EXPIRED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR ANALYSIS AND THERE WAS NO EVIDENCE OF FOAM DEGRADATION OBSERVED. THE VENTILATOR WAS FOUND TO OPERATE AND ALARM AS DESIGNED AND DID NOT CONTRIBUTE TO THE PATIENT'S DEATH. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04722 | 2021-09-30 | 2021-08-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04729 | 2021-09-30 | 2019-08-25 | D, e, a, t, h | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT EXPIRED WHILE USING A VENTILATOR. THE MANUFACTURER IS CURRENTLY INVESTIGATING THIS EVENT AND A FOLLOW-UP REPORT WILL BE FILED WHEN THE INVESTIGATION IS COMPLETE. | IT WAS INITIALLY REPORTED, THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT EXPIRED WHILE USING A VENTILATOR. THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER FOR EVALUATION. AN OUTSIDE SOURCE, PLAINTIFF COUNSEL, CONFIRMED THE MANUFACTURER HAS BEEN DISMISSED FROM THIS CASE AS "THERE IS NOTHING IN ANY REPORT FROM ANY MEDICAL PROVIDER DISCOVERED IN THE 2020 ACTION RELATING TO OR MENTIONING A FAILURE OF ANY VENTILATOR OR ITS ALARM SYSTEM." | Y | N | HEALTH PROFESSIONAL |  | 2021-09
MW5104315 | 2021-09-30 | 2014-06-18 | I, n, j, u, r, y | Voluntary report | CPAP REMSTAR PRO CFLEX +_ | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | USED CPAP PHILIPS SINCE (B)(6) 2012 - FOUND NOSE CANCER IN (B)(6) 2014 SINCE, SINUSITIS, HEADACHES, ETC. JUST LEARN THAT THESE CPAP HAVE PROBLEM (CPAP REGISTERED TODAY AT PHILIPS AFTER RECEIVED LETTER. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04710 | 2021-09-30 | 2021-08-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN INITIAL REPORT SECTION B5 WAS INCOMPLETELY MENTIONED AND THE UPDATED SECTION B5 SHOULD BE- THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING OF SEEING PARTICLES IN TUBING/CHAMBER AND HAVING RESPIRATORY PROBLEMS. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO EVIDENCE OF CONTAMINATION. THE MANUFACTURER COMPLETED AN INTERNAL VISUAL INSPECTION. THE MANUFACTURER FOUND EVIDENCE OF WATER INGRESS AND AN UNKNOWN WHITE PARTICLE CONTAMINATION IN THE FRONT PANEL OF THE CASE, ISO PORT, INLET PORT, THE BLOWER BOX AND BLOWER MOTOR. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS LOGGED. THE DEVICE WAS APPLIED POWER AND THE DEVICE OPERATED PROPERLY. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH WATER INGRESS AND AN UNKNOWN WHITE POWDER PARTICLE CONTAMINATION INCONSISTENT WITH THE SOUND ABATEMENT FOAM. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. SECTION D9, G3, H3 AND H6 HAS BEEN UPDATED / CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04639 | 2021-09-30 | 2021-07-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. SECTION B5 SHOULD BE PREVIOUSLY REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED NOTICED BLACK PARTICLES AND HEADACHE. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT IS SCRAPPED. SECTION H6 (HEALTH EFFECT - CLINICAL CODE) HAS BEEN CORRECTED AND (TYPE OF INVESTIGATION,INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS) HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04492 | 2021-09-30 | 2021-08-03 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500S11 | DSX500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG NODULES. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIUOSLY RECEIVED INFORMATION ALLEGING A PATIENT TO DEVELOP LUNG NODULES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04706 | 2021-09-30 | 2021-07-07 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11C | DSX400H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CHRONIC KIDNEY DISEASE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY SUBMITTED MDR 2518422-2021-04706-1 WITH INCORRECT SECTIONS B1, B2, H1, H6. CORRECTIONS TO PREVIOUS MDR ARE MADE IN THIS REPORT AS FOLLOWS. SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY PRODUCT PROBLEM WAS CHECKED IN PREVIOUS MDR) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY IT WAS BLANK) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6- HEALTH IMPACT CODE WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04491 | 2021-09-30 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP PRO W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX600H11 | DSX600H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SORES IN THEIR NOSE. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 07, 2024, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SORES IN THEIR NOSE. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM (ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR). SECTION B2 WAS CHANGED TO BLANK (PREVIOUSLY IT WAS OTHER SERIOUS). SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04490 | 2021-09-30 | 2021-08-06 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THE PATIENT RECEIVED CHEMOTHERAPY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04653 | 2021-09-30 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK PARTICLES COMING FROM DEVICE.THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO CONTAMINATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER AND FOUND DUST/DIRT CONTAMINATION INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM WAS OBSERVED THROUGHTOUT DEVICE ENCLOSURE, HUMIDIFIER ENCLOSURE AND AIRPATH. THERE WERE ALSO SIGNS OF WATER INGRESS IN THE BLOWER AND ON THE BLOWER BOX. DISCOLORATION OBSERVED ON BRASS NUT ON BLOWER. BLACK CONTAMINATION CONSISTENT WITH BIO-CONTAMINATION AND INCONSISTENT SOUND ABATMENT FOAM WAS OBSERVED ON FLIP LID SEAL OF THE HUMIDIFIER. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND ONE ERORR. THE MANUFACTURER CONCLUDES THE DEVICE HAS DUST/DIRT CONTAMINATION CONSISTENT WITH WATER INGRESS. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04655 | 2021-09-30 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP AUTO SV | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS,INC | DSX900H11C | DSX900H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND OBSERVED THAT THE BASE UNIT WAS MISSING THE ACCESSORY FLIP DOOR. UNKNOWN DUST/DIRT CONTAMINATION PRESENT ON ALL SURFACES OF DEVICE. INTERIOR INVESTIGATION OBSERVED AN UNKNOWN DUST/DIRT CONTAMINATION PRESENT THROUGHOUT DEVICE. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED IN THE BASE UNIT THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS. THE MANUFACTURER CONCLUDES THE PRESENCE OF CONTAMINATION IN THE AIRPATH. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04698 | 2021-09-30 | 2021-08-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER PERFORMED VISUAL INSPECTED THE DEVICE. THE MANUFACTURER OPENED THE DEVICE TO CHECK FOR ANY CONTAMINATION/BLACK PARTICLES WITHIN THE FLOW PATHS. THERE ARE NO SIGNS OF CONTAMINATES ENTERING THE FLOW PATH FROM OUTSIDE THE DEVICE. THE MANUFACTURER FOUND NO SIGNS OF FOAM DEGRADATION. THE DEVICE POWERED ON AND PROVIDED THERAPY. THE ERROR LOG OF THE DEVICE WAS DOWNLOADED AND REVIEWED. THE DEVICE WAS TESTED AND THE MANUFACTURER CONFIRMED THE DEVICE OPERATES AS DESIGNED. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04567 | 2021-09-30 | 2021-08-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | RECERT - DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | UDSX500S11 | UDSX500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE NASAL/THROAT IRRITATION OR SORENESS AND INFECTIONS. THE PATIENT RECEIVED MEDICAL INTERVENTION IN THE FORM OF ANTIBIOTICS. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE NASAL/THROAT IRRITATION OR SORENESS AND INFECTIONS. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM (ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR). SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04493 | 2021-09-30 | 2021-08-06 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/WIFI, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11W | DSX500H11W |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG NODULES. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGED LUNG NODULES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED NASAL/THROAT IRRITATION OR SORENESS, LUNG NODULES ON BOTTOM OF RIGHT LUNG, DIFFICULTY BREATHING, ASTHMA AND THROAT FEELS LIKE CONTINUOUS PHLEGM. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL PART OF THE DEVICE AND FOUND THERE WAS NO UNUSUAL WEAR OR TEAR AND VERY SMALL AMOUNTS OF UNIDENTIFIED "CONTAMINATION" IN THE RECESS AND JOINTS OF THE ENCLOSURE. POWERED ON THE DEVICE USING A KNOWN GOOD POWER SUPPLY AND POWER CHORD WHICH SHOWED THE BASE UNIT PROVIDED AIRFLOW. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND FOUND THERE WERE SOME SLIGHT DIRT AND DUST AROUND THE BLOWER BOX. A SCREW AND TWO BLACK PIECES OF PLASTIC APPEARED TO HAVE BROKEN OFF OF THE BLOWER AND WERE FOUND LOOSE IN THE BLOWER BOX. THE BLOWER SEAL SHOWED FINE DARK CONTAMINATION THAT APPEARS CONSISTENT WITH THE KERATIN CONTAMINATION OBSERVED IN ER (B)(4). THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERRORS. UNIT HAS LOGGED 12,023 MACHINE HOURS, AND 12,023 BLOWER HOURS. LOG IS ATTACHED. THE MANUFACTURER APPLIED POWER TO THE DREAMSTATION. VERIFIED DEVICE POWERS ON, PROVIDES AIRFLOW. RISK TAGS (ER (B)(4) V20) ASSOCIATED WITH THIS MODULE OF FAILURE INCLUDE: IRRIT01, IRRIT02, INFECT01, DEGRAD04, CHOKE02. THE RESULT OF THIS INVESTIGATION DO NOT IMPACT THE CALCULATED RISKS. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS OBSERVED IN THE BASE UNIT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04571 | 2021-09-30 | 2021-08-10 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE PARTICLES IN THEIR LUNGS, HEADACHES AND DIZZINESS. THE PATIENT DID NOT REPORT RECEIVING ANY MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY SUBMITTED MDR 2518422-2021-04571-1 WITH INCORRECT SECTIONS B1, B2, H1, H6. CORRECTIONS TO PREVIOUS MDR ARE MADE IN THIS REPORT AS FOLLOWS. SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY PRODUCT PROBLEM WAS CHECKED IN PREVIOUS MDR) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY IT WAS BLANK) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6- HEALTH IMPACT CODE WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04654 | 2021-09-30 | 2021-09-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | 751P | 751P |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04588 | 2021-09-30 | 2021-09-24 | D, e, a, t, h | Manufacturer report | BIPAP S/T, C-SERIES, CHINA | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | CN1061423 | CN1061423 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO EXPIRE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04421 | 2021-09-30 | 2021-09-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTOSV ADV SYSTEM ONE ASSY,INTL | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS INC | 951P | 951P |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04561 | 2021-09-30 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11W | DSX500H11W |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK SPECKS/PARTICLES IN TUBING/CHAMBER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04568 | 2021-09-30 | 2021-08-12 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE GLAUCOMA, DRY MOUTH, AND CHEST CONGESTION/IRRITATION. THE PATIENT DID HAVE SURGERY TO TREAT GLAUCOMA AS A FORM OF MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE GLAUCOMA, DRY MOUTH, AND CHEST CONGESTION/IRRITATION. THE PATIENT DID HAVE SURGERY TO TREAT GLAUCOMA AS A FORM OF MEDICAL INTERVENTION. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL AND EXTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND BLACK, BROWN AND WHITE UNKNOWN CONTAMINATION (DUST LIKE) AND SOME VERY SMALL POTENTIAL HAIRS AT THE AIR INPUT OF THE BLOWER BOX AND BLOWER MOTOR. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. THE DEVICE'S EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS LOGGED. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND ARE CONSISTENT WITH DUST AND DIRT, DIFFERENT COLORED FIBER OR HAIR CONTAMINATE AND KERATIN, CONTAMINATION INCONSISTENT WITH THE SOUND ABATEMENT FOAM. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. SECTION D9, G3, H3 AND H6 HAS BEEN UPDATED / CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04765 | 2021-10-01 | 2021-07-07 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A STROKE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM AND AND CAUSED A PATIENT TO DEVELOP A STROKE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04745 | 2021-10-01 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE A NODULE IN THE LUNG, CHEST INFECTION AND DIZZINESS. THE PATIENT WENT TO RECEIVE MEDICAL INTERVENTION AND WAS DIAGNOSED WITH NODULE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 14, 2024, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING EXPERIENCE A NODULE IN THE LUNG, CHEST INFECTION AND DIZZINESS THAT ARE RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04751 | 2021-10-01 | 2021-08-04 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE STUFFINESS IN RIGHT NOSTRIL AND PRESSURE IN RIGHT EYE. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF MEDICATION TO TREAT CONGESTION IN NOSTRIL. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A PATIENT EXPERIENCED STUFFINESS IN RIGHT NOSTRIL AND PRESSURE IN RIGHT EYE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF MEDICATION TO TREAT CONGESTION IN NOSTRIL. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04744 | 2021-10-01 | 2021-08-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE A COUGH, CHEST TIGHTNESS AND A FEVER. THE PATIENT WAS HOSPITALIZED TO RECEIVE MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 14, 2024, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE A COUGH, CHEST TIGHTNESS AND A FEVER. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM (ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR) SECTION H1 WAS CHANGED FROM SERIOUS INJURY.TO MALFUNCTION. SECTION H6 HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04747 | 2021-10-01 | 2021-08-08 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400T11C | DSX400T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE CANCER. THE PATIENT DID NOT REPORT RECEIVING ANY MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER PREVIOUSLY SUBMITTED MDR 2518422-2021-04747-1 WITH INCORRECT SECTIONS B1, B2, H1, H6. CORRECTIONS TO PREVIOUS MDR ARE MADE IN THIS REPORT AS FOLLOWS. SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY PRODUCT PROBLEM WAS CHECKED IN PREVIOUS MDR). SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY IT WAS BLANK). SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6- HEALTH IMPACT CODE WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04749 | 2021-10-01 | 2021-08-07 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200H11C | DSX200H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE IRRITATED EYES AND SINUS ISSUES CAUSING ISSUES WITH BREATHING AND NOSEBLEEDS. THE PATIENT DID RECEIVE SURGERY AS MEDICAL INTERVENTION FOR EYE IRRITATION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04743 | 2021-10-01 | 2021-08-06 | I, n, j, u, r, y | Manufacturer report | RECERT - DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | UDSX500S11 | UDSX500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE A TUMOR IN THEIR LUNG. THE PATIENT DID NOT REPORT RECEIVING ANY MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE A TUMOR IN THEIR LUNG. THE PATIENT DID NOT REPORT RECEIVING ANY MEDICAL INTERVENTION. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. DURING THE LAB INVESTIGATION THERE WAS NO SOUND ABATEMENT FOAM DEGRADATION. SECTION D8, D9 AND H6 HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04712 | 2021-10-01 | 2021-09-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04711 | 2021-10-01 | 2021-09-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, A VENTILATOR FAILED TEST STEPS DURING TESTING. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04713 | 2021-10-01 | 2021-09-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04805 | 2021-10-01 | 2021-09-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS NOT PROVIDING PRESSURE. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR WAS NOT PROVIDING PRESSURE. THERE WAS NO HARM OR INJURY REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-10
MW5104331 | 2021-10-01 | 2018-07-01 | D, e, a, t, h | Voluntary report | PHILIPS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | MY HUSBAND USED A PHILIPS DREAMSTATION CPAP MACHINE EVERY NIGHT FOR ABOUT A YEAR BEFORE HE BEGAN EXPERIENCING SHOULDER PAIN. AFTER GOING TO SEVERAL DOCTORS, HE WAS DIAGNOSED WITH LUNG CANCER IN THE FALL OF 2018 AND DIED IN (B)(6) OF 2019. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-10
MW5104334 | 2021-10-01 | 2021-07-31 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS SLEEP AND RESPIRATORY CARE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC |  |  |  | PHILIPS C-PAP RECALLED SERIAL NUMBER (B)(4). AFTER STARTING TO USE C-PAP I BEGAN HAVING COUGH AND BLOODY NOSES. ON (B)(6) 2021 I BEGAN HAVING EYE REDNESS, DIZZINESS, NAUSEA AND VOMITING. ON (B)(6) 2021 I HAD A COVID TEST RAN AT (B)(6) AND IT WAS NEGATIVE, I WAS SEEN AT URGENT CARE ON (B)(6) 2021 AND HAD ANOTHER NEGATIVE COVID TEST. I WAS SEEN BY MY DOCTOR ON (B)(6) 2021 AND PRESCRIBED STEROIDS AND EYE DROPS. I THEN RECEIVED THE PHILIPS C-PAP RECALL LETTER AND STOPPED USING. I AM STILL HAVING NAUSEA AND DIZZINESS. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-10
MW5104335 | 2021-10-01 | 2021-04-14 | I, n, j, u, r, y | Voluntary report | RESPIRONICS DREAMSTATION CPAP WITH HEATED HUMIDIFIER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I HAVE BEEN USING MY CPAP EVERY NIGHT SINCE FALL OF 2020. I HAVE HAD CONSTANT HEADACHES, DIZZINESS, ASTHMA, TROUBLE BREATHING, NAUSEA, DEPRESSION, EYE IRRITATION, AND A FAIRLY CONSTANT COUGH. I NOTICED SOMETIME IN (B)(6) THAT MY AIRWAY HOSE HAD BLACK STUFF IN IT, I ASSUMED I HAD MADE A MISTAKE AND THROUGH THE HOSE AWAY, BUT KEPT USING THE CPAP CLEANING MORE REGULARLY. THE EYE IRRITATION HAPPENS OFTEN WHEN I WAKE UP, BUT IS RELIEVED HALF HOUR AFTER I GET OUT OF BED. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-10
MW5104337 | 2021-10-01 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I WAS ISSUED A PHILIPS DREAMSTATION CPAP DEVICE FOR MY SLEEP APNEA. THE DEVICE WAS RECALLED FOR DEFECTIVE INSULATING FOAM. PHILIPS HAS YET TO REPLACE THE UNIT. I HAVE RENTED AND EVENTUALLY PURCHASED A USED RESMED 10 CPAP TO REPLACE THE DEFECTIVE PHILIPS MACHINE DUE TO THE LENGTHY PERIOD OF INACTION ON THE PART OF PHILIPS. I HAVE REQUESTED THAT PHILIPS PAY (B)(6) FOR THE COSTS OF 3 MONTHS RENTAL AND PURCHASE OF THE REPLACEMENT MACHINE. I DO NOT WANT THE PHILIPS REPLACEMENT MACHINE, NOR WILL I EVER USE ANOTHER PHILIPS PRODUCT. PHILIPS HAS DECLINED TO PAY FOR THE RESMED 10. I DO NOT BELIEVE THAT THIS IS AN UNREASONABLE REQUEST AND HAS GOTTEN ME BACK ON MY THERAPY MUCH MORE RAPIDLY THAN WAITING AROUND FOR THE PHILIPS REPLACEMENT. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | I |  | 2021-10
MW5104355 | 2021-10-01 | 2019-12-15 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAM STATION CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | BESIDES RESPIRATORY ISSUES I WAS DIAGNOSED WITH BREAST CANCER (B)(6) 2019. I HAD TWO SURGERIES AND A MONTH OF RADIATION. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-10
MW5104359 | 2021-10-01 | 2021-07-26 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | THE DREAMSTATION AUTOBIPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX700 | DSX700T11C |  | RECEIVED A NOTICE FROM PHILIPS RESPIRONICS ABOUT THE SLEEP APNEA MACHINE I USE, THE DREAMSTATION #SN(B)(4) TELLING ME THAT PARTICULAR MODEL IS BEING RECALLED DUE TO DANGEROUS PARTICULATES THAT ARE FROM THE DISINTEGRATION FROM A FOAM SEAL AND THAT THOSE PARTICULATES THEN TRAVEL THROUGH THE HOSE AND GO DIRECTLY TO THE LUNGS CAUSING IRRITATION AND OTHER MORE SERIOUS PROBLEMS. NOT USING IT AND I STOP BREATHING MANY TIMES A MINUTE ALL THROUGH THE NIGHT. AS A PATIENT WITH APNEA, COPD, AND CONGESTIVE HEART FAILURE THIS MACHINE HELPS PROLONG THE QUALITY AND LENGTH OF MY LIFE. THE SECONDARY PROBLEM OF NOT HAVING THAT BEDSIDE AIDE IS THAT PHILIPS HAS NO SATISFACTORY ANSWER OF HOW LONG BEYOND A YEAR WILL I AND MILLIONS OF OTHERS BE WITHOUT THIS DEVICE. I DON'T SEE WHY THEY CAN'T MANUFACTURE A NEW FOAM SEAL AND TRAIN MANY PEOPLE IN EVERY STATE TO CHANGE THEM OUT FOR USERS OR PROVIDE OLDER MACHINES, (BTW, THE MACHINE IS IN PERFECT WORKING ORDER OTHER THAN THOSE PESKY PARTICULATES), AT NO COST TO USERS THAT WORKED FINE BEFORE THE DREAMSTATION WAS CREATED. I SPENT THE BETTER PART OF A DAY BEFORE I REACHED YOUR OFFICES TRYING TO GET INFORMATION, AND ALL I GOT WAS SHUFFLED TO THE NEXT PHONE NUMBER AND OFFICE AND NO ONE HAD ANSWERS OR INFORMATION. THIS IS CPAP. PLEASE HAVE THE COURTESY TO ANSWER ME THROUGH AN EMAIL BEFORE THE END OF NEXT WEEK. A QUOTE FROM THE PHILIPS DREAMSTATION HANDBOOK: "RESPIRONICS, INC WARRANTS THAT THE SYSTEM SHALL BE FREE FROM DEFECTS OF WORKMANSHIP AND MATERIALS AND WILL PERFORM IN ACCORDANCE WITH THE PRODUCT SPECIFICATION FOR A PERIOD OF (2) YEARS FROM THE DATE OF SALE BY RESPIRONICS, INC. TO THE DEALER. IF THE PRODUCT FAILS TO PERFORM IN ACCORDANCE WITH THE PRODUCTS SPECIFICATIONS, RESPIRONICS, INC. WILL REPAIR OR REPLACE- AT ITS OPTION-THE DEFECTIVE MATERIAL OR PART. RESPIRONICS, INC. WILL PAY CUSTOMARY FREIGHT CHARGES FROM RESPIRONICS, INC. TO THE DEALER LOCATION ONLY". MY MACHINE WAS BOUGHT 1+1/2 YEARS AGO. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-10
MW5104363 | 2021-10-01 | 2021-05-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION1 | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DS1 |  |  | USED PHILIPS CPAP MACHINE FOR 20 YEARS AND LEARNED BOTH OF THEM WERE RECALLED. COUGHING, BURNING, HEADACHES. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-10
MW5104366 | 2021-10-01 | 2021-09-01 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DREAMSTATION |  |  | I'VE HAD ONGOING SERIOUS, DEBILITATING HEADACHES AND NAUSEA WHILE REGULARLY USING PHILIPS DREAMSTATION CPAP DEVICE. I ALSO HAD SKIN IRRITATION AROUND MY NOSE WHERE THE MASK RESTS ON MY FACE. RECEIVED A RECALL NOTICE AND IMMEDIATELY STOPPED USING THE DEVICE. THE HEADACHES HAVE SUBSIDED, BUT I HAVE SERIOUS CONCERNS ABOUT LONG-TERM SIDE EFFECTS INCLUDING POTENTIAL INCREASED CANCER RISK. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04589 | 2021-10-01 | 2021-08-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP NODULES ON THEIR LUNGS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM AND CAUSED THE PATIENT TO HAVE NODULES ON LUNGS. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENT OF NODULES IN THE PATIENT'S LUNGS AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. THE PATIENT HAS SWAPPED OUT THE DEVICE AT THE DISTRIBUTOR AND AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04731 | 2021-10-01 | 2021-08-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SINUS ISSUES. THE PATIENT HAD SURGERY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO EXPERIENCE HEADACHES, EYE ISSUES, DRY MOUTH, SYMPTOMS WORSENING, SINUS SURGERY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENT OF HEADACHES, EYE ISSUES, DRY MOUTH, SYMPTOMS WORSENING, SINUS SURGERY AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04763 | 2021-10-01 | 2020-01-31 | I, n, j, u, r, y | Manufacturer report | BIPAP AUTOSV ADV,SYS ONE,ASSY,60 SRS,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | 960P | 960P |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103298) ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ALLERGIES. THE PATIENT REQUIRED IMMUNOTHERAPY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04735 | 2021-10-01 | 2021-09-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | 761P | 761P |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04293 | 2021-10-01 | 2021-07-08 | I, n, j, u, r, y | Manufacturer report | BIPAP AVAPS C SERIES CORE PACKAGE, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DS1160HS | DS1160HS |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED DIFFICULTY BREATHING AND SHORTNESS OF BREATH. THE PATIENT DID RECEIVE MEDICAL INTERVENTION AND WAS HOSPITALIZED TWICE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN DIFFICULTY BREATHING AND SHORTNESS OF BREATH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED HOSPITALIZATION THE PATIENT DID RECEIVE MEDICAL INTERVENTION AND WAS HOSPITALIZED TWICE. AFTER THE SECOND ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, THE CUSTOMER STATED THAT THE DEVICE IS AVAILABLE TO RETURN AS ONLY AFTER RECEIVEING REPLACEMENT. THE MANUFACTURER IS SUBMITTING A FOLLOW UP REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FOLLOW UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04800 | 2021-10-01 | 2021-08-10 | I, n, j, u, r, y | Manufacturer report | BIPAP A30, CANADA, HEATED TUBE CORE PKG | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | 1111148 | 1111148 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP KIDNEY DISEASE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP KIDNEY DISEASE, COUGH, HEADACHE AND SHORTNESS OF BREATH. THESE EVENTS WERE ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP KIDNEY DISEASE, COUGH, HEADACHE AND SHORTNESS OF BREATH. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 CORRECTED (ADDED 1880, 1816 AND 4457) AND UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04761 | 2021-10-01 | 2021-09-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S OUTPUT BACTERIA FILTER WAS TURNING DARK. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04798 | 2021-10-01 | 2021-08-11 | I, n, j, u, r, y | Manufacturer report | BIPAP SYSTEM ONE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103251) ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP COLON CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED TO HAVE BURNING EYES, DRY SORE THROAT, COUGHING, HEADACHES, DIZZINESS, RUNNY NOSE, UPPER AIRWAY IRRITATION, RECENT DIAGNOSED COLON CANCER WITH UNIDENTIFIED SPOTS ON LIVER YET AT THIS TIME.THE PATIENT DID NOT REPORT TO RECEIVE ANY MEDICAL INTERVENTION. THERE IS NO CUSTOMER INFORMATION AND WE CANNOT REACH OUT TO THE CUSTOMER, HENCE, ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION CAN NOT BE MADE AT THIS TIME. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS UPDATED REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04764 | 2021-10-01 | 2021-08-04 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11 | DSX400H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY SUBMITTED MDR 2518422-2021-04764-1 WITH INCORRECT INFORMATION IN SECTIONS B1, B2, H1, H6. CORRECTIONS TO PREVIOUS MDR ARE MADE IN THIS REPORT AS BELOW. SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY PRODUCT PROBLEM WAS CHECKED IN PREVIOUS MDR) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY IT WAS BLANK) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6- HEALTH IMPACT CODE WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04602 | 2021-10-01 | 2021-08-04 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | UDSX500S11 | UDSX500S11 |  | THE MANUFACTURER BECAME AWARE OF INFORMATION ALLEGING A HOSPITALIZATION DUE TO HER CPAP DEVICE BEING NO LONGER FUNCTIONAL. USER WAS HOSPITALIZED FOR 8 DAYS DUE TO A CHEST INFECTION AND AN EXACERBATION OF COPD. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION USER WAS HOSPITALIZED FOR 8 DAYS DUE TO A CHEST INFECTION AND AN EXACERBATION OF COPD RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04809 | 2021-10-01 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103279) ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER PREVIOUSLY RECEIVED INFOMATION AN ALLEGING SHORTNESS OF BREATH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. AFTER THE FINAL ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, CUSTOMER DECLINED TO RESPOND TO THE GFE RELATED QUESTIONS AND TERMINATED THE CALL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04750 | 2021-10-01 | 2021-08-04 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE BRONCHITIS, EAR INFECTION, COUGH AND ASTHMA. THE PATIENT DID NOT REPORT RECEIVING ANY MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT ALLEGED TO EXPERIENCE BRONCHITIS, EAR INFECTION, COUGH AND ASTHMA. THE PATIENT DID NOT REPORT RECEIVING ANY MEDICAL INTERVENTION. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION G1 CORRECTED. SECTION H6 TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04807 | 2021-10-01 | 2021-08-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC |  |  |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103312) ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04773 | 2021-10-01 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID/HT, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | CAX400T12 | CAX400T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04739 | 2021-10-01 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04737 | 2021-10-01 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF BLACK PARTICLES IN TUBING/CHAMBER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04736 | 2021-10-01 | 2021-09-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO AUTO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSG500S11 | DSG500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04738 | 2021-10-01 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING/CHAMBER/BLACK DEBRIS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04734 | 2021-10-01 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11C | DSX400H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGING SHORTNESS OF BREATH IN THE MIDDLE OF THE NIGHT. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 (CLINICAL SIGNS) CORRECTED AND SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04774 | 2021-10-01 | 2021-01-12 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP DECREASED LUNG FUNCTION. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04772 | 2021-10-01 | 2021-07-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION BIPAP AUTOSV | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | DSX900T11C | DSX900T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP PULMONARY FIBROSIS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT ALLEGED TO DEVELOP PULMONARY FIBROSIS. NO MEDICAL INTERVENTION WAS REPORTED. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. SECTION D1, D2 CORRECTED. SECTION D4 MODEL #, CATALOG # , SERIAL #, UNIQUE IDENTIFIER (UDI) # UPDATED. SECTION G1 UPDATED. SECTION G4 PMA/510(K) UPDATED. SECTION D8, D9, H2, H3 AND H6 WERE UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04771 | 2021-10-01 | 2021-08-04 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY SUBMITTED MDR 2518422-2021-04771-1 WITH INCORRECT SECTIONS B1, B2, H1, H6. CORRECTIONS TO PREVIOUS MDR ARE MADE IN THIS REPORT AS FOLLOWS. SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY PRODUCT PROBLEM WAS CHECKED IN PREVIOUS MDR) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY IT WAS BLANK) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6- HEALTH IMPACT CODE WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04775 | 2021-10-01 | 2021-08-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP/HUM/HT/CELL,CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | CAX700T12C | CAX700T12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED VISUALIZATION OF PARTICLES. THE PATIENT ALSO ALLEGED HAVING BOTHERED EYES, COUGH, AND HEADACHES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL AND EXTERNAL VISUAL INSPECTION OF DEVICE WERE COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND UNKNOWN WHITE AND BLACK CONTAMINATION (DUST-LIKE), INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM, AT THE AIR INPUT OF THE BLOWER BOX, TINY UNKNOWN BLACK (INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM), BROWN AND WHITE SPECS OF CONTAMINATION ON THE BOTTOM THE BLOWER BOX. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERROR LOGGED. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE MANUFACTURER ALSO CONFIRMS THERE WAS NO VISIBLE DAMAGE OR FUNCTIONALITY FAILURES OF THE DEVICE, WHICH SUGGEST THAT THE SOURCE OF CONTAMINATION WAS EXTERNAL TO THE DEVICE. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04622 | 2021-10-01 | 2021-09-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2200X11B | DS2200X11B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE DEVICE'S FRONT DISPLAY WAS FOUND TO BE DAMAGED. THE DEVICE'S DISPLAY SCREEN AND DISPLAY BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04801 | 2021-10-01 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX600H11C | DSX600H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DRY MOUTH, VERY SORE DRY SCRATCHY THROAT, POWDERY GRAY SUBSTANCES RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DRY MOUTH, VERY SORE DRY SCRATCHY THROAT, POWDERY GRAY SUBSTANCES RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6, G1 WAS CORRECTED, UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04693 | 2021-10-01 | 2021-08-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED HEADACHES, NAUSEA AND PATIENT HAD A HEART ATTACK. THE PATIENT DID NOT REPORT RECEIVING ANY MEDICAL INTERVENTIONS. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO EXPERIENCE A HEART ATTACK, HEADACHES, AND NAUSEA. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENT OF HEART ATTACK AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04531 | 2021-10-01 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED RESPIRATORY ISSUES, COUGH, SINUS PROBLEMS, HEADACHES AND COUGHED UP A FOREIGN BODY THE SIZE OF A PEA. THE PATIENT DID RECEIVE MEDICAL INTERVENTION WITH ANTIBIOTICS. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED RESPIRATORY ISSUES, COUGH, SINUS PROBLEMS, HEADACHES AND COUGHED UP A FORIEGN BODY THE SIZE OF A PEA. THE REPORTED EVENT OF SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED RESPIRATORY ISSUES, COUGH, SINUS PROBLEMS, HEADACHES AND COUGHED UP A FORIEGN BODY THE SIZE OF A PEA AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. THE PATIENT HAS SWAPPED OUT THE DEICE AT THE DISTRIBUTOR AND AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04558 | 2021-10-01 | 2021-08-06 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11 | DSX400H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED SINUS PROBLEMS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION AND WAS PUT ON MEDICATION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING SINUS PROBLEMS,NAAL/THROAT IRRITATION RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM.THE PATIENT DID RECEIVE MEDICAL INTERVENTION AND WAS PUT ON MEDICATION. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION.UNIT WAS SCRAPPED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04651 | 2021-10-01 | 2021-08-05 | D, e, a, t, h | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER WAS MADE AWARE OF INFORMATION ALLEGING A DEATH RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE USER PASSED AWAY AFTER USING HIS DEVICE THE NIGHT BEFORE. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04557 | 2021-10-01 | 2021-08-10 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION BIPAP ST30 H/C, DS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX1030H11C | DSX1030H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED DIFFICULTY BREATHING, SLEEPING AND BARELY ABLE TO FUNCTION. THE PATIENT DID RECEIVE MEDICAL INTERVENTION WITH HOSPITALIZATION FOR ALLEGED ISSUES. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04790 | 2021-10-01 | 2021-09-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO AUTO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSG500S11 | DSG500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04791 | 2021-10-01 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING HEAVINESS IN HER CHEST, BLACK SPECS IN THE RESEVOIR, DIFFICULTY BREATHING/SHORT OF BREATH AND VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. SECTION H6 HEALTH EFFECT - CLINICAL CODE IS UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04793 | 2021-10-01 | 2021-08-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM.THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION.THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04792 | 2021-10-01 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. SECTION B5 SHOULD BE PREVIOUSLY REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED LUNG AND HEART PROBLEMS, HAS BLACK SPECKS ON LUNGS AND LITTLE BLACK SPECKS GETTING INTO EYES CAUSING ABRASION IN EYES, EYE SWELLING, AND HEADACHES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. DURING INVESTIGATION NOTED THAT THE DC JACK POWER CABLE CONNECTOR (P4) WAS HEAVILY CORRODED AND THAT THE BOTTOM ENCLOSURE WAS CRACKED AROUND WHERE P4 WAS CONNECTED. DURING EXTERIOR INVESTIGATION OBSERVED CORROSION ON THE POWER PORT AND THE BOTTOM ENCLOSURE CRACKED AT P4 CONNECTOR. THERE WAS ALSO A BROWNISH YELLOW RESIDUE UNDER UI KNOB. IN INTERIOR INVESTIGATION OBSERVED A DEAD INSECT WITHIN THE BOTTOM ENCLOSURE. OBSERVED RUST BUILDUP AT P4 CONNECTOR. UNKNOWN LIQUID INGRESS AND MINERAL DEPOSITS PRESENT ON BLOWER HOUSING AND MOTOR CASING. OBSERVED A KERATIN-LIKE SUBSTANCE PRESENT AROUND THE BLOWER HOUSING OUTLET. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED IN THE BASE UNIT. IN SECONDARY FINDINGS A KERATIN-LIKE SUBSTANCE WAS OBSERVED AROUND THE BLOWER BOX OUTLET. THE OCCURRENCE OF A KERATIN-LIKE SUBSTANCE OBSERVED AROUND THE BLOWER BOX HAS BEEN PREVIOUSLY ADDRESSED. THE UNKNOWN DUST/DIRT CONTAMINATION AND FIBERS FOUND ON THE BLOWER CASING/IMPELLER AND BLOWER BOX WAS INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM CONTAMINATION. THE PRESENCE OF CONTAMINATION THROUGHOUT THE AIRPATH SUGGESTS A SOURCE EXTERNAL TO THE DEVICE. MINERAL SPOTS CONSISTENT WITH WATER INGRESS WERE FOUND WITHIN BLOWER BOX, MOTOR CASING, BOTTOM ENCLOSURE, AND BLOWER BOX HOUSING. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND 13 ERRORS. THE MANUFACTURER CONCLUDES IS UNABLE DIRECTLY ADDRESS THE SYMPTOMS OUTLINED IN THE COMPLAINT AND CONFIRM THAT THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN OBSERVED IN THE BASE UNIT. ALSO CONFIRM THE PRESENCE OF CONTAMINATION IN THE AIRPATH. SECTION H6 (HEALTH EFFECT - CLINICAL CODE) HAS BEEN CORRECTED AND (TYPE OF INVESTIGATION,INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS) HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04694 | 2021-10-01 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP W/HUM/HT/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX200T11C | DSX200T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED DIFFICULTY BREATHING, LIGHTHEADEDNESS, DIZZINESS, HEADACHES AND SINUS INFECTIONS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION AND WAS ON SEVERAL PRESCRIPTIONS. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 15 , 2024 AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED DIFFICULTY BREATHING, LIGHTHEADEDNESS, DIZZINESS, HEADACHES AND SINUS INFECTIONS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION AND WAS ON SEVERAL PRESCRIPTIONS.THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 WAS CORRECTEDE AND UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04799 | 2021-10-01 | 2018-08-14 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103238) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A HEART ATTACK. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | THE MANUFACTURER PREVIOUSLY RECEIVED A VOLUNTARY MEDWATCH (MW5103238) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A HEART ATTACK. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL AND INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND BROWN AND WHITE UNKNOWN CONTAMINATION AND SOME VERY SMALL POTENTIAL HAIRS AT THE AIR INPUT OF THE BLOWER BOX AND UNKNOWN BROWN AND BLACK CONTAMINANTS WERE ON THE BOTTOM OF THE BLOWER BOX. ALSO FOUND, BLACK CONTAMINATION CONSISTENT WITH KERATIN AT THE AIR OUTLET OF THE BLOWER BOX. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND 32 ERROR. THE MANUFACTURER CONCLUDES CONTAMINATION FOUND WERE CONSISTENT WITH CONTAMINATION OF BROWN, WHITE, BROWN AND BLACK UNKNOWN CONTAMINATION AND SOME VERY SMALL POTENTIAL HAIRS AT THE BLOWER BOX AND BLACK CONTAMINATION CONSISTENT WITH KERATIN AT THE AIR OUTLET OF THE BLOWER BOX. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04806 | 2021-10-01 | 2020-09-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103277) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP INNER EAR DAMAGE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 11, 2024, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103277) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP INNER EAR DAMAGE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS INITIAL MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04767 | 2021-10-01 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500S11 | DSX500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING 3MM DOT POSSIBLY BEEN A PARTICLE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING 3MM DOT POSSIBLY BEEN A PARTICLE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04703 | 2021-10-01 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER WAS MADE AWARE OF AN ALLEGED THERMAL EVENT TO A DREAMSTATION AUTO CPAP WHILE IN USE. THE USER ALLEGES OBSERVING SMOKE AND FLAMES AT THE TIME OF THE EVENT. THERE WAS NO PATIENT HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED OBSERVING SMOKE AND FLAMES AT THE TIME OF THE EVENT. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. ADDITIONAL INFORMATION WAS RECEIVED AND ADDED TO THE REPORT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER ON DATE 03/27/2023 FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED ON DATE 06/29/2023. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. UPON EXTERNAL/INTERNAL EXAMINATION, BLACK CONTAMINATION AT THE BLOWER SEAL OF THE BLOWER BOX AND IT IS CONSISTENT WITH THE KERATIN CONTAMINATION OBSERVED. A BROWN/RED CONTAMINATION WAS OBSERVED NEAR THE POWER INPUT PORT AND ON THE BACKSIDE OF THE BOTTOM ENCLOSURE. PIL FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. PIL FOUND NO EVIDENCE OF A THERMAL EVENT OCCURRED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. NO ERRORS WERE FOUND. THE MANUFACTURER CONCLUDES THAT THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION FOUND WITHIN THE DEVICE. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04766 | 2021-10-01 | 2021-07-07 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED LUNG CANCER. THE PATIENT DID RECEIVE MEDICAL INTERVENTION AND WAS HOSPITALIZED. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE /RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM THAT BECAME DEGRADED AND CAUSED THE PATIENT TO HOSPITALIZE AND WAS DIAGNOSED WITH LUNG CANCER. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04770 | 2021-10-01 | 2021-09-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED. SECTION B5 SHOULD BE CORRECTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP,AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES IN THE WATER CHAMBER,MADE ME SICK WITH ALLERGIC SYMPTOMS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM.THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 (HEALTH EFFECT - CLINICAL CODE) HAS BEEN CORRECTED AND (TYPE OF INVESTIGATION,INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS) HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04768 | 2021-10-01 | 2021-08-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY IN THE INITIAL REPORT SECTION B5 WAS NOT UPDATED, THE CORRECT B5 SHOULD BE - THE MANUFACTURER RECEIVED VOLUNTARY MEDWATCH (MW5103319) .THE PATIENT HAS ALLEGED BLACK PARTICLES IN THE DEVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO EVIDENCE OF CONTAMINATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION ON BOTH SIDES OF THE BLOWER BOX. A HUMIDIFIER WAS ALSO RETURNED TO THE MANUFACTURER ALONG WITH DEVICE. AN INTERNAL VISUAL INSPECTION OF THE HUMIDIFIER WAS COMPLETED BY THE MANUFACTURER AND FOUND AN UNKNOWN CONTAMINATION ON THE BOTTOM OF THE INSIDE OF THE HUMIDIFIER. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS LOGGED.THE DEVICE WAS HOOKED UP TO POWER SUPPLY, AIRFLOW WAS VERIFIED AND THE DEVICE OPERATED PROPERLY. THE MANUFACTURER CONCLUDES AN UNKNOWN CONTAMINATION ON THE BOTTOM OF THE INSIDE OF THE HUMIDIFIER.INCONSISTENT WITH THE SOUND ABATEMENT FOAM. THE MANUFACTURER CONFIRMED THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04769 | 2021-10-01 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF BLACK PARTICLES ON RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS UPDATED REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04675 | 2021-10-04 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | GBX500S15 | GBX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK PARTICLES OBSERVED IN THE DEVICE, HUMIDIFIER, MASK AND TUBING. THE PATIENT HAS BEEN HAVING MIGRAINES FOR THE PAST YEAR. THERE WAS NO REPORT OF SERIOUS HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. THE EXTERNAL PART OF THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO SIGNS OF CONTAMINATION. THE INTERNAL PART OF THE DEVICE WAS VISUALLY INSPECTED BY THE MANUFACTURER AND FOUND EVIDENCE OF CONTAMINATION INSIDE THE BLOWER BOX. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE WAS APPLIED POWER AND THE MANUFACTURER VERIFIED THE DEVICE PROVIDED AIRFLOW. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER AND FOUND ONE E-22 LOGGED. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION WITHIN THE DEVICE. THE PREVIOUS REPORT DID NOT MENTION THE PATIENT HAVING MIGRAINES. SECTION H6 UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04677 | 2021-10-04 | 2021-09-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | GBX500S15 | GBX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DARK SEDIMENT OBSERVED IN BOTH WATER CHAMBERS. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR FURTHER EVALUATION. THE DEVICE WAS VISUALLY INSPECTED BY THE MANUFACTURER AND FOUND EVIDENCE OF WHITE AND BLACK CONTAMINATES IN THE AIR INPUT OF THE BLOWER BOX, AND BLOWER MOTOR. BROWN FIBERS AND HAIR FOUND IN THE BLOWER BOX AND BLACK CONTAMINATION CONSISTENT WITH KERATIN FOUND IN THE AIR OUTLET OF THE BLOWER BOX, AND BLOWER MOTOR SEAL. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE HUMIDIFIER WAD NOT RETURNED FOR EVALUATION. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS LOGGED. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH KERATIN AND HAIR. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION FOUND WITHIN THE DEVICE. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04676 | 2021-10-04 | 2021-09-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | GBX500S15 | GBX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK PARTICLES WERE OBSERVED. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. THE EXTERNAL PART OF THE DEVICE WAS VISUALLY INSPECTED. THE MANUFACTURER FOUND NO SIGNS OF CONTAMINATION ON THE DEVICE. THE INTERNAL PART OF THE DEVICE WAS VISUALLY INSPECTED AND FOUND EVIDENCE OF WATER INGRSS IN THE BLOWER BOX AND THE BLOWER MOTOR. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE WAS APPLIED POWER AND THE MANUFACTURER VERIFIED AIRFLOW. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED AND FOUND ONE ERROR E-130. THE MANUFACTURER CONCLUDES THE DEVICE HAS EVIDENCE OF CONTAMINATION CONSISTENT WITH WATER INGRESS. THERE WAS NO REPORT OF SOUND ABATEMENT FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04674 | 2021-10-04 | 2021-09-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | GBX500S15 | GBX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK PARTICLES OBSERVED. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR FURTHER EVALUATION. THE MANUFACTURER FOUND EVIDENCE OF BLACK CONTAMINATION, CONSISTENT WITH KERATIN, AT THE AIR OUTLET OF THE BLOWER BOX AND THE BLOWER MOTOR SEAL. WHITE CONTAMINATION WAS FOUND IN THE AIR INPUT OF THE BLOWER BOX , ON TOP OF THE BLOWER BOX AND ON TOP OF THE BLOWER MOTOR. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER AND FOUND 3 ERRORS LOGGED. THE MANUFACTURER CONCLUDES THE CONTAMINATION FOUND WAS CONSISTENT WITH KERATIN. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION FOUND WITHIN THE DEVICE. | N | Y | LAY USER/PATIENT |  | 2021-10
MW5104385 | 2021-10-04 |  | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS DREAMSTATIN AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I HAVE THE PHILIPS DREAMSTATION AUTO CPAP. IT IS ONE OF THOSE RECALLED, I DON'T REMEMBER THE EXACT DAY IT HAPPEN BUT I DID FIND BLACK SPECKS IN MY HOSE AND MASK. AT THE TIME I THOUGHT IT WAS JUST DIRT PICKED UP FROM THE FLOOR OR DUST FROM WHERE I LEAVE IT DURING THE DAY. I FLUSHED OUT THE HOSE AND CLEANED THE MASK. I DON'T USE AN OZONE CLEANING DEVICE, JUST PLAIN SOAP AND WATER. IF I HAD KNOWN ABOUT THIS PROBLEM AT THE TIME I WOULD HAVE GOTTEN A SAMPLE OF THE SPECKS. I STOPPED USING THE CPAP. I HAVE A MUCH OLDER CPAP I'M USING AGAIN UNTIL I CAN GET A REPLACEMENT CPAP. BECAUSE WHAT HAS BEEN MENTIONED I HAVE DECIDED TO SEE A RESPIRATORY DOCTOR. THE APPOINTMENT IS NOT UNTIL (B)(6) 2021. IF ANYTHING IS FOUND, I WILL FOLLOW UP ON THIS NOTICE. SINCERELY, (B)(6). FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04827 | 2021-10-04 | 2021-09-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO, CN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | 767P | 767P |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04678 | 2021-10-04 | 2021-09-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP A40 | VENTILATOR, CONTINUOUS, MINIMAL VENTILATORY SUPPORT, FACILITY USE |  | 1111169 | 1111169 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-10
MW5104382 | 2021-10-04 |  | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS RESPIRONICS DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSXHCJ |  |  | I RECEIVED NOTICE FROM PHILIPS RESPIRONICS TODAY, (B)(6) 2021, THAT MY CPAP DREAMSTATION WAS RECALLED BY THE FDA DUE TO UNSAFE DEVICE MATERIALS. I HAVE BEEN PAYING FOR AND USING THIS DEVICE SINCE (B)(6) 2021. I HAD CONTACTED THE PRESCRIBING PHYSICIAN ABOUT THE ISSUE AND THEY INFORMED ME THAT THE DEVICE WAS SAFE AND NOT RECALLED! I REGISTERED MY CPAP ((B)(4)) WITH PHILIPS.COM TODAY. I AM VERY CONCERNED THAT MY RESPIRATORY AND CARDIOVASCULAR HEALTH HAS BEEN JEOPARDIZED. PLEASE CONTACT ME AT (B)(6) TO DISCUSS THIS FURTHER. THANK YOU. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | I |  | 2021-10
MW5104383 | 2021-10-04 | 2015-02-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS, INC. |  |  |  | I HAVE HAD CHRONIC SLEEP APNEA SINCE 2011. I STARTED USE IG A RESUMED CPAP UNTIL MOTOR GAVE OUT, IN 2016 I WAS GIVEN A DREAMSTATION AND I USED THAT UNTIL IT MALFUNCTIONED. I WAS ALWAYS SICK WITH RESPIRATORY ISSUES INCLUDING PNEUMONIA SEVERAL TIMES THAT I NEEDED HOSPITALIZATION. I DEVELOPED COPD AS WELL, I WAS GIVING A NEW DREAMSTATION WHICH LASTED TILL 2020, IN 2017 I WAS DIAGNOSED OF SQUAMOUS SMALL CELL CANCER, I HAVE HAD NUMEROUS TEST, AND RADIATION THERAPY. CANCER CAME BACK THIS PAST MAY, I HAD TO HAVE RADIATION THERAPY AGAIN. I RECEIVED A DIFFERENT MODEL NOPAL IN 2020 AND HAVEN'T HAD ANY RESPIRE ISSUES SINCE STOP USING THE DREAMSTATION. I'M STILL TAKING MEDS FOR COPD AND ASTHMA AND STILL BEING MONITORED FOR THE LUNG CANCER, I HAVE SUFFERED SO MUCH DURING THIS TIME NOW I'M SURE MY HEALTH ISSUES ARE FROM USING THE RECALLED DEVICE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-10
MW5104384 | 2021-10-04 | 2021-04-05 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DreamStation Auto CPAP w/Humid, w/H Tube, w/Cell, US | DSX500T11C |  | AFTER MONTHS OF USING THE PHILIPS DREAMWEAR CPAP MACHINE, I BEGAN HAVING ISSUES WITH MY RESPIRATORY TRACT. I HAVE BECOME EXTREMELY NASALLY, DIFFICULTY BREATHING THROUGH NOSE, YELLOW MUCUS. WAS JUST NOTIFIED OF THE RECALL ON THIS MACHINE TODAY (B)(6) 2021, AND AFTER READING ALL THE INFORMATION ON THIS RECALL, I REALIZE THAT I HAVE INDEED BEEN AFFECTED BY THIS PROBLEM. I WILL BE CONTACTING MY DOCTOR TO DISCUSS. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-10
MW5104387 | 2021-10-04 |  | I, n, j, u, r, y | Voluntary report | SYSTEM ONE AND DREAMSTATION | SYSTEM ONE AND DREAMSTATION | PHILIPS / RESPIRONICS, INC. | DSX500T11 DREAMSTATION AUTOCPAP |  |  | I HAVE USED TWO OF THE RECALLED PHILIPS RESPIRONICS DEVICES FOR OVER 10 YEARS. DURING THAT TIME I BECAME COPD AND REDUCED LUNG FUNCTION NOT UNDERSTAND WHY. I WOULD WAKE UP COUGHING USING PHILLIPS RESPIRONICS RECALLED DEVICES. I WAS PUT ON INHALERS AND OXYGEN AS NEEDED DURING THE USE OF THE RECALLED PRODUCTS OVER THE PAST 10 YEARS AS MY HEALTH DETERIORATED. THIS TEST WAS REQUESTED BY MY DOCTOR IN RESPONSE TO THE PHILIPS RESPIRONICS RECALL TO DETERMINE ANY DAMAGE. PRIOR TO THIS MY CHEST WAS XRAYED PRIOR TO USE OF THE RECALLED DEVICES AND HAD NO MASSES. I CAN SHOW BOTH BEFORE AND AFTER USE OF THE DEVICES XRAYS AND CT SCANS. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04756 | 2021-10-04 | 2021-08-12 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE ATRIAL FIBRILLATION. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF AN ABLATION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE ATRIAL FIBRILLATION. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE EVENT IS CURRENTLY UNKNOWN. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04803 | 2021-10-04 | 2021-09-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING BLACK PARTICLES IN THE WATER TANK AND MASK RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. AT THE TIME OF FIRST ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, RECEIVED AN AUTOMATIC MESSAGE STATING THE NUMBER THAT IS NO LONGER IN SERVICE. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04838 | 2021-10-04 | 2021-07-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04839 | 2021-10-04 | 2021-07-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11W | DSX500T11W |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DEVICE HAS STRANGE ODOR AND PARTICLES IN AIRWAY RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED SKIN IRRITATION. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DEVICE HAS STARNGE ODOR AND PARTICLES IN AIRWAY RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED SKIN IRRITATION. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04826 | 2021-10-04 | 2021-09-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04824 | 2021-10-04 | 2021-08-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11C | DSX400H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS INITIALLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN INITIAL REPORT SECTION B5 WAS MENTIONED INCORRECTLY. THE CORRECT B5 SHOULD BE- THE PATIENT ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO CONTAMINATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER AND FOUND UNKNOWN CONTAMINATE INSIDE OF THE BLOWER BOX AND ON THE BLOWER MOTOR. THE HUMIDIFIER WAS RETURNED FOR EVALUATION. THE INTERNAL INVESTIGATION OF THE HUMIDIFIER SHOWED THAT THERE WERE MULTIPLE SPOTS OF CONTAMINATION INSIDE OF THE HUMIDIFIER. THERE WAS ALSO CONTAMINATION ON THE RESERVOIR SEAL. THE LATCH SPRING FOR THE FLIP LID WAS NOTED TO BE BROKEN AS WELL. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND E-200 STOP ERROR. THE DEVICE WAS APPLIED POWER AND THE DEVICE OPERATED PROPERLY. THE MANUFACTURER CONCLUDES THE DEVICE HAS CONTAMINATION CONSISTENT WITH WATER INGRESS. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. IN THIS REPORT, SECTION D9, G3, H3, H6 HAS BEEN UPDATED OR CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04822 | 2021-10-04 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS INITIALLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE - THE PATIENT ALLEGED PARTICLES IN TUBING/CHAMBER. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO CONTAMINATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04821 | 2021-10-04 | 2021-08-06 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION BIPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX600T11C | DSX600T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON OCTOBER 09, 2023, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING/CHAMBER, ALLEGING BREATHING IN WHITE PARTICLES, HEART TROUBLES (A-FIBRILATION, PACEMAKER, DEFIBRILATOR ),EXPERIENCING NOSE IRRITATION, RED, WATERY EYES AND LIVER ISSUES RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGATION OF DIZZINESS AND/OR HEADACHE, NAUSEA / VOMITING, CANCER MELANOMA. THE SECTION E1 WAS UPDATED AND SECTIONS B1, B2, H1 AND H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR BOTH ADVERSE EVENTS AND PRODUCT PROBLEM (ONLY THE PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS REPORT.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECTS - CLINICAL CODES AND HEALTH EFFECTS - IMPACT CODE WAS CORRECTED AND UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04823 | 2021-10-04 | 2021-08-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. IT WAS ALSO REPORTED TAHT "PATIENT REPORTED SEEING BLACK PARTICLES IN TUBING/CHAMBER. NASAL/THROAT IRRITATION OR SORENESS. PATIENT GETS LOTS OF SINUS INFECTIONS AFTER HE USES HIS MACHINE. THE PATIENT IS HAVING PROBLEMS WITH HIS KIDNEY IN THE LAST FEW MONTHS". THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED DUE TO AGE. IN SECTION H6 SORE THROAT AND UNSPECIFIED KIDNEY OR URINARY PROBLEM WAS MISSED IN INITIAL REPORT, HENCE CAPTURED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04841 | 2021-10-04 | 2021-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. CORRECTED SECTION B5 SHOULD BE REPORTED AS : THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER ALSO RECEIVED INFORMATION ALLEGING ABOUT RESPIRATORY INFECTION. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. (CLINICAL CODE UPDATED FOR RESPIRATORY TRACT INFECTION). | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04825 | 2021-10-04 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX600T11C | DSX600T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED LUNGS HURT, SHORTNESS OF BREATH AND RESIDUE IN THROAT. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED LUNGS HURT, SHORTNESS OF BREATH AND RESIDUE IN THROAT. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04802 | 2021-10-04 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED COUGH, DRY MOUTH AND THROAT. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. UNIT SCRAPPED. SECTION H6 IS UPDATED WITH HEALTH EFFECT - CLINICAL CODES FOR COUGH AND DRY MOUTH. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04619 | 2021-10-04 | 2021-08-17 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103347) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING ASTHMA RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE EVENT IS CURRENTLY UNKNOWN. THE REPORTED EVENT OF ASTHMA AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THE EVENT IS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04620 | 2021-10-04 | 2019-02-15 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103376) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP GASTRIC LYMPHOMA. THE PATIENT UNDERWENT RADIATION THERAPY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103376) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP GASTRIC LYMPHOMA. THE PATIENT UNDERWENT RADIATION THERAPY IN RESPONSE TO THE REPORTED EVENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO EVIDENCE OF CONTAMINATION. THE MANUFACTURER COMPLETED AN INTERNAL VISUAL INSPECTION. THE MANUFACTURER FOUND EVIDENCE OF CONTAMINATION IN THE INLET PORT, BOTTOM PANEL OF THE CASE, DRY BOX SEAL AND IN THE BLOWER MOTOR. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS LOGGED. THE DEVICE WAS APPLIED POWER AND THE DEVICE OPERATED PROPERLY. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04618 | 2021-10-04 | 2019-01-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103376) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SEVERE SKIN IRRITATION AND LETHARGY. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND IN INITIAL REPORT SECTION B5 MENTIONED INCOMPLETE, SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO LETHARGY, SKIN IRRITATION, SORE THROAT AND HEAT IN THE FACE. THE REPORTED EVENTS OF LETHARGY, SKIN IRRITATION, SORE THROAT AND HEAT IN THE FACE AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENTS ARE ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURER CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FOLLOW UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04617 | 2021-10-04 | 2021-08-02 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103341) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP BRONCHITIS. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED A VOLUNTARY MEDWATCH (MW5103341) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP BRONCHITIS. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04616 | 2021-10-04 | 2019-01-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103342) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING ASTHMA RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE REPORTED EVENT AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THE EVENT IS RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04615 | 2021-10-04 | 2021-08-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX200H11C | DSX200H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103407) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A SCARRED LUNG. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 18 , 2023 AND SECTION H10 SHOULD BE REPORTED AS: HE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103407) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A SCARRED LUNG. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME..THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04804 | 2021-10-04 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES IN TUBING CHAMBER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION.THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04672 | 2021-10-04 | 2021-09-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | GBX500S15 | GBX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | IN THE PREVIOUS REPORT SECTION D8 WAS MISSED TO CAPTURE, CORRECTED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04752 | 2021-10-04 | 2021-08-10 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE DIFFICULTY BREATHING AND WHEEZING. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF AN ASTHMA INHALER. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04837 | 2021-10-04 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED. SECTION B5 SHOULD BE CORRECTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP,AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN AIRWAY FROM DEVICE,SAID HE SAW BLACK PARTICLES IN THE MASK RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM.THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04673 | 2021-10-04 | 2021-09-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | GBX500S15 | GBX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED AND SECTION H10 SHOULD BE CORRECTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED THE BLACK BITS OBSERVED. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING THE SUBSEGENT TESTING OF THE DEVICE, THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGARDATION. THE MANUFACTURER USED A FITT (FOAM INTEGRITY TEST TOOL) AND WAS ABLE TO CONFIRM THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION H6 TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04754 | 2021-10-04 | 2021-08-06 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE A SINUS INFECTION. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF AN ANTIBIOTIC. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. ADDITIONAL INFORMATION WAS RECEIVED, SECTION B5 WAS CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE A SINUS INFECTION, NASAL/THROAT IRRITATION OR SORENESS, FEELING LIKE NOT GETTING ENOUGH OXYGEN. THE PATIENT HAS BEEN USING AN ANTIBIOTIC TO HELP HIS SINUS INFECTIONS. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED IN THE BASE UNIT. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WAS NO ERROR FOUND. LIGHT AMOUNT OF DUST/ DIRT FOUND AROUND THE AIR INLET. WATER INGRESS LOCATED IN THE BLOWER BOX. PIL CAN CONFIRM THE PRESENCE OF A DUST/DIRT CONTAMINANT AROUND THE AIR INLET. PIL CAN ALSO CONFIRM THE PRESENCE OF WATER INGRESS IN THE BLOWER BOX. PIL FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION AND BREAKDOWN. SECTION H6 WERE UPDATED AND CORRECTED BY ADDING HEALTH EFFECT - CLINICAL CODES IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04755 | 2021-10-04 | 2021-08-10 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED EYE IRRITATION AND BLISTERING. THE PATIENT DID RECEIVE A STEROID EYE DROP AS A FORM OF MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04748 | 2021-10-04 | 2021-08-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AND CORRECTED AS THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING NASAL/THROAT IRRITATION OR SORENESS,DIFFICULTY BREATHING/SHORTNESS OF BREATH,NAUSEA, BAD HEADACHE,DIZZINESS, CARDIAC ISSUES,DEVICE IS NOISY AND HAD BLACK PARICLES,DEVICE GET HOT TO TOUCH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. CORRECTED INFORMATION WAS PROVIDED IN SECTION B3 AND G1. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04753 | 2021-10-04 | 2021-08-10 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE NAUSEA, HEADACHES, LUNG ISSUES AND WAS DIAGNOSED WITH CANCER. THE PATIENT WAS HOSPITALIZED AND RECEIVED MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON MAY 14, 2024, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE NAUSEA, HEADACHES, AND LUNG ISSUES. HE WAS DIAGNOSED WITH STOMACH CANCER AND A SPOT AROUND THE PANCREAS. THE PATIENT WAS HOSPITALIZED AND RECEIVED MEDICAL INTERVENTION. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGED DIZZINESS, CANCER, AND PANCREATIC NEUROENDOCRINE TUMOR. AFTER THE FIRST ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, CUSTOMER STATED THE DEVICE WON'T AVAILABLE FOR RETURN. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04840 | 2021-10-04 | 2021-08-28 | D, e, a, t, h | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON 10/09/2023, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING TO ANXIETY, FATIGUE , BRAIN FOG, HEADACHE, LIGHTHEADEDNESS, DISORDERED BREATH, COUGH, NAUSEA AND SEEING BLACK PARTICLES RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED ABOUT ALLEGED CANCER AND DEATH. SECTION B2, H1 AND H6 HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04757 | 2021-10-04 | 2021-08-12 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO BE DIAGNOSED WITH LUNG CANCER. THE PATIENT DID NOT REPORT RECEIVING ANY MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04671 | 2021-10-04 | 2021-09-22 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DEX500S13 | DEX500S13 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED. THE PATIENT HAS INDICATED BEING TREATED FOR A PERSISTENT, RECURRING COLD AND NASAL INFLAMMATION. THE PATIENT IS RECEIVING CORTISONE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED. THE PATIENT HAS INDICATED BEING TREATED FOR A PERSISTENT, RECURRING COLD AND NASAL INFLAMMATION. THE PATIENT IS RECEIVING CORTISONE. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL PART OF THE DEVICE AND FOUND NO DISTINCT WEAR, TEAR, OR CONTAMINATION ON THE ENCLOSURE. EXAMINATION OF THE AC AND DC POWER CABLES FOUND EVIDENCE OF CRIMPING, NO EXPOSED CONDUCTORS OR TEARS IN CABLE SHEATHS. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND FOUND MINIMAL DUST CONTAMINATION IN THE BLOWER. NO EVIDENCE OF SOUND ABATEMENT FOAM BREAKDOWN WAS OBSERVED, AND THE FOAM APPEARS INTACT. THE MANUFACTURER VERIFIED THE UNITS DO POWER-UP, PROVIDE FLOW. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION BUT DUST CONTAMINATION WAS OBSERVED IN THE BLOWER, WHICH IS CONSISTENT WITH BEING RELATED TO THE ENVIRONMENTAL CONDITIONS. SECTION D4 UPDATED/CORRECTED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04758 | 2021-10-04 | 2021-09-20 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO BE DIAGNOSED WITH CANCER. THE PATIENT DID NOT REPORT RECEIVING ANY MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-10
MW5104404 | 2021-10-05 | 2021-05-02 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DREANSTNBIPAPPROHUMHTCELLDOM |  |  | ON (B)(6) 2021 WIFE (B)(6) NEARLY DIED. SHE WAS ADMITTED TO (B)(6) HOSPITAL (B)(6) FOR MANY PROBLEMS INCLUDING RESPIRATORY FAILURE. SINCE THEN SHE EXHIBITS MANY OF THE SYMPTOMS, THAT ARE LISTED ON THE RECALL NOTICE SENT BY PHILIPS. THEY INCLUDE FREQUENT HEADACHES THAT DO NOT RESPOND TO CONVENTIONAL HEADACHE MEDICINES, CONSISTENT AND WORSENING COUGH AND SINUS PROBLEMS. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | I |  | 2021-10
2518422-2021-04874 | 2021-10-05 | 2021-09-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04723 | 2021-10-05 | 2021-09-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2200X11B | DS2200X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED A TEST STEP. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR FAILED A TEST STEP. THERE WAS NO HARM OR INJURY REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04893 | 2021-10-05 | 2023-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO VENTILATOR'S SOUND ABATEMENT FOAM. ALSO, ALLEGED HEADACHE, COUGH AND PARTICLES IN TUBING CHAMBER. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO PARTICLES IN TUBING CHAMBER AND UNIT IS SCRAPPED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04625 | 2021-10-05 | 2021-09-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BROKEN. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04626 | 2021-10-05 | 2021-09-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
MW5104422 | 2021-10-05 |  | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAMSTATION AUTO CPAP HUM DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DREAM STATION AUTO HUM DOM |  |  | FDA RECALL BUT IF I STOP USING THIS DEVICE I COULD WITH HIGH LIKELIHOOD STOP BREATHING IN MY SLEEP. I DO NOT HAVE MEANS OF REPLACEMENTS, AND PHILIPS DOESN'T HAVE ANY TO REPLACE IT WITH? FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04875 | 2021-10-05 | 2021-09-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
MW5104419 | 2021-10-05 | 2021-08-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I USE THE PHILLIPS DREAMSTATION CPAP MACHINE AND I HAVE HAD A SORE THROAT AND HEADACHES FOR THE LAST 8 TO 10 WEEKS. I'VE ALSO HAD THE ISSUE WITH MY LUNGS HURTING WHEN I BREATHE AND I HAVE GOTTEN A SINUS INFECTION. I'VE BEEN TO THE DOCTOR TWICE AND HAVE BEEN THROUGH 2 ROUNDS OF ANTIBIOTICS AND STILL HAVE THE SAME ON GOING ISSUES. I JUST SAW THAT THE DREAMSTATION HAS BEEN RECALLED, AND I THINK THIS MAY BE WHY I HAVE BEEN HAVING HEALTH ISSUES. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-10
MW5104414 | 2021-10-05 | 2021-09-25 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS RESPIRONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | IP22 | 00606959025639 |  | WHEN DOING WEEKLY CLEANING, MY FILTER WAS BLACK AND BLACK PARTICLES WERE ON ALL RUBBER PIECES INSIDE MACHINE. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-10
MW5104407 | 2021-10-05 | 2016-10-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP HUMID HT DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX200T11 |  |  | I BEGAN USE OF A PHILIPS RESPIRONICS CPAP WHICH HAS BEEN RECALLED AND NOT REPLACED. I AM A RENAL CELL CARCINOMA SURVIVOR WITH SECONDARY LUNG NODULES. I CANNOT GET PHILIPS TO SEND ME A NEW MACHINE, AN CANNOT RISK LUNG CANCER AS A RESULT OF THEIR NEGLIGENCE. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04868 | 2021-10-05 | 2021-09-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04876 | 2021-10-05 | 2021-09-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1032804 | 1032804 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04796 | 2021-10-05 | 2020-09-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04878 | 2021-10-05 | 2021-09-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1032804 | 1032804 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICES LCD DISPLAY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04879 | 2021-10-05 | 2021-09-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG AND THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A MALFUNCTION OF THE DEVICE'S KEYPAD WAS OBSERVED. THE DEVICES FRONT PANEL BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04871 | 2021-10-05 | 2021-09-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2200X11B | DS2200X11B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, AN INTERNAL OXYGEN LEAK WAS OBSERVED. THE DEVICE'S OXYGEN CHECK VALVE O-RING WAS REPLACED AND THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04869 | 2021-10-05 | 2021-09-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | RJP2100X16B | RJP2100X16B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S EXTERNAL FLOW SENSOR FAILED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE EVALUATION HAS NOT YET BEEN COMPLETED BY THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR'S EXTERNAL FLOW SENSOR FAILED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S EXTERNAL FLOW SENSOR AND EXTERNAL FLOW SENOR CABLE WERE REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04873 | 2021-10-05 | 2021-09-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | RJP2100X16B | RJP2100X16B |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, A 'VENTILATOR INOPERATIVE" ALARM CONDITION WAS OBSERVED. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S SYSTEM BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04795 | 2021-10-05 | 2021-09-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04872 | 2021-10-05 | 2021-09-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A HIGH TEMPERATURE ALARM OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OUTLET FLOW PATH THERMISTOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04797 | 2021-10-05 | 2020-09-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
MW5104406 | 2021-10-05 | 2021-06-25 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | WOKE UP ONE MORNING WITH EVIDENCE OF BLOODY NOSE HAVING OCCURRED DURING THE NIGHT. HAD A BROKEN TOOTH A FEW NIGHTS LATER. BLOODY NOSES BEGAN TO BECOME FREQUENT ALONG WITH INCREASED POST NASAL DRIP AND CHRONIC SINUS INFECTION. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-10
MW5104408 | 2021-10-05 | 2020-04-12 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500H11C |  |  | BOTH FILTERS IN MY PHILIPS DREAMSTATION CPAP MACHINE WERE BLACK AND COVERED WITH AN UNKNOWN SUBSTANCE AFTER BEING USED OVERNIGHT. THIS OCCURRED ON TWO DIFFERENT OCCASIONS, SEVERAL MONTHS APART. (B)(6) 2020 WAS THE MOST RECENT EPISODE. AFTER IT OCCURRED, I HELD ON TO THE FILTERS FOR MANY MONTHS IN CASE THE PROBLEM BECAME FREQUENT. SINCE IT ONLY OCCURRED TWICE, AND DID NOT RECUR AFTER THE SECOND EPISODE, I THREW THE FILTERS OUT A YEAR LATER. IN THE LAST SEVERAL WEEKS, I'VE LEARNED THAT THE FOAM IN THE MACHINE IS KNOWN TO DEGRADE AND CAUSE PROBLEMS PROMPTING A RECALL. I FEAR THAT THE BLACK RESIDUE CLOGGING THE FILTERS WAS THE FOAM THAT DEGRADED IN THE MACHINE AND WAS LIKELY INHALED DURING SLEEP. PHILIPS REPORTED A CORRELATION BETWEEN UNAUTHORIZED CLEANING METHODS AND THIS PROBLEM. I HAVE NEVER USED OZONE OR ANY OTHER UNAPPROVED CLEANING METHOD. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04925 | 2021-10-05 | 2021-09-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO, CN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | 767P | 767P |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04157 | 2021-10-05 | 2021-09-17 | I, n, j, u, r, y | Manufacturer report | BIPAP AUTO, CN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | 767P | 767P |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE LUNG CANCER. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE REPORTED EVENT IS CURRENTLY UNKNOWN. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04780 | 2021-10-05 | 2021-08-15 | I, n, j, u, r, y | Manufacturer report | BIPAP S/T C SERIES CORE PACKAGE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DS1060HS | DS1060HS |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED LUNG AND BLADDER CANCER. THE PATIENT DID RECEIVE MEDICAL INTERVENTION THAT RESULTED IN SEVEN SURGERIES. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED LUNG AND BLADDER CANCER. THE PATIENT DID RECEIVE MEDICAL INTERVENTION THAT RESULTED IN SEVEN SURGERIES. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04877 | 2021-10-05 | 2021-09-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
MW5104413 | 2021-10-05 |  | I, n, j, u, r, y | Voluntary report | DREAMSTNAUTOCPAPHUMCELL DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500H11C |  |  | HAD TOTAL ALOPECIA AFTER STARTED USING THE PHILIPS DREAMSTATION. AFTER THE RECALL, I STOPPED USING THE CPAP. HAIR STARTED TO RETURN, PHILIPS SENT A REPLACEMENT CPAP MACHINE, IN TWO WEEKS WHAT HAIR HAD GROWN BACK, FELL OUT. HAVE SEEN DERMATOLOGIST AND 2 RHEUMATOLOGIST WAS TOLD "SOMETIMES THIS JUST HAPPENS". FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-10
MW5104397 | 2021-10-05 | 2021-05-11 | D, e, a, t, h | Voluntary report | PHILIPS DREAM STATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | NOSEBLEEDS LASTING HOURS. CANCER LESION AT SINUS LOCATION. CANCER IN SEVERAL LOCATIONS OF BODY INCLUDING LUNGS AND LIVER. EARLY REPORTS OF NODULES IN LUNGS. DROOPING EYELID ALMOST TO CLOSURE UNTIL FINALLY CLOSED. SEVERE HEADACHES. TOO MANY TO NAME OVER MANY MONTHS. POSITIVE CANCER AND OTHER PROBLEMS. TOO MANY PRESCRIPTIONS AT DEATH. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04860 | 2021-10-05 | 2021-08-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04896 | 2021-10-05 | 2021-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX600H11C | DSX600H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF BLACK PARTICLES IN AIRWAY RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE MANUFACTURER EVALUATED THE DEVICE EXTERNALLY AND UNKNOWN DUST/DIRT CONTAMINATION IS PRESENT ON SURFACES ON DEVICE. THE MANUFACTURER EVALUATED THE DEVICE INTERNALLY AND UNKNOWN DUST/DIRT CONTAMINATION ,FIBERS FOUND ON THE BLOWER CASING /IMPELLER AND BLOWER BOX WAS INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM CONTAMINATION.THE PRESENCE OF CONTAMINATION THROUGHOUT THE AIRPATH SUGGESTS A SOURCE EXTERNAL TO THE DEVICE. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE IS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS OBSERVED IN THE BASE UNIT AND THEY CONFIRMED THE PRESENCE OF CONTAMINATION IN THE AIRPATH. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04895 | 2021-10-05 | 2021-08-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING COUGH,SINUS ISSUES,PARTICLES IN TUBING AND DEBRIS/MOLD IN RESERVOIR RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM.THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE MANUFACTURER EVALUATED THE DEVICE EXTERNALLY /INTERNALLY AND OBSERVED UNKNOWN DUST/DIRT CONTAMINATE ON EXTERIOR AND UNKNOWN WHITE COLORED RESIDUE IN BLOWER BOX HOUSING AND ON BLOWER,EVIDENCE OF UNKNOWN LIQUID INGRESS ,UNKNOWN MINOR DUST/DIRT CONTAMINATION ON THE BLOWER IMPELLER. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THEY OBSERVED DUST/DIRT CONTAMINATION IN THE AIRPATH AN UNKNOWN WHITE RESIDUE AND EVIDENCE OF UNKNOWN LIQUID INGRESS FROM A SOURCE OF EXTERNAL TO THE DEVICE ,THERE IS NO EVIDENCE OF DEGRADED SOUND ABATEMENT FOAM. SECTION H6 WAS MISSED TO CAPTURE IN PREVIOUS MDR,NOW IT IS CORRECTED AND UPDATED IN HTIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04858 | 2021-10-05 | 2021-07-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN INITIAL REPORT SECTION B5 WAS INCOMPLETELY MENTIONED AND THE UPDATED SECTION B5 SHOULD BE- THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED VISUALIZATION OF PARTICLES IN TUBING/CHAMBER. THE PATIENT ALSO ALLEGED OF HAVING DIFFICULTY BREATHING/SHORT OF BREATH. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO EVIDENCE OF CONTAMINATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF CONTAMINATION INSIDE OF THE HUMIDIFIER BUT NO EVIDENCE OF CONTAMINATION INSIDE OF THE DREAMSTATION. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS LOGGED. THE DEVICE WAS APPLIED POWER AND THE DEVICE OPERATED PROPERLY. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. SECTION D9, G3 AND H6 HAS BEEN UPDATED / CORRECTED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04919 | 2021-10-05 | 2021-07-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED PNEUMONIA AND DEVELOPMENT OF COPD. IT IS UNKNOWN WHAT TYPE OF MEDICAL INTERVENTION WAS REQUIRED. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT WAS ALLEGING TO DEVELOP PNEUMONIA, DEVELOPMENT OF COPD AND BRONCHITIS. NO MEDICAL INTERVENTION WAS REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION E1 EMAIL ADDRESS WHICH WAS MISSED IN PREVIOUS REPORT WAS CAPTURED. SECTION G1 UPDATED. SECTION H6 HEALTH EFFECT - CLINICAL CODE ADDED IN THIS REPORT. SECTION H6 TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04918 | 2021-10-05 | 2021-07-22 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A FOREIGN PARTICLE TO BECOME LODGED IN THE UPPER GUM, RESULTING IN THE NEED FOR TOOTH EXTRACTION AND A BONE GRAFT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A FOREIGN PARTICLE TO BECOME LODGED IN THE UPPER GUM, RESULTING IN THE NEED FOR TOOTH EXTRACTION AND A BONE GRAFT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04922 | 2021-10-05 | 2021-08-17 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | CORRECTION TO PREVIOUS REPORT 2518422-2021-04922: THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE WITH THE DEVICE'S SOUND ABATEMENT FOAM. THE USER ALSO ALLEGED DEVELOPMENT OF LUNG NODULES. NO MEDICAL INTERVENTION WAS REPORTED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04842 | 2021-10-05 | 2021-08-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN THE PREVIOUSLY SUBMITTED REPORT SECTION B5 WAS INCOMPLETE. SECTION B5 WILL BE- THE PATIENT ALSO ALLEGED VISUALIZATION OF PARTICLES IN TUBING/CHAMBER AND HAVING DRY MOUTH. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF LIQUID INGRESS WITH MATERIAL CONSISTENT OF MINERAL DEPOSITS INSIDE THE ISO PORT ON THE REAR PANEL OF THE DEVICE. AN INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF LIQUID INGRESS TO THE BLOWER AND BLOWER BOX, AN UNKNOWN BLACK CONTAMINANT WAS ON THE BOTTOM OF THE BLOWER BOX. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. THE HUMIDIFIER WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR EVALUATION. THE MANUFACTURER FOUND EVIDENCE OF AN UNKNOWN CONTAMINANT WAS OBSERVED ON THE FLIP LID SEAL. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF LIQUID INGRESS WITH MATERIAL CONSISTENT OF MINERAL DEPOSITS INSIDE THE ISO PORT ON THE REAR PANEL OF THE DEVICE, EVIDENCE OF LIQUID INGRESS TO THE BLOWER AND BLOWER BOX, A BLACK CONTAMINANT WAS ON THE BOTTOM OF THE BLOWER BOX, A CONTAMINANT WAS OBSERVED ON THE FLIP LID SEAL OF THE HUMIDIFIER. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04848 | 2021-10-05 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION, SEEING BLACK SPECS IN THE WATER TANK AND SHE HAS SORES IN HER NOSE CURRENTLY AS WELL THAT ARE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGED NASAL/THROAT IRRITATION OR SORENESS. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. SECTION D9, G3, H2, H3 AND H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04849 | 2021-10-05 | 2021-08-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE LUNG ISSUES. THE PATIENT HAS BEEN USING OXYGEN AS A FORM OF MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUNDABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED LUNG ISSUES. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENT OF LUNG ISSUES WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED DUE TO AGE. SECTIONS B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04850 | 2021-10-05 | 2021-08-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE A COUGH. THE PATIENT RECEIVED A NEBULIZER, INHALER AND ANTIBIOTICS AS A FORM OF MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 17 , 2024 AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING COUGH. THE PATIENT RECEIVED A NEBULIZER, INHALER AND ANTIBIOTICS AS A FORM OF MEDICAL INTERVENTION RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04843 | 2021-10-05 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | ADDITIONAL INFORMATION WAS RECEIVED ON APR 15, 2025 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING SHORTNESS OF BREATH, FOAM STUFF IN TANK RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04846 | 2021-10-05 | 2021-08-10 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE DIFFICULTY BREATHING, SHORTNESS OF BREATH, WHEEZING AND RESPIRATORY FAILURE. THE PATIENT WAS HOSPITALIZED AND RECEIVED MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DIFFICULTY BREATHING, SHORTNESS OF BREATH, WHEEZING AND RESPIRATORY FAILURE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED (SYMPTOMS). THE PATIENT DID NOT REPORT TO RECEIVE MEDICAL INTERVENTION. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION.THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE.THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. SECTION H6 WERE UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04851 | 2021-10-05 | 2021-08-12 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11C | DSX400H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO BE DIAGNOSED WITH A LUNG TUMOR. THE PATIENT DID HAVE MEDICAL INTERVENTION IN THE FORM OF AN CHEST X-RAY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING TO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED LUNG TUMOR. THE PATIENT DID HAVE MEDICAL INTERVENTION IN THE FORM OF AN CHEST X-RAY. THIS NOTIFICATION WAS REVIEWED BY THE PMS CLINICAL EXPERT DUE TO THE PATIENT STATED AT FIRST HE THOUGHT HE HAD A COLD LATER TO GO TO A DOCTORS AND THE DOCTOR THOUGHT IT WAS BRONCHITIS UNTIL HE HAD ANOTHER VISIT AND GOTTEN A CHEST X RAY AND THEY FOUND A TUMOR IN HIS LUNGS. THEY DO BELIEVE IT WAS CAUSED BY THE DEVICE. ATTEMPTS HAVE BEEN MADE TO OBTAIN ADDITIONAL INFORMATION AND HAVE BEEN UNSUCCESSFUL. THE PATIENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION RELEVANT TO THE HARM REPORTED. HENCE, THE EVENT IS ASSESSED AS NOT RELATED TO THE RECALLED DEVICE (I.E. A FOAM DEGRADATION CONCERN) IN THIS CASE REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04844 | 2021-10-05 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. SECTION H6 HEALTH EFFECT NOT CAPTURED PREVIOUSLY HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04852 | 2021-10-05 | 2021-08-09 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE A COUGH, ASTHMA, SPOT IN LUNG AND WAS DIAGNOSED WITH ATRIAL FIBRILLATION. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A CATSCAN. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON MAY 23, 2024, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING COUGH, ASTHMA, SPOT IN LUNG AND WAS DIAGNOSED WITH ATRIAL FIBRILLATION RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A CATSCAN. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGED EYE IRRITATION, DIZZINESS AND/OR HEADACHE, ASTHMA (NEW OR WORSENING). | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04847 | 2021-10-05 | 2021-09-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04845 | 2021-10-05 | 2020-01-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO BE DIAGNOSED WITH BILATERAL BREAST CANCER. THE PATIENT DID NOT REPORT RECEIVING ANY MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO BE DIAGNOSED WITH BILATERAL BREAST CANCER. THE PATIENT DID NOT REPORT RECEIVING ANY MEDICAL INTERVENTION. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND FOUND EVIDENCE OF CONTAMINATION ON INLET PORT. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND NO ERRORS LOGGED. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF CONTAMINATION ON INLET PORT. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION D9, G3 AND H6 HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04898 | 2021-10-05 | 2021-09-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | RJPX500S16 | RJPX500S16 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION HOWEVER DUST/DIRT CONTAMINATION INCOSISTENT WITH DEGRADED SOUND ABATEMENT FOAM WAS OBSERVED THROUGHOUT DEVICE ENCLOSURE, AND AIR PATH, SUGGESTING A SOURCE EXTERNAL TO DEVICE. SLIGHT BLACK CONTAMINATION AT BLOWER SEAL OF THE BLOWER BOX IS CONSISTENT WITH THE KERATIN CONTAMINATION OBSERVED NO ERROR WERE LOGGED. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CONTAMINATION. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
MW5104421 | 2021-10-05 | 2021-08-09 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | MYSELF AND MY HUSBAND NOTICED I WAS COUGHING A LOT (IN THE MORNING) AFTER USING MY DREAMSTATION CPAP MACHINE. I CLEANED THE HUMIDITY CHAMBER (BUT NEVER NOTICED ANY SLICKNESS INSIDE) AND CHANGED OUT TUBING. EVEN AFTER THAT, THE MORNING COUGH CONTINUED. I HAD AN UPCOMING DOCTORS APPOINTMENT ((B)(6) 2021) AND MENTIONED TO DR (B)(6) I HAD A COUGH THAT HAD BEGUN ABOUT 3 WEEKS BEFORE MY VISIT. SHE LISTENED TO MY LUNGS AND SAID I WAS ALRIGHT, BUT TO LET HER KNOW IF IT WORSENED. IT'S BEEN CONSISTENT (I'M A NON SMOKER) SO I THOUGHT I SHOULD REPORT MY EXPERIENCE. ABOUT ONE OR TWO WEEKS LATER I SAW THE RECALL ON (B)(6) AND REPORTED MY MACHINE. A PERIOD OF TIME LATER, I RECEIVED A LETTER WITH INFORMATION FROM PHILIPS. HENCE WHY I AM GIVING YOU MY EXPERIENCE THAT CONTINUES. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-10
2518422-2021-04894 | 2021-10-05 | 2021-07-09 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200H11C | DSX200H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 20, 2024, AND SECTION H11 SHOULD BE REPORTED AS THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED THAT SHE HAD SURGERY DUE TO SINUS/INFECTION ISSUES, NOSE HAS SORES, RESPIRATORY PROBLEMS, AND PARTICLES IN TUBING/CHAMBER. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM (PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR). SECTION B2 WAS MADE OTHER SERIOUS (PREVIOUSLY IT WAS BLANK). SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04867 | 2021-10-05 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04828 | 2021-10-05 | 2021-09-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID, EU | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | EUX400H15 | EUX400H15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG NODULES. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON AUG 29, 2023 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG NODULES. THERE WAS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE THIRD PARTY SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED. THERE WERE NO ERRORS FOUND. THE THIRD PARTY SERVICE CENTER CONCLUDES THAT THERE WAS NO VISIBLE FOAM DEGRADATION. DEVICE WAS SCRAPPED. SECTIONS B1, B2, D8, D9, H1, H3, H6 HAS BEEN CORRECTED AND UPDATED IN THIS REPORT. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 WAS CORRECTED AND UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04854 | 2021-10-05 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP NODES IN THEIR RIGHT LUNG. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | ADDITIONAL INFORMATION WAS RECEIVED ON 07 NOV, 2024 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING TO DEVELOP NODES IN THEIR RIGHT LUNG,SINUS PROBLEMS,NAAL/THROAT IRRITATION RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM.THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04856 | 2021-10-05 | 2021-08-04 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION BIPAP PRO W/HUM/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX600H11C | DSX600H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THE ABOVE REPORTED EVENT IS ASSESSED AS POSSIBLY RELATED TO THE DEVICE IN THIS CASE. HENCE, THE DEVICE MAY HAVE CAUSE OR CONTRIBUTE TO THE ALLEGED SERIOUS INJURY. THE PATIENT DID NOT REPORT TO RECEIVE MEDICAL INTERVENTION. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04855 | 2021-10-05 | 2021-08-04 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500S11 | DSX500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP BLACK MOLD ON HER BREAST IMPLANTS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04897 | 2021-10-05 | 2020-09-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | PATIENT HAS EXPERIENCED COUG AND SORE THROAT WHICH MISSED TO MENTIONED IN PREVIOUS REPORT. AND B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM AND CAUSED COUGH AND SORE THROAT. THERE WAS NO REPORT SERIOUS OR OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6-CLINICAL CODE HAS BEEN CORRECTED/UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
MW5104416 | 2021-10-05 | 2021-04-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | I AM NOT CERTAIN OF THE ORIGINAL DATE OF SINUS INFECTION, COUGH, AND PNEUMONIA. IT WAS THE MONTH OF APRIL. I USE THE PHILIPS DREAMSTATION BIPAP MACHINE. I BEGAN EXPERIENCING BREATHING DIFFICULTIES, UPPER AIRWAY INFECTIONS, COUGH. OBTAINED CHEST X-RAY, ANTIBIOTICS. CONTINUED SIGNS OF COUGH AND DIFFICULTY BREATHING. HISTORY OF CANCER AND CHEST X-RAY WAS NOT DEFINITIVE. NECESSITATED CT SCAN SHOWING ATELECTASIS AND OPACITIES IN BOTH LUNGS. TREATED WITH ANTIBIOTICS. ALSO USED THE SOCLEAN DEVICE FOR CLEANING THE MACHINE. NOW UNDERSTAND THIS IS A PROBLEM RELATING TO DEGRADATION OF THE UNIT. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04832 | 2021-10-05 | 2021-06-15 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103792) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP AN UNDETERMINED CARCINOMA. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND BATEMANT FOAM. SECTION B5 WAS INCORRECT IN THE PREVIOUS REPEORT, IT SHOULD HAVE BEEN REPORTED AS : THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION IN A VOLUNTARY MEDWATCH (MW5103792), ALLEGING A PATIENT REPORTED TUMOR IN BOTH SIDES OF THE NECK DIAGNOSED AS UNDETERMINED CARCINOMA WITHOUT A PRIMARY CANCER SOURCE AND NOSE IRRITATION RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE REPORTED EVENT OF NOSE IRRITATION IS ASSESSED AS NON-SERIOUS INJURY AND THE REPORTED EVENT OF TUMOR IN BOTH SIDE OF THE NECK IS ASSESSED AS SERIOUS BUT NOT RELATED TO THE DEVICE. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-10
MW5104402 | 2021-10-05 | 2021-01-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I HAVE DEVELOPED A SEVERE, CONSTANT COUGH THAT SEEMS TO BE GETTING WORSE. SOME MORNINGS, MY LUNGS FELT INFLAMED. REPORTING IT TO MY DOCTOR AND TRYING TO DO EVERYTHING I COULD THINK OF TO REDUCE ALLERGIES (THINKING THAT WAS THE PROBLEM), IT WASN'T HELPING. I LEARNED YESTERDAY THAT MY CPAP MACHINE WAS BEING RECALLED BY PHILIPS RESPIRONICS AND ONE OF THE SIDE EFFECTS WAS A COUGH. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04857 | 2021-10-05 | 2021-07-07 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11 | DSX700T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO FALL AND LOSE THEIR RIGHT EYE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE PATIENT TO FALL AND LOSE THEIR RIGHT EYE.THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE EVENT IS CURRENTLY UNKNOWN. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER EVALUATED THE DEVICE EXTERNALLY /INTERNALLY AND DARK DUST OR DIRT ON THE OUTSIDE SURFACES,CRACKS,CREVICES,AROUNF THE SD AREA AND MODEM AREA,IN AND AROUND THE AIR INLET CANAL AND INSIDE SURFACE OF THE SD FLIP DOOR ,YELLOW DISCOLORATION OF BLOWER BOX,WATER SPOTS FOUND ON THE BLOWET BOX AND BOTTOM SURFACE OF THE BLOWER ALSO FOUND HAIR /PARTICLES FOUND ON THE BLOWER BOX BELOW WHERE THE BLOWER RESTS. THE MANUFACTURER OBSERVED DUST /DIRT ON THE INSIDE SURFACE OF THE BOTTOM ENCLOSURE WHERE THE BLOWER SEAL RESTS ON THE BLOWER BOX AND WHERE THE GROMMET SITS ON THE BLOWER BOX. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WAS ONE ERROR ( E-22,ERR_MOTO_FLUX_MAGNITUDE) FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THEY CONFIRMED THE PRESENCE CONTAMINANTS IN THE BLOWER BOX AND THE INSIDE SURFACE OF THE BOTTOM ENCLOSURE AND THERE IS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04777 | 2021-10-05 | 2021-07-22 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION EXPERT W/HUM/HT, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX501T12 | CAX501T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED CANCER. THE PATIENT DID NOT REPORT RECEIVING MEDICAL INTERVENTIONS REGARDING THE CANCER DIAGNOSIS. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED CANCER. THE PATIENT DID NOT REPORT RECEIVING MEDICAL INTERVENTIONS REGARDING THE CANCER DIAGNOSIS. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FOLLOW-UP REPORT WILL BE FILED. CORRECTION INFORMATION WAS PROVIDED IN SECTION D4. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04776 | 2021-10-05 | 2021-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION EXPERT W/HUM/HT, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX501T12 | CAX501T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED DIFFICUTY BREATHING. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED DIFFICUTY BREATHING. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. CORRECTED INFORMATION PROVIDED IN SECTION D4. SECTION H6 CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04779 | 2021-10-05 | 2021-07-09 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED CANCER. THE PATIENT DID NOT REPORT RECEIVING ANY MEDICAL INTERVENTIONS. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 15, 2024, AND SECTION H11 SHOULD BE REPORTED AS THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO EXPERIENCE A CANCER. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM (PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR). SECTION B2 WAS MADE OTHER SERIOUS (PREVIOUSLY IT WAS BLANK). SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECT - CLINICAL CODE AND HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04865 | 2021-10-05 | 2021-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700S11 | DSX700S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAD BREATHING PROBLEM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04863 | 2021-10-05 | 2021-08-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING /CHAMBER,CONGESTED,ASTHMA FLARES UP AFTER USING THE MACHINE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. A CORRECTION TO B5 WAS MADE AND SHOULD BE THE MANUFACTURER RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING /CHAMBER,CONGESTED,ASTHMA FLARES UP AFTER USING THE MACHINE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION H6 WAS CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04830 | 2021-10-05 | 2019-12-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC |  |  |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103744) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP AN INFECTION. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED DEVELOP AN INFECTION. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THE REPORTED EVENT OF DEVELOP AN INFECTION AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04829 | 2021-10-05 | 2019-09-11 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103753) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP PNEUMONIA. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING HEADACHE, DEEP CHEST COUGH, SORE THROAT, FEVER, PNEUMONIA RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. SECTION H6 (CLINICAL CODE) CORRECTED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04977 | 2021-10-06 | 2021-09-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLANK. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF A VENTILATOR'S LCD SCREEN BEING BLANK. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT COULD NOT BE DUPLICATED. THE DEVICE WAS FOUND TO OPERATE AND ALARM AS DESIGNED. | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04970 | 2021-10-06 | 2021-09-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR STOPPED WORKING. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN EVALUATED BY THE THIRD PARTY SERVICE CENTER. AT THIS TIME, THEY ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE THIRD PARTY INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR STOPPED WORKING. THE VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04968 | 2021-10-06 | 2021-09-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S INTERNAL BATTERY WAS FAULTY. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN EVALUATED BY THE THIRD PARTY SERVICE CENTER. AT THIS TIME, THEY ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE THIRD PARTY INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR'S INTERNAL BATTERY WAS FAULTY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04978 | 2021-10-06 | 2021-09-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S INTERNAL BATTERY WAS DEFECTIVE. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04969 | 2021-10-06 | 2021-09-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S VOLUME DELIVERY WAS LOW AND THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE VENTIALTOR WAS RECALIBRATED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR WAS RECALIBRATED TO ADDRESS A LOW VOLUME DELIVERY AND A FAILED TEST STEP DURING TESTING. DURING ADDITIONAL EVALUATION THE DEVICE'S SENSOR BOARD WAS FOUND TO BE FAULTY AND WAS REPLACED TO ADDRESS THE ISSUES. | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04979 | 2021-10-06 | 2021-09-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040000 | 1040000 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WOULD NOT DELIVER OXYGEN. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF A VENTILATOR THAT WOULD NOT DELIVER OXYGEN. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUES. | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04949 | 2021-10-06 | 2021-09-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTOCLES IN TUBING/CHAMBER, NASAL/THROAT IRRITATION OR SORENESS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 HEALTH EFFECT- CLINICAL CODE (MISSED TO UPDATE IN THE INITIAL MDR), TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04951 | 2021-10-06 | 2021-08-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NASAL/THROAT IRRITATION OR SORENESS SOMETIMES WAKES UP WITH HEADACHE AND ALSO NOTICED LINT LIKE MATERIAL, RESEMBLING DRIED UP WATER AND FLAKES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING NASAL/THROAT IRRITATION OR SORENESS SOMETIMES WAKES UP WITH HEADACHE AND ALSO NOTICED LINT LIKE MATERIAL, RESEMBLING DRIED UP WATER AND FLAKES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTION G1, H6 WAS CORRECTED AND H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04930 | 2021-10-06 | 2016-06-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103383) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP COSTOCHONDRITIS. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY SUBMITTED MDR 2518422-2021-04930-2 WITH INCORRECT SECTIONS B1, B2, H1, H6. CORRECTIONS TO PREVIOUS MDR ARE MADE IN THIS REPORT AS FOLLOWS. SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY PRODUCT PROBLEM WAS CHECKED IN PREVIOUS MDR) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY IT WAS BLANK) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6- HEALTH IMPACT CODE WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04794 | 2021-10-06 | 2021-09-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1032800 | 1032800 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR WAS ALARMING FOR A VENTILATOR INOPERATIVE ALARM CONDITION. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2021-10
MW5104455 | 2021-10-06 | 2020-04-20 | I, n, j, u, r, y | Voluntary report | DREAMSTATION ASV | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | ISSUED THE PHILIPS DREAMSTATION ASV FOR SEVERE SLEEP APNEA I BEGAN TO HAVE COUGHING FITS, LUNG IRRITATION, LUNG CONGESTION AND MUCOUS BUILD UP IN LUNGS. EVERY MORNING POST BIPAP USAGE. WHEREUPON I WOULD ABSTAIN FROM USAGE AND THE SYMPTOMS WOULD SUBSIDE. I STOPPED USING THE MACHINE IN (B)(6) 2021 AFTER BEING NOTIFIED OF THE RECALL BY MY DENTAL HYGIENIST. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-10
MW5104452 | 2021-10-06 | 2016-12-12 | I, n, j, u, r, y | Voluntary report | CPAP SYSTEM ONE AND DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DREAM STATION |  |  | I'VE BEEN USING THE PHILIPS CPAP MACHINES FOR OVER 8 YEARS, BOTH OF MY UNITS ARE SUBJECT TO THE RECALL. STARTING BACK SOMETIME IN 2016 I NOTICED SOME STRANGE ODORS COMING FROM MY CPAP AFTER STARTING THERAPY, AND AGAIN IN THE MIDDLE OF THE NIGHT AFTER BEING AWAKENED BY HEADACHES AND SOMETIMES FEELING NAUSEOUS. THIS WAS NOT ALL THE TIME, BUT STARTED TO INCREASE AS TIME WENT ON. I STARTED EXPERIENCING A LOT OF HEADACHES, AND RUNNY NOSES. THIS SEEMED TO SLOW DOWN WHEN I GOT MY NEW DREAM STATION IN 2018, AFTER ABOUT A YEAR OR SO THE SYMPTOMS STARTED TO RETURN, MORE FREQUENTLY. THE SYMPTOMS NEVER REALLY STOPPED. I'VE ALSO EXPERIENCED INCREASED SINUS INFECTIONS, UPPER RESPIRATORY INFECTIONS, UNCONTROLLABLY SNEEZING, RUNNY EYES AND HAD CONTACT DERMATITIS AROUND MY NOSE AND CHEEKS. I'M VERY CONCERNED ABOUT THE POSSIBLE AND POTENTIAL RISKS OF CHEMICAL EXPOSURES, AND TOXIC CARCINOGENIC AFFECTS DUE TO OFF-GASSING, BECAUSE OF THE MANY YEARS OF THERAPY THAT I'VE BEEN USING THESE 2 CPAP MACHINES. WAITING ON INSURANCE PERMISSION TO HAVE THE 2ND CT SCAN OF MY SINUSES. BOTH TEST WERE ORDERED ON THE SAME PRESCRIPTION REQUEST MADE BY MY DOCTOR. WAS ONLY ALLOWED THE FIRST TEST, BY THE INSURANCE COMPANY. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-10
MW5104449 | 2021-10-06 | 2021-08-01 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | RESPIRONICS DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | THE NOISE FROM MY CPAP MACHINE (HEARING MYSELF BREATHING), IS MUCH, MUCH LOUDER THAN NORMAL. I'VE HAD INCREASED AMOUNT OF HEADACHES, AND I HAD A COUPLE DAYS OF UNUSUAL DIZZINESS. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-10
MW5104441 | 2021-10-06 | 2021-05-07 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I BECAME VERY SICK THIS WEEKEND AND WAS TOLD BY EMERGENCY ROOM I HAD BRONCHOPNEUMONIA. I CONTINUE TO HAVE WORSENING HEADACHES, CHEST PRESSURE, SHORTNESS OF BREATH AND COUGH SINCE (B)(6). I WAS TOLD IN (B)(6) MY CHEST XRAY HAD CHANGES. I HAVE SEVERE SLEEP APNEA AND HAVE USED THE PHILIPS DREAMSTATION BIPAP FOR SEVERAL YEARS AND IT IS NOW ON RECALL. I HAVE NEVER BEEN DIAGNOSED WITH COPD. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04987 | 2021-10-06 | 2021-09-25 | I, n, j, u, r, y | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040004 | 1040004 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT'S BLOOD OXYGEN SATURATION LEVEL DECREASED WHILE USING A VENTILATOR. THE PATIENT WAS PLACED ON A NASAL CANNULA WITHOUT INCIDENT. THE MANUFACTURER IS CURRENTLY INVESTIGATING THIS EVENT AND A FOLLOW-UP REPORT WILL BE FILED WHEN THE INVESTIGATION IS COMPLETE. | IT WAS INITIALLY REPORTED, THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT'S BLOOD OXYGEN SATURATION LEVEL DECREASED WHILE USING A VENTILATOR. THE PATIENT WAS PLACED ON A NASAL CANNULA WITHOUT INCIDENT. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. A REVIEW OF THE DEVICE'S DOWNLOADED EVENT LOG INDICATES A PATIENT CIRCUIT DISCONNECT ALARM SOUNDED FOR 1527 SECONDS, 25.45 MINUTES, ON THE REPORTED DATE OF THE EVENT, 09/25/2021, DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. | Y | N | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04967 | 2021-10-06 | 2021-09-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04976 | 2021-10-06 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DEVICE HAD WHITE THINGS IN THE WATER DEVICE SMELLS LIKE OLD WOOD BURNING, FEELS LIKE LUNGS IS SORE AND IS NOT GETTING ENOUGH AIR FROM THE DEVICE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEVICE HAD WHITE THINGS IN THE WATER DEVICE SMELLS LIKE OLD WOOD BURNING, FEELS LIKE LUNGS IS SORE AND IS NOT GETTING ENOUGH AIR FROM THE DEVICE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04971 | 2021-10-06 | 2021-08-30 | D, e, a, t, h | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT EXPIRED WHILE USING A VENTILATOR. THE MANUFACTURER IS CURRENTLY INVESTIGATING THIS EVENT AND A FOLLOW-UP REPORT WILL BE FILED WHEN THE INVESTIGATION IS COMPLETE. | IT WAS INITIALLY REPORTED, THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT EXPIRED WHILE USING A VENTILATOR. DURING THE EVALUATION OF THE DEVICE AT THE REPORTING FACILITY, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE NEEDS REPLACED TO ADDRESS THE ISSUE. THE REPORTING FACILITY CONFIRMED THE PATIENT DEATH WAS NOT RELATED TO THE DEVICE. THE EVENT WAS RELATED TO COMPLICATIONS FROM COVID-19. | Y | N | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04870 | 2021-10-06 | 2021-09-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WOULD NOT POWER ON. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER FOR EVALUATION. THE DURABLE MEDICAL COMPANY STATES THEY HAVE NOT RECEIVED ANY COMPLAINTS AGAINST THIS DEVICE. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04788 | 2021-10-06 | 2021-07-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/HT/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400T11C | DSX400T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04786 | 2021-10-06 | 2021-07-08 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/HT/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400T11C | DSX400T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED COUGH, CONGESTION AND HEADACHE. THE PATIENT DID RECEIVE MEDICAL INTERVENTION RESULTING IN AN OFFICE VISIT TO A PHYSICIAN AND A PRESCRIBED INHALER. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED COUGH,CONGESTION AND HEADACHE.. THE PATIENT DID NOT REPORT TO RECEIVE MEDICAL INTERVENTION. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04510 | 2021-10-06 | 2021-09-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID/HT, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AUX400T15 | AUX400T15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION B5 WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT SHOULD BE REPORTED AS FOLLOWS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT REPORTED PARTICLE MATTER ENTERING NOSTRILS WHILE IN USE AT NIGHT. PATIENT ALSO REPORTED UPPER AIRWAY IRRITATION AND ITCHINESS. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. SECTION H6 HEALTH EFFECT - CLINICAL CODE WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT. IT HAS BEEN CORRECTED OR UPDATED IN THIS REPORT. UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04784 | 2021-10-06 | 2021-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED DIZZINESS, VOMITING, HEADACHE, LARYNGITIS AND WAS DIAGNOSED WITH SPASTIC DYSPHONIA. THE PATIENT DID NOT REPORT RECEIVING ANY MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT ALLEGED TO CAUSED DIZZINESS, VOMITING, HEADACHE, LARYNGITIS AND WAS DIAGNOSED WITH SPASTIC DYSPHONIA. NO OTHER CLINICAL INFORMATION WAS REPORTED. ADDITIONAL INFORMATION WAS RECEIVED AND ADDED TO THE REPORT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY ON (07/11/2023) FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED ON (07/18/2023). EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED IN THE BASE UNIT. DUST/DIRT CONTAMINATION INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM WAS OBSERVED THROUGHOUT DEVICE ENCLOSURE, AIRPATH, AND HUMIDIFIER ENCLOSURE SUGGESTING A SOURCE EXTERNAL TO THE DEVICE. SLIGHT BLACK CONTAMINATION AT THE BLOWER SEAL OF THE BLOWER BOX IS CONSISTENT WITH THE KERATIN CONTAMINATION OBSERVED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. ONE INSTANCE OF REBOOT ERROR WAS LOGGED. INVESTIGATION WAS NOT ABLE TO CONFIRM THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM BUT COULD CONFIRM THE PRESENCE OF CONTAMINATION IN THE AIRPATH. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04862 | 2021-10-06 | 2021-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04861 | 2021-10-06 | 2021-10-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04814 | 2021-10-06 | 2021-09-17 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION BIPAP AUTOSV H/HT/C, DS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX900T11C | DSX900T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP EYE ISSUES. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED ADDITIONAL INFORMATION THAT PATIENT ALSO HAVING EYE PROBLEM,AND HEADACHE . THE MEDICAL INTERVENTION WAS NOT SPECIFIED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04924 | 2021-10-06 | 2021-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11 | DSX700H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DEVICE MAKES NOISES AND HAS BLACK STUFF EVERY WHERE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEVICE MAKES NOISES AND HAS BLACK STUFF EVERY WHERE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. AFTER THE FINAL ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, CUSTOMER DECLINED TO RESPOND TO THE GFE RELATED QUESTIONS AND TERMINATED THE CALL. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04556 | 2021-10-06 | 2021-09-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AUX500T15 | AUX500T15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION B5 WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT SHOULD BE REPORTED AS FOLLOWS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED DRY COUGH AND SINUS ISSUES. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. SECTION H6 HEALTH EFFECT - CLINICAL CODE WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT. IT HAS BEEN CORRECTED OR UPDATED IN THIS REPORT. UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04668 | 2021-10-06 | 2021-07-14 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | RGBX500S15 | RGBX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS INDICATED EXPERIENCING SERIOUS FATIGUE, SHORTNESS OF BREATH, LOSS OF TEETH, HYPERSENSITIVITY, HEADACHES, DIZZINESS, EYE INFECTIONS/DRYNESS AND DEPRESSION. IT IS UNKNOWN, AT THIS TIME, IF THE PATIENT REQUIRED MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | (B)(4). | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04905 | 2021-10-06 | 2021-08-06 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP STAGE 3 KIDNEY DISEASE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING TO DEVELOP STAGE 3 KIDNEY DISEASE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURER CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REPORTED BY THE PATIENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04810 | 2021-10-06 | 2020-10-19 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103532) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP METHEMOGLOBINEMIA. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING TO DEVELOP METHEMOGLOBINEMIA,HEADACHE,DIZZINESS,TROUBLE IN BREATHING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE EVENT IS CURRENTLY UNKNOWN. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HEALTH IMPACT-CLINICAL CODE WAS MISSED TO CAPTURE IN PREVIOUS MDR,NOW IT IS CORRECTED AND UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04864 | 2021-10-06 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11 | DSX400H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. ADDITIONAL INFORMATION WAS RECEIVED, SECTION B5 WAS CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT ALLEGED BLACK SUD IN THE DEVICE, COUGH AND SORE THROAT DUE TO THE DEVICE. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04926 | 2021-10-06 | 2021-09-28 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO BE DIAGNOSED WITH LYMPHATIC CANCER. THE PATIENT DID NOT REPORT RECEIVING ANY MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04923 | 2021-10-06 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX600H11 | DSX600H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF BLACK PARTICLES RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT ALLEGED EXPERIENCING BREATHING ISSUE AND HAS LUNG PROBLEMS. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION G1 CONTACT OFFICE (AND MANUFACTURING SITE FOR DEVICES) OR COMPOUNDING OUTSOURCING FACILITY CORRECTED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04974 | 2021-10-06 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES, SORE THROAT RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 (CLINICAL CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS) UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04929 | 2021-10-06 | 2021-09-08 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103214) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SHORTNESS OF BREATH. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED A VOLUNTARY MEDWATCH (MW5103214) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SHORTNESS OF BREATH. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04900 | 2021-10-06 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS INITIALLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE - THE PATIENT ALLEGED NASAL/THROAT IRRITATION, MOLD IN WATER, WHITISH GRAYISH PARTICLES IN WATER, COUGHING. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO CONTAMINATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER AND FOUND A FOREIGN WHITE POWDERY SUBSTANCE ON BLOWER MOTOR OUTLET. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04989 | 2021-10-06 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO W/BT, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | AUG400S15 | AUG400S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION B5 WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT SHOULD BE REPORTED AS FOLLOWS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION PATIENT'S FILTERS HAVE BEEN BLACK-. FOUND INSIDE THE DEVICE. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. SECTION H6 CODES CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT HAS BEEN CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04902 | 2021-10-06 | 2021-07-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. THE MANUFACTURER PREVIOUSLY NOT CAPTURED CORRECTLY SECTION G4, AND IT HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04866 | 2021-10-06 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK AND RED STAINS IN THE TUBING AND DEVELOPED A COUGH RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK AND RED STAINS IN THE TUBING AND DEVELOPED A COUGH RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION G1 WAS CORRECTED AND H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04901 | 2021-10-06 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04903 | 2021-10-06 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11C | DSX400H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. IN INITIAL REPORT, SECTION G3 WAS WRONGLY CAPTURED AS 04/10/2021, INSTEAD OF 08/16/2021 AND SECTION G3 WAS NOT CAPTURED CORRECTLY, CORRECTED AND UPDATED THESE FIELDS IN THIS REPORT. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04899 | 2021-10-06 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11W | DSX500T11W |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04944 | 2021-10-06 | 2021-08-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | IN PREVIOUS REPORT SECTION D9 WAS INCORRECTLY CAPTURED AS: 02/24/2023 BUT IT SHOULD BE REPORTED AS: 02/09/2023. SECTION D9 HAS BEEN CORRECTED/UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04945 | 2021-10-06 | 2021-07-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING CHAMBER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGING SINUS INFECTIONS AND COUGH. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING CHAMBER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGING SINUS INFECTIONS AND COUGH. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04988 | 2021-10-06 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO W/BT, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | AUG400S15 | AUG400S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION B5 WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT SHOULD BE REPORTED AS FOLLOWS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION MACHINE ON INITIAL START AIR FELT DUSTY, WHEN HUMIDIFIER WAS DRY BLACK PARTICLES WERE FOUND AT BOTTOM. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. SECTION H6 CODES CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT HAS BEEN CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04785 | 2021-10-06 | 2021-07-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11 | DSX400H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A COUGH, NASAL/THROAT IRRITATION AND SORENESS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION RESULTING IN MEDICATIONS AND A LUNG X-RAY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A COUGH, NASAL/THROAT IRRITATION AND SORENESS. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. BASED ON INFORMATION AVAILABLE AT THE TIME OF THIS REVIEW, THERE IS NO ALLEGATION OF ANY SERIOUS INJURY. THE REPORTED EVENT OF NASAL/THROAT IRRITATION AND COUGH SINCE 2019 THAT DOES NOT IMPROVE WITH MEDICATION WAS ASSESSED AS POSSIBLY RELATED TO THE PRODUCT PROBLEM WITH THE DEVICE. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04973 | 2021-10-06 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04787 | 2021-10-06 | 2021-08-09 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/HT/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400T11C | DSX400T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED BILATERAL PULMONARY EMBOLISMS, BURNING SYMPTOMS IN EYES AND NOSE, DIZZINESS AND DEPLETED OXYGEN LEVELS. THE PATIENT DID NOT REPORT RECEIVING MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED BILATERAL PULMONARY EMBOLISMS, BURNING SYMPTOMS IN EYES AND NOSE, DIZZINESS AND DEPLETED OXYGEN LEVELS. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM (PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR) SECTION B2 WAS CHECKED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS (BLANK IN THE PREVIOUS MDR) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECT - CLINICAL AND CODE HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04972 | 2021-10-06 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED BLACK PARTICLES IN MOUTH, BAD COUGH OF MUCUS PHLEGM, DEVICE RUNS HOT SOMEDAYS, AND DEVICE RUNS OUT OF WATER QUICKLY. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND DUST/DIRT IN AND AROUND THE AIR INLET, AROUND THE ISO PORT, IN THE BLOWER, BEIGE CONTAMINANT ON THE TOP ENCLOSURE, BLOWER BOX IS DISCOLORED. PIL CAN CONFIRM THE PRESENCE OF A DUST/DIRT CONTAMINANT IN AND AROUND THE AIR INLET AND ISO PORT, TOP ENCLOSURE AND FLIP DOOR,FOUND THE BLOWER BOX DISCOLORED. PIL CAN ALSO CONFIRM THE PRESENCE OF A DARK GREY CONTAMINANT AT THE BLOWER SEAL OF THE BLOWER BOX THAT APPEARS CONSISTENT WITH POTENTIAL KERATIN OBSERVED IN ER (B)(4) (V1). PIL FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION AND BREAKDOWN. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND 4 INSTANCES OF REBOOT ERROR E-130 ERR_TASK_WDOG_TIMEOUT. PER (B)(6) V20, RECOMMENDED CORRECTIVE ACTION IS TO REPLACE PCA AND TEST. ERROR LOG SHOWS 7 INSTANCES OF E-53 ERR_COMP_LOG_SEM_ TIMEOUT. PER (B)(6) V20, RECOMMENDED CORRECTIVE ACTION IS TO REPLACE PCA AND TEST. 5888.7 BLOWER HOURS, 5836.7 THERAPY HOURS, AND 5888.7 MACHINE HOURS WERE LOGGED. THE MANUFACTURER CONCLUDES THE DEVICE HAS CONTAMINATION AROUND THE AIR INLET, AROUND THE ISO PORT, IN THE BLOWER. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04975 | 2021-10-06 | 2021-08-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. ADDITIONAL INFORMATION WAS RECEIVED, SECTION B5 WAS CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING ISSUES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT ALLEGED DEVICE WILL NOT TURN ON/DEVICE NOT FUNCTIONING, DEVICE/POWER CORD OR SUPPLY TOO HOT TO TOUCH, PARTICLES IN AIRWAY FROM DEVICE, SOMETIMES LOUDER THAN USUAL (PATIENT NOTICED IT DUE TO DAILY USE), IT WOULD THEN DISAPEAR, PATIENT WOKE UP AT NIGHT CHOCKING, AND FELT LIKE HE HAD PARTICLES IN HIS THROAT, AND HAD THROAT IRRITATION, DEVICE WAS SOMETIMES HOTTER THAN USUAL, BUT IT WOULD GO AWAY. PATIENT'S THROAT IS STILL IRRITATED AFTER USE. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. AFTER THE FINAL ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, THE CUSTOMER HUNG UP AND TERMINATED THE CALL BEFORE ASKING GFE RELATED QUESTIONS. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 WAS UPDATED AND CORRECTED BY ADDING HEALTH EFFECT - CLINICAL CODE WHICH WAS MISSED IN INITIAL REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04907 | 2021-10-07 | 2021-08-05 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11W | DSX500T11W |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | IN THE PREVIOUS REPORT, THE MANUFACTURER MENTIONED INCORRECT VERBIAGE IN SECTION H11, AS THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTION H11 HAS BEEN CORRECTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DIZZINESS, CHRONIC SORE THROAT, NASAL PASSAGES WERE SWOLLEN, NOSE IS STUFFY ALL THE TIME LIKE ALLERGIES; DIFFICULTY BREATHING, NASAL/THROAT IRRITATION OR SORENESS, DRY THROAT AND DRY NOSE, ALONG WITH OTHER FLU-LIKE SYMPTOMS, PARTICLES IN TUBING/CHAMBER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGATION OF CANCER. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05017 | 2021-10-07 | 2021-08-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION. . SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05018 | 2021-10-07 | 2021-07-09 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 21, 2024. SECTION H11 SHOULD BE REPORTED AS THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING OF HAVING NASAL/THROAT IRRITATION OR SORENESS. PATIENT ALSO ALLEGED OF HAVING CANCER, FILTERS TURNED BLACK, SORE THROAT AND RUNNING FEVER. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM ( PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR) SECTION B2 WAS CHANGED TO OTHER SERIOUS (PREVIOUSLY IT WAS BLANK) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY SECTION H6 HEALTH EFFECT - CLINICAL CODE, HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05019 | 2021-10-07 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURYTHIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04955 | 2021-10-07 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP AUTOSV H/HT/C, DS | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS,INC | DSX900T11C | DSX900T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. AFTER SECOND ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WAS UNSUCCESSFUL AND CUSTOMER DECLINED TO RESPOND TO THE GFE RELATED QUESTIONS AND TERMINATED THE CALL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04992 | 2021-10-07 | 2021-07-08 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX500T12 | CAX500T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP THROAT CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED SEVERE APNEA. PATIENT WAS TREATED FOR THROAT CANCER IN 2019. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. DURING THE EXTERIOR INVESTIGATION, UNKNOWN DUST OR DIRT CONTAMINATION WAS PRESENT ON ALL SURFACES OF THE BASE UNIT. THERE IS AN UNKNOWN DUST CONTAMINATE PRESENT AT THE AIR INLET WHERE THE FILTER WOULD BE AND AT THE ISO PORT ENTRANCE. DURING THE INTERIOR INVESTIGATION, THERE WAS UNKNOWN DEBRIS IN THE TRACKS OF THE UI PANEL. THERE IS AN UNKNOWN WHITE DUST CONTAMINATION FOUND ON THE BOTTOM ENCLOSURE, BLOWER BOX HOUSING, BLOWER MOTOR, BLOWER IMPELLER, AND REAR PANEL O-RING. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN WAS NOT OBSERVED IN THE BASE UNIT. A KERATIN-LIKE SUBSTANCE WAS OBSERVED AROUND THE BLOWER BOX OUTLET. THE UNKNOWN DUST/DIRT CONTAMINATION AND FIBERS FOUND ON THE BLOWER CASING/IMPELLER AND BLOWER BOX WERE INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM CONTAMINATION. THE PRESENCE OF CONTAMINATION THROUGHOUT THE AIRWAY SUGGESTS A SOURCE EXTERNAL TO THE DEVICE. MINERAL SPOTS CONSISTENT WITH WATER INGRESS WERE FOUND WITHIN THE BLOWER BOX AND ON THE MOTOR CASING. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS. THE MANUFACTURER CONCLUDES UNABLE TO DIRECTLY ADDRESS THE SYMPTOMS OUTLINED IN THE COMPLAINT. CONFIRMED THAT THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN OBSERVED IN THE BASE UNIT. CONFIRMED THE PRESENCE OF CONTAMINATION IN THE AIRPATH. SECTION D4 WAS CAPTURED INCORRECTLY IN THE PREVIOUS REPORT. IT HAS BEEN CORRECTED IN THIS REPORT. SECTION H6 HEALTH EFFECT: CLINICAL CODE FOR SEVERE APNEA, WHICH WAS MISSED TO CAPTURE IN PREVIOUS REPORTS, HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04834 | 2021-10-07 | 2021-09-03 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103765) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A RESPIRATORY INFECTION. THE PATIENT WAS PRESCRIBED ANTIBIOTICS IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON SEPTEMBER 25, 2023, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DEVELOP A RESPIRATORY INFECTION RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGED OF ASTHMA ( NEW OR WORSENING), LUNG DISEASE. SECTION E1 WAS UPDATED IN THIS REPORT. SECTION H6 HEALTH EFFECTS - CLINICAL CODES WAS UPDATED IN THE REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04991 | 2021-10-07 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/HT/CELL, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AUX400T15C | AUX400T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04953 | 2021-10-07 | 2021-07-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING/ CHAMBER, BLACK BUILD UP IN THE PLASTIC/ RESERVOIR TRA/ BLACK DEBRIS IN THE TUBING, NASAL DRIP, ALLERGIES, SNEEZING, DRY COUGH, HEADACHE, SINUS, HARD TO BREATH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING/ CHAMBER, BLACK BUILD UP IN THE PLASTIC/ RESERVOIR TRA/ BLACK DEBRIS IN THE TUBING, NASAL DRIP, ALLERGIES, SNEEZING, DRY COUGH, HEADACHE, SINUS, HARD TO BREATH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04950 | 2021-10-07 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION /BREAKDOWN. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS LOGGED. THE DEVICE WAS APPLIED POWER AND THE DEVICE OPERATED PROPERLY. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. SECTION D9, G3, H3 AND H6 HAS BEEN UPDATED / CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05009 | 2021-10-07 | 2021-08-17 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A STAPH INFECTION. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04947 | 2021-10-07 | 2021-09-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO AUTO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AUG500S15 | AUG500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05031 | 2021-10-07 | 2021-08-16 | I, n, j, u, r, y | Manufacturer report | BIPAP AUTO BIFLEX W/HTHUM SYSONE60SRSDOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DS760TS | DS760TS |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CHRONIC OBSTRUCTIVE PULMONARY DISEASE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION CAUSED A PATIENT TO DEVELOP CHRONIC OBSTRUCTIVE PULMONARY DISEASE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILE. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05007 | 2021-10-07 | 2021-07-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP BREAST CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED BREAST CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE DEVICE WAS EVALUATED. UNKNOWN WHITE, BROWN AND BLACK CONTAMINATION, INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM AND SOME TINY GRAY AND BLACK POTENTIAL HAIRS OR FIBERS AT THE AIR INPUT OF THE BLOWER BOX. UNKNOWN LAYER OF DUST LIKE CONTAMINATION ON TOP OF THE BLOWER BOX. LIGHT UNKNOWN DUST LIKE CONTAMINATION ON TOP OF THE BLOWER MOTOR AND BLOWER MOTOR SEAL. POTENTIAL MINERAL DEPOSIT SPOTS ON THE BOTTOM OF THE BLOWER BOX, INDICATING POTENTIAL WATER INGRESS. UNKNOWN WHITE, BROWN AND BLACK, SPECS OF CONTAMINATION AT THE BOTTOM OF THE BLOWER BOX. PIL OBSERVED BLACK CONTAMINATION, CONSISTENT WITH ER 2243857. TWO DAMAGED ANTI SLIP RUBBER PADS AT THE BOTTOM OF THE HUMIDIFIER ENCLOSURE. SLIGHT MINERAL DEPOSITS IN THE WATER TANK. POTENTIAL BLACK BIO-CONTAMINATION AND POTENTIAL FIBERS OR HAIR ON THE HUMIDIFIER LID SEAL. UNKNOWN BROWN CONTAMINATION IN THE HUMIDIFIER ENCLOSURE RIM AND ON THE OUTSIDE OF THE DRY BOX AND DRY BOX SEAL. MULTIPLE UNKNOWN BROWN AND BLACK CONTAMINANTS, INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM AND SOME POTENTIAL BLACK HAIRS OR FIBERS INSIDE OF THE HUMIDIFIER ENCLOSURE. THERE IS NO VISIBLE DAMAGE OR FUNCTIONALITY FAILURES OF THE DEVICE, WHICH SUGGEST THAT THE SOURCE OF CONTAMINATION WAS EXTERNAL TO THE DEVICE. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER AND FOUND 2 ERROR LOGGED. THE MANUFACTURER CONCLUDES THE DEVICE HAS CONTAMINATION CONSISTENT WITH WATER INGRESS. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05006 | 2021-10-07 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED BLACK STUFF SMOKE COMING OUT OF THE FILTER, IT IS CAUSING HER HEADACHE AND SORE THROAT,FEELS SICK AFTER USING CPAP . THERE WAS NO REPORT OF PATIENT HARM OR INJURY. AFTER THE FINAL ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, RECIEVED A SYSTEM GENERATED MESSAGE THAT NUMBER IS NO LONGER IN SERVICE. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05005 | 2021-10-07 | 2021-07-09 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SHORTNESS OF BREATH AND DIZZINESS. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. A CORRECTION TO THE B5 SHOULD BE : THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT DEVELOPED SHORTNESS OF BREATH AND DIZZINESS, HEADACHES, NASAL / THROAT IRRITATION / SORENESS, KIDNEY ISSUE, BLACKED OUT SIXTH TIMES IN TWO DAYS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04948 | 2021-10-07 | 2021-09-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-10
MW5104486 | 2021-10-07 | 2020-12-24 | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP WITH HUMID | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX200H11 |  |  | I WAS ISSUED A PHILIPS DREAM STATION CPAP MACHINE FOR MY SLEEP APNEA AND WAS NOTIFIED OF A RECALL ON THE MACHINE. I USED THE MACHINE FOR APPROXIMATELY 4 YEARS AND STARTED TO HAVE SEVER SINUS ISSUES. I HAVE REGISTERED MY DEVICE ON THE PHILIPS RECALL WEBSITE CONFIRMATION NUMBER (B)(4). FOR THE PAST 10 MONTHS I HAVE BEEN DEALING WITH SEVERAL SINUS ISSUES WHICH INCLUDE ON GOING SINUS INFECTIONS AND CYST GROWTHS IN MY SINUSES. I HAVE SEEN AN EAR, NOSE AND THROAT SPECIALIST FOR THE ONGOING ISSUES AND DON'T SEEM TO BE IMPROVING AT ALL. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04835 | 2021-10-07 | 2021-09-04 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION BIPAP ST30 H/C, DS | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | PHILIPS RESPIRONICS, INC. | DSX1030H11C | DSX1030H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103794) ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP AN ISSUE WITH THEIR TONSILS. THE PATIENT REQUIRES SURGERY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED A VOLUNTARY MEDWATCH (MW5103794) ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP AN ISSUE WITH THEIR TONSILS. THE PATIENT REQUIRES SURGERY IN RESPONSE TO THE REPORTED EVENT. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-10
MW5104470 | 2021-10-07 | 2018-09-01 | I, n, j, u, r, y | Voluntary report | DREARNSTATION AUTO CPAP DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I HAVE USED PHILIPS RESPIRONICS CPAP MACHINES FOR MANY YEARS. AS TIME WORE ON, I BEGAN HAVING SERIOUS REACTIONS TO USE OF THE MACHINES. THE PROBLEMS GRADUALLY BEGAN WITH VERY EXCESS MUCOUS BEING PRODUCED WHILE USING THE MACHINES. THAT PROBLEM PROGRESSED INTO IRRITATION OF THE NASAL PASSAGE AND EVENTUALLY BLEEDING IN THAT AREA AND IN MY NOSE. I ALSO BEGAN HAVING SERIOUS NASAL DISCHARGE DURING THE DAY. THE PROBLEM BECAME SO BAD THAT I HAVE HAD TO STOP USING MY CPAP. I HAVE TRIED SEVERAL TIMES TO RESUME USING MY CPAP BY TRYING DIFFERENT PHILIPS NASAL MASKS AND HUMIDITY LEVELS BUT THE PROBLEMS DID NOT STOP. |  | Y | N | LAY USER/PATIENT |  | 2021-10
MW5104472 | 2021-10-07 |  | I, n, j, u, r, y | Voluntary report | DREAM STATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | FOLLOWING AN ACUTE RESPIRATORY INFECTION, I FOUND OUT MY CPAP WAS ON RECALL. I WAS INSTRUCTED TO DISCONTINUE USE AS I HAD USED AN OZONE CLEANER WHICH RESULTED IN BLACK DEPOSITS ON FILTER. I ENDED UP IN INTENSIVE CARE ON 8 LITERS OF OXYGEN AND VIRAL PNEUMONIA WITH ACUTE RESPIRATORY FAILURE AND HYPOXIA SENT HOME ON OXYGEN, AND NOW ON MEDICATION TO BREATHE. EVER SINCE I STARTED USING THE DREAM STATION I BEGAN TO SUFFER FROM RECURRING RESPIRATORY INFECTIONS, I TOLD MY PC THAT I FELT IT WAS THE MACHINE THAT WAS THE CULPRIT, I USED AN OZONE CLEANER OUT OF DESPERATION AS I THOUGHT IT MIGHT BE DIRTY AND WAS CAUSING ME TO BE SICK, I STARTED USING THE CLEANER AND WAS SETTING A BLACK DISCHARGE ON MY FILTER WHICH I MISTAKENLY THOUGHT WAS DIRT OR MOLD AS I HAVE A HEATED HUMIDIFIER ON THE UNIT. I STILL WAS HAVING IRRITATION AND STARTED USING ALBUTEROL TO ALLEVIATE MY SYMPTOMS. MY LAST INFECTION WAS IN (B)(6) 2021 THAT IS WHEN I FOUND OUT ABOUT THE RECALL OF THIS MACHINE. PHILIPS TOLD ME TO STOP USING IT, MY INFECTION PROGRESSED INTO ACUTE RESPIRATORY FAILURE AND HOSPITALIZATION BEFORE I COULD SEE MY SLEEP APNEA DOCTOR. I NOW HAVE DAMAGE TO MY RESPIRATORY SYSTEM AS WELL AS COPD. MY MACHINE WAS ONE MONTH SHORT OF BEING 5 YEARS OLD. |  | Y | N | LAY USER/PATIENT |  | 2021-10
MW5104473 | 2021-10-07 | 2014-09-19 | I, n, j, u, r, y | Voluntary report | PHILIPS BIPAP AUTO (BI-FLEX) SYSTEM ONE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | REF 750P |  |  | WHILE USING MY PHILIPS BIPAP MACHING I DEVELOPED CHRONICALLY INFECTED NASAL PASSAGES WHICH MY PHYSICIAN TREATED WITH INHALANT STEROIDS-NASOCORT. IN THE PAST 2 YEARS I DEVELOPED A RASH ACROSS THE BRIDGE OF MY NOSE WHICH AT TIMES ULCERATED AND WAS TREATED WITH TOPICAL ANTIBIOTICS AND BANDAGE UNTIL THE SKIN HEALED. THIS PROCESS WAS REPEATED EVERY 2 MONTHS OR SO. THE TREATMENT DATES AND TYPES FOR THIS CONDITION ARE ON FILE AT (B)(6) WHERE (B)(6) WAS A PATIENT OF DR. (B)(6) AT THE TIME OF THE CHRONIC SINUSITIS. |  | Y | N | LAY USER/PATIENT |  | 2021-10
MW5104481 | 2021-10-07 | 2021-01-06 | I, n, j, u, r, y | Voluntary report | PHILIPS CPAP DEVICE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | J268482024B75 |  |  | IN JANUARY AND FEBRUARY I SMELLED SOMETHING THAT WAS LIKE OIL, ALSO NOTICED IT WOULD GET LOUD AT TIMES. I HAVE A KIDNEY DISEASE AND NOT SURE IF THE PHILIPS CPAP MACHINE HAS WORSEN MY MEDICAL CONDITION. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05008 | 2021-10-07 | 2021-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP AN UNSPECIFIED RESPIRATORY ISSUE. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT: IN THE PREVIOUS REPORT, SECTION B1 WAS MARKED ADVERSE EVENT AND PRODUCT PROBLEM INCORRECTLY, WHICH HAS BEEN CORRECTED TO REFLECT THE PRODUCT PROBLEM. IN THE PREVIOUS REPORT, SECTION B2 WAS MARKED HOSPITALIZATION (INITIAL OR PROLONGED), WHICH HAS BEEN UNMARKED TO REFLECT THE PRODUCT PROBLEM. IN THE PREVIOUS REPORT, SECTION H1 WAS MARKED SERIOUS INJURY, WHICH HAS BEEN CORRECTED TO REFLECT MALFUNCTION. SECTION H6 HEALTH EFFECT - IMPACT CODE HAS ALSO BEEN UPDATED TO REFLECT THE PRODUCT PROBLEM. SECTION D1 BRAND NAME AND G4 PREMARKET IDENTIFICATION WERE UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-10
MW5104482 | 2021-10-07 | 2021-06-14 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAMSTATIONGO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS INC. | DSG500S11 |  |  | I WAS ADVISE BY PHILIPS TO REGISTER MY CPAP PORTABLE DEVICE DUE TO ITS RECALL AND STOP USING IT. I REGISTERED ON 06/25/2021 AND GOT THE CONFIRMATION NUMBER (B)(4). I HAVE CALLED PHILIPS 3 TIMES ASKING FOR AN UPDATE AND THEY ARE UNRESPONSIVE. I NEED A SOLUTION TO FIX MY MACHINE, REPLACE WITH AN ANOTHER BRAND OR REIMBURSE ME, SO I CAN PURCHASE ANOTHER BRAND. I STOP BREATHING W/OUT A CPAP MACHINE AND I CANNOT AFFORD TO BUY ANOTHER ONE. THEY ARE EXPENSIVE AND NOT COVERED BY MY INSURANCE. I HAVE SPENT ALMOST (B)(6) WITH THE PHILIPS DREAMSTATIONGO. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05021 | 2021-10-07 | 2021-08-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | UDSX500S11 | UDSX500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05057 | 2021-10-07 | 2021-10-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11 | DSX700T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04912 | 2021-10-07 | 2021-09-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04906 | 2021-10-07 | 2021-09-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE AND FOUND DUST IN THE BLOWER AND BLUE PIECE OF PLASTIC IN BOTTOM OF BLOWER BOX. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND 1 ERROR LOGGED. THE MANUFACTURER CONCLUDES NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. THE MANUFACTURER ALSO OBSERVED DUST CONTAMINATION AND FOREIGN PLASTIC PIECE PRESENT IN THE AIRPATH. IN THE INITIAL REPORT ALLEGATION OF VISUALIZATION OF BLACK PARTICLES, NASAL/THROAT IRRITATION OR SORENESS, COUGHING, PARTICLES INSIDE LUNGS WERE NOT MENTIONED. IN THIS REPORT WHICH HAS BEEN UPDATED. SECTION D9, G3, H6 HAS BEEN UPDATED / CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05003 | 2021-10-07 | 2021-09-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-05000 | 2021-10-07 | 2021-09-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A FAILURE OF THE DEVICE TO POWER ON WAS OBSERVED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04999 | 2021-10-07 | 2021-09-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO OXYGEN FLOW WAS OBSERVED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND MANIFOLD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-05002 | 2021-10-07 | 2021-09-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS OUT OF CALIBRATION. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04889 | 2021-10-07 | 2021-09-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04910 | 2021-10-07 | 2021-08-12 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO BE DIAGNOSED WITH ASTHMA. THE PATIENT DID NOT REPORT RECEIVING ANY MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING ASTHMA RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND DUST/DIRT CONTAMINATION FOUND ON THE BOTTOM CASING, MOTOR, IMPELLERS, AND BLOWER BOX WAS INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM CONTAMINATION. THE PRESENCE OF CONTAMINATION THROUGHOUT THE AIRPATH, SUGGESTS A SOURCE EXTERNAL TO THE DEVICE. SLIGHT BLACK CONTAMINATION AT THE BLOWER SEAL OF THE BLOWER BOX APPEARS CONSISTENT WITH THE KERATIN CONTAMINATION OBSERVED . THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGREDATION. THE DEVICE DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERROR. THE MANUFACTURER CONCLUDES THE DEVICE HAS DUST/DIRT CONTAMINATION FOUND ON THE BOTTOM CASING, MOTOR, IMPELLERS, AND BLOWER BOX WAS INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM CONTAMINATION. THE PRESENCE OF CONTAMINATION THROUGHOUT THE AIRPATH, SUGGESTS A SOURCE EXTERNAL TO THE DEVICE. SLIGHT BLACK CONTAMINATION AT THE BLOWER SEAL OF THE BLOWER BOX APPEARS CONSISTENT WITH THE KERATIN CONTAMINATION OBSERVED. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04888 | 2021-10-07 | 2021-09-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04886 | 2021-10-07 | 2021-09-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1032804 | 1032804 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICES LCD DISPLAY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04885 | 2021-10-07 | 2021-09-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICES LCD DISPLAY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04884 | 2021-10-07 | 2021-09-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04883 | 2021-10-07 | 2021-09-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04881 | 2021-10-07 | 2021-09-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04880 | 2021-10-07 | 2021-09-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICES LCD DISPLAY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04697 | 2021-10-07 | 2021-09-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A SERVICE REQUIRED ALARM OCCURRED RELATED TO LOW CIRCUIT LEATK. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE ALSO FAILED A TEST STEP. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. | THIS MDR IS BEING SUBMITTED AS PART OF A BATCH SUBMISSION OF PREVIOUSLY CLOSED COMPLAINTS THAT WERE REASSESSED AS REPORTABLE FOAM DEGRADATION COMPLAINTS; DISCOVERED AS PART OF A RETROSPECTIVE REMEDIATION REVIEW. THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A SERVICE REQUIRED ALARM OCCURRED RELATED TO LOW CIRCUIT LEATK. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE ALSO FAILED A TEST STEP. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. CONTAMINATION WAS OBSERVED ON THE FLOW SENSOR ASSEMBLY, BLOWER, AND STIRRING FAN FOAM. THE FLOW SENSOR ASSEMBLY, BLOWER AND STIRRING FAN FOAM WERE REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04887 | 2021-10-07 | 2021-09-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICES LCD DISPLAY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-05060 | 2021-10-07 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING CONGESTION IN THROAT MAYBE FROM POOR AIR QUALITY IN AREA OR THE MACHINE, SHE ALSO HAS DIZZINESS WHEN SHE SITS UP IN THE MORNING IN HER BED RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04928 | 2021-10-07 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4) | THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN THE PREVIOUSLY SUBMITTED REPORT SECTION B5 WAS INCORRECT, CORRECT B5 SECTION SHOULD BE- THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED VOLUNTARY MEDWATCH (MW5103365). THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND NO EVIDENCE OF CONTAMINATION ON THE OUTSIDE OF THE DEVICE. AN INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE CONTAMINATION ON THE BACK PANEL OF THE CASE. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE HUMIDIFIER WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR EVALUATION. THE MANUFACTURER FOUND EVIDENCE OF CONTAMINATION INSIDE OF THE HUMIDIFIER. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND NO ERRORS LOGGED. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF CONTAMINATION ON THE BACK PANEL OF THE CASE, AND INSIDE OF THE HUMIDIFIER. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. SECTION D9, D10, G3 AND H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04911 | 2021-10-07 | 2021-08-09 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE CONTACT DERMATITIS, RASH, EYE IRRITATION AND ASTHMA. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF MEDICATION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 13, 2024, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO A TUMOR IN THE NECK WHILE USING THE DEVICE.THIS NOTIFICATION WAS REVIEWED BY THE PMS CLINICAL EXPERT DUE PATIENT REPORTING OF ASTHMA WHICH IS ASSESSED AS SERIOUS INJURY. THE PATIENT DID NOT REPORT TO RECEIVE MEDICAL INTERVENTION. SECTION B1 WAS CORRECTED FOR BOTH ADVERSE EVENT AND PRODUCT PROBLEM. ( PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION. TO SERIOUS INJURY. SECTION H6 WAS CORRECTED AND UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04936 | 2021-10-07 | 2021-08-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX500S11 | DSX500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON MAR 21 , 2023 AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING ONGOING SINUS & THROAT PROBLEM, SMALL BLACK FLECKSHAVE BEEN IN MY MUCUSFOR 2 YEARS, REPORTED TO MY DOCTOR WITH NO SOURCE FOUND WITH ONE KNOW ALLERGY CEDAR RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM.THE PATIENT DID NOT REPORT TO RECEIVE MEDICAL INTERVENTION. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF DEVICE WAS INSPECTED. THE DEVICE POWERED ON, AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY MANUFACTURER. THERE WERE NO ERRORS FOUND.THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT WAS SCRAPPED. IN THIS REPORT, SECTIONS D9, G3,H2, H3, AND H6 HAVE BEEN UPDATED OR CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04935 | 2021-10-07 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION.THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 HAS BEEN UPDATED IN THIS REPORT . | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04938 | 2021-10-07 | 2021-10-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL, AU | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AUX500T15C | AUX500T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION B5 WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT SHOULD BE REPORTED AS FOLLOWS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED UNWELL, HEADACHES AND HAS SOME SINUS ISSUES. PATIENT ALSO REPORTED BLACK SPOTS NOTED ON THE HUMIDIFIER. THERE WAS NO PATIENT HARM OR INJURY. NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. SECTION H6 HEALTH EFFECT - CLINICAL CODE WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT. IT HAS BEEN CORRECTED OR UPDATED IN THIS REPORT. UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04932 | 2021-10-07 | 2021-08-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK LIKE MELTED PLASTIC FILM ON HUMIDIFIER SEAL COULD HAVE INHALED THAT BLACK JUNK FOR A WHILE WITHOUT KNOWING IT RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04934 | 2021-10-07 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP PRO W/HUM/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX600H11C | DSX600H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04927 | 2021-10-07 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP W/HUMID | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. CORRECTIVE INFORMATION WAS PROVIDED IN B.5 SECTION (IN PREVIOUS MDR, THE "PATIENT HAS BEEN WAKING UP WITH HEADACHES, BROWN COLOR CONDENSATION COMING FROM THE MASK, DRY NOSE, PATIENT ALSO WENT INTO AFIB AND STRANGE ODOUR CANNOT DESCRIBE" WAS NOT INCLUDED) | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05024 | 2021-10-07 | 2021-09-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04908 | 2021-10-07 | 2021-08-05 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE A RASH AND UNSPECIFIED SINUS ISSUE. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF SINUS SURGERY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE EVENT IS CURRENTLY UNKNOWN. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05020 | 2021-10-07 | 2021-08-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO EVIDENCE OF CONTAMINATION. THE MANUFACTURER COMPLETED AN INTERNAL VISUAL INSPECTION. THE MANUFACTURER FOUND EVIDENCE OF CONTAMINATION ON THE DRY BOX AS WELL AS ON THE RESERVOIR SEAL. THE MANUFACTURER FOUND NO SIGNS OF CONTAMINATION INSIDE OF THE BLOWER BOX. THE HUMIDIFIER WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR EVALUATION. THE MANUFACTURER FOUND EVIDENCE OF CONTAMINATION IN THE BOTTOM OF THE INSIDE OF THE HUMIDIFIER. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS LOGGED. THE DEVICE WAS APPLIED POWER AND THE DEVICE OPERATED PROPERLY. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. SECTION D9, D10, G3, H3 AND H6 HAS BEEN UPDATED / CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05025 | 2021-10-07 | 2021-09-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04995 | 2021-10-07 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. PLACED CALL TO CUSTOMER FOR UNDERLYING HEALTH CONDITIONS, DEVICE RETURN AND ADDITIONAL INFORMATION REGARDING THE THERMAL ISSUES. CUSTOMER HUNG UP AND TERMINATED THE CALL BEFORE I CAN ASK GFE RELATED QUESTIONS. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04994 | 2021-10-07 | 2021-10-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05029 | 2021-10-07 | 2021-08-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP AN UNSPECIFIED INFECTION, AND SWELLING IN THEIR NECK. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP AN UNSPECIFIED INFECTION, AND SWELLING IN THEIR NECK. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO EVIDENCE OF CONTAMINATION. THE MANUFACTURER COMPLETED AN INTERNAL VISUAL INSPECTION. THE MANUFACTURER FOUND EVIDENCE OF CONTAMINANT ON THE BLOWER BOX, BLOWER, BLOWER SEAL, AND ON THE UI PANEL. THE MANUFACTURER ALSO FOUND EVIDENCE OF KERATIN-LIKE CONTAMINATION WAS OBSERVED AROUND THE BLOWER OUTPUT SEAL. THE MANUFACTURER FOUND EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS LOGGED. THE DEVICE WAS APPLIED POWER AND THE DEVICE OPERATED PROPERLY. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH AN UNKNOWN CONTAMINATION AND KERATIN-LIKE CONTAMINATION. THE MANUFACTURER CONFIRMED THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05030 | 2021-10-07 | 2021-08-17 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SPOTS ON THEIR LUNGS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04943 | 2021-10-07 | 2021-08-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED BY THE MANUFACTURER. THE MANUFACTURER FOUND NO SIGNS OF CONTAMINATION ON THE OUTSIDE OF THE DEVICE. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND SIGNS OF CONTAMINATION IN THE INLET PORT. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS LOGGED.THE DEVICE WAS HOOKED UP TO POWER SUPPLY, AIRFLOW WAS VERIFIED AND THE DEVICE OPERATED PROPERLY. THE MANUFACTURER CONCLUDES NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. IN THE INITIAL REPORT ALLEGATION OF HEADACHE WAS NOT MENTIONED WHICH HAS BEEN UPDATED IN THIS REPORT. SECTION D9, G3, H6 HAS BEEN UPDATED / CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04913 | 2021-10-07 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WAS 1 ERROR FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. SECTION H6 WERE UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05023 | 2021-10-07 | 2021-09-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05001 | 2021-10-07 | 2021-09-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD AND MACHINE FLOW SENSOR WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-05068 | 2021-10-08 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION COUGHING AND DRY EYES, BLACK SOOT IN WATER TANK RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
MW5104532 | 2021-10-08 | 2021-09-01 | I, n, j, u, r, y | Voluntary report | DREAM STATION BIPAP VENTILATOR MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I HAVE BEEN USING A NON-CONTINUOUS VENTILATOR (CPAP/BIPAP) FOR THE LAST 13 YEARS, AND HAVE USED A PHILIP RESPIRONICS DREAM STATION BRAND FOR THE LAST 3 YEARS. ONE MORNING THE MACHINE STARTED MAKING A LOUD "WHOOSHING" SOUND, AS IF IT WAS LEAKING A LOT OF AIR, AND WHICH WOKE ME UP AT 6 AM. I USED THE MACHINE FOR THE NEXT 2 NIGHTS, AND IT WAS "WHOOSHING" ON AND OFF ALL THE TIME. I CALLED MY MEDICAL SUPPLIER, AND WAS TOLD ABOUT THE PHILIPS RECALL, AND STOPPED USING IT IMMEDIATELY. MY DOCTOR AND SUPPLIER ARE WORKING ON GETTING ME A REPLACEMENT UNIT. I RECEIVED A LETTER FROM PHILIPS, ADVISING ME OF THE RECALL, AND THE PROBLEMS THEY HAD HAD WITH THE MACHINES. THEIR LETTER INCLUDED POSSIBLE SIDE-EFFECTS, AND THAT THE RECALL INVOLVED PARTICLES OF FOAM AND OFF-GASSING OF CERTAIN CHEMICALS. PHILIPS ALSO SUGGESTED THAT I REPORT ANY PROBLEMS WITH THE FDA MEDWATCH. I HAVE NOT SHOWN SYMPTOMS OF ANY DEFINITE ADVERSE REACTIONS OR ACTIVE PROBLEMS. HOWEVER, I FELT I SHOULD REPORT TO YOU THAT FOR APPROXIMATELY 3 MONTHS PRIOR TO THIS SUDDEN EVENT, I HAD WOKEN UP IN THE MORNINGS WITH A GUMMY, STICKY FEELING IN MY MOUTH, LIKE MY TEETH WERE STICKING TO MY GUMS. I RINSED WITH WATER AND IT WENT AWAY. ALSO, SINCE ABOUT THE END OF (B)(6) 2021, SINCE I STOPPED USING THE MACHINE, I HAVE FELT BETTER AND MORE ENERGIZED EVERY DAY, AND AM NOT SLEEPING 10 HOURS A NIGHT, AS I DID WHEN I WAS USING THE MACHINE. IT'S AS IF IT HAD A DEPRESSANT EFFECT ON ME. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04956 | 2021-10-08 | 2021-07-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE SECTION B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS REPORT.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04957 | 2021-10-08 | 2023-01-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO, GER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DEX400S13 | DEX400S13 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER RECEIVED AN ISSUE RELATED TO A CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES SOUND ABATEMENT FOAM BECOME DEGRADED AND ALLEGED CHRONIC SINUSITIS. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. DESPITE OF MULTIPLE ATTEMPTS THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER RECEIVED NEW AND RELEVANT INFORMATION ON 01/04/2022 PATIENT REFUSES TO RETURN THE DEVICE FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT. IN ADDITION, APPROPRIATE HEALTH EFFECT - CLINICAL CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS HAS BEEN ADDED ON H6 SECTION. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05027 | 2021-10-08 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN THE PREVIOUSLY SUBMITTED REPORT SECTION B5 WAS INCOMPLETE, B5 SECTION WILL BE- THE PATIENT ALLEGED VISUALIZATION OF PARTICLES. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF UNKNOWN CONTAMINANT ON THE UI PANEL, KERATIN-LIKE SUBSTANCE AROUND THE BLOWER BOX OUTLET AND AN UNKNOWN DUST CONTAMINATION ON THE REAR PANEL O-RING, BLOWER, BLOWER SEAL, INSIDE THE BLOWER, AND IN THE BLOWER BOX. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND NO ERRORS LOGGED. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF CONTAMINATION ON THE UI PANEL, KERATIN-LIKE SUBSTANCE AROUND THE BLOWER BOX OUTLET AND AN DUST CONTAMINATION ON THE REAR PANEL O-RING, BLOWER, BLOWER SEAL, INSIDE THE BLOWER, AND IN THE BLOWER BOX. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. SECTION D9, G3 AND H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04958 | 2021-10-08 | 2021-09-28 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, EU | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS INC | EUX500H15 | EUX500H15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP KIDNEY CANCER. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE EVENT IS CURRENTLY UNKNOWN. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05052 | 2021-10-08 | 2021-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING NASAL/THORAT IRRITATION RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. IN INITIAL REPORT, SECTION H6 CLINICAL CODES ARE NOT CAPTURED, CORRECTED AND UPDATED IN THIS REPORT. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04682 | 2021-10-08 | 2021-09-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | RGBX500S15 | RGBX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED. THERE WAS NO PATIENT INJURY OR HARM REPORTED. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING THAT A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED. THERE WAS NO PATIENT INJURY OR HARM REPORTED. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF WHITE AND BLACK UNKNOWN CONTAMINATION IN THE AIR INPUT OF THE BLOWER BOX, A WHITE, BROWN, CREAM COLORED AND BLACK UNKNOWN CONTAMINATION IN THE BOTTOM OF THE BLOWER BOX. THE MANUFACTURER OBSERVED MANY POTENTIAL MINERAL DEPOSITS ON THE BLOWER MOTOR SIDE, BLOWER MOTOR BOTTOM, BLOWER MOTOR ISOLATORS, AND IN THE BLOWER BOX, INDICATING POTENTIAL WATER INGRESS. THE MANUFACTURER FOUND EVIDENCE OF BLACK CONTAMINATION, CONSISTENT WITH KERATIN, AT THE AIR OUTLET OF THE BLOWER BOX, UNKNOWN LIGHT BROWN STAINS ON THE INSIDE ENCLOSURE. THE HUMIDIFIER WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR EVALUATION. THE MANUFACTURER FOUND MULTIPLE UNKNOWN BROWN AND BLACK CONTAMINANTS AND DARK-COLORED HAIR STRANDS INSIDE OF HUMIDIFIER BOX, SOME POTENTIAL MINERAL DEPOSITS ON IT AND SOME WHITE CONTAMINATION ON HEATER PLATE AND DRY BOX SEAL, POTENTIAL MINERAL DEPOSITS WERE ON THE BOTTOM RIM OF THE HEATER PAN JUST OUTSIDE OF THE HEATER PLATE GASKET. THE MANUFACTURER FOUND NO EVIDENCE OF DEGRADED SOUND ABATEMENT FOAM. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND NO ERRORS LOGGED. THE MANUFACTURER POWERED ON THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH BLACK CONTAMINATION CONSISTENT WITH KERATIN AND WHITE AND BLACK UNKNOWN CONTAMINATION, WHITE, BROWN, CREAM COLORED AND BLACK UNKNOWN CONTAMINATION WERE INCONSISTENT WITH THE SOUND ABATEMENT FOAM. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION D9, G3 AND H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04669 | 2021-10-08 | 2020-06-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | GBX500H15 | GBX500H15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED. THE PATIENT HAS INDICATED BEING DIAGNOSED WITH BIBASAL SMALL PLATE ATELECTASIS, SCARRING OF THE LUNGS AND A PARAMEDIASTINAL MASS ON THE LUNGS. THE PATIENT CONTINUES TO EXPERIENCE TIREDNESS, IRRITABILITY, LACK OF ENERGY AND SHORTNESS OF BREATH. IT IS UNKNOWN WHAT MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING THAT A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED. THE PATIENT HAS INDICATED BEING DIAGNOSED WITH BIBASAL SMALL PLATE ATELECTASIS, SCARRING OF THE LUNGS AND A PARAMEDIASTINAL MASS ON THE LUNGS. THE PATIENT CONTINUES TO EXPERIENCE DIFFICULTLY OF BREATHING/SHORTNESS OF BREATH, CHEST PAIN, TIRED, IRRITABLE AND LOW ENERGY. NO MEDICAL INTERVENTION WAS REPORTED. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. INVESTIGATION OBSERVED AN UNKNOWN DUST CONTAMINANT INSIDE THE BLOWER BOX AT THE AIR INLET. AN UNKNOWN CONTAMINANT WAS OBSERVED INSIDE THE ISO PORT OF THE REAR PANEL, ON THE TOP ENCLOSURE, AND FRONT PANEL EXTERIORS. AN UNKNOWN DUST CONTAMINANT WAS OBSERVED ON THE INTERIOR OF THE FRONT PANEL, REAR PANEL, BOTTOM ENCLOSURE, BLOWER, BLOWER SEAL, AND BLOWER BOX AND DUST IN THE AIRPATH WAS CONFIRMED. EVIDENCE OF LIQUID INGRESS WITH A DUST CONTAMINANT CONSISTENT WITH MINERAL DEPOSITS WAS OBSERVED TO THE BLOWER AND BLOWER BOX. A KERATIN-LIKE SUBSTANCE WAS OBSERVED ON THE BLOWER BOX OUTLET. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED. INVESTIGATION CONFIRMED NO PRESENCE OF DEGRADED SOUND ABATEMENT FOAM USING A FITT AND THE ASSOCIATED PROCEDURE. INVESTIGATION CAN CONFIRM THAT THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN OBSERVED IN THE BASE UNIT. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05033 | 2021-10-08 | 2020-05-08 | I, n, j, u, r, y | Manufacturer report | RESPIRONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103632) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP GASTRITIS. THE PATIENT REQUIRED BIOPSIES IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO DEVELOP GASTRITIS. THE PATIENT REQUIRED BIOPSIES IN RESPONSE TO THE REPORTED EVENT. THE REPORTED EVENT OF GASTRITIS AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05035 | 2021-10-08 | 2021-06-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103717) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A SINUS INFECTION. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05014 | 2021-10-08 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SCARRING OF THEIR LUNG. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 07, 2024, AND SECTION H11 SHOULD BE REPORTED AS THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SCARRING OF THEIR LUNG. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM (ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR). SECTION B2 WAS MADE BLANK (PREVIOUSLY OTHER SERIOUS WAS CHECKED). SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05053 | 2021-10-08 | 2021-08-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED (B)(4). | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05079 | 2021-10-08 | 2021-09-21 | I, n, j, u, r, y | Manufacturer report | RESPIRONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | FRX500S14 | FRX500S14 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP PROSTATE CANCER AND EYE ISSUES. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP PROSTATE CANCER AND EYE ISSUES. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT.REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05054 | 2021-10-08 | 2021-08-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING SEEING PARTICLES IN THE HUMIDIFIER, HEADACHES, NAUSEA, DIZZINESS, LIGHTEDNESS (LIGHT HEADEDNESS), NASAL CONGESTION, NASAL DRAINAGE, WHEEZING IN LUNGS THROUGHOUT THE DAY RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. CONTAMINATION APPEARING CONSISTENT WITH INSECT CONTAMINATION WAS OBSERVED ON THE INTERIOR OF THE DEVICE ENCLOSURE. DARK CONTAMINATION APPEARING CONSISTENT WITH KERATIN CONTAMINATION OBSERVED. PIL IS UNABLE TO DIRECTLY ADDRESS THE SYMPTOMS ADDRESSED IN THE COMPLAINT. PIL IS UNABLE TO ASSESS THE COMPLAINT AGAINST THE HUMIDIFIER AS NO HUMIDIFIER WAS RETURNED WITH THE DEVICE. PIL CAN CONFIRM THE PRESENCE OF CONTAMINATION IN THE AIRPATH CONSISTENT WITH KERATIN SUGGESTING A SOURCE EXTERNAL TO THE DEVICE. PIL FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS LOGGED. THE DEVICE WAS HOOKED UP TO POWER SUPPLY, AIRFLOW WAS VERIFIED AND THE DEVICE OPERATED PROPERLY. THE MANUFACTURER CONCLUDES NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. THE MANUFACTURER ALSO CONFIRMED THE PRESENCE OF CONTAMINANTS. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04933 | 2021-10-08 | 2021-08-07 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A SORE THROAT WITH MUCOUS AND INFLAMMATION. THE PATIENT DID RECEIVE MEDICAL INTERVENTION RESULTING IN PRESCRIBED MEDICATIONS. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 07, 2024 AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A SORE THROAT WITH MUCOUS AND INFLAMMATION. THE PATIENT DID RECEIVE MEDICAL INTERVENTION RESULTING IN PRESCRIBED MEDICATIONS. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 WAS CORRECTED AND UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04981 | 2021-10-08 | 2021-10-06 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID/HT, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AUX400T15 | AUX400T15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED IRRITATION TO THE SKIN, EYES AND RESPIRATORY TRACT, HEADACHES, INCREASED KIDNEY LEVELS, SINUS SYMPTOMS, NAUSEA AND DIZZINESS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION BY SEEING A PHYSICIAN AND TAKING PRESCRIBED ANTIBIOTICS AND EYE DROPS. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04941 | 2021-10-08 | 2021-06-08 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A HEART ATTACK AND WAS DIAGNOSED WITH CORONARY HEART INSUFFICIENCY. THE PATIENT DID RECEIVE MEDICAL INTERVENTION THAT REQUIRED HOSPITALIZATION, INSERTION OF A PACEMAKER, QUADRUPLE BYPASS SURGERY AND CARDIAC REHABILITATION. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY SUBMITTED 2518422-2021-04941-1 WITH INCORRECT SECTIONS B1, B2, H1, H6. CORRECTIONS TO PREVIOUS MDR ARE MADE IN THIS REPORT AS FOLLOWS. SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY PRODUCT PROBLEM WAS CHECKED IN PREVIOUS MDR) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY IT WAS BLANK) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6- HEALTH IMPACT CODE WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05065 | 2021-10-08 | 2021-08-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS INITIALLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE - THE PATIENT ALLEGED PARTICLES IN HUMIDIFIER WATER CHAMBER AND CHRONIC DRY MOUTH. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO CONTAMINATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER AND FOUND CONTAMINATION IN THE ISO PORT. THE DUST/ DIRT CONTAMINATION FOUND ON THE MOTOR CASING/ IMPELLERS, BOTTOM PANEL, BACK OF LCD PANEL, AND BLOWER BOX, WAS INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM CONTAMINATION. THERE WAS SLIGHT BLACK CONTAMINATION AT THE BLOWER SEAL OF THE BLOWER BOX AS WELL. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERORRS. THE MANUFACTURER CONCLUDES THAT THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION FOUND WITHIN THE DEVICE. SECTION B5 AND H6 WERE CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05066 | 2021-10-08 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT ALSO ALLEGED PARTICALS IN THE WATER, HEADACHE. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL PART OF THE DEVICE AND FOUND MINIMAL SIGNS OF DUST CONTAMINATION IN THE HUMIDIFIER CHAMBER. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND FOUND MINIMAL DUST CONTAMINATION IN THE BLOWER ASSEMBLY. THE HUMIDIFIER SHOWED A SMALL AMOUNT OF LIQUID INGRESS UNDER THE BOTTOM PLATE. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND ONE ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION, BUT A SMALL AMOUNT OF LIQUID INGRESS AND DUST CONTAMINATION WERE OBSERVED AND ARE CONSISTENT WITH BEING FROM AN EXTERNAL SOURCE. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05056 | 2021-10-08 | 2021-10-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05067 | 2021-10-08 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION B5 WAS CORRECTED IN THIS REPORT AND REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED THEY HAVE FOUND PARTICLES IN THE LUNG AND EXPERIENCED BREATHING ISSUE. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05032 | 2021-10-08 | 2017-05-09 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103644) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP AN EYE INFECTION, AND EXERCISE INDUCED ASTHMA. THE PATIENT WAS PRESCRIBED PRESCRIPTION EYE DROPS IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED A TO CPAP DEVICE'S SOUND ABATEMENT FOAM. IN THE PREVIOUS REPORT, SECTION B5 WAS INCORRECT, IT SHOULD HAVE BEEN REPORTED AS : THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION IN A VOLUNTARY MEDWATCH (MW5103644), ALLEGING A PATIENT DEVELOPED LEFT EYE INFECTION, EXERCISE INDUCED ASTHMA, LIGHTHEADNESS AND DIZZINESS, SORE THROAT RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT WAS PRESCRIBED ANTIBIOTIC EYE DROPS IN RESPONSE TO THE REPORTED EVENT. THE REPORTED EVENT THAT THE PATIENT WAS DIAGNOSED WITH "EXERCISE INDUCED ASTHMA" IS ASSESSED AS A SERIOUS INJURY BUT NOT RELATED TO THE DEVICE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURER CONCLUDES NO FURTHER ACTION IS NECESSARY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-10
MW5104513 | 2021-10-08 | 2020-10-20 | I, n, j, u, r, y | Voluntary report | PHILIPS TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | PHILIPS RESPIRONICS, INC. |  |  |  | THE TRILOGY 100 IS ON RECALL. MY FAMILY MEMBER (B)(6) IS FULLY DEPENDENT ON THIS MACHINE TO SUSTAIN HER LIFE. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | I |  | 2021-10
2518422-2021-05026 | 2021-10-08 | 2021-09-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05012 | 2021-10-08 | 2021-08-04 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A MYOCARDIAL INFARCTION. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY SUBMITTED MDR 2518422-2021-05012-1 WITH INCORRECT SECTIONS B1, B2, H1, H6. CORRECTIONS TO PREVIOUS MDR ARE MADE IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY PRODUCT PROBLEM WAS CHECKED IN PREVIOUS MDR). SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY IT WAS BLANK). SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6- HEALTH EFFECT - IMPACT CODE WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-10
MW5104533 | 2021-10-08 | 2020-10-01 | I, n, j, u, r, y | Voluntary report | DREAMSTAAUTOCPAPHUMCELL DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500H11C |  |  | PERSISTENT MORNING COUGH; I SWITCHED TO A SLEEP APNEA MACHINE FROM PHILIPS IN (B)(6) 2019. AFTER USING FOR ABOUT A YEAR, I DEVELOPED A MORNING COUGH WITH LOTS OF CONGESTION. I BELIEVE IT IS RELATED TO A DEFECT IN THE SLEEP APNEA MACHINE WITH THE SOUND ABATEMENT FOAM THAT I INHALED AND THE OFF-GAS CHEMICALS FROM THE USE OF THE SO-CLEAN MACHINE. THERE IS NO EXPLANATION FOR THE COUGH THAT MY DOCTOR COULD FIND WITHOUT COSTLY FURTHER TESTING. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-10
MW5104541 | 2021-10-08 | 2019-09-01 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | AFTER USING THE PHILIPS DREAMSTATION I WAS HAVING BREATHING PROBLEMS. DOCTOR DIAGNOSED PLEURISY. I HAVE NOTICED SMALL BLACK DEBRIS IN MY MASK ON MANY OCCASIONS BUT I DIDN'T THINK IT WAS COMING FROM THE MACHINE. I WOULD JUST CLEAN MY MASK OUT AND CONTINUE TO USE. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05028 | 2021-10-08 | 2021-10-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO SERIES | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | 751P | 751P |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-04629 | 2021-10-08 | 2021-09-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE OXYGEN BLENDING MODULE BOARD WAS OBSERVED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04627 | 2021-10-08 | 2021-09-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. THE DEVICE'S LCD SCREEN WAS OBSERVED TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04630 | 2021-10-08 | 2021-09-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04628 | 2021-10-08 | 2021-09-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE TO POWER ON WAS OBSERVED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-05075 | 2021-10-08 | 2021-09-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS AUTO CYCLING. THE DEVICE WAS NOT IN PATIENT USE. THE MANUFACTURER IS CURRENTLY INVESTIGATING THIS EVENT AND A FOLLOW-UP REPORT WILL BE FILED WHEN THE INVESTIGATION IS COMPLETE. | IT WAS INITIALLY REPORTED, THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS AUTO CYCLING. THE DEVICE WAS NOT IN PATIENT USE. REPEATED ATTEMPTS FOR ADDITIONAL INFORMATION OR TO HAVE THE DEVICE RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04882 | 2021-10-08 | 2021-09-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-05013 | 2021-10-08 | 2021-08-04 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200S11 | DSX200S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP AN UNSPECIFIED HEART ISSUE. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING HEART ISSUES, AND SEVERE HEADACHES RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION.DUST/DIRT CONTAMINATION INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM WAS OBSERVED THROUGHOUT DEVICE ENCLOSURE, AND AIRPATH, SUGGESTING A SOURCE EXTERNAL TO THE DEVICE.EVIDENCE OF WATER INGRESS ON THE BLOWER AND BLOWER BOX.THREAD LIKE CONTAMINANT FOUND IN BLOWER BOX,SLIGHT BLACK CONTAMINATION AT THE BLOWER SEAL OF THE BLOWER BOX IS CONSISTENT WITH THE KERATIN CONTAMINATION OBSERVED IN ER 2243857. UPON EXTERNAL AND INTERNAL ASPECT OF THE DEVICE.THERE WAS NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT UNABLE TO ADDRESS THE SYMPTOMS DESCRIBED. PIL CAN CONFIRM THE PRESENCE OF CONTAMINATION IN THE AIRPATH. SECTION H6 WERE UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05039 | 2021-10-08 | 2021-09-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-05043 | 2021-10-08 | 2021-10-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-05040 | 2021-10-08 | 2021-09-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR A LOW INSPIRATORY PRESSURE ALARM CONDITION. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05038 | 2021-10-08 | 2021-09-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SENSOR BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-05037 | 2021-10-08 | 2021-09-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-05042 | 2021-10-08 | 2021-09-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S POWER MANAGEMENT BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-05036 | 2021-10-08 | 2021-09-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN EVALUATED BY THE THIRD PARTY SERVICE CENTER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE THIRD PARTY INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR WAS ALARMING FOR A VENTILATOR INOPERATIVE ALARM CONDITION. REPEATED ATTEMPTS TO HAVE THE DEVICE EVALUATED WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-05063 | 2021-10-08 | 2021-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND NO ISSUES. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND A BLACK PLASTIC-LIKE PARTICLE ON BOTTOM ENCLOSURE, AN UNKNOWN CONTAMINANT CONSISTENT WITH FOAM ABATEMENT BREAKDOWN INSIDE THE BLOWER BOX, A STICKY CONGEALED CONTAMINANT WAS OBSERVED ON THE BLOWER BOX, BLOWER SEAL, AND AROUND THE BLOWER OUTPUT, LIQUID INGRESS ON THE BLOWER. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND THE FOLLOWING ERRORS E-53 , E-191 , E-130, E-101. THE MANUFACTURER CONCLUDES EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. THE MANUFACTURER ALSO CONFIRMED CONTAMINATION CONSISTENT WITH FOAM ABATEMENT BREAKDOWN IN THE DEVICE. IN THE INITIAL REPORT ALLEGED SYMPTOMS OF HEADACHES, SINUS INFECTIONS, WATERY EYES, SHORTNESS OF BREATH, NAUSEA WERE NOT MENTIONED WHICH HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05010 | 2021-10-08 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, AHP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | AHX500H11 | AHX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP EMPHYSEMA. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO EXPERIENCE SOAR THROAT AND NOISY DEVICE. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENT OF COPD EMPHYSEMA AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05011 | 2021-10-08 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CHRONIC PULMONARY OBSTRUCTIVE DISEASE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO HAVE DEVELOP CHRONIC PULMONARY OBSTRUCTIVE DISEASE AND COUGHING OUT BLOOD. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENT OF COPD AND COUGHING OF BLOOD AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. IN INITIAL REPORT, SECTION G3 WAS WRONGLY CAPTURED AS (B)(6) 2021, SHOULD BE (B)(6)2021, SECTION H6 CLINICAL CODE WAS NOT CAPTURED, CORRECTED AND UDPATED THESE FEILDS IN THIS REPORT. SECTION G1 WAS UPDATED IN THIS REPORT. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-10
2518422-2021-05041 | 2021-10-08 | 2021-09-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS DIM. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD SCREEN NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-10
2518422-2021-04940 | 2021-10-08 | 2021-07-08 | I, n, j, u, r, y | Manufacturer report | BIPAP AUTO BIFLEX W/HTHUM SYSONE60SRSDOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DS760TS | DS760TS |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED HEADACHES, STOMACH ISSUES AND HIGH LEVELS IN THE PATIENT'S LIVER. THE PATIENT DID RECEIVE MEDICAL INTERVENTION RESULTING IN AN ADMISSION TO THE HOSPITAL AND MEDICAL TESTING. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED HEADACHES, STOMACH ISSUES AND HIGH LEVELS IN THE PATIENT'S LIVER. THE PATIENT DID RECEIVE MEDICAL INTERVENTION RESULTING IN AN ADMISSION TO THE HOSPITAL AND MEDICAL TESTING. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN VISUAL INSPECTION WAS PERFORMED BY THE MANUFACTURER. THE MANUFACTURER FOUND CONTAMINATION ON THE AIR INLET. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS LOGGED. THE MANUFACTURER CONCLUDES NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. SECTION B2, D9, G3, H6 HAS BEEN UPDATED / CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-10

## Monthly Summary
Month | Count | Date
2021-04 | 251 | 2021-04-01
2021-05 | 327 | 2021-05-01
2021-06 | 416 | 2021-06-01
2021-07 | 736 | 2021-07-01
2021-08 | 1093 | 2021-08-01
2021-09 | 1073 | 2021-09-01
2021-10 | 336 | 2021-10-01

## Event Types
Event Type | Count
M, a, l, f, u, n, c, t, i, o, n | 3121
I, n, j, u, r, y | 1066
D, e, a, t, h | 44
N, o, , a, n, s, w, e, r, , p, r, o, v, i, d, e, d | 1

## Brand Names
Brand Name | Count
DREAMSTATION AUTO CPAP | 691
TRILOGY 100 | 687
TRILOGY EVO | 314
TRILOGY 202 | 251
TRILOGY O2 | 222
DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | 188
DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | 147
TRILOGY EV300 | 113
DREAMSTATION AUTO BIPAP | 99
DREAMSTATION | 71
DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | 68
DREAMSTATION CPAP | 64
DREAMSTATION AUTO CPAP, ND | 51
DREAMSTATION CPAP PRO | 50
TRILOGY 200 | 47
DREAMSTATION AUTO CPAP W/HUMID, DOM | 41
DREAMSTATION AUTO BIPAP W/HUM/CELL, DOM | 28
PHILIPS DREAMSTATION | 23
DREAMSTATION AUTO BIPAP/HUM/HT/CELL,DOM | 20
DREAMSTATION BIPAP PRO | 18
DREAMSTATION AUTO CPAP, DOM | 17
DREAMSTATION CPAP W/HUMID, DOM | 16
DREAMSTATION AUTO CPAP W/HUM/HT/CELL, AU | 15
DREAMSTATION CPAP PRO W/HUM/CELL, DOM | 15
DREAMSTATION CPAP PRO W/HUMID/HT, AUS | 14
DREAMSTATION AUTO BIPAP W/HUMID/HT, DOM | 14
DREAMSTATION GO AUTO | 14
DREAMSTATION AUTO CPAP, JP | 13
PHILIPS RESPIRONICS DREAMSTATION | 13
PHILIPS DREAMSTATION CPAP | 12
DREAMSTATION AUTO CPAP W/HUMID/HT, CAN | 12
DREAMSTATION CPAP PRO W/HUM/HT/CELL, DOM | 12
DREAMSTATION CPAP PRO W/HUM/HT/CELL, AUS | 12
TRILOGY EVO O2 | 11
CPAP | 11
DREAMSTATION AUTO CPAP, GER | 10
RESPIRONICS DREAMSTATION | 9
DREAM STATION | 9
DREAMSTATION AUTO BIPAP W/HUMID, DOM | 9
DREAMSTATION BIPAP AUTOSV H/HT/C, DS | 9
DREAMSTATION AUTO CPAP, INTL | 9
TRILOGY 02 | 9
DREAMSTATION CPAP PRO W/HUMID, DOM | 9
DREAMSTATION CPAP PRO W/HUMID/HT, CAN | 8
PHILLIPS RESPIRONICS DREAMSTATION | 8
PHILIPS RESPIRONICS CPAP | 8
DREAMSTATION BIPAP | 8
BIPAP AUTO, INTL | 8
DREAMSTATION BIPAP AUTOSV | 7
PHILIPS RESPIRONICS DREAMSTATION CPAP | 7
DREAMSTATION CPAP PRO, GER | 7
PHILIPS DREAMSTATION CPAP MACHINE | 7
DREAMSTATION EXPERT W/HUM/HT/CELL, CAN | 7
DREAMSTATION BIPAP AUTOSV H/C, DS | 7
DREAMSTATION AUTO, FR | 7
DREAM STATION CPAP | 6
DREAMSTATION AUTO CPAP W/HUMID/HT, AUS | 6
DREAMSTATION GO AUTO W/BT, DOM | 6
DREAMSTATION AUTO CPAP W/HUM/HT/WIFI,DOM | 6
PHILIPS CPAP | 6
RESPIRONICS DREAMSTATION CPAP | 6
BIPAP AUTO | 6
PHILLIPS DREAMSTATION | 6
DREAMSTATION CPAP PRO W/HUMID/HT, DOM | 6
DREAMSTATION CPAP MACHINE | 6
PHILLIPS DREAMSTATION CPAP | 6
DREAMSTATION AUTO CPAP, GB | 6
DREAMSTATION CPAP PRO W/HUM/HT/CELL, CAN | 5
DREAMSTATION BIPAP PRO W/HUM/CELL,DOM | 5
DREAMSTATION EXPERT W/HUM/HT, CAN | 5
BIPAP AUTO SERIES ASSY, INT | 5
DREAMSTATION GO AUTO W/BT | 5
BIPAP AVAPS C SERIES INTL 30, FLOW GEN | 5
CPAP DREAMSTATION | 4
BIPAP AUTO, DOM | 4
PHILIPS DREAMSTATION BIPAP | 4
DREAMSTATION AUTO CPAP W/HUMID, AHP | 4
BIPAP AUTO BIFLEX | 4
DREAMSTATION EXPERT | 4
BIPAP AVAPS C SERIES | 4
DREAMSTATION EXPERT W/HUM/CELL, CAN | 4
BIPAP AUTO SERIES ASSY, DOM | 3
DREAMSTATION AUTO BIPAP, GER | 3
DREAMSTATION ASV | 3
DREAMSTATION AUTO CPAP MACHINE | 3
DREAMSTATION BIPAP PRO W/HUMID, DOM | 3
BIPAP AUTO SERIES ASSY | 3
DREAMSTATION AUTO CPAP W/HUM/HT/CELL,CAN | 3
DREAMSTATION 1 | 3
DREAMSTATION BIPAP AUTO SV | 3
PHILIPS DREAM STATION CPAP | 3
PHILIPS DREAMSTATION AUTO CPAP | 3
DREAMSTATION CPAP DEVICE | 3
DREAMSTATION CPAP W/HUMID/HT, DOM | 3
DREAMSTATION AUTO CPAP HUM DOM | 3
CPAP DREAM STATION | 3
DREAMSTATION AUTO BIPAP/HUM/HT/CELL,CAN | 3
PHILIPS RESPIRONICS DREAM STATION CPAP | 3
DREAM STATION AUTO CPAP | 3
BIPAP AUTO, CN | 3
BIPAP S/T C SERIES | 3
DREAMSTATION AUTO CPAP W/HUM/WIFI, DOM | 3
DREAMSTATION BIPAP ST30 | 3
PHILIPS RESPIRONICS TRILOGY 100 | 3
PHILLIPS RESPIRONICS DREAMSTATION BIPAP | 3
DREAMSTATION CPAP W/HUM/CELL, DOM | 3
DREAMSTATION CPAP PRO W/HUMID | 3
DREAMSTATION AUTO CPAP W/HUMID | 3
BIPAP A40 | 3
BIPAP PRO BIFLEX | 3
DREAMSTATION AUTO CPAP W/HUM/CELL | 3
BIPAP AVAPS CSERIES | 2
DREAMSTATION CPAP AUTO | 2
DREAM MACHINE | 2
BIPAP S/T C SERIES INTL 30, FLOW GEN | 2
DREAMSTATION CPAP WITH HUMIDIFIER | 2
TRILOGY | 2
PHILIPS DREAM MACHINE BIPAP | 2
DREAMSTATION BIPAP PRO W/HUM/HT/CELL,DOM | 2
DREAMSTATION EXPERT W/HUM/HT/CELL | 2
DREAMSTATION AUTO CPAP W/HUMID, GB | 2
DREAMSTATION CPAP W/HUMID, CAN | 2
DREAMSTATION AUTO BIPAP W/HUMID/HT, CAN | 2
PHILLIPS DREAMSTATION CPAP MACHINE | 2
DREAMSTATION ST30 H/HT, CN-NB | 2
DREAMSTATION CPAP W/HUMID | 2
TRILOGY EV 300 | 2
DREAMSTATION GO AUTO CPAP | 2
BIPAP AUTO SERIES | 2
CPAP MACHINE DREAM STATION | 2
PHILLIPS DREAM STATION | 2
PHILIPS DREAMSTATION AUTO CPAP HUMCELL DOM | 2
DREAMSTNAUTOCPAPHUMHTCELL DOM | 2
2
CPAP MACHINE | 2
BIPAP PRO | 2
BIPAP AUTOSV ADV | 2
TRILOGY 100 VENTILATOR | 2
DREAMSTATION GO AUTO W/BT, AUS | 2
DREAMSTATION CPAP W/HUM/HT/CELL, DOM | 2
DREAMSTATION BIPAP AVAPS30 H, DS | 2
PHILIPS RESPIRONICS BIPAP | 2
DREAMSTNAUTOCPAPHUMCELL DOM | 2
BIPAP S/T, C-SERIES, CHINA | 2
BIPAP AUTOSV ADV W/HUM,SYS ONE,60SRS,DOM | 2
BIPAP AUTOSV | 2
DREAMSTATION AUTO CPAP W/HUMID, BL | 2
BIPAP A40, INTERNATIONAL | 2
DREAMSTATION BIPAP PRO, DOM | 2
DREAMSTATION CPAP PRO, CAN | 2
CONTINUOUS VENTILATOR | 2
DREAMSTATION S/T | 2
RENTAL-DREAMSTATION AUTO CPAP | 2
DREAMSTATION EXPERT W/HUM, CAN | 2
DREAMSTATION GO W/BT, AUS | 2
DREAMSTATION AUTO BIPAP HUM DOM | 2
PHILIPS RESPIRONICS SYSTEM ONE | 2
PHILLIPS RESPIRONICS DREAMSTATION AUTO CPAP | 2
PHILIPS RESPERONICS DREAMSTATION | 2
RESPIRONICS BIPAP | 2
BIPAP AUTO BIFLEX W/HTHUM SYSONE60SRSDOM | 2
PHILIPS DREAM STATION CPAP MACHINE | 2
BIPAP AVAPS C SERIES INTL 30 | 2
PHILLIPS RESPIRONICS DREAM STATION | 2
DREAMSTATION BIPAP AVAPS30 | 2
DREAMSTATION AUTO BIPAP/HUM/HT/CELL | 2
PHILIPS RESPIRONICS CPAP MACHINE | 2
DREAMSTATION C-PAP MACHINE | 2
DREAMSTATIONGO | 2
PHILIPS RESPIRONICS | 2
PHILIPS RESPIRONICS DREAM STATION | 2
VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | 2
DREAMSTATION 2 | 2
PHILIPS DREAMSTATION 1 | 2
DREAMSTATION BIPAP ST30 H/C, DS | 2
DREAMSTATION HUMID | 2
RECERT - DREAMSTATION AUTO CPAP | 2
PHILIPS DREAMSTATION CPAP DEVICE | 2
DREAMSTATION AUTO CPAP W/HUMID/HT, INTL | 2
CPAP MACHINE DREAMSTATION | 2
DREAM STATION 1 CPAP | 1
DREAMSTATION BIPAP AVAPS30 AAM H/HT/C,CA | 1
DREAMSTATION CPAP W/HUM/HT/CELL | 1
REMSTAR PRO, -RESPIRONICS INC.- SYSTEM ONE | 1
PHILIPS RESPIRONICS DREAMSTATION ASV CPAP | 1
DREAMSTATION HUM CARE PACK DOM | 1
PHILIPS DREAM STATION C-PAP | 1
DREAM STATION BIPAP PRO | 1
DREAMSTAAUTOCPAP | 1
DREAM STATION AUTO BIPAP | 1
PHILIPS RESPERONICS DREAMSTATION 1 CPAP | 1
DREAMSTAAUTOCPAPHUMCELL DOM | 1
PHILIPS RESPERONICS DREAM STATION C-PAP | 1
DREAMSTATION AUTO CPAP HUM DON | 1
REMSTAR PRO C-FLEX + | 1
PHILIPS / RESPIRONICS, INC. | 1
PHILIPS TRILOGY 100 | 1
PHILIPS RESPIRONIC CPAP | 1
SYSTEM ONE REMSTAR AUTO A-FLEX | 1
DREAMSTATION AUTO BIPAP W/HUMID, AHP | 1
TRILOGY MACHINE | 1
DREAMSTATION CPAP PRO, INTL | 1
RESPIRONICS SYSTEM ONE | 1
DREAMSTATION AUTOBIPAP | 1
PHILIPS CPAP MACHINES | 1
BIPAP MACHINE | 1
PHILIP'S RESPIRONICS DREAMSTATION AUTO CPAP | 1
PHILIPS DREAM STATION | 1
DREAMSTATION AUTOCPAPHUMCELL | 1
PHILIPS RESPIRONICS REMSTAR PRO C-FLEX+ | 1
PHILIPS REMSTAR AUTO A-FLEX CPAP | 1
BIPAP S/T CSERIES | 1
DREAMSTATION AUTO CPAP, ES | 1
DREAMSTATION HUM CORE PACK DOM | 1
REMSTAR AUTO A-FLEX | 1
RENTAL-DREAMSTATION AUTO CPAP, GB | 1
DREAMSTATION CPAP V1.0.2690 | 1
DREAMSTATION AUTO BIPAP W/HUMID, CAN | 1
PHILLIPS DREAMSTATION CPAP CPAP PRO AUTO CPAP | 1
RESPIRONICS DREAM STATION AUTO CPAP | 1
BIPAPAUTOSVADSYS0NE60SRSDOM | 1
PHILIPS RESPIRONICS REMSTAR SYSTEM ONE | 1
DREAMSTATION AUTO CPAP W/HUM/HT/WIFI,CAN | 1
DREAM MAKER CPAP | 1
PHILIPS DREAM STATION WITH HUMIDIFIER AND WIFI | 1
RESPIRONICS REMSTAR CPAP MACHINE | 1
DREAMSTATION CPAP PRO, DOM | 1
DREAMSTATION GO AUTO CPAP W/HUMID, CAN | 1
SYSTEM ONE AND DREAMSTATION | 1
BIPAP A40 PRO | 1
PHILIPS SYSTEM ONE | 1
DREAMSTATION BIPAP AUTOSV, DS | 1
CPAP SYSTEM ONE AND DREAMSTATION | 1
DREAMSTATION BIPAP ST30 H, DS | 1
SYSTEM ONE BIPAP AUTOSV ADV | 1
BIPAP AUTO BIFLEX W/SD CARD, HUMID, US | 1
PHILIPS RESPIRONICS SYSTEMONE 60 CPAP REMSTAR PRO DOM | 1
REMSTAR AUTO | 1
PHILIPS RESPIRONICS DREAMSTATION BIPAP | 1
RESPIRONICS REMSTAR AUTO M SERIES A FLEX | 1
PHILIPS RSPIRONCS DREAM STATION CPAP | 1
REMSTAR AUTO A-FLEX CPAP MACHINE | 1
B-PAP | 1
DREAMSTATION BY PHILIPS RESPRONICS | 1
DREAMSTATION AUTO CPAP DOM | 1
DREAMSTATION CPAP WITH HUMID | 1
BIPAP A30 | 1
DREARNSTATION AUTO CPAP DOM | 1
PHILIPS BIPAP AUTO (BI-FLEX) SYSTEM ONE | 1
BIPAP AUTO BIFLEX, SYS ONE 60 SRS, INT | 1
PHILIPS CPAP DREAMSTATION | 1
PHILIPS DREAMSTATION C-PAP | 1
DREAMSTATION DEVICE | 1
PHILIPS RESPIRTRONICS CPAP | 1
DREAMSTATION AUTO CPAP W/HUMID, EU | 1
BIPAP AUTOSV ADV,SYS ONE,ASSY,60 SRS,DOM | 1
DREAMSTATION CPAP HUMID HT DOM | 1
PHILIPS DREAMSTATIN AUTO CPAP | 1
BIPAP S/T C SERIES CORE PACKAGE | 1
DREAMSTATION CPAP PRO W/HUMID, EU | 1
TRILOGY 202 | 1
BIPAP SYSTEM ONE | 1
BIPAP A30, CANADA, HEATED TUBE CORE PKG | 1
TRILOGY 100 | 1
BIPAP AVAPS C SERIES CORE PACKAGE, DOM | 1
DREAMSTATION1 | 1
PHILIPS RESPIRONICS DREAMSTATION AUTOCPAPHUMHT DOM | 1
THE DREAMSTATION AUTOBIPAP MACHINE | 1
RESPIRONICS DREAMSTATION CPAP WITH HEATED HUMIDIFIER | 1
PHILIPS RESPIRONICS SLEEP AND RESPIRATORY CARE | 1
BIPAP AUTOSV ADV SYSTEM ONE ASSY,INTL | 1
CPAP REMSTAR PRO CFLEX +_ | 1
DREAMSTATION CPAP PRO W/HUM/CELL, CAN | 1
DREAMSTATION AUTO CPAP W/HUMID, CAN | 1
TRILOGY VENTILATOR | 1
BIPAP S/T | 1
DREAMSTATION BIPAP ST | 1
DREAMSTATION BIPAP PRO W/HUMID, LA | 1
RESMED AIRCURVE 10 VAUTO CPAP MACHINE | 1
DREAMSTATION AUTO CPAP WITH HEATED HUMIDIFIER | 1
PHILLIPS RESPIRONICS DREAMSTATION CPAP MACHINE | 1
DREAM MACHINE. C-PAP | 1
DREAMSTATION AUTO CPAP. | 1
PHILIPS RESPIRONICS - SYSTEM ONE | 1
PHILIPS - DREAM STATION | 1
DREAMSTATION AUTO CPAP, BR | 1
DREAMSTATION BIPAP AVAPS/ST | 1
PHILIPS DREAMSTATION AUTO CPAP HUM DOM | 1
PHILIPS CPAP DEVICE | 1
RESPIRONICS REMSTAR SE DOM | 1
SYSTEM ONE CPAP | 1
DREAM STATION BIPAP VENTILATOR MACHINE | 1
DREAMSTATION AUTO CPAP, KR | 1
DREAMSTATION AUTO CPAP, W/HUMID/HT, KR | 1
DREAMSTATION HUMCORC PACK DOM | 1
DREAMSTATION BIPAP AVAPS30 H/C, DS | 1
DREAMSTATION AUTO CPAP W/HUM/CELL, CAN | 1
REMSTAR PRO C-FLEX | 1
REMSTAR SE AUTO | 1
PHILIPS RESPIRONICS REMSTAR AUTO | 1
CPAP BRAND PHILLIPS RESPIRONIC DREAMSTATION | 1
DREAMSTATION CPAP, AR | 1
PHILIPS DREAMSTATION AUTO CPAP W/HUMIDIFER/HEATER | 1
PHLLIPS DREAM STATION CPAP | 1
DREAMSTATION BI-LEVEL PAP DEVICE | 1
RESPIROINCS PHILIPS TRILOGY EVO | 1
SYSTEM ONE Q-SERIES | 1
REMSTAR PRO C-FLEX SYSTEM ONE | 1
DREAMSTATION AUTO BI-PAP | 1
SYSTEMONE ASV4 | 1
BIPAP PRO, INTL | 1
TRILOGY EV300, USA | 1
DREAMSTATION CPAP PRO W/HUMID, CAN | 1
DREAMSTAATION CPAP | 1
SYSTEM ONE 60 SERIES BIPAP AUTO BIFLEX | 1
DREAMSTATION CPAP / BIPAP MACHINE | 1
V60 BIPAP | 1
PHILIPS RESPIRONICS DREAMSTATION AUTO-CPAP | 1
REMSTAR AUTO A-FLEX CPAP | 1
DREAMSTATION AUTO | 1
DREAMSTATION AUTO CPAP W/HUMID/HT, CN | 1
RESPIRONICS PHILIPS DREAMSTATION | 1
BIPAP AVAPS C SERIES , CHINA, FLOW GEN | 1
BIPAP AUTO BIFLEX, W/SD CARD, US | 1
PHILIPS RESPIRONICS DREAM STATION WITH HEATED HUMIDIFIER | 1
PHILLIPS CPAP DREAM STATION | 1
RESPIRONICS BIPAP AUTO SV | 1
BIPAP PRO BIFLEX, SYS ONE 60 SRS, DOM | 1
REMSTAR AUTO, DOM | 1
DERAMSTATION CPAP | 1
CPAP NAME DREAMSTATION AUTO CPAP HUM DOM | 1
RESPIRONICS BIPAP AUTO, SYSTEM ONE | 1
BIPAP PRO BIFLEX, W/SD CARD, HUMID, US | 1
PHILIPS RESPIRONICS BIPAP PRO | 1
DREAMSTATION CELLULAR MODEM | 1
PHILIPS DREAMMACHINE | 1
DREAMSTATION AUTOCPAPHUMHT DOM | 1
DREAMSTATION AUTOSV GB | 1
PHILLIPS RESPRONICS DREAMSTATION | 1
DREAMSTATION ACPAP | 1
DREAM STATION C-PAP MACHINE | 1
DREAM SM AUTO CPAP HUM CU-DOM | 1
DREAMSTN AUTO CPA | 1
DREAMSTATION BIPAP AVAPS | 1
PHILIPS RESPIRONICS DREAM MAKER CPAP | 1
PHILIPS DREAMSTATION BIPAP MACHINE | 1
DREAMSTATIONSTNAUTOCPAPHUMCELL DOM DREAMSTATION STN AUTO CPAP HUM CELL DOM | 1
"DREAMSTATION" CPAP | 1
PHILIPS RESPIRONICS DREAM MACHINE FOR SLEEP APNEA | 1
DREAMSTATION BIPAP ASV | 1
DREAMSTATION HUMIDIFIER | 1
DREAMSTATION CPAP W/HUMID/HT | 1
PHILLIPS DREAM STATION CPAP | 1
REMSTAR SE AUTO CPAP | 1
RESPIRONICS DREAM STATION | 1
DREAM STATION AUTO CPAP WITH HUMIDIFIER | 1
DREAMSTNAUTOCPAPHUMHTWIFIDOM | 1
REMSTAR AUTO A-FLEX, SYS ONE, 60 SRS,GER | 1
BIPAP S/T C SERIES INTL 30, FLOW GEN | 1
BIPAP AUTO BIFLEX, SYS ONE 60 SRS, DOM | 1
CPAP MACHINE (DREAMSTATION) | 1
PHILIPS BREATHING MACHINE | 1
PHILIPS SLEEP CPAP, DREAMSTATION | 1
PHILIPS RESPIRONICS SYSTEM ONE CPAP MACHINE | 1
PHILIPS DREAMSTATION PAP DEVICE | 1
PHILIPS DREAM STATION AUTO BIPAP | 1
SO CLEAN.2 | 1
DREAMSTATION AUTO CPAP WITH HUM DOM | 1
REMSTAR PRO C-FLEX + PHILIPS RESPIRONICS CPAP MACHINE | 1
DREAMSTATION GO W/BT | 1
BIPAP VISION | 1
OMNILAB ADVANCED PLUS | 1
PHILIPS RESPIRONICS CPAP BY PAP | 1
DREAMSTATION AUTO CPAPHUMCELL DOM | 1
DREAM STATION AUTO CPAP HUM HT CELL DOM | 1
DREAMSTATION 1 CPAP | 1
TRILOGY SR | 1
PHILIPS RESPIRONICS SYSTEM ONE CPAP | 1
DREAMSTATION AUTO BIPAP, DOM | 1
DREAMSTATION AUTO CPAP HUM HT DOM DX500T11 | 1
PHILIPS DREAMSTATION AUTO-CPAP | 1
DREAMSTATION (CPAP MACHINE) | 1
REMSTAR AUTO AFLEX ASSY | 1
DREAMSTATION1 CPAP | 1
BIPAP AVAPS C SERIES, INTERNATIONAL 30 | 1
TRILOGY SPIKED ACET SHELL 54MM | 1
PHILIPS V60 BIPAP | 1
BIPAP PRO SERIES ASSY | 1
DREAMSTATION RESPIRONICS | 1
PHILIPS RESPIRONICS SYSTEM 1 BIPAP | 1
PHILIPS | 1
PHIIIPS RESPIRONIC BIPAP MACHINE | 1
DREAMSTATION GO PORTABLE AUTO CPAP SYSTEM WITH HEATED HUMIDIFIER AND DREAMWEAR F | 1
PHILIPS RESPIRONICS DREAMSTATION CPAP DEVICE | 1
DREAMSTATION 1 CPAP MACHINE | 1
PHILIPS DREAMSTATION BI-LEVEL PAP | 1
PHILIPS DREAM MAKER CPAP MACHINE | 1
PHILLIPS DREAMSTATION 1 | 1
BIPAP PRO, DOM | 1
BIPAP S/T C SERIES, INTL 30 | 1
DREAM STATION AUTO-CPAP MACHINE | 1
REMSTAR PRO CPAP | 1
PHILIPS RESPIRONICS DREAM STATION BIPAP PRO | 1
DREAMSTATIONS | 1
DREAMSTATION AUTO BIPAP W/HUMID | 1
SYSTEM ONE PHILIPS RESPIRONICS | 1
DREAM STATION CPAP MACHINE | 1
DREAMSTATION BIPAP MACHINE | 1
AUTO CPAP | 1
RESPERONICS DREAM STATION | 1
SLEEP APNEA MACHINE | 1
DREAM STATION AUTO PHILIPS RESPIRONICS CPAP | 1
DREAMSTATION BIPAP AVAPS30 H/HT/C, DS | 1
PHILIPS CPAP DREAM STATION | 1
BIPAP PRO BIFLEX W/HUM SYS ONE 60SRS,DOM | 1
DREAMSTATION AUTO BIPAP W/HUM/CELL | 1
DREAMSTATION AUTO CPAP W/HUM/HT/CELL | 1
DREAMSTATION CPAP PRO W/HUM/CELL | 1
DREAMSTATION AUTO CPAP HUMHT CELL | 1
PHILIPS DREAM STATION 1 RESPIRONICS | 1
PHILIPS RESPIRONICS REMSTAR PLUS C-FLEX | 1
DREAMWEAR | 1
DREAMSTATION AUTO BIPAP/HUM/HT/WIFI, DOM | 1
DREAMSTATION AUTO CPAP W/ HUM/CELL, DOM | 1
PHILIPS DREAM STATION CPAP MACHINE. | 1
PHILLIPS RESPIRONICS REMSTAR AUTO DOM | 1
PHILLIPS DREAMSTATION BIPAP MACHINE | 1
PHILIPS RESPIRONICS REMSTAR PRO | 1
DREAMSTATION ST30, GB | 1
SERIES 1 BIPAP MACHINE | 1
DREAMSTATION AUTO BIPAP, INTL | 1
DREAMSTATION AUTO BIPAP W/HUMID/HT | 1
DREAMSTATION AUTO CPAP W/HUM/HT/WIFI | 1
DREAMSTATION AUTO CPAP | 1
PHILIPS RESPIRONICS BIPAP DREAMSTATION | 1
DREAMSTATION AUTO CPAP, BL | 1
DREAMSTATION CPAP, AUTO CPAP, BIPAP | 1
DREAM MACHINE CPAP | 1
DREAM STATION AUTO CPAP WITH HUM | 1
CPAP DREAMMACHINE | 1
BIPAP ST, C SERIES, INTERNATIONAL | 1
PHILIPS DREAMSTATION AUTO CPAP HUMHT DOM | 1
BIPAP | 1
DREAMSTATION GO AUTO W/BT, BRAZIL | 1
DREAMSTATION BIPAP AUTOSV H, DS | 1
CPAP DREAMSTATION AUTO CPAP | 1
RESPIRONICS SYSTEM ONE CPAP MACHINE | 1
PHILIPS BIPAP SYSTEM ONE | 1
PHILIPS RESPIRONICS DREAM STATION CPAP DEVICE | 1
DREAM STATION C-PAP | 1
PHILIPS BIPAP DEVICE | 1
DREAMSTATION EXPERT, FR | 1
CPAP DEVICE DREAM STATION | 1
PHILIPS RESPIRONICS CPAP RNTCPAP5 | 1
PHILLIPS DREAMMACHINE CPAP | 1
PHILIPS DREAMSTATION BI-PAP | 1
DREAMSTATION CPAP SLEEP THERAPY SYSTEM WITH HUMIDIFIER | 1
DREAMSTATION CPAP/BIPAP | 1
DREAM STATION CI-PAP/ BI-PAP | 1
UNK TRILOGY LINER | 1
UNKNOWN TRILOGY SCREW | 1
UNK TRILOGY SHELL | 1
PHILIPS RESPIRONICS CPAP DREAMSTATION | 1
DREAMSTATION ONE | 1
PHILIPS DREAMMAKER | 1
DREAM STATION AUTOCPAPHUMHT | 1
PHILIPS RESPIRONICS DREAMSTATION ASV | 1
DREAMSTATION BIPAP ST30 H/HT/C, DS | 1
DREAMSTATION GO AUTO W/BT, SWITZ | 1
PHILIPS RESPERONICS DREAMSTATION AUTO BIPAP WITH HUM CEL | 1
BIPAP PRO BI-FLEX | 1
BI PAP MACHINE | 1
PHILLIPS DREAM MACHINE, SLEEP APNEA MACHINE | 1
DREAMSTATION AUTOCPAP | 1
PHILLIPS DREAMSTATION CPAP DEVICE | 1
PHILIPS RESPIRONICS BI-PAP DREAMSTATION | 1
PHILLIPS RESPIRONIC DREAMSTATION | 1
DREAMSTATION CPAP PRO W/HUM/WIFI, DOM | 1
PHILIPS DREAMSTARION C-PAP | 1
PHILIPS RESPIRONICS DREAM STATION CPAP MACHINE | 1
CPAP PHILLIPS RESPIRONICS SYSTEM ONE HTD HUMID | 1
DREAMSTATION BIPAP AVAPS30 AAM H/HT, CAN | 1
DREAMSTATION AUTO BIPAP HUMIDIFIER | 1
PHILLIPS DREAMSTATION RESPIRONICS | 1
DREAMSTATION (2 CPAP DEVICES | 1
BIPAP S/T C SERIES, INTERNATIONAL | 1
PHILIPS BI-LEVEL CPAP DEVICE | 1
RENTAL DREAMSTATION EXPERT, FR | 1
PHILLIPS RESTMORE CPAP | 1
BIPAP AUTOSV ASSY 60 SRS 30CM,DOM | 1
DREAMSTATIONAUTOCPAP | 1
BIPAP A30-S SILVER SERIES | 1
PHILIPS DREAMSTATION ONE | 1
DREAMSTATION CPAP, DOM | 1

## Case Metadata
Case Name | MDL Number | MDL Filing Date | Court | MAUDE Start Date | MAUDE End Date | Total Records | Date Range
Philips CPAP | 3014 | 2021-10-08 | W.D. Pennsylvania | 2021-04-08 | 2021-10-08 | 4232 | 2021-04-08 00:00:00 to 2021-10-08 00:00:00