## All Reports
report_number | date_received | date_of_event | event_type | report_source_code | brand_name | generic_name | manufacturer_d_name | model_number | catalog_number | device_class | event_description | manufacturer_narrative | adverse_event_flag | product_problem_flag | device_operator | manufacturer_g1_name | year_month
2518422-2021-00997 | 2021-04-08 | 2021-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01001 | 2021-04-08 | 2021-03-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY, SENSOR BOARD, SYSTEM BOARD, POWER MANAGEMENT BOARD, INTERFACE BOARD AND FRONT PANEL/LCD BOARD WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-00987 | 2021-04-08 | 2021-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, CHECK CIRCUIT ALARMS OCCURRED. THE DEVICE'S OVERHEAD BLOWER FILTER WAS REPLACED DUE TO DIRT AND DUST CONTAMINATION. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-00988 | 2021-04-08 | 2021-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-00995 | 2021-04-08 | 2021-04-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S POWER MANAGEMENT BOARD AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-00996 | 2021-04-08 | 2021-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-00999 | 2021-04-08 | 2021-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S BLOWER MOTOR NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01000 | 2021-04-08 | 2021-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S BLOWER MOTOR WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-00998 | 2021-04-08 | 2021-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01025 | 2021-04-09 | 2021-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01024 | 2021-04-09 | 2021-03-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED STEPS DURING TESTING. THE DEVICE'S SENSOR BOARD AND OXYGEN BLENDING MODULE BOARD WERE REPLACED TO ADDRESS THE ISSUES. | IT WAS INITIALLY REPORTED, THE DEVICE'S SENSOR BOARD AND OXYGEN BLENDING MODULE BOARD WERE REPLACED TO ADDRESS THE ISSUES. THE DEVICE'S SYSTEM BOARD WAS ALSO REPLACED. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01026 | 2021-04-09 | 2021-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2200X11B | DS2200X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD, PROPORTIONAL VALVE, SOLENOID VALVE, BLOWER CHASSIS TUBING AND SYSTEM BOARD TUBING WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-00975 | 2021-04-09 | 2021-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-00976 | 2021-04-09 | 2021-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG AND THE DEVICE FAILED STEPS DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND SENSOR BOARD WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01016 | 2021-04-09 | 2021-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040000 | 1040000 |  | A VENTILATOR WAS EVALUATED BY A FIELD SERVICE ENGINEER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE BY THE FIELD SERVICE ENGINEER, THE DEVICE FAILED STEPS DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-00977 | 2021-04-09 | 2021-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01033 | 2021-04-09 | 2021-03-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | DS2200X11B | DS2200X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01015 | 2021-04-09 | 2021-03-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A FAILURE OF A VENTILATOR'S LCD SCREEN WAS OBSERVED. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLANK DURING USE AND WAS NOT VIEWABLE. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S LCD BACKLIGHT CABLE WAS REPLACED TO ADDRESS THE ISSUE. THERE WAS EVIDENCE OF TAMPERING. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01023 | 2021-04-09 | 2021-03-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR A VENTILATOR INOPERATIVE CONDITION. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01008 | 2021-04-09 | 2021-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A FAILURE OF THE VENTILATOR'S EXTERNAL FLOW SENSOR. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN EVALUATED BY THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A FAILURE OF A VENTILATOR'S EXTERNAL FLOW SENSOR. THE DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S FLOW SENSOR CABLE ASSEMBLY WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01011 | 2021-04-09 | 2021-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR A HIGH INTERNAL OXYGEN ALARM CONDITION. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01006 | 2021-04-09 | 2021-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2110X15B | IN2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR A LOW OXYGEN ALARM CONDITION. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS UNABLE TO BE CONFIRMED DUE TO THE DEVICE FAILING TO OPERATE. A VENTILATOR INOPERATIVE ALARM CONDITION INDICATING A FAULTY SYSTEM BOARD AND DISPLAY BOARD WAS OBSERVED IN THE DEVICE'S DOWNLOADED ERROR LOG. THE DEVICE'S MACHINE FLOW SENSOR WAS FOUND TO BE CONTAMINATED WITH DUST AND DIRT CAUSING A VENTILATOR INOPERATIVE ALARM CONDITION. NO REPAIRS WERE DONE TO THE DEVICE, THE DEVICE WAS SCRAPPED. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01004 | 2021-04-09 | 2021-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A FAILURE OF THE VENTILATOR'S EXTERNAL FLOW SENSOR. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN EVALUATED BY THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR'S EXTERNAL FLOW SENSOR MALFUNCTIONED. THE DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S EXTERNAL FLOW SENSOR CABLE WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01007 | 2021-04-09 | 2021-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S TIDAL VOLUME DELIVERY WAS LOW. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN EVALUATED BY THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR'S TIDAL VOLUME DELIVERY WAS LOW. THE DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER AND THE CUSTOMER'S COMPLAINT WAS NOT CONFIRMED. THE DEVICE'S SETTINGS WERE CHANGED TO ADDRESS THE PATIENTS NEEDS. NO MALFUNCTION OF THE DEVICE WAS NOTED. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01009 | 2021-04-09 | 2021-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A FAILURE OF THE VENTILATOR'S EXTERNAL FLOW SENSOR. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN EVALUATED BY THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A FAILURE OF A VENTILATOR'S EXTERNAL FLOW SENSOR. THE DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S EXTERNAL FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01010 | 2021-04-09 | 2021-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A FAILURE OF THE VENTILATOR'S OXYGEN SENSOR. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN EVALUATED BY THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR'S OXYGEN SENSOR MALFUNCTIONED. THE DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER AND THE CUSTOMER'S COMPLAINT CONFIRMED. THE DEVICE'S FIO2 SENSOR WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01003 | 2021-04-09 | 2021-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A FAILURE OF THE VENTILATOR'S EXTERNAL FLOW SENSOR. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN EVALUATED BY THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR'S EXTERNAL FLOW SENSOR MALFUNCTIONED. THE DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S EXTERNAL FLOW SENSOR CABLE WAS RESEATED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01005 | 2021-04-09 | 2021-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A FAILURE OF THE VENTILATOR'S OXYGEN SENSOR. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN EVALUATED BY THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR'S EXTERNAL FLOW SENSOR MALFUNCTIONED. THE DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S EXTERNAL FLOW SENSOR WAS FOUND TO BE BLOCKED WITH CONTAMINATION AND WAS CLEANED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01002 | 2021-04-09 | 2021-03-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S EXTERNAL FLOW SENSOR FAILED TO CONNECT. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN EVALUATED BY THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR'S EXTERNAL FLOW SENSOR FAILED TO CONNECT. THE DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S EXTERNAL FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01014 | 2021-04-09 | 2021-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01029 | 2021-04-12 | 2021-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1032800 | 1032800 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR HAD A VALVE LEAK. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01027 | 2021-04-12 | 2021-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE TO CHARGE ITS INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01028 | 2021-04-12 | 2021-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2200X11B | DS2200X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S PROPORTIONAL VALVE, SOLENOID VALVE AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01050 | 2021-04-12 | 2021-03-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WOULD NOT POWER ON. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S POWER SUPPLY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01049 | 2021-04-12 | 2021-03-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BLOWER MOTOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01048 | 2021-04-12 | 2021-04-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WOULD NOT POWER ON. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S POWER SUPPLY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01045 | 2021-04-12 | 2021-04-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01013 | 2021-04-12 | 2021-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE VENTILATOR'S SOFTWARE WAS RE-LOADED TO ADDRESS THE ISSUE. DURING THE EVALUATION ERROR CODES RELATED TO THE POWER SUPPLY AND SENSOR BOARD WERE OBSERVED. THE DEVICE'S SENSOR BOARD AND POWER SUPPLY WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01012 | 2021-04-12 | 2021-04-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD, OXYGEN BLENDING MANIFOLD AND OXYGEN BLENDING MODULE HARNESS WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01044 | 2021-04-12 | 2021-04-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01047 | 2021-04-13 | 2021-03-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, JP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | RJPX500S16 | RJPX500S16 |  | THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF SPARKING AND DAMAGE OCCURRING TO AN AC POWER CORD OF A CONTINUOUS POSITIVE AIRWAY DEVICE (CPAP) DEVICE. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER RECEIVED PHOTOGRAPHS ONLY FROM THE DISTRIBUTOR FOR INVESTIGATION. THE MANUFACTURER DOES CONFIRM THERE IS EVIDENCE OF PHYSICAL DAMAGE TO THE AC CORD. DURING PHOTOGRAPHIC EVALUATION OF THE AC POWER CORD, THE MANUFACTURER CONFIRMED EVIDENCE OF EXPOSED AC WIRES AND SEPARATION AT THE AC PLUG. THERE WAS NO EVIDENCE OF THERMAL DAMAGE. THE DREAMSTATION USER MANUAL STATES: WARNING: PERIODICALLY INSPECT ELECTRICAL CORDS AND CABLES FOR DAMAGE OR SIGN OF WEAR. DISCONTINUE USE AND REPLACE IF DAMAGED. THIS DEVICE MEETS UL AND IEC REQUIREMENTS FOR FLAMMABILITY. THE MANUFACTURER CONCLUDES THAT NO FURTHER ACTION IS NECESSARY AT THIS TIME. |  | N | Y | LAY USER/PATIENT |  | 2021-04
11655468 | 2021-04-13 | 2021-02-17 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | PHILIPS V60 BIPAP | VENTILATOR, CONTINUOUS, MINIMAL VENTILATORY SUPPORT, FACILITY USE | RESPIRONICS CALIFORNIA, INC. | V60 |  |  | RECEIVED CALL FROM MONITOR TECH THAT PT IN THE 40% SPO2. WHEN I ARRIVED PT WAS BARELY RESPONSIVE AND SAW THE BIPAP/VENT MACHINE SCREEN, WHICH READ ERROR WITH BLACK SCREEN. THE BIPAP FACE MASK WAS STILL APPLIED AND NOT REMOVED. I HAD VOCERA BUT WAS NOT ABLE TO MAKE ANY CALL DUE TO MISCONNECTION, SO I PRESSED CODE BLUE BUTTON FOR FURTHER ASSISTANCE. COMPRESSIONS WAS STARTED DUE TO UNRESPONSIVE STATE. INITIAL HEART RHYTHM WAS SINUS BRADYCARDIA PULSELESS ELECTRICAL ACTIVITY (SB-PEA), CARDIOPULMONARY RESUSCITATION (CPR) AND BAG VALVE MASK VENTILATION (BVM) INITIATED FOR 1 MINUTE AFTER LEARNING THE PATIENT WAS LIMITED CODE BIPAP ONLY. PATIENT BOUNCE BACK AND HEART RATE IS AT 130 AND OXYGEN SATS IS 94% ON BVM. PATIENT BACK ON THE BIPAP BUT REMAIN UNRESPONSIVE. SITTER AT BEDSIDE. |  | N | Y | I |  | 2021-04
2518422-2021-01063 | 2021-04-13 | 2021-04-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S POWER MANAGEMENT BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01062 | 2021-04-13 | 2021-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01061 | 2021-04-13 | 2021-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THIRD PARTY SERVICE CENTER, A "SERVICE REQUIRED" CODE RELATED TO THE INTERNAL BATTERY WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01060 | 2021-04-13 | 2021-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01046 | 2021-04-13 | 2021-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S DISPLAY WAS PARTIALLY BLANK AND UNABLE TO BE VIEWED. THE DEVICE WAS NOT IN PATIENT USE. THE MANUFACTURER RECEIVED THE DEVICE FOR EVALUATION, AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. THERE WAS EVIDENCE OF PHYSICAL DAMAGE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01058 | 2021-04-13 | 2021-03-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A FAILURE OF THE VENTILATOR'S EXTERNAL FLOW SENSOR ADAPTER. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S FLOW SENSOR CABLE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01031 | 2021-04-13 | 2021-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01030 | 2021-04-13 | 2021-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S MACHINE FLOW SENSOR WAS REPLACED DUE TO DUST AND DIRT CONTAMINATION. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01053 | 2021-04-13 | 2021-04-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1032800 | 1032800 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01052 | 2021-04-13 | 2021-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01056 | 2021-04-13 | 2021-04-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S SENSOR BOARD AND ACTIVE EXHALATION CONTROL MODULE WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01055 | 2021-04-13 | 2021-04-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040005 | 1040005 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. THE DEVICE FAILED A TEST STEP DURING TESTING AND AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01057 | 2021-04-13 | 2021-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2110X15B | IN2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S DISPLAY WAS DAMAGED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S DISPLAY WAS REPLACED DUE TO PHYSICAL DAMAGE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01059 | 2021-04-13 | 2021-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1032804 | 1032804 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICES LCD DISPLAY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01073 | 2021-04-14 | 2021-03-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TO CALIBRATE. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGINGATION THAT A VENTILATOR FAILED TO CALIBRATE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT COULD NOT BE CONFIRMED OR DUPLICATED. THE DEVICE PERFORMED TO MANUFACTURER'S SPECIFICATIONS. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01040 | 2021-04-14 | 2021-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD AND POWER MANAGEMENT BOARD NEED TO BE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01037 | 2021-04-14 | 2021-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040005 | 1040005 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01034 | 2021-04-14 | 2021-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILS TO CHARGE ITS INTERNAL BATTERY. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01035 | 2021-04-14 | 2021-03-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE DEVICE FAILED TO CALIBRATE. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01083 | 2021-04-14 | 2021-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | SP2100X26B | SP2100X26B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS NOT REACHING THE SET TIDAL VOLUME. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | IT WAS INITIALLY REPORTED, THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS NOT REACHING THE SET TIDAL VOLUME. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BLOWER MOTOR WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01082 | 2021-04-14 | 2021-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S BREATH RATE DELIVERY WAS INCREASED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S FLOW SENSOR ASSEMBLY WAS FOUND TO BE DAMAGED AND NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01065 | 2021-04-14 | 2021-04-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01038 | 2021-04-14 | 2021-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S DETACHABLE BATTERY CABLE WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01041 | 2021-04-14 | 2021-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S POWER MANAGEMENT BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01076 | 2021-04-14 | 2021-04-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR A "HIGH PRIORITY" ALARM CONDITION. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD NEED TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01039 | 2021-04-14 | 2021-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01067 | 2021-04-14 | 2021-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR HAD A "BOARD ISSUE". THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. SERVICE REQUIRED CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG AND THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S POWER MANAGEMENT BOARD AND OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01068 | 2021-04-14 | 2021-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01036 | 2021-04-14 | 2021-04-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S BLOWER MOTOR WAS REPLACED TO ADDRESS THE ISSUE. THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01069 | 2021-04-14 | 2021-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE TO POWER ON WAS OBSERVED. THE DEVICE'S SYSTEM BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. THE DEVICE HAD EVIDENCE OF WATER INGRESS. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01064 | 2021-04-14 | 2021-04-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01054 | 2021-04-14 | 2021-04-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, THE VENTILATOR FAILED TO DELIVER FLOW. THE DEVICE'S EXTERNAL FLOW SENSOR WAS ADJUSTED AND THE DEVICE WAS REBOOTED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR'S EXTERNAL FLOW SENSOR WAS ADJUSTED AND THE DEVICE WAS REBOOTED TO ADDRESS A FAILURE TO DELIVER FLOW ISSUE. ADDITIONAL INFORMATION RECEIVED FROM THE THIRD PARTY SERVICE CENTER STATES THE SYSTEM BOARD AND OXYGEN BLENDING MODULE BOARD WERE REPLACE TO ADDRESS THIS ISSUE ALSO. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01077 | 2021-04-14 | 2021-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01081 | 2021-04-14 | 2021-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S DISPLAY WAS BLANK. THERE WAS NO HARM OR INJURY REPORTED. THE MANUFACTURER IS CURRENTLY INVESTIGATING THIS ISSUE AND A FOLLOW-UP REPORT WILL BE FILED WHEN THE INVESTIGATION IS COMPLETE. | IT WAS INITIALLY REPORTED, THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S DISPLAY WAS BLANK.THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER THE CUSTOMER'S COMPLAINT COULD NOT BE CONFIRMED OR DUPLICATED. THE DISPLAY OPERATED AS DESIGNED. THE DEVICE'S AC CONNECTOR WAS FOUND TO BE PHYSICALLY DAMAGED. THE AC CONNECTOR WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01075 | 2021-04-14 | 2021-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2200X15B | IN2200X15B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD, OXYGEN BLENDING MODULE ASSEMBLY AND WIRING HARNESS WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01072 | 2021-04-14 | 2021-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040004 | 1040004 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT'S BLOOD OXYGEN SATURATION DECREASED WHILE ON A VENTILATOR. THE PATIENT WAS SWITCHED TO ANOTHER VENTILATOR. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER BUT THE INVESTIGATION IS STILL ONGOING. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON 04 NOV, 2024 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A PATIENT'S BLOOD OXYGEN SATURATION DECREASED WHILE ON A VANTILATOR. THE PATIENT WAS SWITCHED TO ANOTHER VENTILATOR. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECTS: HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01078 | 2021-04-14 | 2021-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING AND A "SERVICE REQUIRED" ALARM CONDITION OCCURRED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01084 | 2021-04-15 | 2021-03-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040000 | 1040000 |  | THE MANUFACTURER RECEIVED A REQUEST FOR SERVICE OF A VENTILATOR. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE BY THE MANUFACTURER'S FIELD SERVICE ENGINEER, AN ISSUE RELATED TO THE OXYGEN BLENDING MODULE BOARD WAS OBSERVED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01020 | 2021-04-15 | 2021-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01042 | 2021-04-15 | 2021-04-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01019 | 2021-04-15 | 2021-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01022 | 2021-04-15 | 2021-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED STEPS DURING TESTING. THE DEVICE'S SENSOR BOARD NEEDS REPLACED TO ADDRESS THE ISSUES. THE DEVICE WAS RETURNED UNREPAIRED TO THE CUSTOMER PER THE CUSTOMER'S REQUEST. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01021 | 2021-04-15 | 2021-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED STEPS DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01043 | 2021-04-15 | 2021-04-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01074 | 2021-04-15 | 2021-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | SP2100X26B | SP2100X26B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOSULY REPORTED AN ALLEGATION OF A VENTILATOR INOPERATIVE CONDITION OCCURRING. THE DEVICE WAS RETURNED ON A DUPLICATE COMPLAINT (B)(4) AND FINAL REPORTING WAS SUBMITTED UNDER MDR 2518422-2022-13606. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01085 | 2021-04-15 | 2021-03-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. THE DEVICE HAD EVIDENCE OF PHYSICAL DAMAGE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01066 | 2021-04-15 | 2021-04-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01094 | 2021-04-15 | 2021-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILS TO CHARGE ITS INTERNAL BATTERY. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01017 | 2021-04-15 | 2021-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED EVENT LOG. THE DEVICE'S SYSTEM BOARD AND SENSOR BOARD NEED REPLACED TO ADDRESS THE ISSUES. | THE MANUFACTURER PREVIOUSLY REPORTED THAT A VENTILATOR HAD "SERVICE REQUIRED" CODES THAT WERE FOUND IN THE VENTILATOR'S DOWNLOADED EVENT LOG AND THE DEVICE'S SYSTEM BOARD AND SENSOR BOARD WERE REPLACED TO ADDRESS THE ISSUES. FURTHER INFORMATION WAS RECEIVED THAT AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01018 | 2021-04-15 | 2021-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED.. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WERE FOUND IN THE VENTILATOR'S DOWNLOADED EVENT LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01099 | 2021-04-16 | 2021-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR FLOW SENSOR ISSUES. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S FLOW SENSOR CABLE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01032 | 2021-04-16 | 2021-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | I |  | 2021-04
2518422-2021-01100 | 2021-04-16 | 2021-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01118 | 2021-04-19 | 2021-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S POWER MANAGEMENT BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01105 | 2021-04-19 | 2021-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WOULD NOT POWER ON. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE DEVICE'S INTERNAL BATTERY WAS RECONNECTED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01095 | 2021-04-19 | 2021-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BLOWER MOTOR WAS FOUND TO BE CONTAMINATED AND NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01104 | 2021-04-19 | 2021-04-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2200X15B | IN2200X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01103 | 2021-04-19 | 2021-04-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01097 | 2021-04-19 | 2021-04-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01116 | 2021-04-19 | 2021-04-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S INTERFACE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01126 | 2021-04-19 | 2021-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO PATIENT HARM OR INJURY REPORTED. DURING THE EVALUATION AT THE THIRD PARTY SERVICE CENTER, "SERVICE REQUIRED" CODES RELATED TO THE INTERNAL BATTERY WERE OBSERVED IN THE DOWNLOADED EVENT LOG. THE DEVICE EVALUATION IS NOT COMPLETE AT THE THIRD PARTY SERVICE CENTER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE THIRD PARTY SERVICE CENTER INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. AND DURING THE EVALUATION AT THE THIRD PARTY SERVICE CENTER, "SERVICE REQUIRED" CODES RELATED TO THE INTERNAL BATTERY WERE OBSERVED IN THE DOWNLOADED EVENT LOG. THE DEVICE EVALUATION WAS NOT COMPLETED. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE SERVICE REQUIRED CODES RELATED TO THE INTERNAL BATTERY WERE OBSERVED IN THE DOWNLOADED EVENT LOG. THE DETACHABLE BATTERY CABLE WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01128 | 2021-04-19 | 2021-04-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01127 | 2021-04-19 | 2021-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR TIDAL VOLUME ISSUES. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE EXTERNAL FLOW SENSOR WAS CLEANED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01098 | 2021-04-19 | 2021-04-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A SERVICE REQUIRED ALARM CONDITION OCCURRED RELATED TO THE FAILURE OF THE DEVICE TO CHARGE ITS INTERNAL BATTERY. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01117 | 2021-04-19 | 2021-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01096 | 2021-04-19 | 2021-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01110 | 2021-04-20 | 2021-04-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040005 | 1040005 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE LCD SCREEN WAS OBSERVED. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01109 | 2021-04-20 | 2021-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WOULD NOT POWER ON. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01113 | 2021-04-20 | 2021-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01114 | 2021-04-20 | 2021-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, ERROR CODES RELATED TO THE CHARGING OF THE INTERNAL BATTERY WERE OBSERVED. THE DEVICE'S INTERNAL BATTERY AND POWER MANAGEMENT BOARD WERE REPLACED TO ADDRESS THE ISSUES. | THE MANUFACTURER PREVIOUSLY REPORTED THE DEVICE'S INTERNAL BATTERY AND POWER MANAGEMENT BOARD WERE REPLACED TO ADDRESS ERROR CODES RELATED TO THE CHARGING OF THE INTERNAL BATTERY. THE PARTS WERE NOT REPLACED. THE DEVICE WAS RETURNED TO THE CUSTOMER UNREPAIRED. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01115 | 2021-04-20 | 2021-04-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01112 | 2021-04-20 | 2021-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE LCD SCREEN WAS OBSERVED. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01111 | 2021-04-20 | 2021-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER HAD PREVIOUSLY REPORTED THAT THE OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS A "SERVICE REQUIRED" CODE IN THE DOWNLOADED ERROR LOG. THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01130 | 2021-04-20 | 2021-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01131 | 2021-04-20 | 2021-04-16 | D, e, a, t, h | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | FX2100X15B | FX2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT BECAME DISCONNECTED FROM A VENTILATOR AND EXPIRED. THE MANUFACTURER IS CURRENTLY INVESTIGATING THIS EVENT AND A FOLLOW UP REPORT WILL BE FILED WHEN THE INVESTIGATION IS COMPLETE. | IT WAS INITIALLY REPORTED, A PATIENT BECAME DISCONNECTED FROM THE DEVICE AND EXPIRED. THE MANUFACTURER RECEIVED THE VENTILATOR'S DOWNLOADED EVENT LOGS FROM THE REPORTER OF THE EVENT. ON THE DAY OF THE EVENT, (B)(6) 2021, A LOW OXYGEN INLET PRESSURE ALARM OCCURRED WITH A DURATION OF 2791 SECONDS, APPROXIMATELY 46 MINUTES. AN OXYGEN REGULATION ALARM ALSO OCCURRED WITH A DURATION OF 2641 SECONDS, APPROXIMATELY 44 MINUTES. THESE ALARMS WERE ACKNOWLEDGED AS EVIDENCED BY AN ALARM SILENCE KEY PRESS. THE VENTILATOR WAS BEING USED WITH 85% OXYGEN. THESE ALARMS INDICATE THE OXYGEN FLOW WAS INADEQUATE FOR THE VENTILATOR TO SUPPLY 85% OXYGEN TO THE PATIENT. THE ALARMS ARE INDICATIVE OF A LARGE LEAK, SUCH AS A PATIENT CIRCUIT DISCONNECT. OTHER ALARMS INCLUDING THE PATIENT CIRCUIT DISCONNECT AND LOW MINUTE VENTILATION WERE TURNED OFF AT THE TIME OF THE EVENT. NO EVENTS WERE LOGGED THAT WOULD INDICATE A DEVICE MALFUNCTION OR FAILURE TO DELIVER THERAPY OCCURRED. THE MANUFACTURER CONCLUDES THE DEVICE OPERATED AND ALARMED AS DESIGNED. | Y | N | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01087 | 2021-04-20 | 2021-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR A "SERVICE REQUIRED" CONDITION RELATED TO THE DEVICE'S OXYGEN BLENDING MODULE. THERE WAS NO HARM OR INJURY REPORTED. THE MANUFACTURER RECEIVED THE DEVICE FOR EVALUATION, AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01119 | 2021-04-20 | 2021-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01120 | 2021-04-20 | 2021-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01121 | 2021-04-20 | 2021-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01122 | 2021-04-20 | 2021-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S DISPLAY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01124 | 2021-04-20 | 2021-04-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01123 | 2021-04-20 | 2021-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01125 | 2021-04-20 | 2021-04-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01107 | 2021-04-20 | 2021-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERNAL BATTERY WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01106 | 2021-04-20 | 2021-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DISPLAY WAS BLANK DUE TO PHYSICAL DAMAGE. THE LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
MW5100912 | 2021-04-20 | 2021-03-31 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | PHILLIPS / RESPIRONICS INC. | TRILOGY EVO | DS2110X11B |  | PATIENT DISCONNECTED FROM THE TRILOGY EVO AND THE VENT DID NOT ALARM FOR CIRCUIT DISCONNECT. WENT TO THE PATIENT'S HOME AND SWAPPED OUT THE VENT. ACCORDING TO MOM, NO HARM TO THE PATIENT. NO HARM TO PATIENT. RETURNING THE VENT FOR FURTHER EVALUATION. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-04
2518422-2021-01129 | 2021-04-20 | 2021-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01088 | 2021-04-21 | 2021-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG AND THE DEVICE FAILED STEPS DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED, AND THE DEVICE'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01086 | 2021-04-21 | 2021-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | IN2200X15B | IN2200X15B |  | A VENTILATOR WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED EVENT LOG. THE DEVICE'S OXYGEN BLENDING MODULE SUBASSEMBLY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01161 | 2021-04-21 | 2021-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | DS2200X11B | DS2200X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR HAD LOW OXYGEN READINGS. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION, AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN BLENDING MODULE HARNESS AND OXYGEN BLENDING MODULE SUBASSEMBLY NEED REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01157 | 2021-04-21 | 2021-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT COULD NOT BE CONFIRMED. THE DEVICE WAS FOUND TO OPERATE AND ALARM AS DESIGNED. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01156 | 2021-04-21 | 2021-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01155 | 2021-04-21 | 2021-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE VENTILATOR'S SOFTWARE WAS RE-LOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01101 | 2021-04-22 | 2021-04-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | PHILIPS RESPIRONICS INC. | 1040000 | 1040000 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED "SERVICE REQUIRED" ALARM CONDITION OCCURRED AND THAT THE SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. THE SENSOR BOARD WAS NOT REPLACED FOR THE REPORTED ISSUE, ONLY THE OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01092 | 2021-04-22 | 2021-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01089 | 2021-04-22 | 2021-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED EVENT LOG. THE DEVICE'S INTERFACE BOARD AND OXYGEN BLENDING MODULE BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01093 | 2021-04-22 | 2021-04-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED A STEP RELATED TO OXYGEN DELIVERY DURING TESTING. THE DEVICE WAS NOT IN PATIENT USE. THE MANUFACTURER RECEIVED THE DEVICE FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01181 | 2021-04-22 | 2021-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01180 | 2021-04-22 | 2021-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S BLOWER MOTOR WAS RECONNECTED TO THE SYSTEM BOARD TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01169 | 2021-04-22 | 2021-04-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2200X11B | DS2200X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01170 | 2021-04-22 | 2021-04-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TESTING. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01168 | 2021-04-22 | 2021-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2110X15B | IN2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01179 | 2021-04-22 | 2021-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERFACE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01178 | 2021-04-22 | 2021-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01177 | 2021-04-22 | 2021-04-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01102 | 2021-04-22 | 2021-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | PHILIPS RESPIRONICS INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE INTERNAL BATTERY WAS DEPLETED ON A VENTILATOR. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR ALLEGEDLY HAVING A DEPLETED INTERNAL BATTERY AND THE INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. THE DEVICE WAS NOT IN PATIENT USE. THE MANUFACTURER CONFIRMS THE INTERNAL BATTERY WAS NOT REPLACED AND THE DEVICE WAS RETURNED UNREPAIRED. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01176 | 2021-04-22 | 2021-04-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
0001822565-2021-01126 | 2021-04-22 | 2021-03-29 | I, n, j, u, r, y | Manufacturer report | TRILOGY SPIKED ACET SHELL 54MM | PROSTHETIC, HIP | ZIMMER BIOMET, INC. | N/A | 00620005423 |  | IT WAS REPORTED THAT THE SURGEON SAID PATIENT WAS HAVING SOME PAIN AND XRAYS CONFIRMED THEY NEEDED A CUP REVISION. ACETABULAR CUP WAS REVISED AND PSEUDO-TUMOR WAS REMOVED. ATTEMPTS HAVE BEEN MADE AND ADDITIONAL INFORMATION ON THE REPORTED EVENT IS UNAVAILABLE AT THIS TIME. | (B)(4). CONCOMITANT MEDICAL PRODUCTS: 00801803603 FEMORAL HEAD STERILE PRODUCT DO NOT RESTERILIZE 12/14 TAPER 60181513. UNKNOWN TRILOGY LINER. THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS DUE TO HOSPITAL POLICY; HOWEVER, AN INVESTIGATION OF THE REPORTED EVENT IS IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETED, A SUPPLEMENTAL MDR WILL BE SUBMITTED. | Y | N | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01108 | 2021-04-22 | 2021-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01133 | 2021-04-22 | 2021-03-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WOULD NOT POWER ON. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR FAILED TO POWER ON. THE VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S POWER SUPPLY WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01132 | 2021-04-22 | 2021-03-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1032800 | 1032800 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED A TEST STEP DURING TESTING. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR FAILED A TEST STEP DURING TESTING. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SENSOR BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01137 | 2021-04-22 | 2021-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01134 | 2021-04-22 | 2021-03-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WOULD NOT POWER ON. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR FAILED TO POWER ON. THE VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S POWER SUPPLY NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01136 | 2021-04-22 | 2021-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01174 | 2021-04-22 | 2021-04-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S INTERFACE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01175 | 2021-04-22 | 2021-04-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01135 | 2021-04-22 | 2021-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SENSOR BOARD AND ACTIVE EXHALATION CONTROL MODULE WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01160 | 2021-04-23 | 2021-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054097 | 1054097 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SENSOR BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01188 | 2021-04-23 | 2021-04-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01187 | 2021-04-23 | 2021-04-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040007 | 1040007 |  | VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01171 | 2021-04-23 | 2021-04-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. THE DEVICE HAD EVIDENCE OF WATER INGRESS. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01172 | 2021-04-23 | 2021-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR HAD A CIRCUIT LEAK ERROR. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01173 | 2021-04-23 | 2021-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BLOWER MOTOR WAS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01184 | 2021-04-23 | 2021-04-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND SENSOR BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01182 | 2021-04-23 | 2021-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A FAILURE VENTILATOR'S EXTERNAL FLOW SENSOR. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S EXTERNAL FLOW SENSOR CABLE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01185 | 2021-04-23 | 2021-04-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S BLOWER MOTOR NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01140 | 2021-04-23 | 2021-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S POWER MANAGEMENT BOARD AND DETACHABLE BATTERY CABLE WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01139 | 2021-04-23 | 2021-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, THE VENTILATOR FAILED TO POWER ON. THE DEVICE'S POWER SUPPLY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01141 | 2021-04-23 | 2021-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01138 | 2021-04-23 | 2021-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01091 | 2021-04-23 | 2021-04-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040000 | 1040000 |  | A VENTILATOR WAS EVALUATED BY A FIELD SERVICE ENGINEER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED EVENT LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01183 | 2021-04-23 | 2021-04-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1032804 | 1032804 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01210 | 2021-04-26 | 2021-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | DS2200X11B | DS2200X11B |  | A VENTILATOR WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE REBOOTING WAS OBSERVED. THE DEVICE'S DISPLAY BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01212 | 2021-04-26 | 2021-03-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S PRESSURE DIDN'T RISE. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR'S PRESSURE DIDN'T RISE. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS RETURNED TO THE MANUFACTURE FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OVERHEAD BLOWER FILTER WAS REPLACED TO ADDRESS THE ISSUE. THERE WAS EVIDENCE OF CONTAMINATION. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01193 | 2021-04-26 | 2021-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01192 | 2021-04-26 | 2021-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01211 | 2021-04-26 | 2021-04-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLANK DURING USE AND WAS NOT VIEWABLE. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01202 | 2021-04-26 | 2021-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, A FAILURE OF THE VENTILATOR TO DELIVER VOLUME WAS OBSERVED. THE DEVICE'S FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01191 | 2021-04-26 | 2021-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND GASKET WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01167 | 2021-04-26 | 2021-02-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND OXYGEN BLENDING MODULE MANIFOLD WERE FOUND TO BE CONTAMINATED AND WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01158 | 2021-04-26 | 2021-01-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR PREVENTIVE MAINTENANCE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01189 | 2021-04-26 | 2021-04-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01190 | 2021-04-26 | 2021-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01159 | 2021-04-26 | 2021-01-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR PREVENTIVE MAINTENANCE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01162 | 2021-04-26 | 2021-04-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR A "SERVICE REQUIRED" CONDITION RELATED TO THE DEVICE'S INTERNAL BATTERY. THERE WAS NO HARM OR INJURY REPORTED. THE MANUFACTURER RECEIVED THE DEVICE FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE ALSO FAILED A STEP DURING TESTING. THE DEVICE'S INTERNAL BATTERY AND POWER MANAGEMENT BOARD NEED REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01164 | 2021-04-27 | 2021-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01163 | 2021-04-27 | 2021-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040004 | 1040004 |  | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLANK DURING USE AND WAS NOT VIEWABLE. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01166 | 2021-04-27 | 2021-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01216 | 2021-04-27 | 2021-04-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR A "SERVICE REQUIRED" CONDITION RELATED TO LOW INSPIRATORY PRESSURE. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION, AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OVERHEAD BLOWER FILTER AND ACTIVE EXHALATION CONTROL MODULE NEED REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01218 | 2021-04-27 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S OVERHEAD BLOWER FILTER WAS OBSERVED. THE DEVICE'S OVERHEAD BLOWER FILTER NEEDS REPLACED TO ADDRESS THE ISSUE. THERE WAS EVIDENCE OF CONTAMINATION. | THE MANUFACTURE PREVIOUSLY REPORTED AN ISSUE RELATED TO A VENTILATOR'S OVERHEAD BLOWER FILTER WAS OBSERVED AND THE OVERHEAD BLOWER FILTER NEEDED REPLACED TO ADDRESS THE ISSUE. DURING ADDITIONAL TESTING, THE DEVICE FAILED A TEST STEP. THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01208 | 2021-04-27 | 2021-04-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE TO CHARGE IT'S INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01219 | 2021-04-27 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01222 | 2021-04-27 | 2021-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BLOWER MOTOR NEEDS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION, A "SERVICE REQUIRED" CODE WAS OBSERVED IN THE DOWNLOADED EVENT LOG. THE DEVICES OXYGEN BLENDING MODULE BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01220 | 2021-04-27 | 2021-04-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLANK. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS EVALUATED AT THE MANUFACTURER'S SERVICE CENTER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01165 | 2021-04-27 | 2021-04-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054655T | 1054655T |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01204 | 2021-04-27 | 2021-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR MALFUNCTIONED WHILE IN PATIENT USE. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE VENTILATOR'S PRESSURE DELIVERY FLUCTUATED. THE DEVICE'S BLOWER MOTOR WAS FOUND TO BE CONTAMINATED AND WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01221 | 2021-04-27 | 2021-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1032804 | 1032804 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, A VENTILATOR'S PRESSURE FAILED TO INCREASE. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01214 | 2021-04-27 | 2021-04-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01215 | 2021-04-27 | 2021-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01213 | 2021-04-27 | 2021-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S AIRFLOW WAS NOT DELIVERING PROPERLY. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR'S AIRFLOW WAS NOT DELIVERING PROPERLY. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT COULD NOT BE DUPLICATED. THE DEVICE WAS FOUND TO OPERATE AND ALARM AS DESIGNED. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01198 | 2021-04-27 | 2021-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01199 | 2021-04-27 | 2021-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S BLOWER MOTOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01200 | 2021-04-27 | 2021-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01201 | 2021-04-27 | 2021-04-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG AND THE DEVICE FAILED A TEST STEP DUIRNG TESTING. THE DEVICE'S INTERFACE BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01205 | 2021-04-27 | 2021-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, A FAILURE OF THE DEVICE TO CHARGE IT'S INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01209 | 2021-04-27 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01207 | 2021-04-27 | 2021-04-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S MACHINE FLOW SENSOR WAS FOUND TO BE CONTAMINATED AND THE SYSTEM BOARD WAS FOUND TO BE FAULTY AND WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01206 | 2021-04-27 | 2021-04-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01203 | 2021-04-27 | 2021-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TO CHARGE ITS INTERNAL BATTERY. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S A/C POWER CORD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01223 | 2021-04-28 | 2021-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR "SERVICE REQUIRED" RELATED TO OXYGEN REGULATION. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01239 | 2021-04-28 | 2021-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED A TEST STEP. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR FAILED A TEST STEP. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE WAS EVALUATED BY A THIRD PARTY FIELD SERVICE ENGINEER AND THE CUSTOMER'S COMPLAINT WAS NOT DUPLICATED. THE TEST FAILURE WAS DUE TO USER ERROR AND NO MALFUNCTION OF THE DEVICE WAS OBSERVED. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01240 | 2021-04-28 | 2021-04-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A FAILURE OF A VENTILATOR'S TOUCHSCREEN. THERE WAS NO PATIENT HARM OR INJURY. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A FAILURE OF A VENTILATOR'S TOUCHSCREEN. THERE WAS NO PATIENT HARM OR INJURY REPORTED. THE DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT COULD NOT BE DUPLICATED. THE DEVICE OPERATED AND ALARMED AS DESIGNED. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01231 | 2021-04-28 | 2021-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, ISSUES RELATED TO THE SENSOR BOARD AND KEYPAD WERE OBSERVED THE DEVICE'S SENSOR BOARD AND KEYPAD WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01234 | 2021-04-28 | 2021-04-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR WAS ALARMING FOR A "VENTILATOR INOPERATIVE" ALARM CONDIITON. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD AND MACHINE FLOW SENSOR NEED TO BE REPLACED TO ADDRESS THE ISSUE. THE DEVICE WAS FOUND TO HAVE TOBACCO CONTAMINATION. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01238 | 2021-04-28 | 2021-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01236 | 2021-04-28 | 2021-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01229 | 2021-04-28 | 2021-04-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | DS2200X11B | DS2200X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATORS TOUCHSCREEN FAILED TO RESPOND. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR'S TOUCHSCREEN FAILED TO RESPOND. THERE WAS NO HARM OR INJURY REPORTED. THE MANUFACTURER'S FIELD SERVICE ENGINEER EVALUATED THE DEVICE AND WAS UNABLE TO CONFIRM THE CUSTOMER'S COMPLAINT. THE DEVICE WAS FOUND TO OPERATE AND ALARM AS DESIGNED. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01230 | 2021-04-28 | 2021-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS FAILING A STEP DURING TESTING. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR FAILED A STEP DURING TESTING. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION, AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01144 | 2021-04-28 | 2021-04-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE DISPLAY WAS OBSERVED TO BE BLANK AND UNREADABLE. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01149 | 2021-04-28 | 2021-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01142 | 2021-04-28 | 2021-04-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01143 | 2021-04-28 | 2021-04-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE DEVICE FAILED CALIBRATION. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01154 | 2021-04-28 | 2021-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01145 | 2021-04-28 | 2021-04-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A FAILURE OF THE DEVICE TO CHARGE ITS INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S POWER MANAGEMENT BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01152 | 2021-04-28 | 2021-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01150 | 2021-04-28 | 2021-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S POWER MANAGEMENT BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01153 | 2021-04-28 | 2021-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01146 | 2021-04-28 | 2021-04-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040005 | 1040005 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01148 | 2021-04-28 | 2021-04-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE TO POWER ON WAS OBSERVED. THE DEVICE'S AC INLET CONNECTOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01228 | 2021-04-28 | 2021-04-21 | D, e, a, t, h | Manufacturer report | DREAMSTATION AUTO CPAP | DREAMSTATION AUTO CPAP | RESPIRONICS, INC. | RJPX500S16 | RJPX500S16 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY WHILE USING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE. NO FURTHER INFORMATION IS AVAILABLE REGARDING THE EVENT AT THIS TIME. THE MANUFACTURER REQUESTED THE RETURN OF THE DEVICE FOR EVALUATION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED THE DEVICE FOR EVALUATION. THE INVESTIGATION BY THE MANUFACTURER FOUND THE DEVICE PASSED ALL TESTING AND WAS FOUND TO OPERATE AS DESIGNED. ADDITIONAL INFORMATION RECEIVED FROM CUSTOMER INDICATING THE DEVICE DID NOT CAUSE NOR CONTRIBUTE TO THE NOTED USER DEATH. THE PHILIPS RESPIRONICS DEVASTATION SYSTEMS DELIVER POSITIVE AIRWAY PRESSURE THERAPY FOR THE TREATMENT OF OBSTRUCTIVE SLEEP APNEA IN SPONTANEOUSLY BREATHING PATIENTS WEIGHING OVER 30 KG (66 LBS). IT IS FOR USE IN THE HOME OR HOSPITAL/INSTITUTIONAL ENVIRONMENT. | Y | Y | LAY USER/PATIENT |  | 2021-04
2518422-2021-01147 | 2021-04-28 | 2021-04-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01151 | 2021-04-28 | 2021-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S POWER MANAGEMENT BOARD AND INTERNAL BATTERY NEED TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01232 | 2021-04-28 | 2021-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040000 | 1040000 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS NOT ALLOWING AN FIO2 SETTING. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER HAD PREVIOUSLY REPORTED AN ALLEGATION OF A VENTILATOR WAS NOT ALLOWING AN FIO2 SETTING TO BE SET. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01226 | 2021-04-29 | 2021-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01250 | 2021-04-29 | 2021-04-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01251 | 2021-04-29 | 2021-04-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01227 | 2021-04-29 | 2021-04-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD DISPLAY WAS FOUND TO BE BLANK. THE RIBBON CABLE CONNECTOR ON THE DISPLAY WAS RESEATED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01247 | 2021-04-29 | 2021-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01225 | 2021-04-29 | 2021-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01224 | 2021-04-29 | 2021-04-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01246 | 2021-04-29 | 2021-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01235 | 2021-04-29 | 2021-04-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TO FUNCTION. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR ALLEGEDLY FAILED TO FUNCTION. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01233 | 2021-04-29 | 2021-04-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2200X11B | DS2200X11B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, THE LCD WAS FOUND TO BE PHYSICALLY DAMAGED. THE DEVICE'S DISPLAY NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01237 | 2021-04-29 | 2021-04-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, A FAILURE OF THE EXHALATION PORTING BLOCK WAS OBSERVED. THE DEVICE'S EXHALATION PORTING BLOCK WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01243 | 2021-04-29 | 2021-04-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. REPEATED ATTEMPTS FOR ADDITIONAL INFORMATION OR TO HAVE THE DEVICE RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01249 | 2021-04-29 | 2021-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, AN ISSUE RELATED TO THE OXYGEN BLENDING MODULE BOARD WAS OBSERVED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01244 | 2021-04-29 | 2021-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED EVENT LOG. THE DEVICE'S INTERNAL BATTERY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01248 | 2021-04-29 | 2021-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S DISPLAY WAS OBSERVED. THE DEVICE'S DISPLAY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01256 | 2021-04-30 | 2021-04-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-04
2518422-2021-01253 | 2021-05-03 | 2021-04-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S POWER MANAGEMENT BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01252 | 2021-05-03 | 2021-04-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BLOWER MOTOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01254 | 2021-05-03 | 2021-04-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S MACHINE FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01255 | 2021-05-03 | 2021-04-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01267 | 2021-05-03 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01266 | 2021-05-03 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, ERROR CODES RELATED TO THE CHARGING OF THE INTERNAL BATTERY WERE OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01265 | 2021-05-03 | 2021-04-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUE. THE DEVICE HAD EVIDENCE OF PHYSICAL DAMAGE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01272 | 2021-05-03 | 2021-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1032800 | 1032800 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED THE FLOW VERIFICATION TEST STEP DURING TESTING. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR FAILED THE FLOW VERIFICATION TEST STEP DURING TESTING. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01273 | 2021-05-03 | 2021-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR WAS ALARMING FOR A "VENTILATOR INOPERATIVE" ALARM CONDITION. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BLOWER MOTOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01268 | 2021-05-03 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01269 | 2021-05-03 | 2021-04-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE AND OVERHEAD BLOWER FILTER WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01271 | 2021-05-03 | 2021-04-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, A FAILURE OF THE DEVICE TO DELIVER FLOW WAS OBSERVED. THE DEVICE'S TRANSITION TUBE WAS FOUND TO BE TWISTED. THE DEVICE'S FLOW SENSOR ASSEMBLY WAS REPLACED TO ADDRESS THE ISSUE. THE DEVICE HAD EVIDENCE OF PHYSICAL DAMAGE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01274 | 2021-05-03 | 2021-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR WAS ALARMING FOR A "VENTILATOR INOPERATIVE" ALARM CONDITION. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01270 | 2021-05-03 | 2021-04-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01217 | 2021-05-04 | 2021-04-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE RELATED TO HIGH TEMPERATURE WAS FOUND IN THE VENTILATOR'S DOWNLOADED EVENT LOG. THE INCREASED AIRSTREAM TEMPERATURE (HIGH TEMPERATURE ALARM) APPEARS TO HAVE BEEN CAUSED BY THE AMBIENT CONDITIONS IN WHICH THE DEVICE WAS OPERATING. NO MALFUNCTIONS WERE NOTED. THE DEVICE WAS FOUND TO AUDIBLY AND VISUALLY ALARM, AS DESIGNED. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01245 | 2021-05-04 | 2021-04-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED EVENT LOG. THE DEVICE'S INTERFACE BOARD AND OXYGEN BLENDING MODULE BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01261 | 2021-05-04 | 2021-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01263 | 2021-05-04 | 2021-04-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01302 | 2021-05-04 | 2021-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040003 | 1040003 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS A "SERVICE REQUIRED" ALARM CONDITION. DURING FURTHER TESTING AT THE THIRD PARTY SERVICE CENTER, THE SENSOR BOARD WAS REPLACED TO ADDRESS AN ADDITIONAL SERVICE REQUIRED ALARM CONDITION. | N | Y |  |  | 2021-05
2518422-2021-01257 | 2021-05-04 | 2021-03-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS HAVING AN INTERNAL BATTERY ISSUE. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERFACE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01258 | 2021-05-04 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01264 | 2021-05-04 | 2021-04-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01260 | 2021-05-04 | 2021-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED THAT THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS AN OBSERVED ISSUE. THE DEVICE LATER FAILED A STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01259 | 2021-05-04 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01194 | 2021-05-04 | 2021-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01195 | 2021-05-04 | 2021-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE DEVICE FAILED STEPS DURING TESTING. THE DEVICE'S DETACHABLE BATTERY AND CONNECTOR WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01304 | 2021-05-04 | 2021-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | GB2110X15B | GB2110X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01262 | 2021-05-04 | 2021-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01303 | 2021-05-04 | 2021-04-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01294 | 2021-05-05 | 2021-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, A HIGH TEMPERATURE ALARM CONDITION WAS OBSERVED. THE INCREASED AIRSTREAM TEMPERATURE (HIGH TEMPERATURE ALARM) APPEARS TO HAVE BEEN CAUSED BY THE AMBIENT CONDITIONS IN WHICH THE DEVICE WAS OPERATING. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01242 | 2021-05-05 | 2021-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ERROR WITH A VENTILATOR'S FLOW SENSOR. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE FLOW SENSOR CABLE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01291 | 2021-05-05 | 2021-04-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, THE DEVICE'S TIDAL VOLUME DELIVERY WAS INACCURATE. THE DEVICE'S EXTERNAL FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01296 | 2021-05-05 | 2021-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED A TEST STEP DURING TESTING. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01292 | 2021-05-05 | 2021-04-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | FX2100X15B | FX2100X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, THE DEVICE'S BLOWER FAILED TO FUNCTION. THE DEVICE'S BLOWER MOTOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01295 | 2021-05-05 | 2021-04-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TO CHARGE ITS INTERNAL BATTERY. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01293 | 2021-05-05 | 2021-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01290 | 2021-05-05 | 2021-04-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WOULD NOT POWER ON. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. THE DEVICE WARRANTY HAS EXPIRED AND THE CUSTOMER IS NOT WILLING TO PAY FOR REPAIRS. NO REPAIRS WILL BE DONE ON THIS DEVICE AT THIS TIME. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01333 | 2021-05-05 | 2021-04-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1032800 | 1032800 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED A STEP DURING TESTING. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR FAILED A STEP DURING TESTING. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR EVALUATION, AND THE CUSTOMER'S COMPLAINT WAS UNABLE TO BE DUPLICATED. THE DEVICE WAS FOUND TO OPERATE AND ALARM AS DESIGNED. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01344 | 2021-05-05 |  | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S TOUCHSCREEN DOES NOT FUNCTION. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | IT WAS INITIALLY REPORTED, THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S TOUCHSCREEN DOES NOT FUNCTION. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE CUSTOMER'S COMPLAINT COULD NOT BE CONFIRMED OR DUPLICATED. THE DEVICE WAS FOUND TO OPERATE AND ALARM AS DESIGNED. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01324 | 2021-05-06 | 2021-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. THE DEVICE HAD EVIDENCE OF PHYSICAL DAMAGE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01326 | 2021-05-06 | 2021-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01327 | 2021-05-06 | 2021-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE RELATED TO THE REMOTE ALARM CONNECTOR WAS OBSERVED. THE DEVICE'S INTERFACE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01325 | 2021-05-06 | 2021-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S SENSOR BOARD AND TUBING KIT WERE REPLACED TO ADDRESS THE ISSUES. THE DEVICE HAD EVIDENCE OF WATER INGRESS. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01346 | 2021-05-06 | 2021-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED EVENT LOG. THE DEVICE'S INTERNAL BATTERY NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01345 | 2021-05-06 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, A VENTILATOR INOPERATIVE CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED EVENT LOG. THE DEVICE'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. A SERVICE REQUIRED CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED EVENT LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01350 | 2021-05-06 | 2021-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR HAD AN OXYGEN DELIVERY ISSUE. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS EVALUATED BY THE MANUFACTURER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01351 | 2021-05-06 | 2021-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR HAD AN OXYGEN DELIVERY ISSUE. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS EVALUATED BY THE MANUFACTURER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01349 | 2021-05-06 | 2021-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01347 | 2021-05-06 | 2021-05-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED STEPS DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01196 | 2021-05-07 | 2021-04-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01197 | 2021-05-07 | 2021-04-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01363 | 2021-05-07 | 2021-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE FAILED TO POWER ON. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE ACTIVE EXHALATION PORTING BLOCK WAS FOUND TO BE PHYSICALLY DAMAGED AND WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01360 | 2021-05-07 | 2021-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BLOWER MOTOR, MACHINE FLOW SENSOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. THE DEVICE HAD EVIDENCE OF WATER INGRESS. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01361 | 2021-05-07 | 2021-04-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR DOES NOT DISPLAY TIDAL VOLUMES. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01359 | 2021-05-07 | 2021-05-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
11789841 | 2021-05-07 | 2021-01-28 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | DS2200X11B | DS2200X11B |  | PATIENT ON NOCTURNAL VENTILATION. TURNED OFF AFTER PATIENT REMOVED DEVICE INTERFACE, TURNED BACK ON TO REINITIATE THERAPY AS PATIENT FELL ASLEEP. UPON POWERING UP DEVICE ALARMED TO SERVICE UNIT - THIS THERAPIST RESET ALARM AND DID BASIC TROUBLESHOOTING. ALARM DID NOT REOCCUR AND THERAPY WAS RESTARTED. 1 HOUR LATER ALARM RESUMED WHILE PATIENT WAS WEARING. PATIENT WAS PLACED ON SUPPLEMENTAL O2 AND DEVICE REMOVED FROM ROOM WITH INTENTION TO REPLACE. |  | N | Y | I |  | 2021-05
2518422-2021-01357 | 2021-05-07 | 2021-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01356 | 2021-05-07 | 2021-04-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION THE MACHINE FLOW SENSOR WAS FOUND TO BE FAULTY AND WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01319 | 2021-05-07 | 2021-04-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11 | DSX700H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE DEVICE HAD A BURNING SMELL DURING USE. THE USER OPENED THE DEVICE AND FOUND A BLACK SUBSTANCE IN THE BLOWER AND SURROUNDING AREA. THE USER STATES HE HAS A COUGH, SORE THROAT AND CHEST PAINS RELATED TO DEVICE USE. THE MANUFACTURER REQUESTED THE RETURN OF THE DEVICE AND ALL ACCESSORIES FOR EVALUATION. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THE END USER INDICATED OZONE OR OTHER UNAPPROVED CLEANING AND DISINFECTION METHOD WAS USED. THIS ACTION WAS REPORTED TO FDA PER 21 CRF PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01362 | 2021-05-07 | 2021-04-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S DISPLAY BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01355 | 2021-05-07 | 2021-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, A MALFUNCTION OF THE TIDAL VOLUME DELIVERY WAS FOUND. THE DEVICE'S FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01358 | 2021-05-07 | 2021-05-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01338 | 2021-05-10 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICES LCD DISPLAY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01339 | 2021-05-10 | 2021-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES RELATED TO THE CHARGING OF THE INTERNAL BATTERY WERE OBSERVED IN THE DOWNLOADED EVENT LOG. THE DEVICE'S INTERNAL BATTERY AND POWER MANAGEMENT BOARD WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01375 | 2021-05-10 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01342 | 2021-05-10 | 2021-04-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICES LCD DISPLAY NEEDS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED THE INVESTIGATION WAS ONGOING AND A FOLLOW UP REPORT WOULD BE SUBMITTED WHEN THE INVESTIGATION WAS COMPLETE. NO FOLLOW UP REPORT IS REQUIRED. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01368 | 2021-05-10 | 2021-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2200X11B | DS2200X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01371 | 2021-05-10 | 2021-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TO DELIVER ACCURATE PRESSURE. THERE WAS NO HARM OR INJURY REPORTED. THE MANUFACTURER RECEIVED THE DEVICE FOR EVALUATION, AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OVERHEAD BLOWER FILTER AND ACTIVE EXHALATION CONTROL MODULE WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01369 | 2021-05-10 | 2021-04-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TO DELIVER PRESSURE. THERE WAS NO HARM OR INJURY REPORTED. THE MANUFACTURER RECEIVED THE DEVICE FOR EVALUATION, AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE DEVICE'S DOWNLOADED ERROR LOG. THE DEVICE'S BLOWER MOTOR NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01337 | 2021-05-10 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01372 | 2021-05-10 | 2021-04-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR A SERVICE REQUIRED CONDITION RELATED TO THE DEVICE'S OXYGEN FLOW. THERE WAS NO HARM OR INJURY REPORTED. THE MANUFACTURER RECEIVED THE DEVICE FOR EVALUATION, AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01370 | 2021-05-10 | 2021-04-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S DISPLAY WAS BLANK AND UNABLE TO BE VIEWED. THERE WAS NO HARM OR INJURY REPORTED. THE MANUFACTURER RECEIVED THE DEVICE FOR EVALUATION, AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S LCD DISPLAY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01374 | 2021-05-10 | 2021-04-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2200X15B | IN2200X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE FOR A VENTILATOR INOPERATIVE CONDITION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01335 | 2021-05-10 | 2021-04-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01336 | 2021-05-10 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THIRD PARTY SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01341 | 2021-05-10 | 2021-04-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01340 | 2021-05-10 | 2021-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1032804 | 1032804 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01241 | 2021-05-11 | 2021-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR), | RESPIRONICS, INC. | RJPX500S16 | RJPX500S16 |  | THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF EXPOSED WIRES ON A DAMAGED AC POWER CORD OF A CONTINUOUS POSITIVE AIRWAY DEVICE (CPAP) DEVICE. THERE WAS NO PATIENT HARM OR INJURY. THE SALES REP CONFIRMED EVIDENCE OF EXPOSED AC WIRES ON THE AC POWER CORD. THE POWER CORD WAS REPLACED FOR THE PATIENT. THERE WAS NO EVIDENCE OF THERMAL DAMAGE. THE DREAMSTATION USER MANUAL STATES: WARNING: PERIODICALLY INSPECT ELECTRICAL CORDS AND CABLES FOR DAMAGE OR SIGN OF WEAR. DISCONTINUE USE AND REPLACE IF DAMAGED. | THE MANUFACTURER RECEIVED THE AC POWER CORD FOR EVALUATION. THE FINDINGS REVEALED THE SHEATH OF THE POWER CORD HAD GOT OUT OF POSITION AND HAD CAUSED THE INTERNAL BLACK AND WHITE WIRES TO BE EXPOSED. THE COPPER CONDUCTING WIRE WAS NOT EXPOSED. LABELING INSTRUCTS THE USER TO PERIODICALLY INSPECT THE AC POWER CORD FOR SIGNS OF WEAR OR DAMAGE AND TO REPLACE THE AC POWER CORD IF WORN OR DAMAGED. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01378 | 2021-05-11 | 2021-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED (PHYSICAL DAMAGE). THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01377 | 2021-05-11 | 2021-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01343 | 2021-05-11 | 2021-04-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2100X15B | IN2100X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION AT THE THIRD PARTY SERVICE CENTER, A MALFUNCTION OF THE EXTERNAL FLOW SENSOR WAS OBSERVED. THE EXTERNAL FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01391 | 2021-05-11 | 2021-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BLOWER MOTOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01390 | 2021-05-11 | 2021-04-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S DISPLAY BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01413 | 2021-05-11 | 2021-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS NOT DELIVERING THE CORRECT FIO2. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01427 | 2021-05-12 | 2021-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2110X15B | IN2110X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01426 | 2021-05-12 | 2021-04-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A MALFUNCTION OF A VENTILATOR'S TOUCHSCREEN. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A FAILURE OF A VENTILATOR'S TOUCHSCREEN. THERE WAS NO PATIENT HARM OR INJURY REPORTED. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT COULD NOT BE DUPLICATED. THE DEVICE WAS FOUND TO OPERATE AND ALARM AS DESIGNED. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01432 | 2021-05-12 | 2021-04-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR SHUT DOWN DURING USE. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR SHUT DOWN DURING USE. THERE WAS NO HARM OR INJURY REPORTED. FURTHER INFORMATION WAS RECEIVED FROM THE MANUFACTURER'S FIELD SERVICE ENGINEER STATING THE ISSUE WITH THE REPORTED ALLEGATION WAS RELATED TO A NONFUNCTIONING OUTLET THE DEVICE WAS PLUGGED INTO. THE VENTILATOR FUNCTIONED AS DESIGNED WHEN CONNECTED TO A KNOWN WORKING POWER OUTLET. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01431 | 2021-05-12 | 2021-04-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR SHUT DOWN DURING USE. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR SHUT DOWN DURING USE. THERE WAS NO HARM OR INJURY REPORTED. FURTHER INFORMATION WAS RECEIVED FROM THE MANUFACTURER'S FIELD SERVICE ENGINEER STATING THE ISSUE WITH THE REPORTED ALLEGATION WAS RELATED TO A NONFUNCTIONING OUTLET THE DEVICE WAS PLUGGED INTO. THE VENTILATOR FUNCTIONED AS DESIGNED WHEN CONNECTED TO A KNOWN WORKING POWER OUTLET. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01394 | 2021-05-12 | 2021-04-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION, AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01393 | 2021-05-12 | 2021-04-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S TOUCHSCREEN FAILED TO RESPOND. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION, AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S DISPLAY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01334 | 2021-05-12 | 2021-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM OCCURRED RELATED TO A MALFUNCTION OF THE VENTILATOR'S TOUCHSCREEN. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S DISPLAY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01383 | 2021-05-12 | 2021-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A HIGH TEMPERATURE ALARM WAS OBSERVED. THE DEVICE'S BLOWER MOTOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01385 | 2021-05-12 | 2021-04-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE LCD SCREEN WAS OBSERVED TO BE BLANK AND UNREADABLE. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01386 | 2021-05-12 | 2021-05-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01388 | 2021-05-12 | 2021-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01389 | 2021-05-12 | 2021-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S INTERFACE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01384 | 2021-05-12 | 2021-04-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01376 | 2021-05-12 | 2021-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S MACHINE FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01387 | 2021-05-12 | 2021-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01406 | 2021-05-12 | 2021-04-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | RIN2110X15B | RIN2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S TIDAL VOLUME WAS LOW. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR'S TIDAL VOLUME WAS LOW. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01429 | 2021-05-12 | 2021-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE WAS RECALIBRATED AND THE OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01373 | 2021-05-12 | 2021-05-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | PHILIPS RESPIRONICS INC. | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01397 | 2021-05-12 | 2021-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR A "SERVICE REQUIRED" CONDITION RELATED TO THE DEVICE'S INTERNAL BATTERY. THERE WAS NO HARM OR INJURY REPORTED. THE MANUFACTURER RECEIVED THE DEVICE FOR EVALUATION, AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S INTERNAL BATTERY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01428 | 2021-05-12 | 2021-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A MALFUNCTION OF A VENTILATOR'S TOUCHSCREEN. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S DISPLAY WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED THIS MALFUNCTION ON MDR 2518422-2021-01334. REPORT, MDR 2518422-2021-01428 IS A DUPLICATE AND WAS FILED IN ERROR. NO DUPLICATE MALFUNCTION OCCURRED AND ALL REPORTING IS CORRECT ON MDR 2518422-2021-01334. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01434 | 2021-05-12 | 2021-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01433 | 2021-05-12 | 2021-04-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS SOUNDING PRESSURE ALARMS. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. THE DEVICE HAD EVIDENCE OF WATER INGRESS. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01430 | 2021-05-12 | 2021-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE AUTO CYCLING WAS OBSERVED. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01443 | 2021-05-14 | 2021-04-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S DISPLAY WAS BLANK. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S DISPLAY WAS REPLACED DUE TO PHYSICAL DAMAGE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01448 | 2021-05-17 | 2021-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, A VENTILATOR FAILED TO CHARGE ITS INTERNAL BATTERY. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S POWER MANAGEMENT BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01411 | 2021-05-17 | 2021-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01398 | 2021-05-17 | 2021-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S POWER MANAGEMENT BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01400 | 2021-05-17 | 2021-05-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED EVENT LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01410 | 2021-05-17 | 2021-05-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01409 | 2021-05-17 | 2021-05-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THIRD PARTY SERVICE CENTER, "SERVICE REQUIRED" CODES WERE OBSERVED IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY AND SENSOR BOARD WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01407 | 2021-05-17 | 2021-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01395 | 2021-05-17 | 2021-04-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS DELIVERING AN INACCURATE TIDAL VOLUME. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION, AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S EXTERNAL FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01396 | 2021-05-17 | 2021-04-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | DS2200X11B | DS2200X11B |  | A VENTILATOR WAS EVALUATED BY A FIELD SERVICE ENGINEER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED EVENT LOG. THE DEVICE'S OXYGEN BLENDING MODULE ASSEMBLY, SYSTEM BOARD AND OXYGEN BLENDING MODULE HARNESS NEED REPLACED TO ADDRESS THE ISSUES. THE DEVICE'S OXYGEN SENSOR FAILED TO CALIBRATE. THE DEVICE'S OXYGEN SENSOR NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01399 | 2021-05-17 | 2021-05-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S BLOWER MOTOR WAS OBSERVED. THE DEVICE'S BLOWER MOTOR WAS REPLACED TO ADDRESS THE ISSUE |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01457 | 2021-05-17 | 2021-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR A "SERVICE REQUIRED' CONDITION RELATED TO HIGH INTERNAL OXYGEN. THE DEVICE WAS NOT IN PATIENT USE. THE MANUFACTURER RECEIVED THE DEVICE FOR EVALUATION, AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN BLENDING MODULE MANIFOLD NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01435 | 2021-05-17 | 2021-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE'S DISPLAY WAS BLANK. THE LCD SCREEN WAS REPLACED DUE TO PHYSICAL DAMAGE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01408 | 2021-05-17 | 2021-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S POWER SUPPLY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01476 | 2021-05-18 | 2021-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01463 | 2021-05-18 | 2021-04-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE MANUFACTURER RECEIVED THE DEVICE FOR EVALUATION, AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BLOWER MOTOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01474 | 2021-05-18 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S PRESSURE DELIVERY WAS INACCURATE. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED AND THE OVERHEAD BLOWER FILTER WAS CLOGGED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE AND OVERHEAD BLOWER FILTER NEED TO BE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01477 | 2021-05-18 | 2021-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01473 | 2021-05-18 | 2021-04-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054097 | 1054097 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PRESSURE VALVE WAS OUT OF TOLERANCE. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR'S PRESSURE VALVE WAS OUT OF TOLERANCE. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01475 | 2021-05-18 | 2021-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01482 | 2021-05-18 | 2021-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01460 | 2021-05-18 | 2021-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01450 | 2021-05-18 | 2021-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR HAD A "HIGH INTERNAL OXYGEN" ALARM . THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01470 | 2021-05-18 | 2021-04-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S POWER MANAGEMENT BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01478 | 2021-05-18 | 2021-04-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR HAD A BLACK LCD SCREEN . THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD AND LCD CABLE WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01458 | 2021-05-18 | 2021-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S PRESSURE DELIVERY WAS NOT ACCURATE. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR WAS FOUND TO HAVE WATER INGRESS AND WAS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S DISPLAY BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01461 | 2021-05-18 | 2021-05-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01459 | 2021-05-18 | 2021-05-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01462 | 2021-05-18 | 2021-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01456 | 2021-05-18 | 2021-04-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S INTERFACE BOARD AND OXYGEN BLENDING MODULE BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01449 | 2021-05-18 | 2021-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WOULD NOT RECOGNIZE THE CIRCUIT . THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S FLOW SENSOR ASSEMBLY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01471 | 2021-05-18 | 2021-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01484 | 2021-05-19 | 2021-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01483 | 2021-05-19 | 2021-05-17 | D, e, a, t, h | Manufacturer report | TRILOGY EVO O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS DELIVERING AN INACCURATE TIDAL VOLUME TO A PATIENT. THE PATIENT DEVELOPED BAROTRAUMA AND SUBSEQUENTLY EXPIRED. THE DATE OF DEATH IS CURRENTLY UNKNOWN. THE MANUFACTURER HAS REQUESTED FURTHER DETAILS REGARDING THE ALLEGATION AND THE RETURN OF THE DEVICE TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LAB FOR EVALUATION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF A VENTILATOR DELIVERING AN INACCURATE TIDAL VOLUME TO A PATIENT. THE PATIENT DEVELOPED BAROTRAUMA AND SUBSEQUENTLY EXPIRED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR EVALUATION AND THE CUSTOMER'S COMPLAINT COULD NOT BE DUPLICATED. THE DEVICE WAS FOUND TO OPERATE AND ALARM AS DESIGNED. | Y | N | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01496 | 2021-05-19 | 2021-05-17 | I, n, j, u, r, y | Manufacturer report | TRILOGY EVO O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS DELIVERING AN INACCURATE TIDAL VOLUME TO A PATIENT. THE PATIENT DEVELOPED BAROTRAUMA AND REQUIRED PROLONGED HOSPITALIZATION. THE MANUFACTURER HAS REQUESTED FURTHER DETAILS REGARDING THE ALLEGATION AND THE RETURN OF THE DEVICE TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LAB FOR EVALUATION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF A VENTILATOR DELIVERING AN INACCURATE TIDAL VOLUME TO A PATIENT. THE PATIENT DEVELOPED BAROTRAUMA AND REQUIRED PROLONGED HOSPITALIZATION. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR EVALUATION AND THE CUSTOMER'S COMPLAINT COULD NOT BE DUPLICATED. THE DEVICE WAS FOUND TO OPERATE AND ALARM AS DESIGNED. | Y | N | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01412 | 2021-05-19 | 2021-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01498 | 2021-05-19 | 2021-05-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. THERE WAS EVIDENCE OF PHYSICAL DAMAGE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01497 | 2021-05-19 | 2021-05-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01485 | 2021-05-19 | 2021-04-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2200X11B | DS2200X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR'S SYSTEM BOARD WAS REPLACED BY A THIRD PARTY SERVICE CENTER TO ADDRESS A VENTILATOR INOPERATIVE ALARM CONDITION. THE DEVICE'S BLOWER MOTOR WAS ALSO REPLACED TO ADDRESS THIS ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01499 | 2021-05-19 | 2021-05-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | RIN2110X15B | RIN2110X15B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO PATIENT HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01486 | 2021-05-19 | 2021-04-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | EU2110X15B | EU2110X15B |  | A PEDIATRIC PATIENT WAS ADMITTED TO AN INTENSIVE CARE UNIT DUE TO WORSENING CONDITION. IT IS ALLEGED THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE (PEEP) INCREASED 12 CMH2O ABOVE THE SET VALUE. THIS ISSUE WAS ALLEGEDLY CORRECTED BY POWERING DOWN AND RESTARTING THE DEVICE. THE MANUFACTURER IS CURRENTLY INVESTIGATING THIS ISSUE AND A FOLLOW UP REPORT WILL BE FILED WHEN THE INVESTIGATION IS COMPLETE. | IT WAS INITIALLY REPORTED, A PEDIATRIC PATIENT WAS ADMITTED TO AN INTENSIVE CARE UNIT DUE TO WORSENING CONDITION. IT IS ALLEGED THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE (PEEP) INCREASED 12 CMH2O ABOVE THE SET VALUE. THIS ISSUE WAS ALLEGEDLY CORRECTED BY POWERING DOWN AND RESTARTING THE DEVICE. THE DEVICE'S SOFTWARE NEEDS RELOADED TO ADDRESS THIS ISSUE. REPEATED ATTEMPTS FOR ADDITIONAL INFORMATION OR TO HAVE THE DEVICE RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | N | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01414 | 2021-05-19 | 2021-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | PHILIPS RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01479 | 2021-05-19 | 2021-04-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS EVALUATED BY A THIRD PARTY FIELD SERVICE ENGINEER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE BY THE THIRD PARTY FIELD SERVICE ENGINEER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01469 | 2021-05-20 | 2021-05-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED EVENT LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01472 | 2021-05-20 | 2021-05-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01508 | 2021-05-20 | 2021-05-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TO POWER ON. THERE WAS NO HARM OR INJURY REPORTED. THE MANUFACTURER RECEIVED THE DEVICE FOR INVESTIGATION, AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BLOWER MOTOR NEEDS REPLACED TO ADDRESS THE ISSUE. THERE WAS EVIDENCE OF CONTAMINATION. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01467 | 2021-05-20 | 2021-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED STEPS DURING TESTING. THE DEVICE'S SENSOR BOARD AND ACTIVE EXHALATION CONTROL MODULE NEED REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01468 | 2021-05-20 | 2021-05-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S SENSOR BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. A "SERVICE REQUIRED CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED EVENT LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. AN ISSUE RELATED TO THE DEVICE OVERHEATING WAS OBSERVED. THE DEVICE'S BLOWER MOTOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01437 | 2021-05-20 | 2021-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A ISSUES RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WERE OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01507 | 2021-05-20 | 2021-05-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S DISPLAY WAS BLANK AND UNABLE TO BE VIEWED. THE DEVICE WAS NOT IN PATIENT USE. THE MANUFACTURER RECEIVED THE DEVICE FOR INVESTIGATION, AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S LCD DISPLAY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01465 | 2021-05-20 | 2021-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS EVALUATED BY A THIRD-PARTY, AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01466 | 2021-05-20 | 2021-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | RIN2100X15B | RIN2100X15B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED EVENT LOG. THE DEVICE'S PROPORTIONAL VALVE AND SOLENOID VALVE NEED REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01464 | 2021-05-20 | 2021-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE MANUFACTURER RECEIVED THE DEVICE FOR EVALUATION, AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD NEED REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01480 | 2021-05-20 | 2021-04-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TO CHARGE ITS INTERNAL BATTERY. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S POWER MANAGEMENT BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01500 | 2021-05-20 | 2021-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S DISPLAY WAS BLANK. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S LCD SCREEN AND CABLE WERE REPLACED TO ADDRESS THE ISSUE. | IT WAS INITIALLY REPORTED, THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S DISPLAY WAS BLANK. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S LCD SCREEN AND CABLE WERE REPLACED TO ADDRESS THE ISSUE. THE DEVICE WAS RETURNED UNREPAIRED AS PER THE CUSTOMER. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01436 | 2021-05-20 | 2021-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01451 | 2021-05-20 | 2021-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILS TO CHARGE ITS INTERNAL BATTERY. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01454 | 2021-05-20 | 2021-05-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01493 | 2021-05-20 | 2021-04-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | EE2110X15B | EE2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SET PRESSURE INCREASED WITHOUT MANUAL ADJUSTMENT. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VARIATION IN THE VENTILATOR'S PEAK INSPIRATORY PRESSURE DELIVERY AND POSITIVE END EXPIRATORY PRESSURE WAS OBSERVED. ADDITIONAL INFORMATION RECEIVED INDICATES THE VENTILATOR'S SOFTWARE WAS RE-LOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01488 | 2021-05-20 | 2021-04-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | EE2110X15B | EE2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SET PRESSURE INCREASED WITHOUT MANUAL ADJUSTMENT. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VARIATION IN THE VENTILATOR'S PEAK INSPIRATORY PRESSURE DELIVERY AND POSITIVE END EXPIRATORY PRESSURE WAS OBSERVED. ADDITIONAL INFORMATION RECEIVED INDICATES THE VENTILATOR'S SOFTWARE WAS RE-LOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01492 | 2021-05-20 | 2021-04-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | EE2110X15B | EE2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SET PRESSURE INCREASED WITHOUT MANUAL ADJUSTMENT. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VARIATION IN THE VENTILATOR'S PEAK INSPIRATORY PRESSURE DELIVERY AND POSITIVE END EXPIRATORY PRESSURE WAS OBSERVED. ADDITIONAL INFORMATION RECEIVED INDICATES THE VENTILATOR'S SOFTWARE WAS RE-LOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01490 | 2021-05-20 | 2021-04-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | EE2110X15B | EE2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SET PRESSURE INCREASED WITHOUT MANUAL ADJUSTMENT. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VARIATION IN THE VENTILATOR'S PEAK INSPIRATORY PRESSURE DELIVERY AND POSITIVE END EXPIRATORY PRESSURE WAS OBSERVED. ADDITIONAL INFORMATION RECEIVED INDICATES THE VENTILATOR'S SOFTWARE WAS RE-LOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01491 | 2021-05-20 | 2021-04-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | EE2110X15B | EE2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SET PRESSURE INCREASED WITHOUT MANUAL ADJUSTMENT. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VARIATION IN THE VENTILATOR'S PEAK INSPIRATORY PRESSURE DELIVERY AND POSITIVE END EXPIRATORY PRESSURE WAS OBSERVED. ADDITIONAL INFORMATION RECEIVED INDICATES THE VENTILATOR'S SOFTWARE WAS RE-LOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01494 | 2021-05-20 | 2021-04-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | EE2110X15B | EE2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SET PRESSURE INCREASED WITHOUT MANUAL ADJUSTMENT. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VARIATION IN THE VENTILATOR'S PEAK INSPIRATORY PRESSURE DELIVERY AND POSITIVE END EXPIRATORY PRESSURE WAS OBSERVED. ADDITIONAL INFORMATION RECEIVED INDICATES THE VENTILATOR'S SOFTWARE WAS RE-LOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01495 | 2021-05-20 | 2021-04-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | EE2110X15B | EE2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SET PRESSURE INCREASED WITHOUT MANUAL ADJUSTMENT. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN INCREASE IN THE VENTILATOR'S SET PRESSURE WITHOUT MANUAL ADJUSTMENT. ADDITIONAL INFORMATION RECEIVED INDICATES THE VENTILATOR'S SOFTWARE WAS RE-LOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01481 | 2021-05-20 | 2021-05-09 | I, n, j, u, r, y | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S PEAK INSPIRATORY PRESSURES DID NOT RISE WHILE IN USE ON A PATIENT. THE PATIENT WAS UNABLE TO BREATHE EASILY AND WAS SWITCHED TO ANOTHER VENTILATOR. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER AND THE CUSTOMER'S COMPLAINT WAS NOT CONFIRMED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING AND SERVICE REQUIRED CODES WERE OBSERVED IN THE DOWNLOADED EVENT LOG. THE DEVICE'S OVERHEAD BLOWER FILTER NEEDS REPLACED TO ADDRESS THE ISSUES. THERE WAS EVIDENCE OF CONTAMINATION. |  | Y | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01452 | 2021-05-20 | 2021-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OUTLET FLOW PATH THERMISTOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01453 | 2021-05-20 | 2021-05-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN MANIFOLD ASSEMBLY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01489 | 2021-05-20 | 2021-04-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | EE2110X15B | EE2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SET PRESSURE INCREASED WITHOUT MANUAL ADJUSTMENT. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VARIATION IN THE VENTILATOR'S PEAK INSPIRATORY PRESSURE DELIVERY AND POSITIVE END EXPIRATORY PRESSURE WAS OBSERVED. ADDITIONAL INFORMATION RECEIVED INDICATES THE VENTILATOR'S SOFTWARE WAS RE-LOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01512 | 2021-05-21 | 2021-05-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01524 | 2021-05-21 | 2021-05-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLANK DURING USE AND WAS NOT VIEWABLE. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01509 | 2021-05-21 | 2021-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01523 | 2021-05-21 | 2021-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED EVENT LOG. THE DEVICE'S INTERNAL BATTERY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01513 | 2021-05-21 | 2021-05-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01522 | 2021-05-21 | 2021-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE FAILED STEPS DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUES. | CORRECTED SECTION G3 FROM (B)(6) 2020 TO CORRECT DATE OF (B)(6) 2021 | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01503 | 2021-05-21 | 2021-05-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE SPEAKERS TO AUDIBLY ALARM WAS OBSERVED. THE DEVICE'S SPEAKER KIT WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01511 | 2021-05-21 | 2021-05-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE DEVICE FAILED STEPS DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD, FLOW SENSOR TUBE AND ACTIVE EXHALATION CONTROL MODULE WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01455 | 2021-05-21 | 2021-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER HAD PREVIOUSLY REPORTED THAT THE DEVICE'S INTERNAL BATTERY HAD BEEN REPLACED TO ADDRESS AN ERROR CODE IN THE EVENT LOG. THIS WAS INCORRECT. THE DEVICE WAS RETURNED UNREPAIRED DUE TO REJECTION OF SERVICE ESTIMATE AND THE INTERNAL BATTERY WAS NOT REPLACED. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01504 | 2021-05-21 | 2021-05-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01501 | 2021-05-21 | 2021-04-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED AND THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01506 | 2021-05-21 | 2021-05-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S REMOTE ALARM CONNECTOR FAILED TO CONNECT. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S INTERFACE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01502 | 2021-05-21 | 2021-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD NEED TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01510 | 2021-05-21 | 2021-05-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2200X11B | DS2200X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S PROPORTIONAL VALVE, SOLENOID VALVE AND SENSOR BOARD NEED REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01520 | 2021-05-24 | 2021-05-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01532 | 2021-05-24 | 2021-05-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERFACE BOARD WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED THE DEVICE'S INTERFACE BOARD WAS REPLACED FOR A "SERVICE REQUIRED" ALARM CONDITION. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS ALSO REPLACED TO ADDRESS THIS ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01514 | 2021-05-24 | 2021-05-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. | IT WAS INITIALLY REPORTED, DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. THE OXYGEN BLENDING MODULE BOARD WAS NOT REPLACED, THE INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01515 | 2021-05-24 | 2021-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01420 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500T11 | DSX500T11 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01299 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX600H11C | DSX600H11C |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01322 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01321 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX600H11C | DSX600H11C |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01276 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01425 | 2021-05-25 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTION H6 TYPE OF INVESTIGATION ,INVESTIGATION FINDINGS,INVESTIGATION CONCLUSIONS WAS NOT CAPTURED IN PREVIOUS MDR,NOW IT IS CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01354 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | CAX500T12 | CAX500T12 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01318 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01309 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER INSPECTED THE EXTERNAL PART AND INTERNAL PART OF THE DEVICE AND FOUND BLOWER FOAM DEGRADATION AND BLOWER FOAM PARTICLE INSIDE. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND FIVE ERROR CODES. THE MANUFACTURER REPLACED PARTS, CLEAR ERROR LOG, RESET TIME CLOCK. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION H6 WERE UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01279 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01307 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01308 | 2021-05-25 | 2021-04-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED NOISY, MASK FIT READING LOW, LEAK INSIDE. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED DUE TO AGE. SECTION H3 WAS CAPTURED INCORRECTLY AS 'NOT RETURNED TO THE MANUFACTURER'. B3 WAS CORRECTED WHICH WAS CAPTURED INCORRECTLY IN THE INITIAL REPORT AND SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01277 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01297 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01402 | 2021-05-25 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01320 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX400T11C | DSX400T11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01300 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01275 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11W | DSX500H11W |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01287 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500T11C | DSX500T11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01314 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX700T11C | DSX700T11C |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01329 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01289 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX400H11 | DSX400H11 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01404 | 2021-05-25 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01301 | 2021-05-25 | 2021-03-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX400H11C | DSX400H11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01278 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400T11 | DSX400T11 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE 32 ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS FAULTY BLOWER AND EVIDENCE OF FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01403 | 2021-05-25 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01280 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX400H11 | DSX400H11 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE THIRD PARTY SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED. THERE WERE NO ERRORS FOUND. THE THIRD PARTY SERVICE CENTER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT WAS SCRAPPED IN THIS REPORT, SECTION D8, D9, H2, H3 AND H6 HAS BEEN UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01418 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX700S11 | DSX700S11 |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01332 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500H11C | DSX500H11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01328 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01365 | 2021-05-25 | 2021-04-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX400T11C | DSX400T11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01313 | 2021-05-25 | 2021-03-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500H11C | DSX500H11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01312 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400T11C | DSX400T11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING ISSUES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 WAS UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01286 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED DEVICE WILL NOT ALLOW THE HUMIDITY SETTING TO CHANGE THE SETTING FROM THE THE PT MODE AND DEVICE WAS NOT IN PATIENT USE. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE 32 ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. SECTION H6 WAS UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01417 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500T11C | DSX500T11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01316 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11 | DSX700T11 |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED SERVICE REQUIRED, BEEPING AND NO AIR FLOW. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. UNIT SCRAPPED DUE TO DATE 2017-04-25. SECTION H6 UPDATED IN THIS REPORT. SECTION A1 WAS CORRECTED BY CAPTURING PATIENT IDENTIFIER WHICH WAS MISSED IN INITIAL REPORT. SECTION E1 WAS CORRECTED BY CAPTURING NAME AND ADDRESS WHICH WAS INCORRECT IN INITIAL REPORT. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01288 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX700H11C | DSX700H11C |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01331 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500T11C | DSX500T11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01285 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT PASSED THE FINAL TEST. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01364 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11W | DSX500T11W |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED TO SEEING PARTICLES IN THE AIR PATH. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERROR FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS VISIBLE FOAM DEGRADATION AND UNIT IS SCRAPPED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01317 | 2021-05-25 | 2021-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01315 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX200H11 | DSX200H11 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01311 | 2021-05-25 | 2021-03-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01284 | 2021-05-25 | 2021-04-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01281 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11W | DSX500H11W |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01401 | 2021-05-25 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX700H11C | DSX700H11C |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01310 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01298 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE, THE DEVICE WAS NOT IN PATIENT USE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01405 | 2021-05-25 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01283 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01419 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX700H11C | DSX700H11C |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01282 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01542 | 2021-05-25 | 2021-05-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO PATIENT HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01539 | 2021-05-25 | 2021-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. THE ESTIMATE WAS REJECTED BY THE CUSTOMER PER THEIR REQUEST AND RETURNED UNREPAIRED. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01416 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTOSV ADV | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS INC. | 960P | 960P |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01381 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | 661P | 661P |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY REPORTED. DURING EVALUATION AT A THIRD PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADATION. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01422 | 2021-05-25 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP PRO BIFLEX | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | IN661S | IN661S |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING EVALUATION AT A THIRD PARTY SERVICE CENTER, THE SOUND ABATMENT FOAM WAS FOUND TO BE DEGRADATION. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01557 | 2021-05-25 | 2021-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE DEVICE WOULD NOT POWER ON. THE DEVICE'S SYSTEM BOARD, POWER MANAGEMENT BOARD AND FRONT PANEL/KEYPAD LED BOARD WERE REPLACED TO ADDRESS THE ISSUE. THE DEVICE HAD EVIDENCE OF INSECT CONTAMINATION. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01558 | 2021-05-25 | 2021-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054097 | 1054097 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS FAILING A TEST STEP. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR WAS FAILING A TEST STEP. THERE WAS NO HARM OR INJURY REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01545 | 2021-05-25 | 2021-05-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD DISPLAY WAS FOUND TO BE BLANK. THE DEVICE'S LCD DISPLAY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01544 | 2021-05-25 | 2021-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01543 | 2021-05-25 | 2021-04-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A LEAK FROM A VENTILATOR'S EXHALATION VALVE. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE ACTIVE EXHALATION CONTROL MODULE AND OVERHEAD BLOWER FILTER NEED REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01330 | 2021-05-25 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX400H11 | DSX400H11 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01541 | 2021-05-25 | 2021-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S BLOWER BELLOWS WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01439 | 2021-05-25 | 2021-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500H11C | DSX500H11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01547 | 2021-05-25 | 2021-05-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TO CHARGE ITS INTERNAL BATTERY. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S INTERNAL BATTERY NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01551 | 2021-05-25 | 2021-05-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S EXTERNAL FLOW SENSOR WAS NOT FUNCTIONING. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S FLOW SENSOR CABLE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01548 | 2021-05-25 | 2021-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01546 | 2021-05-25 | 2021-04-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, THE DEVICE FAILED TEST A STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01550 | 2021-05-25 | 2021-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01540 | 2021-05-25 | 2021-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1032804 | 1032804 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01549 | 2021-05-25 | 2021-05-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01444 | 2021-05-26 | 2021-05-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO BIFLEX | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | IN761S | IN761S |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY REPORTED. DURING EVALUATION AT A THIRD PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADATION. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01446 | 2021-05-26 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO AUTO W/BT | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | GBG502S15 | GBG502S15 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING EVALUATION AT A THIRD PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADATION. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01572 | 2021-05-26 | 2021-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | KR2110X15B | KR2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S TOUCHPAD WAS FAULTY. THE DEVICE WAS NOT IN PATIENT USE. THE MANUFACTURER IS CURRENTLY INVESTIGATING THIS ISSUE AND A FOLLOW-UP REPORT WILL BE FILED WHEN THE INVESTIGATION IS COMPLETE. | IT WAS INITIALLY REPORTED, THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S TOUCHPAD WAS FAULTY. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION AND THE TOUCHPAD WAS REPLACED. ONLY THE TOUCHPAD WAS RETURNED TO THE MANUFACTURER FOR INVESTIGATION. DURING THE MANUFACTURER'S INVESTIGATION IT WAS OBSERVED THE TOUCHPAD LOCK WAS ACTIVATED. THE LOCK WAS DE-ACTIVATED AND THE TOUCHPAD PASSED ALL TESTING AND OPERATED AS DESIGNED. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01526 | 2021-05-26 | 2021-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S POWER MANAGEMENT BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01525 | 2021-05-26 | 2021-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN OPEN CONDITION WAS FOUND WITHIN THE DEVICE'S POWER CORD. THE DEVICE'S POWER CORD WAS REPLACED TO ADDRESS THE ISSUE. A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S BLOWER MOTOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01527 | 2021-05-26 | 2021-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER HAD PREVIOUSLY REPORTED THAT THE INTERNAL BATTERY WAS REPLACED TO ADDRESS AN ERROR CODE RELATED TO THE BATTERY CHARGING. THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01528 | 2021-05-26 | 2021-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01531 | 2021-05-26 | 2021-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WOULD NOT PASS SERVICE TESTING. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE AND INTERNAL TUBING WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01538 | 2021-05-26 | 2021-04-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS DISPLAYING A "HIGH INTERNAL OXYGEN" ALARM. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN MANIFOLD WAS REPLACED AND WAS CALIBRATED TO ADDRESS THE ISSUE. | THE MANUFACTURER HAD PREVIOUSLY REPORTED THAT A VENTILATORS OXYGEN MANIFOLD WAS REPLACED AND THE DEVICE RECALIBRATED TO ADDRESS A "HIGH INTERNAL OXYGEN" ALARM. THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01415 | 2021-05-26 | 2021-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | PHILIPS RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S POWER MANAGEMENT BOARD, OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01447 | 2021-05-26 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO AUTO W/BT | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | GBG502S15 | GBG502S15 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING EVALUATION AT A THIRD PARTY SERVICE CENTER, THE SOUND ABATMENT FOAM WAS FOUND TO BE DEGRADATION. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01348 | 2021-05-26 | 2021-04-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AVAPS C SERIES | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS INC | 1161T | 1161T |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING EVALUATION AT A THIRD PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADATION. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01519 | 2021-05-26 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO SERIES ASSY | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | 751P | 751P |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING EVALUATION AT A THIRD PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADATION. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01353 | 2021-05-26 | 2021-04-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO, INTL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | 761P | 761P |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING EVALUATION AT A THIRD PARTY SERVICE CENTER, THE SOUND ABATMENT FOAM WAS FOUND TO BE DEGRADATION. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01380 | 2021-05-26 | 2021-04-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP S/T C SERIES | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS INC | IN1061X | IN1061X |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01530 | 2021-05-26 | 2021-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01517 | 2021-05-26 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING BLOWER ISSUES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01529 | 2021-05-26 | 2021-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01445 | 2021-05-26 | 2021-05-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO AUTO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | HKG502S15 | HKG502S15 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY REPORTED. DURING EVALUATION AT A THIRD PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADATION. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01518 | 2021-05-26 | 2021-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY REPORTED. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. SECTION B5 SHOULD BE PREVIOUSLY REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF BLACK SUBSTANCE IN WATER WELL AND INSIDE TUBE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. IT WAS ALSO ALLEGED THAT THE BLOWER IS MAKING LOUD NOISES. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION G1 UPDATED IN THIS REPORT. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01560 | 2021-05-26 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV 300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S TIDAL VOLUME DELIVERY WAS NOT WORKING. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S EXTERNAL FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01552 | 2021-05-26 | 2021-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED A TEST STEP DURING TESTING AND WAS ALARMING FOR A LOW INSPIRATORY PRESSURE ALARM CONDITION. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SENSOR BOARD TO PORTING BLOCK TUBING WAS FOUND TO BE DISCONNECTED. THE DEVICE'S TUBING WAS RECONNECTED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01555 | 2021-05-26 | 2021-05-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE RELATED TO THE REMOTE ALARM CONNECTOR WAS OBSERVED. THE DEVICE'S INTERFACE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01571 | 2021-05-26 | 2021-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS NOT PROVIDING FLOW. THERE WAS NO HARM OR INJURY REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR WAS NOT PROVIDING FLOW. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BLOWER MOTOR WAS REPLACED TO ADDRESS THE ISSUE. THERE WAS EVIDENCE OF CONTAMINATION. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01569 | 2021-05-26 | 2021-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S DISPLAY WAS BLANK. THERE WAS NO HARM OR INJURY REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR'S DISPLAY WAS BLANK. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER AND THE CUSTOMER'S COMPLAINT COULD NOT BE DUPLICATED. THE DEVICE WAS FOUND TO OPERATE AND ALARM AS DESIGNED. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01570 | 2021-05-26 | 2021-04-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SHUT DOWN WHILE IN USE. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED VENTILATOR'S SHUT DOWN WHILE IN USE. THERE WAS NO PATIENT HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, THE CUSTOMER'S COMPLAINT COULD NOT BE CONFIRMED. THE DEVICE WAS FOUND TO BE CONTAMINATED WITH PET HAIR. NO TESTING WAS ABLE TO BE PERFORMED DUE TO THE CONTAMINATION. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01568 | 2021-05-26 | 2021-04-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S WAS FAILING A TEST STEP. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR'S WAS FAILING A TEST STEP. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS EVALUATED BY A THIRD PARTY FIELD SERVICE ENGINEER AND THE COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND SENSOR BOARD NEED REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01576 | 2021-05-26 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01561 | 2021-05-26 | 2021-05-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR A HIGH INSPIRATORY PRESSURE ALARM CONDITION. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. DURING THE EVALUATION THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S SENSOR BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01573 | 2021-05-26 | 2021-05-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500T11C | DSX500T11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01559 | 2021-05-26 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV 300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S TIDAL VOLUME DELIVERY WAS NOT WORKING. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S EXTERNAL FLOW SENSOR AND FLOW SENSOR CABLE WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01553 | 2021-05-26 | 2021-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE WAS FOUND TO BE INFESTED WITH INSECTS AND WAS RETURNED TO THE CUSTOMER UNREPAIRED. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01554 | 2021-05-26 | 2021-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE TO POWER ON WAS OBSERVED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01567 | 2021-05-26 | 2021-04-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2200X11B | DS2200X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S TOUCHSCREEN WAS UNRESPONSIVE. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR'S TOUCHSCREEN WAS UNRESPONSIVE. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS EVALUATED BY A THIRD PARTY FIELD SERVICE ENGINEER AND THE CUSTOMER'S COMPLAINT COULD NOT BE CONFIRMED. THE DEVICE OPERATED AND ALARMED AS DESIGNED. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01595 | 2021-05-27 | 2021-04-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SET POSITIVE END EXPIRATORY PRESSURE INCREASED WITHOUT MANUAL ADJUSTMENT. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR'S SET POSITIVE END EXPIRATORY PRESSURE INCREASED WITHOUT MANUAL ADJUSTMENT. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE VENTILATOR'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01578 | 2021-05-27 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AVAPS C SERIES, INTERNATIONAL 30 | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC | IN1161X | IN1161X |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING EVALUATION AT A THIRD PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADATION. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01596 | 2021-05-27 | 2021-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01594 | 2021-05-27 | 2021-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2110X15B | IN2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SET PRESSURE INCREASED WITHOUT MANUAL ADJUSTMENT. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR'S SET PRESSURE INCREASED WITHOUT MANUAL ADJUSTMENT. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE VENTILATOR'S SOFTWARE WAS RE-LOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01564 | 2021-05-27 | 2021-05-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01592 | 2021-05-27 | 2021-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054097 | 1054097 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "VENTILATOR SERVICE REQUIRED" ALARM RELATED TO A VENTILATOR'S MALFUNCTION OF A VENTILATOR'S OXYGEN FLOW SENSOR. THERE WAS NO PATIENT HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON MAR 13, 2024 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A "VENTILATOR SERVICE REQUIRED" ALARM RELATED TO A VENTILATOR'S MALFUNCTION OF A VENTILATOR'S OXYGEN FLOW SENSOR. THERE WAS NO PATIENT HARM OR INJURY REPORTED. THE DEVICE WAS RETURNED TO THE THIRD PARTY SERVICE CENTER FOR FURTHER EVALUATION. TTHE DEVICE WAS EVALUATED. DURING THE EVALUATION, DUST WAS OBSERVED IN THE BLOWER BOX. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED. THERE WERE NO ERRORS FOUND. THE THIRD PARTY SERVICE CENTER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION. DEVICE WAS CORRECTED. SECTIONS D8, D9, H2, H3 AND H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01591 | 2021-05-27 | 2021-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A MALFUNCTION OF A VENTILATOR'S TOUCHSCREEN. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A MALFUNCTION OF A VENTILATOR'S TOUCHSCREEN. THE VENTILATOR WAS RETURNED OT THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S DISPLAY BOARD WAS REPALCED TO ADDRESS THE ISSUE. THE DISPLAY WAS ALSO REPLACED TO ADDRESS THE ISSUE. CODING FOR THE BOARD WAS INADVERTENLY OMITTED ON THE PREVIOUS SUBMITTED REPORT. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01566 | 2021-05-27 | 2021-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2100X15B | IN2100X15B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "VENTILATOR INOPERATIVE" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01577 | 2021-05-27 | 2021-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500H11C | DSX500H11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01593 | 2021-05-27 | 2021-05-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | EE2110X15B | EE2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. THERE WAS EVIDENCE OF CONTAMINATION. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01565 | 2021-05-27 | 2021-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01563 | 2021-05-27 | 2021-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE'S EXTERNAL FLOW SENSOR WAS FOUND TO BE DEFECTIVE. THE DEVICE'S EXTERNAL FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01562 | 2021-05-27 | 2021-05-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED A TEST STEP DURING TESTING. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01608 | 2021-05-28 | 2021-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500S11 | DSX500S11 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01603 | 2021-05-28 | 2021-05-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S NAVIGATIONAL BUTTON WAS OBSERVED. THE DEVICE'S FRONT PANEL/KEYPAD LED BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01612 | 2021-05-28 | 2021-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | GBG502S15 | GBG502S15 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY REPORTED. DURING EVALUATION AT A THIRD PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADATION. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01614 | 2021-05-28 | 2021-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO AUTO W/BT | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | HKG502S15 | HKG502S15 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING EVALUATION AT A THIRD PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADED. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01613 | 2021-05-28 | 2021-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO AUTO W/BT | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | HKG502S15 | HKG502S15 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING EVALUATION AT A THIRD PARTY SERVICE CENTER, THE SOUND ABATMENT FOAM WAS FOUND TO BE DEGRADED. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A DEVICE'S THE SOUND ABATMENT FOAM WAS FOUND TO BE DEGRADED. THE INCORRECT SERIAL NUMBER WAS REPORTED IN THE PREVIOUS REPORT. THE CORRECT SERIAL NUMBER IS (B)(6) . | N | Y | LAY USER/PATIENT |  | 2021-05
2518422-2021-01600 | 2021-05-28 | 2021-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01601 | 2021-05-28 | 2021-05-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01556 | 2021-05-28 | 2021-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01366 | 2021-05-28 | 2021-05-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TO READ THE RETURN VOLUME. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. THE DEVICE WAS FOUND TO BE CONTAMINATED WITH DIRT AND DUST. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01580 | 2021-05-28 | 2021-05-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01602 | 2021-05-28 | 2021-05-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUES. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR'S SENSOR BOARD WAS REPLACED TO ADDRESS FAILED TEST STEPS DURING TESTING. THE SENSOR BOARD WAS NOT REPLACED. THE DEVICE WAS RETURNED TO THE CUSTOMER UNREPAIRED. | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01598 | 2021-05-28 | 2021-04-27 | I, n, j, u, r, y | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR STOPPED WORKING WHEN A BACTERIA FILTER WAS CHANGED. THE PATIENT'S BLOOD OXYGEN SATURATION DECREASED AND THE PATIENT HAD TO BE SWITCHED TO A BACK UP VENTILATOR. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE VENTILATOR'S DOWNLOADED EVENT LOGS WERE REVIEWED BY THE MANUFACTURER. THE DEVICE MAINTAINS THE EVENT LOGS IN NON-VOLATILE MEMORY. AN EVENT, AS LOGGED BY THE DEVICE, INCLUDES EVERY KEYPRESS, ALARM OR FAILURE OF THE DEVICE TO REMAIN WITHIN SPECIFICATIONS. THERE WERE NO EVENTS LOGGED RELATED TO DEVICE PERFORMANCE OR DEVICE MALFUNCTIONS. IF A DEVICE MALFUNCTION OCCURRED, THE EVENT LOGS WOULD HAVE CONTAINED ENTRIES RELATED TO THE MALFUNCTION. THE DEVICE PASSED ALL TESTING. THE MANUFACTURER CONCLUDES THE VENTILATOR OPERATED AND ALARMED AS DESIGNED AND DID NOT CAUSE OR CONTRIBUTE TO THE REPORTED EVENT. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR STOPPED WORKING WHEN A BACTERIA FILTER WAS EXCHANGED. DURING FINAL DEVICE TESTING, THE DEVICE FAILED A TEST STEP. ALTHOUGH THE DEVICE FAILED A TEST STEP, THE CUSTOMERS ALLEGATION OF THE DEVICE NOT WORKING WAS NOT DUPLICATED. THE DEVICE WAS FOUND TO OPERATE AND ALARM TO DESIGN SPECIFICATIONS DURING TESTING. THE FAILED TEST STEP WOULD NOT HAVE CAUSED OR CONTRIBUTED TO THE NOTED EVENT. THE DEVICE NEEDS RECALIBRATED TO ADDRESS THE ISSUE | Y | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01599 | 2021-05-28 | 2021-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR A HIGH PRIORITY ALARM CONDITION. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. A "VENTILATOR INOPERATIVE ALARM CONDITION WAS OBSERVED. THE DEVICE'S BLOWER ASSEMBLY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01579 | 2021-05-28 | 2021-05-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-05
2518422-2021-01533 | 2021-06-01 | 2021-05-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DISPLAY SCREEN WAS BLANK DUE TO PHYSICAL DAMAGE. THE LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TO POWER ON. THE DEVICE'S POWER MANAGEMENT BOARD WAS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE'S NURSE CALL CONNECTOR WAS FOUND TO BE BROKEN. THE DEVICE'S INTERFACE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01534 | 2021-06-01 | 2021-05-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S POWER MANAGEMENT BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | I |  | 2021-06
2518422-2021-01606 | 2021-06-01 | 2021-05-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERFACE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01607 | 2021-06-01 | 2021-05-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR WAS RETURNED FOR PREVENTATIVE MAINTENANCE AND DUING TESTING, THE DEVICE FAILED A TEST STEP. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. ON THE PREVIOUS REPORT, SECTION D4, THE SERIAL NUMBER WAS INOCRRECTLY PLACED FOR THE UDI NUMBER. THE CORRECT SERIAL NUMBER AND UDI NUMBER ARE PLACED IN THIS REPORT. SECTION G2: "FOREIGN" WAS NOT CHECKED. SECTION H4: THE WRONG MANUFACTURE DATE WAS ENTERED. THE CORRECT DATE SHOULD BE 11/21/2014. ALL CORRECTIONS ARE REFLECTED IN THIS REPORT. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01620 | 2021-06-01 | 2021-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01615 | 2021-06-01 | 2021-05-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01624 | 2021-06-01 | 2021-05-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01622 | 2021-06-01 | 2021-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01617 | 2021-06-01 | 2021-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01621 | 2021-06-01 | 2021-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01618 | 2021-06-01 | 2021-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01619 | 2021-06-01 | 2021-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01535 | 2021-06-01 | 2021-05-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MANIFOLD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01604 | 2021-06-01 | 2021-05-19 | I, n, j, u, r, y | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR "HIGH INSPIRATORY PRESSURE". THE PATIENT'S BLOOD OXYGEN SATURATION DECREASED AND THE PATIENT HAD TO BE SWITCHED TO A BACK UP VENTILATOR. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. THE DEVICE PERFORMED TO MANUFACTURER'S SPECIFICATIONS. |  | Y | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01623 | 2021-06-01 | 2021-05-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S INTERNAL BATTERY WAS FAULTY. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01616 | 2021-06-01 | 2021-05-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICES LCD DISPLAY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01605 | 2021-06-01 | 2021-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE AND OVERHEAD BLOWER FILTER WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01635 | 2021-06-02 | 2021-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT HAD DIFFICULTY BREATHING AND THE DEVICE FAILED TO TRIGGER SPONTANEOUS BREATHS. THE PATIENT WAS PLACED ON AN ALTERNATE DEVICE. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 18, 2024, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING HAD DIFFICULTY BREATHING AND THE DEVICE FAILED TO TRIGGER SPONTANEOUS BREATHS. THE PATIENT WAS PLACED ON AN ALTERNATE DEVICE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT COULD NOT BE CONFIRMED OR DUPLICATED RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTIONS B1, B2, H1 AND H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: B1 WAS CORRECTED FOR ONLY THE PRODUCT PROBLEM (BOTH ADVERSE EVENTS AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. | Y | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01628 | 2021-06-02 | 2021-05-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S TOUCHSCREEN WAS FROZEN. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION THAT A VENTILATOR'S TOUCHSCREEN WAS FROZEN. A FOLLOW REPORT WAS SENT TO THE FDA ON 06/02/2021 THAT FAILED TO INCLUDE THE MANUFACTURER'S FINDINGS. THE CODING WAS INCLUDED IN THE REPORT. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION, AND THE CUSTOMER'S COMPLAINT COULD NOT BE DUPLICATED. NO MALFUNCTION OF THE TOUCHSCREEN WAS OBSERVED. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01589 | 2021-06-02 | 2021-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01587 | 2021-06-02 | 2021-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A FAILURE OF THE DEVICE TO CHARGE ITS INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01588 | 2021-06-02 | 2021-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A FAILURE OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01590 | 2021-06-02 | 2021-05-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01629 | 2021-06-02 | 2021-05-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A "VENTILATOR INOPERATIVE" ALARM CONDITION OCCURRED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. THE DEVICE WAS FOUND TO BE CONTAMINATED WITH WHITE POWDER FROM AN EXTERNAL SOURCE. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01630 | 2021-06-02 | 2021-05-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. THE DEVICE HAD EVIDENCE OF TOBACCO CONTAMINATION. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
MW5101667 | 2021-06-02 | 2019-01-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. |  |  |  | I HAVE HAD ONGOING HEADACHES SINCE USING THE DREAMSTATION AUTOCPAP. I ALSO HAVE SIGNIFICANT NASAL CONGESTION AFTER USE NIGHTLY AND CHRONIC RUNNY NOSE. SORE THROAT ON AND OFF. I HAVE BEEN BLAMING OTHER CAUSES UP UNTIL RECENTLY. I HAVE NOW COME TO THE CONCLUSION THAT THIS ALL STARTED WHEN I STARTED CPAP. IT IS THIS DEVICE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01610 | 2021-06-02 | 2021-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | PHILIPS RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE, PORTING BLOCK AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01636 | 2021-06-02 | 2021-05-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2200X11B | DS2200X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS NOT MEASURING OXYGEN LEVELS. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS EVALUATED BY A THIRD PARTY FIELD SERVICE ENGINEER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01583 | 2021-06-02 | 2021-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01585 | 2021-06-02 | 2021-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, ERROR CODES RELATED TO THE CHARGING OF THE INTERNAL BATTERY WERE OBSERVED. THE DEVICE'S INTERNAL BATTERY NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01582 | 2021-06-02 | 2021-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01574 | 2021-06-02 | 2021-05-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS DISPLAYING BAD OXYGEN READINGS. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01581 | 2021-06-02 | 2021-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, ERROR CODES RELATED TO THE CHARGING OF THE INTERNAL BATTERY WERE OBSERVED. THE DEVICE'S DETACHABLE BATTERY CABLE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01586 | 2021-06-02 | 2021-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01575 | 2021-06-02 | 2021-05-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR SHUT OFF DURING USE. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | SECTION B5 SHOULD BE PREVIOUSLY REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR SHUT OFF DURING USE AND BLUE SCREEN APPEARS WHEN IT REBOOTS. THERE WAS NO HARM OR INJURY REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01584 | 2021-06-02 | 2021-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054097 | 1054097 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01633 | 2021-06-02 | 2021-05-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SCREEN WAS BLANK AND UNABLE TO BE VIEWED. THERE WAS NO HARM OR INJURY REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01643 | 2021-06-03 | 2021-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE INTERNAL PORTING BLOCK ADAPTOR AND PASSIVE EXHALATION PORTING BLOCK WERE FOUND TO BE MISSING. THE COMPONENTS WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01647 | 2021-06-03 | 2021-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01632 | 2021-06-03 | 2021-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED THE DATE OF EVENT (B3) AS (B)(6) 2012 INCORRECTLY. THE CORRECT DATE OF EVENT (B3) SHOULD READ (B)(6) 2021. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01536 | 2021-06-03 | 2021-05-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01631 | 2021-06-03 | 2021-05-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01644 | 2021-06-03 | 2021-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01645 | 2021-06-03 | 2021-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERFACE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01642 | 2021-06-03 | 2021-05-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040005 | 1040005 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, THE LCD WAS FOUND TO BE BLANK. THE DEVICE'S LCD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01537 | 2021-06-03 | 2021-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01638 | 2021-06-03 | 2021-05-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLACK. THERE WAS NO HARM OR INJURY REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01639 | 2021-06-03 | 2021-05-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SET PRESSURE INCREASED WITHOUT MANUAL ADJUSTMENT. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER HAD PREVIOUSLY REPORTED AN ALLEGATION THAT A VENTILATOR'S SET PRESSURE INCREASED WITHOUT MANUAL ADJUSTMENT. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE VENTILATOR'S SOFTWARE WAS RE-LOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01650 | 2021-06-03 | 2021-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO THE DEVICE'S SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01641 | 2021-06-03 | 2021-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AVAPS C SERIES INTL 30, FLOW GEN | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | 1161T | 1161T |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING EVALUATION AT A THIRD PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADED. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01640 | 2021-06-03 | 2021-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SET PRESSURE INCREASED WITHOUT MANUAL ADJUSTMENT. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01521 | 2021-06-04 | 2021-04-07 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION FROM A DURABLE MEDICAL EQUIPMENT (DME) SUPPLIER INQUIRING ABOUT THE DEVICE'S SOUND ABATEMENT FOAM. AN END USER STATED SHE HAS EXPERIENCED HEADACHES, A SORE THROAT AND WAS RECENTLY TREATED WITH ANTIBIOTICS. THE MANUFACTURER REQUESTED THE RETURN OF THE DEVICE AND ALL ACCESSORIES FOR EVALUATION. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THE END USER INDICATED OZONE OR OTHER UNAPPROVED CLEANING AND DISINFECTION METHOD WAS USED. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01652 | 2021-06-04 | 2021-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AVAPS C SERIES INTL 30, FLOW GEN | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | 1161T | 1161T |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO PATIENT HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, AN ISSUE RELATED TO SOUND ABATEMENT FOAM DEGRADATION WAS OBSERVED. THE DEVICE'S SOUND ABATEMENT FOAM WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01487 | 2021-06-04 | 2021-04-09 | D, e, a, t, h | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION OCCURRED ON (B)(6) 2021 WHILE ON A PATIENT AND THE PATIENT'S OXYGEN SATURATION DROPPED. THE PATIENT WAS THEN PLACED ON A DIFFERENT DEVICE. IT WAS REPORTED THAT ON (B)(6) 2021 AT 6:50 A.M. EST THE PATIENT AGAIN DECOMPENSATED AND PASSED AWAY AT 7:02 A.M. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PERVIOUSLY RECEIVED INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION OCCURRED ON (B)(6) 2021 WHILE ON A PATIENT AND THE PATIENT'S OXYGEN SATURATION DROPPED. THE PATIENT WAS THEN PLACED ON A DIFFERENT DEVICE. IT WAS REPORTED THAT ON (B)(6) 2021 AT 6:50 A.M. EST THE PATIENT AGAIN DECOMPENSATED AND PASSED AWAY AT 7:02 A.M. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE INTERNAL AND EXTERNAL PART OF THE DEVICE. THE MANUFACTURER FOUND THE E-349 ERROR FROM THE DEVICE AFTER FURTHER INVESTIGATION. THE MANUFACTURER CONCLUDES THEY COULD CONFIRM THE CUSTOMER ALLEGATION BUT THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS OBSERVED IN THIS DEVICE. | Y | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01648 | 2021-06-04 | 2021-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01649 | 2021-06-04 | 2021-05-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01691 | 2021-06-07 | 2021-06-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED STEPS DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01686 | 2021-06-07 | 2021-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01689 | 2021-06-07 | 2021-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01688 | 2021-06-07 | 2021-05-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 054655 | 054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICE'S LCD SCREEN NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01690 | 2021-06-07 | 2021-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED STEPS DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01684 | 2021-06-07 | 2021-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01658 | 2021-06-07 | 2021-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01687 | 2021-06-07 | 2021-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. AN ISSUE RELATED TO THE DEVICE'S FLOW SENSOR WAS OBSERVED. THE DEVICE'S FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01685 | 2021-06-07 | 2021-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, AN ISSUE RELATED TO THE CHARGING OF THE DEVICE'S INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01682 | 2021-06-07 | 2021-05-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR A "SERVICE REQUIRED" CONDITION RELATED TO OXYGEN REGULATION. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE WAS EVALUATED BY A FIELD SERVICE ENGINEER, AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN BLENDING MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01670 | 2021-06-07 | 2021-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01660 | 2021-06-07 | 2021-05-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01666 | 2021-06-07 | 2021-05-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01653 | 2021-06-07 | 2021-05-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WOULD NOT POWER ON. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. NO REPAIR WAS PERFORMED DUE TO INSECT INFESTATION. THE DEVICE WAS RETURNED TO THE CUSTOMER UNREPAIRED. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01657 | 2021-06-07 | 2021-06-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" ALARM CONDITION INDICATING AN ISSUE WITH THE DEVICE'S TOUCHSCREEN WAS OBSERVED. THE DEVICE'S RIBBON CABLE WAS FOUND TO BE SEATED IMPROPERLY. THE RIBBON CABLE WAS RESEATED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01659 | 2021-06-07 | 2021-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040005 | 1040005 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED STEPS DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01655 | 2021-06-07 | 2021-05-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLANK. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS EVALUATED AT THE MANUFACTURER'S SERVICE CENTER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01656 | 2021-06-07 | 2021-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01667 | 2021-06-07 | 2021-05-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMNG FOR A LOW CIRCUIT LEAK ALARM CONDITION. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S SENSOR BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01668 | 2021-06-07 | 2021-05-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S TIDAL VOLUME DELIVERY WAS INACCURATE AND THE DEVICE FAILED TEST STEPS DURING TESTING. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUES. | THE MANUFACTURER PREVIOUSLY REPORTED THAT DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S TIDAL VOLUME DELIVERY WAS INACCURATE AND THE DEVICE FAILED TEST STEPS DURING TESTING. THE ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. DURING FURTHER EVALUATION THE OVERHEAD BLOWER FILTER WAS FOUND TO BE CLOGGED WITH CONTAMINATION. THE DEVICE'S OVERHEAD BLOWER FILTER WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01671 | 2021-06-07 | 2021-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01673 | 2021-06-07 | 2021-06-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01672 | 2021-06-07 | 2021-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01669 | 2021-06-07 | 2021-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01696 | 2021-06-08 | 2021-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD NEED REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
MW5101772 | 2021-06-08 | 2020-10-30 | I, n, j, u, r, y | Voluntary report | DREAMSTATION1 CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS INC. | DSX500H11C | DSX500H11C |  | I WAS EXPERIENCING SWOLLEN LYMPH GLANDS IN MY NECK (NECK FELT THICK), PRESSURE IN EARS, RINGING IN EARS, REDUCTION IN NECK MOTION, HEADACHE, AND EVENTUALLY HIGH BLOOD PRESSURE. I FIRST MADE AN APPT. WITH MY PA/PHYSICIAN GROUP. SHE RAN BLOOD WORK AND TOOK AN X-RAY, BUT NOTHING OUTSTANDING WAS FOUND. NEXT CAME AN ULTRASOUND OF MY NECK. NO SIGNIFICANT FINDINGS, EXCEPT SMALL THYROID NODULES. NEXT WAS A CAT SCAN WITH CONTRAST. AGAIN NOTHING OUTSTANDING WAS FOUND. NOW WE WERE AT (B)(6) AND I AM WORSENING, FEELING MISERABLE, AND CANNOT SLEEP. I DROVE MYSELF TO THE ER QUICKCARE IN (B)(6), AND NEXT STOP WOULD BE THE ER IF NO ASSISTANCE. I WAS EXAMINED AND IT WAS NOTED I HAD "SYMPTOMS WITH AN UNCERTAIN CAUSE", AND WAS PUT ON A STEROID (MEDROL) AND Z-PACK (AZITHROMYCIN). MY BP WAS VERY HIGH THAT DAY. HIGHEST WAS 170/99. I WAS VERY AGITATED AS THIS HAD BEEN GOING ON FOR 2 MONTHS. THE MEDICINE STARTED TO WORK WITHIN 24-48 HOURS AND MY NECK STARTED FEELING BETTER. MY NEXT APPOINTMENT WAS WITH AN ENT ON (B)(6) DR. (B)(6) LOOKED ME OVER THOROUGHLY AND NOTED MY SYMPTOMS WERE GETTING BETTER WITH TREATMENT. MY EAR PRESSURES WERE MEASURED AND WERE BOTH HIGHER THAN NORMAL. HE CONCLUDED, "I WAS ALLERGIC TO SOMETHING". I ASKED HIM WHAT IT COULD BE, SINCE I HAVE HAD VERY FEW ALLERGIC REACTIONS OR ALLERGIES THROUGHOUT MY LIFE (DOB (B)(6) HE WASN'T SURE, BUT PUT ME ON A REGIMEN OF FLONASE AND ZYRTEC DAILY. MY SYMPTOMS IMPROVED, BUT THE RINGING IN MY EARS REMAINS AND MY GLANDS SEEMED TO STAY SWOLLEN AND SENSITIVE TO THE TOUCH. A FOLLOW UP VISIT TO ENT SHOWED HIGH PRESSURES IN MY EARS REMAINED. I USE A PHILIPS RESPIRONICS DREAMSTATION CPAP EVERY NIGHT (SINCE (B)(6) 2018). I TOLD THIS TO MY ENT DOCTOR, BUT NO CORRELATION WAS MENTIONED. NOT UNTIL LATE (B)(6) 2021, WHEN I HEARD OF A PRODUCT HOLD ON THIS CPAP PRODUCT, DID I PUT TWO AND TWO TOGETHER TO HIGHLY SPECULATE THAT MY "ALLERGY" COULD BE THE DEGRADED FOAM BEING INVESTIGATED BY PHILIPS AND THE FDA. I HAVE BEEN WAITING FOR FURTHER INFORMATION AND CLARIFICATION ON RESEARCH, BUT TO DATE I HAVE NO MORE INFORMATION. I TOOK MYSELF OFF THE CPAP IMMEDIATELY AND REQUESTED A DIFFERENT UNIT FROM APRIA. NO ONE HAS BEEN ABLE TO GET ME A NEW DEVICE, SO I HAVE BEEN USING A NEW CPAP GO (PHILIPS RESPIRONICS ALSO) THAT I HAD BOUGHT, BUT HADN'T USED YET. FROM THE INFORMATION I HAVE FROM THE PHILIPS WEBSITE, THE CAUSE FOR THE DEGRADING FOAM COULD BE OZONE CLEANING, HIGH HUMIDITY AND/OR AGE. MY UNIT IS 3.5 YEARS OLD, I LIVE IN FL AND I USED OZONE CLEANING. I WOULD LIKE TO GET MY UNIT TESTED, BUT HAVEN'T FOUND A RELIABLE COMPANY YET. I WOULD LIKE TO FIND OUT IF THIS DEVICE HAS CAUSED MY REACTION AND WHAT LONG TERM AFFECTS IT COULD CAUSE. YOUR INPUT WOULD BE GREATLY APPRECIATED. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01677 | 2021-06-08 | 2021-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01661 | 2021-06-08 | 2021-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01663 | 2021-06-08 | 2021-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUES. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S BLOWER BELLOWS WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01662 | 2021-06-08 | 2021-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S SENSOR BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01681 | 2021-06-08 | 2021-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01692 | 2021-06-08 | 2021-06-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S POWER MANAGEMENT BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01678 | 2021-06-08 | 2021-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | RIN2110X15B | RIN2110X15B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01694 | 2021-06-08 | 2021-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | REMSTAR AUTO AFLEX ASSY | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | IN551HS | IN551HS |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK PARTICLES WERE OBSERVED DURING SERVICING OF THE DEVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR EVALUATION. THE MANUFACTURER FOUND THE BLOWER TO BE EMITTING NOISE AND ERRORS LOGGED RELATED TO THE MAIN PCA. THE MANUFACTURER FOUND THE AIR INLET FILTERS AND AIRWAY TO BE CONTAMINATED WITH DIRT. THE MANUFACTURER FOUND THE DEVICE'S SOUND ABATEMENT FOAM TO HAVE DEGRADED. THE BLOWER, MAIN PCA AND THE SOUND ABATEMENT FOAM WILL NEED TO BE REPLACED TO ADDRESS THE ISSUES. THE DEVICE WAS RETURNED UNREPAIRED BY THE REQUEST OF THE CUSTOMER. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01693 | 2021-06-08 | 2021-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01676 | 2021-06-08 | 2021-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1032804 | 1032804 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLANK. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE WAS EVALUATED AT THE MANUFACTURER'S SERVICE CENTER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S LCD SCREEN NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01675 | 2021-06-08 | 2021-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | RIN2110X15B | RIN2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A MALFUNCTION OF A VENTILATOR'S TOUCHSCREEN. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S DISPLAY WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A FAILURE OF A VENTILATOR'S TOUCHSCREEN AND THE DEVICE'S DISPLAY WAS REPLACED TO ADDRESS THE ISSUE. AN INCORRECT DATE WAS ENTERED FOR B.4 OF 28-JUN-2021. THE CORRECT DATE FOR B.4 IS 08-JUN-2021. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01654 | 2021-06-08 | 2021-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE'S FLOW SENSOR ASSEMBLY WAS FOUND TO BE BLOCKED WITH CONTAMINATION. THE FLOW SENSOR ASSEMBLY NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01680 | 2021-06-08 | 2021-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01699 | 2021-06-09 | 2021-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500T11C | DSX500T11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING EVALUATION, AN ISSUE RELATED TO THE SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01664 | 2021-06-09 | 2021-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040004 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01705 | 2021-06-09 | 2021-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S OXYGEN BLENGING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01674 | 2021-06-09 | 2021-05-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040004 | 1040004 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING AND THE OXYGEN LEVEL COULD NOT BE ADJUSTED. THE PATIENT'S BLOOD OXYGEN SATURATION DECREASED AND THE PATIENT WAS SWITCHED TO ANOTHER VENTILATOR. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUE. | SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH THE PRODUCT PROBLEM AND ADVERSE EVENT WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK . (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MAL FUNCTION. SECTION H6 HEALTH EFFECTS: IMPACT CODE WAS UPDATED. | Y | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01683 | 2021-06-09 | 2021-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR A "SERVICE REQUIRED" CONDITION. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE WAS EVALUATED BY A THIRD-PARTY, AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SOLENOID VALVE NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01698 | 2021-06-09 | 2021-05-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S FLOW SENSOR CABLE WAS OBSERVED. THE DEVICE'S FLOW SENSOR CABLE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01707 | 2021-06-10 | 2021-05-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S FLOW SENSOR FAILED. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE WAS EVALUATED BY A THIRD-PARTY SERVICE CENTER, AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S FLOW SENSOR NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01701 | 2021-06-10 | 2021-06-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01700 | 2021-06-10 | 2021-06-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE NURSE CALL CONNECTOR (BROKEN) WAS OBSERVED. THE DEVICE'S INTERFACE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01702 | 2021-06-10 | 2021-06-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERFACE BOARD WAS REPLACED TO ADDRESS THE ISSUE. | IT WAS INITIALLY REPORTED, A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERFACE BOARD WAS REPLACED TO ADDRESS THE ISSUE. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS ALSO REPLACED TO ADDRESS THE ISSUE. THE FAULT OF THE OXYGEN BLENDING MODULE BOARD WAS OBSERVED DURING FUTHER INVESTIGATION OF THE COMPONENT IN THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01626 | 2021-06-10 | 2021-05-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A VENTILATOR . THERE WAS NO REPORT OF PATIENT HARM OR INJURY. AFTER THE FINAL ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, CUSTOMER DECLINED TO RESPOND TO THE GFE RELATED QUESTIONS AND TERMINATED THE CALL. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01708 | 2021-06-10 | 2021-05-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP AUTOSV H/C, DS | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | DSX900H11C | DSX900H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL PART OF THE DEVICE AND FOUND NO EVIDENCE OF THERMAL DAMAGE OR FAILURE. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND FOUND NO EVIDENCE OF THERMAL DAMAGE. THE MANUFACTURER DID OBSERVE EVIDENCE OF CONTAMINANTS PRESENT IN THE DREAMSTATION/HUMIDIFIER. THE MANUFACTURER FOUND THERE ARE BLACK SPECKS OF THE FOAM MATERIAL IN THE FRESH AIR INTAKE PORT OF THE DREAMSTATION, ON THE OUTLET PORT OF THE HUMIDIFIER, THE HUMIDIFIERS WATER CHAMBER HAS A FILM COATING CONTAMINATE WITH A SWEET FRAGRANT ODOR WITH THE BLACK CONTAMINATE (FOAM MATERIAL) IN A FEW SPOTS. ON THE INSIDE OF THE DREAMSTATION, FOUND THE FOAM MATERIAL IS MISSING IN A PORTION OF THE SOUND ABATEMENTS ENCLOSURE, THE BLOWER MOTOR/BOX HAS THE OILY DEGRADED FOAM MATERIAL THROUGHOUT, ALSO IN THE BLOWER BOXES OUTLET PORT AND IN THE DREAMSTATIONS REAR PANELS OUTLET PORT. THE AIR PATH OF THE DEVICE WAS COMPROMISED AND THE CONTAMINATE WOULD HAVE BEEN THROUGHOUT THE DEVICES AIR PATH. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND TO CONTAIN 9 ERROR CODES. THE MANUFACTURER VERIFIED THE UNITS DO POWER-UP, PROVIDE FLOW. THE MANUFACTURER DID OBSERVE BLOWER MOTOR NOISE. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION AND THE ODORS ARE LIKELY RELATED TO THE ENVIRONMENTAL CONDITIONS THE UNITS WERE BEING USED IN, WHICH RESULTED IN THE SOUND ABATEMENT FOAM TO DEGRADE. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01703 | 2021-06-11 | 2021-06-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SENSOR BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. THE DEVICE WAS RETURNED UNRAIRED AS PER THE CUSTOMER. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01704 | 2021-06-11 | 2021-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S POWER MANAGEMENT BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01709 | 2021-06-11 | 2021-05-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2200X11B | DS2200X11B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, A "VENTILATOR INOPERATIVE" ALARM CONDITION WAS OBSERVED AND THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01715 | 2021-06-11 | 2021-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR A SERVICE REQUIRED CONDITION RELATED TO OXYGEN REGULATION. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION, AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN BLENDING MODULE ASSEMBLY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01712 | 2021-06-11 | 2021-06-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01711 | 2021-06-11 | 2021-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01710 | 2021-06-11 | 2021-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. THE DEVICE HAD EVIDENCE OF WATER INGRESS. | THE MANUFACTURER PREVIOUSLY REPORTED THE MACHINE FLOW SENSOR NEEDED TO BE REPLACED TO ADDRESS A VENTILATOR INOPERATIVE ALARM CONDITION. THE DEVICE'S MACHINE FLOW SENSOR WAS REPLACED AND ALSO THE SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01714 | 2021-06-11 | 2021-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040000 | 1040000 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR A SERVICE REQUIRED CONDITION RELATED TO THE DEVICE'S OXYGEN BLENDING MODULE. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION, AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER HAD PREVIOUSLY REPORTED AN ALLEGATION THAT THE DEVICE WAS DISPLAYING AN ALARM RELATED TO THE OXYGEN BLENING MODULE BOARD AND THAT THE INTERFACE BOARD AND OXYGEN BLENDING MODULE BOARD WERE REPLACED TO ADDRESS THE ISSUE. THIS IS INCORRECT. THE INTERFACE BOARD WAS NOT REPLACED. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01751 | 2021-06-14 | 2021-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SENSOR BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. THERE WAS EVIDENCE OF CONTAMINATION. THE DEVICE WAS RETURNED UNREPAIRED PER CUSTOMER'S REQUEST. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01750 | 2021-06-14 | 2021-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01749 | 2021-06-14 | 2021-06-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01747 | 2021-06-14 | 2021-06-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01746 | 2021-06-14 | 2021-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1032804 | 1032804 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01745 | 2021-06-14 | 2021-06-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01679 | 2021-06-14 | 2021-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION, AN ISSUE RELATED TO THE UNIVERSAL PORTING BLOCK WAS OBSERVED DUE TO PHYSICAL DAMAGE. THE UNIVERSAL PORTING BLOCK WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01733 | 2021-06-14 | 2021-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED STEPS DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01735 | 2021-06-14 | 2021-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED EVENT LOG. THE DEVICE'S INTERNAL BATTERY NEEDS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01734 | 2021-06-14 | 2021-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED EVENT LOG. THE DEVICE'S INTERNAL BATTERY NEEDS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01732 | 2021-06-14 | 2021-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01731 | 2021-06-14 | 2021-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED EVENT LOG. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01716 | 2021-06-14 | 2021-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED EVENT LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01637 | 2021-06-14 | 2021-05-21 | I, n, j, u, r, y | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR PATIENT CIRCUIT DISCONNECT CONDITIONS AND THE PATIENT WAS TAKEN TO A HOSPITAL. ACCORDING TO THE REPORTER OF THE EVENT, THE VENTILATOR WAS ALARMING FOR PATIENT DISCONNECT CONDITIONS ON (B)(6) 2021. THE PATIENT WAS MANUALLY VENTILATED AND TAKEN TO A HOSPITAL. THE PATIENT REMAINED IN THE HOSPITAL FOLLOWING THIS EVENT, BUT THE REPORTER OF THE EVENT CONFIRMED IT WAS NOT DUE TO THIS EVENT. THE PATIENT HAS A POOR PROGNOSIS WITH ADVANCED STAGE CANCER. THE REPORTER OF THE EVENT CONFIRMED THE VENTILATOR WILL NOT BE RETURNED TO THE MANUFACTURER FOR FURTHER INVESTIGATION. THE REPORTER OF THE EVENT BELIEVES THE EVENT WAS CAUSED BY THE PATIENT'S CONDITION AND THE VENTILATOR SETTINGS. THE PATIENT DATA LOGS WERE DOWNLOADED FROM THE DEVICE AND SENT TO THE MANUFACTURER. REVIEW OF THE PATIENT DATA LOGS INDICATES PATIENT DISCONNECT CONDITIONS DID OCCUR ON THE REPORTED DATE OF THE EVENT AS EVIDENCED BY SEVERAL LARGE SPIKES IN LEAK VALUES. NO OTHER ANOMALIES WERE OBSERVED. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURER BELIEVES THE DEVICE OPERATES AND ALARMS AS DESIGNED. |  | Y | N | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01697 | 2021-06-14 | 2021-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01713 | 2021-06-14 | 2021-06-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD, OXYGEN BLENDING MANIFOLD, INTERFACE BOARD AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01736 | 2021-06-14 | 2021-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040005 | 1040005 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED STEPS DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE NEEDS REPLACED TO ADDRESS THE ISSUES. |  | N | N | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01721 | 2021-06-15 | 2021-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01755 | 2021-06-15 | 2021-06-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01791 | 2021-06-15 | 2021-05-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A MALFUNCTION OF A VENTILATOR'S TOUCHSCREEN. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01756 | 2021-06-15 | 2021-05-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1032804 | 1032804 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A HIGH TEMPERATURE ALARM OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S THERMISTOR NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01760 | 2021-06-15 | 2021-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2200X11B | DS2200X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION OCCURRED RELATED TO A FAILURE OF THE BACKLIGHT IN THE DISPLAY. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS BEEN EVALUATED BY A THIRD PARTY FIELD SERVICE ENGINEER, HOWEVER, THE INVESTIGATION IS YET TO BE COMPLETED. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE THIRD PARTY FIELD SERVICE ENGINEER INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION OCCURRED RELATED TO A FAILURE OF THE BACKLIGHT IN THE DISPLAY. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS EVALUATED BY A THIRD PARTY FIELD SERVICE ENGINEER AND THE CUSTOMERS COMPLAINT WAS CONFIRMED. THE DEVICE'S DISPLAY BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01761 | 2021-06-15 | 2021-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01754 | 2021-06-15 | 2021-06-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01720 | 2021-06-15 | 2021-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01762 | 2021-06-15 | 2021-05-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200T11 | DSX200T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01723 | 2021-06-15 | 2021-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2200X11B | DS2200X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S MACHINE FLOW SENSOR NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01742 | 2021-06-15 | 2021-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01741 | 2021-06-15 | 2021-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01743 | 2021-06-15 | 2021-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01740 | 2021-06-15 | 2021-06-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE TO POWER ON WAS OBSERVED. THE DEVICE'S POWER MANAGEMENT BOARD AND POWER SUPPLY WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01739 | 2021-06-15 | 2021-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR HAD A HIGH LEAK. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED AND A "VENTILATOR INOPERATIVE" ALARM CONDITION WAS OBSERVED. THE DEVICE'S MACHINE FLOW SENSOR WAS FOUND TO BE CONTAMINATED WITH DIRT AND DUST AND WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01738 | 2021-06-15 | 2021-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BLOWER MOTOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01748 | 2021-06-15 | 2021-06-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01722 | 2021-06-15 | 2021-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01725 | 2021-06-16 | 2021-05-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01764 | 2021-06-16 | 2021-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTOSV | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS INC | IN951S | IN951S |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO THE DEVICE'S SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01727 | 2021-06-16 | 2021-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01726 | 2021-06-16 | 2021-05-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01744 | 2021-06-16 | 2021-05-17 | I, n, j, u, r, y | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT'S CO2 LEVELS WERE ELEVATED WHILE USING A VENTILATOR. THE PATIENT WAS TAKEN TO A HOSPITAL AND PLACED ON ANOTHER DEVICE. THE PATIENT FULLY RECOVERED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED STEPS DURING TESTING. THE DEVICE'S SENSOR BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. |  | Y | N | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01730 | 2021-06-16 | 2021-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01718 | 2021-06-16 | 2021-06-11 | D, e, a, t, h | Manufacturer report | BIPAP AVAPS CSERIES | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | PHILIPS RESPIRONICS INC. | CN1167HS | CN1167HS |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BIPAP AVAPS CSERIES DEVICE WAS NOT OPERATING PROPERLY AND THE PATIENT EXPIRED. THE REPORTER OF THE EVENT CONFIRMS THE PATIENT'S FAMILY IS NOT RELEASING THE DEVICE. THE DEVICE IS NOT RETURNING FOR EVALUATION. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. |  | Y | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01766 | 2021-06-16 | 2021-06-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE DEVICE WAS RETURNED TO THE MANUFACTURER QUALITY PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. DURING THE INITIAL VISUAL INVESTIGATION OF THE EXTERIOR OF THE DEVICE, FOUND EVIDENCE OF BLACK FOAM LIKE MATERIAL IN THE HUMIDIFIERS WATER CHAMBER AND IN THE DEVASTATION UNIT'S AIRFLOW OUTLET PORT AND THE HUMIDIFIERS AIR FLOW INLET PORT. INTERNAL INVESTIGATION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND THE DEVICE TO HAVE CONTAMINATION. THE MANUFACTURER FOUND THERE ARE BLACK SPECKS OF THE FOAM MATERIAL ON THE INSIDE OF THE UNIT. THE MANUFACTURER ALSO FOUND A PORTION OF THE FOAM MATERIAL MISSING IN A PORTION OF THE SOUND ABATEMENT FOAM ENCLOSURE. THE MANUFACTURER ALSO FOUND THE BLOWER MOTOR/BOX HAS THE OILY DEGRADED FOAM MATERIAL. THE MANUFACTURER CONCLUDES THAT SOUND ABATEMENT FOAM HAS EVIDENCE OF DEGRADATION WITHIN THE DEVICE. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01765 | 2021-06-16 | 2021-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION EXPERT | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | CAX501T12C | CAX501T12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING BLACK SUBSTANCES EVERYWHERE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL PART OF THE DEVICE AND FOUND BLACK SUBSTANCES EVERYWHERE MOOG BLOWER, BLOWER DRY BOX- TOP AND BOTTOM, DIRT FROM THE EXTERNAL SOURCE. THE MANUFACTURER INSPECTED THE DEVICE INTERNALLY AND FOUND CORRISSION/RUST, DISCOLORATION TO THE PCA POWER CONNECTOR P4 DUE TO THE WATER INGRESS TO THE POWER CONNECTORS. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE FIVE ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED. SECTION D8, D9 AND H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01763 | 2021-06-16 | 2021-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED PARTICLES INSIDE WATER CHAMBER. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONFIRMED THE CUSTOMER'S ALLEGATION AND THERE WAS VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED. IN THE PREVIOUS REPORT SECTION A1 (PATIENT IDENTIFIER) WAS NOT CAPTURED, CORRECTED IN THIS REPORT. SECTION H6 WERE UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01724 | 2021-06-16 | 2021-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01767 | 2021-06-16 | 2021-05-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TO CHARGE ITS INTERNAL BATTERY. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S POWER MANAGEMENT BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01729 | 2021-06-16 | 2021-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01775 | 2021-06-16 | 2021-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. THE OXYGEN BLENDING MODULE BOARD ONLY WAS RETURNED TO THE MANUFACTURER'S QUALITY ASSURANCE LABORATORY FOR FURTHER INVESTIGATION. DURING THE INVESTIGATION THE ROOT CAUSE OF THE FAILURE WAS DUE TO THE OXYGEN AND AIR FLOW SENSOR'S VOLTAGE BEING OUT OF SPECIFICATION. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01728 | 2021-06-16 | 2021-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01768 | 2021-06-16 | 2021-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040000 | 1040000 |  | A VENTILATOR WAS EVALUATED BY A FIELD SERVICE ENGINEER. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE, THE DEVICE FAILED STEPS DURING TESTING. THE DEVICE'S BLOWER MOTOR NEEDS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01774 | 2021-06-16 | 2021-05-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S REMOTE ALARM CONNECTOR FAILED TO ALARM. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR'S REMOTE ALARM CONNECTOR FAILED TO ALARM. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01634 | 2021-06-16 | 2021-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040000 | 1040000 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TO POWER ON. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION, AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. A REVIEW OF THE DEVICE'S DOWNLOADED EVENT LOG REVEALED VENTILATOR INOPERATIVE CODES. THE DEVICE'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUES. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR FAILED TO POWER ON AND THE DEVICE'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. DURING FURTHER EVALUATION OF THE DEVICE, A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01776 | 2021-06-16 | 2021-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, A FAILURE OF THE DEVICE'S POWER SUPPLY WAS OBSERVED. THE DEVICE'S POWER SUPPLY WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR'S POWER SUPPLY WAS REPLACED TO ADDRESS A POWER SUPPLY FAILURE. ADDITIONAL EVALUATION WAS DONE AND THE VENTILATOR FAILED TEST STEPS DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUES. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
MW5101965 | 2021-06-17 | 2021-01-18 | I, n, j, u, r, y | Voluntary report | DREAMSTATION (CPAP MACHINE) | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. | J29194494F845 | DSX500T11C |  | I STARTED USING THE PHILIPS DREAM MACHINE (CPAP MACHINE) AT THE END OF (B)(6) 2019. I DID NOT NOTICE ANY SIDE EFFECTS. I STARTED GETTING NEW DAILY CHRONIC HEADACHES IN (B)(6), WITH NAUSEA AND DIZZINESS. I DIDN'T THINK IT COULD BE RELATED TO MY CPAP MACHINE. I THOUGHT IT WAS MAYBE DUE TO STRESS. I SAW MY DOCTOR AT THE END OF (B)(6) WHEN MY HEADACHES WEREN'T RESPONDING WELL TO ADVIL/TYLENOL. SHE GAVE ME TRIPTAN MEDICATION TO TRY AND ALSO PUT ME ON FLUOXETINE. THE HEADACHES DECREASED SOMEWHAT, BUT STILL OCCURRED DAILY (A DULL HEADACHE UPON WAKENING THAT PERSISTED THROUGHOUT THE DAY). THE NAUSEA AND DIZZINESS STILL OCCUR ALMOST DAILY. MY DOCTOR THEN HAD ME DO AN MRI, WHICH CAN BACK FINE. YESTERDAY, ON THE NEWS, I SAW THAT MY CPAP MACHINE HAS A RECALL DUE TO DETERIORATING FOAM THAT CAN CAUSE BREATHING ISSUES AND HEADACHES. THESE FOAM PARTICLES ARE ALSO CARCINOGENIC, ACCORDING TO THE REPORT. I STOPPED USING MY MACHINE LAST NIGHT AND CONTACTED MY RESPIRATORY THERAPIST AND DOCTOR. MY RESPIRATORY THERAPIST ADVISED ME TO FILL OUT AN ADVERSE REACTION FORM. I AM ALSO GOING TO REGISTER MY MACHINE WITH PHILIPS. I AM STILL WAITING TO HEAR BACK FROM MY DOCTOR. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-06
MW5101975 | 2021-06-17 | 2021-06-14 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION AUTO-CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I OWN TWO PHILIPS DREAMSTATION AUTO-CPAP MACHINES, AND ONE DREAMSTATION GO AUTO-CPAP MACHINES. I HAVE SEVERE SLEEP APNEA AND REQUIRE THESE MACHINES TO AVOID A HEART ATTACK AND DEATH. I HAVE DISCOVERED TODAY THAT PHILIPS HAS DECIDED TO MAKE SOME KIND OF "LAZY RECALL" OF THESE DEFECTIVE DEVICES. THEY ANNOUNCED A RECALL BUT THERE IS NO WAY TO REGISTER THE DEFECTIVE MACHINE WITH THEM. (THIS IS STUPID ANYWAY; THE MACHINES GENERALLY REQUIRE USE OF THE PHILIPS DREAMMAPPER APP ON A COMPUTER OR IPHONE, AND THEY HAVE ALL THE OWNER REGISTRATION DETAILS ALREADY). IF THESE MACHINES ARE DANGEROUS AND BEING RECALLED I NEED MY MACHINES REPLACED. THE PHILIPS "PRESS RELEASE" SAYS "THE BENEFIT OF CONTINUED USAGE OF THESE THERAPY DEVICES MAY OUTWEIGH THE RISKS IDENTIFIED IN THE RECALL NOTIFICATION" IS ABSURD IN THE EXTREME. I'M SUPPOSED TO CHOOSE BETWEEN CANCER OR A HEART ATTACK? THIS IS BEING HORRIBLY HANDLED. TIME TO PROTECT THE CONSUMERS. THESE ARE MEDICAL DEVICES. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | I |  | 2021-06
MW5101976 | 2021-06-17 | 2021-06-15 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO CPAP HUM HT DOM DX500T11 | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS / RESPIRONICS, INC. | DSX500T11 |  |  | I AM A PHYSICIAN. I USE THE PHILIPS RESPIRONICS DREAMSTATION CPAP DSX500T11 FOR SLEEP APNEA. I RECEIVED NOTIFICATION OF RECALL FOR EMISSION OF VOLATILE ORGANIC COMPOUNDS. I HAVE BEEN USING THE DEVICE SINCE (B)(6) 2019. I HAVE BEEN EXPERIENCING PROFOUND FATIGUE, COUGH, UPPER AIRWAY IRRITATION, SORE THROAT FOR THE PAST SIX MONTHS OR MORE. PHILIPS WEBSITE FOR INITIATING REPLACEMENT OF THE DEVICE IS NOT FUNCTIONAL. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-06
MW5101978 | 2021-06-17 | 2000-05-05 | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS, INC. | DSX500T11C |  |  | I HAVE BEEN USING A CPAP FOR MANY MANY YEARS. THIS DREAMSTATION IS ABOUT 2-3 YEARS SLID. I HAVE NOT BEEN FEELING WELL AT ALL AFTER WAKING UP- BAD HEADACHES, WHICH GO IN FOR WEEKS, SINUS ISSUES, HYPERSENSITIVE TO SOUNDS, NAUSEA, SORE THROAT AND EYE AND NOSE IRRITATION MANY DOCTORS VISITS FOR HEADACHES AND MIGRAINES- INCLUDING A THREE DAY INFUSION. I ALSO HAVE COMMON VARIABLE IMMUNE DEFICIENCY, SO I AM VARY CONCERNED ABOUT POTENTIAL CHEMICALS HARMING ME. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-06
MW5101989 | 2021-06-17 | 2021-05-20 | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS / PHILIPS ELECTRONICS LTD | DSX500T11C | DSX500T11C |  | I JUST BEGAN USING PHILIPS RESPIRONICS DREAMSTATION CPAP DEVICE. DEVELOPED UNKNOWN BURNS AND BLISTERS TO MY FACE ABOVE LIP, CHIN, CHEEKS, AND THROAT IRRITATION. I ALSO SINCE USE WAS DIAGNOSED AT STAGE 3A KIDNEY DISEASE. READ TODAY ON PHILIPS WEBSITE THAT GASES RELEASED FROM POLY-ESTER POLYURETHANE (PE-PUR) CAN LEAD TO SKIN AND AIRWAY IRRITATIONS, LIVER AND KIDNEY DAMAGE, AND CARCINOGENIC EFFECTS. MY NEUROLOGIST AND DERMATOLOGIST WERE BAFFLED AS TO WHAT WAS CAUSING THE BURNS TO MY FACE. THIS RELEASE BY PHILIPS IS MORE LIKELY THAN NOT THE CAUSATION FOR MY INJURIES. ADDITIONAL DOCTOR VISITS TO THE DERMATOLOGIST. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-06
MW5101987 | 2021-06-17 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS ELECTRONICS LTD | DSX500T11C |  |  | ADVISED BY NEWS PROGRAMS THAT DREAMSTATION AUTO CPAP UNDER RECALL. PHILIPS REP ADVISED PATIENT TO D/C USE. DEVICE SUPPLIER, VERYS, ADVISED TO CONTINUE USE OF DEVICE. PULMONOLOGIST GAVE NO RECOMMENDATION EITHER WAY. PATIENT'S LABILE HYPERTENSION MAY BE AFFECTED BY INTERRUPTION IN CPAP USE. AT THIS TIME NO IMMEDIATE REMEDY IS IN PLACE. PHILIPS, VERYS AND TREATING PHYSICIAN HAVE NO INFORMATION OR REMEDY FOR IMMEDIATE REPAIR/REPLACEMENT. PARTIES HAVE CONFLICTING OR NO INFORMATION REGARDING CONTINUED USE OF DEVICE. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | I |  | 2021-06
2518422-2021-01780 | 2021-06-17 | 2021-05-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A MALFUNCTION OF THE OXYGEN BLENDING MODULE BOARD. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUES. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE NAVIGATIONAL BUTTON FAILED TO RESPOND. THE DEVICE'S FRONT PANEL/KEYPAD LED BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01781 | 2021-06-17 | 2021-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01758 | 2021-06-17 | 2021-05-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S ACTIVE EXHALATION CONTROL MODULE WAS LEAKING. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OVERHEAD BLOWER FILTER WAS REPLACED TO ADDRESS THE ISSUE. THERE WAS EVIDENCE OF CONTAMINATION. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR'S ACTIVE EXHALATION CONTROL MODULE WAS LEAKING. THE DEVICE WAS NOT IN PATIENT USE. THE ACTIVE EXHALATION CONTROL MODULE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY AND THE ROOT CAUSE FOR THE FAILURE WAS THE PROPORTIONAL VALVE DID NOT CLOSE PROPERLY. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01759 | 2021-06-17 | 2021-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054097 | 1054097 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A MALFUNCTION OF A VENTILATOR'S ACTIVE EXHALATION CONTROL MODULE. THERE WAS NO PATIENT HARM OR INJURY REPORTED. THE DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER BUT THE EVALUATION HAS YET TO BEGIN. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01783 | 2021-06-17 | 2021-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01782 | 2021-06-17 | 2021-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TO CHARGE ITS INTERNAL BATTERY. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S POWER SUPPLY WAS REPLACED TO ADDRESS THE ISSUE. | THIS MDR IS BEING SUBMITTED AS PART OF A BATCH SUBMISSION OF PREVIOUSLY CLOSED COMPLAINTS THAT WERE REASSESSED AS REPORTABLE FOAM DEGRADATION COMPLAINTS; DISCOVERED AS PART OF A RETROSPECTIVE REMEDIATION REVIEW. THE MANUFACTURER PREVIOUSLY REPORTED RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TO CHARGE ITS INTERNAL BATTERY. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S POWER SUPPLY WAS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION, THE INLET AIR PATH ASSEMBLY, REMOVABLE AIR PATH FOAM, 43 MICRON FILTER AND POLLEN FILTER WERE REPLACED DUE TO DIRT CONTAMINATION. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01785 | 2021-06-18 | 2021-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICES LCD DISPLAY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01786 | 2021-06-18 | 2021-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICES LCD DISPLAY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01784 | 2021-06-18 | 2021-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01810 | 2021-06-21 | 2021-05-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR ALARMED FOR A VENTILATOR INOPERATIVE CONDITION. THE DEVICE WAS NOT IN PATIENT USE. THE MANUFACTURER RECEIVED THE DEVICE FOR EVALUATION, AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01811 | 2021-06-21 | 2021-05-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLANK DURING USE AND WAS NOT VIEWABLE. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S LCD SCREEN NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01813 | 2021-06-21 | 2021-05-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1032804 | 1032804 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, A VENTILATOR'S LCD WAS OBSERVED TO BE BLANK. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S LCD DISPLAY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01770 | 2021-06-21 | 2021-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED EVENT LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01814 | 2021-06-21 | 2021-06-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01772 | 2021-06-21 | 2021-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED STEPS DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01773 | 2021-06-21 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED STEPS DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01812 | 2021-06-21 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED STEPS DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01771 | 2021-06-21 | 2021-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE'S DISPLAY WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LAB FOR EVALUATION AND WAS FOUND TO MALFUNCTION. THE DISPLAY WAS SCRAPPED PER MANUFACTURER PROTOCOL. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01752 | 2021-06-21 | 2021-04-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION FROM SOCIAL MEDIA ALLEGING A THERMAL EVENT OCCURRED WITH A DREAMSTATION DEVICE. NO ADDITIONAL INFORMATION WAS RECEIVED. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. |  | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01769 | 2021-06-21 | 2021-05-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR ALARMED FOR A VENTILATOR INOPERATIVE CONDITION. THE DEVICE WAS NOT IN PATIENT USE. THE MANUFACTURER RECEIVED THE DEVICE FOR EVALUATION, AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. THERE WAS EVIDENCE OF WATER INGRESS TO THE DEVICE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01788 | 2021-06-22 | 2021-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY NEEDS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01787 | 2021-06-22 | 2021-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DISPLAY SCREEN WAS BLANK. THE DEVICE'S DISPLAY WAS REPLACED DUE TO PHYSICAL DAMAGE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01830 | 2021-06-22 | 2021-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILS TO CHARGE ITS INTERNAL BATTERY. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S POWER MANAGEMENT BOARD WAS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01757 | 2021-06-22 | 2021-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS EVALUATED BY A THIRD PARTY FIELD SERVICE ENGINEER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE THIRD PARTY FIELD SERVICE ENGINEER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERNAL BATTERY NEED REPLACED TO ADDRESS THE ISSUES. | THE MANUFACTURER PREVIOUSLY SUBMITTED A REPORT WITH AN INCORRECT DATE FOR SECTION B4 AS (B)(6) 2021. THE CORRECT DATE SHOULD BE (B)(6) 2022. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01789 | 2021-06-22 | 2021-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01816 | 2021-06-22 | 2021-06-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR RETURNING FOR SERVICE AND 'SERVICE REQUIRED" CODES OBSERVED IN THE DOWNLOADED EVENT LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUE. SECTION G.3 WAS ALSO MISSING BUT A CORRECTION WAS FILED UNDER MDR 2518422-2021-01816-1. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01790 | 2021-06-22 | 2021-06-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01817 | 2021-06-22 | 2021-06-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICES LCD DISPLAY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01819 | 2021-06-22 | 2021-05-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" ALARM CONDITION OCCURRED RELATED TO "LOW CIRCUIT LEAK". THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED A SERVICE REQUIRED CODE FOUND IN A VENTILATORS DOWNLOADED EVENT LOG DURING SERVICE AND THE SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. B.3 WAS OMITTED ON THE FIRST REPORT. THE DATE SHOULD BE (B)(6) 2021 AND IS REFLECTED IN THIS REPORT. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01820 | 2021-06-22 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01808 | 2021-06-22 | 2021-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01833 | 2021-06-22 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE MACHINE FLOW SENSOR WAS OBSERVED. THE DEVICE'S MACHINE FLOW SENSOR WAS REPLACED DUE TO DIRT AND DUST CONTAMINATION. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
MW5102022 | 2021-06-22 | 2021-02-01 | I, n, j, u, r, y | Voluntary report | PHILLIPS DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500T11 | DSX500T11 |  | CHRONIC COUGH AND SINUS CONGESTION. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01826 | 2021-06-22 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11 | DSX500T11 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO THE DEVICE'S SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01837 | 2021-06-22 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01828 | 2021-06-22 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700S11 | DSX700S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01838 | 2021-06-22 | 2021-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WAS ONE ERROR FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. UNIT SCRAPPED. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01824 | 2021-06-22 | 2021-05-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | CAX500T12C | CAX500T12C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LAB FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LAB FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THE PATIENT ALSO HAD BEEN GETTING DRY NOSE AND THAT CAUSED HIM TO HAVE A NOSEBLEED FROM THE DRYNESS. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL PART OF THE DEVICE AND FOUND NO EVIDENCE OF THERMAL OR PHYSICAL DAMAGE. THE MANUFACTURER DID FIND THE WATER HAS BEEN ON THE HUMIDIFIERS HEATER PLATE. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND FOUND NO EVIDENCE OF THERMAL DAMAGE. THE MANUFACTURER DID OBSERVE EVIDENCE OF CONTAMINANTS PRESENT IN THE DREAMSTATION/HUMIDIFIER. THE MANUFACTURER FOUND THE SOUND ABATEMENT FOAM HAS STARTED TO DETERIORATE AND THERE ARE BLACK SPECKS OF THE FOAM MATERIAL ON THE INSIDE OF THE DREAMSTATION. ALSO FOUND THE FOAM MATERIAL IS MISSING IN A PORTION OF THE SOUND ABATEMENTS ENCLOSURE, THE BLOWER MOTOR/BOX HAS THE TACKY/OILY DEGRADED FOAM MATERIAL AND THE EVIDENCE OF WATER INGRESS. THE AIR PATH OF THE DEVICE WAS COMPROMISED AND THE CONTAMINATE WOULD HAVE BEEN IN THE DEVICES AIR PATH. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND NO EVIDENCE OF ERROR. THE MANUFACTURER APPLIED POWER TO THE DEVICES. THE DEVICES DO POWER UP PROVIDE FLOW AND THE HUMIDIFIER HEATER PLATE AND TUBING HEATS. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DETERIORATION. THE DEVICE HAS EVIDENCE OF THE PRESENCE OF CONTAMINANTS AND EVIDENCE OF WATER INGRESS. SECTION A1, A2, A3, B3, D4, D9, H6 WERE UPDATED/CORRECTED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01815 | 2021-06-22 | 2021-06-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG AND THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND ACTIVE EXHALATION CONTROL MODULE WERE REPLACED TO ADDRESS THE ISSUES. DURING THE EVALUATION, THERE WAS PHYSICAL DAMAGE OBSERVED TO THE UNIVERSAL PORTING BLOCK. THE UNIVERSAL PORTING BLOCK WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01831 | 2021-06-22 | 2021-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01805 | 2021-06-22 | 2021-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED A "SERVICE REQUIRED" ALARM CONDITION THAT WAS ADDRESSED BY RELOADING THE SOFTWARE. IN THE ORIGINAL REPORT G3 WAS REPORTED AS 06/16/2018. THE CORRECT DATE FOR G3 IS 06/16/2021. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01809 | 2021-06-22 | 2021-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2100X15B | IN2100X15B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE EXTERNAL FLOW SENSOR WAS OBSERVED. THE DEVICE'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01821 | 2021-06-22 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO SERIES ASSY, INT | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | IN751S | IN751S |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-06
MW5102051 | 2021-06-22 | 2020-03-25 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO CPAP HUM DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS INC. | DSX500H11 |  |  | ABOUT 30 - 45 DAYS AFTER BEGINNING USE OF PHILIPS DREAMSTATION CPAP MACHINE A SORE DEVELOPED IN MY NOSTRIL. I WOULD USE SALINE GEL AND ANTIBIOTIC OINTMENT ON IT. IT WOULD HEAL BRIEFLY BUT IT WOULD COME RIGHT BACK IN A DAY OR TWO. ON (B)(6) 2021 WHEN THIS CPAP MACHINE WAS RECALLED I WENT TO MY MD AND PURCHASE A NEW DREAMSTATION 2 THAT WAS NOT PART OF THE RECALL. THE SORE IN MY NOSTRIL HAS HEALED AND HAS NOT RETURNED. I ALWAYS CLEAN AND CARE FOR MY MACHINE, HOSE, AND MASK AS ADVISED. I DID NOT INITIALLY LINK THE SORE TO THE CPAP. I HAD THOUGHT IT WAS ALL THE KN95 MASK I HAVE BEEN WEARING AS A HEALTHCARE WORKER CAUSING ME TO HAVE MORE SINUS DRAINAGE. HOWEVER, I AM STILL A HEALTHCARE WORKER AND I WEAR KN95 MASK EACH DAY AT WORK STILL. THE ONLY DIFFERENCE NOW IS THE NEW CPAP MACHINE. THUS, IT IS MY BELIEF THAT THE PARTICLES OR FUMES FROM THE SOUND ABATEMENT FOAM IN THE ORIGINAL (NOW RECALLED) PHILIPS DREAMSTATION HAS BEEN CAUSING THIS LONG TERM IRRITATION OF MY NASAL PASSAGES AND THE SORE IN MY NOSTRIL WAS THE MOST PROMINENT RESULT OF THIS IRRITATION. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-06
MW5102031 | 2021-06-22 | 2021-05-24 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500T11C |  |  | SEVERE COUGHING; STARTING IN (B)(6) 2020 I HAD AN COUGH THROUGH OUT THE DAY AND NIGHT, IT HAS INCREASED TO A POINT I HAVE LOST MY VOICE, HAVING SEVERE COUGHING FITS THAT MAKES IT DIFFICULT TO BREATH AND GAGING. I AM VERY WORRIED ABOUT LONG TERM AFFECTS THAT HAVE NOT BEEN DETECTED YET. USED THE DEVICE FOR 3 YEARS AND STARTED GETTING AFFECTS IN (B)(6) 2020 GRADUALLY GOT WORSE AND WAS SEEING A DOCTOR FOR MY HEALTH ISSUES, THEN FOUND OUT ABOUT A RECALL ON (B)(6) 2021. THAT LIST THE EFFECTS I HAVE BEEN HAVING. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-06
MW5102027 | 2021-06-22 | 2018-08-15 | I, n, j, u, r, y | Voluntary report | DREAM STATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I AM A PHYSICIAN (NEPHROLOGIST). MY WIFE STARTED USING CPAP (PHILIPS DREAM STATION) ABOUT (B)(6) 2016. SHE DEVELOPED RIGHT SIDED CHF ABOUT (B)(6) 2016. RIGHT-SIDED CARDIAC CATH THEN SHOWED MILD PULMONARY HYPERTENSION AND DIASTOLIC DYSFUNCTION. THIS IS WELL-CONTROLLED WITH DIURETICS AND ANTI-HYPERTENSIVES. SHE HAS HAD MILD ASTHMA FOR DECADES AND WAS VIRTUALLY ASYMPTOMATIC ON FLOVENT 110 QD AND PRILOSEC 20 QD. IN LATE 2018 HER ASTHMA BEGAN TO WORSEN- IT HAS CONTINUED TO GRADUALLY BECOME MORE SEVERE SINCE DESPITE INTENSIFICATION OF HER MEDICAL REGIMEN. BOTH HER PULMONOLOGIST AND CARDIOLOGIST CONSIDER HER EUVOLEMIC. SHE SAW A PULMONOLOGIST ABOUT (B)(6) 2019 WHO DID NOT CONSIDER HER SYMPTOMS OF A PRODUCTIVE COUGH AND WHEEZING RELATED AT ALL TO SIX MONTHS OF AMIODARONE SHE TOOK WHICH ENDED ABOUT (B)(6) 2018. PFTS DONE IN BOTH 2019 AND 2020 WERE NORMAL. DUE TO THE RECALL OF DREAM STATION SHE DID NOT USE HER CPAP THE PAST TWO EVENINGS (B)(6). HER ASTHMA BEGINNING ON AWAKENING (B)(6) WAS VIRTUALLY GONE AND HAS CONTINUED TO BE VIRTUALLY ABSENT. SHE HAD USED HER DREAM STATION EVERY NIGHT SINCE SHE STARTED USING IT IN 2016 UNTIL NOW. WHEN HOSPITALIZED SHE ALWAYS USED HER OWN CPAP MACHINE. SHE DEVELOPED REFRACTORY PERSISTENT ATRIAL FIBRILLATION IN (B)(6) 2016. DESPITE AN ECHO (UNCHANGED IN LATE 2019) SHOWING ONLY MILD LVH AND A NORMAL STRESS TEST THIS HAS CONTINUED DESPITE THREE ATRIAL ABLATIONS (FIRST ONE (B)(6) 2016) AND MULTIPLE CARDIOVERSIONS. HER ELECTROPHYSIOLOGIST HAS CONSIDERED HER FAILURE TO RESPOND TO THE THREE ABLATIONS VERY UNUSUAL. SHE WAS HYPERTHYROID FOR ABOUT THREE MONTHS IN MID-2018 FROM AMIODARONE. SHE HAS MAINLY USED VINEGAR FOR HER DAILY CLEANING OF HER CPAP MACHINE. SHE HAS OCCASIONALLY USED MILD DISH SOAP. SHE HAS NEVER USED OZONE. SHE HAS NEVER NOTED ANY BLACK PARTICLES IN HER CPAP TUBING OR MASK. FOLLOWING ANKLE SURGERY IN (B)(6) 2020 DUE TO HER IMMOBILITY SHE DID SPRAY CLEANSMART DAILY SURFACE CLEANER TO CLEAN HER CPAP MACHINE FOR ABOUT TWO MONTHS. SHE GOT A NEW CPAP MACHINE (AGAIN PHILIPS DREAM STATION) ABOUT A MONTH AND A HALF AGO. THE REASON FOR THAT WAS THAT THE HUMIDIFICATION CHAMBER STOPPED WORKING. TO SUMMARIZE, IT APPEARS THAT THE MARKED WORSENING OF HER MILD ASTHMA WAS PROBABLY DUE TO THE DREAM STATION CPAP MACHINE AS ON STOPPING USING IT HER ASTHMA IMMEDIATELY ALMOST COMPLETELY RESOLVED. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | I |  | 2021-06
2518422-2021-01719 | 2021-06-22 | 2021-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LEAK VALUE WAS HIGH. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE WILL NEED RECALIBRATED OR THE SENSOR BOARD WILL NEED TO BE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
MW5102024 | 2021-06-22 | 2021-06-14 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAM STATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSXZ700 |  |  | CANNOT USE EQUIPMENT; STOP USE OF CPAP MACHINE BECAUSE OF DEFECT THAT RESULTS IN PARTICLES BEING BREATHED IN. ALSO PRESENCE OF CARCINOGENS HAVE BEEN REPORTED IN USING THE EQUIPMENT. HAVE BEEN TOLD TO STOP IMMEDIATE USE OF THE DEVICE. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-06
MW5102021 | 2021-06-22 | 2015-04-27 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS SYSTEM ONE CPAP | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | PHILIPS RESPIRONICS / RESPIRONICS, INC. | REM STAR AUTO DOM 560P |  |  | IN MID (B)(6) 2015, I WAS HAVING SYMPTOMS OF AN EAR OR SINUS INFECTION. SEEKING SOME SOLUTIONS ON MY OWN, I PURCHASED A NEW CPAP TO SEE IF PERHAPS IT WAS CONTRIBUTING TO THE ISSUE. I RECEIVED THE NEW PHILLIPS CPAP ON OR ABOUT THE (B)(6) 2015 AND SYMPTOMS GREW WORSE. MY SYMPTOMS WERE BAD ENOUGH TO GO TO THE DOCTOR AND TO HAVE MY DOCTOR SIGN A "DO NOT FLY" ORDER WHICH RESULTED IN MY BEING TEMPORARILY DISABLED (COLLECTING DISABILITY) I CONTINUE TO HAVE SYMPTOMS, HAVE GONE TO AN ENT AND HAVE MY EARS AND THROAT CHECKED EACH TIME I GO TO THE DOCTOR. I HAVE SINCE RETIRED AS I WAS NO LONGER WILLING TO PUT MY BODY THROUGH THE EXTREME DISCOMFORT OF FLYING AND ALL OF THE STRESS RELATED TO THE PROCESS OF DECOMPRESSION AND COMPRESSION WITH FLYING. THOUGH I AM NOT DISABLED, I DO SUFFER RECURRING (EVERY OTHER DAY) FEELING OF FULLNESS AND IN DIZZINESS IN MY HEAD FOR WHICH THERE HAS BEEN NO EXPLANATION. WITH THIS CURRENT RECALL, I AM EAGERLY AWAITING THE OPPORTUNITY TO TRY A DIFFERENT CPAP MACHINE TO DETERMINE IF THIS IS THE SOURCE OF MY ONGOING DISCOMFORT. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-06
MW5102017 | 2021-06-22 |  | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAMSTATION AUTO BIPAP HUM DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS / PHILIPS ELECTRONICS LTD. | DSX700H11 | DSX700H11 |  | I LEARNED OF THE VOLUNTARY RECALL FOR THE PHILIPS DREAMSTATION CPAP (SN (B)(4)). I HAVE ATTEMPTED TO CONTACT THEM VIA PHONE. I GET A RECORDING THEN A LANGUAGE SELECTION OPTION. AFTER I SELECT I AM A PATIENT. THE RECORDING STATES THEY ARE HAVING A HIGH CALL VOLUME AND TO GO TO THEIR WEBSITE, THEN THE CALL IS DISCONNECTED. I WENT TO THEIR WEBSITE TO REGISTER MY PRODUCT. I ENTERED THE SERIAL NUMBER AND FOR OVER 20 MINUTES A SET OF "GEARS" ARE SPINNING ON THE PAGE AND NOTHING IS HAPPENING. PER THEIR RECALL NOTICE I AM NOT TO USE MY CPAP AND SHOULD CONTACT MY PHYSICIAN, WHICH I HAVE DONE. I AM AWAITING A CALL BACK. I AM DEPENDENT ON MY CPAP AND I DO NOT KNOW HOW TO PROCEED. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | I |  | 2021-06
2518422-2021-01832 | 2021-06-22 | 2021-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01806 | 2021-06-22 | 2021-06-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01827 | 2021-06-22 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | ITX400S15 | ITX400S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT OBSERVED BLACK PARTICLES INSIDE THE MASK AND THE FILTERS. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DESPITE OF THREE ATTEMPTS (B)(6) 2022, (B)(6) 2022, (B)(6) 2022 ; THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. THERE HAS BEEN NO RESPONSE FROM THE CUSTOMER ON THE RETURN OF THE DEVICE. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01804 | 2021-06-22 | 2021-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2017865-2021-23951 | 2021-06-23 |  | D, e, a, t, h | Manufacturer report | TRILOGY SR | IMPLANTABLE PACEMAKER PULSE GENERATOR | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2260L |  |  | RELATED MANUFACTURING REFERENCE NUMBER: 2017865-2021-23952. IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED. | FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED. UDI NOT AVAILABLE AS PRODUCT WAS MANUFACTURED ON OR BEFORE 23 SEP 2014. THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. | Y | N | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01844 | 2021-06-23 | 2021-05-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLANK. THERE WAS NO PATIENT HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLANK. THERE WAS NO PATIENT HARM OR INJURY REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
MW5102123 | 2021-06-23 | 2021-05-23 | I, n, j, u, r, y | Voluntary report | DREAMSTATION 1 CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. |  |  |  | PATIENT CALLED TO REPORT NOTICE OF RECALL FOR HIS DREAMSTATION CPAP DEVICE. PATIENT STATED HE'S HAD A CHRONIC COUGH FOR THE LAST 4 WEEKS, COUGHED UP BLOOD FOR TWO OF THOSE WEEKS, HAS HEADACHES, FATIGUE AND IS NOT SLEEPING WELL. PATIENT STATED HE WAS MADE AWARE OF THE RECALL LAST MONDAY AND WAS TOLD BY HIS DOCTOR TO DISCONTINUE USE OF THE DEVICE. PATIENT STATED THE RECALL WAS RELATED TO A FOAM INSULATION THAT WAS OXIDIZING AND PUSHING PARTICLES INTO THE LUNGS. |  | Y | Y | I |  | 2021-06
MW5102108 | 2021-06-23 | 2021-06-21 | I, n, j, u, r, y | Voluntary report | PHILLIPS RESPIRONICS DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS RESPIRONICS/ RESPIRONICS, INC. |  |  |  | PATIENT REPORTED COUGH ATTRIBUTED TO PHILLIPS RESPIRONICS DREAMSTATION AUTO CPAP USE. FDA SAFETY REPORT ID # (B)(4). |  | Y | N | LAY USER/PATIENT |  | 2021-06
2518422-2021-01859 | 2021-06-23 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTRE FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT WAS SCRAPPED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01850 | 2021-06-23 | 2021-06-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DS460HS | DS460HS |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED DUE TO TOBACCO ODOR. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01848 | 2021-06-23 | 2021-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP AUTOSV | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC | DSX900T11C | DSX900T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO THE DEVICE'S SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01849 | 2021-06-23 | 2021-06-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11 | DSX400H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01835 | 2021-06-23 | 2021-06-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01841 | 2021-06-23 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO THE DEVICE'S SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01840 | 2021-06-23 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGED HEADACHE, VOMITING, SICK TO THE STOMACH AND SHE ALSO STATED THAT SHE HAD A CHEMICAL SMELL COMING FROM DEVICE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. CORRECTED INFORMATION WAS PROVIDED IN H.6. SECTION (IN PREVIOUS MDR HEADACHE AND VOMITING WAS NOT INCLUDED). | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01852 | 2021-06-23 | 2021-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | PHILIPS RESPIRONICS, INC. | 1032800 | 1032800 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. THE MANUFACTURER RECEIVED INFORMATION STATING THE DEVICE WAS RETURNED TO THE CUSTOMER UNREPAIRED. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01857 | 2021-06-23 | 2021-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | PHILIPS RESPIRONICS INC. | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR TIDAL VOLUME WAS NOT FUNCTIONING. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S EXTERNAL FLOW SENSOR WAS CLEANED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
MW5102113 | 2021-06-23 | 2021-06-14 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ / RESPIRONICS INC | DREAMSTATION |  |  | ADMITTED CARCINOGENIC EFFECT; I USED THE PHILIPS RESPIRONICS DREAMSTATION FOR THE PAST 4 YEARS AND MOST THEY HAVE COME OUT WITH A RECALL REGARDING THEIR FOAM BREAKDOWN AND OFF-GASSING OF CARCINOGENIC CHEMICALS (FOAM IS ACTUALLY IN THE AIRWAY PATH). PRIOR TO THIS RECALL, IN THE PAST 3 MONTHS I HAVE BEGUN TO EXPERIENCE CHEST PAIN AND BURNING SENSATIONS AFTER USING THIS PRODUCT-MOST NOTABLE UPON WAKING AND THEN WANING OVER THE COURSE OF THE DAY. RESPIRONICS CLAIM THE FDA IS NOT MANDATING A RECALL. I WOULD LIKE TO KNOW WHY THERE IS NOT A MANDATED RECALL WITH A MANUFACTURER'S ADMITTED KNOWN CARCINOGENIC EFFECTS IN A PRODUCT DESIGNED TO IMPROVE HEALTH. THEY CLAIM A LOW RATE OF INCIDENCE-THAT IS PROBABLY BECAUSE NO ONE IN THEIR RIGHT MIND WOULD THINK TO ASSOCIATE THESE NEGATIVE EFFECTS WITH A DEVICE DESIGNED TO IMPROVE THEIR HEALTH. ONLY AFTER THEY CREATED A RECALL DID I PUT TWO AND TWO TOGETHER. IT IS UNCONSCIONABLE TO NOT MANDATE A RECALL WHERE THE MANUFACTURER IS SAYING DO NOT USE ANY FURTHER. I DEMAND TO KNOW WHAT ACTION IS BEING TAKEN WITH REGARDS TO ALLOWING THIS OCCUR. HOW WAS THIS PRODUCT EVEN APPROVED? FOAM WILL BREAKDOWN. WHAT WAS THE MANUFACTURER'S CLAIMED LIFESPAN OF THE FOAM? WHAT WAS THE KNOWN CHEMICAL COMPOSITION OF THE FOAM AND WHY DID THE FDA APPROVE IT? I AM GOING TO HIRE A LAWYER AND SEEK THIS INFORMATION AS I BELIEVE THE FDA IS ALSO LIABLE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-06
MW5102119 | 2021-06-23 | 2021-05-28 | I, n, j, u, r, y | Voluntary report | DREAM STATION AUTO CPAP HUM HT CELL DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC | DREAM STATION AUTO CPAP HUMHTCEL | UNKNOWN |  | I INSERTED A LONG METAL BOBBY PIN INTO MY HAIR AND USED THE MACHINE. BEFORE I FELL ASLEEP, I FELT A JOLT OF ELECTRICITY SEEMINGLY TO COME FROM MY HEART THROUGH MY RIGHT BREAST AND THROUGH THE NIPPLE. MY AT HOME EKG SHOWED WILD ACTIVITY. EVEN THOUGH I HAVE AFIB, THE ELECTRICAL SIGNALS WERE NEVER THAT WILD. I WAS SORE FOR ABOUT A WEEK, AND THE SYMPTOMS ALL WENT AWAY FINALLY . I REPORTED THIS TO THE HEALTHLINE OFFICE WHERE I GOT THE MACHINE, AS WELL AS TO MY CARDIOLOGIST. I HAVE NOT SLEPT WITH ANY PINS OR ANYTHING ELSE IN MY HAIR SINCE. NO FURTHER PROBLEMS OF THAT NATURE TO REPORT. I UNDERSTAND THERE IS A RECALL IN EFFECT, AND I HAVE STOPPED USING THE MACHINE. I AM HOPING TO GET ANOTHER MACHINE NOT AFFECTED BY THE RECALL. I BELIEVE THE RECALL INVOLVES FOAM MATERIAL? THANK YOU. (B)(6). THE MACHINE IS UNDER RECALL, AND THE INFORMATION IS AS FOLLOWS; PHILIPS RESPIRONICS DREAM STATION AUTO CPAP HUM HT CELL DOM. REFERENCE # DSX 500 TIIC, WITH DREAM STATION HUM CORE, PACK DOM, REFERENCE #DSXHCP. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01842 | 2021-06-23 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01834 | 2021-06-23 | 2021-06-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01846 | 2021-06-23 | 2021-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BLOWER MOTOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01797 | 2021-06-23 | 2021-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED AN OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS A "SERVICE REQUIRED" ERROR CODE. THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S INTERFACE BOARD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01795 | 2021-06-23 | 2021-06-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S FLOW SENSOR ASSEMBLY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01796 | 2021-06-23 | 2021-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01851 | 2021-06-23 | 2021-06-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS NOT DELIVERING PRESSURE. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR WAS NOT DELIVERING PRESSURE. THERE WAS NO HARM OR INJURY REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01798 | 2021-06-23 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01803 | 2021-06-23 | 2021-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01800 | 2021-06-23 | 2021-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01793 | 2021-06-23 | 2021-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S POWER MANAGEMENT BOARD AND SYSTEM BOARD NEED TO BE REPLACED TO ADDRESS THE ISSUE. AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01801 | 2021-06-23 | 2021-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01802 | 2021-06-23 | 2021-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE TO POWER ON WAS OBSERVED. THE DEVICE'S SYSTEM BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01799 | 2021-06-23 | 2021-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01843 | 2021-06-23 | 2021-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLANK. THERE WAS NO PATIENT HARM OR INJURY REPORTED. THE DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLANK. THERE WAS NO PATIENT HARM OR INJURY REPORTED. THE DEVICE WAS EVALUATED BY A THIRD PARTY SERVICE CENTER AND THE CUSTOMER'S COMPLAINT COULD NOT BE DUPLICATED. THE DEVICE WAS FOUND TO OPERATE AND ALARM AS DESIGNED. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01845 | 2021-06-23 | 2021-05-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2200X11B | DS2200X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A FAILURE OF A VENTILATOR'S OXYGEN SENSOR. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A FAILURE OF A VENTILATOR'S OXYGEN SENSOR. THE DEVICE WAS NOT IN PATIENT USE. THE OXYGEN SENSOR WAS AN FIO2 SENSOR ASSEMBLY AND IS AN ACCESSORY TO THE VENTILATOR. THIS ACCESSORY WOULD NOT AFFECT THE VENTILATOR TO DELIEVER THERAPY AS DESIGNED. THERE WAS NO MALFUNCTION RELATED TO THE VENTILATOR. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01794 | 2021-06-23 | 2021-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040004 | 1040004 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01792 | 2021-06-23 | 2021-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BLOWER MOTOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01860 | 2021-06-24 | 2021-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S ALARMS STOPPED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUNDABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S ALARMS STOPPED RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. SECTION H9 UPDATED IN THIS REPORT. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01861 | 2021-06-24 | 2021-05-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S TOUCHSCREEN WAS DAMAGED. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION THAT A VENTILATOR'S LCD SCREEN WAS DAMAGED. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01865 | 2021-06-24 | 2021-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD AND MACHINE FLOW SENSOR WERE REPLACED TO ADDRESS THE ISSUE. THERE WAS EVIDENCE OF WATER INGRESS. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01872 | 2021-06-24 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01856 | 2021-06-24 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01870 | 2021-06-24 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01869 | 2021-06-24 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S BLOWER MOTOR WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY A VENTILATOR RETURNING FOR PREVENTATIVE MAINTENANCE AND A SERVICE REQUIRED CODE WAS FOUND IN THE DOWNLOADED EVENT LOG AND THE DEVICE'S BLOWER MOTOR WAS REPLACED TO ADDRESS THE ISSUE. DURING ADDITIONAL TESTING, THE DEVICE FAILED A TEST STEP. THE DEVICES SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01853 | 2021-06-24 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01854 | 2021-06-24 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE TUBING, INTERNAL BATTERY CONNECTOR AND SPEAKER CONNECTORS WERE RECONNECTED TO ADDRESS THE ISSUES. THE DEVICE APPEARS TO HAVE BEEN TAMPERED WITH. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S BLOWER BELLOWS WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01863 | 2021-06-24 | 2021-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE OXYGEN BLENDING MODULE BOARD WAS OBSERVED. THE OXYGEN BLENDING MODULE BOARD WAS CLEANED TO ADDRESS THE ISSUE. THERE WAS EVIDENCE OF CONTAMINATION. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01866 | 2021-06-24 | 2021-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, A VENTILATOR'S LCD WAS OBSERVED TO BE BLANK. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S LCD DISPLAY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01855 | 2021-06-24 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S BLOWER BELLOWS WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01864 | 2021-06-24 | 2021-06-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS FAILING A TEST STEP. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01880 | 2021-06-24 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
MW5102148 | 2021-06-24 | 2021-05-03 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS / RESPIRONICS, INC. | DSX400H11 |  |  | PATIENT WAS USING A PHILIPS RESPIRONICS DREAMSTATION CPAP. THEY HAD A BLACK SUBSTANCE LEAKING OUT OF HUMIDIFIER CHAMBER AND TUBING. PATIENT HAS HAD AN ON GOING SINUS INFECTION AND PROBLEMS FOR A FEW MONTHS NOW. WE REPLACED THE TUBING AND HUMIDIFIER CHAMBER. ABOUT A MONTH LATER ALL THE DREAMSTATION CPAP'S HAVE BEEN RECALLED. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-06
MW5102132 | 2021-06-24 | 2020-09-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO CPAPHUMCELL DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | ABOUT 9 MONTHS AGO I BEGAN TO HAVE A SORE THROAT ON THE LEFT TONSIL/AREA. IT EBBED AND FLOWED BUT GRADUALLY GOT WORSE. I WENT TO THE ENT FIRST BUT I WAS ALREADY ON AN ANTIBIOTIC FROM SOMETHING ELSE SO THEY COULD NOT CULTURE IT. A FEW MONTHS PASSED AND I WENT TO MY PRIMARY AND HAD A CULTURE TAKEN. IT CAME POSITIVE WITH BACTERIA BUT COULD NOT IDENTIFY IT EXACTLY. WAS PUT ON ANTIBIOTIC. IT SEEMED TO WORK BUT A MONTH LATER IT CAME BACK. SAW MY PRIMARY AGAIN. HAD A SWAB/CULTURE TAKEN. CAME BACK NEGATIVE BUT SHE REFERRED ME TO ANOTHER ENT. HE SAID HE'LL TAKE A CULTURE AND IF IT COMES BACK NEGATIVE MAYBE WE SHOULD DISCUSS HAVING TONSILS REMOVED. IT DID COME BACK NEGATIVE BUT HE WROTE ME FOR ANTIBIOTIC AS SOMETIMES CULTURES DO NOT RECOGNIZE OTHER BACTERIA OR VIRUSES. I AM NOW ON 7TH DAY OF 14 DAY DOSAGE OF ANTIBIOTIC AND HAVE NO IMPROVEMENT. I AM SUPPOSED TO BE SCHEDULED FOR AN LONG AWAITED AND OVERDUE TOTAL HIP REPLACEMENT. I AM CONCERNED THAT MY SWOLLEN/INFLAMED LEFT TONSIL MAY PREVENT ME FROM EITHER HAVING THE HIP REPLACEMENT OR CREATE PROBLEMS DURING THE HIP SURGERY AS I WILL BE INTUBATED OR, MAY HAVE TO POSTPONE HIP SURGERY UNTIL THEY BIOPSY MY TONSIL AREA AND/OR REMOVED THEM COMPLETELY. AND, NOW FEW DAYS AGO I RECEIVED EMAIL FROM MY HEALTH INSURANCE COMPANY SAYING MY PHILIPS CPAP MAY BE ON A RECALL LIST AND TO REGISTER IT ON LINK FOR PHILIPS TO SEE IT IF IS ON RECALL. WELL, IT IS ON RECALL. MY CONCERN IS MY THROAT ISSUE IS ONE OF THE SYMPTOMS OF A MALFUNCTIONING MACHINE AND/OR CPAP CLEANER THAT CAUSED IT TO MALFUNCTION. I HAVE BEEN USING A DEVICE CALLED SOCLEAN. THE PHILIPS RECALL LISTS THINGS LIKE POSSIBLE CARCINOGENS FROM DEGRADED SOUND PROOFING FOAM. THEY SAY THE HIGH HEAT FROM SOCLEAN OZONE CLEANER MAY HAVE DEGRADED THE SOUND PROOFING FOAM. ALSO OTHER THINGS LIKE BLACK DEBRIS IN THE HUMIDIFIER PART OF CPAP (I HAVE SEEN THAT DEBRIS FLOATING IN THE WATER CHAMBER VERY OFTEN, I ALWAYS REMOVE IT AND WASH IT OUT). OTHER SYMPTOMS ARE COUGHING, SINUS DRAINAGE LIKE NEVER BEFORE THAT IS EXTREMELY UNCOMFORTABLE AND UNUSUALLY VISCOUS. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-06
MW5102130 | 2021-06-24 |  | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS CPAP BY PAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DREAM STATION AUTO CPAP DOM |  |  | RECALL ON CPAP MACHINE DUE TO FOAM CONTAMINATION FROM MACHINE THAT CAUSED SINUS AND LUNG AND POSSIBLY KIDNEY AND BLADDER ISSUES .QUIT USING AND REVERTED TO OLD MACHINE AND FILED CLAIM WITH MANUFACTURER FOR REPLACEMENT IN THERE DETAILS I HAVE ALL SYMPTOMS OF POSSIBLE CAUSE OF FOAM PARTICLES BEING INGESTED FAMILY DOCTOR TREATMENT OF ANTIBIOTICS. SINUS HEADACHE SINUS INFECTION BRONCHITIS AND KIDNEY INFECTION. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-06
12056952 | 2021-06-24 | 2021-06-14 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | OMNILAB ADVANCED PLUS | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | PHILIPS RESPIRONICS, INC. | OMNI LAB ADVANCED PLUS |  |  | THE OMNI LAB ADVANCED PLUS MACHINES WERE RECALLED FROM PHILIPS RESPIRONICS. WE HAD TO CANCEL 11 PATIENTS AND ARE NOT SCHEDULING ANY CPAP (CONTINUOUS POSITIVE AIRWAY PRESSURE THERAPY)/BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) PATIENTS UNTIL FURTHER NOTICE. MANUFACTURER RESPONSE FOR CIPAP - BIPAP, (BRAND NOT PROVIDED) (PER SITE REPORTER). MANUFACTURER MADE US AWARE OF RECALL. |  | N | Y | I |  | 2021-06
2518422-2021-01868 | 2021-06-24 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01818 | 2021-06-24 | 2021-05-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | FR2110X14B | FR2110X14B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01878 | 2021-06-24 | 2021-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX600H11C | DSX600H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO HAVE A LUNG NODULE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER HAS ATTEMPTED TO ARRANGE FOR THE RETURN OF THE DEVICE TO THE MANUFACTURER'S PRODUCT INVESTIGATION LAB. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO HAVE A LUNG NODULE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THAT TIME. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND UNKNOWN DUST CONTAMINANT ON THE FRONT PANEL, BOTTOM ENCLOSURE, BLOWER, BLOWER SEAL, AND BLOWER BOX. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND ERROR E-193. THE MANUFACTURER CONCLUDES NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. AFTER FURTHER REVIEW, THE MANUFACTURER CONFIRMED THAT THE PREVIOUSLY SUBMITTED REPORT WAS INCORRECTLY FILED AS ADVERSE EVENT, WHICH SHOULD HAVE BEEN FILED AS PRODUCT PROBLEM ONLY. THE CORRECT B5 SECTION SHOULD BE- THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED. THERE IS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. SECTION B1, B2, B7, H1, D9, G3, H6 HAS BEEN UPDATED / CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01879 | 2021-06-24 | 2021-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO HAVE LUNG NODULES. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER HAS ATTEMPTED TO ARRANGE FOR THE RETURN OF THE DEVICE TO THE MANUFACTURER'S PRODUCT INVESTIGATION LAB. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 11, 2024, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICESTHE MANUFACTURER RECEIVED INFORMATION ALLEGING TO HAVE LUNG NODULES RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTIONS B1, B2, H1 AND H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR ONLY THE PRODUCT PROBLEM (BOTH ADVERSE EVENTS AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01875 | 2021-06-24 | 2021-05-24 | D, e, a, t, h | Manufacturer report | BIPAP VISION | VENTILATOR, CONTINUOUS, MINIMAL VENTILATORY SUPPORT, FACILITY USE |  | 582059 | 582059 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A EXPIRED WHILE USING A BIPAP VISION. THE INVESTIGATION IS STILL ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER HAS COMPLETED THE INVESTIGATION. | THE MANUFACTURER PREVIOUSLY REPORTED A PATIENT EXPIRED WHILE USING A BIPAP VISION. THE DEVICE IS NOT RETURNING TO THE MANUFACTURER FOR EVALUATION. THE REPORTING FACILITY CONFIRMED THE PATIENT WAS NOT ON THE DEVICE AT THE TIME OF THE REPORTED EVENT. THE REPORTING FACILITY CONFIRMED THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO THE PATIENT'S DEATH. | Y | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01858 | 2021-06-24 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY CONFIRMED THE CUSTOMER'S ALLEGATION AND THERE WAS VISIBLE FOAM DEGRADATION AND UNIT SCRABBED DUE TO AGE. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01881 | 2021-06-24 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01882 | 2021-06-24 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01883 | 2021-06-24 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01884 | 2021-06-24 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERFACE BOARD AND OXYGEN BLENDING MODULE BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01862 | 2021-06-24 | 2021-06-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. ATTEMPTS TO RETRIEVE FURTHER INFORMATION REGARDING THE DEVICE HAVE BEEN UNSUCCESSFUL. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | CORRECTION: MANUFACTURER PREVIOUSLY REPORTED ON MDR 2518422-2021-01706. | N | Y | LAY USER/PATIENT |  | 2021-06
MW5102154 | 2021-06-25 | 2021-05-15 | I, n, j, u, r, y | Voluntary report | CPAP MACHINE DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. | DREAMSTATION |  |  | LONG LASTING HEADACHES, EYE IRRITATION, MILD CHEST PAIN, IMBALANCE, DIZZINESS; HEADACHES LASTING FOR 35 MINUTES ON TOP OF HEAD, SLEPT WELL PREVIOUS NIGHT BUT WAS VERY TIRED FROM NOON TILL 5 PM, FELT UNSTEADY. FELL AGAINST WALLS AT HOME, DIZZINESS. WENT TO PERFORMANCE PHYSICAL THERAPY FOR VERTIGO TREATMENT. AFTER 5-6 PT SESSIONS VERTIGO PROBLEM RESOLVED. STARTING ABOUT (B)(6) 2021, HEADACHES LASTING LONGER-MORE INTENSE CONTINUOUS ALL DAY. I HAVE AN APPOINTMENT AT (B)(6) 2021. MANUFACTURED DATE 11/24/2021. |  | Y | N | LAY USER/PATIENT |  | 2021-06
2518422-2021-01822 | 2021-06-25 | 2021-06-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S PRESSURE DELIVERY WAS LOW. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR WAS FOUND TO BE FAULTY AND CONTAMINATED WITH DUST AND DIRT. THE MACHINE FLOW SENSOR NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01898 | 2021-06-25 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED LUNG DEGRADATION. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE PATIENT REFUSED TO PROVIDE DETAILS ON THE SERIAL NUMBER OF THE DEVICE OR CONTACT INFORMATION. THE MANUFACTURER IS UNABLE TO ARRANGE FOR THE RETURN OF THE DEVICE FOR EVALUATION DUE TO A LACK OF INFORMATION PROVIDED BY THE PATIENT. IF FURTHER INFORMATION REGARDING THE DEVICE BECOME AVAILABLE, IT WILL BE SUBMITTED IN THE FOLLOW-UP REPORT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01899 | 2021-06-25 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE RESPIRATORY ISSUES. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A DOCTOR'S ASSESSMENT AND PRESCRIBED MEDICATION. THE MANUFACTURER HAS ATTEMPTED TO ARRANGE FOR THE RETURN OF THE DEVICE TO THE MANUFACTURER'S PRODUCT INVESTIGATION LAB. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUNDABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING ANISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO EXPERIENCING HEADACHES, SEVERE SLEEP APNEA, RESPIRATORY ISSUES, NAUSEA, AND PNEUMONIA. THE REPORTED EVENTS OF HEADACHES, SEVERE SLEEP APNEA, RESPIRATORY ISSUES, NAUSEA, AND PNEUMONIA AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURER CONCLUDES NO FURTHER ACTION IS NECESSARY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL AND INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND OBSERVED THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM. THE ISSUE OF DEGRADED SOUND ABATEMENT FOAM IS ADDRESSED BY CAPA 7211. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER POWERED UP THE DEVICE AND PROVIDES AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS PRESENCE OF DEGRADED SOUND ABATEMENT FOAM. SECTIONS B1, B2 HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, MEDICAL DEVICE PROBLEM CODE,TYPE OF INVESTIGATION,INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-06
MW5102159 | 2021-06-25 |  | I, n, j, u, r, y | Voluntary report | PHILLIPS RESPIRONICS DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS / RESPIRONICS, INC. |  |  |  | PATIENT REPORTED TO EXPERIENCE CHEST TIGHTNESS, CHEST PAIN FOR DUE TO INFLAMMATION FELT TO BE CAUSED BY HER PHILLIPS RESPIRONICS DREAMSTATION AUTO CPAP, WHICH HAS RECENTLY BEEN RECALLED. SHE WENT TO ED FOR CHEST PAIN ON (B)(6) 2021 FOR CHEST PAIN, NO CARDIAC CAUSE FOUND. FDA SAFETY REPORT ID# (B)(4). |  | Y | N | LAY USER/PATIENT |  | 2021-06
2518422-2021-01823 | 2021-06-25 | 2021-06-17 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AUX500T15 | AUX500T15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN END USER IS INVESTIGATING IF A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE HAS CAUSED A LUMP IN HER THROAT RELATED TO THE SOUND ABATEMENT FOAM. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. A FINAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED THE DEVICES FOR EVALUATION. THE MANUFACTURER FOUND NO EVIDENCE OF THERMAL, PHYSICAL DAMAGE OR ODOR. THE MANUFACTURER OPERATED THE DREAMSTATION AND HUMIDIFIER WITH A BACTERIA FILTER IN LINE DURING TESTING. THERE WAS NO EVIDENCE OF ANY BLACK DEBRIS ON THE BACTERIA FILTER, NOR ANY EVIDENCE OF FOAM MATERIAL OR ODORS BEING EMITTED DURING TESTING. THE INTERNAL INSPECTION OF THE DEVICES FOUND NO EVIDENCE OF LOOSE SOUND ABATEMENT FOAM OR ODOR IN THE BLOWER MOTOR/BOX OR IN THE AIRPATH OF THE DEVICE. THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CRF PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01911 | 2021-06-25 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID/HT, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | AUX400T15 | AUX400T15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT SAID THAT BLACK PARTICLES IN SWIVEL TUBE CONNECTOR TO MACHINE AND TUBING. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. SECTION H6 CODES CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT HAS BEEN CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01839 | 2021-06-25 | 2021-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID/HT, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AUX400T15 | AUX400T15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. SECTION H6 MEDICAL DEVICE PROBLEM CODE WAS CODED INCORRECTLY ON THE PREVIOUS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01912 | 2021-06-25 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/HT/CELL, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AUX400T15C | AUX400T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01706 | 2021-06-25 | 2021-06-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION FROM A DURABLE MEDICAL EQUIPMENT (DME) SUPPLIER STATING THAT AN END USER WOKE UP AND FOUND BLACK FOAM MATERIAL ALL OVER HIS PILLOW WHILE USING HIS CPAP DEVICE. THE END USER HAS BEEN IN AND OUT OF THE HOSPITAL WITH LUNG ISSUES AND HAS BEEN PRESCRIBED INHALERS. THE MANUFACTURER REQUESTED THE RETURN OF THE DEVICE WITH ALL ACCESSORIES FOR EVALUATION. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON MARCH 7, 2024, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED AN ALLEGATION OF AN ISSUE RELATED TO A DURABLE MEDICAL EQUIPMENT (DME) SUPPLIER STATING THAT AN END USER WOKE UP AND FOUND BLACK FOAM MATERIAL ALL OVER HIS PILLOW WHILE USING HIS CPAP DEVICE. THE PATIENT HAS BEEN IN AND OUT OF THE HOSPITAL WITH LUNG ISSUES AND HAS BEEN PRESCRIBED INHALERS. PATIENTS ARE ALSO EXPERIENCING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS OR HEADACHE, NAUSEA, VOMITING, ASTHMA (NEW OR WORSENING), AND LUNG DISEASES. SECTION H6 HEALTH EFFECTS: CLINICAL CODES HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01836 | 2021-06-25 | 2021-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BLOWER MOTOR NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01909 | 2021-06-25 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | RJPX500S16 | RJPX500S16 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE REPORTER ALSO ALLEGED VISUALIZATION OF BLACK DUST INSIDE A TUBE. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL PART OF THE DEVICE AND FOUND DUST AND DIRT IN THE AIR OUTLET, INCLUDING BLACK PARTICULATE THAT APPEARED TO BE FIBROUS IN NATURE. EXAMINATION OF THE AIR INLET PORT FOUND WHITISH CONTAMINATION BENEATH THE POLLEN FILTER. THE POLLEN FILTER APPEARED TO BE UNUSED AND LABELED WITH A MARKER. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND OBSERVED DISCOLORATION ON THE SD CARD INTERFACE ASSEMBLY INCLUDING WHAT APPEARS TO BE A PARTIAL FINGERPRINT. THE MANUFACTURER ALSO OBSERVED THAT THE SILVER COATING OF THE BLUETOOTH ANTENNA TRACE APPEARED FADED. CORROSION WAS OBSERVED ON THE TOP OF THE DC INPUT CONNECTOR OF THE P4. EXAMINATION OF THE BLOWER BOX ASSEMBLY FOUND DUST CONTAMINATION THROUGHOUT THE AIRPATH, ON THE BLOWER, ON THE BLOWER OUTPUT SEAL, AND ON THE BLOWER ISOLATORS. THE MANUFACTURER OBSERVED NO EVIDENCE OF SOUND ABATEMENT FOAM DAMAGE OR DEGRADATION. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER WAS ABLE TO CONFIRM THE COMPLAINT. CONTAMINATION WAS FOUND IN THE AIR INLET PORT SUGGESTING AN EXTERNAL SOURCE FOR THE WHITE PARTICULATE. THE BLACK DUST OBSERVATION IS MORE CONSISTENT WITH TEXTILE FIBER THAN WITH SOUND ABATEMENT FOAM DEGRADATION. FURTHERMORE, NO BLACK PARTICULATE WAS OBSERVED IN THE BLOWER BOX, ONLY WHITE PARTICULATE. NO BLOWER HOUR DATA AND NO CARE ORCHESTRATOR DATA CONFIRMS THAT THE DEVICE WAS EITHER RESET PRIOR TO THE MANUFACTURER RECEIVING THE DEVICE, OR THE PATIENT HAS NOT USED THIS DEVICE. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION BUT CONTAMINATIONS THROUGHOUT THE DEVICE. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01889 | 2021-06-25 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AUX500T15C | AUX500T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | SECTION B5 WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT SHOULD BE REPORTED AS FOLLOWS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED COUGH. THERE WAS NO PATIENT HARM OR INJURY. NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. IN SECTION D4 ADDITIONAL UNIQUE DEVICE IDENTIFIER (UDIS) NUMBER WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT. IT HAS BEEN CORRECTED OR UPDATED IN THIS REPORT. SECTION H6 HEALTH EFFECT - CLINICAL CODE WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT. IT HAS BEEN CORRECTED OR UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01913 | 2021-06-25 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL, AU | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | AUX500T15C | AUX500T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | SECTION B5 WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT SHOULD BE REPORTED AS FOLLOWS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION PATIENT REPORTS SOME SINUS IRRITATION AND BLACK PARTICLES ON MASK AND IN TUBING. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. SECTION H6 CODES CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT HAS BEEN CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01897 | 2021-06-25 | 2021-06-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT WAS DIAGNOSED WITH LUNG CANCER. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE MANUFACTURER HAS ATTEMPTED TO ARRANGE FOR THE RETURN OF THE DEVICE TO THE MANUFACTURER'S PRODUCT INVESTIGATION LAB. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT WAS DIAGNOSED WITH LUNG CANCER. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE REPORTED EVENT AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THE EVENT IS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01908 | 2021-06-28 | 2021-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING THROAT ISSUE (FEELS LIKE A LUMP IN THROAT), SHORTNESS OF BREATH, COUGHING, ANXIETY, CHEST TIGHTNESS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. AFTER THE FINAL ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, CUSTOMER DECLINED TO RESPOND TO THE GFE RELATED QUESTIONS AND TERMINATED THE CALL. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. CORRECTED INFORMATION PROVIDED IN SECTIONS A1 (PATIENT INDENTIFIER ADDED), D4 (UDI NUMBER ADDED, IT WAS WRONG IN INITIAL REPORT), H6 (HEALTH EFFECT CLINICAL CODE). SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01894 | 2021-06-28 | 2021-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE SINUS INFECTIONS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTION MEDICATIONS. THE MANUFACTURER HAS ATTEMPTED TO ARRANGE FOR THE RETURN OF THE DEVICE TO THE MANUFACTURER'S PRODUCT INVESTIGATION LAB. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 06 , 2024 AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING THAT THE PATIENT HAD A SINUS INFECTION RELATED TO THE CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGATION OF NOSE IRRITATION AND MULTIPLE SINUS INFECTIONS. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. | Y | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01907 | 2021-06-28 | 2021-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING UNEXPLAINED WHEEZING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY AFTER THE SECOND ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, CUSTOMER REPLIED THAT THE DEVICE WAS NOT AVAILABLE FOR RETURN AS IT WAS IN THE GARBAGE. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01903 | 2021-06-28 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01918 | 2021-06-28 | 2021-06-23 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND HAS CAUSED A NODULE IN THE PATIENT'S LUNGS. THE DETAILS OF MEDICAL INTERVENTION THE PATIENT RECEIVED ARE CURRENTLY UNKNOWN. THE MANUFACTURER HAS ATTEMPTED TO ARRANGE FOR THE RETURN OF THE DEVICE FOR EVALUATION TO THE MANUFACTURER'S PRODUCT INVESTIGATION LAB. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECOME DEGRADED AND CAUSED THE PATIENT TO DEVELOP SINUS INFECTIONS. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE EVENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER QUALITY PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. DURING THE INITIAL VISUAL INSPECTION OF THE DEVICE FOUND NO EVIDENCE OF DISTINCT WEAR, TEAR, OR CONTAMINATION ON THE ENCLOSURE. RED MARKS WERE OBSERVED ON THE UI PANEL MIDWAY BETWEEN THE UI KNOB AND DISPLAY. THE DEVICE WAS DISASSEMBLED FOR INTERNAL VISUAL INSPECTION. THE MANUFACTURER OBSERVED NO DISTINCT WEAR OR TEAR. DUST/DANDER CONTAMINATION WAS OBSERVED ON THE TOP OF THE BLOWER OF THE DEVICE. THE MANUFACTURER ALSO OBSERVED AN OBJECT, APPEARING CONSISTENT WITH THE RETAINING CLIP OF A NON-PHILIPS POLLEN FILTER (COMPARISON PHOTO ATTACHED), IN THE BOTTOM OF THE BLOWER BOX BENEATH THE BLOWER AIR INTAKE. NO EVIDENCE OF SOUND ABATEMENT FOAM BREAKDOWN WAS OBSERVED WITHIN THE DEVICE, AND THE FOAM APPEARS INTACT. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND CONFIRMED THE DEVICE PROVIDED FLOW. THE MANUFACTURER CONCLUDES THERE IS NO EVIDENCE OF SOUND ABATEMENT FOAM BREAKDOWN WITHIN THE DEVICE. THE DUST CONTAMINATION IS MOST LIKELY FROM THE EXTERNAL TO THE DEVICE. | Y | N | LAY USER/PATIENT |  | 2021-06
2518422-2021-01901 | 2021-06-28 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | IN PREVIOUS REPORT SECTION H10 WAS CAPTURED INCORRECTLY AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. BUT IT SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGED NAUSEA AND MILD HEADACHE. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01877 | 2021-06-28 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DIZZINESS AND HEADACHE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HEALTH CLINICAL CODE HAS BEEN CORRECTED AND TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01896 | 2021-06-28 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED SINUS INFECTION, HEADACHE, PARTICLES IN TUBING/CHAMBER, NASAL/THROAT IRRITATION OR SORENES. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED SECTION D4 HAS BEEN UPDATED AND CORRECTED IN THIS REPORT. SECTION H6 AND CLINICAL CODEHAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01876 | 2021-06-28 | 2021-06-15 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON MAY 02, 2024, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING FATIGUE,HEADACHE,STOMACH ACHE,BAD COUGHING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGED NOSE IRRITATION, SKIN IRRITATION AND LUNG DISEASE. THE SECTIONS B1, B2, H1, AND H6 HAVE BEEN UPDATED OR CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR ADVERSE EVENTS (THE PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECTS: CLINICAL CODES AND HEALTH EFFECTS - IMPACT CODE WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01917 | 2021-06-28 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO AUTO W/BT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSG500S11 | DSG500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED NODULES IN THE PATIENT'S LUNGS. THE DETAILS OF THE MEDICAL INTERVENTION THE PATIENT RECEIVED ARE CURRENTLY UNKNOWN. THE MANUFACTURER HAS ATTEMPTED TO ARRANGE FOR THE RETURN OF THE DEVICE FOR EVALUATION TO THE MANUFACTURER'S PRODUCT INVESTIGATION LAB. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON OCT 09, 2023, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGATION ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAVE ALLEGED OF HAVING NODULES IN PATIENT'S LUNGS. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGATION OF DIZZINESS AND HEADACHE. THE SECTION E1 WAS UPDATED AND SECTIONS B1, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECTS: CLINICAL CODES AND HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01873 | 2021-06-28 | 2021-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | CAX400H12C | CAX400H12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. ATTEMPTS TO RETRIEVE FURTHER INFORMATION REGARDING THE DEVICE HAVE BEEN UNSUCCESSFUL. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO THE DEVICE'S SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01895 | 2021-06-28 | 2021-06-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11 | DSX700H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. ADDITIONAL INFORMATION WAS RECEIVED AND ADDED TO THE REPORT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER ON 03/17/2023 FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED ON 05/24/2023. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01916 | 2021-06-28 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN THE INITIAL REPORT SECTION B5 WAS NOT UPDATED, THE CORRECT B5 SHOULD BE THE PATIENT HAS ALLEGED RESPIRATORY ISSUES,MUCUS PRODUCTION, TIGHTNESS, AND HYPERSENSITIVITY WITH SEVERE SLEEP APNEA. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF LIGHT DUST CONTAMINATION ON THE BLOWER, DISCOLORATION AND CORROSION FROM WATER INGRESS AT P4 DC INPUT CONNECTOR.THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS LOGGED. THE MANUFACTURER CONCLUDES CONTAMINATES FOUND WERE CONSISTENT WITH DUST CONTAMINATION ON THE BLOWER, DISCOLORATION AND CORROSION FROM WATER INGRESS AT P4 DC INPUT CONNECTOR. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION D9,G3,H6 HAS BEEN UPDATED/CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01919 | 2021-06-28 | 2021-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING ASTHMA HAS BEEN ACTING UP ALONG WITH SEVERE SINUS HEADACHES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING ASTHMA HAS BEEN ACTING UP ALONG WITH SEVERE SINUS HEADACHES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FOLLOW-UP REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01885 | 2021-06-28 | 2021-06-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN THE INITIAL REPORT SECTION B5 WAS NOT UPDATED, THE CORRECT B5 SHOULD BE - THE PATIENT HAS ALLEGED SNEEZING, EYE BURN, DRY MOUTH, RUNNY NOSE, ITCHY EYES, TAKEN ALLERGIC MEDICATION, VISITED EYE DOCTOR. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER AND FOUND NO EVIDENCE OF CONTAMINATION. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. . THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERROR LOGGED. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01915 | 2021-06-28 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01888 | 2021-06-28 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | GBX500S15 | GBX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01900 | 2021-06-28 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DRY COUGH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED IN SECTION D4 UNIQUE IDENTIFIER (UDI) NUMBER UPDATED IN THIS REPORT. IN SECTION H6 HEALTH EFFECT - CLINICAL CODE HAS BEEN CORRECTED AND TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01910 | 2021-06-28 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 04, 2024, AND SECTION H11 SHOULD BE REPORTED AS THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT ALLEGED TO DEVELOP LUNG ISSUES. NO MEDICAL INTERVENTION WAS REPORTED. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM (ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR). SECTION B2 WAS CHANGED TO BLANK (PREVIOUSLY IT WAS OTHER SERIOUS). SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01953 | 2021-06-28 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1032804 | 1032804 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01950 | 2021-06-28 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01934 | 2021-06-28 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01929 | 2021-06-28 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01926 | 2021-06-28 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01957 | 2021-06-28 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01952 | 2021-06-28 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, ERROR CODES RELATED TO THE CHARGING OF THE INTERNAL BATTERY WERE OBSERVED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01933 | 2021-06-28 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01949 | 2021-06-28 | 2021-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01904 | 2021-06-28 | 2021-06-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01956 | 2021-06-28 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01951 | 2021-06-28 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01954 | 2021-06-28 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S POWER MANAGEMENT BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01931 | 2021-06-28 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01932 | 2021-06-28 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01955 | 2021-06-28 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01925 | 2021-06-28 | 2021-06-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01988 | 2021-06-29 | 2021-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S FRONT PANEL COULD NOT BE ACCESSED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S FRONT PANEL BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01924 | 2021-06-29 | 2021-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WOULD NOT POWER ON. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S POWER SUPPLY AND INTERNAL BATTERY NEED TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01987 | 2021-06-29 | 2021-06-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLANK. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN EVALUATED BY THE THIRD PARTY SERVICE CENTER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE THIRD PARTY INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR'S LCD WAS BLANK. WE WERE UNABLE TO OBTAIN ANY FURTHER INFORMATION AT THAT TIME. THE DEVICE HAS BEEN RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01983 | 2021-06-29 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICES LCD DISPLAY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01981 | 2021-06-29 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01980 | 2021-06-29 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED IN SECTION B.4, DATE OF THIS REPORT AS (B)(6) 2021. THE CORRECT DATE IS (B)(6) 2021. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01982 | 2021-06-29 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG AND THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD, POWER MANAGEMENT BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01989 | 2021-06-29 | 2021-06-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR AC POWER DISCONNECTION WHILE PLUGGED INTO POWER. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S POWER SUPPLY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01936 | 2021-06-29 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01928 | 2021-06-29 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01930 | 2021-06-29 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2200X11B | DS2200X11B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S TUBING WAS RESEATED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01867 | 2021-06-29 | 2021-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, PHYSICAL DAMAGE WAS OBSERVED TO THE ACTIVE EXHALATION CONTROL MODULE. THE ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01975 | 2021-06-29 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, A LOW LEAK WAS OBSERVED WITH A VENTILATOR. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01976 | 2021-06-29 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS,INC. | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICES LCD DISPLAY NEEDS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED THIS DEVICE ON MDR 2518422-2021-01976-1. PLEASE DISREGARD MDR 2518422-2021-01976-1 AS IT WAS FILED IN ERROR. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01977 | 2021-06-29 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01978 | 2021-06-29 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01979 | 2021-06-29 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01994 | 2021-06-29 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | SX500H11C | SX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED PARTICLES BLOWING OUT FROM BLOWER. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING PARTICLES BLOWING OUT FROM BLOWER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. SECTION G1-MANUFACTURING CONTACT OFFICE DETAILS WERE CORRECTLY UPDATED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01937 | 2021-06-29 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01902 | 2021-06-29 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DETAILS OF THE PATIENT'S FIRST NAME AND LAST NAME IS UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01920 | 2021-06-29 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL, AU | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | AUX500T15C | AUX500T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. CORRECT B5 SHOULD BE : THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF BLACK SPECS IN THE CORNER OF THE WATER TANK AND ALSO ALLEGED HAVING COUGH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DESPITE OF MULTIPLE ATTEMPTS, THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION AS THE DEVICE HAS BEEN DISCARDED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION H6 HAS BEEN UPDATED / CORRECTED IN THIS REPORT. IN SECTION D4, UNIQUE IDENTIFIER (UDI) HAS BEEN CORRECTED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01941 | 2021-06-29 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11 | DSX500T11 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING EVALUATION OF THE DEVICE AT A THIRD-PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01939 | 2021-06-29 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11 | DSX500T11 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING EVALUATION OF THE DEVICE AT A THIRD-PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01940 | 2021-06-29 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500S11 | DSX500S11 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING EVALUATION OF THE DEVICE AT A THIRD-PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED BALCK PARTICLES BLOWING OUT FORM DEVICE. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT WAS SCRAPPED. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01943 | 2021-06-29 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11 | DSX500T11 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING EVALUATION OF THE DEVICE AT A THIRD-PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01942 | 2021-06-29 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11 | DSX500T11 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING EVALUATION OF THE DEVICE AT A THIRD-PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01886 | 2021-06-29 | 2021-01-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO HAVE LUNG CANCER. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF CT SCAN. NO FURTHER DETAILS OF THE PATIENT'S MEDICAL INTERVENTION ARE KNOWN AT THIS TIME. ATTEMPTS ARE BEING MADE TO HAVE THE DEVICE RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LAB FOR EVALUATION. THE INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 05 , 2024 AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING CANCER AND LUNG ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 WAS CORRECTED AND UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01887 | 2021-06-29 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED HEADACHES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE 18 ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. UNIT SCRAPPED DUE TO MULTIPLE E53 ERROR CODES. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01891 | 2021-06-29 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS INITIALLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING USING OF OZONE OR OTHER UNAPPROVED CLEANING METHOD AND DISINFECTION METHOD AND ALSO ALLEGED BLACK SOOT INSIDE THE WATER CHAMBER, HAVING EYE AND BREATHING ISSUE AND OIL GLANDS ARE BLOCKED RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL AND INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND OBSERVED AN UNKNOWN DUST CONTAMINANT ON THE REAR PANEL, BOTTOM ENCLOSURE, BLOWER, BLOWER SEAL, AND BLOWER BOX. PIL WAS ABLE TO CONFIRM THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM RESIDUE IN THE BLOWER BOX. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND 2 ERRORS WERE LOGGED. SECTION A1 AND D4 WERE CORRECTED AND SECTION H6 WAS UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01737 | 2021-06-29 | 2021-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION FROM A CUSTOMER THAT AN END USER OBSERVED A THERMAL EVENT RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S POWER CORD. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER RECEIVED THE DEVICE FOR EVALUATION, AND OBSERVED THE DEVICE'S AC POWER CORD HAD PHYSICAL DAMAGE, THERMAL DAMAGE AND WAS SEVERED. AN EVALUATION OF THE DEVICE'S AC POWER CORD REVEALED EXPOSURE OF THE AC CONDUCTING WIRING. THE MANUFACTURER CONCLUDES THE PHYSICAL AND THERMAL DAMAGE ARE LIKELY A RESULT OF THE DEVICE'S AC POWER CORD BECOMING SEVERED DUE TO BEING EXPOSED TO FORCES BEYOND ITS INTENDED DESIGN. THE MANUFACTURER CONCLUDES NO FURTHER ACTION IS NECESSARY AT THIS TIME. | DATE OF REPORT: SHOULD BE 06/29/2021. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01984 | 2021-06-29 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01935 | 2021-06-29 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A HIGH TEMPERATURE ALARM CONDITION WAS OBSERVED. THE DEVICE'S OUTLET FLOWPATH THERMISTOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01923 | 2021-06-29 | 2021-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WOULD NOT POWER ON AND THE LCD DISPLAY WAS BLANK. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD AND LCD SCREEN WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01922 | 2021-06-29 | 2021-06-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1032800 | 1032800 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S AC POWER FAILED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR'S INTERNAL BATTERY WAS REPLACED TO ADDRESS A CHARGING ISSUE. DURING THE EVALUATION A "SERVICE REQUIRED" ALARM WAS OBSERVED IN THE DEVICE'S DOWNLOADED ERROR LOG. THE DEVICE'S POWER MANAGEMENT BOARD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01927 | 2021-06-29 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE'S PRESSURE FLUCTUATED AND THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S BLOWER MOTOR, OXYGEN BLENDING MODULE MANIFOLD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01921 | 2021-06-29 | 2021-06-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1032800 | 1032800 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S OXYGEN DELIVERY WAS LOW. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN CONNECTOR WAS FOUND TO BE BROKEN AND WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-02001 | 2021-06-29 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP PRO SERIES ASSY | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | 651P | 651P |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING EVALUATION OF THE DEVICE AT A THIRD-PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-02002 | 2021-06-29 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | 661P | 661P |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING EVALUATION OF THE DEVICE AT A THIRD-PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | N | Y | LAY USER/PATIENT |  | 2021-06
MW5102216 | 2021-06-30 | 2021-04-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION ASV | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. |  |  |  | STARTING USING PHILIPS RESPIRONICS DREAMSTATION ASV IN (B)(6) 2021. IN (B)(6) 2021 HAVING SEVERE IRRITATION OF ESOPHAGUS, CHEST PRESSURE, DIZZINESS, NAUSEA AND HYPERSENSITIVITY AND IT'S ONGOING. WAS INFORMED YESTERDAY, (B)(6) 2021 OF THE MEDICAL DEVICE RECALL. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01963 | 2021-06-30 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01970 | 2021-06-30 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01965 | 2021-06-30 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S SENSOR BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01966 | 2021-06-30 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01959 | 2021-06-30 | 2021-06-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | TV119040142 | TV119040142 |  | DURING A CHECK-OUT PROCEDURE OF A VENTILATOR AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE TO POWER ON WAS OBSERVED. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01964 | 2021-06-30 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01958 | 2021-06-30 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, ERROR CODES RELATED TO THE CHARGING OF THE INTERNAL BATTERY WERE OBSERVED. THE DEVICE'S INTERNAL BATTERY AND POWER MANAGEMENT BOARD NEED TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-02009 | 2021-06-30 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-02010 | 2021-06-30 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-02017 | 2021-06-30 | 2021-06-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED A TEST STEP DURING TESTING AND THE LCD SCREEN WAS BLANK. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR FAILED A TEST STEP DURING TESTING. THE DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BLOWER MOTOR NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. THE BLOWER MOTOR WAS FOUND TO BE CONTAMINATED WITH DUST. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01938 | 2021-06-30 | 2021-06-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2200X11B | DS2200X11B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS OBSERVED. THE DEVICE'S BLOWER MOTOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-02008 | 2021-06-30 | 2021-06-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-02019 | 2021-06-30 | 2021-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S DISPLAY WAS BLANK. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR'S DISPLAY WAS BLANK. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-02018 | 2021-06-30 | 2021-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S BATTERY WAS DEPLETED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER BUT THE EVALUATION IS NOT YET COMPLETE. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR'S BATTERY WAS DEPLETED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE POWER SUPPLY WAS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION, A "SERVICE REQUIRED" CODE WAS OBSERVED IN THE DOWNLOADED EVENT LOG. THE DEVICE'S SENSOR BOARD AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01969 | 2021-06-30 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "HIGH TEMP" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S BLOWER AND OUTLET FLOW PATH THERMISTOR NEED TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01960 | 2021-06-30 | 2021-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01993 | 2021-06-30 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE.. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01962 | 2021-06-30 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040005 | 1040005 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01893 | 2021-06-30 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED DUE TO AGE". | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01998 | 2021-06-30 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | PHILIPS RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE'S LCD SCREEN WAS FOUND TO BE CRACKED AND UNABLE TO BE VIEWED. THE LCD SCREEN WILL NEED TO BE REPLACED. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01991 | 2021-06-30 | 2021-06-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE.. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01967 | 2021-06-30 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE NEES TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01947 | 2021-06-30 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO SERIES ASSY | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | IN751S | IN751S |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING EVALUATION OF THE DEVICE AT A THIRD-PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE . | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO THE DEVICE'S SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01992 | 2021-06-30 | 2021-06-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S POWER MANAGEMENT BOARD WAS REPLACED TO ADDRESS THE ISSUE.. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-02016 | 2021-06-30 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO, INTL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | 761P | 761P |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING THE EVALUATION OF THE DEVICE AT A THIRD-PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01874 | 2021-06-30 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AUX500T15C | AUX500T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO HAVE HEADACHES AND VOMITING. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTIONS FOR THE MEDICATIONS STEMITIL AND TRAMADOL. THE MANUFACTURER HAS ATTEMPTED TO ARRANGE FOR THE RETURN OF THE DEVICE FOR EVALUATION TO THE MANUFACTURER'S PRODUCT INVESTIGATION LAB. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS : THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED TO EXPERIENCE HEADACHES, VOMITING, ALLERGIC REACTION. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTIONS FOR THE MEDICATIONS STEMITIL AND TRAMADOL. THE REPORTED EVENT OF HEADACHES AND VOMITING, ALLERGIC REACTION AND ITS REPORTED SEVERITY WAS REV IEWED BY THE MANUFACTURER'S CLINICAL EXPERT. BASED ON THE INFORMATION AVAILABLE, THE MANUFACTURER CONCLUDES NO FURTHER ACTION IS NECESSARY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION B1, B2 HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT-IMPACT CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-02021 | 2021-06-30 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO W/BT | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | AUG400S15 | AUG400S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. SECTION H6 CODES CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT HAS BEEN CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01971 | 2021-06-30 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, INTL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | INX500S15 | INX500S15 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING THE EVALUATION OF THE DEVICE AT A THIRD-PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO THE DEVICE'S SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01972 | 2021-06-30 | 2021-06-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, INTL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | INX500T15 | INX500T15 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING THE EVALUATION OF THE DEVICE AT A THIRD-PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO THE DEVICE'S SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01892 | 2021-06-30 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED OF FINDING BLACK SPECS IN THE ELBOW TUBE AND IN ONE OF THE FILTERS. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. AFTER THE THIRD ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, CUSTOMER DECLINED TO RESPOND TO THE GFE RELATED QUESTIONS AND TERMINATED THE CALL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS UPDATED REPORT WILL BE FILED. SECTION H6 AND G1 HAS BEEN UPDATED IN THIS REPORT. SECTION D9 HAS BEEN CAPTURED INCORRECTLY IN INITIAL MDR 2518422-2021-01892, IT HAS BEEN CORRECTED IN THIS REPORT. SECTION G4 THE COMBINATION PRODUCT WAS WAS INCORRECTLY CAPTURED AS YES IN MDR 2518422-2021-01892-1, IT HAS BEEN CORRECTED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01995 | 2021-06-30 | 2021-06-29 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AUX500T15 | AUX500T15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A BLOOD CLOT IN THEIR LUNG. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A DOCTOR'S ASSESSMENT AND A PRESCRIBED BLOOD THINNER MEDICATION. THE MANUFACTURER IS ATTEMPTING TO HAVE THE DEVICE RETURNED TO ITS PRODUCT INVESTIGATION LAB. THE INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOVEMBER 12, 2024 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGED TO DEVELOP A BLOOD CLOT IN THEIR LUNG. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A DOCTOR'S ASSESSMENT AND A PRESCRIBED BLOOD THINNER MEDICATION. THE REPORTED EVENT OF BLOOD CLOT IN THEIR LUNG RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTIONS B1, B2, H1 AND H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR BOTH ADVERSE EVENTS AND PRODUCT PROBLEM (ONLY THE PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECTS - CLINICAL CODES AND HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-06
MW5102217 | 2021-06-30 | 2018-03-28 | I, n, j, u, r, y | Voluntary report | DREAMSTATION RESPIRONICS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I STARTED USING THE DREAMSTATION CPAP MACHINE IN (B)(6) 2017. FROM THEN UNTIL LAST WEEK I WOULD WAKE UP IN THE MORNINGS WITH THROAT AND NASAL IRRITATION VERY OFTEN. I'M (B)(6) 2018 I WAS DIAGNOSED WITH NON-SMALL CELL LUNG CANCER/LARGE CELL AND AM FIGHTING IT STILL. I WAS TOLD IT IS A VERY RARE STRAIN AND THERE'S NO CURE FOR IT. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-06
2518422-2021-01961 | 2021-06-30 | 2021-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01968 | 2021-06-30 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, ERROR CODES RELATED TO THE CHARGING OF THE INTERNAL BATTERY WERE OBSERVED. THE DEVICE'S POWER MANAGEMENT BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-01990 | 2021-06-30 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-06
2518422-2021-02007 | 2021-07-01 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID/HT, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | AUX400T15 | AUX400T15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. SECTION H6 CODES CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT HAS BEEN CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02023 | 2021-07-01 | 2021-01-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO HAVE A MYOCARDIAL INFARCTION AND VENTRICULAR TACHYCARDIA. THE PATIENT DID RECEIVE MEDICAL INTERVENTION AND WAS HOSPITALIZED FOR THE MYOCARDIAL INFARCTION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 11, 2024 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION CAUSED A PATIENT TO HAVE A MYOCARDIAL INFARCTION AND VENTRICULAR TACHYCARDIA. THE PATIENT DID RECEIVE MEDICAL INTERVENTION AND WAS HOSPITALIZED FOR THE MYOCARDIAL INFARCTION. THE REPORTED EVENT OF A MYOCARDIAL INFARCTION ANDVENTRICULAR TACHYCARDIA AND ITS REPORTED SEVERITY RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02027 | 2021-07-01 | 2021-06-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR DISPLAYED A "HIGH TEMPERATURE" ALARM. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER HAD PREVIOUSLY REPORTED THE ALLEGATION OF A HIGH TEMPERATURE ALARM OCCURRING. THE BLOWER MOTOR WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S BLOWER MOTOR WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02012 | 2021-07-01 | 2021-06-24 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION BIPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX600T11C | DSX600T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND RESULTED IN THE PATIENT'S LUNG COLLAPSING. THE DETAILS OF THE MEDICAL INTERVENTION THE PATIENT RECEIVED ARE CURRENTLY UNKNOWN. THE INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND RESULTED IN THE PATIENT'S LUNG COLLAPSING. THE DETAILS OF THE MEDICAL INTERVENTION THE PATIENT RECEIVED ARE CURRENTLY UNKNOWN. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL PART OF THE DEVICE AND FOUND EVIDENCE OF WHITE COLORED TALC-LIKE CONTAMINANT IN THE IN THE DREAMSTATION AIR INLET PORT. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND FOUND NO EVIDENCE OF THE WHITE COLORED TALC-LIKE CONTAMINANT IN THE DREAMSTATION BLOWER MOTOR OR BOX. THE MANUFACTURER FOUND NO EVIDENCE OF THE SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND TO CONTAIN 0 ERROR CODES. THE MANUFACTURER VERIFIED THE UNITS DO POWER-UP, PROVIDE FLOW. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION BUT FOUND CONTAMINANTS IN THE DREAMSTATION AIR INLET PORTS INDICATES THE SOURCE OF THE CONTAMINANTS IS EXTERNAL TO THE DEVICE. ON THE PREVIOUSLY SUBMITTED REPORT, SECTION D4 HAD INCORRECT UNIQUE IDENTIFIER (UDI), WHICH WAS UPDATED/CORRECTED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102277 | 2021-07-01 | 2018-03-29 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS / RESPIRONICS, INC. | DSX500T11C |  |  | I STARTED USING A PHILIPS DREAMSTATION CPAP MACHINE IN (B)(6) 2018. AT MY (B)(6) 2018 APPOINTMENT I QUESTIONED MY SLEEP DOCTOR IF THE MACHINE COULD CAUSE ISSUES, AS I WAS EXPERIENCING SHORTNESS OF BREATH. AT THE TIME MY QUESTION WAS ASSOCIATED TO POSSIBLE LUNG INFECTION RESULTING IN SHORTNESS OF BREATH. THE DOCTOR STATED THE DEVICE COULD NOT DO THIS AND THAT I SHOULD FOLLOW UP WITH PRIMARY CARE PHYSICIAN. I HAVE CONFIRMED THIS CONVERSATION IS DOCUMENTED IN THE DOCTORS VISIT NOTES. IN THE LAST 3.5 YEARS I HAVE HAD PULMONARY FUNCTION TESTING, CARDIAC TESTING, BLOOD/LAB TESTING, ETC CONSTANTLY ALL ATTEMPTING TO IDENTIFY THE REASON FOR THE BREATHLESSNESS. THEN I SAW THE RECALL OF THE PHILIPS DREAMSTATION CPAP DEVICE AND NOTED THE ABATEMENT FOAM RELEASING 'TOXIC AND CARCINOGENIC' GASES. AFTER DISCUSSION WITH MY PULMONARY PHYSICIAN IT HAS BEEN DETERMINED TO IMMEDIATELY STOP USE OF THE PHILIPS DREAMSTATION CPAP TO DETERMINE IF THE SYMPTOMS ARE CONNECTED TO THE DEVICE. I ATTEMPTED TO REPORT THIS ADVERSE EVENT TO PHILIPS AND THE ONLY WAY I WAS ABLE TO PROVIDE ON RECORD WAS A 'REQUEST FOR COMPENSATION' EVEN THOUGH I STATED I WISHED TO REPORT THE ADVERSE EVENTS ASSOCIATED TO USE OF THE DEVICE. THEY TOOK MY NUMBER 3 DAYS AGO, BUT HAVE YET TO FOLLOW UP FOR MORE DETAILS. BOTH MY PULMONARY PHYSICIAN AND I FEEL THERE IS A HIGH LIKELIHOOD MY BREATHLESSNESS IS DIRECTLY CONNECTED TO USE OF THE PHILIPS DREAMSTATION DEVICE GIVING TIMING OF ONCOMING SYMPTOMS AND CONTINUED TESTING WITH NO ASSOCIATED CAUSE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102272 | 2021-07-01 | 2018-05-15 | I, n, j, u, r, y | Voluntary report | PHILIPS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500H11C |  |  | TOO MANY TO DESCRIBE; I HAVE HAD ISSUES FOR YEARS, DIABETIC, KIDNEY DISEASE, STOMACH PROBLEMS, SINUS ISSUES ,RUNNY NOSE, HEADACHES, TROUBLE SLEEPING SORE THROAT, LOST TASTE AND SMELL, LUNG ISSUES, COPD, SWELLING, ANXIETY, WEAKNESS IN LEGS HEART STENT, CHF, AFIB. I HAVE HAD SO MANY TEST DONE THAT I WOULDN'T KNOW WHERE TO BEGAN. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-07
MW5102263 | 2021-07-01 | 2021-06-27 | I, n, j, u, r, y | Voluntary report | DREAMSTATION BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS ELECTRONICS LTD |  |  |  | INCREASING BP, SOB AND SEVERE HEADACHES- PRODUCT RECALLED PHILIPS HAS NOT STARTED REPLACEMENT/REPAIR PROGRAM- REFUSES TO STATE WHEN A NEW DEVICE WILL BE SUPPLIED. I HAVE BEEN WITHOUT PHILIPS BIPAP MACHINE FOR OVER TWO WEEKS- INCREASED BP, SEVERE HEADACHES, SHORTNESS OF BREATH. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-07
MW5102243 | 2021-07-01 | 2021-05-03 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. | ? REFERENCE #. DSX500H11C |  |  | CHEST PRESSURE OCCURRED AT NIGHT (B)(6) 2021, LIKE SOMEONE WAS SITTING ON MY CHEST; WAS NAUSEATED, COULD NOT LAY DOWN OR PRESSURE WOULD INCREASE. WENT TO LOCAL ER AT (B)(6) HOSPITAL. WAS ADMITTED FOR CARDIAC OBSERVATION; NUMEROUS TESTS DONE. ALL NEGATIVE. PRESSURE CONTINUED, FINALLY LESSENING TOWARD 5PM, BUT STILL THERE. NEVER HAD ANY EXPERIENCE LIKE THIS IN MY LIFE. DISCHARGED WITH CELEBREX, TOOK 3 OVER NEXT 2 DAYS. HAVE BEEN USING AN AUTO CPAP, SERIAL NUMBER#(B)(4), SINCE 2018 AND A PURIFY O3 SANITIZER MODEL# 190-6000, FOR 2YEARS I THINK. I RECEIVED AN ALERT FROM THE PHARMACY THAT SUPPLIES MY EQUIPMENT REGARDING A RECALL FOR MY MACHINE BY PHILIPS RESPIRONICS, HENCE MY REPORT. HAD APPOINTMENT WITH MY PULMONOLOGIST, ALREADY SCHEDULED, BEFORE THIS EVENT. HAD PULMONARY FUNCTION TESTS WHICH CORRELATED WITH TESTS I HAD DONE A YEAR AGO. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-07
MW5102279 | 2021-07-01 | 2017-06-01 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS DREAM STATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC |  |  |  | I STARTED USING A PHILIPS DREAMSTATION CPAP MACHINE IN APPROXIMATELY (B)(6) OF 2016. BY 2017 I NOTICED THAT MY NOSE AND SINUS PASSAGES WERE ALWAYS SWOLLEN AND IRRITATED AND I HAD AN OCCASIONAL WHEEZE WHEN BREATHING. I SAW THE RECALL FOR THE DREAMSTATION CPAP A FEW WEEKS AGO AND READ THAT MY SYMPTOMS ALONG WITH INCREASED RISK OF CANCER WERE THE POSSIBLE SIDE EFFECTS FROM THIS PRODUCT. I HAVE USED THIS PRODUCT EVERY DAY FOR OVER 5 YEARS AND I AM VERY CONCERNED WITH MY LONG TERM HEALTH BECAUSE OF IT. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102233 | 2021-07-01 | 2021-06-25 | I, n, j, u, r, y | Voluntary report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | PHILIPS / RESPIRONICS, INC. |  |  |  | WE REPORTED TO (B)(6) 1ST THAT WE WERE UNABLE TO BLEED OXYGEN THROUGH MY TRILOGY 100. THEN REPORTED THE MACHINE DIDN'T FEEL RIGHT AND I KEPT HAVING OXYGEN DESATS. HEALTHCARE SOLUTIONS BLAMED ANXIETY FOR THE DESATS AND SAID TO USE MY NASAL CANNULA WITH THE DEVICE WHICH TAKES AWAY THE EFFECTIVENESS. A COUPLE DAYS LATER MY DAUGHTER HAD TO CALL 911 FOR DECREASED MENTATION, AND I WAS HOSPITALIZED WITH INFLAMMATION OF AIRWAY, HIGH CO2 AND LOW BODY PH. THEN I HAD A SEIZURE AT THE ICU AT (B)(6) HOSPITALS AND HAD TO BE INTUBATED AND TRANSFERRED TO (B)(6) HOSPITAL FOR SEIZURE MONITORING. I HAVE NO HISTORY OF SEIZURES. THANKFULLY I WAS ABLE TO BE EXTUBATED TODAY, BUT AM STILL IN CRITICAL CONDITION AT ICU. (B)(6) HAS BEEN VERY RUDE AND THEY DIDN'T NOTIFY ME OF THE TRILOGY 100 RECALL, EVEN KNOWING I WAS HAVING ISSUES WITH THE DEVICE. PLEASE LEAVE ME AN EMAIL AND I WILL SEND OVER ALL LABS AND IMAGING RESULTS. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | 0 |  | 2021-07
MW5102230 | 2021-07-01 | 2021-06-01 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | J309794839ECB |  |  | VESTIBULAR NEURITIS; STARTED WITH SEVERE HEADACHES, NASAL CONGESTION, SINUS PAIN, VERTIGO, DIZZINESS, SINUS INFECTION. REQUIRING MULTIPLE DR APPOINTMENTS. ANTIBIOTICS, STEROIDS X2 MISSED WORK. UNABLE TO DRIVE ACTIVE RECALL ON CPAP MACHINE WITH GAS EXPOSURE LISTING ABOVE SYMPTOMS. STARTED USING THIS DEVICE 1 MONTH PRIOR TO THE ABOVE ISSUES. HAVE BEEN FIGHTING THESE SYMPTOMS FOR 25 DAYS AND WORSE IN MORNING AFTER CPAP USE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102229 | 2021-07-01 | 2021-03-31 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | J28279845817 |  |  | I NOTICED THE BLACK PARTICLES WHEN I WAS WASHING THE HOSE OF MY PHILIPS RESPIRONICS CPAP MACHINE. I HAVE ALSO EXPERIENCED CHEST PAIN, SINUS PRESSURE AND DIZZINESS UPON WAKING WHICH I HAVE BEEN ADVISED ARE SYMPTOMS OF BEING EXPOSED TO THE CARCINOGENIC MATERIAL. I HAVE BEEN ADVISED TO CEASE USE OF MY CPAP MACHINE BUT I HAVE SEVERE SLEEP APNEA SO I AM WAITING FOR FURTHER INSTRUCTIONS FROM MY DOCTORS OFFICE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102270 | 2021-07-01 |  | I, n, j, u, r, y | Voluntary report | PHILLIPS RESPIRONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS / RESPIRONICS, INC. | DSX900T11C |  |  | EXTREME HEADACHES, RESPIRATORY AND INFLAMMATION ISSUES, HYPERSENSITIVITY, LUNG CANCER 1 NODULE RIGHT LUNG THAT ONCOLOGIST CAN'T DETERMINE. IRRITATION AND CANNOT SLEEP. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-07
2518422-2021-02014 | 2021-07-01 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | CNX700T17 | CNX700T17 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND A PATIENT TO DEVELOP CANCER CELLS IN THE BLOOD. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A COMPREHENSIVE EXAMINATION AND LAB TESTING. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO THE DEVICE'S SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02026 | 2021-07-01 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID/HT, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | AUX400T15 | AUX400T15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | SECTION B5 WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT SHOULD BE REPORTED AS FOLLOWS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION PATIENT HAVEING PATIENT COUGHING, EYE IRRITATION AND SINUS ISSUES WHEN USING THE DEVICE. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. SECTION H6 CODES CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT HAS BEEN CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102235 | 2021-07-01 |  | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAMSTATION BIPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/RESPIRONICS, INC. |  |  |  | WHY IS THERE NO INFORMATION ON THE PHILLIPS RESPIRONICS RECALL ON YOUR WEBSITE? THIS RECALL AFFECTS MANY AND MOST OF US HAVE RECEIVED VERY LIMITED INFO AND I CERTAINLY DIDN'T GET ANYTHING FROM THE MANUFACTURER. I WANT TO KNOW MORE ABOUT THE FILTER AND VOC COMPOUND GASES RELEASED FROM THE FOAM. DOES NO ONE TEST FOR THIS BEFORE APPROVAL? HOW DO I KNOW THAT RESMED DOESN'T HAVE VOC AND FILTER PROBLEMS ALSO? DO WE NEED LUNG X-RAYS AND LONG TERM MONITORING NOW? MY DOCTORS OFFICE SENT EMAIL WITH THE COMPANY'S RECALL NOTICE BUT REALLY NOT WHAT I SHOULD DO EXCEPT REGISTER MY MACHINE WITH THE COMPANY FOR THE RECALL AND THE COMPANY SHOULD ALREADY HAVE HAD THAT INFO ACTUALLY. WHERE DOES THAT LEAVE EVERYONE WHO USED THE MACHINES? IS THE FDA LOOKING INTO THIS? FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02041 | 2021-07-02 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AVAPS C SERIES | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | 1161T | 1161T |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING EVALUATION OF THE DEVICE AT A THIRD-PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02042 | 2021-07-02 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AVAPS C SERIES | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | 1161T | 1161T |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING EVALUATION OF THE DEVICE AT A THIRD-PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02031 | 2021-07-02 | 2021-06-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR A HIGH PRESSURE CONDITION AND THE DEVICE POWERED OFF. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR WAS ALARMING FOR A HIGH INSPIRATORY PRESSURE ALARM CONDITION AND POWERED OFF. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02034 | 2021-07-02 | 2021-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | THE MANUFACTURER RECEIVED INFORMATION FROM A THIRD PARTY SERVICE CENTER, ALLEGING AN ISSUE WITH A VENTILATOR'S OXYGEN BLENDING MODULE BOARD. THE COMPONENT ONLY WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE COMPONENT, FAILED TEST STEPS DURING TESTING WERE OBSERVED. THE ROOT CAUSE OF THE FAILURE WAS FOUND TO BE CAUSED BY A FAULTY AIRFLOW SENSOR ON THE OXYGEN BLENDING MODULE BOARD. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
MW5102302 | 2021-07-02 | 2019-06-01 | I, n, j, u, r, y | Voluntary report | DREAM STATION CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | COMPLAINANT CALLED TO REPORT SHE HAS BEEN HAVING RESPIRATORY ISSUES FOR THE LAST TWO YEARS. SHE RECENTLY HAD A CHEST XRAY AND HER LUNGS ARE CLEAR BUT SHE HAS DEVELOPED ASTHMA. SHE HAS BEEN USING A CPAP MACHINE SINCE 2016 AND HAD ISSUES WITH HER CURRENT DREAM STATION CPAP MACHINE. SHE HAS NOTICED BLACK SPECS IN THE HOSES AND LATER SAW THAT PHILIPS RESPIRONICS INITIATED A RECALL OF THIS DEVICE DUE TO THE PE-PUR FOAM WHICH MAY DEGRADE INTO PARTICLES WHICH MAY ENTER THE DEVICE VIA THE AIR PATHWAY AND BE INGESTED OR INHALED BY THE USER AND THE PE-PUR FOAM MAY OFF-GAS CERTAIN CHEMICALS. THIS INFORMATION WAS OBTAINED FROM THE RECALL NOTICE ISSUED BY THE MANUFACTURER. HER HUSBAND HAS THE SAME MACHINE BUT HE NOTICED LESS SPECKS AND SHE BELIEVES THE DIFFERENCE MAY BE BECAUSE SHE USED A SOCLEAN MACHINE AFTER HAVING RESPIRATORY ISSUES.(WILL ENTER COMPLAINT FOR SO CLEAN MACHINE ALSO). SHE NEED TO HAVE HER MACHINE REPLACED AS SOON AS POSSIBLE BUT HER INSURANCE COMPANY WILL NOT APPROVE THE REPLACEMENT. ASTHMA AND OTHER RESPIRATORY ISSUES BELIEVED TO BE CAUSED BY RECALLED DREAM STATION CPAP. DIFFICULTY BREATHING, SYSTEM AFFECTED: RESPIRATORY, ONSET TIME: MONTHS, DURATION: PERSISTS, DEVELOPED ASTHMA. RECALLED RES 88058. |  | Y | Y | I |  | 2021-07
2518422-2021-02033 | 2021-07-02 | 2021-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR POWERED OFF WHILE IN PATIENT USE. THE PATIENT BECAME BRADYCARDIC AND THEIR BLOOD OXYGEN SATURATION DECREASED. THE PATIENT WAS MANUALLY VENTILATED AND RECOVERED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR POWERED OFF WHILE IN PATIENT USE. THE PATIENT BECAME BRADYCARDIC AND THEIR BLOOD OXYGEN SATURATION DECREASED. THE PATIENT WAS MANUALLY VENTILATED AND RECOVERED. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. DURING THE TESTING OF THE VENTILATOR THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS FOUND TO BE FAULTY AND WILL NEED TO BE REPLACED TO ADDRESS THE ISSUE. THE DEVICE MAINTAINS THE EVENT LOGS IN NON-VOLATILE MEMORY. AN EVENT, AS LOGGED BY THE DEVICE, INCLUDES EVERY KEYPRESS, ALARM OR FAILURE OF THE DEVICE TO REMAIN WITHIN SPECIFICATION. A REVIEW OF THE DEVICE'S DOWNLOADED EVENT LOG CONFIRMS THE DEVICE DID NOT SPONTANEOUSLY POWER OFF WHILE IN PATIENT USE. | Y | Y | HEALTH PROFESSIONAL |  | 2021-07
MW5102301 | 2021-07-02 |  | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILLIPS RESPIRONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500H11 |  |  | DREAM STATION CPAP RECALL; HAVE BEEN COUGHING FOR SEVERAL MONTHS AND ALSO SINUS DRAINAGE FOR MONTHS. HAVE BEEN USING ONE OF THE RECALLED CPAP MACHINES. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | I |  | 2021-07
MW5102299 | 2021-07-02 | 2021-06-30 | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. | DSX400T11C | DSX400T11C |  | I WAS USING THE "SOCLEAN" DEVICE TO CLEAN MY CPAP UNTIL I NOTICED AN UNPLEASANT ODOR AND TASTE IN MY DREAMSTATION CPAP WISP NASAL CUSHION AND TUBING. I STOPPED USING THE SOCLEAN AND RETURNED TO SOAP AND WATER TO CLEAN MY CPAP, WHICH HAS NOW BEEN RECALLED BY "PHILIPS RESPIRONICS". MY DOCTORS OFFICE SENT ME AN EMAIL ALERTING ME OF THE RECALL. I REGISTERED MY CPAP MACHINE ON THE WEBSITE TODAY-(B)(6) 2021. I AM WONDERING IF THE "SOCLEAN" WILL ALSO ISSUE A RECALL? I EXPERIENCED HEADACHES AND SINUS ISSUES ON AND OFF WHILE USING THE "SOCLEAN". IT MAY BE WHY OUR CPAP MACHINES START DETERIORATING. I HAVE "OBSTRUCTIVE SLEEP APNEA" DIAGNOSED IN 2012. I HAVE REACHED OUT TO MY DOCTOR FOR ADVICE ON CONTINUING TO USE MY CPAP DURING THE PHILIPS RESPIRONICS RECALL. I AM VERY CONCERNED THAT PERMANENT LUNG DAMAGE OR CANCER CAN OCCUR DUE TO THIS PROBLEM. I HOPE THEY REPLACE OUR CPAP QUICKLY TO AVOID PROBLEMS CONTINUING OR PEOPLE DISCONTINUING USE OF THEIR CPAP AND HAVING OBSTRUCTIVE SLEEP APNEA TAKE THEIR LIVES. (B)(6). FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102298 | 2021-07-02 | 2017-08-10 | I, n, j, u, r, y | Voluntary report | SYSTEM ONE PHILIPS RESPIRONICS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | REM STAR AUTO |  |  | HEADACHES, STOMACH ACHES, AND SINUS PRESSURE ISSUES. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02035 | 2021-07-02 | 2021-06-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02032 | 2021-07-02 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS A "SERVICE REQUIRED" ALARM CONDITION. THE DEVICE'S OXYGEN BLENDING MANIFOLD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02030 | 2021-07-02 | 2021-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-01986 | 2021-07-02 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, INTL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | INX500S15 | INX500S15 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING THE EVALUATION OF THE DEVICE AT A THIRD-PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO THE DEVICE'S SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-01985 | 2021-07-02 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11 | DSX500T11 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING THE EVALUATION OF THE DEVICE AT A THIRD-PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO THE DEVICE'S SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02005 | 2021-07-02 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED JUNE 30, 2021 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING SYMPTOMS OF COUGH, HEADACHE, AND THROAT IRRITATION RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGED LIGHTHEADED FEELING AND DIZZINESS. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN THIS REPORT, SECTION H6 HEALTH EFFECT - CLINICAL CODE HAS BEEN UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-01779 | 2021-07-02 | 2021-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX400H11C | DSX400H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP A LUNG NODULE. THE DETAILS OF THE MEDICAL INTERVENTION THE PATIENT HAS RECEIVED ARE CURRENTLY UNKNOWN. THE MANUFACTURER HAS ATTEMPTED TO ARRANGE FOR THE RETURN OF THE DEVICE FOR EVALUATION TO THE MANUFACTURER'S PRODUCT INVESTIGATION LAB. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 06 , 2024 AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED LUNG NODULE. THE REPORTED EVENT OF LUNG NODULE AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE EVENT IS CURRENTLY UNKNOWN.. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02015 | 2021-07-02 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUMID | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11 | DSX700H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED ON THIS DEVICE IN MDR 2518422-2021-01895. THIS REPORT WAS SUBMITTED AS A DUPLICATE REPORT OF THE PREVIOUSLY SUBMITTED REPORT. SECTION D1, D4 AND H4 HAS BEEN CORRECTED/UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02013 | 2021-07-02 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DEVICE GETTING LOUDER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02011 | 2021-07-02 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO THE DEVICE'S SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02022 | 2021-07-02 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, JP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | RJPX500S16 | RJPX500S16 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE REPORTER ALSO ALLEGED BLACK POWDER CAME OUT FROM THE DEVICE. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL PART OF THE DEVICE AND FOUND DUST CONTAMINATION IN THE HOSE CONNECTION PORT. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND FOUND NO EVIDENCE OF FOAM DEGRADATION OR BLACK PARTICLES. HEAVY DUST CONTAMINATION WAS OBSERVED IN THE BLOWER AND ESPECIALLY ON THE BLOWER IMPELLER. UPON OPENING THE SEALED BAGS CONTAINING THE MASK AND TUBING, A HEAVY MILDEW ODOR WAS OBSERVED. THE MANUFACTURER THEN EXAMINED THE 15MM TUBING RETURNED WITH THE UNIT, FINDING NO EVIDENCE OF BLACK PARTICULATE MATTER. LIGHT DIRT AND FIBER BUILDUP WAS OBSERVED IN BOTH CONNECTION ENDS, BUT NOTHING WAS OBSERVED IN THE TUBING. CLOSE EXAMINATION OF THE MASK DID CONFIRM THE PRESENCE OF DARK/BLACK CONTAMINATION, AS WELL AS LINT/FIBER CONTAMINATION, AND A LIGHTER COLORED CONTAMINANT. THE MANUFACTURER EXAMINED ALL CONTAMINANTS UNDER A MICROSCOPE, CONCLUDING THE DARK CONTAMINANT TO BE ORGANIC IN NATURE AND MOST LIKELY MOLD. THE MANUFACTURER WAS UNABLE TO IDENTIFY THE LIGHT-COLORED CONTAMINANT WITH ANY CERTAINTY, BUT A BODY FLUID SEEMS LIKELY. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CAN CONFIRM DARK CONTAMINATION WAS PRESENT IN THE MASK, MOST LIKELY MOLD AND A LIGHT-COLORED CONTAMINANT ON THE MASK. THE MANUFACTURER WAS ALSO ABLE TO CONFIRM DIRT/LINT CONTAMINATION IN PORTIONS OF THE AIRPATH, WITH LACK OF A POLLEN FILTER DURING OPERATION AS THE MOST LIKELY SOURCE OF THIS CONTAMINANT. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION BUT CONTAMINATIONS THROUGHOUT THE DEVICE THAT ARE FROM AN EXTERNAL SOURCE. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-01777 | 2021-07-02 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM ALLEGEDLY BECAME DEGRADED. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT ALSO ALLEGED HAVING DIZZINESS, IRRITATION AND BURNING EYE AND NASAL AND THROAT ISSUES, HEADACHES. THE DEVICE WAS RETURNED TO THE MANUFACTURER QUALITY PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. DURING THE INITIAL VISUAL INSPECTION THE MANUFACTURER, THE DEVICE WAS DISASSEMBLED FOR INTERNAL VISUAL INSPECTION. THE MANUFACTURER FOUND NO SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WITHIN THIS DEVICE. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND TO CONTAIN 3 ERROR CODES. THE MANUFACTURER CONCLUDES THERE IS NO EVIDENCE OF SOUND ABATEMENT FOAM BREAKDOWN WITHIN THE DEVICE. SECTION H6 HAS BEEN UPDATED/CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02024 | 2021-07-02 | 2021-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02006 | 2021-07-06 | 2021-06-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT ATTHIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102317 | 2021-07-06 | 2021-04-15 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS / RESPIRONICS |  |  |  | MY PHILLIPS DREAMSTATION CPAP EMITTED STRONG ODORS WHILE USING IT AT NIGHT. I NOTICED THIS ON SEVERAL OCCASIONS, I WOULD REMOVE IT AND THE BEDROOM SMELLED NORMAL. THE NEXT MORNING I WOKE WITH A SEVERE HEADACHE WHICH LASTED MOST OF THE DAY. THIS WENT ON FOR MONTHS. I THINK THE CPAP MACHINE HAS HAD AN ADVERSE HEALTH EFFECT ON ME. I HAVE REGISTERED MY CPAP WITH PHILLIPS USING THE RECALL REGISTRATION SITE, WHICH DID VERIFY MY CPAP WAS RECALLED. SINCE CPAP THERAPY IS A LIFELONG TREATMENT THAT CANNOT JUST BE STOPPED AS RECOMMENDED BY PHILLIPS I AND MANY OTHERS ARE LEFT TO FEND FOR OURSELVES. SO FAR I HAVE NOT RECEIVED ANY COMMUNICATION FROM PHILLIPS ABOUT THIS INCREDIBLY URGENT ISSUE. CAN THE FDA SOMEHOW MAKE PHILLIPS REPLACE OUR DEFECTIVE DEVICES IMMEDIATELY. I HOPE THAT THE FDA CAN FORCE SOME ACTION TO GET THE PROCESS MOVING. AND PLEASE INVESTIGATE HOW PHILLIPS KNEW ABOUT THIS AS EARLY AS APRIL 2021 AND DIDN'T NOTIFY THE PUBLIC UNTIL MID JUNE 2021. THAT IS A LONG TIME TO BE BREATHING TOXIC FUMES UNNECESSARILY. THIS MATTER IS A GREAT CONCERN TO MYSELF AND PROBABLY MANY OTHERS. THANK YOU (B)(6). FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-01807 | 2021-07-06 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/HT/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX400T11C | DSX400T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 06, 2024, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED ALLEGATION OF GENERAL SICKNESS, BURNING SENSATION IN THROAT, INFLAMMATORY RESPONSE, ASTHMA, PULMONARY DAMAGE (UNSPECIFIED). THIS RECORD WAS REVIEWED BY THE PMS CLINICAL EXPERT DUE TO THE CUSTOMER REPORTING NO OTHER RELEVANT CLINICAL INFORMATION WAS PROVIDED. BASED ON INFORMATION AVAILABLE AT THE TIME OF THIS REVIEW, THERE IS SUFFICIENT EVIDENCE TO PRONOUNCE A SERIOUS INJURY THE PATIENT DID NOT REPORT TO RECEIVE MEDICAL INTERVENTION. SECTION B1 WAS CORRECTED FOR BOTH ADVERSE EVENT AND PRODUCT PROBLEM. (PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION. TO SERIOUS INJURY. SECTION H6 WAS CORRECTED AND UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02040 | 2021-07-06 | 2020-01-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP CANCER. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF CHEMOTHERAPY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP CANCER. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF CHEMOTHERAPY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02045 | 2021-07-06 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | RGBX500S15 | RGBX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO SERIOUS OR PERMANENT INJURY. DESPITE MULTIPLE ATTEMPTS, THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102337 | 2021-07-06 |  | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAMSTATIONS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. | DSX500T11W |  |  | I HAVE A PHILIPS DREAMSTATION CPAP MACHINE. WHEN I CALLED MY PROVIDER, THEY TOLD ME THE PRODUCT WAS UNDER RECALL AND SHOULD STOP USING IT. I CALLED MY NEUROLOGIST AND HE TOLD ME TO STOP USING THE MACHINE, BUT HAD NO OTHER SUGGESTIONS. I HAVE USED A CPAP MACHINE FOR 9 YEARS AND HAVE A SIGNIFICANT TROUBLE SLEEPING WITHOUT IT. I CANNOT FIND ANYONE, INCLUDING YOUR OFFICE, WHO CAN TELL ME WHEN I CAN GET A REPLACEMENT OR WHAT I SHOULD EXPECT IN TERMS OF RESOLVING THE PROBLEM. I WOULD AT LEAST LIKE INFORMATION BESIDE THE USELESS, GENERIC INFO ON THE PHILIPS WEBSITE. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | I |  | 2021-07
MW5102336 | 2021-07-06 | 2020-10-08 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS DREAM STATION BIPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. | DSX700 |  |  | FIRST OF THREE SEVERE SINUS INFECTIONS, AFTER STARTING SLEEP THERAPY WITH PHILIPS BIPAP FOR SLEEP APNEA. SECOND INFECTION (B)(6) 2020, THIRD INFECTION (B)(6) 2021. CT SCAN OF SINUS (B)(6) 2021, AND SINUS SURGERY (B)(6) 2021. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102328 | 2021-07-06 | 2019-11-24 | I, n, j, u, r, y | Voluntary report | REMSTAR PRO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | 1005960 |  |  | RESPIRATORY; I HAD STROKE SYMPTOMS WHICH TURNED OUT TO BE TIA, ENDED UP WITH STAGE 4 RENAL FAILURE, AND HAD TO BE PUT ONTO BREATHING TREATMENTS TO GET MY LUNGS CORRECTED FROM A LUNG INFECTION. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102326 | 2021-07-06 | 2021-04-22 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAM STATION AUTO-CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500H11C |  |  | PHILIPS RESPIRONIC DREAM STATION STARTED MAKING HIGH PITCHED NOISE AND INCREASING THE PRESSURE FROM 7 PSI TO OVER 20 PSI RAPIDLY. THIS CONTINUED TO HAPPEN NIGHTLY. THE NORMAL PSI FOR MY SLEEP APNEA CORRECTION IS 7-8 PSI. THE RAISING PSI WOKE ME UP AND LEFT ME WITH HEADACHES. THE NOISE CAUSED MY HUSBAND TO WAKE ME UP SEVERAL TIMES TO CHECK MY MACHINE. I USED MY DREAM STATION GO WHILE I DETERMINED HOW TO FIX THE ISSUE. I ORDERED NEW FILTERS INCASE THEY GENERIC ONES WERE NOT SITTING IN PLACE CORRECTLY. ORDERING PHILIPS RESPIRONIC FILTERS WAS DIFFICULT TO FIND. THE FILTERS DID NOT SOLVE THE NOISE OR THE PRESSURE ISSUE. I LOOKED UP HOW TO SET THE MIN AND MAX PRESSURES. I SET THE MAX PRESSURE TO 11PSI TO PREVENT THE SLEEP DISRUPTION. THE HEADACHES CONTINUED OFF AND ON. I USE MY MACHINE NIGHTLY FOR 7-8 PLUS HOURS WITHOUT FAIL. THE MACHINE STARTED MAKING THE NOISE AGAIN IN (B)(6) AND HITS THE MAX PSI SET OF 11 AND STAYS AT 11PSI UNTIL IT WAKES ME UP AND I PUSH THE RESET TO STARTING PSI. LOSING SLEEP AND HAVING HEADACHES HAS NOT BEEN THE NORM UNTIL RECENTLY SINCE THE FIRST MONTHS OF USE. I WAS SEARCHING FOR A REASON WHY THE AWFUL NOISE AND WILDLY HIGH PRESSURES WAS HAPPENING BACK IN (B)(6) AND FOUND NOTHING. TODAY, I FOUND THE RECALL. I DON'T KNOW IF THE ISSUES WITH MY MACHINE ARE LINKED TO THE RECALL OR NOT. I DO NOT USE ANYTHING OTHER THAN SOAP AND WATER TO CLEAN MY MACHINES. THE MACHINE WAS PURCHASED IN (B)(6) 2018 AND CAUSED SEVERE HEADACHES TO THE POINT I STOPPED USING IT AND WAITED TO SEE A SPECIALIST. THE SPECIALIST DID NOT HAVE ANY SOLID ANSWERS FOR ME OTHER THAN "YOU HAVE SLEEP APNEA, YOU NEED TO USING THE MACHINE". THE TEST RESULT BELOW IS FROM THE SLEEP STUDY. THERE IS A MEDICAL NEED FOR MY CPAP MACHINE TO WORK AND NOT MALFUNCTION. A MEDICAL DEVICE SHOULD NOT BE CAUSING ME MORE HARM THAN GOOD. GIVEN THE NOW FOUND RECALL, I HAVE STOPPED USING MY DREAM STATION AND MY DREAM STATION GO. I HAVE REGISTERED MY MACHINES WITH PHILIPS PER THE RECALL. GIVEN THE NOISE ISSUE AND PRESSURE INCREASING ISSUE, ANY FILTER "FIXES" ARE NOT APPLICABLE; MY ONLY OPTION IS TO BUY A NEW MACHINE. WHILE I WAIT FOR INSTRUCTIONS FROM PHILIPS OR THE FDA, I ORDERED A DIFFERENT BRAND OF MACHINE. I WAS DIAGNOSED WITH SLEEP APNEA IN (B)(6) 2018; 80 SLEEP APNEA EVENTS IN 320 MINUTES WERE NOTED. AHI OF 7 AND RDI OF 15. THERE WERE REDUCED OXYGEN SATURATION RATES LESS THAN 88% WITH LOWEST AT 80%. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102325 | 2021-07-06 | 2021-02-02 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS /RESPIRONICS, INC. |  |  |  | NOTICED AN ODOR AT FIRST ONSET OF USING CPAP NIGHTLY. MUCH DIFFERENT EXPERIENCE THAN WAS ACCUSTOMED TO DURING PREVIOUS YEARS OF USE. ALSO NOTICED WHITE GRANULAR PARTICLES BUILDING UP ON THE CHARCOAL FILTER SO MUCH SO THAT THE AIR WAS NOT GETTING THROUGH THE CHARCOAL FILTER TO THE DISPOSABLE FILTER. COULD NOT USE A DISPOSABLE FILTER DUE TO LACK OF AIR FLOW. THIS CONTINUED ON AND BECAME THE NEW NORMAL IN USING MY DREAM STATION. I KNEW SOMETHING WAS NOT QUITE RIGHT AND MENTIONED IT TO MY PULMONOLOGIST ON (B)(6) 2021 DURING A ROUTINE FOLLOW UP VISIT. HE NEVER MENTIONED ANYTHING ABOUT A RECALL AT THAT TIME BUT DID INFORM ME THAT IF THE SITUATION WITH THE CPAP MACHINE CONTINUES TO BE BOTHERSOME TO LET HIM KNOW. NOW, TODAY (B)(6) 2021 I AM INFORMED BY MY MEDICAL EQUIPMENT SUPPLY PROVIDER DURING A PHONE CALL TO REORDER SUPPLIES THAT THE MAKER OF MY CPAP MACHINE HAS A MAJOR RECALL. WILL THIS UNFAMILIAR ODOR AND PARTICLES NOW CAUSE ME HEALTH ISSUES? HOW WILL I BE ABLE TO USE MY CPAP MACHINE NIGHTLY WHEN IT IS NOW A HEATH RISK? SOMETHING THAT IS SUPPOSED TO HELP ME NIGHTLY HAS NOW BECOME A POTENTIAL HAZARD. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102334 | 2021-07-06 | 2021-06-14 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS / RESPIRONICS INC. | DSX500H11C |  |  | I AM A USER OF THE PHILLIPS DREAMSTATION C-PAP MACHINE AND WAS TOLD BY PHILLIPS TO DISCONTINUE USE OF THE MACHINE AND NOT GIVEN ANY DIRECTION AS FAR AS WHAT TO DO UNTIL THE SITUATION WAS CORRECTED. THIS IS A THERAPY STARTED 8 YEARS AGO BY PRESCRIPTION FROM MY PULMONARY DR. TO HELP ME SLEEP THROUGH THE NIGHT. THIS IS LIKE SOMEONE TAKING AWAY YOUR WHEELCHAIR AND TELLING YOU THERE IS NOTHING THEY CAN SUBSTITUTE FOR YOU. I HAVE LOST MANY NIGHTS OF SLEEP DUE TO THIS RECALL. PLEASE HELP. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-01996 | 2021-07-06 | 2021-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLANK. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS EVALUATED AT THE MANUFACTURER'S SERVICE CENTER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02047 | 2021-07-06 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP S/T C SERIES, INTL 30 | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS INC | IN1061X | IN1061X |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A BI-LEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED BRONCHITIS. DESPITE MULTIPLE ATTEMPTS, THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. CODES IN SECTION H6 WERE UPDATED OR CORRECTED AND UNIQUE IDENTIFIER (UDI) # IN SECTION D4 WAS INCORRECTLY ADDED IN THE INITIAL REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02056 | 2021-07-06 | 2021-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S TIDAL VOLUME WAS INCREASING. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SENSOR BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02057 | 2021-07-06 | 2021-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE WITH A VENTILATOR'S PRESSURE DELIVERY WAS OBSERVED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02058 | 2021-07-06 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02059 | 2021-07-06 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02060 | 2021-07-06 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG AND THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS CLEANED TO ADDRESS THE ISSUE. THERE WAS EVIDENCE OF CONTAMINATION. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02061 | 2021-07-06 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02062 | 2021-07-06 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A VENTILATORI INOPERTATIVE CODE WAS OBSERVED IN THE DOWNLOADED EVENT LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02063 | 2021-07-06 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02065 | 2021-07-06 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02064 | 2021-07-06 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02069 | 2021-07-06 | 2021-06-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02070 | 2021-07-06 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02071 | 2021-07-06 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02072 | 2021-07-06 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02073 | 2021-07-06 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG AND THE LCD WAS OBSERVED TO BE BLANK. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND LCD DISPLAY WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-01974 | 2021-07-06 | 2021-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATORS TIDAL VOLUME WAS INCREASING. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE, FLOW SENSOR, OVERHEAD BLOWER FILTER AND SENSOR BOARD WERE REPLACED TO ADDRESS THE ISSUE. THERE WAS EVIDENCE OF CONTAMINATION. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02004 | 2021-07-06 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP PRO, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | 660P | 660P |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING HEART PROBLEM AND A BREATHING ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 HEALTH EFFECT- CLINICAL CODE.TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02068 | 2021-07-06 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S KEYPAD WAS OBSERVED. THE DEVICES FRONT PANEL BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
MW5102316 | 2021-07-06 |  | I, n, j, u, r, y | Voluntary report | PHILLIPS DREAMSTATION 1 | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS RESPIRONICS, INC. |  |  |  | RESTRONIC PHILLIPS DREAMSTATION 1. I HAVE DEVELOPED CHRONIC KIDNEY DISEASE STAGE 3, PNEUMONIA LAST 6 MONTHS, SPOTS ON LUNGS LAST 2 YEARS. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | I |  | 2021-07
MW5102348 | 2021-07-07 | 2021-06-14 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAM MAKER CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | SEVERE EAR PAIN WITH CANAL SWELLING COMPLETELY SHUT. STILL DEALING WITH ISSUES TODAY (B)(6) 2021. |  | Y | N | I |  | 2021-07
MW5102373 | 2021-07-07 | 2021-06-18 | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS / RESPIRONICS, INC. |  |  |  | CHEST PAIN, CONGESTION, BRONCHITIS, ASTHMA FLARE UPS, CHRONIC COUGH, BREATHLESSNESS, MUCUS, CONTINUED ILLNESS AFTER USING RESPIRONICS DREAMSTATION AT HIGH HUMIDITY LEVELS. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-07
MW5102369 | 2021-07-07 | 2021-06-05 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C |  |  | HAD STARTED NOTICING A STRANGE ODOR THAT SMELLED LIKE BURNT FIBERGLASS FOR A COUPLE OF WEEKS BUT DID NOT ASSOCIATE IT WITH MY CPAP. A FEW DAYS LATER I WOKE GASPING FOR AIR. I WAS BREATHING BUT FELT LIKE I WAS SUFFOCATING. IT SEEMED TO TAKE FOREVER BEFORE IT FELT LIKE I WAS GETTING AIR AGAIN. PROBABLY 10 OR 15 MINUTES. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102349 | 2021-07-07 | 2021-04-12 | I, n, j, u, r, y | Voluntary report | RESPIRONICS BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | DSX700H11C |  |  | HAVE A PHILIPS RESPIRONICS BIPAP THAT IS ON RECALL. I HAVE HAD CHRONIC SINUS CONGESTION WITH PERIODIC NASAL BLEEDING FOR GOING APPROACHING 3 MONTHS. NO MEDICATIONS HAVE HELPED THAT I HAVE BEEN USING EITHER PRESCRIBED OR OTC. I HAVE REGISTERED MY MACHINE WITH PHILIPS RESPIRONICS AND DISCONTINUED USING BIPAP AS OF LAST NIGHT WHEN (B)(6) ADVISED AT APPOINTMENT YESTERDAY DURING ANNUAL PULMONARY FOLLOW UP. THIS SUBMISSION IS IN RESPONSE TO INSTRUCTION RECEIVED FROM VA ABOUT RECALL ON MY VERIFIED BIPAP. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102354 | 2021-07-07 | 2021-05-07 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION BI-LEVEL PAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS /RESPIRONICS INC | PHILIPS DREAMSTATION BI-LEVEL P |  |  | I STARTED URINATING BLOOD ON (B)(6) 2021,AND SAW MY UROLOGIST DR. (B)(6) THE NEXT DAY. THE DOCTOR ORDERED TESTS, BUT ON (B)(6) 2021 I WENT TO THE EMERGENCY ROOM AND WAS ADMITTED TO THE HOSPITAL DUE TO EXCESSIVE BLEEDING, AND A BLOOD CLOT STOPPED MY URINE FLOW. ON (B)(6) 2021 I WAS DIAGNOSED WITH CANCER IN MY RIGHT KIDNEY. I HAD A 10 CENTIMETER BY 7 CENTIMETER CANCEROUS GROWTH. I HAD ROBOTIC SURGERY TO REMOVE MY ENTIRE RIGHT KIDNEY ON (B)(6) 2021. I WAS DISCHARGED FROM (B)(6) ON (B)(6) 2021. I AM CURRENTLY BEING TREATED FOR MY CANCER BY ONCOLOGIST DR. (B)(6). I HAVE BEEN USING A PHILIPS DREAMSTATION BI-LEVEL PAP SINCE 2015, AND BECAME AWARE OF THE RECALL 2 DAYS AFTER MY SURGERY. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-07
MW5102377 | 2021-07-07 |  | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAMSTATION 1 CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS / RESPIRONICS, INC. |  |  |  | I'VE BEEN USING THE DREAMSTATION 1 CPAP MACHINE. NO ONE HAS BOTHERED TO REACH OUT AND NOTIFY ME OR OTHERS OF THE SO CALLED RECALL ON THESE MACHINES WHICH IS CAUSING NUMEROUS ISSUES IN PEOPLE. NOTHING IS EVEN CURRENTLY BEING DONE FAST ENOUGH TO RECTIFY THIS SITUATION EITHER. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102358 | 2021-07-07 |  | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500T11C | SEE PHOTOS OF LABEL |  | I LIVE IN A HUMID CLIMATE AND HAVE BEEN USING THE PHILLIPS DREAMSTATION CPAP MACHINE SINCE (B)(6) 2016. IN A NORMAL FOLLOW UP VISIT FOR PFT WITH MY PULMONOLOGIST ON (B)(6) I HAD COMPLAINED OF DIZZINESS, COUGH, AND WHAT FELT LIKE A SINUS INFECTION. I HAD NO FEVER BUT SOME OF THESE SYMPTOMS PERSISTED FOR QUITE SOME TIME (EVEN YEARS BEFORE) OFF AND ON. MOST RECENTLY I HAD BEEN FEELING FATIGUED EVEN WITH NORMAL CPAP USE. HE SAID AFTER EXAMINING I HAD BLOOD ON TOP OF MY DONE IN BOTH EARS, AND GAVE ME A ZPAK. I HAD NOT HEARD OF THE RECALL YET NOR DID HE INFORM ME WHILE I WAS IN THE OFFICE. I PROMPTLY STARTED THE MEDICATION THAT DAY STILL FEELING SYMPTOMS. THE NEXT MORNING I ACTUALLY CALLED THE OFFICE AND ASKED THE STAFF AGAIN IF THERE WAS ANY WAY TO CLEAN THE INSIDE PIPES OF THE UNIT AS I SUSPECTED IT COULD BE SOMETHING TRAPPED INSIDE. I NEVER LIKED THAT THERE WAS NO WAY TO DO THIS. THE LADY SAID THERE WAS ACTUALLY A RECALL ON MY CPAP BUT ONLY FOR PEOPLE THAT USED SOCLEAN OZONE DEVICES AND SENT ME A LINK TO A FORM ON EMAIL THAT DIDN'T WORK WHEN I CLICKED ON IT. I CONTACTED MY INSURANCE AFTERWARDS, REFERRED TO THEIR SUPPLIER, THEN TOLD BY (B)(6) TO CONTACT BAYCARE FOR MY SPECIAL SITUATION AND SHE COULD NOT HELP ME AFTER SPEAKING TO HER MANAGER. BAYCARE HAD BEEN A HORRIBLE EXPERIENCE FOR ME IN THE PAST AS THEY FRAUDULENTLY SENT ME EQUIPMENT I NEVER ASKED FOR AND CHARGED MY MRA EVEN AFTER ITEMS I DIDN'T WANT WERE RETURNED. SOME PEOPLE WERE RUDE AND ONE I RECALL EVEN COMPLAINED ABOUT THEIR EMPLOYER TO ME PATIENT. I DID NOT GET A HOLD OF ANYONE BEFORE THEY CLOSED AND TODAY IS SATURDAY. I'M ON DAY TWO OF MEDS, STILL DRY COUGH, SOME BREATHING ISSUES, AND DIZZINESS. LAST NIGHT I WAS TOLD BY A FRIEND THAT I SHOULD LOOK INTO THE RECALL AS HE HEARD IT COULD CAUSE LUNG CANCER. THIS MORNING (SATURDAY) I FOUND THIS FDA WARNING WHICH MAKES ME ANGRY THAT MY DOCTOR DID NOT TELL ME ABOUT IT, AND WHY DIDN'T HE TELL ME TO IMMEDIATELY STOP USING THE MACHINE AND FIND AN ALTERNATIVE KNOWING MY SYMPTOMS DIRECTLY LINK TO THESE PROBLEMS I WAS EXPERIENCING. HAD I NOT READ YOUR ARTICLE I WOULD HAVE CONTINUED USE THE CPAP MORE. I WILL TRY YOUR LINK BUT I NEED A CPAP THAT WON'T CONTINUE TO MAKE ME ILL. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102362 | 2021-07-07 | 2017-10-16 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS DREAMSTATION CPAP DEVICE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS / RESPIRONICS, INC. |  |  |  | ONE YEAR AFTER BEGINNING CPAP THERAPY USING THE PHILIPS RESPIRONICS DREAMSTATION CPAP DEVICE, I WAS DIAGNOSED WITH NON-SMALL CELL LUNG CANCER. I AM CURRENTLY A PATIENT AT (B)(6). MEDICAL RECORDS AVAILABLE UPON REQUEST. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102367 | 2021-07-07 |  | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500T11C |  |  | I HAVE HAD WORSENING CONGESTION, COUGH, DIZZINESS AND FOUND OUR ON (B)(6) 2021 THAT THE CPAP DREAMSTATION THAT I HAVE USED SINCE (B)(6) 2019 HAS BEEN RECALLED DUE TO BAD FOAM WHICH CAN CAUSE HEALTH PROBLEMS. I HAVE REGISTERED THE DEVICE ON PHILIPS RECALLED WEBPAGE AND HAVE HAD NO RESPONSE FROM THE COMPANY. YOU WOULD LIKE MY MACHINE REPLACED WITH ANOTHER BRAND AND/OR FULL REFUND. IF MY HEALTH PROBLEMS CONTINUE I WOULD LIKE PHILIPS TO PAY FOR TREATMENT. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102371 | 2021-07-07 | 2021-06-18 | I, n, j, u, r, y | Voluntary report | DREAMSTATION GO PORTABLE AUTO CPAP SYSTEM WITH HEATED HUMIDIFIER AND DREAMWEAR F | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS | DSG500S11 | DSG500S11 |  | ON (B)(6) 2021, I PURCHASED FOR OVER (B)(6) AND BEGAN USING A DREAMSTATION GO AUTO-ADJUSTING CPAP MACHINE FOR SEVERE SLEEP APNEA DIAGNOSED A FEW DAYS EARLIER. FROM THE BEGINNING OF USE, I WOKE UP WITH A RUNNY NOSE, CONGESTION, SORE THROAT, AND A FEELING OF DISCOMFORT IN MY LUNGS THAT FELT AS IF MY LUNGS WERE RAW. AFTER BEING AWAKE AN HOUR OR TWO, THE SYMPTOMS RESOLVED, SO I KNEW I DID NOT HAVE A COLD. SINCE I HAD NEVER USED A CPAP BEFORE, I THOUGHT THESE SYMPTOMS WERE DUE TO BEING UNUSED TO CPAP TREATMENT. ON (B)(6) 2021, I FOUND A SMALL BIT OF BLACK DEBRIS IN MY MASK. ON (B)(6) 2021, I RECEIVED WORD THAT PHILLIPS RESPIRONICS HAD RECALLED MY DREAMSTATION GO BECAUSE THE PEPUR SOUND ABATEMENT FOAM RELEASED HARMFUL GASES AND COULD DECOMPOSE UNDER NORMAL USE, ALLOWING DEBRIS IN THE AIR PATH. I STOPPED USING IT IMMEDIATELY. A FEW DAYS LATER, I PURCHASED AN AUTO-ADJUSTING CPAP BY ANOTHER COMPANY. SINCE THEN, I HAVE EXPERIENCED NONE OF THE PREVIOUS SYMPTOMS. THE RUNNY NOSE, CONGESTION, SORE THROAT, AND LUNG DISCOMFORT I HAD EXPERIENCED FOR OVER TWO MONTHS WAS COMPLETELY GONE. I BELIEVE THOSE SYMPTOMS WERE CAUSED BY THE FOAM IN THE DREAMSTATION GO. I HAVE REGISTERED MY DREAMSTATION GO IN PHILIPS RESPIRONICS' RECALL, BUT HAVE RECEIVED NO FURTHER COMMUNICATION WITH THEM. I BELIEVE THIS VOLUNTARY RECALL SHOULD BE TURNED INTO A MANDATORY ONE AND THAT PHILIPS RESPIRONICS SHOULD BE REQUIRED TO REFUND THE COST OF THE MACHINES AND ACCESSORIES WHICH WILL ONLY WORK WITH THOSE MACHINES. I HAVE A PRODUCT LESS THAN THREE MONTHS OLD, WHICH IS DANGEROUS TO USE, ALONG WITH A HUMIDIFIER AND TRAVEL CASE THAT WILL ONLY WORK WITH THIS MACHINE. PHILIPS RESPIRONICS HAS GIVEN US NO INDICATION OF WHEN THEY WILL TAKE CARE OF THIS, FORCING ME TO PURCHASE A NEW MACHINE TO PROTECT MY HEALTH. I AM ALSO CONCERNED BECAUSE OF THE STATED RISK OF CANCER DUE TO EXPOSURE TO THIS PE-PUR FOAM. I HAVE A FAMILY HISTORY OF CANCER AND WOULD HAVE NEVER EXPOSED MYSELF TO THIS RISK WILLINGLY. PLEASE MAKE PHILIPS RESPIRONICS DO WHAT IS RIGHT AND PAY US BACK FOR WHAT WE'VE WASTED BECAUSE OF THEM AND SET ASIDE FUNDS TO TREAT ANY CONDITIONS THAT COME OUT OF THIS. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102364 | 2021-07-07 | 2021-05-28 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS RESPIRONICS / RESPIRONICS, INC. | DSX500H11C |  |  | APPROXIMATELY 6 WEEKS AGO AT THE END OF (B)(6) 2021, I NOTICED AND DEVELOPED A CHRONIC COUGH WHILE USING THE DREAM STATION SLEEP APNEA MACHINE AND ALSO AFTER CLEANING IT WITH THE SOCLEAN MACHINE. THE COUGH HAS PERSISTED AND HAS NOW AFFECTED MY LUNGS, AND BREATHING AND I HAVE REQUIRED MEDICAL ATTENTION. I HAVE NOT RESOLVED THIS IN SPITE OF ANTIBIOTICS, AN INHALER AND A STEROID PACK. I WILL BE PURSUING FURTHER MEDICAL ATTENTION NEXT WEEK. DISCOID ATELECTASIS LEFT LUNG, CONSTRICTIVE AIRWAY PER PHYSICIAN. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102372 | 2021-07-07 | 2018-07-01 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION 1 | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DREAMSTNAUTOCPAPHUMCELL DOM |  |  | GALL BLADDER REMOVAL; I WAS GETTING DIZZY AND VERY UNCOORDINATED. I ALSO LOST MY HEARING ON THE LEFT SIDE. FOUND OUT I HAD A VERY FAST GROWING BRAIN TUMOR (ACOUSTIC NEUROMA) THAT HAD TO BE REMOVED. RECENTLY, MY LIVER HAS STARTED TO DETERIORATE AND I AM HAVING GALL BLADDER ISSUES, (B)(6) 2021 DATE FOR REMOVAL. NOW WHILE THE ACOUSTIC NEUROMA MIGHT NOT HAVE BEEN CAUSED BY THE DREAMSTATION 1 I HAD BEEN USING SINCE 2017, THE FAST GROWTH OF THE TUMOR COULD HAVE BEEN ATTRIBUTED. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102374 | 2021-07-07 |  | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | CPAP RECALLED. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | I |  | 2021-07
MW5102383 | 2021-07-07 | 2020-12-20 | I, n, j, u, r, y | Voluntary report | PHIIIPS RESPIRONIC BIPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | STARTING (B)(6) 2020, THEN TWO MORE TIMES, HAD SEVERE UNEXPLAINED NAUSEA. EACH TIME. I STAYED IN BED FOR THE DAY AND USED OVER-THE-COUNTER MEDICATIONS TO RELIEVE THE NAUSEA. |  | Y | N | LAY USER/PATIENT |  | 2021-07
MW5102386 | 2021-07-07 | 2020-03-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION 1 | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | 1 |  |  | STARTED HAVING NOSE BLEEDS AND NASAL CONGESTION AROUND THIS DATE. STOPPED USING THE DREAMSTATION CPAP AND BOUGHT A RESMED CPAP. PROBLEM IS GETTING BETTER. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102387 | 2021-07-07 | 2021-07-05 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS RESPIRONICS SYSTEM 1 BIPAP | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | PHILIPS / RESPIRONICS, INC. |  |  |  | PHILIPS RESPIRONICS BIPAP, PARTICLES ARE VISIBLE. COMPANY SAID TO STOP USING, BUT OFFERED NO REMEDY. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102388 | 2021-07-07 | 2021-05-24 | I, n, j, u, r, y | Voluntary report | REMSTAR PRO C-FLEX + PHILIPS RESPIRONICS CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. |  |  |  | UPPER RESPIRATORY INFECTION; I WAS USING THE REMSTAR PRO C-FLEX PLUS PHILIPS RESPIRONICS C-PAP MACHINE (SN (B)(4)) AND CAUGHT A REALLY BAD COLD WITH CONGESTION IN MY SINUSES AND THROAT, WHICH LASTED FOR ABOUT FIVE WEEKS. I WAS SO CONGESTED THAT I COULD NOT USE MY CPAP MACHINE BECAUSE I COULD NOT BREATHE THROUGH MY NOSE. ONCE MY SINUSES CLEARED UP, I STARTED USING THE MACHINE AND THE FIRST NIGHT I USED IT, I DEVELOPED CONGESTION AGAIN. I STOPPED USING THE MACHINE ONCE I LEARNED OF THE RECALL. MY THROAT AND SINUSES CLEARED UP. I RECENTLY LEARNED OF THE MACHINE'S RECALL BY ITS MANUFACTURER (PHILIPS), SO I AM NOT USING THE MACHINE AT ALL RIGHT NOW BECAUSE I AM CONCERNED ABOUT THE SOUND ABATEMENT FOAM IN THE MACHINE. PHILIPS HAS NOT INDICATED WHEN OR HOW THE FOAM IN THE MACHINE WILL BE REPLACED. I AM WRITING WITH HOPES THAT THE FDA CAN PRESSURE PHILIPS TO EXPEDITE THE PROCESS, AS I HAVE SEVERE SLEEP APNEA AND NEED THE MACHINE TO SLEEP MORE COMFORTABLY AND BETTER MANAGE MY BLOOD PRESSURE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102392 | 2021-07-07 | 2021-06-01 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPERONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | MY PHILIPS CPAP HAS MADE MY THROAT STAY SORE AND MY NOSE BLEED PERIODICALLY APPROXIMATELY TWO WEEKS BEFORE MY DOCTOR APPOINTMENT ON OR AROUND STARTING AROUND (B)(6) 2021. THE MAJORITY OF THE TIME, I HAVE A PERIODIC ONGOING COUGH AND SORE THROAT. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102393 | 2021-07-07 | 2020-09-01 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. | DREAMSTATION AUTO CPAPHUMCELL DO |  |  | I HAVE A PHILIPS DREAM STATION CPAP THAT HAS BEEN LISTED IN THE RECALL. MACHINE # (B)(4). I HAVE NOTICED LAST YEAR I WAS HAVING MORE PROBLEMS WITH NASAL, THROAT AND LUNG CONGESTION IN THE MORNING, AFTER USING THE CPAP. IT HAS BEEN GETTING WORSE OVER THE LAST FEW MONTHS AND I WOULD HAVE HEADACHES AND FELT LOUSY AFTER USING THE MACHINE. I THOUGHT IT WAS RELATED TO ALLERGIES OR THE MASK SLIPPING OFF AT NIGHT. SINCE I HAVE STOPPED USING THE MACHINE AND THE HEADACHES AND LOUSY FEELING HAVE STOPPED, AND THE CONGESTION IN THE MORNING IS MUCH LESS. I AM USING AN OLD RESMED MACHINE UNTIL I GET A REPLACEMENT. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02044 | 2021-07-07 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO SERIES ASSY, INT | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | IN751S | IN751S |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING THE EVALUATION OF THE DEVICE AT A THIRD-PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102357 | 2021-07-07 | 2021-05-15 | I, n, j, u, r, y | Voluntary report | PHILLIPS RESPIRONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500T11C |  |  | USING CPAP MACHINE THAT WAS ISSUED BY THE VA, I AM NOW SUFFERING FROM DEGRADED FOAM EXPOSURE. I HAVE NAUSEA, MIGRAINES, SINUS PROBLEMS, ITCHY EYES, WATERING EYES, SHORTNESS OF BREATH. I FEEL LIKE MY LUNGS ARE AT 1/3 CAPACITY FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102363 | 2021-07-07 | 2021-04-01 | I, n, j, u, r, y | Voluntary report | PHILLIPS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS RESPIRONICS / RESPIRONICS, INC. |  |  |  | I'VE BEEN HAVING FREQUENT HEADACHES WHICH HAVE NEVER HAPPENED BEFORE. NOTICED WHEN IT WAS TIME TO CHANGE DREAMSTATION FILTERS THEY LOOK DARK GRAY. I CHECKED ON PHILLIPS WEBSITE TODAY AND MY MACHINE IS ONE EFFECTED. I HAVE USED A CLEANER YOU HOOK TO THE END OF HOSE BUT I DID NOT KNOW. I'VE HAD THIS UNIT FOR OVER A YEAR. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02098 | 2021-07-07 | 2021-04-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | CAX400H12 | CAX400H12 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. CORRECTED DATA: B3 DATE OF EVENT: PREVIOUSLY REPORTED AS 06/09/2021 ; UPDATED TO (B)(6)2021 D8 WAS THIS DEVICE SERVICED BY A THIRD PARTY?: PREVIOUSLY REPORTED AS NO ; UPDATED TO UNKNOWN E1 INITIAL REPORTER CONTACT (NAME) & ADDRESS E4 INITIAL REPORTER ALSO SENT REPORT TO FDA?: PREVIOUSLY REPORTED AS NO ; UPDATED TO UNKNOWN G2 REPORT SOURCE: PREVIOUSLY REPORTED AS COMPANY REPRESENTATIVE ; UPDATED TO FOREIGN/USER FACILITY | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102356 | 2021-07-07 | 2020-09-07 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO CPAP WITH HUM DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS RESPIRONICS / RESPIRONICS, INC. |  |  |  | STARTED HAVING HEADACHES, THROAT IRRITATION, AND COUGHING IN THE MORNINGS AFTER USING MY CPAP MACHINE (DREAMSTATION) MADE BY PHILLIPS RESPIRONICS. DID NOT GO TO THE DOCTOR AS COVID RESTRICTIONS WERE IN PLACE AND IT WAS DIFFICULT TO GET APPOINTMENTS. ALSO I WAS ABLE TO MANAGE THE SYMPTOMS ON MY OWN. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02049 | 2021-07-07 | 2021-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, A LOW CIRCUIT LEAK ALARM OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE FAILED A TEST STEP FOR FLOW DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02054 | 2021-07-07 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02048 | 2021-07-07 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. THE MACHINE FLOW SENSOR WAS FOUND TO BE CONTAMINATED WITH DIRT AND DUST. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02053 | 2021-07-07 | 2021-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02055 | 2021-07-07 | 2021-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR A HIGH PRIORITY ALARM CONDITION. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR WAS ALARMING FOR A HIGH PRIORITY ALARM CONDITION. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. A "VENTILATOR INOPERATIVE" ALARM CONDITION OCCURRED. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. THE DEVICE HAD EVIDENCE OF TOBACCO CONTAMINATION THROUGHOUT THE DEVICE. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02050 | 2021-07-07 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02052 | 2021-07-07 | 2021-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02051 | 2021-07-07 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02093 | 2021-07-07 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02092 | 2021-07-07 | 2021-06-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED A TEST STEP DURING TESTING. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR FAILED A TEST STEP DURING TESTING. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. A "VENTILATOR INOPERATIVE" ALARM CONDITION OCCURRED. THE DEVICE'S OXYGEN BLENDING MODULE MANIFOLD, OXYGEN BLENDING MODULE FLOW ELEMENT, OXYGEN BLENDING MODULE MIXING ELEMENT WERE REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02091 | 2021-07-07 | 2021-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | JP2100X16B | JP2100X16B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S TOUCHSCREEN FAILED TO RESPOND. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN EVALUATED BY THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR'S TOUCHSCREEN FAILED TO RESPOND. THE VENTILATOR WAS NOT RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT COULD NOT BE CONFIRMED. THE DEVICE PERFORMED TO MANUFACTURER'S SPECIFICATIONS. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02025 | 2021-07-07 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, INTL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | INX500S15 | INX500S15 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING THE EVALUATION OF THE DEVICE AT A THIRD-PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO THE DEVICE'S SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02088 | 2021-07-07 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR | RESPIRONICS INC | RJPX500S16 | RJPX500S16 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S POWER CORD HAD EXPOSURE TO THE AC CONDUCTING WIRING. THERE WAS NO PATIENT HARM OR INJURY REPORTED. NO SERIAL NUMBER OR DATE OF MANUFACTURE FOR THE DEVICE IS AVAILABLE. THE DEVICE IS NOT BEING RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. |  | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102359 | 2021-07-07 | 2021-05-15 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILLIPS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS / RESPIRONICS, INC. |  |  |  | I NOTICED BLACK SPECKS FLOATING IN MY WATER TANK OF MY DREAMSTATION MACHINE. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02075 | 2021-07-07 | 2021-06-25 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A RESPIRATORY INFECTION AND ASTHMA. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A DOCTOR'S ASSESSMENT AND A PRESCRIBED STEROID, ANTIBIOTICS AND INHALER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING COUGHING BLOOD, CAUSED A PATIENT TO DEVELOP A RESPIRATORY INFECTION AND ASTHMA, BLEEDING FROM THE LUNGS, HEADACHES AND VOMITING UP BLOOD. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A DOCTOR'S ASSESSMENT AND A PRESCRIBED STEROID, ANTIBIOTICS AND INHALER. IN SECTION H6 HEALTH EFFECT - CLINICAL CODE HAS BEEN CORRECTED / UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02076 | 2021-07-07 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AUX500T15 | AUX500T15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. B5 SHOULD HAVE BEEN REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED TO EXPERIENCING WAKING UP WITH COLD BODY TEMP, TREMORS, SOMETIMES PAIN IN THE STOMACH AND CHEST. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. SECTION H6 WAS UPDATED WITH THE APPROPRIATE HEALTH EFFECT - CLINICAL CODES AND THE TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAVE BEEN CORRECTED TO REFLECT THAT THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02078 | 2021-07-07 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11 | DSX500H11 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102385 | 2021-07-07 | 2020-01-01 | I, n, j, u, r, y | Voluntary report | SO CLEAN.2 | DISINFECTANT, MEDICAL DEVICES | SOCLEAN, INC. | SC1200 | 187293000860 |  | WHILE USING MY RESPIRONICS DREAMSTATION CPAP NIGHTLY FOR THE PAST YEAR I HAD MANY NIGHTLY COUGHING SPELLS THROUGHOUT MY DISRUPTED SLEEP CYCLES WHILE USING THE CPAP. I HAD USED THE SO-CLEAN OZONE SANITATION DEVICE DAILY PRIOR TO USE OF THE CPAP. ALTHOUGH THIS COMPLAINT FORM ONLY ALLOWS A SINGLE DATE MY ISSUES HAVE BEEN WELL OVER A YEARS TIME. I ONLY FOUND OUT ABOUT THE FDA ALERT AND THE PHILLIPS CPAP RECALL NOTICE THIS WEEK. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02037 | 2021-07-07 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | RENTAL-DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | RGBX500S15 | RGBX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO THE DEVICE'S SOUND ABATE FOAM.THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102411 | 2021-07-08 | 2021-07-02 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. |  |  |  | EMERGENCY ROOM VISIT SEVERE HEADACHE ETC. AFTER LEARNING OF THE RECALL MY DOCTOR ADVISED TO CALL THE EQUIPMENT DISTRIBUTOR IN WHICH OUR DISTRIBUTOR ADVISED THEY HAVE BEEN TELLING PEOPLE TO CALL THEIR DOCTOR. THEY ADVISED TO CONTINUE USE BUT DO NOT USE OZONE CLEANER. I HAVE ONLY BEEN USING FOR A FEW MONTHS, BUT I HAVE CONSTANT HEADACHES EXCESSIVE PHLEGM AND HAVE WEIRD EXHAUST LIKE SMELL AND TASTE IN THE MORNINGS. (B)(6) WHAT THE EMERGENCY ROOM PRESUMED AS STOMACH VIRUS. I HAD SEVERE MIGRAINE LIKE HEADACHE, 102.8 FEVER, CHEST PAIN AND BURNING, EARS THROAT HURTING AND BURNING, AND SWOLLEN LYMPH NODES IN MY NECK. POSITIVE FOR STAPHYLOCOCCUS. FDA SAFETY REPORTED ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02087 | 2021-07-08 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SENSOR BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02090 | 2021-07-08 | 2021-06-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054097 | 1054097 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S FLOW SENSOR FAILED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02086 | 2021-07-08 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, PHYSICAL DAMAGE WAS OBSERVED BY THE TECHNICIAN. THE DEVICE'S PORTING BLOCK ADAPTOR AND OXYGEN BLENDING MODULE INLET ADAPTOR WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02085 | 2021-07-08 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02089 | 2021-07-08 | 2021-06-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040000 | 1040000 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WOULD NOT FUNCTION. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON FEB 24, 2025, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A VENTILATOR WOULD NOT FUNCTION. THERE WAS NO HARM OR INJURY REPORTED. AFTER THE FIRST ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02114 | 2021-07-08 | 2021-06-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2100X15B | IN2100X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, A VENTIALTOR FAILED A CALIBRATION. THE DEVICE'S OXGYEN BLENDING MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
MW5102426 | 2021-07-08 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS RESPIRONICS DREAM MAKER CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DREAMMAKER |  |  | DIRT AND DEBRIS IN AIR MASK AND TUBING OF CPAP MACHINE. MACHINE HAS BEEN RECALLED FOR FAILURE OF FOAM NOISE CONTROL FOAM. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102422 | 2021-07-08 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500T11C |  |  | RECEIVED RECALL NOTICE FROM PHILIPS ABOUT THE CDPAP DEVICE I WAS PRESCRIBED AS BEING RECALL, AS THE SOUND ABATEMENT FOAM AND SOME BACTERIA MAY BREAK OFF DURING USE AND CAUSE PERMANENT INJURY. THE NOTICE FURTHER TELLS ME TO GO TO MY DOCTOR AND DISCUSS WHETHER I SHOULD CONTINUE USE, BASED ON MY CLINICAL NEEDS/SITUATION. MY DOCTOR WILL NOT CONTACT ME, AND INSTEAD - SENT AN EMAIL THAT CLOSELY RESEMBLES THE SUGGESTED COMMUNICATION FOR HEALTHCARE PROVIDERS ON PHILIPS' WEBSITE ADDRESS, DEDICATED TO THE RECALL. THE FDA ACKNOWLEDGED THIS ISSUE FORMALLY ON JUNE AND STILL NO RECOURSE OR WAY TO RECTIFY THE SITUATION, OTHER THAN GETTING ANOTHER CDPAP MACHINE (THESE COST (B)(6) AND INSURANCE WILL NOT PAY FOR ONE THAT IS LESS THAN 5 YEARS OLD). THIS IS AN OUTRAGE - I NEED THIS DEVICE TO SLEEP, OR I STOP BREATHING, CAUSING CARDIAC ISSUES (HENCE THE REASON FOR IT). PHILIPS HAS PUT THE ONUS ON THE PHYSICIAN TO MAKE A CLINICAL CALL AS TO THE RISK OF USING THE MACHINE AS IS, AND THE HEALTHCARE PROVIDERS ARE POINTING RIGHT BACK AT PHILIPS. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | I |  | 2021-07
MW5102412 | 2021-07-08 | 2021-07-06 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | DSX400H11W |  |  | PHILIPS CPAP SN (B)(4) HAS BLACK PARTICLES COMING OUT OF THE UNIT. IT IS PART OF THE RECALL. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102407 | 2021-07-08 |  | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I UNDERSTAND THAT THERE IS A RECALL OUT FOR PHILIPS CPAP DREAMSTATION. THE MACHINES ARE FOUND TO BREAK DOWN RELEASE SMALL FOAM PARTIALS WITHIN THE DEVICE. THIS BREAKDOWN CAN BE WORSE IF A OZONE CLEANING DEVICE IS USED. I HAVE ALWAYS USED THE (B)(6) TO SANITIZE MY CPAP. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02038 | 2021-07-08 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | INX500S15 | INX500S15 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING EVALUATION OF THE DEVICE AT A THIRD-PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO THE DEVICE'S SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102403 | 2021-07-08 | 2017-01-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I HAVE ASTHMA COPD LUPUS MIGRAINES. APPROXIMATELY A YEAR AFTER STARTING TO USE THE MACHINE, I STATED WAKING UP WITH MUCUS IN MY THROAT AND CHEST TIGHTNESS. I HAS TO START TAKING 2 PUFFS OF VENTALIN TO HELP. SOMETIMES IT WORKED BUT MOSTLY DID NOT. I STARTED TO HAVE SWALLOWING ISSUES LAST YEAR. I HAD AN ENDOSCOPY IN THE PAST FEW MONTHS AND HAD TO HAVE MY ESOPHAGUS WIDENED THOUGH I DIDN'T WHY. IT HELPED BUT LAST MONTH, THE CHOKING BEGAN AGAIN. I HAVE REQUESTED A CHANGE IN MACHINES BUT MY DOCTOR HASN'T APPROVED. TO DATE, MY MEDICAL SUPPLIES COMPANY HAS DONE NOTHING. I AM (B)(6) AND IN DETERIORATING HEALTH. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02095 | 2021-07-08 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED DUE TO AGE. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02003 | 2021-07-08 | 2020-10-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED, CAUSING AN ALLERGIC REACTION. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A DOCTOR'S ASSESSMENT AND PRESCRIPTIONS FOR STEROIDS AND ANTIBIOTICS. THE PATIENT DID NOT PROVIDE DETAILS ON THE SERIAL NUMBER OF THE DEVICE OR CONTACT INFORMATION. THE MANUFACTURER IS UNABLE TO ARRANGE FOR THE RETURN OF THE DEVICE FOR EVALUATION DUE TO A LACK OF INFORMATION PROVIDED BY THE PATIENT. IF FURTHER INFORMATION REGARDING THE DEVICE BECOMES AVAILABLE, IT WILL BE SUBMITTED IN THE FOLLOW-UP REPORT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 12, 2024. SECTION H11 SHOULD BE REPORTED AS THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND ALLERGIC REACTION. THE PATIENT REPORT TO RECEIVE MEDICAL INTERVENTION IN THE FORM OF A DOCTOR'S ASSESMENT AND PRESCRIPTIONS FOR STEROID AND ANTIBIOTIC. THERE IS NO REPORT OF PATIENT HARM OR INJURY. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM (ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR) SECTION B2 WAS CHANGED TO BLANK (PREVIOUSLY IT WAS OTHER SERIOUS) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION SECTION H6 HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02066 | 2021-07-08 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | CAX500T12 | CAX500T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02028 | 2021-07-08 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/HT/CELL, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | AUX400T15C | AUX400T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | SECTION B5 WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT SHOULD BE REPORTED AS FOLLOWS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION PATIENT HAVING PATIENT COUGHING, EYE IRRITATION AND SINUS ISSUES WHEN USING THE DEVICE. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. SECTION H6 CODES CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT HAS BEEN CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02074 | 2021-07-08 | 2021-06-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500H11C | DSX500H11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02094 | 2021-07-08 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING SINUS INFECTION RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE - THE PATIENT ALLEGED SINUS INFECTION. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED AND CORRECTED CLINICAL CODE IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02101 | 2021-07-08 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP AVAPS | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | DSX1130H11 | DSX1130H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT REFUSED TO PROVIDE DETAILS ON CONTACT INFORMATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED DHEADACHE AND VISUALIZATION OF BLACK PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. ADDITIONAL INFORMATION WAS RECEIVED AND ADDED TO THE REPORT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER ON 05/01/2023 FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED ON 05/26/2023 . THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED EXTERNAL AND INTERNALLY BROWN SUBSTANCE STAINS ON THE INSIDE OF THE FRONT OF THE ENCLOSURE AT THE RIM AND OUTSIDE OF THE ENCLOSURE BACK DOOR PANELS AND UNKNOWN TINY BLACK AND WHITE CONTAMINATION DUST LIKE INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM ALSO OBSERVED ON HTE BOTTOM OF THE ENCLOSURE WITH SOME POTENTIAL BLACK HAIRS OR FIBERS. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THEY CONFIRM THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM AND THERE IS NO EVIDENCE OF DAMAGE OF THE DEVICE WHICH SUGGEST THAT THE SOURCE OF CONTAMINATION WAS EXTERNAL TO THE SOURCE AND THEY CAN CONFIRM THE BLACK TINY PARTICLES CONTAMINATION BUT NO CONTAMINATION CONSISTNET WITH DEGRADED FOAM. SECTION D8, D9, H2, H3 AND H6 WERE UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02081 | 2021-07-08 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR HIGH INTERNAL OXYGEN. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02100 | 2021-07-08 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN AIRWAY FROM DEVICE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. AFTER THE SECOND ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, CUSTOMER DECLINED TO RESPOND TO THE GFE RELATED QUESTIONS AND TERMINATED THE CALL. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02099 | 2021-07-08 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES, DIFFICULTY BREATHING AND DEPRESSION RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM.THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. THE MANUFACTURER PREVIOUSLY NOT CAPTURED, VISUALIZATION OF PARTICLES,DIFFICULTY BREATHING AND DEPRESSION IN SECTION B5 AND IT HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102405 | 2021-07-08 | 2020-07-25 | I, n, j, u, r, y | Voluntary report | DREAMSTN AUTO CPA | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500T11C |  |  | RESPIRATORY INFECTIONS AND SINUS INFECTION THAT RECURRED FOR MORE THAN A YEAR. ENT PRESCRIBED FLONASE, WHICH DID NOT IMPROVE RECURRENCES. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02046 | 2021-07-08 | 2021-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A FAILURE OF ITS INTERNAL BATTERY. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, ERROR CODES RELATED TO THE CHARGING OF THE INTERNAL BATTERY WERE OBSERVED. THE DEVICE'S INTERNAL BATTERY AND POWER MANAGEMENT BOARD NEED TO BE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
MW5102430 | 2021-07-08 | 2020-12-31 | I, n, j, u, r, y | Voluntary report | DREAM SM AUTO CPAP HUM CU-DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/RESPIRONICS INC. | DS500H11C | DNX500H11C |  | REPORTER STATED THAT HE HAS BEEN USING HIS DREAM SM AUTO CPAP SINCE (B)(6) 2020. SOMETIMES IN (B)(6) 2020 HE STARTS COUGHING OUT BLOOD DURING USE OF THE MACHINE. HE CALLED THE MANUFACTURER ON (B)(6) 2021 AFTER HEARING ABOUT THE RECALL OF THE DEVICE. HE GAVE THE MODEL NUMBER OF THE MACHINE AND WAS TOLD HIS DEVICE IS ONE OF THE RECALLS. HE WAS ADVISED TO CALL FDA AND GIVE A REPORT. |  | N | Y | I |  | 2021-07
2518422-2021-02043 | 2021-07-08 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AVAPS C SERIES INTL 30, FLOW GEN | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS INC | 1161T | 1161T |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING THE EVALUATION OF THE DEVICE AT A THIRD-PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02104 | 2021-07-08 | 2021-06-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02102 | 2021-07-08 | 2021-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | CA1032800B | CA1032800B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED A TEST STEP DURING TESTING. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. THE VENTILATOR'S INTERNAL TUBING WAS FOUND TO BE PINCHED, WHICH WOULD CAUSE THE DEVICE TO FAIL TESTING. THE DEVICE'S TUBING KIT WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR FAILED A TEST STEP DURING TESTING. THE DEVICE'S TUBING KIT WAS FOUND TO BE PINCHED AND WAS REPLACED TO ADDRESS THE ISSUE. DURING ADDITIONAL EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS AGAIN DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE NEEDS TO BE REPLACED TO ADDRESS THE ISSUES. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02082 | 2021-07-08 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS FAILING LEAK TESTING. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BLOWER CHASSIS PLATE WAS REPLACED TO ADDRESS THE ISSUE. AN ISSUE RELATED TO THE TOUCHSCREEN WAS OBSERVED. THE DEVICE'S DISPLAY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02103 | 2021-07-08 | 2021-06-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | KO1032800 | KO1032800 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A POSITIVE END-EXPIRATORY PRESSURE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A POSITIVE END-EXPIRATORY PRESSURE MALFUNCTION. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02079 | 2021-07-08 | 2021-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040000 | 1040000 |  | ON SITE REMOTE TECHNICIAN SERVICE WAS REQUESTED FOR A DEVICE. THE OXYGEN BLENDING MODULE BOARD WAS REPLACED. THERE WAS NO HARM OR INJURY REPORTED. THE OXYGEN BLENDING MODULE BOARD WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. DURING THE INVESTIGATION, THE BOARD FAILED STEPS DURING TESTING. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02083 | 2021-07-08 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02084 | 2021-07-08 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
MW5102419 | 2021-07-08 | 2017-05-17 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILPS / RESPIRONICS, INC. | J18590103151A |  |  | I HAD 2 STROKES IN 2017 AND PUT ON A PHILIPS DREAMSTATION BIPAP. I HAVE HAD SHORTNESS OF BREATH AND A SQUEEZING HEAVY CHEST SENSATION SINCE. BREATHING THROUGH THE MASK HAD A PECULIAR SMELL WHICH I THOUGHT WAS THE PLASTIC FROM THE NEW MASK. AFTER THE RECALL MY DME COMPANY ATTACHED A FILTER TO THE MACHINE BETWEEN THE MACHINE AND TUBING WHICH HAS TAKEN THE SMELL AWAY AS A TEMPORARY FIX WAITING FOR A PERMANENT SOLUTION. I DO NOT FEEL CONFIDENT IN PHILIPS REMEDY AND WANT A REPLACEMENT RESMED MACHINE. AS MY MACHINE IS NOT 5 YEARS OLD THEY ARE INDICATING I HAVE TO PAY (B)(6) OUT OF POCKET BECAUSE THEY HAVE NO REMEDY. I AM REALLY DISTURBED AS NO ONE ADVISED ME OF THIS PROBLEM!!! PHILIPS NOR THE DME CO, ADVISED OF THIS ISSUE. I HAVE BEEN EXCESSIVELY SHORT OF BREATH FOR YEARS ON THE MACHINE WITHOUT CAUSE. I HAVE BREATHED A STRANGE SMELL THROUGH THE MASK FOR 2 YEARS THAT I THOUGHT WAS A NEW MASK SMELL AND IT WAS THE OFF GAS FROM THE FILTER BREAKING DOWN. NO ONE TIL THIS DAY HAS SENT THE USERS A LETTER ADVISING US OF THIS. NOT THE DR, SUPPLIER OR PHILIPS. THIS US A DISGRACE. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02080 | 2021-07-08 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS CONSTANTLY ALARMING. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
MW5102437 | 2021-07-09 | 2021-07-07 | I, n, j, u, r, y | Voluntary report | DREAM STATION C-PAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. |  |  |  | NOTIFIED BY MY PULMONOLOGIST THAT MY C-PAP MACHINE HAS BEEN RECALLED. I AM HAVING RESPIRATORY ISSUES: CHRONIC COUGH, DRY THROAT, RASPY VOICE. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102436 | 2021-07-09 | 2021-01-01 | I, n, j, u, r, y | Voluntary report | DREAM STATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS INC |  |  |  | I HAVE A CHRONIC COUGH FOR APPROXIMATELY 6 1/2 MONTHS NOW. IN EARLY (B)(6) I WAS ON COUGH MEDICINE FOR 14 DAYS TRYING TO GET RID OF IT DIDN'T WORK, (B)(6) I WENT TO THE DOCTOR HE PUT ME ON TWO DIFFERENT INHALERS PROVENTIALL AND CORTISONE FOR SIX WEEKS DIDN'T HELP. ABOUT TWO MONTHS AGO HE PUT ME ON PREDNISONE FOR 11 DAYS AND NEW INHALER DIDN'T HELP EITHER. I HAD AN X-RAY AND MY LUNGS ARE FINE, THEY SAID, DOCTOR DOESN'T KNOW WHY I KEEP COUGHING. I STILL HAVE THE COUGH I USED THE CPAP WITH THE SO CLEAN DEVICE AND TRULY BELIEVE THIS IS MY PROBLEM ,I HAD TO START ON CONSTANT USE OF THE INHALERS I DO HAVE ASTHMA BUT I HAVE NOT USED INHALERS FOR AT LEAST FIVE OR SEVEN YEARS NOW ALL OF A SUDDEN IT'S A DAILY THING. THEY HELP ME SLIGHTLY BUT I'M STILL COUGHING THAT HARD IT MAKES MY HEAD HURT AND IT MAKES OTHER PEOPLES EARS HURT. IT RUINED MY RESIMED CPAP THAT I HAD FOR 4 YEARS AND THEN I GOT A NEW DREAM STATION MACHINE AND IT BROKE AFTER 11 MONTHS NOW I'M ON THE THIRD DREAM STATION MACHINE AND IT STARTED TO ACT UP SO I AM IN THE PROCESS OF GETTING ANOTHER CPAP MACHINE. I JUST FOUND OUT TODAY (B)(6) 2021 THAT THIS IS CAUSING PROBLEMS WITH PEOPLES LUNGS AND THE CPAP MACHINES ,BECAUSE OF THE OZONE IT USES, I TRULY BELIEVE NOW THAT THIS IS WHAT MY PROBLEM HAS COME FROM AND I DON'T KNOW HOW I'LL GET RID OF THIS CHRONIC COUGH. NONE OF THE COMPANIES REPORTED THE PROBLEM TO ME I FOUND OUT ON (B)(6). THERE ARE TIMES WHEN MY THROAT FEELS LIKE IT IS CLOSING AND WHEN I BREATHE IN AIR FROM THE AIR CONDITIONER OR SOMETIMES JUST AIR IT SEEMS TO DO SOMETHING IN MY THROAT THAT WILL BRING ON THIS COUGH. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102438 | 2021-07-09 | 2021-06-10 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS DREAM STATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | AFTER USING THE PHILIPS RESPIRONICS DREAM STATION CPAP FOR ALMOST 2 MONTHS NIGHTLY FOR AT LEAST 6 HOURS A NIGHT,I STARTED HAVING SEVERE CONGESTION IN THE MORNING AFTER WAKING UP. THE SINUS CONGESTION WAS SO SEVERE THAT I WAS UNABLE TO BREATH CAUSING GASPING FOR BREATH AND EVENTUAL VOMITING. THIS OCCURRED DAILY FOR ALMOST TWO WEEKS UNTIL I WAS INFORMED TO STOP USING THE PHILIPS RESPIRONICS DREAM STATION CPAP DUE TO A RECALL DUE TO TOXIC CHEMICALS AND FOAM. I HAVE DISCONTINUED USING THE DEVICE BUT STILL HAVE SOME EPISODES OF CONGESTION AND VOMITING EVEN AFTER GETTING DECONGESTION MEDICINE FROM MY DOCTOR. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102434 | 2021-07-09 | 2020-11-15 | I, n, j, u, r, y | Voluntary report | DREAMSTATION ACPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS / RESPIRONICS, INC. | DREAMSTATION | DSX500S11 |  | ONGOING CONGESTION ISSUES, HEADACHES, LUNG INFLAMMATION, HIGH BLOOD CO2, AND LOW O2 SATURATION WITH PHILIPS DREAMSTATION. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02113 | 2021-07-09 | 2021-06-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1134711 | 1134711 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A MALFUNCTION OF A VENTILATOR'S EXTERNAL FLOW SENSOR CAUSING DECREASED TIDAL VOLUMES. THERE WAS NO HARM OR INJURY REPORTED. THE COMPONENT HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A MALFUNCTION OF A VENTILATOR'S EXTERNAL FLOW SENSOR CAUSING DECREASED TIDAL VOLUMES. THERE WAS NO HARM OR INJURY REPORTED. THE COMPONENT WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY AND THE CUSTOMER'S COMPLAINT WAS NOT DUPLICATED. THE COMPONENT WAS FOUND TO OPERATE AND ALARM AS DESIGNED. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
MW5102439 | 2021-07-09 |  | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS, INC. |  |  |  | PHILIPS IS REFUSING TO REPLACE RECALLED CPAP MACHINE. WHEN YOU CALL PHILIPS THEY EITHER HANG UP ON YOU OR DO NOT ANSWER THE CALL. BEING WITHOUT MY CPAP MACHINE IS NEGATIVELY AFFECTING MY HEALTH. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | I |  | 2021-07
MW5102446 | 2021-07-09 | 2021-07-08 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSXHCP |  |  | PHILIPS AUTO BIPAP POTENTIALLY CAUSING PULMONARY FIBROSIS. TWO YEAR USE, SYMPTOMS APPEARING AFTER ONE YEAR USE AND PROGRESSING. NO CONTRAST, AREAS OF SUBPLEURAL RETICULATION AT THE LUNG BASES. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02116 | 2021-07-09 | 2021-07-02 | D, e, a, t, h | Manufacturer report | DREAMSTATION AUTO CPAP, GB | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | GBX500S15 | GBX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE DEATH OF A PATIENT. THE MANUFACTURER HAS ATTEMPTED TO ARRANGE FOR THE RETURN OF THE DEVICE FOR EVALUATION TO THE MANUFACTURER'S PRODUCT INVESTIGATION LAB. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02112 | 2021-07-09 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02110 | 2021-07-09 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02111 | 2021-07-09 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02108 | 2021-07-09 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02106 | 2021-07-09 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02105 | 2021-07-09 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02107 | 2021-07-09 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02118 | 2021-07-09 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX500T12C | CAX500T12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE RESPIRATORY AND SINUS INFECTIONS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTION MEDICATIONS. THE MANUFACTURER HAS ATTEMPTED TO ARRANGE FOR THE RETURN OF THE DEVICE TO THE MANUFACTURER'S PRODUCT INVESTIGATION LAB. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON MAY 02, 2023 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE RESPIRATORY AND SINUS INFECTIONS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTION MEDICATIONS.. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE THIRD PARTY SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED. THERE WERE NO ERRORS FOUND. THE THIRD PARTY SERVICE CENTER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. DEVICE WAS SCRAPPED. SECTIONS B1, B2, D8, D9, H1, H3, H6 HAS BEEN CORRECTED AND UPDATED IN THIS REPORT. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02119 | 2021-07-09 | 2021-06-18 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | CAX500T12W | CAX500T12W |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT WAS DIAGNOSED WITH BLADDER CANCER AND REQUIRED SURGERY. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE MANUFACTURER HAS ATTEMPTED TO ARRANGE FOR THE RETURN OF THE DEVICE TO THE MANUFACTURER'S PRODUCT INVESTIGATION LAB. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT WAS DIAGNOSED WITH BLADDER CANCER AND REQUIRED SURGERY. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. IN THE INITIAL REPORT, B3 DATE WAS INCORRECT, WHICH IS UPDATED IN THIS REPORT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL AND INTERNAL VISUAL INSPECTION OF DEVICE WERE COMPLETED BY THE MANUFACTURER AND FOUND EVIDENCE OF UNKNOWN CONTAMINANT ON THE UI KNOB ON THE FRONT PANEL, ON THE ISO PORT, ON THE REAR PANEL, AND ON THE BOTTOM ENCLOSURE EXTERNALLY. INTERNAL INVESTIGATION SHOWS UNKNOWN CONTAMINANT ON THE TOP ENCLOSURE AND FRONT PANEL, A DUST CONTAMINANT ON THE FRONT PANEL, AN UNKNOWN CONTAMINANT ON THE BLOWER CAP, AN UNKNOWN DUST CONTAMINANT ON THE BOTTOM ENCLOSURE, AN UNKNOWN CONTAMINANT ON THE BLOWER SEAL, LIQUID INGRESS WITH MINERAL DEPOSITS ON THE BLOWER AND BLOWER BOX, KERATIN-LIKE SUBSTANCE ON THE BLOWER BOX OUTLET. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS LOGGED. THE DEVICE WAS APPLIED POWER AND THE DEVICE OPERATED PROPERLY. THE MANUFACTURER CONCLUDES MULTIPLE CONTAMINATES FOUND WERE INCONSISTENT WITH SOUND ABATEMENT FOAM DEGRADATION AND A DUST CONTAMINANT ON THE FRONT PANEL, LIQUID INGRESS WITH MINERAL DEPOSITS ON THE BLOWER AND BLOWER BOX, KERATIN-LIKE SUBSTANCE ON THE BLOWER BOX OUTLET WERE CONSISTENT WITH THE SOUND ABATEMENT FOAM. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02109 | 2021-07-09 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02165 | 2021-07-12 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICES LCD DISPLAY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02162 | 2021-07-12 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLANK. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE WAS EVALUATED AT THE MANUFACTURER'S SERVICE CENTER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. | THE PREVIOUS SUBMITTED REPORT WAS MISSING THE SERIAL NUMBER IN SECTION D4. THE INFORMATION IS CORRECT ON THIS REPORT. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
MW5102489 | 2021-07-12 | 2021-02-28 | I, n, j, u, r, y | Voluntary report | RESPIRONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS / RESPIRONICS INC. |  |  |  | DEVELOPED A BAD COUGH, SHORT BREATH, DIFFICULTY BREATHING. DIAGNOSED WITH INTERMITTENT ASTHMA. GIVEN 2 ROUNDS OF PREDNISONE, 2 INHALERS AND MEDICATION. NOTHING HAPPENED. WITHIN 3 DAYS OF STOPPING USE OF SOCLEAN SYMPTOMS EASED. WITHIN A WEEK ALL SYMPTOMS DISAPPEARED. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02161 | 2021-07-12 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S SENSOR BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02164 | 2021-07-12 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO PATIENT HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS OBSERVED IN THE DOWNLOADED EVENT LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02163 | 2021-07-12 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS OBSERVED IN THE DOWNLOADED EVENT LOG AND THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02128 | 2021-07-12 | 2021-07-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1032804 | 1032804 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02177 | 2021-07-12 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE RELATED TO THE INTERNAL BATTERY WAS OBSERVED IN THE DOWNLOADED EVENT LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02168 | 2021-07-12 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICE'S LCD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02169 | 2021-07-12 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02170 | 2021-07-12 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02120 | 2021-07-12 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02121 | 2021-07-12 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD AND MACHINE FLOW SENSOR WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02125 | 2021-07-12 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S OXYGEN PERCENTAGE DELIVERY WAS INACCURATE. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S INTERFACE BOARD AND OXYGEN BLENDING MODULE BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02122 | 2021-07-12 | 2021-06-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2100X15B | IN2100X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02124 | 2021-07-12 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02178 | 2021-07-12 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02123 | 2021-07-12 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD AND MACHINE FLOW SENSOR WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02127 | 2021-07-12 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TO RUN ON AC POWER. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S POWER CORD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02171 | 2021-07-12 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02172 | 2021-07-12 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S DISPLAY SCREEN TURNED WHITE. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S LCD SCREEN NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
MW5102464 | 2021-07-12 | 2020-12-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP MACHINE | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. |  |  |  | SEVERE DIZZINESS, NAUSEA, VOMITING AND SEVERE FATIGUE. COULD THIS BE FROM MY CPAP MACHINE? IT IS ON THE RECALL LIST AND HAVE HAD IT OVER A YEAR. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | I |  | 2021-07
2518422-2021-02174 | 2021-07-12 | 2021-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02175 | 2021-07-12 | 2021-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE RELATED TO THE INTERNAL BATTERY WAS OBSERVED IN THE DOWNLOADED EVENT LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02173 | 2021-07-12 | 2021-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02176 | 2021-07-12 | 2021-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE OBSERVED IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD, SENSOR BOARD, AND SENSOR BOARD CABLE WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02126 | 2021-07-12 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A FAILURE OF THE DEVICE TO CHARGE ITS INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S POWER SUPPLY BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
MW5102459 | 2021-07-12 | 2020-01-15 | I, n, j, u, r, y | Voluntary report | PHILLIPS RESPRONICS DREAMSTATION | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | PHILLIPS / RESPIRONICS INC. | DEX500T11C |  |  | I STARTED HAVING LUNG PAIN AND SHORTNESS OF BREATH JUST OVER A YEAR AND A HALF AGO. I WENT TO THE HOSPITAL A FEW TIMES OVER THE LAST YEAR. I HAD TEST AND X-RAYS AND THEY FOUND INFECTION AND VERY SMALL NODULES IN MY CT SCANS. I SAW THAT THERE IS NOW A RECALL ON THE PHILLIPS CPAP MACHINES I HAVE USED THE LAST 5 YEARS. ONE LASTED UNTIL JAN OF THIS YEAR AND THEN I GOT A REPLACEMENT. THE RECALL IS BASED UPON A "SMALL CHANCE" OF THE ISSUES I HAVE BEEN HAVING. I REPORTED IT YESTERDAY TO MY DOCTOR, AND HOPE THEY REPORT IT TO YOU. I ALSO PRAY THAT THERE IS NO LONG TERM ISSUES FROM THESE MACHINES. I HAVE HAD SHORTNESS OF BREATHS, LUNG PAIN, AND COUGHING SPELLS FOR A YEAR. I REGISTERED BOTH DEVICES AND BOTH WERE A PART OF THE RECALL. I NEVER USED AN OZONE CLEANER OR ANYTHING NON VERIFIED BY THE MANUFACTURER. I WANTED TO REPORT TO YOU. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02166 | 2021-07-12 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2100X15B | IN2100X15B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER SERVICE CENTER, THE DEVICE FAILED TO POWER ON. THE DEVICE'S POWER SUPPLY BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02146 | 2021-07-12 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AVAPS C SERIES INTL 30 | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC | 1161T | 1161T |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING EVALUATION OF THE DEVICE AT A THIRD-PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02145 | 2021-07-12 | 2021-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500S11 | DSX500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02148 | 2021-07-12 | 2021-06-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING HEADACHE AND SORE THROAT AND MACHINE GAVE AN ERROR CODE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING HEADACHE AND SORE THROAT AND MACHINE GAVE AN ERROR CODE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION A1, B5 AND H6 WAS CORRECTED( CODINGS WERE NOT UPDATED IN PREVIOUS REPORTS) AND ALSO H6 WAS UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02149 | 2021-07-12 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTOSV GB | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC | GBX900S15 | GBX900S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A BI-LEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. DESPITE MULTIPLE ATTEMPTS, THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. CODES IN SECTION H6, UNIQUE IDENTIFIER (UDI) # IN SECTION D4 WERE UPDATED OR CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02147 | 2021-07-12 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AVAPS C SERIES INTL 30 | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC | 1161T | 1161T |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING EVALUATION OF THE DEVICE AT A THIRD-PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102486 | 2021-07-12 | 2021-06-20 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSXHCP | H239136091FF4 |  | PHILIPS ADVISE ME THAT MY DREAM STATION MACHINE IS BEING RECALLED I HAVE NOT BEEN ADVISED BY THEM OF HOW IT EFFECTED ME. I JUST HAVE A PROBLEM WITH THE DREAM STATION SYSTEM WHICH HAS A HUMIDIFIER. PHILIPS RECENTLY RECALLED THE DREAM STATION UNIT I HAVE BEEN USING FOR 5 YEARS. I HAVE VERY DRY MOUTH DURING THE NIGHT WHEN I WAKE UP AND STARTING IN EARLY 2020 I NOTICED I HAVE A WRENCHING COUGH EVERY MORNING FOR BETWEEN 2 AND 4 MINUTES. I THOUGHT IT MIGHT BE ALLERGIES BUT I WONDER IF IT HAS TO DO WITH THE DEFECT WHICH TRIGGERED A RECALL. THE SERIAL NUMBER OF MY UNIT IS (B)(4). I REGISTERED IT WITH PHILLIPS #2021062201205041. THE MACHINE WAS DELIVERED AND SERVICED BY LINCARE AT (B)(4). IT HAS A HUMIDIFIER TOGETHER THEY FORM "? ?" THE HUMIDIFIER IS DREAMSTATION "? ?" SERIAL # (B)(4), REF # DSXHCP. |  | Y | Y | I |  | 2021-07
MW5102472 | 2021-07-12 | 2021-05-06 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS DREAMMACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS, INC. | DREAM MACHINE |  |  | BLACK SPECS STARTED TO APPEAR WHEN I BLEW MY NOSE IN THE AM. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | I |  | 2021-07
2518422-2021-02167 | 2021-07-12 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2100X15B | IN2100X15B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER SERVICE CENTER, THE DEVICE FAILED TO POWER ON. THE DEVICE'S POWER SUPPLY BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
MW5102506 | 2021-07-12 |  | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | RESPIRONICS BIPAP AUTO SV | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | PHILIPS / RESPIRONICS, INC. |  |  |  | PHILIPS AUTO SV BIPAP AND CPAP RECALLED - POSSIBLE RELEASE OF CARCINOGENIC GASES DUE TO DEFECT IN NOISE REDUCING FOAM. AFFECTED CPAP USED 12/2018 TO 6/2019. AFFECTED BIPAP PURCHASED (B)(6) 2019 TO REPLACE CPAP DUE TO MORE EXTENSIVE SLEEP APNEA. SQUAMOUS CELL CARCINOMA (BASE OF TONGUE) DIAGNOSED (B)(6) 2020. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102527 | 2021-07-13 | 2021-07-01 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500T11W |  |  | I LEARNED THROUGH A BOSTON GLOBE ARTICLE THAT THE PHILIPS CPAP DREAMSTATION I'VE BEEN USING FOR MONTHS IS PART OF A RECALL. THE FOAM IS BREAKING DOWN AND I'VE BEEN BREATHING THE PARTICLES EVERY NIGHT. NO ONE FROM PHILIPS CONTACTED ME. NO ONE FROM (B)(6). NO ONE FROM MY SLEEP DISORDER CLINIC AT (B)(6) MEDICAL CENTER. THIS IS UNCONSCIONABLE. I HAVE HAD AN ONGOING COUGH FOR WEEKS. PHILIPS CLAIMS ON THEIR WEBSITE THAT FEW PEOPLE ARE AFFECTED. THEY HAVE NO CLUE HOW MANY PEOPLE ARE AFFECTED BECAUSE THERE IS NO PLACE ON THEIR WEBSITE WHERE I CAN REPORT MY SYMPTOMS. WHO IS RESPONSIBLE FOR NOTIFYING PATIENTS? WHO IS RESPONSIBLE FOR OFFERING IMMEDIATE REPLACEMENTS? FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02197 | 2021-07-13 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
MW5102523 | 2021-07-13 | 2020-12-01 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. | DSX50011C |  |  | I HAVE BEEN HAVING MUCUS IN MY LUNG WHICH I HAVE TO FORCE A COUGH TO GET OUT OF MY LUNG SEVERAL TIMES A DAY. X-RAY CONFIRMED A SPOT ON MY LUNG. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | I |  | 2021-07
2518422-2021-02144 | 2021-07-13 | 2021-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING HEADACHE, NAUSEA, SORE THROAT RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02159 | 2021-07-13 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WOULD NOT POWER ON. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. AN ISSUE RELATED TO THE TOUCHSCREEN WAS OBSERVED. THE DEVICE'S DISPLAY BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02158 | 2021-07-13 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED.. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02129 | 2021-07-13 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02160 | 2021-07-13 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TESTING. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SENSOR BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02188 | 2021-07-13 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S TOUCH SCREEN WOULD NOT FUNCTION. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S TOUCH SCREEN/DISPLAY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02195 | 2021-07-13 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WOULD NOT POWER ON. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S DISPLAY BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
MW5102524 | 2021-07-13 | 2021-07-11 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS RESPIRONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I HAVE A PHILLIPS DREAMSTATION CPAP MACHINE. THIS MORNING I TOOK THE HOSE, FILTER, AND WATER CHAMBER INTO THE KITCHEN TO CLEAN THEM, AND NOTICED THAT BOTH FILTERS WERE BLACK. THE FILTERS ARE SUPPOSED TO BE WHITE AND HAVE ALWAYS BEEN WHITE SINCE I'VE OWNED THE MACHINE. I KNOW THERE'S A RECALL WITH THE MACHINE, SO I REGISTERED MY MACHINE WITH THE MANUFACTURER AND IT IS INDEED AN AFFECTED/RECALLED UNIT. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | I |  | 2021-07
2518422-2021-02115 | 2021-07-13 | 2021-06-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING RESPIRATORY PROBLEMS AND VISUALIZATION OF COUPLE PIECES OF BLACK SPECS FLOATING AROUND IN WATER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT WAS SCRAPPED. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02036 | 2021-07-13 | 2021-06-17 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX400T11C | DSX400T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND RESULTED IN THE PATIENT DEVELOPING PNEUMONIA AND ABSCESS ON LEFT LUNG. THE DETAILS OF THE MEDICAL INTERVENTION THE PATIENT RECEIVED ARE CURRENTLY UNKNOWN. THE INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND RESULTED IN THE PATIENT DEVELOPING PNEUMONIA AND ABSCESS ON LEFT LUNG. THE DETAILS OF THE MEDICAL INTERVENTION THE PATIENT RECEIVED ARE CURRENTLY UNKNOWN. THE REPORTED EVENT AND ITS SEVERITY WERE REVIEWED BY THE MANUFACTURER'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. HENCE, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO THE ALLEGED SERIOUS INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02150 | 2021-07-13 | 2021-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | CAX400H12C | CAX400H12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | IN THE PREVIOUS REPORT, SECTION B5 WAS NOT UPDATED, THE CORRECT B5 SHOULD BE- THE PATIENT HAS ALLEGED HEADACHES, SORE THROAT, DIZZINESS, LIGHTHEADED, SHORTNESS OF BREATH, NASAL SORENESS, RESPIRATORY ISSUES, INFLAMMATION AND CHEST PAIN. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. IN THIS REPORT, SECTION H6 HAS BEEN UPDATED/CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02143 | 2021-07-13 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A NODE IN THE PATIENT'S LUNG. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE NODE IN THE PATIENT'S LUNGS. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE EVIDENCE OF WATER INGRESS ON BLOWER AND BLOWER BOX OBSERVED IN THE BASE UNIT AND DUST/DIRT CONTAMINATION INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM WAS OBSERVED THROUGHOUT DEVICE ENCLOSURE, AND AIRPATH, SUGGESTING A SOURCE EXTERNAL TO THE DEVICE. SLIGHT BLACK CONTAMINATION AT THE BLOWER SEAL OF THE BLOWER BOX IS CONSISTENT WITH THE KERATIN CONTAMINATION OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY CONFIRM THE CUSTOMER'S ALLEGATION AND THERE IS DUST / DIRT CONTAMINATION AND THE PRESENCE OF CONTAMINATION IN THE AIRPATH, SOURCE OF CONTAMINATIONS WERE EXTERNAL TO THE DEVICE. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02202 | 2021-07-13 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102520 | 2021-07-13 | 2021-06-14 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILLIPS RESPIRONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS / RESPIRONICS | DNX500H11C | 1120833 |  | I READ THAT THE CPAP DEVICE THAT I USE MIGHT BE ONE THAT IS BEING RECALLED. WHEN I CONTACTED PHILLIPS THEY TOLD ME THEY DID NOT HAVE A TIMELINE FOR REPLACING MY MACHINE AND THAT IT WAS UP TO ME AND MY PHYSICIAN TO DETERMINE WHETHER I SHOULD CONTINUE USE OF MY CPAP. I USE A CPAP BECAUSE I HAVE SLEEP APNEA. I CANNOT SLEEP WITHOUT IT. A MONTH LATER AND AFTER REGISTERING MY MACHINE ONLINE, NO ONE CAN TELL ME WHEN I WILL GET A NEW MACHINE. MY MACHINE IS NOT 5 YRS OLD SO UNDER NORMAL CIRCUMSTANCES I DO NOT NEED TO REPLACE IT. IN THE RECALL ANNOUNCEMENT PHILLIPS STATED THAT I SHOULD "DISCONTINUE USE OF AFFECTED UNITS AND CONSULT WITH PHYSICIANS TO DETERMINE THE BENEFITS OF CONTINUING THERAPY AND POTENTIAL RISKS." WHAT DOES THAT MEAN? (I HAVE HAD (AND CONTINUES TO HAVE) HEADACHES AND SINUS ISSUES FOR A YEAR, BUT JUST CHALKED IT UP TO A SIDE EFFECT OF USING A CPAP). IF IT IS DANGEROUS FOR ME TO USE MY MACHINE AND I CANNOT SLEEP WITHOUT IT, WHAT ARE MY OPTIONS? I DO NOT THINK IT IS MY RESPONSIBILITY TO BUY A NEW MACHINE. THIS ISSUE WAS BROUGHT TO THE ATTENTION OF PHILLIPS IN APRIL OF '21. IT IS NOW 3 MOS LATER AND THUS FAR HAVE DONE NOTHING TO RECTIFY THIS SITUATION. PLEASE HELP ME, AND THE MILLIONS OF OTHERS HERE IN THE US WHO NEED CPAP'S TO TREAT OUR SLEEP APNEA. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | I |  | 2021-07
MW5102529 | 2021-07-13 |  | I, n, j, u, r, y | Voluntary report | DREAMSTATION BIPAP AUTO SV | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | PHILLIPS / RESPIRONICS | D5X900H11C |  |  | MY BIPAP MACHINE IS ONE OF THE PHILLIPS RESPIRONICS MACHINES UNDER THE HUGE RECALL. I NEVER KNEW THAT'S WHY I KEPT GETTING A RASH ON MY FACE EVERY NIGHT FROM WEARING MY FACE MASK. UNTIL I JUST RECENTLY HEARD ABOUT THE TEA CALL AND PART OF IT CAN CASE SKIN, EYE, AND RESPIRATORY TRACT IRRITATION. NOW I KNOW WHY MY FACE KEEPS BREAKING OUT EACH NIGHT FROM MY MACHINE. I TRIED ALL KINDS OF CREAMS AND ONE DOCTOR THOUGHT IT WAS ECZEMA AND TOLD ME TO TRY A CERTAIN LOTION THAT DID NOT WORK. THE RASH IS NOT PAINFUL BUT IS ANNOYING AND LOOKS UGLY BECAUSE IT'S LITTLE RED SPORTS ALL OVER MY FACE WHICH IS EMBARRASSING. IT'S ON MY CHEEKS MOSTLY. NO THAT I HAVE THE RECALL PAPER I FINALLY KNOW WHAT IT IS. MY DOCTOR DID SAY TO KEEP USING THE MACHINE THOUGH UNTIL PHILLIPS DECIDES TO EITHER REPLACE OR FIX MY MACHINE. MY DOCTOR SAID THAT USING THE MACHINE OUT WAYS THE RISKS OF A RASH AND TO KEEP AND I ON IT OR IF I SEE ANY PARTICLES IN MY TUBING (WHICH IS HARD WHEN I AM ASLEEP). SO I STILL USE IT. THE PAPER I RECEIVED FROM THE COMPANY THAT HANDLES MY MACHINES THINGS LIKE MY MASKS, TUBING, CUSHIONS, ECT. NAMED APRIA STATED THAT IF WE HAD ANY SIDE EFFECTS FROM THE MACHINES THAT'S ON THE RECALL LIST TO REPORT IT HERE. SO I AM. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | I |  | 2021-07
2518422-2021-02220 | 2021-07-14 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S PARTICULATE FILTER WAS FOUND TO BE BLOCKED AND NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02152 | 2021-07-14 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AVAPS C SERIES | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS INC | 1161T | 1161T |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING THE EVALUATION OF THE DEVICE AT A THIRD-PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02194 | 2021-07-14 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE'S KEYPAD RESPONSE WAS INADEQUATE. THE DEVICE'S FRONT PANEL/KEYPAD LED BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
MW5102552 | 2021-07-14 | 2021-07-09 | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | DATE GIVEN IS WHEN I LEARNED OF THE RECALL FROM ANOTHER PATIENT, NOT FROM THE MANUFACTURER WHO HAS NEVER NOTIFIED ME. MOST APPARENT PROBLEM IS THAT MY LIFE LONG ASTHMA HAS GOTTEN MUCH WORSE OVER THE LAST 2 TO 3 YEARS WHILE USING THE CPAP MACHINE. ALSO, IT HAS BEEN SUGGESTED THAT THE VOC FUMES FROM THE DETERIORATING PE PUR FOAM MAY HAVE PLAYED A ROLE IN MY PROSTATE CANCER AND IN MY NON-HODGKINS LYMPHOMA. HAVE STOPPED USING CPAP DEVICE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102550 | 2021-07-14 | 2019-06-01 | I, n, j, u, r, y | Voluntary report | DREAM STATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I HAVE BEEN USING A PHILIPS DREAM STATION CPAP FOR 3 YEARS. I HAVE SINCE BEEN DIAGNOSED WITH ASTHMA. I HAVE HEADACHES, DIZZINESS, COUGH, SORE THROAT, RED/IRRITATED THROAT, ASTHMA, CHEST PRESSURE, TROUBLE BREATHING, STUFFY NOSE, HYPERSENSITIVITY TO SMELLS AND CHEMICALS. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102542 | 2021-07-14 |  | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS RESPIRONICS DREAM STATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. | DSX500H11C |  |  | FOR CRITICAL APNEA I USED 2 PHILLIPS RESPIRONICS DREAM STATION CPAPS OVER PAST 5 YRS. A NEW PHILIPS MACHINE WAS PURCHASED (B)(6) 2021 . DURING THIS TIME I HAD A CHRONIC COUGH AND NEED TO CLEAR MY THROAT EVERY MORNING . SINCE RECALL OF PHILIPS RESPIRONICS CPAP 6/14/2021. I PURCHASED OUT OF POCKET ((B)(6) 2021)A RESMED AIR SENSE 10 CPAP. I HAVE NOT BEEN COUGHING OR CLEARING MY THROAT IN THE MORNING SINCE USING RES MED AIR SENSE 10. ALSO ON 2 OCCASIONS WITH THE PHILIPS RESPIRONICS CPAP DREAM MACHINE PURCHASED IN (B)(6) OF THIS YEAR THERE WAS DEBRIS FOUND IN WATER CHAMBER. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | I |  | 2021-07
2518422-2021-02191 | 2021-07-14 | 2021-07-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02192 | 2021-07-14 | 2021-07-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02200 | 2021-07-14 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD AND MACHINE FLOW SENSOR WERE REPLACED TO ADDRESS THE ISSUE. THE DEVICE HAD EVIDENCE OF WATER INGRESS. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02219 | 2021-07-14 | 2020-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE AND OVERHEAD BLOWER FILTER NEED TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02193 | 2021-07-14 | 2021-07-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02130 | 2021-07-14 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02133 | 2021-07-14 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02135 | 2021-07-14 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A FAILURE OF THE DEVICE TO POWER ON WAS OBSERVED. THE DEVICE'S POWER SUPPLY WAS REPLACED TO ADDRESS THE ISSUE. A FAILURE OF THE DEVICE TO CHARGE ITS INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02136 | 2021-07-14 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S SENSOR BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02137 | 2021-07-14 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS PROVIDING LOW FLOW. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S INTERNAL TUBING WAS RECONNECTED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02134 | 2021-07-14 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02139 | 2021-07-14 | 2021-07-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02189 | 2021-07-14 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02218 | 2021-07-14 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | RIN2110X15B | RIN2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VARIATION IN THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE DELIVERY WAS OBSERVED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN EVALUATED BY THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VARIATION IN THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE DELIVERY WAS OBSERVED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE VENTILATOR'S SOFTWARE WAS RE-LOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02221 | 2021-07-14 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02217 | 2021-07-14 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | RIN2110X15B | RIN2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VARIATION IN THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE DELIVERY WAS OBSERVED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN EVALUATED BY THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VARIATION IN THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE DELIVERY WAS OBSERVED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE VENTILATOR'S SOFTWARE WAS RE-LOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02225 | 2021-07-14 | 2021-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR HAD AN OXYGEN LEAK. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN BLENDING MODULE NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02224 | 2021-07-14 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE MANUFACTURER IS CURRENTLY INVESTIGATING THIS EVENT AND A FOLLOW-UP REPORT WILL BE FILED WHEN THE INVESTIGATION IS COMPLETE. | IT WAS INITIALLY REPORTED, A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE MANUFACTURER IS CURRENTLY INVESTIGATING THIS EVENT AND A FOLLOW-UP REPORT WILL BE FILED WHEN THE INVESTIGATION IS COMPLETE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERNAL BATTERY WERE REPLACED TO ADDRESS THE ISSUES. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02222 | 2021-07-14 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02131 | 2021-07-14 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02216 | 2021-07-14 | 2021-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | RIN2110X15B | RIN2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VARIATION IN THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE DELIVERY WAS OBSERVED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN EVALUATED BY THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VARIATION IN THE VENTILATOR'S PRESSURE DELIVERY. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02227 | 2021-07-14 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | 760P | 760P |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02190 | 2021-07-14 | 2021-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
MW5102551 | 2021-07-14 | 2020-10-13 | I, n, j, u, r, y | Voluntary report | DREAMSTATION CELLULAR MODEM | MEDICAL DEVICE DATA SYSTEM | PHILIPS / RESPIRONICS, INC. | J291549965318 |  |  | STOPPED USING DEVICE; I HAVE BEEN USING A CPAP FOR SEVERE SLEEP APNEA SINCE (B)(6) 2020. OVER THE MONTHS I HAVE HAD INCREASED HEADACHES, THROAT IRRITATION AND COUGHING. THERE HAS ALSO BEEN UNEXPLAINED DIZZINESS INVOLVED. IT IS ONLY IN (B)(6) 2021 THAT WE LEARNED OF THE RECALL ON THE CPAP DREAM STATION DEVICE. MY SERVICE PROVIDER, WHERE DEVICE WAS PURCHASED DID NOT INFORM ME OF THIS UNTIL I CALLED THEM PERSONALLY AND TOLD THEM I LEARNED OF IT THROUGH INTERNET. I HAVE NOT HAD ANY TESTS DONE AS YET. I HAD A TELEPHONE CONFERENCE WITH MY GP ON (B)(6) AND I'M AWAITING CALL FROM DOCTOR SPECIALIST IN SLEEP APNEA. TESTS TO FOLLOW. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102535 | 2021-07-14 | 2020-06-06 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | BLACK SPOTS WERE NOTICED ON A SILICONE SEAL IN MY ALMOST-FIVE-YEAR-OLD RESPIRONICS DREAMSTATION AUTO CPAP MACHINE. I THOUGHT IT WAS BLACK MOLD, BUT I COULDN'T REMOVE IT WITH LIGHT PRESSURE FROM A QTIP. THE SPOTS WERE LOCATED ON THE CORNERS OF THE SILICONE SEAL ON THE UNDERSIDE OF THE HUMIDIFIER. AFTER I GOT A NEW DREAMSTATION CPAP MACHINE (B)(6) 2021, I PULLED OFF THE SILICONE SEAL FROM THE OLD CPAP MACHINE AND WASHED IT WITH BABY SHAMPOO AND WATER, AND IT EVENTUALLY THE BLACK SPOTS WERE REMOVED. I STILL HAVE THIS OLD CPAP MACHINE AND KEPT IT AS A BACKUP MACHINE. IN (B)(6) 2021 I WAS NEWLY DIAGNOSED WITH A RARE FORM OF BLOOD CANCER, POLYCYTHEMIA VERA, BY A HEMATOLOGIST ONCOLOGIST AT THE (B)(6). JAK2 TEST ORDERED BY A DIFFERENT DOCTOR THE LATTER PART OF 2020, AND HE REFERRED ME TO THE HEMATOLOGIST WHO DIAGNOSED POLYCYTHEMIA VERA. I'M GIVING APPROXIMATE MONTH AND DAY DATES AS I'D HAVE TO LOOK LATER FOR THE EXACT DATE OF THE LAB RESULT, OR EXACT DAY I SAW THE HEMATOLOGIST. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102534 | 2021-07-14 |  | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS RESPIRONICS, INC. |  |  |  | WAS USING THE PHILLIPS RESPIRONICS CPAP DREAM STATION FOR 3 YEARS AND JUST GOT NOTIFIED OF A RECALL FOR THE MACHINE. HAVE OFTEN GOTTEN A COUGH AND COOLER AIRWAY IRRITATION WHEN USING. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-07
2518422-2021-02196 | 2021-07-14 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02155 | 2021-07-14 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE REPORTER ALSO ALLEGED THAT PATIENT HAVING NASAL/THROAT IRRITATION AND NOSE BLEEDS. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND EVIDENCE OF 3 ERRORS OCCURRING. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION A1, A2, A3, D9, H3 AND H6 WERE UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102547 | 2021-07-14 |  | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. |  |  |  | PATIENT IS THROUGH THE (B)(6), THE (B)(6) SENT HIM A NEW MOUTH PIECE AND HOSE AND NOW THE MACHINE HAS CRACKED. THERE IS NO REPLACEMENT AVAILABLE YET AS HE DOESN'T SEE THE DOCTOR UNTIL THE (B)(6) FROM THE (B)(6) AND HE NEEDS A NEW MACHINE. SEEN AT (B)(6) AND DISCHARGED EVEN WITH BROKEN MACHINE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02208 | 2021-07-14 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02156 | 2021-07-14 | 2021-06-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AHX500H11 | AHX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT ALLEGED OF HAVING BURNING SENSATION, SNEEZING, NASAL OR THROAT IRRITATION. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL PART OF THE DEVICE AND FOUND TWO BLACK MARKS ON THE UPPER ENCLOSURE/SD CARD FLIP COVER, AND ONE BLACK MARK ON THE BOTTOM ENCLOSURE. SLIGHT YELLOWISH DISCOLORATION WAS OBSERVED ON THE ISO PORT/AIR OUTLET. BLACK PARTICLES WERE OBSERVED ON THE INTERIOR OF THE ISO PORT. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND FOUND MORE BLACK PARTICLES IN THE BLOWER BOX AND ON THE BLOWER. A SMOKY ODOR SIMILAR TO BURNED/MELTED PLASTIC WAS OBSERVED. THE BLOWER WAS FOUND TO HAVE WHAT APPEARS TO BE A DARK CONGEALED MASS OF SEMI-LIQUID MATERIAL AROUND THE INTAKE. THE IMPELLER WAS ABLE TO ROTATE. EXAMINATION OF THE SOUND ABATEMENT FOAM IN THE AIRPATH SHOWED A SECTION OF FOAM APPROXIMATELY 2.5CM-3CM HAD BROKEN OFF AND WAS MISSING. THE CONGEALED MATERIAL FOUND IN THE BLOWER, THE UNUSUAL ODOR, THE PATIENT'S STATEMENT REGARDING THE BLOWER MOTOR STUTTERING, AND THE E-22 SUGGEST THAT THE MISSING SECTION OF FOAM WAS CAUGHT BETWEEN THE BLOWER ENCLOSURE AND IMPELLER. BEING CAUGHT IN THE BLOWER MOST LIKELY WOULD HAVE TRIGGERED THE E-22 TO BE GENERATED, POTENTIALLY RESULTING IN THE ODOR AND CONGEALED MATERIAL OBSERVED. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND TO CONTAIN 1 ERROR CODE. THE MANUFACTURER VERIFIED THE UNITS DO POWER-UP, PROVIDE FLOW. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION AND CONTAMINATION THROUGHOUT THE DEVICE, WHICH IS LIKELY FROM EXTERNAL SOURCE. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02153 | 2021-07-14 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS UPDATED REPORT WILL BE FILED. CORRECTED INFORMATION PROVIDED IN A2 AND A3 (PATIENTS AGE AND GENDER) WAS NOT INCLUDED IN INITIAL REPORT. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02198 | 2021-07-14 | 2021-06-18 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION EXPERT W/HUM/HT/CELL, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX501T12C | CAX501T12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED PNEUMONIA IN A PATIENT. THE PATIENT RECEIVED MEDICAL INTERVENTION FROM THEIR DOCTOR IN THE FORM OF AN X-RAY AND ANTIBIOTICS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT ALLEGED PNEUMONIA WHILE USING THE DEVICE. PATIENT ALSO ALLEGED VISUALIZATION OF PARTICLES IN THE AIRWAY. MEDICAL INTERVENTION WAS PRESCRIPTION ANTIBIOTICS AND AN X-RAY. THE MANUFACTURER CONTACT INFORMATION HAS ALSO BEEN UPDATED. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED IN THIS DEVICE. PRODUCT INVESTIGATION LAB VISUALLY OBSERVED MINIMAL BEIGE DUST OR DIRT CONTAMINATE OR COATING, AND DARK FIBER CONTAMINATE WITH THE UNIT AND AIRPATH. PRODUCT INVESTIGATION LAB VISUALLY OBSERVED 1 DARK HAIR OR FIBER CONTAMINATE. THE HAIR OR FIBER CONTAMINATE WAS NOT LOCATED INSIDE THE AIRPATH. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND 0 LOGGED ERRORS. THE MANUFACTURER IS UNABLE TO ADDRESS THE COMPLAINT OR SYMPTOMS. THE MANUFACTURER FINDS THAT THE OBSERVED CONTAMINATES WERE SOURCED FROM EXTERNAL ENVIRONMENTAL CONDITIONS. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. PLEASE SEE SECTIONS D9, G1, H3 AND H6 FOR THIS UPDATED INFORMATION AND CODING. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02154 | 2021-07-14 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | AUX500T15C | AUX500T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. SECTION H6 CODES CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT HAS BEEN CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02207 | 2021-07-14 | 2021-06-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | DSX900S11 | DSX900S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02245 | 2021-07-15 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02243 | 2021-07-15 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR WAS ALARMING FOR A VENTILATOR INOPERATIVE CONDITION. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS NOT CONFIRMED. A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02138 | 2021-07-15 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS HAVING A TOUCHSCREEN ISSUE. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A VENTILATOR WAS HAVING A TOUCHSCREEN ISSUE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02223 | 2021-07-15 | 2021-06-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS HAVING A TOUCHSCREEN ISSUE. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF A VENTILATOR HAVING A TOUCHSCREEN ISSUE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02180 | 2021-07-15 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02151 | 2021-07-15 | 2021-07-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | INX700S15 | INX700S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02181 | 2021-07-15 | 2021-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CNX500T17 | CNX500T17 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE LUNG NODULES. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02211 | 2021-07-15 | 2021-06-17 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11 | DSX700T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP HEADACHES AND SPOTS IN LUNGS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING TO DEVELOP HEADACHES AND SPOTS IN LUNGS,HEADACHE,CONGESTED,COUGH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM.THE PATIENT ALSO ALLEGED TO SEVERE SINUS AND THROAT DRYNESS, HEADACHES, DIZZINESS AND KIDNEY ISSUES WAS REVIEWED BY THE PMS CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE.THE PATIENT DID RECEIVE MEDICAL INTERVENTION AND REPORTED TO HAVE SURGERY ON THE STOMACH. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02182 | 2021-07-15 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | AUX500T15C | AUX500T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | SECTION B5 WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT SHOULD BE REPORTED AS FOLLOWS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION PATIENT NOTICED PARTICLES IN WATER CHAMBER. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. SECTION H6 CODES CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT HAS BEEN CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02242 | 2021-07-15 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED A TEST STEP DURING TESTING. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN EVALUATED BY THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR FAILED A TEST STEP DURING TESTING. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02117 | 2021-07-15 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING STOMACH PAIN RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING STOMACH PAIN RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. AFTER THE SECOND ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, THE CUSTOMER STATED THAT THE DEVICE WAS ALREADY RETURNED AND ALSO RECEIVED REPLACEMENT. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. CORRECTED INFORMATION WAS PROVIDED IN SECTION A2 AND E1 (TELEPHONE NUMBER). SECTION H6 WAS CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02184 | 2021-07-15 | 2021-06-29 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON OCTOBER 12, 2023, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING LUNG PROBLEMS AND DIFFICULTY BREATHING WHILE USING THE DEVICE. PATIENT ALSO ALLEGED VISUALIZATION OF BLACK PARTICLES/SPECS IN THE WATER CHAMBER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGATION OF RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, ASTHMA (NEW OR WORSENING), KIDNEY DISEASE/TOXICITY, LUNG DISEASE, OTHER, SINUS INFECTION. THE SECTION E1 WAS UPDATED AND SECTIONS B1, B2, H1 AND H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR BOTH ADVERSE EVENTS AND PRODUCT PROBLEM (ONLY THE PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS REPORT.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECTS - CLINICAL CODES AND HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02204 | 2021-07-15 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION EXPERT | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX501H12C | CAX501H12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102572 | 2021-07-15 | 2021-06-29 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS BIPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS / RESPIRONICS, INC. | DREAM STATION PRO | UNK |  | EARLY ON THE MORNING OF (B)(6) WHILE SLEEPING WITH MY BI-PAP MACHINE, I AWOKE AND FELT AS IF I WAS HAVING A HEART ATTACK. I IMMEDIATELY TOOK THE BI-PAP MASK OFF BUT WAS UNABLE TO TURN OVER DUE TO THE PAIN IN MY LEFT CHEST AND ARM. I GUESS I PASSED OUT FROM THE PAIN. WHEN I CAME TO I TOLD MY HUSBAND THAT I COULDN'T TAKE A DEEP BREATH. I DIDN'T FEEL SO GOOD. SOMETHING WAS WRONG. HE ASKED IF I WANTED TO GO TO THE DOCTOR OR HOSPITAL. I FELT THAT THE DOCTOR WOULD SEE ME QUICKER. MY DOCTOR DID AN EKG AND DETERMINED THAT MY HEART WAS GOOD. HE DID LAB WORK AND THEN AN X-RAY. HE SAID I HAD A COLLAPSED LUNG AND I NEEDED TO GO IMMEDIATELY TO THE EMERGENCY ROOM THEY WOULD NOTIFY THEM I WAS ON MY WAY. AT THE EMERGENCY ROOM THEY TOOK ANOTHER X-RAY AND A CAT SCAN. THEY PUT A TUBE IN MY RIGHT SIDE TO RE-INFLATE MY LEFT LUNG. THE DOCTOR IN THE EMERGENCY ROOM SAID HE FELT THAT THE PROBLEM WAS CAUSED BY THE BI-PAP MACHINE SINCE THERE WAS NO TRAUMATIC EVENT LIKE A FALL OR ACCIDENT OR PHYSICAL EVENT. THEY KEPT ME OVERNIGHT. THEY TOOK ANOTHER X-RAY THAT MORNING AND SAID IT WASN'T COMPLETELY FILLED BACK UP AND WANTED TO KEEP ME AND DO ANOTHER X-RAY AFTER LUNCHTIME. I WENT HOME THE NEXT MORNING. PRODUCT: PHILIPS RESPIRONICS MODEL DREAM STATION BI-PAP MACHINE WITH HUMIDIFIER. FIRST TEST X-RAY WAS AT DOCTORS OFFICE. SECOND TEST X-RAY WAS AT HOSPITAL ER. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102599 | 2021-07-15 | 2018-07-23 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | J17648938CB46 |  |  | HAD TO ADD OXYGEN AT NIGHT ONLY WITH CPAP UNIT. NOTICED A LAYER OF GUNK ON FILTER EVEN AFTER CHANGING IT. INFECTION IN MY EYES, SINUSES, HEADACHE. I NEED ANOTHER UNIT. AS YOU KNOW YOU CANNOT SLEEP WHEN YOU DON'T HAVE A CPAP. WHAT AM I TO DO? THE ONE I WAS USING IS HARMFUL. MY EMAIL IS (B)(6). FDA SAFETY REPORT ID# (B)(4). |  | N | Y | 0 |  | 2021-07
2518422-2021-02244 | 2021-07-15 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERFACE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02183 | 2021-07-15 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP PRO BIFLEX, W/SD CARD, HUMID, US | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DS650HS | DS650HS |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED ASTHMA IN A PATIENT. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTION MEDICATIONS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED ASTHMA. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT.THE REPORTED COLLAPSED ASTHMA AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION,THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTION MEDICATIONS. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTIRON B7 HAS BEEN CORRECTED/UPDATED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102576 | 2021-07-15 | 2016-10-31 | I, n, j, u, r, y | Voluntary report | RESPIRONICS BIPAP AUTO, SYSTEM ONE | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | PHILLIPS RESPIRONICS / RESPIRONICS, INC. | SYSTEM ONE |  |  | I WAS SOLD A RESPIRONICS SYSTEM ONE BIPAP AUTO WITH HUMIDIFIER BY SLEEP TECHNOLOGIES ON (B)(6) 2016. I HAVE RELIGIOUSLY USED IT EVERY NIGHT. I NOW HAVE A PARTIALLY COLLAPSED RIGHT LOWER LUNG, PERSISTENT COUGH FOR 4 PLUS MONTHS, SORES PERIODICALLY ON MY NOSTRILS, SINUS CONGESTION/INFECTIONS, SOB SEVERELY LIMITING WALKING, WEIGHT GAIN, WEAKNESS, ASTHMA (NEWLY DX BY PULMONOLOGIST), POSSIBLE HEART ISSUES (STILL BEING EVALUATED). I PURCHASED THE SOCLEAN SYSTEM APPROXIMATELY 1-2 YEARS AGO AND USE IT DAILY IN A WELL VENTILATED LARGE ROOM AS A METHOD OF CLEANING. THE ATELECTASIS WAS RECENTLY DIAGNOSED VIA CHEST CT, I REALIZED THAT I HAVE HAD A SLIGHT CRINKLING SOUND DURING BREATHING AS LONG AS TWO OR THREE YEARS AGO WHICH IS INTERMITTENT AND ABSENT PRESENTLY. THE SINUS CONGESTION/INFECTIONS HAVE BEEN GOING ON FOR OVER FOUR YEARS. THE SOB HAS STARTED WITHIN THE LAST FIVE YEARS BUT HAS BECOME EXTREME IN THE LAST FOUR MONTHS PROMPTING SIGNIFICANT MEDICAL TESTING. I HAVE WALKED FOR EXERCISE 5-6 TIMES PER WEEK FOR DECADES OFTEN GOING 3-4 MILES PER WALK BUT SINCE THE SEVERE SHORTNESS OF BREATH I HAVE BEEN HARDLY ABLE TO WALK A MILE, NEEDING FREQUENT RESTS TO CATCH MY BREATH. RECENTLY I STARTED MONITORING MY WALKS WITH A PULSE OX. ON MORE THAN ONE OCCASION, AFTER GOING ONLY 1/8TH OF A MILE, MY OX DROPPED TO 89 BUT PARADOXICALLY MY PULSE WAS ONLY 55 DESPITE SEVERE HUFFING AND PUFFING. YET AS SOON AS I STOPPED AND SAT DOWN MY PULSE JUMPED TO 129 AND MY OX SAT RATE SLOWLY RECOVERED TO 96 WITHIN 1 OR 2 MINUTES. THIS HAS SEVERELY AND ADVERSELY EFFECTED MY OVERALL HEALTH. MY SOB IS SIGNIFICANTLY WORSE WHEN I TRY TO GARDEN OR BEND OVER. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102565 | 2021-07-15 | 2021-07-01 | I, n, j, u, r, y | Voluntary report | PHILLIPS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS RESPIRONICS, INC |  |  |  | I CAN'T EXACTLY PIN POINT WHAT DATE, BUT I DID EXPERIENCE CHEST PAIN AND TIGHTNESS. I FELT LIKE I WAS GONNA PASS OUT AND COULD NOT BREATH, I HAVE HAD A COUGH THAT WILL NOT GO AWAY FOR WEEKS, IT FEELS AS IF I HAVE SOMETHING DRY AND GRITTY IN MY THROAT. MY BLOOD PRESSURE HAS BEEN HIGHER THAN USUAL AND I HAVE HAD HEADACHES MORE OFTEN. WHEN I WOULD USE THE MACHINE, I WOULD STOP HALFWAY THROUGH THE NIGHT BECAUSE I FELT LIKE I COULDN'T BREATH AND SOMETHING WAS IN MY THROAT. MY NOSE HAS BEEN CONSTANTLY PLUGGED UP, MAKING IT MORE DIFFICULT TO BREATH AT TIMES. I ALWAYS FEEL LIKE I HAVE A SINUS INFECTION THE NEXT MORNING AFTER USING MY CPAP. A FEW NIGHTS AGO, I WAS DOING MY SCHOOL WORK, WHEN ALL OF SUDDEN I GOT REALLY DIZZY, LIGHT HEADED, I FELT LIKE SOMEONE WAS SITTING ON MY CHEST, I HAD PAIN GOING DOWN MY LEFT ARM AND MY NECK. MY LEFT SIDE GOT VERY WEAK, AND MY CHEST WAS HURTING BAD. I TOOK MY NITRO PILL, AND WAITED ABOUT AN HOUR FOR THE SYMPTOMS TO GO AWAY, BUT I WAS VERY TIRED FOR THE NEXT FEW DAYS, IN FACT MY LEFT ARM IS STILL PRETTY WEAK, IT TIRES OUT EASILY. I HAVE HAD A TIA IN THE PAST, AND I AM ALMOST CERTAIN THAT IS WHAT HAPPENED. I STILL HAVE CHEST PAIN, SHORTNESS OF BREATH, HEADACHE, DIZZINESS AND THROAT IRRITATION WITH A COUGH RIGHT NOW. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102574 | 2021-07-15 | 2019-05-14 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS RESPIRONICS / RESPIRONICS, INC. | DSX500T11C |  |  | SINUS SURGERY; I HAVE BEEN TRYING TO STOP A CHRONIC COUGH AND ONGOING SINUSITIS SINCE (B)(6) 2018. SURGICAL INTERVENTION (TURBINATE REDUCTION) FOR CHRONIC SINUSITIS/RHINITIS ON (B)(6) 2019 WITH NO IMPROVEMENT. NO ALLERGIES WHEN TESTED, REGULAR VISITS TO ENT DOCTOR FOR HELP. AFTER MANY ROUNDS OF STEROID AND ANTIBIOTIC TREATMENT ALONG WITH MULTIPLE MEDICATIONS TO TREAT RESPIRATORY ISSUES, THE CHRONIC COUGH AND SINUSITIS PERSISTS THROUGH TODAY, (B)(6) 2021. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102570 | 2021-07-15 |  | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS RESPIRONICS CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. |  |  |  | I OFTEN WAKE UP WITH A SORE THROAT, DRY MOUTH. IT HAS A DIFFERENT SMELL AT TIMES. I REPORTED THIS RIGHT AFTER I RECEIVED THE NEW UNIT. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102568 | 2021-07-15 | 2017-06-20 | I, n, j, u, r, y | Voluntary report | CPAP NAME DREAMSTATION AUTO CPAP HUM DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. |  | DSX500H11 |  | I WAS DIAGNOSED WITH A SLEEP APNEA AND THE DOCTOR ORDERED A CPAP MACHINE. I PURCHASED A DREAMSTATION AUTO CPAP DOM SERIAL NUMBER (B)(4). I STARTED USING IT AND SLEPT WELL. AFTER A WHILE I WAS GIVEN AN MRI AND A SPICULATED NODULE WAS DISCOVERED IN MY RIGHT UPPER LUNG. I DON'T SMOKE AND AM NOT AROUND ANY SMOKE. I DON'T WORK IN A HOME GARDEN OF DO ANYTHING WITH SOIL. A SMALL BIOPSY TURNED UP NO CANCER CELLS. A PET SCAN SHOW INCREASED CELLULAR ACTIVITY IN THE NODULE AND SOME LYMPH NODES. IT WAS DECIDED TO REMOVE A WEDGED SHAPED PIECE OF LUNG THAT CONTAINED THE NODULE IN MY RIGHT UPPER LUNG LOBE. I WAS IN THE HOSPITAL FOR SEVERAL DAYS TO RECOVER AND RELEASED ON (B)(6) 2017 (MY BIRTHDAY). SAMPLES OF THE SPECULATED NODULE WERE SENT TO SEVERAL PATHOLOGY LABS AND THE PROBLEM COULD NOT BE IDENTIFIED. IT WAS NEVER OFFICIALLY ESTABLISHED. I TOLD THE DOCTOR THAT THE ONLY THING DIFFERENT IN MY LIFE WAS THE CPAP. AT THE TIME HE DOUBTED IT COULD BE THAT. I WAS WORRIED ENOUGH I STOPPED USING IT. ON (B)(6) 2017 I HAD BRAIN SURGERY FOR TRIGEMINAL NEURALGIA. I STILL DID NOT USE A CPAP AGAIN UNTIL MY BIRTHDAY ON (B)(6) 2018 WHEN MY SON PURCHASED A SOCLEAN DEVICE THAT CLEANS CPAP MACHINES. MY SUBSEQUENT MRIS SHOW SOME ACTIVITY IN THE AREA OF SURGERY AND MY DOCTOR THINKS IT'S JUST POST SURGERY CHANGES. I SAW THE CPAP RECALL ON THE TV NEWS AND IMMEDIATELY STOPPED USING THE CPAP MACHINE. I CALLED MY DOCTOR, MY CPAP SUPPLIER AND NO OF THEM KNEW ANYTHING ABOUT THE RECALL. I CALLED PHILIPS AND THEY ACKNOWLEDGED THE RECALL BUT SAID I COULD NOT REGISTER MY MACHINE WITH THEM UNTIL THEIR WEBSITE WAS SET UP AND IT WOULD BE THE FOLLOWING WEEK. I WAS ABLE TO REGISTER THE FOLLOWING WEEK AND THE SITE INDICATED MY SERIAL NUMBER WAS AMONG THOSE BEING RECALLED. I HAVE HAD A NUMBER OF DIFFERENT PHYSICAL SYMPTOMS THAT MAY BE RELATED TO THIS CPAP MACHINE. I DON'T KNOW. BUT I CAN'T THINK OF ANYTHING ELSE AS A CAUSE. I DON'T RECALL THE DATES OF MY BIOPSY, PET SCAN, OR EXACT LUNG SURGERY DATE. CPAP STOPPED DUE TO RECALL. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102564 | 2021-07-15 | 2021-05-05 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO BIPAP HUM DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. |  |  |  | RECALL ON BIPAP MACHINE. WAS HAVING SEVERE COUGHING, EYE, NOSE, IRRITATION, DIZZINESS. FEEL DRAINED, NO ENERGY. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02241 | 2021-07-15 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TO MAINTAIN SET PRESSURE SETTINGS. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR FAILED TO MAINTAIN SET PRESSURE SETTINGS. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
MW5102614 | 2021-07-16 | 2021-07-09 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | PHILIPS / RESPIRONICS, INC. | TRILOGY 100 |  |  | MY WIFE HAD ABDOMINAL SURGERY ON MARCH 10TH. SINCE THAT TIME SHE HAS BEEN THROUGH 4 ATTEMPTS AND REHABILITATION TO TRY TO GET HER HOME SAFELY. IN MARCH WHILE AT (B)(6) HOSPITAL SHE WAS DIAGNOSED WITH COPD EXASPERATED FROM A LUNG INFECTION. ON MARCH 30 SHE WAS PRESCRIBED PHILLIPS TRILOGY 100 BIPAP MACHINE. WE WENT FROM (B)(6) HOSPITAL TO (B)(6) HOSPITAL. WE WERE DISCHARGED FROM THERE TO HOME WHERE WE STARTED USING THIS PHILIPS PRODUCT. AFTER A SHORT STAY AT HOME WE WERE READMITTED TO (B)(6) HOSPITAL EMERGENCY ROOM FOR LOW BLOOD PRESSURE DUE TO DEHYDRATION. WE SPENT 6 DAYS AT (B)(6) HOSPITAL BEFORE BEING ADMITTED TO (B)(6) FACILITY. AFTER 10 DAYS AT THAT FACILITY WE WERE ADMITTED TO (B)(6) HOSPITAL EMERGENCY WITH A DRUG-RESISTANT UTI. AFTER A 14 DAY COURSE SPECIALIZED ANTIBIOTICS WE WERE DISCHARGED BACK TO(B)(6). ON JULY 9, AFTER SHE HAD SEVERAL DAYS A PHYSICAL DETERIORATION AND AN INABILITY TO PARTICIPATE IN PHYSICAL THERAPY CULMINATING IN POOR OXYGEN LEVELS IN THE BLOOD WE WERE RETURNED BACK TO (B)(6) EMERGENCY DUE TO BLOOD OXYGEN LEVELS IN THE LOW 80S TO MID 70%. CAT SCAN AT (B)(6) SHOWED A BLOOD CLOT IN HER RIGHT LUNG. THE BLOOD CLOT IS BEING TREATED WITH BLOOD THINNERS AND AFTER STABILIZATION WE WERE TRANSFERRED TO A DIFFERENT REHABILITATION FACILITY. ON MONDAY JULY 12 WE RECEIVED WRITTEN NOTICE THAT THE BIPAP MACHINE HAS BEEN RECALLED ON JUNE 14. THIS MACHINE HAS BEEN USED EVERY NIGHT SINCE WE WERE DISCHARGED TO OUR HOME. IT HAS BEEN MY OPINION THAT THE MACHINE DID HELP COMBAT THE EFFECTS OF COPD AND RESULTING CARBON DIOXIDE BUILD UP. HOWEVER THIS RECALL STATING THAT THEIR MACHINE CAN BE SLOUGHING OFF PARTICLES AND OUTGASSING CHEMICALS IS VERY TROUBLING. WHEN CALLING THE LOCAL SUPPLIER WE WERE TOLD THAT IT IS A PHILLIPS ISSUE. WHEN CALLING PHILLIPS TO FIND OUT EXACTLY WHAT AND WHEN THEY ARE GOING TO DO ABOUT THIS MACHINE AND IT'S RECALL WE WERE TOLD THAT IT'S UP TO THE FDA TO PROVIDE A REPAIR OR REPLACE SOLUTION. IN RECALLING OUR LOCAL SUPPLIER WE ARE TOLD OUR OPTIONS ARE WAIT FOR PHILLIPS SOLUTION OR GO TO A DOCTOR AND GET A NEW PRESCRIPTION WRITTEN FOR ANOTHER MACHINE. OF COURSE THIS IS ALL AT OUR EXPENSE. AND EVEN IF I COULD PULL OFF GETTING THAT DONE MY SUSPICION IS IT WILL TAKE A MONTH. I DO NOT BELIEVE MY WIFE HAS A MONTH TO DEAL WITH BUREAUCRATS AND COMPANY SALES PEOPLE CALL BLAMING EACH OTHER FOR THE DELAY IN HER TREATMENT. THERE ARE ALREADY CLASS ACTION SUITS AGAINST PHILLIPS FOR THIS ISSUE ALL CLAIMING KNEW THE PROBLEM EXISTED WAY BEFORE THE RECALL NOTIFICATION. AND THE RECALL NOTIFICATION SIMPLY WAS TIMED FOR THEIR NEW MODEL ROLL OUT. MY SUSPICION IS PHILIPS KNEW THIS WAS AN ISSUE BEFORE THIS UNIT WAS PRESCRIBED FOR MY WIFE. IF THIS IS TRUE I HAVE A HUGE PROBLEM WITH THE FDA ASSISTING IN THIS KINDLE OF BEHAVIOR BY AN INTERNATIONAL CORPORATION. OUR QUESTION FOR YOUR ORGANIZATION ARE YOU HOLDING UP THEIR SOLUTION TO THE RECALL AND IF YOU ARE WHEN MIGHT ONE EXPECT HAVE ANSWER AS TO WHEN THE SOLUTION WILL COME FORTH? FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102616 | 2021-07-16 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DERAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX700H11C |  |  | CAN'T CONTACT DEVICE MAKER; THERE IS A VOLUNTARY RECALL ON PHILLIPS RESPIRONICS CPAP MACHINES, THIS NEED TO BE MANDATORY. I TRIED TO REGISTER MY MACHINE FOR THE RECALL AND I AM UNABLE TO DO SO. THE PHILLIPS RECALL SITE WILL NOT ACCEPT MY INFORMATION. THE PHONE NUMBER PROVIDED HAS LIMITED HOURS AND A MESSAGE THAT SAYS HIGH CALL VOLUME. THIS NEEDS TO BE A MANDATORY RECALL BASED ON PHILLIPS IS NOT ADDRESSING THE PROBLEM. I HAVE BEEN REGISTERED WITH THEM, YET THEY FAILED TO NOTIFY ME MY MAIL OR EMAIL. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | I |  | 2021-07
MW5102618 | 2021-07-16 | 2020-12-01 | I, n, j, u, r, y | Voluntary report | CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | CAX400T12 |  |  | I BEGAN TO HAVE THROAT IRRITATION AND VISUAL REDNESS AT THE BACK AND SIDES OF MY THROAT IN (B)(6), 2020. I WENT TO A WALK-IN CLINIC IN (B)(6), 2021, AND A DOCTOR PRESCRIBED 5 DAYS OF ANTIBIOTICS TO TRY TO CLEAR UP ANY INFECTIONS. IT DIDN'T WORK. THE SWELLING AND SORENESS IN MY THROAT INCREASED. I WENT TO THE WALK-IN CLINIC AGAIN IN (B)(6), 2021, AND A DIFFERENT DOCTOR PRESCRIBED NETIPOT USE, AND AN ANTIHISTAMINE. I USED THESE FOR TWO WEEKS, AND THERE WAS NO CHANGE TO MY CONDITION. I HAD A WELLNESS TELEPHONE CONVERSATION WITH MY GP, AND HE ADVISED ME TO STOP DOING ANYTHING TO MY THROAT FOR ONE WEEK, AND THEN SEE IF THERE WAS ANY CHANGE. THERE WASN'T. I AM NOW SEEING AN EAR, NOSE, AND THROAT SPECIALIST WHO HAS ORDERED A CT SCAN ON MY THROAT IN ORDER TO FIND OUT WHAT MIGHT BE GOING WRONG WITH MY THROAT. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102620 | 2021-07-16 | 2021-07-13 | I, n, j, u, r, y | Voluntary report | REMSTAR AUTO, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. |  |  |  | I WAS DIAGNOSED WITH SLEEP APNEA IN (B)(6) OF 2015 AND WAS PRESCRIBED A CPAP MACHINE AND HAVE BEEN USING THE SAME MACHINE REGULARLY FOR YEARS. I ALSO HAVE A HISTORY OF CANCER IN MY LUNG. I'VE BEEN CANCER FREE FOR 4 YEARS AND GO FOR REGULAR CHECK UPS. I HAD A CT SCAN (B)(6) 2020 AND MY MOST RECENT SCAN WAS (B)(6) 2021. THE RESULTS OF MY (B)(6) SCAN OF MY LUNGS SHOWED "PLEURAL-PARENCHYMAL THICKENING IN THE LATERAL BASE OF THE LOWER LEFT LOBE. ALONG THE MEDICAL ASPECT OF THE LEFT LINGULA BY THE MEDIASTINUM, A 3.1 CM OVOID, LOBULATED, PLEURAL-BASED DENSITY IS PRESENT...". A PET SCAN WAS DONE AND IT TESTED NEGATIVE FOR CANCER. I WAS NOT SICK THIS PAST WINTER, SO THERE IS NO REASON I CAN THINK OF THAT MY LUNG WOULD SHOW NEW DAMAGE OR ADDITIONAL SCARRING UNLESS MY CPAP MACHINE HAS CAUSED THIS. I FOUND OUT (B)(6) 2021 ABOUT THE RECALL ON MY CPAP MACHINE AND I BELIEVE THAT IT IS POSSIBLE THAT IT CAUSED OR CONTRIBUTED TO NEW DAMAGE TO MY LUNG. FOR SEVERAL MONTHS NOW I HAVE EXPERIENCED INCREASED HEADACHES, COUGHING, TIGHTNESS IN MY CHEST, AND MORE DIFFICULTY BREATHING. I ALSO NOTICED SOME CHANGES WITH MY VOICE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02237 | 2021-07-16 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP PRO BIFLEX, SYS ONE 60 SRS, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DS660S | DS660S |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING/CHAMBER AND NASAL/THROAT IRRITATION OR SORENESS RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02261 | 2021-07-16 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02205 | 2021-07-16 | 2021-07-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | AUX500T15C | AUX500T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | SECTION B5 WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT SHOULD BE REPORTED AS FOLLOWS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION PATIENT STATE THAT HEADACHES AFTER USING THE DEVICE TIRED, VERY DRY THROAT AND COUGH . THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. SECTION H6 CODES CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT HAS BEEN CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02203 | 2021-07-16 | 2021-06-18 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED AN INFECTION. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A DOCTOR'S ASSESSMENT AND PRESCRIBED MEDICATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02234 | 2021-07-16 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED TO SEEING THE PARTICLES IN AIR PATH. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. SECTION A1 HAS BEEN CORRECTED/UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02248 | 2021-07-16 | 2017-03-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX400S12 | CAX400S12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED BLEEDING, SCAR TISSUE AND LUNG ISSUES. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON (B)(6)2024, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED BLEEDING, SCAR TISSUE AND LUNG ISSUES. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY THE PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02215 | 2021-07-16 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP AUTOSV H/C, DS | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | DSX900H11C | DSX900H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02239 | 2021-07-16 | 2021-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER COMPLETED AN INTERNAL VISUAL INSPECTION OF THE DEVICE. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERROR LOGGED. THE HUMIDIFIER WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR EVALUATION. THE MANUFACTURER FOUND EVIDENCE OF DIRT/DUST/KERATIN INSIDE OF THE HUMIDIFIER CHAMBER. THE MANUFACTURER CONCLUDES NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. SECTION D9, G3, H6 HAS BEEN UPDATED / CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02201 | 2021-07-16 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | CAX500T12C | CAX500T12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED ON THIS DEVICE IN MDR 2518422-2021-02118. THIS REPORT WAS SUBMITTED AS A DUPLICATE REPORT OF THE PREVIOUSLY SUBMITTED REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02251 | 2021-07-16 | 2021-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | KRX400T15 | KRX400T15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND ALLEGEDLY CAUSED BRONCHIECTASIS IN A PATIENT. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE SECTIONS B1, B2, H1 AND H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM. (IN THE PREVIOUS MDR BOTH ADVERSE EVENTS AND PRODUCT PROBLEM WAS CHECKED.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CORRECTED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECTS - CLINICAL CODES AND HEALTH EFFECTS - IMPACT CODE WERE CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02213 | 2021-07-16 | 2021-06-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP PRO, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX600S11 | DSX600S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP HEADACHES AND KIDNEY CANCER. THE PATIENT RECEIVED MEDICAL INTERVENTION IN THE FORM OF SURGERY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE /RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM THAT BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP HEADACHES AND KIDNEY CANCER, NASAL/THROAT IRRITATION, SORENESS, FEELING LIGHTHEADEDNESS, DIZZINESS, DIFFUCITY BREATHING, AND SHORTNESS OF BREATH, ALSO HE UNDERWENT SURGERY. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HEALTH EFFECT- CLINICAL CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02260 | 2021-07-16 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S BATTERIES WERE DEPLETED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S POWER SUPPLY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02250 | 2021-07-16 | 2021-06-25 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX400T12C | CAX400T12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM DEGRADED AND THE PATIENT WAS DIAGNOSED WITH CANCER. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. SECTION B5 SHOULD BE PREVIOUSLY REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT WAS DIAGNOSED WITH CANCER AND COMPLETED HER RADIATION TREATMENT. THE EVENT OF CANCER IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURER CONCLUDES NO FURTHER ACTION IS NECESSARY. PATIENT ALSO ALLEGED DIFFICULTY BREATHING, SHORTNESS OF BREATH, FATIGUE, NAUSEA, VOMITING, LIGHTHEADEDNESS, DIZZINESS, LOTS OF PAIN, HEADACHES, COUGHING, SNEEZING, AND HYPERSENSITIVITY TO SOUND. NO OTHER MEDICAL INTERVENTION WAS REPORTED BY PATIENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED. SECTION D4 UNIQUE IDENTIFIER (UDI) # CORRECTED. SECTION H6 HEALTH EFFECT - CLINICAL CODE WHICH WAS MISSED IN PREVIOUS REPORTS WAS CAPTURED. SECTION D8, D9, H2, H3 AND H6 TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102627 | 2021-07-19 | 2021-01-01 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION BIPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | REPORTER STATES THAT PT CAME TO HER PRACTICE IN (B)(6) 2021, FEELING UNWELL AND STATES THAT HIS SYMPTOMS WORSENED IN (B)(6) 2021. REPORTER STATES THAT PT WAS HOSPITALIZED TWICE AND DX WITH BILATERAL LUNG ABSCESSES. REPORTER STATES THAT DESPITE DX, DOCTORS HAVE NOT BEEN ABLE TO UNDERSTAND THE CAUSE AND AN APPROPRIATE TREATMENT PLAN FOR THE ABSCESSES. REPORTER ALSO NOTES THAT PT WAS SEEN BY 5 DOCTORS AT (B)(6) AND THE PT WAS D/C WITHOUT AN ANSWER TO THE ORIGIN OF HIS HEALTH ISSUES. REPORTER STATES THAT PT PRESENTED TO HER WITH COUGH, GENERAL UNWELL FEELING, HEAD PRESSURE AND IRRITATION IN ADDITION TO THE LUNG ABSCESSES. REPORTER STATES SHE WENT THROUGH PT HISTORY WITH THE WIFE AND AFTER LEARNING ABOUT HIS 75 POUND WEIGHT LOSS, NEGATIVE BRONCHOSCOPY, SUDDEN SYMPTOM ONSET AND THE RECALLED DEVICE, SHE BELIEVES THERE IS A POSSIBILITY THE HEALTH ISSUES ARE A RESULT OF THE BIPAP MACHINE. REPORTER STATES RESEARCHING THE RECALL AND NOTING THIS DEVICE HAD BEEN ASSOCIATED WITH FOAM BREAKDOWN AND BLACK PARTICULATE CONTAMINATION WHICH SHE BELIEVES COULD BE CAUSE OF THE LUNG ABSCESSES. OF NOTE, REPORTER STATES THAT THE DEVICE WAS RECALLED AND REPLACED BY THE DME COMPANY FRESH AIR. REPORTER STATES THAT PT IS HAVING PERSISTING SYMPTOMS EVEN AFTER DEVICE REPLACEMENT. REPORTER IS CURRENTLY CONCERNED WITH FINDING AN APPROPRIATE TREATMENT PROTOCOL FOR THE PATIENT. |  | Y | Y | I |  | 2021-07
2518422-2021-02253 | 2021-07-19 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/RESPIRONICS, INC | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02274 | 2021-07-19 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO, INTL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | 761P | 761P |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | IN THE PREVIOUS FOLLOW-UP MDR SUBMISSION, IN SECTION B.2 OF THE 3500A FORM THE "REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE" BOX WAS CHECKED IN ERROR. THERE WAS NO REQUIRED INTERVENTION REPORTED. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02308 | 2021-07-19 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICES LCD DISPLAY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02132 | 2021-07-19 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S BLOWER MOTOR AND FLOW SENSOR ASSEMBLY WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02277 | 2021-07-19 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE HEADACHE,DRY MOUTH,NAUSEA,SKIN IRRITATION. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE EVENT IS CURRENTLY UNKNOWN. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 (HEALTH IMPACT CODE), B7 WAS MISSED TO CAPTURE IN INITIAL/FOLLOW UP MDR ,NOW IT IS CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02278 | 2021-07-19 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN INITIAL REPORT SECTION B5 WAS MENTION INCORRECTLY, CORRECT B5 SHOULD BE - THE PATIENT ALLEGED WHITE PARTICLES IN AIR PATH, PHLEM. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INPSECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND BLACK CONTAMINANT CONSISTENT WITH KERATIN IN THE AIR OUTLET. AN INTERNAL VISUAL INPSECTION WAS COMPLETED BY THE MANUFACTURER AND FOUND AN UNKNOWN DARK BROWN CONTAMINATION AT THE BOTTOM OF THE BLOWER BOX AND POTENTIAL MINERAL DEPOSITS ON THE BLOWER MOTOR AND THE BLOWER BOX, INDICATING POTENTIAL WATER INGRESS. ALSO FOUND BROWN AND WHITE CONTAMINANTS AND A POTENTIAL DARK-COLORED HAIR INSIDE OF HUMIDIFIER BOX. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND 19 ERRORS LOGGED WHILE WAS POWERING DEVICE ON. THE MANUFACTURER CONCLUDES A MINERAL DEPOSTIE CONTAMINATION FOUND WERE CONSISTENT WITH WATER INGRESS AND PRESENCE OF MULTIPLE CONTAMINANTS THAT ARE NOT CONSISTENT WITH DEGRADED SOUND ABATEMENT. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. IN THIS REPORT, SECTION D9, D10, G3, H3, H6 HAS BEEN UPDATED OR CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02287 | 2021-07-19 | 2021-05-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED LOUD NOISE COMING FROM THE MACHINE ALMOST SOUNDS LIKE RATTLING SOUND. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. ADDITIONAL INFORMATION WAS RECEIVED AND ADDED TO THE REPORT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER ON 05/14/2021 FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED ON 05/18/2021. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED AND OBSERVED WATER INGRESS. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE 4 ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. UNIT SCRAPPED A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING LOUD NOISE COMING FROM THE MACHINE ALMOST SOUNDS LIKE RATTLING SOUND RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION D8, D9, H2, H3 AND H6 WERE UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02214 | 2021-07-19 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED ASTHMA AND RESPIRATORY ISSUES IN A PATIENT. THE PATIENT RECEIVED MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTION MEDICATIONS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 12, 2024, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICESTHE MANUFACTURER RECEIVED INFORMATION ALLEGING ASTHMA AND RESPIRATORY ISSUES IN A PATIENT, RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTIONS B1, B2, H1 AND H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR ONLY THE PRODUCT PROBLEM (BOTH ADVERSE EVENTS AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02280 | 2021-07-19 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A NODULE ON THEIR THYROID AND A SPOT ON THEIR LUNG. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 12 , 2024 AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING TO DEVELOP A NODULE ON THEIR THYROID AND A SPOT ON THEIR LUNG RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE.THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02240 | 2021-07-19 | 2021-06-18 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02238 | 2021-07-19 | 2021-06-18 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THE PATIENT IS UNDERGOING AN UNSPECIFIED TREATMENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP CANCER. THE PATIENT IS UNDERGOING AN UNSPECIFIED TREATMENT. THE REPORTED EVENT OF HAVING CANCER WAS REVIEWED BY THE PMS CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02249 | 2021-07-19 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | INX500S15 | INX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02290 | 2021-07-19 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING BLACK PARTICLES IN CHAMBER AND TUBING RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02289 | 2021-07-19 | 2021-06-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02256 | 2021-07-19 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING BLACK PARTICLES IN THE WATER IN THE RESERVOIR AND HOSE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK PARTICLES IN THE WATER IN THE RESERVOIR AND HOSE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. AFTER THE FINAL ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, CUSTOMER DECLINED TO RESPOND TO THE GFE RELATED QUESTIONS AND TERMINATED THE CALL. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTIONS A1 NOT CAPTURED IN THE INITIAL REPORT WAS CAPTURED AND SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02252 | 2021-07-19 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500S11 | DSX500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02316 | 2021-07-19 | 2021-07-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S KEYPAD WAS OBSERVED. THE DEVICES FRONT PANEL BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02315 | 2021-07-19 | 2021-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICES LCD DISPLAY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02314 | 2021-07-19 | 2021-07-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S OVERHEAD BLOWER FILTER WAS REPLACED TO ADDRESS THE ISSUE. THERE WAS EVIDENCE OF CONTAMINATION. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02313 | 2021-07-19 | 2021-07-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1032804 | 1032804 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02312 | 2021-07-19 | 2021-07-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE OVERHEAD BLOWER FILTER WAS REPLACED TO ADDRESS THE ISSUE. THERE WAS EVIDENCE OF CONTAMINATION. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02311 | 2021-07-19 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR FAILED A TEST STEP AND THE SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. THE SYSTEM BOARD WAS NOT REPLACED. THE SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02310 | 2021-07-19 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02309 | 2021-07-19 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICES LCD DISPLAY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02307 | 2021-07-19 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION OCCURRED AND THE INTERNAL BATTERY DID NOT CHARGE. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S POWER MANAGEMENT BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02263 | 2021-07-19 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, A LEAK WAS OBSERVED FROM THE EXHALATION VALVE. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02265 | 2021-07-19 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S POWER MANAGEMENT BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02247 | 2021-07-19 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TO POWER OFF. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DISPLAY FAILED TO RESPOND. THE DEVICE'S DISPLAY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02264 | 2021-07-19 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02262 | 2021-07-19 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR POWERED OFF WHEN PLUGGED INTO THE AC OUTLET. THE DEVICE WAS NOT IN USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02141 | 2021-07-19 | 2021-07-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED A DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS A "SERVICE REQUIRED" CODE FOUND IN THE VENTILATORS DOWNLOADED ERRROR LOG. THE REPORT HAD "DEATH" SELECTED IN SECTION B.2. OF THE MDR REPORT. THE WAS SELECTED IN ERROR. THERE WAS NO DEATH ASSOCIATED WITH THIS OCCURRENCE. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02140 | 2021-07-19 | 2021-07-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER HAD PREVIOUSLY REPORTED THAT A "SERVICE REQUIRED" ERROR CODE WAS FOUND IN THE DEVICE'S EVENT LOG AND THE OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. SECTION B.2 HAD AN INCORRECT SELECTION OF "DEATH". THERE WAS NO ADVERSE EVENT OR DEATH ASSOCIATED WITH THIS OCCURRENCE. THE OUTCOME OF DEATH WAS SELECTED IN ERROR. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02288 | 2021-07-19 | 2021-05-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11 | DSX500T11 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CRF PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02349 | 2021-07-20 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED THE INCORRECT 510K NUMBER IN SECTION G4. THE CORRECT NUMBER IS K083526 AND IS REFLECTED IN THIS REPORT. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02324 | 2021-07-20 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02269 | 2021-07-20 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500S11 | DSX500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO THE DEVICE'S SOUND ABATE FOAM. THE END USER INDICATED OZONE OR OTHER UNAPPROVED CLEANING AND DISINFECTION METHOD WAS USED. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02291 | 2021-07-20 | 2021-06-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO THE DEVICE'S SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02292 | 2021-07-20 | 2021-06-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11 | DSX400H11 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT ALSO ALLEGED BLACK STUFF IN MACHINE AND HAVING NASAL/THROAT BURNING AND IRRITATION. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL PART OF THE DEVICE. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION H6 WERE UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02267 | 2021-07-20 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER PREVIOUSLY RECEIVED INFOMATION AN ALLEGING LUMP ON HIS BODY, CHEST PAIN, SINUS PROBLEM, DIZZINESS, HEADACHE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTIONS A1, A3 AND B7 WERE CORRECTED IN THIS REPORT. SECTION H6 CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02266 | 2021-07-20 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECOME DEGRADED AND THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT ALSO ALLEGED HAVING HEADACHE. THE DEVICE WAS RETURNED TO THE MANUFACTURER QUALITY PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. DURING THE INITIAL VISUAL INSPECTION OF THE DEVICE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN IN THIS DEVICE. THE DEVICE WAS DISASSEMBLED FOR INTERNAL VISUAL INSPECTION. THE MANUFACTURER FOUND EVIDENCE OF EXTERNAL CONTAMINATES PRESENT IN THE UNIT. THE MANUFACTURER APPLIED POWER TO THE UNIT AND PROVIDED THE FLOW. THE MANUFACTURER REVIEWED THE DEVICE ERROR LOG. THE MANUFACTURER CONCLUDES THAT THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE CONTAMINATION FOUND WITHIN THE DEVICE IS CONSISTENT WITH BEING FROM AN EXTERNAL SOURCE. SECTION D8, D9, H3, H6 WERE UPDATED / CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02370 | 2021-07-20 | 2021-07-15 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT WAS DIAGNOSED WITH CANCER. THE DETAILS OF THE MEDICAL INTERVENTION THAT THE PATIENT REQUIRED ARE CURRENTLY UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02233 | 2021-07-20 | 2021-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRABBED DUE TO AGE". | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02271 | 2021-07-20 | 2021-06-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT ALSO ALLEGED BLACK STUFF IN MACHINE AND HAVING NASAL/THROAT BURNING AND IRRITATION. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL PART OF THE DEVICE AND FOUND MINIMAL SIGNS OF DUST CONTAMINATION IN THE HUMIDIFIER CHAMBER. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND FOUND MINIMAL DUST CONTAMINATION IN THE BLOWER ASSEMBLY. THE HUMIDIFIER SHOWED A SMALL AMOUNT OF LIQUID INGRESS UNDER THE BOTTOM PLATE. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION BUT A SMALL AMOUNT OF LIQUID INGRESS AND DUST CONTAMINATION WERE OBSERVED AND ARE CONSISTENT WITH BEING FROM AN EXTERNAL SOURCE. SECTION E1 AND H6 WERE UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02272 | 2021-07-20 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | SECTION B5 WAS INCORRECTLY REPORTED IN THE INITIAL REPORT. IT HAS TO BE THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE BLEEDS, HEADACHE, DIFFICULTY BREATHING/SHORT OF BREATH AND STRANGE ODOR FROM THE DEVICE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF HARM OR INJURY. SECTION H10 IN THE PREVIOUS REPORT WAS INCORRECT. IT HAS TO BE THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED NOSE BLEEDS, HEADACHE, DIFFICULTY BREATHING/SHORT OF BREATH, STRANGE ODOR FROM THE DEVICE. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THERE WAS UNKNOWN DUST CONTAMINANT ON THE TOP ENCLOSURE. THE REAR ENCLOSURE APPEARED TO HAVE HAIR-LIKE PARTICLES ON IT. THE BOTTOM ENCLOSURE WAS OBSERVED TO HAVE A DARK UNKNOWN CONTAMINANT ON IT. AN UNKNOWN DUST CONTAMINANT WAS OBSERVED ON THE BLOWER BOX. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT DIRECTLY ADDRESS THE SYMPTOMS OR COMPLAINTS AND THERE WAS NO VISIBLE FOAM DEGRADATION, BUT DUST CONTAMINATION WAS FOUND ON TOP ENCLOSURE, BLOWER BOX AND REAR ENCLOSURE APPEARED TO HAVE HAIR-LIKE PARTICLES, BOTTOM ENCLOSURE WAS OBSERVED TO HAVE A DARK UNKNOWN CONTAMINANT ON IT. SECTIONS H6 (HEALTH EFFECT-CLINICAL CODE, MEDICAL DEVICE PROBLEM CODE, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS) CORRECTED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02327 | 2021-07-20 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION EXPERT W/HUM/HT/CELL, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | CAX501T12C | CAX501T12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102635 | 2021-07-20 | 2021-02-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS RESPIRONICS INC |  |  |  | CONTINUING NASAL CONGESTION, PERSISTENT COUGH. (REGISTERED/CONFIRMED WITH PHILLIPS AS PART OF RECALL). FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102647 | 2021-07-20 | 2020-12-18 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | I BEGAN EXPERIENCING HEADACHES, NAUSEA, AND VOMITING IN (B)(6) 2020. I HAD NOT PREVIOUSLY EXPERIENCED THESE SYMPTOMS WITH THIS FREQUENCY-WEEKLY-A FEW DAYS EACH WEEK. I AM BEGINNING TREATMENT TO ABATE THESE SYMPTOMS. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102648 | 2021-07-20 | 2021-03-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS RESPIRONICS / RESPIRONICS | DSX500T11 | DSX500T11 |  | I HAVE BEEN USING THE PHILLIPS DREAMSTATION CPAP FOR YEARS. SINCE MARCH I HAVE BEEN EXPERIENCING BLOATING, GAS, AND ESOPHAGEAL DISCOMFORT. I HAD AN UPPER GI DONE, AND GASTRITIS WAS FOUND. I HAVE NOT BEEN IMPROVING. I JUST FOUND OUT ABOUT THE RECALL, AND I AM CONCERNED THAT THE FOAM HAS BEEN HARMFUL FOR ME. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102656 | 2021-07-20 | 2020-07-23 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILPS / RESPIRONICS, INC. | DREAMSTATION |  |  | I BEGAN USING MY FAULTY MACHINE, THAT IS REGISTERED WITH PHILLIPS NOW, (B)(6) 2020. APX 2 MONTHS LATER I WAS DIAGNOSED WITH STAGE LV ESOPHAGEAL CANCER. UNTIL NOW, THE CAUSE HAS BEEN UNKNOWN. I HAVE BEEN TOLD THAT MY CANCER IS TREATABLE BUT NOT CURABLE. I WAS GIVEN A GRIM OUTLOOK OF 50% OF PEOPLE WITH THIS CANCER, DIE WITHIN 6 MONTHS, AND 10% MAKE IT TO A YEAR. I AM THANKFUL THAT THE TRIAL DRUG TREATMENT I AM ON IS HELPING. GOD IS GOOD TO ME. MY DR SAYS I NEED THE MACHINE, SO NOT TO STOP USING IT, BUT TO GET SOME TYPE FILTER FOR IT. I HAVE A PROTEIN IN MY BODY THAT FEEDS THIS TYPE CANCER, HER2 +3. IT IS THE REASON IT HAS SPREAD SO FAST. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102659 | 2021-07-20 | 2021-07-12 | I, n, j, u, r, y | Voluntary report | DREAMSTATIONSTNAUTOCPAPHUMCELL DOM DREAMSTATION STN AUTO CPAP HUM CELL DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS / RESPIRONICS INC | DreamStation Auto CPAP w/Humidifier, w/Cell, US | DSX500H11C |  | HARSH COUGH NONPRODUCTIVE, COARSE VOICE, INABILITY TO CLEAR THROAT FROM MUCUS. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-07
MW5102671 | 2021-07-20 | 2021-07-18 | I, n, j, u, r, y | Voluntary report | "DREAMSTATION" CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | I USE THE RESPIRONICS CPAP "DREAM STATION" WHICH HAS BEEN RECALLED. I NOTICED THAT MY ASTHMA HAD GOTTEN WORSE, AND I HAVE BEEN EXPECTORATING SPUTUM. PLEASE NOTE THAT THE COMPANY I GOT THE MACHINE FROM, (B)(6), HAS OBSTRUCTED MY ATTEMPTS TO GET INFORMATION OR A REPLACEMENT MACHINE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-07
MW5102672 | 2021-07-20 | 2021-07-05 | I, n, j, u, r, y | Voluntary report | PHILLIPS DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS / RESPIRONICS, INC. | DSX500H11C |  |  | ENDED IN URGENT CARE WITH ACUTE BACTERIAL SINUSITIS AND SEVERE HEADACHE. FLUID IN THE EARS. WENT TO DOCTORS 2 DAYS IN A ROW DUE TO VISION DISTURBANCE, HEADACHE AND DIZZINESS. STILL HAVING DIZZINESS SPELLS. CURRENTLY USE A PHILLIPS DREAMSTATION CPAP MACHINE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102675 | 2021-07-20 | 2020-10-31 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | DIZZINESS PAINFUL RED SWOLLEN ITCHING SKIN BREAKOUT, CLOGGED BREATHING PASSAGES. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-01871 | 2021-07-20 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | RIN2110X15B | RIN2110X15B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION BUT THE EVALUATION IS NOT COMPLETE. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR WAS RETURNED FOR SERVICE. DURING THE EVALUATION, A SERVICE REQUIRED CODE WAS OBSERVED IN THE DOWNLOADED EVENT LOG. THE ERROR WAS CONFIRMED IN THE LOG BUT NO ABNORMALITY OR MALFUNCTION WAS OBSERVED WITH THE VENTILATOR. THERE WAS EVIDENCE OF WATER INGRESS. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02317 | 2021-07-20 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02318 | 2021-07-20 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICES LCD DISPLAY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02319 | 2021-07-20 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICES LCD DISPLAY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02346 | 2021-07-20 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02347 | 2021-07-20 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02348 | 2021-07-20 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02268 | 2021-07-20 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DEVICE MAKING NOISE, GETTING COUGH, DIZZINESS AND COLD SYMPTOMS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. IN INITIAL REPORT, SECTION A1 WAS NOT CAPTURED, SECTION G1 WAS WRONGLY CAPTURED, SECTION G3 WAS CAPTURED AS (B)(6) 2021 SHOULD BE (B)(6) 2021 AND SECTION H6 CLINICAL CODES WERE NOT CAPTURED, CORRECTED ALL THESE FIELDS AND UPDATED IN THIS REPORT. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02258 | 2021-07-20 | 2021-05-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE 1 ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS VISIBLE FOAM DEGRADATION. SECTION H6 WERE UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02282 | 2021-07-20 | 2021-06-24 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A RASH ON THE FACE WHICH CAUSED A SCAR AND COLON CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING TO DEVELOP A RASH ON THE FACE WHICH CAUSED A SCAR AND COLON CANCER,NASA/THROAT IRRITATION AND SORENESS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE EVENT IS CURRENTLY UNKNOWN. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HEALTH EFFECT- CLINICAL CODE WAS MISSED TO CAPTURE IN PREVIOUS MDR,NOW IT IS CORRECTED AND UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02342 | 2021-07-20 | 2021-06-14 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP PNEUMONIA AND CONGESTIVE HEART FAILURE. THE REPORTED INFORMATION STATES THE PATIENT WAS HOSPITALIZED TWICE IN RESPONSE TO THE REPORTED EVENTS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT TO DEVELOP PNEUMONIA AND CONGESTIVE HEART FAILURE. THE REPORTED INFORMATION STATES THE PATIENT WAS HOSPITALIZED TWICE IN RESPONSE TO THE REPORTED EVENTS. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF MULTIPLE CONTAMINANTS THROUGHOUT DEVICE AND WATER INGRESS IN THE AIR PATHWAY. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE MANUFACTURER WAS UNABLE TO READ ERROR LOG DUE TO CHECK POWER SUPPLY ERROR. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH WATER INGRESS AT P4 PORT AND CONTAMINATION. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION D9, G3 AND H6 HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02320 | 2021-07-20 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR POWERED OFF DURING USE. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE DOWNLOADED EVENT LOG SHOWS THE DEVICE REBOOTED. THE DEVICE'S DISPLAY BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02322 | 2021-07-20 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR WAS ALARMING FOR A HIGH INTERNAL OXYGEN ALARM CONDITION. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02321 | 2021-07-20 | 2021-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
MW5102636 | 2021-07-20 | 2021-06-01 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS RESPIRONICS DREAM MACHINE FOR SLEEP APNEA | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS INC | DSX500H11C |  |  | DRY THROAT, DIZZINESS, COUGH; I HAVE A RECALLED PHILIPS DREAM MACHINE FOR SLEEP APNEA. I HAVE HAD A DRY MOUTH, DRY COUGH AND RASPY VOICE , DIZZINESS AND HEADACHES SINCE THE FIRST OF (B)(6). I JUST FOUND OUT BY ACCIDENT ABOUT THE RECALL. WE CALLED PHILIPS AND THEY WERE VERY RUDE AND NO HELP AT ALL. I HAVE A DOCTORS APPOINTMENT (B)(6) FOR A CHECK UP AND TALK ABOUT MY THROAT AND VOICE. I KNOW THE MACHINE HAS DONE THIS BECAUSE I STOPPED USING IT 2 DAYS AGO AND MY THROAT FEELS A BIT BETTER. I HAVE READ MANY COMPLAINTS ABOUT THIS SAME PROBLEM. I ALSO FOUND BLACK RESIDUE ON MY HOSE AND ON THE WATER CHAMBER. I AM A CANCER SURVIVOR AND AM SO AFRAID OF WHAT THIS HAS DONE TO ME. I ALSO HAVE ASTHMA. MY INSURANCE WON'T GET ME A NEW MACHINE BECAUSE IT'S ONLY 2.5 YEARS OLD. MY HUSBAND ALSO HAS ONE AND WE ARE BOTH ON(B)(6) AND OVER 65. I STOP BREATHING UP TO 28 TIMES AN HOUR AND NOW I HAVE NO DEVICE TO USE. THIS IS VERY WRONG THAT WE WERE NEVER NOTIFIED AT ALL. WE ARE GOING TO HAVE TO SUFFER MEDICAL PROBLEMS AND POSSIBLE BRAIN DAMAGE BECAUSE OF NOT USING THIS MACHINE. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102638 | 2021-07-20 | 2021-05-21 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAM STATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS INC |  |  |  | I HAVE USED THE PHILIPS DREAM STATION AUTO BIPAP SINCE (B)(6) 2017. ABOUT 6-7 MONTHS AGO I SAW SOME OCCASIONAL BLACK RESIDUE NEAR THE WATER RESERVOIR AND AT THE BASE OF THE TUBING CONNECTOR. I WASHED IT OFF WITH SOAPY WATER AND VINEGAR AS DIRECTED. IT CONTINUED TO COME BACK. IN MID (B)(6) 2021 I BEGAN WAKING UP IN THE MIDDLE OF THE NIGHT WITH NAUSEA, DIZZINESS AND A HEADACHE. I WOULD GO OUT ONTO A RECLINER TO SLEEP WITHOUT MY BIPAP AND SYMPTOMS IMPROVED. AS SOON AS MY DME COMPANY INFORMED ME OF THE PHILIPS RECALL IN (B)(6) I REGISTERED MY MACHINE SERIAL NUMBER WITH PHILIPS AND THE WEBSITE CONFIRMED MINE IS RECALLED. I WAS ABLE TO GET A LOANER REAMED BIPAP AND AM PAYING (B)(6) PER MONTH WHICH PHILIPS SAID THEY WILL NOT REIMBURSE WHEN I CALLED THEM. I STILL HAVE HEARD NOTHING FROM PHILIPS ABOUT THE RECALL. MY ASTHMA IS WORSE THAN IT EVER HAS BEEN, I CONTINUE TO HAVE SOME OTHER VAGUE SYMPTOMS THAT I WILL ADDRESS WITH MY DR NEXT WEEK. I AM A RESPIRATORY THERAPIST SO I KEEP VERY GOOD TRACK OF MY ASTHMA. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102640 | 2021-07-20 | 2021-07-12 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. | DREAMSTATION |  |  | ER VISIT; I HAVE BEEN USING THE PHILIPS DREAMSTATION CPAP MACHINE FOR 3 MONTHS AND HAVING MANY HEALTH ISSUES. I HAVE A RECORD OF MANY DR/HOSPITAL VISITS BECAUSE OF THE CPAP MACHINE. I AM FINDING OUT NOW THAT THE MACHINE HAS BEEN RECALLED. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102642 | 2021-07-20 | 2021-06-16 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION PAP DEVICE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS / RESPIRONICS | DSX500T11C |  |  | THE COMPANY HAS ISSUED A RECALL BUT HAS DONE NOTHING TO FIX THE ISSUE OR PROVIDE A DATE TO FIX THE ISSUE. THEY ARE COMPLETELY AND UTTERLY LOST. I NEED THE MACHINE FOR SLEEP AND THIS RECALL HAS AFFECTED MY DAY TO DAY LIFE AND THEY HAVE AGAIN PROVIDED NOT DATE ON HOW TO FIX IT. I WOULD JUST LIKE A FULL REFUND AND RETURN IT SO THAT I CAN GET ANOTHER COMPANY MACHINE BECAUSE THESE GUYS ARE LOST. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-07
MW5102644 | 2021-07-20 | 2020-09-30 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION 1 | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | SINUSITIS WITH NASAL OBSTRUCTION AND POSTNATAL DRIP, IMPAIRED VISION AFTER RISING IN THE MORNING, COUGH SPUTUM, THESE SYMPTOMS IMPROVED AFTER I SWITCHED TO A RESMED DEVICE. EXPERIENCED WITH USE OF PHILIPS DREAMSTATION 1. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-07
MW5102646 | 2021-07-20 | 2021-07-16 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS SYSTEM ONE CPAP MACHINE | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | PHILIPS / RESPIRONICS, INC. | SYSTEM ONE |  |  | I USE A PHILLIPS CPAP MACHINE FOR SLEEP APNEA AND WAS INFORMED BY MY MEDICAL EQUIPMENT SUPPLIER BY EMAIL ON (B)(6) 2021 THAT MY CPAP MACHINE WAS PART OF THE RECALL. I HAVE USED A CPAP MACHINE EVERY NIGHT SINCE 1997. I HAVE SYMPTOMS OF RECURRING SINUS INFECTIONS, INFLAMMATION, CHEST PAIN, SHORTNESS OF BREATH, TIGHTNESS IN MY CHEST, HEADACHES, EYE ISSUES, LIVER AND KIDNEY ISSUES, HEART PALPITATIONS, DIZZINESS, NAUSEA AND VOMITING. I AM GOING TO MAKE AN APPOINTMENT TO SEE MY PRIMARY CARE DOCTOR TO DISCUSS MY SYMPTOMS. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102653 | 2021-07-20 |  | I, n, j, u, r, y | Voluntary report | PHILIPS SLEEP CPAP, DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | PROBLEM BREATHING WITH MACHINE USAGES. BUILD UP OF MUCUS IN THE THROAT OVER SEVEN YEARS. PROBLEM CONTINUED TO WORSEN. REPORTED THE ISSUE TO SEVERAL DOCTORS BUT NOTHING WAS RESOLVED. I TOOK MYSELF OFF THE MACHINE ABOUT FOUR MONTHS AGO. MACHINE WAS SMELLING FUNNY AND CAUSING ME TO "BREATH" IN FUMES. WHICH TASTED BAD. I BEGAN TO HAVE HEART ISSUES AND ENDED UP IN THE HOSPITAL WITH PNEUMONIA. I HAD COVID-19 SYMPTOMS. I EXPLAINED IT WAS ME USING THE BREATHING MACHINE. I HAD BEGAN TO HAVE A FUNNY TASTE IN MY MOUTH WHILE ON THE MACHINE. CONSTANTLY TALKING IN MY SLEEP. I STILL HAVE THE MACHINE. (B)(6). FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102655 | 2021-07-20 | 2020-04-15 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION BIPAP | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | PHILIPS / RESPIRONICS, INC. |  |  |  | I USE A PHILIPS DREAMSTATION BIPAP MACHINE FOR OBSTRUCTIVE SLEEP APNEA AND FROM (B)(6) 2020 FORWARD HAVE CONSISTENTLY HAD PROBLEMS WITH THE MACHINE. IT SEEMS LIKE THE HUMIDIFIER PLATE IS BURNING, AS THE MACHINE OFTEN GIVES OFF A BURNING PLASTIC SMELL. I'VE HAD CHRONIC UPPER RESPIRATORY AND SINUS PROBLEMS SINCE THAT DATE AND OCCASIONALLY WAKE UP SICK TO MY STOMACH FOLLOWING OVERNIGHT USAGE. I REPLACED THE MACHINE WITH A NEW MODEL OF THE SAME MACHINE IN EARLY 2021 AND HAVE CONTINUED TO EXPERIENCE SYMPTOMS. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102657 | 2021-07-20 | 2021-03-09 | I, n, j, u, r, y | Voluntary report | PHILIPS BREATHING MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | UNK |  |  | I HAD BEEN COMPLAINING ABOUT MY BREATHING MACHINE STARTING TO MAKE ME FEEL BAD. I BECAME VERY RESTLESS EVERY NIGHT WITH MACHINE ON. IT BEGAN TO BECOME VERY DIFFICULT TO SLEEP WITH MACHINE ON. BREATHING BECAME WORSE. COUGHING INCREASED. MUCUS BUILD UP BECAME WORSE IN AIRWAY. MADE SEVERAL REPORTS TO DOCTORS BUT TO NO AVAIL. GOING TO EMERGENCY ROOM MORE OFTEN. AT ONE POINT I HAD A FUNNY SMELL COMING FROM WITHIN MY MACHINE. KEPT MACHINE AS CLEAN AS I COULD. THE MORE I COMPLAINED TO THE DOCTOR ABOUT WANTING TO STOP THE USE OF THE MACHINE. THE MORE I WAS ENCOURAGED TO CONTINUE ON. I BEGAN TO HAVE BADLY BLURRED VISION. PRESCRIBED EYE WARE DIDN'T HELP. BEGAN TO HAVE CHRONIC HEADACHES. INCREASED CHEST PAIN AND HEART BEAT. ENDED UP BEING HOSPITALIZED FOR HEART ISSUES, PNEUMONIA, HIGH FEVER, AND HALLUCINATIONS. IT WAS THOUGHT TO BE COVID-19. ALL THREE COVID TEST. CAME BACK NEGATIVE. I DISCONTINUED THE USAGE OF MY BREATHING MACHINE AFTERWARDS. THIS VERY DAY. ALL PHYSICIANS I'M UNDER. INSIST I CONTINUE THE USE OF MY MACHINE. UNTIL THIS VERY DAY. I CONTINUE TO SUFFER WITH CHRONIC HEADACHES. BLURRY VISION. IRREGULAR HEARTBEATS. SUGAR LEVELS OFF THE CHART. VISION IS CHRONIC. GODLY PAIN HAS INCREASED TO A LEVEL THAT'S HARD TO DEAL WITH. I'M UNDER GOING MRI X-RAY AND CT SCANS. ON (B)(6) 2021 I'M HAVING AN MRI BRAIN TEST DID FOR MY HEADACHES. I'M TAKING MONTHLY INJECTIONS FOR HEADACHES. AND WILL BEGAN BOTOX INJECTIONS FOR MIGRAINES. MY ISSUES WITH MY MACHINE BEGAN IN 2019. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02323 | 2021-07-20 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
MW5102658 | 2021-07-20 | 2016-02-02 | I, n, j, u, r, y | Voluntary report | DREAM STATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | HAD SEVERE DIFFICULTY BREATHING, ENDED UP HOSPITALIZED WITH PNEUMONIA FROM (B)(6) 2016 I BELIEVE DUE TO THE CPAP AS WELL AS ON GOING INCREASED SINUS INFECTIONS AND GENERAL UPPER RESPIRATORY INFECTIONS SINCE STARTING TO USE THE CPAP. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102661 | 2021-07-20 | 2019-07-01 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | RESPIRONICS DREAMSTATION CPAP MACHINE THAT CAUSED HEADACHES AND CONTINUOUS SINUS INFECTIONS. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102664 | 2021-07-20 | 2021-04-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DNX500H11C |  |  | PHILIPS RESPIRONICS DREAMSTATION CPAP RECALL. I WAS EXPOSED TO DEGRADED SOUND ABATEMENT FOAM WHICH RESULTED IN SEVERE HEADACHES AND DIFFICULTY BREATHING. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102666 | 2021-07-20 | 2019-08-12 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS, INC. | KRX500T15 |  |  | I JUST NOTICED MY CPAP DID NOT WORK AS GOOD AS IT DID WHEN I FIRST GOT IT. I TOOK IT IN TO CHECK THE SETTINGS AND WAS TOLD IT WAS FINE. STILL HAVE PROBLEMS WITH BEING SLEEPY AND MY NOSE AND BREATHING ARE STARTING TO HURT. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102673 | 2021-07-20 | 2019-04-01 | I, n, j, u, r, y | Voluntary report | CPAP MACHINE (DREAMSTATION) | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | SINCE APPROX ABOVE DATE, REOCCURRING PAINFUL NASAL IRRITATION, RUNNY NOSE, UNDER-EYE RASHES. I HAVE BEEN USING PHILIPS DREAMSTATION CPAP FOR 5.5 YEARS. IT IS APPALLING TO NOT HAVE REASONABLE REPLACEMENT OPTIONS DROP DEAD TIRED WITHOUT CPAP THERAPY. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102674 | 2021-07-20 |  | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I HAVE BEEN HAVING HEADACHES FOR 2-3 WEEKS EVERY MORNING. I JUST FOUND OUT THAT MY PHILIPS DREAM STATION HAS A RECALL AND HEADACHES ARE ONE OF THE SYMPTOMS. I FOUND OUT WHEN A FRIEND POSTED ON (B)(6) AND THEN I RESEARCHED IT. I STOPPED USING THE DEVICE AS OF LAST NIGHT. I AM UPSET THAT THE PLACE WHERE I BOUGHT THE MACHINE AND MY DOCTOR DID NOT INFORM ME WHEN THIS FIRST CAME OUT. IT COULD HAVE SAVED ME FROM THE HEADACHES SOONER. I CAN ONLY HOPE THAT THERE ISN'T ANY OTHER HEALTH ISSUES FOR ME FROM THIS DEVICE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02235 | 2021-07-20 | 2021-06-17 | I, n, j, u, r, y | Manufacturer report | BIPAP AUTO BIFLEX, SYS ONE 60 SRS, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DS760S | DS760S |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP HEADACHES AND LESIONS ON BOTH LUNGS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP HEADACHES AND LESIONS ON BOTH LUNGS. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE EVENT IS CURRENTLY UNKNOWN. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION G1-MANUFACTURER CONTACT OFFICE INFORMATION WAS CORRECTY UPDATED. SECTION G4 WAS PREVIOUSLY INCORRECT AND HAS NOW BEEN CORRECTLY UPDATED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02335 | 2021-07-20 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP S/T C SERIES INTL 30, FLOW GEN | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS,INC | 1061T | 1061T |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING BLACK PARTICLES BLOWING OUT OF THE DEVICE AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND DEVICE WAS SCRAPPED. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02336 | 2021-07-20 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP S/T C SERIES INTL 30, FLOW GEN | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | 1061T | 1061T |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02231 | 2021-07-20 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING COUGHING,DIZZY,SHORTNESS OF BREATH,DIZZINESS,WHITE PARTICLES IN MASK,FATIGUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE MANUFACTURER EVALUATED THE DEVICE EXTERNALLY AND OBSERVED A CRACK IN THE TOP ENCLOSURE CROSSING THE AMBIENT LIGHT SENSOR PASS-THROUGH, CONTAMINATION AND OTHER PRESENCE OF LIGHT WEAR AND TEAR OBSERVED. THE MANUFACTURER EVALUATED THE DEVICE INTERNALLY AND OBSERVED CONTAMINATION IN THE BLOWER BOX AROUND THE BLOWER SEAL APPEARING CONSISTENT WITH THE KERATIN CONTAMINATION WITH DEGRADED SOUND ABATEMENT FOAM,EVIDENCE OF LIQUID INGRESS AND DUST/DIRT CONTAMINATION OBSERVED IN THE BLOWER. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THEY CONFIRMED THE PRESENCE OF CONTAMINATION IN THE BLOWER BOX APPEARING CONSISTENT WITH KERATIN CONTAMINATION AND THEY CONFIRMED THE EVIDENCE OF DEGRADED SOUND ABATEMENT FOAM WAS OBSERVED ALSO CONFIRMED THE CRACK IN THE TOP ENCLOSURE. SECTION H6 HEALTH IMPACT CODE WAS MISSED TO CAPTURE IN PREVIOUS MDR,NOW IT IS CORRECTED AND UPDATED IN HTIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02232 | 2021-07-20 | 2021-06-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON DEC 16, 2021, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES PARTICLES IN THE CHAMBER/TUBING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGATION OF STOMACH IRRITATION. STATING THAT POSSIBLY COULD BE FROM INGESTING THE PARTICULATES IN THE TUBING AND CHAMBER AND DEVICE HAS STRANGE ODOR. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTION E1 AND H6 HEALTH EFFECTS: CLINICAL CODES HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02254 | 2021-07-20 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND RESULTED IN THE PATIENT HAVING STAGE FOUR LUNG CANCER. THERE IS CURRENTLY NO FURTHER INFORMATION REPORTED ON THE MEDICAL INTERVENTION THAT THE PATIENT MAY HAVE RECEIVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT ALLEGED DIAGNOSED WITH STAGE FOUR LUNG CANCER. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURER CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REPORTED BY THE PATIENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02281 | 2021-07-20 | 2020-12-05 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/HT/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX400T11C | DSX400T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP AN ABSCESS ON THEIR LEFT LUNG, AND BACTERIAL PNEUMONIA. THE REPORTED INFORMATION STATES THE PATIENT WAS HOSPITALIZED FOR FIVE DAYS IN RESPONSE TO THE REPORTED EVENTS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING THAT A PATIENT DEVELOPED AN ABSCESS ON THEIR LEFT LUNG, AND BACTERIAL PNEUMONIA RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE REPORTED INFORMATION STATES THE PATIENT WAS HOSPITALIZED FOR FIVE DAYS IN RESPONSE TO THE REPORTED EVENTS. THE REPORTED EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. HENCE, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO THE ALLEGED SERIOUS INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIAGATION LAB FOR FURTHER EVALUATION. THE EXTERNAL AND INTERNAL ASPECT OF THE DEVICE WAS EVALUATED AND EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS OBSERVED IN THIS DEVICE BY THE MANUFACTURER. THE MANUAFCTURER OBSERVED MOIST SPOTS ON THE SOUND ABATEMENT FOAM AT THE AIR OUTPUT SIDE OF THE BLOWER BOX. BLOWER MOTOR MAKES HIGH PITCH WHINY SOUND WHILE PROVIDING THERAPY. UNKNOWN WHITE AND BLACK CONTAMINATION (DUST LIKE), INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM, AND SOME POTENTIAL VERY SMALL HAIRS AND OR FIBERS AT THE AIR INPUT OF THE BLOWER BOX WAS FOUND. UNKNOWN WHITE POWDERY SUBSTANCE WAS OBSERVED AT THE AIR INPUT OF THE BLOWER BOX AND THROUGHOUT THE BLOWER BOX AND ON THE TOP OF THE BLOWER MOTOR, THE BLOWER MOTOR SEAL AND INSIDE OF THE BLOWER MOTOR SEAL AND THE BLOWER MOTOR. SLIGHT BLACK CONTAMINATION, CONSISTENT WITH KERATIN WAS OBSERVED AT THE AIR OUTLET OF THE BLOWER BOX WAS BY THE MANUFACTURER. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER AND NINE ERRORS WERE OBSERVED. THE MANUFACTURER CONCLUDES THAT THEY CONFIRMED THE PRESENCE OF SOUND ABATEMENT FOAM DEGRADATION IN THE DEVICE. SECTION D8, D9, H2, H3 AND H6 WAS UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102660 | 2021-07-20 | 2021-05-24 | I, n, j, u, r, y | Voluntary report | PHILLIPS RESPIRONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | YES |  |  | I WAS USING THE DREAM STATION AND CONSTANTLY FELT LIKE SOMETHING WAS HITTING MY FACE WHEN I WOULD BE WEARING THE MASK. I THEN BEGAN TO DEVELOP VERY IRRITATED AND PAINFUL SORES JUST INSIDE BOTH NOSTRILS OF MY NOSE. I HAD SINUS SURGERY IN (B)(6) OF 2020 AND JUST CONTRIBUTED THE SORES TO THAT SURGERY BUT DIDN'T UNDERSTAND WHY I WOULD HAVE THEM NOW AFTER 8 MONTHS. WHEN I CALLED MY CPAP SUPPLIER TO ORDER SUPPLIES, I WAS TOLD ABOUT THE RECALL AND TO STOP USING MY CPAP. WHY NO ONE BOTHERED TO CALL TO TELL ME OF THE RECALL PRIOR TO THIS IS BEYOND ME. I THEN TOLD THE LADY I WAS SPEAKING WITH ABOUT THE PROBLEM I WAS HAVING WITH MY NOSE AND THAT UPON HEARING ABOUT THE RECALL I BEGAN TO THINK THAT MAYBE IT COULD BE CAUSED FROM THE CPAP AND SHE THEN STATED THAT SHE DID NOT KNOW IF IT WAS RELATED, SHE HAD VERY LITTLE INFO. I SHOULD ALSO MENTION THAT NOT ONCE HAVE I RECEIVED ANY CALLS FROM MY DOCTOR THAT ORDERED ME TO START USING THE CPAP. ONCE I STOPPED USING MY CPAP MY NOSE IMMEDIATELY STARTED HEALING UP AND AFTER APPROXIMATELY 2 WEEKS I NO LONGER HAVE THE IRRITATED AND PAINFUL SORES. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102649 | 2021-07-20 | 2021-07-04 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAM STATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | HEADACHE, URGENT CARE VISIT. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102711 | 2021-07-21 | 2021-01-01 | I, n, j, u, r, y | Voluntary report | CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS, INC. | DSX500T11 |  |  | FOR AT LEAST THE PAST 6 OR 7 MONTHS, I HAVE BEEN HAVING HEADACHES OFF AND ON, WITHOUT ANY REASON (LIKE POLLEN OR ALLERGIES). MY BIGGEST CONCERN IS MY LOWER BACK ON THE RIGHT SIDE. I KEEP HAVING A NAGGING PAIN TO SHARP PAIN JUST ABOVE MY RIGHT KIDNEY. INITIALLY I THOUGHT IT WAS HOW I WAS SITTING IN MY OFFICE CHAIR BUT I CHANGED POSITIONS MANY TIMES. MY RIGHT SIDE LOWER BACK CONTINUALLY HURTS SITTING OR STANDING. I ALSO COUGH MORE AND HAVE A STUFFY NOSE MORE THAN USUAL. SINCE I RECEIVED THE LETTER ABOUT MY PHILIPS RESPIRONICS CPAP MACHINE, IT NOW SEEMS TO MAKE SENSE WHY MY BACK HURTS AND I GET HEADACHES NOW. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02398 | 2021-07-21 | 2021-04-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | REMSTAR AUTO A-FLEX, SYS ONE, 60 SRS,GER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | DE562S | DE562S |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ACUTE KIDNEY FAILURE. IT IS UNKNOWN WHAT MEDICAL INTERVENTION WAS RECEIVED BY THE PATIENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER REPORTED A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM ALLEGEDLY BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ACUTE KIDNEY FAILURE. IT IS UNKNOWN WHAT MEDICAL INTERVENTION WAS RECEIVED BY THE PATIENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. THE DEVICE WAS VISUALLY INSPECTED. THE EXTERNAL INVESTIGATION FOUND NO EVIDENCE OF DAMAGE OR CONTAMINATION. THE MANUFACTURER VERIFED THAT POWER CORDS AND POWER SUPPLIES WERE FUNCTIONAL BY POWERING ON THE BASE UNIT AND INITIATING THERAPY. THE DEVICE'S ERROR LOG WAS REVIEWED AND THE MANUFACTURER CONFIRMED NO ERRORS WERE LOGGED. AN INTERNAL INVESTIGATION WAS PERFORMED ON THE DEVICE. THE MANUFACTURER CONFIRMED THE DEVICE HAD NO EVIDENCE OF FOAM DEGRADATION. THE MANUFACTURER FOUND EVIDENCE OF MODERATE DUST AND DIRT ON THE SOUND ABATEMENT FOAM, IN THE BLOWER BOX, ON THE BLOWER AND ON THE BLOWER IMPELLER. THE MANUFACTURER FOUND DISCOLORATION TO THE BLOWER WIRE SEAL AND THE AIR INTAKE SEAL CONSISTENT WITH TOBACCO CONTAMINATION. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02353 | 2021-07-21 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THIRD PARTY SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG AND THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02352 | 2021-07-21 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02402 | 2021-07-21 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED A TEST STEP DURING TESTING. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02406 | 2021-07-21 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S TOUCHSCREEN WAS UNRESPONSIVE. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR'S TOUCHSCREEN WAS UNRESPONSIVE. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02404 | 2021-07-21 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION RELATED TO THE DEVICE'S OXYGEN BLENDING MODULE BOARD OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A "SERVICE REQUIRED" ALARM CONDITION RELATED TO THE OXYGEN BLENDING MODULE OCCURRED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02405 | 2021-07-21 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S FLOW DELIVERY WAS INACCURATE. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S PROXIMAL PRESSURE PORT OF THE EXHALATION PORTING BLOCK WAS FOUND TO BE CLOGGED BY A FOREIGN SUBSTANCE, AND WAS CLEANED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02403 | 2021-07-21 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION RELATED TO THE DEVICE'S OXYGEN BLENDING MODULE BOARD OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A "SERVICE REQUIRED" ALARM CONDITION RELATED TO THE OXYGEN BLENDING MODULE OCCURRED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
MW5102704 | 2021-07-21 | 2021-03-26 | I, n, j, u, r, y | Voluntary report | DREAMSTNAUTOCPAPHUMHTWIFIDOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS | DSX500T11W | DSX500T11W |  | I WAS NOTIFIED JULY 14, 2021 TO STOP USING MY PHILLIPS RESPIRONIC DREAMSTNAUTOCPAPHUMHTWIFIDOM MACHINE FOR SLEEP APNEA WHICH I STARTED USING FOR THE FIRST TIME IN THE MIDDLE OF (B)(6) 2021. (DREAMSTATION TYPE) THIS IS A 4 MILLION PERSON RECALL. THERE ARE NO REPLACEMENT PARTS, OPTIONS OR EXCHANGES DUE TO THE FOAM THAT LINES THE MACHINE WHICH HAS BEEN DEGRADING AND RELEASING CARCINOGENS RESULTING IN SUCH SYMPTOMS AS COUGH, CHEST PRESSURE, NAUSEA, ETC. I WAS ADMITTED TO THE HOSPITAL (B)(6) 2021 FOR ATYPICAL CHEST PRESSURE AND A COUGH THAT WAS NOT PRODUCTIVE. A HEART ATTACK WAS RULED OUT, AND A POSSIBLE DIAGNOSIS, ALTHOUGH NOT PROVEN DIAGNOSTICALLY, WAS OF GASTROINTESTINAL ORIGIN. THE INFORMATION REGARDING THIS RECALL BY PHILIPS RESPIRONICS REQUESTING IMMEDIATE STOPPAGE OF USE, DID NOT BECOME AVAILABLE TILL RECENTLY (APRIL 2021), BUT PHILIPS WAS NOTIFYING PEOPLE AS OF JUNE 14, 2021. I DO NOT KNOW FOR SURE IF MY HOSPITALIZATION WAS CONNECTED, BUT I THINK IT IS SURELY WORTH PROVIDING THE INFORMATION AND POSSIBLY PREVENTING OTHERS FROM HARM. IN THE MEANTIME, I HAVE NOT BEEN OFFERED ANY ALTERNATIVES BUT TO DISCONTINUE USE. COMING TO THE DECISION TO ACCEPT USING C-PAP OR AUTO PAP, AS MINE IS CALLED, WAS DIFFICULT ENOUGH. NOW I MUST STOP USING IT WITH NO RECOURSE OR REPLACEMENT. SINCERELY, (B)(6). I HAD STRESS TESTS, CARDIAC ENZYMES, EKG, CONTINUOUS HEART MONITORING, BLOOD WORK ETC WHICH WAS ALL NEGATIVE FOR HEART ATTACK. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02407 | 2021-07-21 | 2021-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLANK. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE WAS EVALUATED AT THE MANUFACTURER'S SERVICE CENTER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02409 | 2021-07-21 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02300 | 2021-07-21 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02408 | 2021-07-21 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02295 | 2021-07-21 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S INTERNAL BATTERY WAS NOT CHARGING. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02303 | 2021-07-21 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A FAILURE OF THE DEVICE TO POWER ON WAS OBSERVED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02299 | 2021-07-21 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02294 | 2021-07-21 | 2021-06-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR HIGH TIDAL VOLUMES. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02296 | 2021-07-21 | 2021-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE TO POWER ON WAS OBSERVED. THE DEVICE'S FRONT PANEL/KEYPAD LED BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02293 | 2021-07-21 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WOULD NOT CALIBRATE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE INTERNAL TUBING WAS OBSERVED TO BE DISCONNECTED. THE DEVICE'S INTERNAL TUBING WAS RECONNECTED TO ADDRESS THE ISSUE. A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S DISPLAY BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02302 | 2021-07-21 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02351 | 2021-07-21 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02369 | 2021-07-21 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS INITIALLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT ALLEGED DEVICE WAS DRY WHEN AWOKEN, PATIENT FEELS NAUSEOUS, TIGHTNESS IN CHEST, SORE THROAT, SINUS INFECTION LIKE SYMPTOMS, RUNNY NOSES, CONGESTED, DEVICE DOES NOT WORK ALL THE TIME, NOSE HAS BEEN BURNING, AND HEADACHES.THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO CONTAMINATION. THERE WERE ALSO SIGNS OF WATER INGRESS IN THE BLOWER AND ON THE BLOWER BOX DUST/DIRT CONTAMINATION INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM WAS OBSERVED THROUGHOUT DEVICE ENCLOSURE, AND AIRPATH, SUGGESTING A SOURCE EXTERNAL TO THE DEVICE.THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERORRS. THE MANUFACTURER CONCLUDES THE DEVICE HAS CONTAMINATION CONSISTENT WITH WATER INGRESS. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. IN THIS REPORT, SECTION D9, G3, H3, H6 HAS BEEN UPDATED OR CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02401 | 2021-07-21 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S DISPLAY WAS WHITE. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | IT WAS INITIALLY REPORTED, THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S DISPLAY WAS WHITE.THERE WAS NO HARM OR INJURY REPORTED. REPEATED ATTEMPTS FOR ADDITIONAL INFORMATION OR TO HAVE THE DEVICE RETURNED FOR EVALUATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02306 | 2021-07-21 | 2021-06-25 | I, n, j, u, r, y | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPORINOC INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE DEVICE'S SOUND ABATEMENT FOAM DEGREDATION CASUED RESPIRATORY INFECTIONS IN A PATIENT. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTION MEDICATIONS. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON JAN 23 , 2025 AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE DEVICE'S SOUND ABATEMENT FOAM DEGREDATION CASUED RESPIRATORY INFECTIONS IN A PATIENT. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTION MEDICATIONS. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WAS NO ERROR FOUND. THE MANUFACTURER CONCLUDES THAT COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS VISIBLE FOAM DEGRADATION AND UNIT IS SCRAPPED. IN THIS REPORT, SECTIONS D9, G3,H2, H3, AND H6 HAVE BEEN UPDATED OR CORRECTED. | Y | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02345 | 2021-07-21 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, GB | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | GBX500H15 | GBX500H15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02199 | 2021-07-21 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID/HT, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AUX400T15 | AUX400T15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | SECTION B5 WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT SHOULD BE REPORTED AS FOLLOWS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED MIGRAINES AND SHORTNESS OF BREATH. THERE WAS NO PATIENT HARM OR INJURY. NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. IN SECTION D4 ADDITIONAL UNIQUE DEVICE IDENTIFIER (UDIS) NUMBER WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT. IT HAS BEEN CORRECTED OR UPDATED IN THIS REPORT. SECTION H6 HEALTH EFFECT - CLINICAL CODE WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT. IT HAS BEEN CORRECTED OR UPDATED IN THIS REPORT. UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02275 | 2021-07-21 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11 | DSX700T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. AFTER THE FINAL ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, CUSTOMER DECLINED TO RESPOND TO THE GFE RELATED QUESTIONS AND TERMINATED THE CALL. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102684 | 2021-07-21 | 2020-04-22 | I, n, j, u, r, y | Voluntary report | PHILLIPS DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS RESPIRONICS, INC. | DSX500H11C |  |  | I USE A PHILLIPS DREAMSTATION CPAP THAT HAS BEEN RECALLED AND INTERMITTENTLY NOTED BLACK PARTICLES IN HUMIDIFIER AND TUBING AND WAS DIAGNOSED WITH STAGE 4 METASTATIC BREAST CANCER ALONG WITH LYMPH NODES AND LEFT LUNG. USED CPAP FOR YEARS. WITH THE BREAST CANCER I HAVE HAD MULTIPLE ECHOCARDIOGRAMS, FULL BODY PET SCANS, CT SCANS, BIOPSY, CHEMO TREATMENTS, MULTIPLE IMMUNOTHERAPY TREATMENTS, AND MULTIPLE OTHER TESTS AS A RESULT. I HAVE ALSO BEEN DIAGNOSED WITH LYMPHEDEMA IN MY LEFT ARM AND HAND WITH LIFELONG TREATMENTS AND LIFELONG WEAR OF COMPRESSION GARMENTS. I CAN PROVIDE A MORE DETAILED LIST OF TESTS IF NEEDED. MULTIPLE TESTS AND CAN PROVIDE A COMPLETE LIST OF TESTS AND TREATMENTS LATER IF NEEDED. TOO MANY TO LIST HERE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-07
MW5102690 | 2021-07-21 | 2021-02-01 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 |  |  | NOTICED STRANGE SMELL IN CPAP, THOUGHT IT WAS JUST BECAUSE IT WAS A NEW DEVICE, HAD HEADACHE AFTER EACH USE. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102696 | 2021-07-21 | 2021-07-15 | I, n, j, u, r, y | Voluntary report | PHILLIPS DREAMSTATION CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS RESPIRONICS / RESPIRONICS INC. | J289742757A30 |  |  | COUGHING, DIZZINESS; PHILLIPS CAPP MACHINE CAUSED NAUSEA, COUGHING FAINTING ETC. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102700 | 2021-07-21 | 2020-12-30 | I, n, j, u, r, y | Voluntary report | PHILLIPS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500H11C |  |  | DEVELOPED LUNG CANCER AFTER USAGE OF PHILLIPS DREAMSTATION MACHINE. WENT IN FOR SURGERY ON (B)(6) 2020. MACHINE WAS IN USE FOR APPROXIMATELY 4.5 YEARS. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02367 | 2021-07-21 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. A CORRECTION TO B5 WAS MADE AND SHOULD BE THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED BLACK SPECS WERE FOUND IN BOTH MACHINES.THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WAS 1 ERROR FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02366 | 2021-07-21 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT IS UNABLE TO PROVIDE DETAILS ON THE SERIAL NUMBER OF THE DEVICE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02301 | 2021-07-21 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02365 | 2021-07-21 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02333 | 2021-07-21 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11W | DSX500T11W |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING WEEZING, NASAL/THROAT DRYNESS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. IN INITIAL REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02332 | 2021-07-21 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102718 | 2021-07-21 | 2021-07-14 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. |  |  |  | NEED CPAP MACHINE TO LIVE; USED A LIFE THREATENING CANCER CAUSING CPAP MACHINE. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | I |  | 2021-07
MW5102712 | 2021-07-21 | 2021-04-19 | I, n, j, u, r, y | Voluntary report | DREAM STATION AUTO CPAP WITH HUMIDIFIER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500H11C |  |  | I AM A CPAP USER WHO WAS USING PHILIPS DREAM STATION 2 CPAP MACHINE. MY SLEEP APNEA IS RATED AS SEVERE, SO I USE THE MACHINE EVERY NIGHT. AROUND THE SAME TIME (APPROXIMATELY (B)(6) 2021), I STARTED WAKING UP WITH VERY PAINFUL HEADACHES AND NOTICED BLACK PARTICLES IN MY HUMIDIFICATION CHAMBER (WHICH I ONLY USED BOTTLED DISTILLED WATER FOR). AS THIS WAS BEFORE THE PHILIPS RECALL NOTICE, I SIMPLY KEPT USING THE MACHINE. WHEN THE RECALL NOTICE CAME OUT, ON THE RECOMMENDATION OF MY DOCTOR - I HAVE STOPPED USING THE CPAP MACHINE COMPLETELY. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102709 | 2021-07-21 | 2020-03-07 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | RESPIRONICS DREAM STATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I BEGAN HAVING AN UNEXPLAINED COUGH. I THOUGHT IT MAY HAVE BEEN RELATED TO MASK USE DURING COVID BECAUSE I STILL HAVE IT, BUT I WEAR A MASK WHILE WORKING. I JUST LEARNT TODAY THAT THE COUGH MAY BE DUE TO A RECALL, SO I REGISTERED MY CPAP WITH PHILIPS AS IT IS A UNIT THAT IS AFFECTED AND WAS INSTRUCTED TO COMPLETE THIS FORM IF I HAD ANY OF THE REPORTED SIDE EFFECTS. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02339 | 2021-07-21 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTIALATOR'S LCD SCREEN WAS DAMAGED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. THE DEVICE HAD EVIDENCE OF PHYSICAL DAMAGE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02340 | 2021-07-21 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BLOWER MOTOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02341 | 2021-07-21 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WOULD NOT POWER ON. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S POWER SUPPLY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02400 | 2021-07-21 | 2021-06-21 | I, n, j, u, r, y | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT'S CARBON DIOXIDE LEVELS WERE INCREASED WHILE USING A VENTILATOR. THE PATIENT WAS HOSPITALIZED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A PATIENT'S CARBON DIOXIDE LEVELS WERE INCREASED WHILE USING A VENTILATOR. THE PATIENT WAS HOSPITALZIED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | N | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02334 | 2021-07-21 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11 | DSX400H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. DURING THE EVALUATION VISIBLE FOAM PARTICLES WERE OBSERVED. PRESERVED CONTAMINATION FINDINGS WERE FOUND. THERE WERE 4 ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS VISIBLE FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02298 | 2021-07-21 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040005 | 1040005 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02350 | 2021-07-21 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02368 | 2021-07-21 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11 | DSX700H11 |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED FOAM DEGRADATION. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE 18 ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION. THE UNIT WAS SCRAPPED AND THERE WAS VISIBLE FOAM DEGRADATION FOUND. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102680 | 2021-07-21 |  | D, e, a, t, h | Voluntary report | REMSTAR SE AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | THIS IS ABOUT MY DAD. HE PASSED AWAY LAST (B)(6) FROM THYROID CANCER. HE NEVER HAD ANY HISTORY OF THYROID PROBLEMS OR ANYTHING RELATED. TWO YEARS BEFORE HE DIED HE BEGAN WITH A COUGH, HEADACHE, DIZZINESS. HE FINALLY WENT TO URGENT CARE REGARDING HIS COUGH. THE RESULT WAS THAT THEY FOUND HE HAD THYROID CANCER AND HAD TO HAVE SURGERY. EVENTUALLY AFTER INITIALLY RECOVERING THE CANCER RETURNED AND RESULTED IN HIS DEATH. WHEN I SAW THE RECALL OF THE PHILIPS CPAP I COULDN'T HELP BUT WONDER IF IT IS RELATED. MY DAD USED ONE OF THE RECALLED DEVICES UP UNTIL HIS CANCER SURGERY AT WHICH TIME THE DOCTOR SAID HE COULDN'T USE IT ANYMORE HAVING HAD NO THROAT, LUNG OR COUGHING ISSUES BEFORE, AND BEING A SWIMMER AND STILL IN ATHLETIC CONDITION (JUMPED ROPE 80 TIMES A DAY IN HIS 80'S), THE WHOLE THING WAS VERY SUDDEN AND SURPRISING. YOU SAID THAT THE PHILIPS CPAP RECALL DID NOT LEAD TO ANY DEATHS BUT I'M NOT SURE THAT IS TRUE. IS THERE A WAY TO INVESTIGATE THROUGH MY DAD'S MEDICAL RECORDS (WHICH I HAVE MOST OF) IF THIS COULD BE RELATED? FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102686 | 2021-07-21 | 2020-11-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS INC |  |  |  | HAVE HAD A CHRONIC COUGH. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102693 | 2021-07-21 | 2021-07-15 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. |  |  |  | I LEARNED TODAY THAT MY PHILLIPS DREAMSTATION CPAP MACHINE HAS BEEN RECALLED. I HEARD ABOUT IT ON TV FROM A LAWYER ADVERTISEMENT. NO ONE FROM PHILLIPS HAD CONTACTED ME NOR ANYONE FROM (B)(6) HOME HEALTH WHICH IS MY PROVIDER. I HAVE BEEN HAVING HEADACHES, DIZZINESS, AND HEAVINESS IN MY CHEST AND THESE ARE SYMPTOMS THAT ARE LISTED ON THE RECALL. I HAVE BEEN BREATHING TOXIC AND CARCINOGENIC COMPOUNDS FROM THIS MACHINE FOR MONTHS. I CONTACTED MY SLEEP APNEA DR AND WAS TOLD TO STOP USING THIS MACHINE IMMEDIATELY. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102703 | 2021-07-21 | 2021-07-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSXHCP |  |  | SYMPTOMS; I HAVE BEEN USING A PHILLIPS RESPIRONICS CPAP MACHINE THAT HAS BEEN RECALLED. I DIDN'T KNOW ABOUT IT UNTIL (B)(6) 2021. I HAVE BEEN HAVING HEADACHES, WHEEZING AND SINUS ISSUES FOR A FEW WEEKS BEFORE I KNEW ABOUT THE RECALL. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102698 | 2021-07-21 |  | I, n, j, u, r, y | Voluntary report | PHILLIPS DREAM STATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DREAMSTNAUTOCPAPHUMCELL DOM |  |  | RECURRENT SINUS INFECTIONS; RECURRENT SINUS INFECTIONS. USE PHILLIPS DREAM STATION CPAP THAT HAS BEEN RECALLED. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-07
2518422-2021-02447 | 2021-07-22 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WOULD NOT POWER ON. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02305 | 2021-07-22 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS PROVIDING WEAK PRESSURE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02446 | 2021-07-22 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WSA NOT MEASURING PRESSURE. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02445 | 2021-07-22 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S TOUCHSCREEN WAS NOT RESPONSIVE. THERE WAS NO HARM OR INJURY REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02440 | 2021-07-22 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02297 | 2021-07-22 | 2021-07-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE TO POWER ON WAS OBSERVED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
MW5102732 | 2021-07-22 | 2019-02-15 | I, n, j, u, r, y | Voluntary report | PHILLIPS RESPIRONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C |  |  | I HAVE BEEN USING A PHILLIPS RESPIRONICS CPAP MACHINE SINCE (B)(6) 2019 FOR SEVERE SLEEP APNEA. DURING THE LAST TWO YEARS I HAVE SUFFERED FROM IRRITATION AND INFLAMMATION OF LUNGS, CHRONIC SEVERE COUGH, SCRATCHY BURNING SENSATIONS IN MY THROAT, CHEST TIGHTNESS AND SHORTNESS OF BREATH, FACIAL SKIN RASH AND SKIN SENSITIVITY, IRREGULAR HEARTBEAT, FATIGUE, MUSCLE WEAKNESS AND SWALLOWING DIFFICULTY. THESE SYMPTOMS HAVE INCREASED OVER TIME AND IN SEVERITY AND NOW ARE SEVERELY AFFECTING MY DAILY QUALITY OF LIFE. I AM NEVER WITHOUT PAIN, COUGH AND DISCOMFORT. I HAVE SOUGHT ANSWERS, MEDICAL ADVICE AND HAD EXTENSIVE TESTING THROUGHOUT THE LAST TWO YEARS AND FOUND NO ANSWERS UNTIL NOW WHEN I LEARNED OF CPAP RECALL. I SUSPECT IT WAS CONSTANT CHEMICAL EXPOSURE TO POLYURETHANE AND CHEMICALS SUCH AS THE ISOCYANATES INSIDE THE MACHINE. I STOPPED USING THE CPAP (B)(6) 2021 WHEN I LEARNED OF THE RECALL. MY SYMPTOMS HAVE NOT IMPROVED. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02439 | 2021-07-22 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S POWER MANAGEMENT BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02438 | 2021-07-22 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICE'S LCD DISPLAY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02304 | 2021-07-22 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "HIGH TEMPERATURE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S EXHAUST FAN ASSEMBLY, BATTERY FAN, AND STIRRING FAN WERE REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY MARKED H10 INCORRECTLY. THIS MDR WAS NOT SUBMITTED AS PART OF A BATCH SUBMISSION OF COMPLAINTS THAT WERE REASSESSED AS REPORTABLE FOAM DEGRADATION COMPLAINTS DISCOVERED AS PART OF A RETROSPECTIVE REMEDIATION REVIEW. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02455 | 2021-07-22 | 2021-07-15 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | JPX500S16 | JPX500S16 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND A PATIENT HAD A CT SCAN DONE THAT REVEALED SPOTS ON THEIR LUNGS. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE REPORTED EVENT IS CURRENTLY UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGATION ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND A PATIENT HAD A CT SCAN DONE THAT REVEALED SPOTS ON THEIR LUNGS. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE REPORTED EVENT IS CURRENTLY UNKNOWN. THE ABOVE REPORTED EVENT IS ASSESSED AS POSSIBLY RELATED TO THE DEVICE IN THIS CASE. HENCE, THE DEVICE MAY HAVE CAUSE OR CONTRIBUTE TO THE ALLEGED SERIOUS INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02448 | 2021-07-22 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S DISPLAY BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02451 | 2021-07-22 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX400H12C | CAX400H12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT DEVELOPED COUGHING, WHEEZING, SHORTNESS OF BREATH AND CONGESTION. THE PATIENT REQUIRED MEDICAL INTERVENTION IN THE FORM OF ANTIBIOTICS AND STEROIDS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO EXPERIENCE COUGHING, WHEEZING, SHORTNESS OF BREATH AND CONGESTION. THERE WAS MEDICAL INTERVENTION REQUIRED BY THE PATIENT IN THE FORM OF ANTIBIOTICS AND STEROIDS. THE REPORTED EVENT OF COUGHING, WHEEZING, SHORTNESS OF BREATH AND CONGESTION AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION AND AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION G3 WAS CAPTURED INCORRECTLY IN THE INITIAL REPORT. SECTION D4 WAS CORRECTED AND SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02372 | 2021-07-22 | 2021-05-10 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, GB | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | GBX500S15 | GBX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A THROAT INFECTION. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A DOCTOR'S ASSESSMENT AND PRESCRIPTION FOR ANTIBIOTICS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A THROAT INFECTION. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A DOCTOR'S ASSESSMENT AND PRESCRIPTION FOR ANTIBIOTICS. THE PATIENT ALSO ALLEGED TRACE OF BLACK MOLD ON TUBE INLET. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL PART OF THE DEVICE AND FOUND NO EVIDENCE OF THERMAL, PHYSICAL DAMAGE, ODORS, FOAM MATERIAL OR MOLD. THE MANUFACTURER DID OBSERVE DIRT/DUST CONTAMINATION ON THE RETURNED FILTERS, THE FRESH AIR INLET PORT, AND IN THE HUMIDIFIERS AIR FLOW OUTLET PORT. THE MANUFACTURER ALSO FOUND THE HUMIDIFIERS FLIP LID RELEASE IS DAMAGED. THE INSPECTION OF THE 2 TUBING'S AND THE PATIENTS MASK FOUND NO EVIDENCE OF MOLD OR DEBRIS. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND FOUND NO EVIDENCE OF A FAILURE OF THE SOUND ABATEMENT FOAM OR ANY MOLD IN THE BLOWER MOTOR/BOX OR IN THE AIR PATH OF THE DEVICES. THE MANUFACTURER DID FIND DIRT/DUST IN THE BLOWER MOTOR WHICH WOULD BE IN THE AIR PATH OF THE DEVICES. THE MANUFACTURER DID NOT OBSERVE ANY EVIDENCE OF DEBRIS OR ANY ODORS BEING EMITTED DURING THE TIME THE DEVICES WERE OPERATED IN THE LAB. THERE WAS NO EVIDENCE OF ANY BLACK DEBRIS ON THE BACTERIA FILTER. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE DEVICE WAS POWERED ON AND WAS FOUND TO OPERATE PROPERLY. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND 0 ERROR CODE. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION BUT FOUND DIRT/DUST CONTAMINATION IN THE AIR PATH OF THE DEVICE, WHICH IS LIKELY FROM EXTERNAL SOURCE. SECTION D4, SECTION D9, SECTION H3, CODES IN SECTION H6 WERE UPDATED/ CORRECTED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02325 | 2021-07-22 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION EXPERT W/HUM/CELL, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX501H12C | CAX501H12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02437 | 2021-07-22 | 2021-06-18 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, BL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | BLX500H15 | BLX500H15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG CANCER. IT IS UNKNOWN WHAT MEDICAL INTERVENTION WAS RECEIVED BY THE PATIENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED THAT A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM ALLEGEDLY BECAME DEGRADED. THE PATIENT HAS INDICATED MEDICAL INTERVENTION WAS REQUIRED FOR DEVELOPING LUNG CANCER. THE DEVICE WAS RETURNED TO THE MANUFACTURER QUALITY PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. DURING THE INITIAL VISUAL INSPECTION OF THE DEVICE, MINOR CONTAMINANTS WERE FOUND TO BE PRESENT AROUND INLET FILTER AREAS, OUTLET PORT AND THE HUMIDIFIER LATCH AREA. THE AIR INLET FILTER CONTAINED A MODERATE AMOUNT OF OFF-WHITE FIBROUS MATERIAL. THE DEVICE WAS DISASSEMBLED FOR INTERNAL VISUAL INSPECTION. THE MANUFACTURER FOUND EVIDENCE OF DUST CONTAMINATION INSIDE THE DEVICE. THE FILTER HOUSING AND INLET AREA CONTAINED MINOR DARK DUST OR DIRT CONTAMINANT, AND A BEIGE DUST OR DIRT CONTAINMENT. THE POLLEN AND ULTRAFINE FILTERS WERE BOTH ALMOST BLACK IN COLOR. DUST OR DIRT CONTAMINATION WERE LOCATED ON THE INSIDE OF THE BLOWER BOX, PRESSURE SENSOR PORT. AN INSECT WAS FOUND BELOW THE BOTTOM LEFT OF THE UI PANEL INSIDE THE BOTTOM ENCLOSURE CHANNEL. MANUFACTURER ALSO FOUND DARK DUST OR DIRT CONTAMINANT, AND DARK FIBROUS MATERIAL THROUGHOUT THE AIRPATH. THE MANUFACTURER CONFIRMS THE DUST AND DIRT CONTAMINATION WERE FROM AN EXTERNAL SOURCE. THE SOUND ABATEMENT FOAM WAS FOUND TO HAVE NO EVIDENCE OF DEGRADATION AND RETURNED TO ITS ORIGINAL SHAPE AFTER BEING COMPRESSED. SOUND ABATEMENT FOAM DID HAVE MORE EVIDENCE OF DUST CONTAMINATION THAN THE REST OF THE DEVICE. THE MANUFACTURER CONCLUDES THERE IS NO EVIDENCE OF FOAM DEGRADATION WITHIN THE DEVICE. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02421 | 2021-07-22 | 2021-07-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/HT/CELL, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AUX400T15C | AUX400T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP TUMORS, THYROID CANCER, AND LEUKEMIA. THE REPORTED INFORMATION STATES THE PATIENT HAD THYROID SURGERY, HAS BEGUN CHEMOTHERAPY, AND WAS PRESCRIBED TEMOXAFIN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02376 | 2021-07-22 | 2020-01-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AUX500T15C | AUX500T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | SECTION B5 WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT SHOULD BE REPORTED AS FOLLOWS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED EYE IRRITATION, SINUS ISSUES, CHEST PRESSURE, AND FINDING IT HARD TO BREATH. THERE WAS NO PATIENT HARM OR INJURY. NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. IN SECTION D4 ADDITIONAL UNIQUE DEVICE IDENTIFIER (UDIS) NUMBER WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT. IT HAS BEEN CORRECTED OR UPDATED IN THIS REPORT. SECTION H6 HEALTH EFFECT - CLINICAL CODE WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT. IT HAS BEEN CORRECTED OR UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02364 | 2021-07-22 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | CAX500T12 | CAX500T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO LOSE HEARING IN THEIR RIGHT EAR. IT IS UNKNOWN WHAT MEDICAL INTERVENTION WAS RECEIVED BY THE PATIENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOVEMBER, 05 2024, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING CAUSED A PATIENT TO LOSE HEARING IN THEIR RIGHT EAR RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTION E1 WAS UPDATED IN THIS REPORT. (GETTING ADDITIONAL INFORMATION) SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM. SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02371 | 2021-07-22 | 2021-07-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, GER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DEX500S13 | DEX500S13 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED HEALTH ISSUES SUCH AS BREATHLESSNESS, COUGH, RED EYES, AND ASTHMA. THE DETAILS OF THE MEDICAL INTERVENTION THAT THE PATIENT REQUIRED ARE CURRENTLY UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED HEALTH ISSUES SUCH AS BREATHLESSNESS, COUGH, RED EYES, AND ASTHMA. THE DETAILS OF THE MEDICAL INTERVENTION THAT THE PATIENT REQUIRED ARE CURRENTLY UNKNOWN. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL PART OF THE DEVICE AND FOUND NO DISTINCT WEAR, TEAR, OR CONTAMINATION ON THE ENCLOSURE. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND FOUND NO DISTINCT WEAR, TEAR OR CONTAMINATION. NO EVIDENCE OF SOUND ABATEMENT FOAM BREAKDOWN WAS OBSERVED, AND THE FOAM APPEARS INTACT. THE MANUFACTURER VERIFIED THE UNITS DO POWER-UP, PROVIDE FLOW. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION D4 UPDATED/CORRECTED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02452 | 2021-07-22 | 2021-01-02 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION EXPERT W/HUM/HT/CELL, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX501T12C | CAX501T12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG ISSUES. THE REPORTED INFORMATION STATES THE PATIENT WAS PRESCRIBED CETIRIZINE, SYMBICORT, AND MOMETASONE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG ISSUES. THE REPORTED INFORMATION STATES THE PATIENT WAS PRESCRIBED CETIRIZINE, SYMBICORT, AND MOMETASONE. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR EXAMINATION. NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN WAS OBSERVED WITHIN THE BASE UNIT. HOWEVER, DUST AND DIRT CONTAMINATION DETERMINED TO BE INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM WAS FOUND THROUGHOUT THE DEVICE ENCLOSURE AND AIRPATH, INDICATING A LIKELY EXTERNAL SOURCE OF CONTAMINATION. THE DEVICE'S EVENT LOG WAS REVIEWED BY THE MANUFACTURER, AND NO ERRORS WERE FOUND. THE MANUFACTURER CONCLUDES THAT IT IS UNABLE TO ADDRESS THE SYMPTOMS DESCRIBED IN THE COMPLAINT. THE INVESTIGATION CONFIRMED THE PRESENCE OF CONTAMINATION WITHIN THE AIRPATH. HOWEVER, THE MANUFACTURER IS UNABLE TO EVALUATE OR ADDRESS ANY CONCERNS RELATED TO THE FILTER, AS IT WAS NOT RETURNED WITH THE DEVICE. IN THIS REPORT, SECTION D9, G3, H3, H6 HAS BEEN UPDATED OR CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02378 | 2021-07-22 | 2021-06-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP W/HUMID/HT | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200T11 | DSX200T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP LUNG ISSUES INCLUDING RESPIRATORY VIRAL INFECTIONS AND LUNG NODULES. THE PATIENT WAS REPORTEDLY HOSPITALIZED IN RESPONSE TO THE EVENTS AND REQUIRED STEROIDS AND AN INHALER FOR TREATMENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 13, 2024, AND SECTION H11 SHOULD BE REPORTED AS THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP ITCHY THROAT, LUNG ISSUES INCLUDING RESPIRATORY VIRAL INFECTIONS AND LUNG NODULES. THE PATIENT WAS REPORTEDLY HOSPITALIZED IN RESPONSE TO THE EVENTS AND REQUIRED STEROIDS AND AN INHALER FOR TREATMENT. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM (PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR). SECTION B2 WAS CHANGED TO HOSPITALIZATION (PREVIOUSLY IT WAS BLANK). SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02375 | 2021-07-22 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | RGBX500S15 | RGBX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02377 | 2021-07-22 | 2021-06-18 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500T11W | DSX500T11W |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE A LUNG SPOT. THERE IS CURRENTLY NO FURTHER INFORMATION REPORTED ON THE MEDICAL INTERVENTION THAT THE PATIENT MAY HAVE RECEIVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN SPOT ON LUNG BECAME BIGGER THAN THE LATEST UPDATE AND HAD IRRITATION TO NOSE AND NASAL PASSAGES. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THEREFORE, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HEALTH EFFECT - CLINICAL CODE HAS BEEN CORRECTED, AND TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02390 | 2021-07-23 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES AND LOUD SOUND DURING THERAPY. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION OF VISUALIZATION OF PARTICLES AND AND CONFIRMED THE CUSTOMER COMPLAINT OF NOISE AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRABBED DUE TO AGE". | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02441 | 2021-07-23 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | RIN2110X15B | RIN2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VARIATION IN THE VENTILATOR'S PRESSURE DELIVERY WAS OBSERVED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN EVALUATED BY THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VARIATION IN THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE DELIVERY WAS OBSERVED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE VENTILATOR'S SOFTWARE WAS RE-LOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02471 | 2021-07-23 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TO CHARGE ITS INTERNAL BATTERY. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR FAILED TO CHARGE IT'S INTERNAL BATTERY. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
MW5102751 | 2021-07-23 | 2021-04-19 | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS / RESPIRONICS INC. | DSX500T11C |  |  | SINCE MID (B)(6) 2021 I HAVE BEEN BATTLING RESPIRATORY ISSUES INCLUDING COUGH, NASAL CONGESTION, CHEST TIGHTNESS, SHORTNESS OF BREATH. I HAVE ALSO BECOME OVERLY SENSITIVE TO ALCOHOL AND ALSO VINEGAR VAPORS. I WAS DIAGNOSED WITH ASTHMA ON (B)(6) 2021. I HAVE SLEEP APNEA AND HAVE BEEN A LONG TIME USER OF CPAP THERAPY. I HAVE BEEN A REGULAR USER OF A PHILIPS RESPIRONICS DREAMSTATION CPAP THAT WAS RECENTLY PART OF A RECALL. I SUSPECT THIS MAY HAVE CAUSED THE MEDICAL ISSUES I HAVE BEEN RECENTLY DEALING WITH. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02388 | 2021-07-23 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF BLACK PARTICLES INSIDE THE LID AREA AND IT STILL PERSISTS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. THE MANUFACTURER PREVIOUSLY NOT CAPTURED, "VISUALIZATION OF BLACK PARTICLES INSIDE THE LID AREA" IN SECTION B5 AND IT HAS BEEN UPDATED IN THIS REPORT. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02387 | 2021-07-23 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. UNIT WAS SCRAPPED. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02477 | 2021-07-23 | 2021-07-20 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | RJPX500S16 | RJPX500S16 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP PNEUMONIA. THE PATIENT WAS REPORTEDLY HOSPITALIZED IN RESPONSE TO THE EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED THAT A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM ALLEGEDLY BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP PNEUMONIA. THE PATIENT WAS REPORTEDLY HOSPITALIZED IN RESPONSE TO THE EVENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER QUALITY PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. THE DEVICE WAS DISASSEMBLED FOR INTERNAL VISUAL INSPECTION. THE MANUFACTURER FOUND EVIDENCE OF DIRT/DUST CONTAMINATION INSIDE THE DEVICE BLOWER MOTOR WITH NO EVIDENCE OF SOUND ABATEMENT FOAM MATERIAL. THE FILTER HOUSING AND INLET AREA CONTAINED MINOR DARK DUST OR DIRT CONTAMINANT, AND A BEIGE DUST OR DIRT CONTAINMENT. THE POLLEN AND ULTRAFINE FILTERS WERE BOTH ALMOST BLACK IN COLOR. THE MANUFACTURER CONFIRMS THE DIRT/DUST CONTAMINATION WERE FROM AN EXTERNAL SOURCE. THE MANUFACTURER REVIEWED OF THE DEVICE'S EVENT LOG AND FOUND 13 LOGGED ERROR CODES. THESE CODES WOULD NOT BE RELATED TO THIS COMPLAINT. THE POWER WAS APPLIED AND THE DEVICE DOES POWER UP PROVIDES FLOW THE HUMIDIFIERS HEATER PLATE HEATS. THE DEVICE DID NOT GENERATE ANY CODES DURING THE TIME THE DEVICE WAS OPERATED IN THE LAB. THE MANUFACTURER OPERATED THE DEVICES FOR AN APPROXIMATELY 2.5 HR. TIME PERIOD WHILE CONNECTED TO AN ASL SERVO LUNG TO SIMULATE A PATIENT BREATHING ON THE DEVICE. THE MANUFACTURER OBSERVED NO DEBRIS OR CONTAMINATES ON THE FILTER AFTER THIS TIME. THE MANUFACTURER CONCLUDES THERE IS NO EVIDENCE OF FOAM DEGRADATION WITHIN THE DEVICE. THE DEVICE WAS CONTAMINATED WITH DUST AND DIRT CONSISTENT FROM AN EXTERNAL SOURCE. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02236 | 2021-07-23 | 2021-06-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LESIONS ON BOTH LUNGS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 12, 2024, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LESIONS ON BOTH LUNGS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02380 | 2021-07-23 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON APR 12, 2024 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING, SHORTNESS OF BREATH AND PATIENT ALSO ALLEGEDLY STATES PARTICLES/BLACK DEBRIS IN THE WATER CHAMBER AND INGESTED PARTICLES FROM THE DEVICE THAT ARE RELATED CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. DURING AN INVESTIGATION, THE MANUFACTURER OBSERVED THAT THERE IS DUST/DIRT CONTAMINATION ON TOP AND BOTTOM ENCLOSURES, UI PANEL, REAR PANEL, SD COVER FLIP DOOR, CONTROL DIAL, KEYPAD, BLOWER, BLOWER BOX, P4 POWER CONNECTOR, DC JACK COLOR INSERT, FLIP LID ASSEMBLY, FLIP LID SEAL, BOTTOM ASSEMBLY, HEATER PLATE, HEAT SHIELD, BOTTOM COVER, DRY BOX. THE DEVICE IS MISSING THE ACCESSORY MODULE FLIP DOOR, WATER TANK, AND DRY BOX SEAL. AN UNKNOWN CONTAMINATE WAS OBSERVED INSIDE THE TOP AND BOTTOM ENCLOSURES, INSIDE THE UI AND REAR PANELS. DISCOLORATION WAS OBSERVED ON THE FLIP LID SEAL. LIQUID INGRESS WAS OBSERVED ON THE P4 POWER CONNECTOR. A CONTAMINATE CONSISTENT WITH KERATIN WAS OBSERVED ON THE BLOWER BOX. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED IN THE BASE UNIT. THE MANUFACTURER WAS NOT ABLE TO CONFIRM THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM.THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER.THE MANUFACTURER FOUND NO ERROR CODES.THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THAT THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION.THE MANUFACTURER OBSERVED DUST/DIRT CONTAMINATION IN THE AIRPATH, LIKELY FROM A SOURCE EXTERNAL TO THE DEVICE AND UNABLE TO ADDRESS THE PATIENT'S SYMPTOMS SECTIONS D9, G3, H2, H3, H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02381 | 2021-07-23 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02382 | 2021-07-23 | 2021-07-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION HUMIDIFIER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | JPXH | JPXH |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY INFORMATION ALLEGING HAVING BLACK TRASH HAS GOTTEN INTO THE MAINBODY ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM.THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE MANUFACTURER OBSERVED THE EVIDENCE OF VISUAL INSPECTION REVEALED THAT BLACK POINT-LIKE DIRT, WHOSE SIZE WAS ABOUT TWO MILLIMETERS IN A DIAMETER, HAD CLUNG TO A GAP BETWEEN THE MAIN BODY AND CLEAR RESIN PART ON THE INNER PART OF THE DREAMSTATION WATER TANK. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. SECTION H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02389 | 2021-07-23 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO FILTER WAS ALWAYS BLACK. CORRECTED INFORMATION PROVIDED IN A1. (PATIENTS IDENTIFIER WAS NOT INCLUDED IN INITIAL MDR). CORRECTED INFORMATION PROVIDED IN B.5. (INFORMATION ALLEGING THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING FILTER WAS ALWAYS BLACK WAS NOT INCLUDED IN INITIAL MDR). | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102758 | 2021-07-23 | 2021-07-01 | I, n, j, u, r, y | Voluntary report | CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | PHILLIPS CPAP MACHINE HAS BEEN CAUSING COUGH ON AND OFF EVER SINCE OWNERSHIP. ALSO STUFFY NOSE A LITTLE WHILE AFTER BEING DONE WITH IT USUALLY WITHIN AN HOUR. COUGHING IS THE WORST THOUGH SOMETIMES IT'S MINIMAL OR NONEXISTENT SOME TIMES IT BAD. MENTIONED TO DOCTOR SINCE OWNED WE CHOCKED IT UP TO NOT CLEANING IT ENOUGH, EVEN THOUGH SOMETIMES I WAS WASHING THE HOSE, MASK AND TRAY EVERY OTHER DAY, OR THE FILTERS BEING DIRTY, OR ME NOT BEING USE TO THE PRESSURE OF THE MACHINE EVEN WHEN IT WAS MILD. NOW IT'S CRANKED TO 14 AND THE LAST FEW WEEK I SOMETIMES HAVE FELT SOMETHING LIGHT FLOATING IN MY FACE MASK, BUT WHEN I TAKE IT OF THERE IS NOTHING THERE. I DO NOT USE HEAT ON MY MACHINE, IT'S JUST STRAIGHT COLD AIR NO WATER. TO MAKE THIS WORSE THE (B)(6) HAS YET TO GET IN CONTACT WITH VETERANS IN MY AREA AND INFORM THEM ABOUT PROBLEMS THIS CAN CAUSE, BUT I HAVE SPOKEN WITH A SLEEP SPECIALIST WHO INSISTED USING THE MACHINE FOR THE SAKE OF COMPLIANCE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102759 | 2021-07-23 | 2021-06-28 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION BIPAP | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | PHILIPS / RESPIRONICS, INC. |  |  |  | PHILIPS DREAMSTATION BIPAP MACHINE AND SO CLEAN OZONE SANITATION MACHINE CAUSED HEADACHES, MEMORY LOSS, FATIGUE, AND NASAL IRRITATION AND RESPIRATORY ISSUES AFTER USE. I USED MACHINE AFTER CLEANING IT AND WAS UNABLE TO BREATHE USING MACHINE BECAUSE NOSE WAS FILLED WITH MUCUS. I REMOVED MASK AND DAY LATER BECAME WORSE. I WENT TO URGENT CARE AT (B)(6) TO FIND OUT WHAT WAS WRONG. I RECEIVED MEDICATION AND HAVEN'T USED MACHINE TEST WERE RUN TO RULE OUT OTHER ILLNESSES AND WAS DETERMINED THAT I SUSTAINED A RESPIRATORY INFECTION. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102762 | 2021-07-23 |  | I, n, j, u, r, y | Voluntary report | DREAMSTATION BIPAP ASV | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | PHILIPS / RESPIRONICS |  |  |  | MY MACHINE WAS RECALLED IN JUNE. AFTER STOPPING USE MY HEADACHES, CONSTANT MORNING COUGH, SINUS ISSUES, DIZZINESS STOPPED. HOWEVER I AM HAVING ISSUES FROM SLEEP APNEA AS I CANNOT USE THE MACHINE DUE TO OFF GASSING SINCE I AM HIGHLY SENSITIVE TO CHEMICALS IT IS ESSENTIAL NOT TO USE IT ON THE OTHER HAND, IT IS MEDICALLY NECESSARY TO USE IT. PHILIPS RESPIRONICS HAS NOT COMMUNICATED ANYTHING TO PATIENTS. THEY BLAME THE FDA FOR THE HOLD UP DUE TO APPROVAL IS NECESSARY FOR SAID "PLAN". TO REPLACE THE MACHINE THAT I USE WILL COST (B)(6) WHICH BEING ON SOCIAL SECURITY I DO NOT HAVE. MEDICARE WILL NOT REPLACE THE MACHINE DESPITE THE RECALL NOR WITH THE PROVIDER. SOMETHING MUST BE DONE, 4 MILLION PEOPLE ARE AFFECTED AND THIS IS VERY STRESSFUL FOR US. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102764 | 2021-07-23 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | PHILIPS DREAM STATION | DSX500T11C |  | PHILIPS ISSUED A RECALL ON THEIR CPAP MACHINES. I WAS INSTRUCTED BY PHILIPS TO DISCONTINUE USE OF MY CPAP. I NEED THE INSTRUMENT TO SLEEP EVERY NIGHT BECAUSE I HAVE SLEEP APNEA, A HEART VALVE PROBLEM, AND A PACEMAKER, I CALLED THE PHILIPS CPAP CUSTOMER INQUIRY LINE AND WAS TOLD IT WAS RECOMMENDED TO STOP USING THE INSTRUMENT BUT PHILIPS HAS NOT MADE A DECISION ON REPLACING THE EQUIPMENT. MY PHYSICIAN FELT IT WAS BETTER TO STOP USE OF MY CURRENT PHILIPS CPAP INSTRUMENT UNDER RECALL AND HE PRESCRIBED ANOTHER CPAP FROM A DIFFERENT MANUFACTURER. I RECEIVED A NEW PRESCRIPTION FROM MY DOCTOR BUT MY INSURANCE (B)(6) IS REFUSING TO PAY FOR THE NEW CPAP INSTRUMENT. I SHOULD NOT BE CAUGHT IN THE MIDDLE OF THIS ISSUE, CIGNA SHOULD ALLOW THE PRESCRIPTION TO GO THROUGH AND THEN SEEK REIMBURSEMENT FROM PHILIPS. WHY SHOULD MY HEALTH BE IMPACTED BECAUSE PHILIPS MANUFACTURED AN INSTRUMENT THAT IS INCOMPATIBLE WITH CERTAIN CLEANING METHODS WHICH I DO NOT USE. THE FDA NEEDS TO REQUEST PHILIPS EXPEDITE AS RESOLUTION TO THIS RECALL. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102765 | 2021-07-23 | 2021-07-22 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAMSTATION AUTOCPAPHUMHT DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500T11 |  |  | BURNING IN CHEST/SORE THROAT; CURRENTLY USING A REQUIRED MEDICAL DEVICE (CPAP) TO PREVENT AFIB. THE CPAP MACHINE HAD PARTICLES IN THE WATER TRAY. I HAVE NOT GOTTEN NOTIFIED OF THE RECALL ON THE DEVICE BUT THE WEBSITE LISTED MY MACHINE AS DEFECTED. DREAM STATION AUTOCPAPHUMHT DOM, REF DSX500T11, S/N: (B)(4). FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02379 | 2021-07-23 | 2021-06-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX600H11 | DSX600H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING HEADACHES FOR YEARS SHORTNESS OF BREATH, NASAL IRRITATION, LIGHTHEADED, DIZZY, PATIENT STATES ODOUR COMING FROM THE MECHINE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING HEADACHES FOR YEARS SHORTNESS OF BREATH, NASAL IRRITATION, LIGHTHEADED, DIZZY, PATIENT STATES ODOUR COMING FROM THE MECHINE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, FOLLOWUP REPORT WILL BE FILED.SECTION G1 WAS CORRECTED AND H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102756 | 2021-07-23 | 2021-03-15 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. | DREAMSTATION AUTO CPAP | DSX500T11C |  | I STARTED EXPERIENCING HEADACHES AND EXTREME DIZZINESS ON A DAILY BASIS STARTING AROUND (B)(6) 2021. I THOUGHT AT FIRST IT WAS MY VISION PRESCRIPTION THAT CHANGED SO I GOT A NEW PRESCRIPTION AND NEW GLASSES. THE PROBLEM REMAINED. ABOUT A MONTH AGO I RECEIVED NOTICE FROM MY SLEEP THERAPY DOCTOR THAT THERE WAS A RECALL ON THE PHILIPS DREAMSTATION CPAP DEVICE I WAS USING. I REGISTERED MY SERIAL NUMBER WITH PHILIPS ON (B)(6) 2021. I KEPT USING THE MACHINE UNTIL A WEEK AGO AFTER I READ THAT HEADACHES AND DIZZINESS WERE SYMPTOMS OF THE FOAM INHALATION. THE NEXT DAY AFTER STOPPING, MY HEADACHE WENT AWAY AND MY DIZZINESS WAS SUBSTANTIALLY LESS. TWO DAYS LATER I WAS HEADACHE AND DIZZINESS FREE. NOW MY PROBLEM IS THAT I DON'T GET A GOOD NIGHT'S SLEEP. PHILIPS SAID THEY ARE WAITING ON REGULATORY APPROVAL AND THERE IS NOTHING THEY CAN DO TO REPLACE OR REPAIR MY DEVICE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02442 | 2021-07-23 | 2021-07-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | RIN2110X15B | RIN2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VARIATION IN THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE DELIVERY WAS OBSERVED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN EVALUATED BY THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VARIATION IN THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE DELIVERY WAS OBSERVED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE VENTILATOR'S SOFTWARE WAS RE-LOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02472 | 2021-07-23 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR SHUT DOWN. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "VENTILATOR INOPERATIVE" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S MACHINE FLOW SENSOR, SYSTEM BOARD AND BLOWER ASSEMBLY WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
MW5102748 | 2021-07-23 | 2019-12-30 | I, n, j, u, r, y | Voluntary report | RESPERONICS DREAM STATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS INC. |  |  |  | AFTER STARTING TO USE MY DREAM STATION CPAP I NOTICED MY MIGRAINES WERE ALMOST DAILY AND MY CONGESTION ISSUES SEEMED TO BE GETTING WORSE. I BLAMED IT ON ALLERGIES AT FIRST. AS TIME WENT ON I THOUGHT IT SMELLED WHEN I USED IT. I CHANGED OUT MY FILTERS AS RECOMMENDED BUT IT STILL SMELLED. I HAVE TRIED 3 DIFFERENT MASKS DUE TO HEADACHES. I HAVE BEEN GOING TO THE ENT, ALLERGIST, GI DR, PAIN DR FOR SHOTS FOR MIGRAINES, AND EYE DR. AT WOLFE EYE CLINIC. MY SYMPTOMS ARE FLUSHING, FLUSHING WITH SWEATING ON OCCASIONS, EPISODES OF RAPID HEARTRATE, NAUSEA, DIARRHEA, SEVERE STOMACH CRAMPING, KERATITIS, EAR ACHES, MIGRAINES, FATIGUE, SINUS CONGESTION, BREATHING ISSUES, FACIAL SWELLING AND ALLERGY ISSUES. THE ALLERGIST HAS CONFIRMED THAT I HAVE HISTAMINE INTOLERANCE. I AM HYPERSENSITIVE. I AM CURRENTLY RECEIVING TRIGGER POINT INJECTIONS EVERY 2-3 WEEKS FOR HEADACHES. ALLERGY MEDS H1 AND H2 ANTIHISTAMINES. RX FOR MY EARS, EYES, NOSE, NAUSEA, DIARRHEA, AND CRAMPING. I HAVE BEEN REFERRED TO THE UNIVERSITY OF IOWA HOSPITAL FOR AN EVALUATION BECAUSE MY GI AND THE WOLFE EYE CLINIC THINK IT MAY BE A SYSTEMIC CONDITION DUE TO GLOBAL INFLAMMATION. I HAVE BEEN HAVING BLOOD WORK DONE BY VARIOUS DOCTORS AS WELL TO RULE OUT OTHER CONDITIONS. MY DOCTORS ARE TREATING SYMPTOMS BUT FEEL I NEED A SYSTEMIC WORKUP TO SEE IF ALL OF THESE ISSUES MAY BE RELATED. I WAS NOTIFIED YESTERDAY AT AROUND 3:30 PM BY MY CPAP MEDICAL SUPPLY COMPANY THAT MY MACHINE HAD A RECALL WHEN THEY CALLED TO TELL ME THEY APPROVED MY REQUEST TO LOWER MY PRESSURES. IT WENT FROM 10-15 DOWN TO 8-12. AFTER READING INFORMATION ON THE FDA RECALL SITE I AM WONDERING IF MY ISSUES ARE DUE TO MY CPAP MACHINE THAT TODAY I DISCOVERED IS ONE RECALLED. I DID REGISTER IT ONLINE TODAY. I HAVE ALSO HAD MY HOME CHECKED FOR MOLD AND OTHER TOXINS WHILE GRASPING FOR ANSWERS TO MY MEDICAL SYMPTOMS. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02410 | 2021-07-23 | 2021-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE WAS SCRAPPED DUE TO INSECT INFESTATION. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02412 | 2021-07-23 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02411 | 2021-07-23 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040005 | 1040005 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE BLOWER MOTOR WAS OBSERVED (CONTAMINATION). THE DEVICE'S BLOWER MOTOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
MW5102752 | 2021-07-23 | 2020-12-01 | I, n, j, u, r, y | Voluntary report | PHILLIPS DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS / RESPIRONICS INC. |  |  |  | STARTED USING MY CPAP MACHINE IN (B)(6) 2020. STARTED NOT FEELING WELL IN (B)(6). I HAVE HAD NUMEROUS SINUS INFECTIONS AND UPPER RESPIRATORY ISSUES IN 2021. I HAVE BEEN ON ANTIBIOTICS THREE TIMES AND ALSO DID THREE WEEKS OF STEROIDS. IN (B)(6) I ENDED UP WITH A STREP INFECTION IN MY NOSE. THE BEGINNING OF (B)(6), I DEVELOPED A BAD COUGH AND BRONCHITIS. I WAS GIVEN ANOTHER ROUND OF ANTIBIOTICS. HOWEVER CURRENTLY I AM STILL CONGESTED AND NOW RUNNING A LOW-GRADE FEVER, HEADACHE AND WHITE CELL COUNT IS A LITTLE ELEVATED. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-07
MW5102750 | 2021-07-23 | 2021-06-01 | I, n, j, u, r, y | Voluntary report | PHILLIPS DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS RESPIRONICS INC. | DREAMSTATION |  |  | ON MY CPAP MACHINE, SOMETIME BEFORE THE RECALL, I NOTICED THAT THE FILTERS WERE TURNING EXTREMELY BLACK VERY QUICKLY. MY ONLY SYMPTOMS WERE MILD THAT WERE SINUS RELATED BUT THE SUDDEN BLACK FILTERS WERE CAUSE FOR ALARM FOR ME. NOW THAT I AM AWARE OF THE RECALL CONCERNING THE BLACK FOAM I AM WONDERING IF THE FOAM IS ENTERING THE FILTER SYSTEM AND I WAS BREATHING IT ALL ALONG. I HAVE KEPT THE LAST FILTER I USED AND AM DISCONTINUING USE OF THE MACHINE AS I AM AFRAID THE UNIT IS BAD. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102746 | 2021-07-23 | 2021-07-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS-RESPIRONICS, INC. | DREAMSTNAUTOCPAPHUMHTCELL DOM | DSX500T11C |  | HAD BEEN USING DREAMSTATION CPAP FOR ABOUT SIX MONTHS. AT FIRST IT WAS QUITE HELPFUL; I FELT BETTER RESTED DURING THE DAY. ABOUT A MONTH AGO, I COULD SENSE TOXIC FUMES IN MY AIRWAY AND LUNGS WHICH IS LIKELY FROM THE SOUND ABATING FOAM DUE TO THE RECALL. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02391 | 2021-07-23 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF BLACK PARTICLES, POST NASAL DRIP AND PLUGGED EARS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED VISUALIZATION OF BLACK PARTICLES IN THE DEVICE AND POST NASAL DRIP AND PLUGGED EARS. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION(S) A2, A3, B5 AND D8 WAS CORRECTED. PREVIOUSLY INCORRECT DATE WAS MENTIONED IN SECTION B3 (NOW CORRECTLY UPDATED) AND G3 THE CORRECT DATE SHOULD HAVE BEEN ENTERED AS 06/28/2021. SECTION H6 (HEALTH EFFECT - CLINICAL CODE CORRECTLY UPDATED) UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102782 | 2021-07-26 | 2021-03-01 | I, n, j, u, r, y | Voluntary report | DREAMWEAR | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. |  |  |  | RESPIRATORY ISSUES LED TO NEW DIAGNOSIS OF ASTHMA AFTER ONE YEAR CPAP USE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102781 | 2021-07-26 | 2020-12-16 | I, n, j, u, r, y | Voluntary report | BI PAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I GOT CO2 POISONING AND WAS HOSPITALIZED WITH LIFE THREATENING SYMPTOMS. MAJOR CONFUSION, WEAKNESS IN MY LEGS (I COULD NOT STAND UP OR WALK ON MY OWN EVEN WITH THE USE OF A ROLATOR), LOSS OF BLADDER CONTROL (I HAD TO CATHETERIZED), DIZZINESS, NAUSEA, DEBILITATING HEADACHES. I LITERALLY DIDN'T KNOW WHERE I WAS FOR 2 WEEKS. I WAS THERE A WEEK. I WAS DISCHARGED AND CAME HOME AND RESUMED USE OF MY BIPAP. I WAS READMITTED 3 DAYS LATER AGAIN WITH CO2 POISONING. I WAS THERE ANOTHER WEEK AND THEN DISCHARGED. I REALLY HAVEN'T USED MY BIPAP AGAIN AND HAVE NOT HAD ANYMORE CO2 POISONING. I HAVE HOWEVER RETAINED WEAKNESS IN MY LEGS (I SOMETIMES HAVE TO USE A CANE OR A WALKER), AND MASSIVE HEADACHES. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102780 | 2021-07-26 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | BIPAP PRO BI-FLEX | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS / RESPIRONICS, INC. |  |  |  | I HAVE BEEN USING A PHILIPS RESPIRONICS "SYSTEM ONE BIPAP PRO" MODEL BIPAP MACHINE FOR TREATING OBSTRUCTIVE SLEEP APNEA, SINCE 2010. I WAS ABOUT TO ADD WATER TO THE HUMIDIFIER SECTION OF THE MACHINE ON THE EVENING OF (B)(6) 2020, AND DISCOVERED BITS OF FOAM INSIDE THE WATER RESERVOIR. I EMPTIED THE RESERVOIR, RINSED AND DRIED IT, AND PUT IT BACK IN THE MACHINE - EMPTY. I TURNED ON THE MACHINE, THEN SHUT IT OFF A COUPLE OF MINUTES LATER. MORE BITS OF FOAM WERE PRESENT IN THE RESERVOIR. I RECALLED HAVING SEEN NOTICE OF A RECALL BY PHILIPS SOME DAYS EARLIER IN A SLEEP APNEA GROUP PAGE ON (B)(6). I IMMEDIATELY SEARCHED FOR, AND FOUND, THE NOTICE. I THEN CHECKED THE PHILIPS WEBSITE AND CONFIRMED (WITH THE SERIAL NUMBER) THAT MY MACHINE IS ON THE RECALL LIST. NEXT, I REGISTERED THIS MACHINE WITH PHILIPS, BUT HAVE HAD NO FURTHER COMMUNICATION WITH THEM, EXCEPT TO REGISTER MY WIFE'S MACHINE, WHICH IS ALSO SUBJECT TO THE RECALL. I HAVE NOT USED MY MACHINE SINCE (B)(6). MY BEST RECOLLECTION IS THAT I LAST ADDED WATER TO THE MACHINE ON (B)(6). NO BITS OF FOAM OR OTHER FOREIGN MATTER WAS SEEN THAT NIGHT. THE MACHINE HAS NEVER BEEN CLEANED WITH OZONE OR ANY OTHER CLEANING AGENTS. I MET WITH MY PULMONOLOGIST ON (B)(6), TO START THE PROCESS OF GETTING ANOTHER MACHINE. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02393 | 2021-07-26 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON DEC 06, 2023 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING OF DEVICE MAKING LOUD NOISE, HEADACHES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, NO SECONDARY FINDINGS WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTRE. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED. SECTION D8, D9,G3,H2,H3,H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02285 | 2021-07-26 | 2021-06-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LIVER CANCER WHICH METASTASIZED TO HIS LUNGS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02286 | 2021-07-26 | 2021-06-18 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP BREAST CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING BREAST CANCER, HEADACHE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. SECTION H6 (CLINICAL CODE) CORRECTED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02356 | 2021-07-26 | 2021-06-24 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SHORTNESS OF BREATH. THE REPORTED INFORMATION STATES THE PATIENT WAS PRESCRIBED INHALERS AND ANTIBIOTICS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM AND CAUSED A PATIENT TO DEVELOP SHORTNESS OF BREATH, HEADACHE AND SORE THROAT. THE REPORTED INFORMATION STATES THE PATIENT WAS PRESCRIBED INHALERS AND ANTIBIOTICS. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL AND INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE UNKNOWN BROWN AND LIGHT GRAY CONTAMINATION AT THE MODEM INPUT SIDE OF THE DEVICE. UNKNOWN WHITE AND BLACK CONTAMINATION (DUST-LIKE), INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM, AT THE AIR INPUT OF THE BLOWER BOX. FILM OF LIGHT GRAY UNKNOWN CONTAMINATION ALL THROUGHOUT INSIDE OF ENCLOSURE AND ON TOP OF BLOWER BOX. LIGHT UNKNOWN DUST-LIKE CONTAMINATION ON TOP OF THE BLOWER MOTOR, BLOWER MOTOR ISOLATORS AND THE BLOWER MOTOR SEAL. BLACK CONTAMINATION, CONSISTENT WITH KERATIN, AT THE AIR OUTLET OF THE BLOWER BOX. TINY UNKNOWN BLACK (INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM), AND WHITE SPECS OF CONTAMINATION ON THE BOTTOM THE BLOWER BOX. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND THREE ERORRS. THE MANUFACTURER CONCLUDES THE DEVICE HAS CONTAMINATION. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02210 | 2021-07-26 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02185 | 2021-07-26 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, INTL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | INX500S15 | INX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON JUNE 05, 2022 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02392 | 2021-07-26 | 2021-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING/CHAMBER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL PART OF THE DEVICE AND FOUND DUST LIKE UNKNOWN WHITE AND BLACK CONTAMINANTS INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM AND POTENTIAL VERY SMALL HAIRS OR FIBERS AT THE AIR INPUT OF THE BLOWER BOX. MOIST SPOTS ON THE SOUND ABATEMENT FOAM IN THE BLOWER BOX. SOME OF THE SOUND ABATEMENT FOAM WAS MISSING ESPECIALLY ON THE AIR OUTPUT SIDE. A FEW BLACK PIECES OF SOUND ABATEMENT FOAM ON THE UNDERSIDE OF BLOWER BOX TOP. THE BLOWER MOTOR HAD UNKNOWN WHITE DUST-LIKE CONTAMINATION ON TOP OF IT. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND FOUND THAT BLOWER BOX HAD PIECES OF BLACK FOAM ALL THROUGHOUT IT. THE BLOWER MOTOR HAD POTENTIAL CONGEALED BLACK DEGRADED SOUND ABATEMENT FOAM INSIDE OF IT. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THAT THERE IS CONFIRMED PRESENCE OF DEGRADED SOUND ABATEMENT FOAM ON BOTH SIDES OF THE BLOWER BOX. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02480 | 2021-07-26 | 2021-05-07 | I, n, j, u, r, y | Manufacturer report | RENTAL-DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | RGBX500S15 | RGBX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP A SINUS INFECTION. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION FOR ANTIBIOTICS AND NASAL SPRAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED HEADACHES. DESPITE MULTIPLE ATTEMPTS, THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. CODES IN SECTION H6 WERE UPDATED OR CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102778 | 2021-07-26 |  | I, n, j, u, r, y | Voluntary report | PHILLIPS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C |  |  | SINCE USING MY PHILLIPS DREAMSTATION I HAVE HAD WORSENING ASTHMA SYMPTOMS, A ODD DRY COUGH, SHORTNESS OF BREATH, CHEST TIGHTNESS AND MIGRAINE / CLUSTER HEADACHES. SOME OF THE HEADACHES HAVE REQUIRES ER VISITS. I HAVE BEEN USING THIS DEVICE FOR YEARS AND HAVE NOTICED A GENERAL DECLINE IN MY OVERALL HEALTH. MY DREAMSTATION WAS REPLACED WITH THE SAME MODEL ABOUT 6 MONTHS AGO AND I BELIEVE MY SYMPTOMS HAVE GOTTEN MUCH WORSE WITH THE NEW UNIT. I KEPT FINDING SMALL BLACK PARTICLES IN THE WATER CHAMBER OF THE OLD UNIT. WHETHER THE PROBLEMS I AM EXPERIENCING ARE JUST FROM THE OLD UNIT OR FROM BOTH SINCE THE NEW ONE IS EXACTLY THE SAME I CAN'T BE SURE. HOWEVER, I HAVE EXPERIENCED LESS SEVERE HEADACHES SINCE DISCONTINUING USE OF THE DREAMSTATION. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02276 | 2021-07-26 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | DSX500T11 | RESPIRONICS INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02209 | 2021-07-26 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, INTL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | INX500S15 | INX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02343 | 2021-07-26 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11C | DSX400H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. IN INITIAL REPORT, SECTION G3 WAS INCORRECTLY UPDATED AS 07/07/2021. THE CORRECT DATE SHOULD BE 06/24/2021. CORRECTED INFORMATION WAS PROVIDED IN SECTION B7. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02283 | 2021-07-26 | 2021-06-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CRF PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02357 | 2021-07-26 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP W/HUM/HT/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX200T11C | DSX200T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. THE DEVICE WAS EVALUATED EXTERNALLY. THE MANUFACTURER FOUND UNKNOWN DUST/DIRT CONTAMINATION PRESENT ON ALL SURFACES OF THE BASE UNIT. THE DEVICE DID NOT RETURN WITH AN ACCESSORY PORT FLIP DOOR, SD CARD, OR FILTER. THERE IS AN UNKNOWN DUST CONTAMINATE PRESENT AT THE AIR INLET WHERE THE FILTER WOULD BE AND AT THE ISO PORT ENTRANCE. THERE IS AN UNKNOWN RESIDUE ON THE UI KNOB. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER OBSERVED UNKNOWN DEBRIS IN THE TRACKS OF THE UI PANEL. THERE IS AN UNKNOWN WHITE DUST CONTAMINATION FOUND ON THE BOTTOM ENCLOSURE, BLOWER BOX HOUSING, BLOWER MOTOR, BLOWER IMPELLER, AND REAR PANEL O-RING. THE PRESENCE OF CONTAMINATION THROUGHOUT THE AIRPATH SUGGESTS A SOURCE EXTERNAL TO THE DEVICE THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. SECTIONS D8, D9, H2, H3 AND H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02359 | 2021-07-26 | 2021-06-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGING IRRITATION TO EYES, CONGESTED NOSE, WEIGHT IN CHEST, FACE AND HANDS TURN BRIGHT RED, AND PARTICLES IN TUBING AND CHAMBER. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. ADDITIONAL INFORMATION WAS RECEIVED AND ADDED TO THE REPORT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER ON 01/30/2023 FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED ON 03/29/2023. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. POWERED UP THE DEVICE AND CONFIRMED AIRFLOW USING A KNOWN GOOD POWER SUPPLY AND POWER CORD. DURING REVIEW OF THE ERROR LOGS, PIL DETERMINED THAT THE DEVICE WAS LIKELY RESET PRIOR TO PIL RECEIPT DUE TO THE MACHINE HAVING 5090.7 HOURS AND 0 BLOWER AND THERAPY HOURS. NO CARE ORCHESTRATOR DATA WAS AVAILABLE FOR DOWNLOAD. DURING THE INVESTIGATION OF THE DEVICE, PIL OBSERVED AN UNKNOWN DUST CONTAMINANT ON THE FRONT PANEL, REAR PANEL, BOTTOM ENCLOSURE, ON THE BLOWER, BLOWER SEAL, AND ON THE BLOWER BOX. A KERATIN-LIKE SUBSTANCE WAS OBSERVED ON THE BLOWER BOX OUTLET. (B)(4) ADDRESSES THE ISSUE OF KERATIN. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT IS SCRAPPED. SECTION D8, D9, H2, H3 AND H6 WERE UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02399 | 2021-07-26 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A FAILURE OF A VENTILATOR'S BATTERY TO CHARGE. THERE WAS NO PATIENT HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A FAILURE OF A VENTILATOR'S BATTERY TO CHARGE. THERE WAS NO PATIENT HARM OR INJURY REPORTED. THE DEIVCE WAS EVALUATED BY A THIRD PARTY FIELD SERVICE ENGINEER AND THE CUSTOMER'S COMPLAINT COULD NOT BE DUPLICATED. THE BATTERY WAS ABLE TO CHARGE AS DESIGNED. DURING THE EVALUATION OF THE DEVICE, THE DEVICE WAS FAILING A TEST STEP. THE DEVICE'S PROXIMAL PRESSURE INTAKE TUBE WAS RECONNECTED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02459 | 2021-07-27 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11C | DSX400H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02460 | 2021-07-27 | 2021-06-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON MAY 19, 2024 , AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO BIPAP DEVICE SOUND ABENDENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGED EYE, NOSE, RESPIRATORY TRACT IRRITATIONS, DIZZINESS, HEADACHE, HYPERSENSITIVITY, ASTHMA, INFLAMMATORY RESPONSE AND LUNG DISEASE. THE SECTIONS B1, B2, H1, H2 AND H6 HAVE BEEN UPDATED OR CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR ADVERSE EVENTS (THE PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECTS - CLINICAL CODES AND HEALTH EFFECTS - IMPACT CODE WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02508 | 2021-07-27 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02509 | 2021-07-27 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02510 | 2021-07-27 | 2021-07-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THERE WAS NO PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN THE AIRWAY FROM THE DEVICE, BLOODY NOSE, DRY BLOOD, COUGHING, WHILE SPEAKING GETTING BOTHERED, BREATH IN THE LUNGS EXPANDING AND BEING IN PAIN ON THE RIGHT SIDE, HEADACHES AND DIFFICULTY IN BREATHING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HEALTH EFFECT - CLINICAL CODE HAS BEEN CORRECTED AND TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED. SECTION A1 HAS BEEN CORRECTED / UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02512 | 2021-07-27 | 2021-07-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE DETAILS OF THE DEVICE'S SERIAL NUMBER ARE CURRENTLY UNKNOWN. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED THROAT IRRITATION, SORENESS, PARTICLES IN AIRWAY, DIFFICULTY BREATHING, SHORT OF BREATHE, NAUSEA, DIZZY, HEADACHE. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED THROAT IRRITATION, SORENESS, PARTICLES IN AIRWAY, DIFFICULTY BREATHING, SHORT OF BREATHE, NAUSEA, DIZZY, HEADACHE IN THE DEVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. SECTION H6 HEALTH EFFECT - CLINICAL CODE CORRECTED AND TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02514 | 2021-07-27 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11C | DSX400H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. ADDITIONAL INFORMATION WAS RECEIVED, SECTION B5 WAS CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING FILTER GETS BLACK FAST, ISSUES WITH BREATHING HAS OCCURED AND DIFFICULTY BREATHING/SHORT OF BREATH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. AFTER THE FINAL ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, THE CUSTOMER STATED THAT THEY DID NOT HAVE ANY COMPLAINTS AND HUNG UP THE CALL. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED AND CORRECTED BY CAPTURING HEALTH EFFECT - CLINICAL CODE WHICH WAS MISSED IN INITIAL REPORT. SECTION A1 WAS CORRECTED BY CAPTURING PATIENT IDENTIFIER WHICH WAS MISSED IN INITIAL REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02464 | 2021-07-27 | 2021-06-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, GB | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | RGBX500S15 | RGBX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02542 | 2021-07-27 | 2021-03-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP PERIPHERAL NEUROPATHY, NAUSEA, AND HEADACHES. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102818 | 2021-07-27 | 2020-10-20 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPERONICS DREAMSTATION AUTO BIPAP WITH HUM CEL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX700H11C |  |  | I HAVE HAD PERSISTENT COUGHING AND SWELLING IN MY THROAT AT NIGHT FOR THE PAST YEAR WHILE USING THE PHILIPS RESPIRONICS DREAMSTATION AUTO BIPAP WITH HUM CELL. I USED IT WITH 1 TO 2 L CONT. O2 ALSO. I HAVE REPORTED TO MY MEDICAL PROVIDER THAT IT SMELLED LIKE CHEMICALS AND THAT MY MASK WAS DEGRADING AND STICKING TO FACE. I HAVE SUFFERED COGNITIVE DECLINE OVER THE PAST YEAR, WITH EXTREME FATIGUE, WEAKNESS AND CONFUSION. I HAVE INTERMITTENT TACHYCARDIA AND FEEL QUITE ILL IN THE MORNING. I FOLLOWED THE CLEANING INSTRUCTION FROM THE VA AND DO NOT USE ANY OTHER CLEANING PROCEDURES. I RECEIVED A RECALL NOTICE ON JULY 2, 2021 AND I DISCONTINUED USE OF THE MACHINE AND SUBMITTED FOR THE RECALL. I AM USING THE CONTINUOUS OXYGEN AT NIGHT UNTIL I RECEIVE A NEW MACHINE. I HAD COMPLAINED ABOUT THESE ISSUES FOR QUITE SOME TIME BEFORE THE RECALL. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-07
2518422-2021-02499 | 2021-07-27 | 2021-07-20 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, GER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DEX500S13 | DEX500S13 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP AN EXACERBATION OF ASTHMA. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION FOR AN INHALED STEROID. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED THAT A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM ALLEGEDLY BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP AN EXACERBATION OF ASTHMA. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION FOR AN INHALED STEROID. THE PATIENT ALSO ALLEGED HAVING EYE REDNESS. LUNG IRRITATION (HAS TO TAKE CORTISONE SPRAY), COUGH AND CHOKING. THE DEVICE WAS RETURNED TO THE MANUFACTURER QUALITY PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. DURING THE INITIAL VISUAL INSPECTION OF THE DEVICE THE MANUFACTURER FOUND EVIDENCE OF CONTAMINANTS PRESENT AROUND AND IN THE INLET FILTERS AREA, OUTLET PORT, AND REAR TOP AND BOTTOM ENCLOSURE SURFACES. THE DEVICE WAS DISASSEMBLED FOR INTERNAL VISUAL INSPECTION. THE MANUFACTURER FOUND MULTIPLE (13) CONTAMINATES THROUGHOUT THE UNIT. THE LARGE AMOUNTS OF POTENTIAL MINERAL DEPOSITS AND BEIGE CONTAMINATES LOCATED THROUGHOUT THE INTERNAL BLOWER BOX AND BLOWER. POTENTIAL MINERAL DEPOSITS WERE ALSO LOCATED ON THE BOTTOM AND SOME OF THE SIDES OF THE SOUND ABATEMENT FOAM. THE POTENTIAL MINERAL DEPOSITS RESEMBLE THE USE OF NON-DISTILLED WATER. THE LARGE AMOUNTS OF POTENTIAL MINERAL DEPOSITS FOUND IN THE BLOWER BOX CAN POTENTIALLY BE CONTRIBUTED TO OVER HUMIDIFICATION, OR THE USER OVERFILLING THE HUMIDIFIER RESERVOIR AND OR MOVING THE UNIT WHILE IN USE. THE INNER CANAL OF THE INLET FILTER AREA CONTAINS A LARGE AMOUNT OF DARK DUST OR DIRT CONTAINMENT, AND A BEIGE DUST OR DIRT CONTAINMENT. THIS LARGE AMOUNT OF CONTAMINATION IN THE INNER CANAL OF THE INLET FILTER AREA CAN POTENTIALLY BE CONTRIBUTED TO NOT CHANGING OUT THE POLLEN FILTER REGULARLY AND ALSO THE NON-USE OF AN ULTRA-FINE FILE. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND CONFIRMED THE DEVICE PROVIDED FLOW. THE MANUFACTURER ACCESSED THE POWER LOG BUT AND NO ERROR LOG WAS FOUND. THE MANUFACTURER CONCLUDES THERE IS NO EVIDENCE OF SOUND ABATEMENT FOAM BREAKDOWN WITHIN THE DEVICE. THE CONTAMINATION FOUND IN THE AS OBSERVED IN THE BLOWER AND AIR PATH. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02458 | 2021-07-27 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT ALLEGED OF HAVING HIGH AST IN LIVER, HEADACHES AND SEEING BLACK PARTICLES IN TUBING AND MASK. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY BUT UNABLE TO ADDRESS THE COMPLAINT OR SYMPTOMS. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND TO CONTAIN 1 ERROR CODE. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION A1, SECTION A3, SECTION D9, CODES IN SECTION H6 WERE UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102815 | 2021-07-27 | 2021-07-26 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAM STATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | PHILLIPS RESPIRONICS DREAMSTATION RECALL, REGISTERED THE ITEM AFTER RECALL ON NEWS BUT DRS OFFICE SAID TO KEEP USING THE MACHINE. NOW FDA SAYS SOMETHING DIFFERENT. DRS OFFICE WAS NOT VERY HELPFUL PREVIOUSLY WHEN CALLING ABOUT IT. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | I |  | 2021-07
2518422-2021-02330 | 2021-07-27 | 2021-06-24 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP A RESPIRATORY INFECTION. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECOME DEGRADED AND CAUSED THE PATIENT TO DEVELOP A RESPIRATORY INFECTION. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE EVENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER QUALITY PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. DURING THE INITIAL VISUAL INSPECTION THE MANUFACTURER, THE DEVICE WAS DISASSEMBLED FOR INTERNAL VISUAL INSPECTION. THE MANUFACTURER FOUND NO SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WITHIN THIS DEVICE. THE MANUFACTURER FOUND DARK AND BEIGE DUST OR DIRT CONTAMINANTS, DARK FIBER CONTAMINANTS, POTENTIAL KERATIN CONTAMINANTS, AND DARK SHEEN CONTAMINANTS IN THE DEVICE. THE MANUFACTURER ALSO FOUND A SMALL FOREIGN SILVER DISK OBJECT WITH DARK PAINT THAT FELL OUT OF THE OPEN UNIT, OBJECT WAS NOT LOCATED INSIDE THE AIRPATH. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND TO CONTAIN 1 ERROR CODE. THE MANUFACTURER CONCLUDES THERE IS NO EVIDENCE OF SOUND ABATEMENT FOAM BREAKDOWN WITHIN THE DEVICE. THE MANUFACTURER CONFIRMED CONTAMINATES WERE FOUND SOURCED FROM EXTERNAL ENVIRONMENTAL CONDITIONS. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02329 | 2021-07-27 | 2021-07-07 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP A THRUSH INFECTION OF THE TRACHEA AND STOMACH ISSUES. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE EVENT AND WAS TREATED WITH ANTIBIOTICS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 18, 2024, AND SECTION H11 SHOULD BE REPORTED AS THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP A THRUSH INFECTION OF THE TRACHEA AND STOMACH ISSUES. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE EVENT AND WAS TREATED WITH ANTIBIOTICS. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM (PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR). SECTION B2 WAS CHANGED TO OTHER SERIOUS (PREVIOUSLY IT WAS BLANK). SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02331 | 2021-07-27 | 2021-06-18 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP SINUS INFECTIONS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION FOR ANTIBIOTICS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 12 , 2024 AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP SINUS INFECTIONS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION FOR ANTIBIOTICS. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.). SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.). SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02361 | 2021-07-27 | 2021-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO AUTO W/BT, SWITZ | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CHG502S15 | CHG502S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGED AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED DEVICE IS NOISY, WATER CONTAINER CAME LOOSE WITH HOSE THEN WATER RAN ONTO BED OR THE FLOOR. THE PATIENT ALSO ALLEGED TO MULTIPLE HEADACHE AND INFLAMED MUCOUS MEMBRANE IN NOSE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02386 | 2021-07-27 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO THE DEVICE'S SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02467 | 2021-07-27 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO, INTL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | 761P | 761P |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | IN THE PREVIOUS FOLLOW-UP MDR SUBMISSION, IN SECTION B.2 OF THE 3500A FORM THE "REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE" BOX WAS CHECKED IN ERROR. THERE WAS NO REQUIRED INTERVENTION REPORTED. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02270 | 2021-07-27 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | INX500S15 | INX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02360 | 2021-07-27 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING/CHAMBER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. IN INITIAL REPORT DATE RECEIVED BY MANUFACTURER WAS WRONGLY CAPTURED. IT SHOULD BE 21JUN2021. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02395 | 2021-07-27 | 2021-07-07 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION BIPAP ST30 H/HT/C, DS | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS,INC | DSX1030T11C | DSX1030T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | IN THE PREVIOUS REPORT, SECTION H10 WAS INCORRECT; THE CORRECT SHOULD BE: THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CENTRAL SLEEP APNEA, ASTHMA, AND A WORSE HEART CONDITION. THE PATIENT ALSO REPORTED THAT HER LIVER HAS BEEN "HURTING HER" SHE HAD SEVERE CHEST PAIN WHEN WAKING UP, STOPPED BREATHING AT A RATE OF 114 PER HOUR, AND IS VERY TIRED AND FRUSTRATED. THE PATIENT ALSO REPORTED THAT HER DEVICE IS ALWAYS RINSED WITH WATER, THAT IT GIVES A WEIRD TASTE IN HER MOUTH, AND THAT BOTH FILTERS WERE TOTALLY BLACK WITHIN 3 WEEKS. HER PHYSICIAN SAYS THE PATIENT IS AT RISK OF A HEART ATTACK AND STROKE IF THE DEVICE IS NOT FIXED OR REPLACED AS SOON AS POSSIBLE. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THE MEDICAL INTERVENTION WAS NOT SPECIFIED. UPON REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION, THE CUSTOMER RESPONDED ON 05/18/2023 AND ADDITIONAL INFORMATION WAS RECEIVED, BUT THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. HOWEVER, WE REQUEST THE DEVICE BACK. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02417 | 2021-07-27 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP PRO W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX600H11 | DSX600H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED PARTICLES IN TUBING/CHAMBER. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD CONFIRMED THE DUST CONTAMINATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02384 | 2021-07-27 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO AUTO W/BT | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | HKG502S15 | HKG502S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO THE DEVICE'S SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02465 | 2021-07-27 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02500 | 2021-07-27 | 2021-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, GER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DEX500S13 | DEX500S13 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02468 | 2021-07-27 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, INTL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | INX500S15 | INX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02436 | 2021-07-27 | 2021-07-20 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED A BLACK SPOT ON HER LUNG. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02434 | 2021-07-27 | 2021-07-20 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11W | DSX500T11W |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A NASAL TUMOR. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02229 | 2021-07-27 | 2021-06-17 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP SINUS AND STOMACH INFECTIONS. THE PATIENT REQUIRED MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION FOR ANTIBIOTICS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING WATERY AND ITCHY EYES,CHOCKING, STOMACH INFECTION, DRY MOUTH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02328 | 2021-07-27 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP THROAT IRRITATION, DIZZINESS, AND SHORTNESS OF BREATH. THE PATIENT REQUIRED MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION FOR ANTIBIOTICS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY SUBMITTED MDR 2518422-2021-02328-3 WITH INCORRECT SECTIONS B1, B2, H1, H6. CORRECTIONS TO PREVIOUS MDR ARE MADE IN THIS REPORT AS FOLLOWS. SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN PREVIOUS MDR) SECTION B2 WAS CORRECTED. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6- HEALTH IMPACT CODE WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02230 | 2021-07-27 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP HEADACHES, NAUSEA, SORE THROAT, AND CONGESTION. THE PATIENT REQUIRED MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION ANTIBIOTICS AND STEROIDS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP HEADACHES, NAUSEA, SORE THROAT, AND CONGESTION. THE PATIENT REQUIRED MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION ANTIBIOTICS AND STEROIDS. IN PREVIOUS REPORT, SECTION B5 WAS MENTIONED INCORRECTLY. CORRECT B5 SHOULD BE - THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING HEADACHES, NAUSEA, LIGHTHEADEDNESS, SORE THROAT, AND CONGESTION. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER AND FOUND UNKNOWN DUST LIKE WHITE AND BLACK CONTAMINATION AND SOME POTENTIAL VERY SMALL HAIRS OR FIBERS AT THE AIR INPUT OF THE BLOWER BOX. POTENTIAL LIGHT DUST ON TOP OF THE BLOWER MOTOR AND BLOWER MOTOR SEAL. TINY SPECS OF WHITE CONTAMINATION ON THE BOTTOM OF THE BLOWER BOX. THE MANUFACTURER FOUND EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND 0 ERRORS LOGGED WHILE WAS POWERING DEVICE ON. THE DEVICE WAS APPLIED POWER AND THE DEVICE OPERATED PROPERLY. THE MANUFACTURER CONCLUDES MULTIPLE CONTAMINATION OF DUST, WHITE AND BLACK DUST, WHITE SPECS FOUND WERE EXTERNAL TO THE DEVICE. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. IN THIS REPORT, SECTION B1, B2, D9, G3, H3, H1, H6 HAS BEEN UPDATED OR CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02456 | 2021-07-27 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02385 | 2021-07-27 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX700T11 | DSX700T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO THE DEVICE'S SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02469 | 2021-07-27 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO AUTO W/BT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSG500S11 | DSG500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02396 | 2021-07-27 | 2021-06-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM AND CAUSED IRRITATION, LIGHT HEADNESS AND DIZZINESS. THERE WAS NO SERIOUS OR PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. SECTION(S) D8, D9 AND H6-CLINICAL CODE HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102813 | 2021-07-27 | 2019-05-09 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | CONFIRM NUMBER 2021072601296523 |  |  | DECREASED SPO2 AND CHRONIC SHORTNESS OF BREATH, IDIOPATHIC PULMONARY FIBROSIS. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02532 | 2021-07-27 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S DISPLAY BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02529 | 2021-07-27 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TESTING.. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02530 | 2021-07-27 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1050007 | 1050007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAIL TESTING. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02498 | 2021-07-27 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG.. THE DEVICE'S INTERFACE BOARD AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02523 | 2021-07-27 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLANK. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS EVALUATED AT THE MANUFACTURER'S SERVICE CENTER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S LCD SCREEN NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
MW5102810 | 2021-07-27 | 2021-07-20 | I, n, j, u, r, y | Voluntary report | RESPIRONICS BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX900T11C | DSX900T11C |  | HAVE NEVER HAD KIDNEY PROBLEMS IN THE PAST. ON (B)(6) 2021 THE ONSET OF BACK ISSUES OR WHAT WAS THOUGHT TO BE BACK ISSUES TURNED OUT TO BE KIDNEY PAIN. UPON ARRIVAL TO THE HOSPITAL FOUND OUT A PORTION OF THE KIDNEY HAD BLOOD CLOTS IN IT. THE HOSPITAL COULD NOT DO ANY FURTHER TESTING DUE TO NOT HAVING THE APPROPRIATE TESTING EQUIPMENT. THE HOSPITAL WAS GOING TO TRANSFER ME TO A DIFFERENT HOSPITAL TO DO FURTHER TESTING BUT COULD NOT FIND AN OPEN BED. WE WENT HOME TO TRY AGAIN THE NEXT DAY. ON (B)(6) 2021 ANOTHER ATTEMPT WAS MADE AT A DIFFERENT HOSPITAL. WERE SENT HOME WITH BLOOD THINNERS, PAIN KILLERS, AND ANTI-NAUSEA PILLS, PENDING FOLLOW UP APPOINTMENT WITH GP TO GET REFERRALS TO KIDNEY SPECIALS. I DO NOT KNOW ALL THE TEST RESULTS BUT IT HAS TO DO WITH KIDNEY BLOOD CLOTS. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102809 | 2021-07-27 | 2020-06-30 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS DREAMSTATION ASV | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | DSX500T11C |  |  | IN (B)(6) OF LAST YEAR I REPLACED MY OLDER CPAP MACHINE WITH A NEW PHILIPS RESPIRONICS DREAMSTATION MACHINE. OVER A LITTLE TIME I STARTED HAVING WHAT I THOUGHT WAS ALLERGY ISSUES - COUGH, CLEARING MY THROAT, AND SINUS CONGESTION. TO THE POINT I REPORTED IT TO MY FAMILY DOCTOR WHO PRESCRIBED OVER THE COUNTER ALLERGY MEDICINE ALONG WITH OMEPRAZOLE FOR WHAT WAS THOUGHT TO BE AN ESOPHAGEAL REFLUX ISSUE. I EVEN GOT A CHEST X-RAY TO MAKE SURE IT WASN'T MY LUNGS. I STARTED TAKING THESE MEDS BUT WAS STILL HAVING ISSUES SO I WENT TO AN EAR, NOSE, AND THROAT SPECIALIST WHERE I HAD TESTING FOR ALLERGIES WHICH WERE COSTLY. A MONTH AGO WE WENT ON A TRIP TO (B)(6) WHERE I DIDN'T TAKE THE MACHINE. AMAZINGLY THE ALLERGY SYMPTOMS PRETTY MUCH CLEARED UP IN ONE WEEK. I JUST RECEIVED THE RECALL NOTICE FROM PHILLIPS IN (B)(6) 2021 ON THIS UNIT BUT STILL HAVEN'T HEARD ANYTHING FROM THEM. I SWITCHED BACK TO MY OLDER CPAP MACHINE SINCE THE DREAMSTATION WAS RECALLED, BUT AM NOW STILL HAVING THE COUGHING, IRRITATED THROAT ETC. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102806 | 2021-07-27 | 2020-03-01 | I, n, j, u, r, y | Voluntary report | DREAM STATION AUTOCPAPHUMHT | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. | DSX500T11 |  |  | SINCE (B)(6) 2016 I HAVE USED A DREAM STATION CPAP WITH HUMIDIFIER FOR SLEEP APNEA. STARTING AROUND (B)(6) 2020 I STARTED HAVING SIGNIFICANT UPPER RESPIRATORY ISSUES (CONGESTION, STUFFY NOSE/RUNNY NOSE, COPIOUS DRAINAGE, COUGH, FATIGUE, RIB AND DIAPHRAGM PAIN FROM COUGHING, AND NAUSEA FROM SWALLOWING SO MUCH POST NASAL DRIP; THAT MADE IT DIFFICULT AND THEN IMPOSSIBLE TO USE MY BIPAP. I PURCHASED A DIFFERENT TYPE OF MASK, BUT WASN'T SUCCESSFUL WITH IT EITHER. I INITIALLY ATTRIBUTED ALL THIS TO A BAD ALLERGY SEASON, BUT WHEN IT DIDN'T RESOLVE, IN (B)(6) I CONTACTED MY DOCTOR. I HAVE SEEN MY DOCTOR, A PULMONOLOGIST, AND AN ENT (11 VISITS IN ONE YEAR) AND THEY HAVE BEEN UNABLE TO DEFINITIVELY IDENTIFY THE SOURCE OF MY RESPIRATORY PROBLEMS. THE PULMONOLOGIST ADJUSTED THE SETTINGS ON MY MACHINE - IT DIDN'T HELP. I HAVE HAD A NEW PULMONARY FUNCTION TEST - NOTHING NEW IDENTIFIED, AND A CT OF MY SINUSES, NOTHING THAT WOULD CAUSE ALL MY PROBLEMS. I GOT SO SICK (B)(6) 2020 - THAT I EVEN WENT AND HAD A COVID TEST (NOT COVID.) I JUST LEARNED ON MY (B)(6) NEWS FEED, ON (B)(6) 2021 ABOUT THE RECALL OF CERTAIN PHILIPS DREAM STATIONS AND OTHER MACHINES. I WENT ONLINE AND MINE IS ONE OF THE MACHINES RECALLED. I REGISTERED ON THEIR RECALL SITE ON (B)(6) 2021 BUT HAVE HEARD NOTHING FROM THEM. I AM UPSET THAT PHILIPS DIDN'T NOTIFY ME, AND MY DME DIDN'T NOTIFY ME UNTIL (B)(6) 2021 - ALL THEY SAID WAS TO REGISTER ON THE PHILIPS WEBSITE. I HAVE TO WONDER IF THE TOXIC GASES EMITTED BY MY MACHINE ARE THE SOURCE OF MY TROUBLES. I NEED TO GET BACK TO USING MY BIPAP OR A CPAP - I DON'T SLEEP WELL WITHOUT IT AND HAVE BEEN TAKING NAPS EVERY DAY TO COMPENSATE FOR THE LOST SLEEP. I WORRY THAT I AM GOING TO HAVE TO LIVE WITH THIS UPPER RESPIRATORY DISTRESS THE REST OF MY LIFE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102805 | 2021-07-27 | 2021-07-23 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS DREAMMAKER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500T11 |  |  | (B)(6) 2019 MY PHILIPS DREAMMAKER AUTO CPAP DEVELOPED SOME BLACK MOLD ON THE GASKETS FOR THE HUMIDIFIER WATER TRAY. I REPLACED THE GASKETS AND RAN OZONE THROUGH THE DEVICE TO DISINFECT IT USING A CPAP HOSE OZONE GENERATOR I HAD. I RECEIVED THE PHILIPS RECALL LETTER THIS MONTH AND NOTICED THAT THE EXHAUST PORT OF MY CPAP IS COVERED WITH WHITE DUST OR PARTICLES SO HAVE STOPPED USING IT. THE BRIGHTER BLACK STREAK IS WHERE I WIPED SOME OF THE WHITE DUST OFF WITH A Q-TIP. MY CONCERN IS THAT I HAD HEADACHES IN THE MORNING FOR THE PAST FEW MONTHS AND WONDER IF THIS DUST IS FROM THE NOISE SILENCING FOAM AND IF IT WAS THE CAUSE. YOU MAY HAVE THE DEVICE TO DISASSEMBLE AND ANALYZE AS I WILL NO LONGER BE USING IT (I HAVE A SECOND DEVICE). ALL I ASK IS THAT YOU CONFIRM WHAT THE DUST IS. IT'S DATE OF MANUFACTURE WAS OCT 27, 2015. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | I |  | 2021-07
MW5102804 | 2021-07-27 | 2017-10-30 | I, n, j, u, r, y | Voluntary report | DREAMSTATION ONE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX700H11C |  |  | RECEIVED MY DREAM STATION IN (B)(6) 2017. HAD SURGERY FOR AN UNRELATED PROBLEM IN (B)(6) 2017. I HAD A LARGE HIATAL HERNIA THAT HAD PULLED OVER HALF OF MY STOMACH THROUGH THE DIAPHRAGM AND IT WAS STUCK INSIDE MY CHEST CAVITY. FOLLOWING THAT SURGERY I NOTED MY ABDOMEN WAS INFLATED AND BLOATED FROM THE AIR FROM MY BIPAP. CHECKED WITH MY PULMONOLOGIST OFFICE WHO AGREED TO REDUCE MY PRESSURE LEVEL. CONTINUED WITH SAME PROBLEM AND AIR PRESSURE WAS REDUCED ONCE AGAIN. IN (B)(6) 2020 I HAD EMERGENCY SURGERY DUE TO A NEW HERNIA ON MY RECENT SURGICAL REPAIR ALONG WITH A BOWEL OBSTRUCTION. THE SURGEON FELT THAT THE PRESSURE FROM THE BIPAP CONTRIBUTED TO THE BOWEL BECOMING STUCK IN MY NEW HERNIA. FOLLOWING THIS OCCURRING I SPOKE AGAIN WITH MY PULMONOLOGIST WHO FURTHER REDUCED THE PRESSURE AND STATED HE HAD NEVER HEARD OF ANYTHING LIKE THIS. MY PRESSURE ON THE DREAM STATION WAS REDUCED TO A NON THERAPEUTIC LEVEL FOR MY SLEEP APNEA BUT IT GREATLY REDUCED THE BLOATING I WAS EXPERIENCING. NOW WITH THE RECENT RECALL OF THE DREAM STATION, I HAVE DISCONTINUED THE USE OF THAT DEVICE. I HAVE MY OLDER MACHINE, A RESMED THAT I AM USING NOW WITH THE HIGHER PRESSURE SETTING. I AM NOT EXPERIENCING THE ABDOMINAL BLOATING THAT HAS BEEN OCCURRING. I HAVE COPIES OF THE MEDICAL REPORTS FROM THE HOSPITALIZATIONS IF NEEDED. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102800 | 2021-07-27 | 2021-07-14 | I, n, j, u, r, y | Voluntary report | CPAP DREAM STATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500T11W |  |  | PHILIPS CPAP DREAM STATION WAS RECALLED. I HAD NOT BEEN MYSELF OVER THE PAST 4-6 WEEKS. TIRED, TROUBLE BREATHING DURING RUNNING. AFTER THE RECALL I FIGURED OUT I HAD BEEN BREATHING FOAM HAD A LARGE CHUNK IN MY MOUTH AND SMALL PIECE INSIDE MY MASK THE LAST MORNING I USED IT. I'VE BEEN TO THE DOCTOR AND I'M WORKING THROUGH WHAT ACUTE OR CHRONIC DAMAGE THIS HAS CAUSED ME. CURRENTLY MY DOCTORS SAYS I NEED TO USE A CPAP PHILIPS, MY INSURANCE AND THE MEDICAL DEVICE PROVIDER WILL NOT REPLACE OR FIX THE DEVICE. THERE IS NO TIMETABLE EITHER. PLEASE CALL ME TO DISCUSS. (B)(6). USING A INHALER NOW. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102798 | 2021-07-27 | 2020-07-08 | I, n, j, u, r, y | Voluntary report | CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I HAD BEEN USING THE PHILIPS RESPIRONICS CPAP MACHINE FOR A COUPLE OF YEARS. I DEVELOPED A CHRONIC, MOSTLY NON PRODUCTIVE COUGH IN THE SUMMER OF 2020. I HAD A FEELING THAT THE MORE I USED THE CPAP THE WORSE THE COUGH GOT, BUT COULDN'T VERIFY THAT FEELING. I SAW DOCTORS IN (B)(6) AND AGAIN IN 2021 TO DISCUSS MY CHRONIC COUGH, INCLUDING PCP, PULMONARY NP, ENT. I HAD MULTIPLE COVID TESTS. I HAD CHEST X RAY AND ENDOSCOPY OF THE NASOPHARYNX WHICH DID NOT REVEAL PATHOLOGY. I HAD A THERAPEUTIC TRIAL OF PPIS FOR THE PRESUMPTIVE DX OF REFLUX BUT THAT DIDN'T HELP. I ORDERED A NEW CPAP IN CASE THERE WAS SOMETHING WRONG WITH THE OLD ONE, BUT THAT DIDN'T HELP. I REQUESTED TELEPHONIC RESPIRATORY THERAPY COACHING AND THEY INSTRUCTED ME ON CLEANING; I CLEANED MORE FREQUENTLY USING JUST SOAP AND WATER, BUT THAT DIDN'T HELP. YESTERDAY I GOT A NOTICE THAT THE MODEL OF CPAP I AM USING HAS BEEN RECALLED. ONE OF THE SYMPTOMS THAT HAS BEEN REPORTED IS COUGH. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102797 | 2021-07-27 | 2021-05-10 | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DREAMSTATION |  |  | BREATHING HAD BEEN GETTING WORSE SINCE CPAP PRESSURE WAS INCREASED. BLACK SPECKS IN MACHINE AND TUBING. PHILIPS HAD NOT ANNOUNCED RECALL YET. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102790 | 2021-07-27 | 2012-09-01 | I, n, j, u, r, y | Voluntary report | CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS INC. | DSX500T11 | PRESCRIBED |  | RESPIRATORY ISSUES; NOT CERTAIN OF THE ABOVE DATE. REPORTING THIS AS I RECEIVED A RECALL OF MY PHILIPS CPAP. AFTER USING MY CPAP I DEVELOPED CHRONIC INFLAMMATION OF MY RESPIRATORY SYSTEM RESULTING IN MULTIPLE DR. VISITS WITH ENT'S AND PULMONOLOGISTS. I WAS DIAGNOSED WITH ASTHMA AND OTHER CONDITIONS WHICH HAVE IMPACTED MY QUALITY OF LIFE PROFOUNDLY. I CONTINUE WITH THERAPY'S WITHOUT A CLEAR DIAGNOSIS FROM ANY OF THE DOCTORS I HAVE SEEN. ULTIMATELY DIAGNOSED WITH ASTHMA AND CHRONIC INFLAMMATION OF THE MUCUS MEMBRANE. MRI WITH DIAGNOSIS OF DIFFUSE MUCOSAL THICKENING BILATERAL ETHMOID SINUSES WITH VARIABLE PARTIAL OPACIFICATION SEEN WITH MUCOSAL THICKENING. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102789 | 2021-07-27 | 2021-01-11 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS CPAP DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500T11C |  |  | SEVERAL DOCTOR VISITS; I BEGAN FEELING VERTIGO ON MY RIGHT SIDE WHEN I GO TO SLEEP AT NIGHT AND WHEN I WAKE UP. I AM THINKING IT IS BECAUSE OF THE PHILIPS RESPIRONICS CPAP DREAM STATION I USE FOR MY SLEEP APNEA. THE VERTIGO STARTED JUST WITH LAYING ON MY RIGHT SIDE WHEN I ROLL ONTO MY RIGHT SIDE. I HAVE ABOUT A 30 SECOND SPIN AND THEN IT STOPS. THROUGH THE NIGHT IF I GO BACK TO MY RIGHT SIDE IN MY SLEEP IT OCCURS AGAIN. IN (B)(6) MY SYMPTOMS INCREASED AND THE WEEK OF (B)(6) 2021 I HAD VERTIGO ON BOTH SIDES OF MY HEAD - DIZZY AND NAUSEA SEVERELY AND NEEDED TO STAY IN BED FOR SEVERAL DAYS. I TAKE BONNINE TO HELP BUT IT TOOK SEVERAL DAYS FOR THE HEADACHES AND DIZZINESS TO LIGHTEN. IN (B)(6) I SAW MY PCP AND SHE ADMINISTERED THE EPLEY MANEUVER WHICH I CONTINUED OVER THE NEXT WEEK TRYING TO GET THE VERTIGO SPIN TO STOP. IT NEVER REALLY WORKED COMPLETELY. THE SYMPTOMS SUBSIDED ON MY LEFT SIDE BUT NOT MY RIGHT. SHE ALSO PRESCRIBED AN ANTI-NAUSEA MEDICATION TO HELP WITH THE SYMPTOMS. THEN THE NAUSEA, HEADACHES AND VERTIGO CAME BACK IN (B)(6) AND I BEGAN THE WHOLE PROCESS AGAIN. IN (B)(6) I SAW A NEUROLOGIST WHO DID THE EPLEY MANEUVER 3 TIMES OVER 2 VISITS AND I STILL HAVE DIZZINESS ON MY RIGHT SIDE WHEN I LAY ON THAT SIDE DURING SLEEP AND WHEN I WAKE UP IN THE MORNING I FEEL A WAVE OF DIZZINESS. TWO WEEKS AGO I LEARNED OF THE CPAP RECALL AND REGISTERED MY UNIT. I HAVE BEEN USING MY OLDER CPAP AND STOPPED USING THE PHILIPS BRAND TO SEE IF THAT IS WHAT IS CAUSING MY SYMPTOMS. I DO FEEL MY SYMPTOMS ARE IMPROVING SLIGHTLY. I DON'T KNOW IF THE CPAP IS WHAT IS CAUSING THE VERTIGO, BUT I WANTED TO BRING IT TO YOUR ATTENTION. OMT = OSTEOPATH MANIPULATION TREATMENT. APPOINTMENTS WITH NEUROLOGIST (B)(6). FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02528 | 2021-07-27 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02478 | 2021-07-27 | 2021-07-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | AUX400T15C | AUX400T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP A NASAL POLYP AND AN INVERTED PAPILLOMA. THE PATIENT REQUIRED SURGICAL REMOVAL OF THE INVERTED PAPILLOMA AND NASAL POLYP IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORED AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO DEVELOP A NASAL POLYP AND AN INVERTED PAPILLOMA. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. AN EXTERNAL VISUAL INSPECTED WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF MODERATE DUST CONTAMINATION ON THE POLLEN FILTER AND A SPOT LEFT OF THE THERAPY BUTTON OF UNKNOWN ORIGIN. POWER WAS APPLIED TO THE DEVICE AND THE MANUFACTURER VERIFIED THAT THE DEVICE PROVIDES THERAPY. THE MANUFACTURER REVIEWED THE DEVICE'S EVENT LOGS AND FOUND NO ERRORS LOGGED. THE DEVICE WAS VISUALLY INSPECTED INTERNALLY AND FOUND EVIDENCE OF DUST AND DIRT CONTAMINATES TO THE BLOWER BOX, BLOWER AND BLOWER IMPELLER. THE HUMIDIFIER WAS INSPECTED AND FOUND CARBONACEOUS DEPOSITS ON THE WATER TANK AND DISCOLORED SPOTS ON THE FLIP LID SEAL, AS WELL AS DUST AND DIRT. A MICROSCOPIC REVIEW OF THE CONTAMINANTS WAS COMPLETED BY THE MANUFACTURER. THE COMTAMINATES WERE CONFIRMED TO BE SMALL INSECTS 2-3MM LONG. THERE WAS NO REPORT OF SOUND ABATEMENT FOAM DEGRADATION. THE MANUFACTURER CONFIRMS THE CONTAMINATION IS THAT OF THE OPERATING ENVIRONMENT AND INSECT CONTAMINATION. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02525 | 2021-07-27 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | RIN2110X15B | RIN2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VARIATION IN THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE DELIVERY AND PEAK INSPIRATORY PRESSURE WAS OBSERVED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN EVALUATED BY THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VARIATION IN THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE DELIVERY WAS OBSERVED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE VENTILATOR'S SOFTWARE WAS RE-LOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02526 | 2021-07-27 | 2021-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02527 | 2021-07-27 | 2021-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
MW5102814 | 2021-07-27 |  | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC | DSX500H11C |  |  | SKIN IRRITATIONS UNDER NOSE (LOOKED LIKE STAPH) BEGAN WITHIN A MONTH OF USING MY DREAM STATION CPAP MACHINE. CALLED DOCTOR'S OFFICE TO SEE IF SWITCHING TYPES OF MASK OR NASAL APPLICATOR WOULD HELP. IT DID NOT. MEDICAL EQUIPMENT PERSON SAID TO TRY CLEANING EQUIPMENT MORE AND BE SURE TO CLEAN FACE BEFORE USING MASK. FINALLY CALLED DERMATOLOGIST WHO RECOMMENDED USING CORTISONE CREAM ON AFFECTED AREA. IT CLEARED IT UP SOMEWHAT. HEADACHES IN THE MORNING AND DURING THE NIGHT WERE COMMON (THOUGHT IT WAS THE HEAD STRAP BEING TOO TIGHT SO WOKE UP FREQUENTLY TO READJUST). NASAL CONGESTION EVERY MORNING AND COUGHING AND DRAINAGE ARE COMMON OCCURRENCE. EYE IRRITATION AND DRYNESS EVERY MORNING DUE TO AIR FLOW IN EYES. HAVE STOPPED USING CPAP MASK BASED ON RECALL NOTICE AND THAT IT WAS VERIFIED THROUGH THE SERIAL NUMBER THAT MY MACHINE WAS AFFECTED. WILL BE CALLING MY SLEEP DOCTOR IN THE MORNING TO FIND OUT WHAT TO DO NOW THAT I CANNOT USE THE CPAP MACHINE PHILIPS DREAMSTATION RESPIRONICS. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-07
2518422-2021-02524 | 2021-07-27 | 2021-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
0001822565-2021-02016 | 2021-07-27 | 2021-07-01 | I, n, j, u, r, y | Manufacturer report | UNK TRILOGY SHELL | PROSTHESIS, HIP | ZIMMER BIOMET, INC. | N/A | N/I |  | IT WAS REPORTED THAT THE PATIENT WAS REVISED DUE TO INFECTION APPROXIMATELY 7 MONTHS POST IMPLANTATION. ALL COMPONENTS WERE REMOVED AND ANTIBIOTIC LOADED CEMENT PROSTALAC WAS IMPLANTED. ATTEMPTS HAVE BEEN MADE AND ADDITIONAL INFORMATION ON THE REPORTED EVENT IS UNAVAILABLE. | (B)(4). REPORT SOURCE: (B)(6). MULTIPLE REPORTS WERE SUBMITTED ALONG WITH THIS REPORT: 0001822565-2021-02017. 0001822565-2021-02018. CONCOMITANT MEDICAL PRODUCTS: UNKNOWN ARCOS CONE BODY. UNKNOWN ARCOS DISTAL STEM. TRILOGY SHELL. TRILOGY LINER . TRILOGY SCREW. CUSTOMER HAS INDICATED THAT THE PRODUCT WILL NOT BE RETURNED TO ZIMMER BIOMET FOR INVESTIGATION. THE INVESTIGATION IS IN PROCESS. ONCE THE INVESTIGATION HAS BEEN COMPLETED, A FOLLOW-UP MDR WILL BE SUBMITTED. | Y | N | HEALTH PROFESSIONAL |  | 2021-07
0001822565-2021-02018 | 2021-07-27 | 2021-07-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN TRILOGY SCREW | PROSTHESIS, HIP | ZIMMER BIOMET, INC. | N/A | N/I |  | IT WAS REPORTED THAT THE PATIENT WAS REVISED DUE TO INFECTION APPROXIMATELY 7 MONTHS POST IMPLANTATION. ALL COMPONENTS WERE REMOVED AND ANTIBIOTIC LOADED CEMENT PROSTALAC WAS IMPLANTED. ATTEMPTS HAVE BEEN MADE AND ADDITIONAL INFORMATION ON THE REPORTED EVENT IS UNAVAILABLE. | (B)(4). SOURCE: (B)(6). MULTIPLE REPORTS WERE SUBMITTED ALONG WITH THIS REPORT 0001822565-2021-02016, 0001822565-2021-02017. CONCOMITANT PRODUCTS: UNKNOWN ARCOS CONE BODY, UNKNOWN ARCOS DISTAL STEM, TRILOGY SHELL, TRILOGY LINER , TRILOGY SCREW. CUSTOMER HAS INDICATED THAT THE PRODUCT WILL NOT BE RETURNED TO ZIMMER BIOMET FOR INVESTIGATION. THE INVESTIGATION IS IN PROCESS. ONCE THE INVESTIGATION HAS BEEN COMPLETED, A FOLLOW-UP MDR WILL BE SUBMITTED. DEVICE NOT RETURNED FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2021-07
0001822565-2021-02017 | 2021-07-27 | 2021-07-01 | I, n, j, u, r, y | Manufacturer report | UNK TRILOGY LINER | PROSTHESIS, HIP | ZIMMER BIOMET, INC. | N/A | N/I |  | IT WAS REPORTED THAT THE PATIENT WAS REVISED DUE TO INFECTION APPROXIMATELY 7 MONTHS POST IMPLANTATION. ALL COMPONENTS WERE REMOVED AND ANTIBIOTIC LOADED CEMENT PROSTALAC WAS IMPLANTED. ATTEMPTS HAVE BEEN MADE AND ADDITIONAL INFORMATION ON THE REPORTED EVENT IS UNAVAILABLE. | (B)(4). SOURCE: (B)(6). MULTIPLE REPORTS WERE SUBMITTED ALONG WITH THIS REPORT : 0001822565-2021-02016, 0001822565-2021-02018. CONCOMITANT PRODUCTS: UNKNOWN ARCOS CONE BODY, UNKNOWN ARCOS DISTAL STEM, TRILOGY SHELL, TRILOGY LINER , TRILOGY SCREW. CUSTOMER HAS INDICATED THAT THE PRODUCT WILL NOT BE RETURNED TO ZIMMER BIOMET FOR INVESTIGATION. THE INVESTIGATION IS IN PROCESS. ONCE THE INVESTIGATION HAS BEEN COMPLETED, A FOLLOW-UP MDR WILL BE SUBMITTED. DEVICE NOT RETURNED FOR EVALUATION. | Y | N | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02413 | 2021-07-27 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2200X15B | IN2200X15B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02415 | 2021-07-27 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02414 | 2021-07-27 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR HAD A SENSOR BOARD FAILURE. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS EVALUATED AT A THIRD PARTY SERVICE CENTER AND THE SENSOR BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02501 | 2021-07-27 | 2021-06-21 | I, n, j, u, r, y | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WAS HOSPITALIZED FOR PNEUMONIA. THE MANUFACTURER IS CURRENTLY INVESTIGATING THIS EVENT AND A FOLLOW-UP REPORT WILL BE FILED WHEN THE INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY MARKED H10 INCORRECTLY. THIS MDR WAS NOT SUBMITTED AS PART OF A BATCH SUBMISSION OF COMPLAINTS THAT WERE REASSESSED AS REPORTABLE FOAM DEGRADATION COMPLAINTS DISCOVERED AS PART OF A RETROSPECTIVE REMEDIATION REVIEW. | Y | N | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02496 | 2021-07-27 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TO WORK PROPERLY. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD AND SENSOR BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02495 | 2021-07-27 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" AND "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02475 | 2021-07-27 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | PHILIPS RESPIRONICS INC. | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A TRILOGY VENTILATOR DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO SERIOUS OR PERMANENT INJURY. DESPITE MULTIPLE ATTEMPTS, THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. CODES IN SECTION H6 WERE UPDATED. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02493 | 2021-07-27 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE AND SENSOR BOARD WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02494 | 2021-07-27 | 2021-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S LCD SCREEN NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02531 | 2021-07-27 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
MW5102811 | 2021-07-27 | 2021-04-08 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | CT 48576432C7C | DSX500T11C |  | SEVERE SINUS INFECTION, VERTIGO, URI. I HAVE BEEN USING PHILLIPS DREAMSTATION CPAP MACHINE FOR 2 YEARS. I HAVE CHECKED ON THEIR WEBSITE AND MY MACHINE'S SERIAL NUMBER IS ONE OF THE MACHINES THAT HAS BEEN RECALLED. I HAVE BEEN HAVING SEVERE URI, SINUS INFECTIONS, VERTIGO; NOT REALIZING THESE SYMPTOMS COULD BE A DIRECT RESULT OF USING MY CPAP MACHINE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02544 | 2021-07-27 | 2021-07-07 | D, e, a, t, h | Manufacturer report | BIPAP S/T C SERIES | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS INC | 1067P | 1067P |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND RESULTED IN A PATIENT'S DEATH. NO FURTHER DETAILS REGARDING THE ALLEGATION ARE CURRENTLY KNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02543 | 2021-07-27 | 2021-06-29 | I, n, j, u, r, y | Manufacturer report | BIPAP S/T C SERIES | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC | 1067P | 1067P |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG CANCER. THE DETAILS OF THE MEDICAL INTERVENTION THAT THE PATIENT REQUIRED ARE CURRENTLY UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02450 | 2021-07-27 | 2021-07-20 | I, n, j, u, r, y | Manufacturer report | BIPAP A40 | VENTILATOR, CONTINUOUS, MINIMAL VENTILATORY SUPPORT, FACILITY USE | RESPIRONICS, INC. | 1116156 | 1116156 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BIPAP A40 DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED SEVERE CHEST PAIN. THE PATIENT WAS ADMITTED TO THE HOSPITAL FOR TWO WEEKS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A BIPAP A40 DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED SEVERE CHEST PAIN. THE PATIENT WAS ADMITTED TO THE HOSPITAL FOR TWO WEEKS. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. DURING INVESTIGATION VISUALLY INSPECTED THE A40 SILVER SERIES AND DID NOT OBSERVE ANY VISUAL DEFECTS. REVIEWED ERROR LOGS AND CONSULTED ENGINEERING. NOTHING ABNORMAL APPEARED, INSTALLED A BACTERIA FILTER ON THE OUTLET OF THE UNIT AND RAN THE UNIT AT THE RECEIVED SETTINGS, ON AN ASL LUNG FOR AN HOUR AT APPROXIMATELY 500ML VOLUME WITHOUT ERROR. REMOVED THE BACTERIA FILTER AND INSPECTED IT. NO FOAM WAS OBSERVED IN THE FILTER. UNIT OPERATES AS EXPECTED. THE APNEA ALARM AND THE CIRCUIT DISCONNECT ALARM WERE CHECKED. BOTH ALARMS OPERATED AS EXPECTED. | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102835 | 2021-07-28 | 2019-01-01 | I, n, j, u, r, y | Voluntary report | CPAP DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | J194145830B89 |  |  | ASTHMA, SINUS INFECTION, COUGH; ABOUT 6 MONTHS AFTER USING MY CPAP ,I DEVELOPED A COUGH. MY ASTHMA BECAME WORSE, A POST-NASAL DRIP WAS DEVELOPED ALSO. AFTER SEEING MY FAMILY DOCTOR SHE REFERRED ME TO A PULMONARY DOCTOR. COVID THEN ARRIVED TO (B)(6). I REMAINED IN MY HOME FEARFUL OF GOING ANYWHERE. ALL DOCTORS WERE AFRAID TO SEE NEW PATIENTS WITH A HISTORY OF COUGH. IN (B)(6) I STARTED ALLERGY SHOTS, THINKING THEY COULD HELP, AND AT THE SAME TIME THE PULMONARY DOCTOR GAVE ME AN APPT. I HAVE BEEN ON ALLERGY SHOTS , NUMEROUS STEROID MAINTENANCE INHALERS AND STEROID NASAL SPRAYS. ALSO WENT TO AN ENT FOR THE POST NASAL DRIP AND WAS ABOUT TO GET NASAL SURGERY WHEN I HEARD OF THE RECALL, YOU CAN IMAGINE MY RELIEF, OF THE TIMING. I HAVE REGISTERED WITH PHILIPS, AND HAVE A MACHINE THAT WAS RECALLED. I HAVE NOT USED IT SINCE I BECAME AWARE OF THE RECALL. STILL ON MY MEDICATION, BECAUSE THE ASTHMA, COUGH AND POST NASAL DRIP AND SINUS INFLAMMATION IS STILL A PROBLEM. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102836 | 2021-07-28 | 2021-07-26 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS / RESPIRONICS, INC. | DSX500T11C |  |  | DIAGNOSED WITH CANCER (B)(6) 2019, WAKE UP EVERY MORNING WHEEZING AND HAVING DIFFICULTY BREATHING. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02571 | 2021-07-28 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02573 | 2021-07-28 | 2021-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1032804 | 1032804 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02474 | 2021-07-28 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260B | 1054260B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DARK DEBRIS IN TUBING AND MASK. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE PATIENT HAS ALLEGED DARK DEBRIS IN TUBING AND MASK. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS UPDATED REPORT WILL BE FILED. SECTION E1 CORRECTLY UPDATED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02481 | 2021-07-28 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, IT WAS OBSERVED THE DEVICE'S DISPLAY WAS BLANK. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02570 | 2021-07-28 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02574 | 2021-07-28 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02578 | 2021-07-28 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02575 | 2021-07-28 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040007 | 10400047 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02572 | 2021-07-28 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION OCCURRED RELATED TO OXYGEN. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02577 | 2021-07-28 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY, POWER MANAGEMENT BOARD AND OXYGEN BLENDING MODULE BOARD WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02569 | 2021-07-28 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, A MALFUNCTION OF A VENTILATOR'S RATE WAS OBSERVED. THE DEVICE WAS NOT IN PATIENT USE. THERE WAS EVIDENCE OF CONTAMINATION. THE DEVICE'S SENSOR BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02563 | 2021-07-28 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S DISPLAY WAS BLANK. THERE WAS NO HARM OR INJURY REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02564 | 2021-07-28 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | IN2110X15B | IN2110X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A FAILURE OF THE DEVICE TO DELIVER 100% FIO2 WAS OBSERVED. THE DEVICE'S INTERNAL TUBING WAS RECONNECTED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02487 | 2021-07-28 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02540 | 2021-07-28 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02491 | 2021-07-28 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE BLOWER MOTOR WAS OBSERVED. THE DEVICE'S BLOWER MOTOR NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02488 | 2021-07-28 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS DISPLAYING A "LOW PRESSURE" ALARM. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S INTERNAL TUBING WAS RECONNECTED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02486 | 2021-07-28 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02484 | 2021-07-28 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02489 | 2021-07-28 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S SENSOR BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02490 | 2021-07-28 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE BLOWER MOTOR WAS OBSERVED. THE DEVICE'S BLOWER MOTOR NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02485 | 2021-07-28 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02483 | 2021-07-28 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR DISPLAYS AN INTERNAL BATTERY ERROR MESSAGE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02482 | 2021-07-28 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02541 | 2021-07-28 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO PATIENT HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02558 | 2021-07-28 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2100X15B | IN2100X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, A FAILURE OF THE EXTERNAL FLOW SENSOR WAS OBSERVED. THE DEVICE'S EXTERNAL FLOW SENSOR CABLE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02565 | 2021-07-28 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2100X15B | IN2100X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD ASSEMBLY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02559 | 2021-07-28 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2100X15B | IN2100X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, A FAILURE OF THE EXTERNAL FLOW SENSOR WAS OBSERVED. THE DEVICE'S EXTERNAL FLOW SENSOR CABLE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02536 | 2021-07-28 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SINUS INFECTIONS. THE REPORTED INFORMATION STATES THE PATIENT WAS PRESCRIBED ANTIBIOTICS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING TO DEVELOP SINUS INFECTIONS, COUGHING UP GREEN PHLEGM AND SEEING WHITE PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE REPORTED INFORMATION STATES THE PATIENT WAS PRESCRIBED ANTIBIOTICS. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM (ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR). SECTION B2 WAS MADE BLANK (PREVIOUSLY OTHER SERIOUS WAS CHECKED). SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02358 | 2021-07-28 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. SECTION B5 SHOULD BE PREVIOUSLY REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED DRY COUGH, COUGHING SPELLS THAT GET VERY BAD AND VERY DRY. SINUS PROBLEM HAVE INCREASED A LOT MORE SINUS INFECTIONS. HAS A HEADACHE IN THE MORNINGS AFTER USING THE CPAP. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 (HEALTH EFFECT - CLINICAL CODE) HAS BEEN CORRECTED AND (TYPE OF INVESTIGATION,INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS) HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02538 | 2021-07-28 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT ALSO ALLEGED PARTICLES IN TUBING/CHAMBER, MORE SINUSES PROBLEMS AND PARTICLES IN THE DEVICES. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL AND INTERNAL INVESTIGATION SHOWED THAT THERE WAS EVIDENCE OF FOAM DEGRADATION ASSOCIATED TO THIS ALLEGATION, CONFIRMED POSITIVE DEPOSIT OF FOAM PARTICULATE ON THE FIT TOOL, EVIDENCE OF LIQUID INGRESS. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN IN THE RETURNED MATERIALS. THERE IS NO VISIBLE DAMAGE OR FUNCTIONALITY FAILURES OF THE DEVICE, WHICH SUGGEST THAT THE SOURCE OF CONTAMINATION WAS EXTERNAL TO THE DEVICE. SECTION H6 WERE UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02535 | 2021-07-28 | 2021-06-18 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, AHP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | AHX500H11 | AHX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP NASAL CONGESTION AND HEADACHE. THE PATIENT WAS PRESCRIBED ANTIBIOTICS DUE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING LEFT NOSTRIL IS ALWAYS CLOGGED AND SHOCKING TYPE OF PAIN ON THE RIGHT SIDE OF THE FOREHEAD AND HAD ANTIBIOTICS. THE PATIENT ALSO REPORTED NASAL CONGESTION AND HEADACHE. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED DUE TO AGE. SECTION H6 HEALTH EFFECT - CLINICAL CODE, HEALTH EFFECT - IMPACT CODE HAS BEEN CORRECTED, AND TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED. SECTIONS D8, D9, H2, H3 WERE UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02534 | 2021-07-28 | 2021-06-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG ISSUES AND A SPOT ON THEIR LIVER. THE PATIENT WAS PRESCRIBED MEDICINE FROM THEIR DOCTOR DUE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO EXPERIENCE A HEADACHE, DIZZINESS, NAUSEA, VOMITING, RINGING IN EARS, PNEUMONIA, TROUBLE BREATHING, SEVERE FATIGUE, WEIGHT LOSS AND VISION PROBLEMS, A SPOT IN LIVER AND LUNG ISSUES. THE PATIENT WAS PRESCRIBED MEDICINE FROM THEIR DOCTOR DUE TO THE REPORTED EVENT. THE REPORTED EVENT OF "I HAVE BEEN SICK SINCE AT LEAST JANUARY OF THIS YEAR, IT WASN'T JUST THE DIZZINESS, IT HAD BEEN NAUSEA, VOMITING, SEVERE HEADACHES, TROUBLE BREATHING, DIZZINESS, TUNNEL VISION, A SPOT IN MY LIVER THAT SHOWED UP ON A CT SCAN BECAUSE OF HAVING BREATHING PROBLEMS, SEVERE FATIGUE, AND WEIGHT LOSS OF 31 POUNDS, I HAVE TO GET A CAT SCAN OF MY HEAD AND ABDOMEN AND PELVIS ALSO NOW" AND RINGING EARS AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THE PATIENT WAS PRESCRIBED MEDICINE FROM THEIR DOCTOR DUE TO THE REPORTED EVENT. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02284 | 2021-07-28 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION EXPERT W/HUM/CELL, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | CAX501H12C | CAX501H12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CRF PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02580 | 2021-07-28 | 2021-06-18 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING ASTHMA RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02521 | 2021-07-28 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DEX500S13 | DEX500S13 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02522 | 2021-07-28 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DEX400S13 | DEX400S13 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02473 | 2021-07-28 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO, GER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DEX400S13 | DEX400S13 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR EVALUATION. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND TO CONTAIN 0 ERROR CODES. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION AND THE UNIT PASSED THE FINAL TEST. IN THE PREVIOUSLY SUBMITTED REPORT SECTION B5 WAS INCOMPLETE, CORRECT SECTION B5 SHOULD BE- THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED OF HAVING PARTICLES IN THE NOSE AND HAVING COUGH. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02546 | 2021-07-28 | 2020-12-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED FILM IN THE DISH THAT CAUSED WEIRD ODOR, DEVICE OVERHEATING NEAR THE HUMIDIFIER, COUGH. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE EXTERNAL AND INTERNAL ASPECTS OF THE DEVICE WERE EVALUATED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. UPON EXTERNAL / INTERNAL EXAMINATION, WINGED DEAD BUG FOUND AT AIR INPUT AREA, UNKNOWN WHITE, TAN AND BLACK CONTAMINATION, INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM, AT THE AIR INPUT OF THE BLOWER BOX. PIECE OF BROWN UNKNOWN CONTAMINATION AND POTENTIAL FIBERS ON THE SIDE OF THE INSIDE OF ENCLOSURE. LIGHT DUST-LIKE CONTAMINATION ON THE TOP OF THE BLOWER MOTOR. BLACK CONTAMINATION, CONSISTENT WITH KERATIN, AT THE AIR OUTLET OF THE BLOWER BOX. TINY PIECE OF POTENTIAL HAIR OF FIBER IN THE BOTTOM OF THE BLOWER BOX. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION, BUT CONTAMINATION WAS OBSERVED. SECTION G4 (COMBINATION PRODUCT) WAS INCORRECTLY CHECKED IN FOLLOW-UP 1 AND 2 REPORTS. SECTION G2 CORRECTED (INCORRECTLY CAPTURED IN FOLLOW-UP 1 AND 2 REPORTS). SECTION B5 WAS INCORRECT IN INITIL REPORT AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING FILM IN THE DISH THAT CAUSED WEIRD ODOR, DEVICE OVERHEATING NEAR THE HUMIDIFIER, COUGH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02454 | 2021-07-28 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | INX500S15 | INX500S15 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING THE EVALUATION OF THE DEVICE AT A THIRD-PARTY SERVICE CENTER, THE SOUND ABATEMENT FOAM WAS FOUND TO BE DEGRADED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102834 | 2021-07-28 | 2021-07-10 | I, n, j, u, r, y | Voluntary report | DREAM STATION CI-PAP/ BI-PAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DreamStation Auto CPAP w/Humidifier, w/Cell, US | DSX500H11C |  | STROKE; STARTED WITH COUGHING THAT WOULDN'T GO AWAY. CHEST PAIN. BODY PAIN. NOW IT'S WORSENING CAN'T WALK LONG DISTANCES WITHOUT SHORTNESS OF BREATH. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102847 | 2021-07-28 |  | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAMSTATION CPAP/BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | FOAM FILTER IS DISSIPATING & NEEDS TO BE REPLACED. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102833 | 2021-07-28 | 2021-06-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | THE CPAP PHILIPS DREAMSTATION MADE ME SICK AND GAVE ME BREATHING PROBLEMS. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102830 | 2021-07-28 | 2021-02-25 | I, n, j, u, r, y | Voluntary report | PHILLIPS RESPIRONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | 00606959025646 | DSX500T11C |  | FROM (B)(6) 2020 TO DATE, I HAVE NOTICED AN INCREASE IN HOW OFTEN I HAVE RESPIRATORY ISSUES. FOR A WHILE NOW, I'VE BEEN SUSPECTING THAT IT MAY BE SOMETHING WITH MY CPAP. I STARTED CLEANING IT MUCH MORE FREQUENTLY. INITIALLY, WE THOUGHT IT MIGHT BE DUST ALLERGIES AND STARTED HAVING THE HOUSE CLEANED MORE FREQUENTLY BUT THE ISSUE HAS NOT ABATED. WE'VE NOTICED THAT I'M ALWAYS HAVING RESPIRATORY ISSUES AND HAVE HAD TO TAKE STEROIDS MUCH MORE FREQUENTLY WHICH OFTEN HELPS A LITTLE BIT THEN I GET SICK AGAIN WITHIN WEEKS AND THE CYCLE CONTINUES. THE RESPIRATORY ISSUES TRIGGER VERY BAD MIGRAINES WHICH HAVE BASICALLY PREVENTED ME FROM DOING NORMAL DAY-TO-DAY ACTIVITIES. I'VE NEEDED MUCH MORE MEDICAL CARE THAN I'VE EVER NEEDED IN MY LIFE. I'M CONSTANTLY HAVING RESPIRATORY ISSUES AS A RESULT OF THIS. EVEN AS I TYPE, I'M HAVING INFLAMMATION ISSUES WITH THIS VERY BAD INCESSANT COUGHS THAT AFFECT MY QUALITY OF LIFE. TODAY, ON A ROUTINE CALL FO MY MEDICAL SUPPLIES STORE, I DISCOVERED THAT A RECALL NOTICE WAS ISSUED FOR MY CPAP MACHINES FOR ISSUES SIMILAR TO WHAT I'VE BEEN DEALING WITH FOR OVER 7 MONTHS. I NEVER RECEIVED A LETTER FROM PHILLIPS RESPIRONICS THE MAKER OF MY DREAMSTATION AUTO CPAP. I HAVE COMPLETED THE REGISTRATION AS REQUESTED ON THEIR WEBSITE BUT I'M QUITE UPSET THAT I WAS NOT INFORMED THERE WAS A RECALL. THIS IS TOTALLY UNACCEPTABLE AS IT'S SEVERELY IMPACTED MYSELF AND MY FAMILY. I'VE BEEN UNABLE TO COMPLETE THE PROGRAM I STARTED BECAUSE I'M CONSTANTLY SICK. I WOULD LOVE FOR THIS TO BE INVESTIGATED AND RESOLVED ASAP. PHILLIPS HAS NOT ACCUSED RESPONSIBLY. THEY CLAIM LETTERS HAVE BEEN SENT OUT TO ALL USERS BUT I HAVEN'T RECEIVED ANY. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102829 | 2021-07-28 | 2021-05-22 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS / RESPIRONICS, INC. | 500X110 |  |  | I USED THIS NEW PHILIPS DREAMSTATION AUTO CPAP MACHINE (B)(6) 2016 THROUGH (B)(6) 2021. I NOTICED A PERSISTENT ODOR, LIKE SOMETHING WAS BURNING, WHEN I USED IT. MY CONCERN IS: WHAT WAS I INHALING? COULD IT HAVE A NEGATIVE IMPACT ON MY HEALTH? I STOPPED USING IT WHEN I LEARNED OF THE RECALL, AND REGISTERED IT ON THE PHILIPS RECALL WEBSITE. I HAVE HEARD NOTHING. THE VENDOR, (B)(6), SAID PHILIPS WON'T LET THEM ISSUE A REFUND OR CREDIT. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02537 | 2021-07-28 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500S11 | DSX500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SINUS INFECTIONS. THE REPORTED INFORMATION STATES THE PATIENT WAS PRESCRIBED ANTIBIOTICS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO EXPERIENCE A HEADACHE, SINUS INFECTIONS, RIGHT EYE MATTER. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENT OF HEADACHE, SINUS INFECTIONS, RIGHT EYE MATTER AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02567 | 2021-07-28 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2100X15B | IN2100X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD ASSEMBLY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02479 | 2021-07-28 | 2021-06-24 | D, e, a, t, h | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02533 | 2021-07-28 | 2021-07-06 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP DIFFICULTY BREATHING. THE PATIENT RECEIVED A PRESCRIPTION FOR STEROIDS IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP DIFFICULTY BREATHING. THE PATIENT RECEIVED A PRESCRIPTION FOR STEROIDS IN RESPONSE TO THE REPORTED EVENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02566 | 2021-07-28 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2100X15B | IN2100X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD ASSEMBLY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02443 | 2021-07-28 | 2021-07-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | RIN2110X15B | RIN2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VARIATION IN THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE DELIVERY WAS OBSERVED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE EVALUATION HAS NOT BEEN COMPLETED BY THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VARIATION IN THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE DELIVERY WAS OBSERVED. ADDITIONAL INFORMATION RECEIVED INDICATES THE VENTILATOR'S SOFTWARE WAS RE-LOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02568 | 2021-07-28 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040005 | 1040005 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLANK. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR'S LCD SCREEN WAS BLANK. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS NOT CONFIRMED. NO MALFUNCTION OF THE DEVICE WAS OBSERVED. THE DEVICE DID HAVE EVIDENCE OF PHYSICAL DAMAGE. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02557 | 2021-07-28 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02453 | 2021-07-28 | 2021-06-18 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SINUS INFECTIONS AND DIFFICULTY BREATHING. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON OCTOBER 09, 2023, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING SINUS INFECTIONS AND DIFFICULTY BREATHING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGED CANCER. THE SECTION E1 WAS UPDATED AND SECTIONS B1, B2, H1 AND H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR BOTH ADVERSE EVENTS AND PRODUCT PROBLEM (ONLY THE PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECTS - CLINICAL CODES AND HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02576 | 2021-07-28 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE TO POWER ON WAS OBSERVED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION OF THE DEVICE, A FAILURE OF THE DEVICE TO CHARGE ITS INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
MW5102846 | 2021-07-28 | 2021-01-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP SLEEP THERAPY SYSTEM WITH HUMIDIFIER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I HAD A SERIES OF MEDICAL OCCURRENCES THAT BEGAN IN EARLY (B)(6) 2020 ALTHOUGH NEITHER I NOR MY DOCTORS EVER SENSED A POSSIBLE CONNECTION TO MY PHILIPS RESPIRONICS DREAMSTATION CPAP UNTIL I STARTED READING ABOUT THE RECALL ISSUED ON (B)(6) 2021. AT THAT TIME, I RECALLED TELLING MY HUSBAND THAT SOMETIMES I WOULD WAKE UP AND HAVE THE REMNANTS OF A FORMALDEHYDE-TYPE CAUSTIC ODOR COMING FROM THE HOSE ON THE CPAP. ON (B)(6) 2020 MY PRIMARY DOCTOR REFERRED ME TO AN OTOLARYNGOLOGIST (WHO I SAW (B)(6) 2020) TO TRY TO FIND A CAUSE FOR MY PERSISTENT HOARSENESS. NO DEFINITE CAUSE WAS FOUND. WITHIN SEVERAL MORE MONTHS I WAS HAVING INTERMITTENT BOUTS OF SHORTNESS OF BREATH WITH CHEST TIGHTNESS FOR NO APPARENT REASON. ON (B)(6) 2020, MY ALLERGIST REQUESTED THAT I SEE A CARDIOLOGIST (SEEN (B)(6) 2021) FOR THE SOB. MY INTERNIST FELT I SHOULD ALSO BE SEEN BY A PULMONOLOGIST (SEEN (B)(6) 2021) TO FIND A CAUSE FOR THE INTERMITTENT SOB. NEITHER THE CARDIOLOGY OR PULMONARY ASSESSMENTS CAME UP WITH A CAUSE FOR THE SOB. THE PULMONOLOGIST ORDERED ANOTHER SLEEP STUDY FOR (B)(6) 2021 THAT, UNFORTUNATELY, I WAS UNABLE TO KEEP BECAUSE I WAS HOSPITALIZED FOR ADDITIONAL PROBLEMS. ON (B)(6) 2021 I WAS SENT TO THE ER FOR POSSIBLE GALLBLADDER SYMPTOMS. SOMETIME IN EARLY 2020 I BEGAN TO NOTICE THE REMNANTS OF AN INTERMITTENT FORMALDEHYDE-TYPE CAUSTIC ODOR COMING THROUGH THE AIR HOSE OF MY PHILIPS RESPIRONICS DREAMSTATION CPAP IF I WOKE DURING THE NIGHT. ALTHOUGH I MENTIONED IT TO MY HUSBAND, I DID NOT PURSUE IT. I JUST THOUGHT IT WAS THE WATER CHAMBER BECOMING DRY DURING THE NIGHT WHICH SOMETIMES HAPPENED. I NOTICED THE ODOR INTERMITTENTLY UPON WAKING UNTIL I STOPPED USING THE MACHINE ON (B)(6) 2021 WHEN I FOUND OUT THE PHILIPS RECALL WAS ISSUED AND LEARNED MY MODEL WAS AMONG THE RECALLS. ON (B)(6) 2020, AT MY NORMAL APPOINTMENT WITH MY INTERNIST, DR (B)(6), I COMPLAINED ABOUT MY PERSISTENT HOARSENESS THAT HAD BEEN GOING ON FOR A FEW WEEKS. I HAD PREVIOUSLY HAD A PROBLEM AFTER THYROID SURGERY THAT WAS RESOLVED WITH SPEECH THERAPY BUT THE EXERCISES WERE NOT HELPING NOW. I WAS REFERRED TO DR (B)(6) FROM ASSOCIATED OTOLARYNGOLOGISTS OF PA FOR ANOTHER EVALUATION. AFTER A VIDEOSTROBE AND EXAM, I WAS REFERRED FOR VOICE THERAPY THAT DID NOT SEEM TO BRING ANY IMPROVEMENT. AFTER 2-1/2 MONTHS THE THERAPIST AGREED I WAS NOT BENEFITTING FROM THE THERAPY AND WE STOPPED. THE SHORTNESS OF BREATH BEGAN INTERMITTENTLY SOMETIME DURING THE EARLY SUMMER MONTHS BUT I DID NOT SEE A DOCTOR TO MENTION IT TO UNTIL MY NORMAL ALLERGY APPOINTMENT ON (B)(6) 2021. MY ALLERGIST (DR (B)(6)) BECAME CONCERNED WHEN I TOLD HER ABOUT TIGHTNESS IN THE UPPER CHEST AND SHORTNESS OF BREATH THAT HAD BEEN GOING ON FOR SEVERAL MONTHS. THE ALLERGIST DID AN IN-OFFICE PULMONARY TEST AND SENT ME FOR A CHEST X-RAY. THE INITIAL PULMONARY TEST WAS BAD BUT IT WAS DETERMINED AT THE NEXT APPOINTMENT THAT I WAS NOT GIVEN CORRECT INSTRUCTIONS INITIALLY ON HOW TO TAKE THE TEST. I WAS SCHEDULED FOR ADDITIONAL OFFICE FOLLOW-UPS ON (B)(6) 2020 AND (B)(6) 2021 AND GIVEN TWO ADDITIONAL PULMONARY TESTS (NORMAL) WITH STILL NO DETERMINATION OF THE CAUSE OF THE SOB. AT THE (B)(6)2021 APPOINTMENT I WAS TOLD I SHOULD BE CHECKED BY A CARDIOLOGIST. I AGREED TO TALK TO MY INTERNIST FOR A REFERRAL TO A CARDIOLOGIST. ON (B)(6) 20211 SAW MY INTERNIST (DR (B)(6)) WHO FELT I SHOULD BE EVALUATED BY BOTH A CARDIOLOGIST AND A PULMONOLOGIST FOR THE BREATHING PROBLEMS. HE ALSO ORDERED A CT OF THE CHEST (SINCE I AM A KIDNEY CANCER SURVIVOR) FOR (B)(6) 2021 THAT GAVE AN IMPRESSION OF "UNREMARKABLE CT CHEST". REFERRALS WERE SENT FOR BOTH CARDIOLOGY AND PULMONOLOGY. MY NORMAL BLOOD WORK DONE A WEEK PRIOR TO THIS APPOINTMENT SHOWED CARBON DIOXIDE ONE POINT OUTSIDE OF THE HIGH RANGE AND A GFR NON-AFRICAN AMERICAN THAT WAS SLIGHTLY LOW. OTHERWISE, THE RESULTS WERE THE SAME AS USUAL. I WAS SEEN BY A CARDIOLOGY PA ((B)(6) PA-C) ON (B)(6) 2021 AND, ALTHOUGH HE FOUND NOTHING IN HIS EXAM, HE ORDERED AN ECG (12-LEAD) AND A NUCLEAR STRESS TEST ((B)(6) 2021) - JUST TO BE SURE. THE ORAL REPORT I RECEIVED WAS THAT THEY DID NOT SEE ANYTHING. UNFORTUNATELY, I HAD TO CANCEL THE FOLLOW-UP VISIT SINCE I WAS HOSPITALIZED. MY NEW FOLLOW-UP VISIT FOR THE RESULTS IS SCHEDULED FOR (B)(6) 2021. I MET WITH A PULMONOLOGIST (DR (B)(6)) ON (B)(6) 2021. HE REQUESTED ADDITIONAL BLOOD WORK (NORMAL RESULTS) AND A FULL PULMONARY FUNCTION TEST AT THE HOSPITAL ((B)(6) 2021). HE ALSO PRESCRIBED A RESCUE INHALATOR (ALBUTEROL SULFATE HFA INHALATION AEROSOL) TO BE CARRIED AT ALL TIMES IN CASE OF NEED . THE ONE TIME I USED THE INHALER, IT DID NOTHING TO RELIEVE THE TIGHTNESS OR SOB. AT THE FOLLOW-UP APPOINTMENT ((B)(6) 2021) HE TOLD ME THERE WAS STILL NO CONCLUSIVE RESULTS TO EXPLAIN THE SHORTNESS OF BREATH. THE PULMONOLOGIST THEN WANTED ME TO DO ANOTHER SLEEP STUDY WHICH, UNFORTUNATELY, I HAD TO DELAY SINCE I WAS HOSPITALIZED ON THE DATE OF THE ORIGINAL TEST. THE NEW SLEEP STUDY DATE IS (B)(6) 2021. INTERESTINGLY, THE SHORTNESS OF BREATH SEEMS TO HAVE STOPPED SINCE I DISCONTINUED USING THE RECALLED CPAP ON (B)(6) 2021. TO SUM IT UP, NONE OF THE SPECIALISTS I HAD SEEN WERE ABLE TO COME UP WITH A CAUSE FOR THE SHORTNESS OF BREATH. HOWEVER, THE MORNING OF (B)(6) 20211 EXPERIENCED EXTREME EPIGASTRIC PAIN REMINISCENT OF PRIOR SEVERE GASTRIC INFLAMMATION FROM MY GERO. I CONSULTED WITH MY GI GROUP (PA GI CONSULTANTS P.C.) AND WAS TOLD TO TAKE A SECOND DEXILANT (USED FOR MY SEVERE GERD). THE PAIN DISSIPATED WITHIN A FEW HOURS. THAT EVENING I BEGAN DRY HEAVES, CHILLS AND A LOW-GRADE FEVER. THE NEXT MORNING ((B)(6) 2021) I AGAIN CALLED MY GI GROUP AND WAS INSTRUCTED BY DR (B)(6) TO GO TO THE ER. LATE MORNING ON (B)(6) 2021, I WENT TO THE ER ((B)(6) HOSPITAL ER, DR (B)(6)) AND THEY IMMEDIATELY BEGAN CHECKING ME FOR A POSSIBLE GALLBLADDER PROBLEM DUE TO THE SYMPTOMS ON (B)(6) 2021, ALTHOUGH NONE OF THOSE SYMPTOMS REMAINED AT THAT TIME. NUMEROUS TESTS WERE PERFORMED (ECG, BLOOD WORK AND CULTURES, US, X-RAYS, CT) WITH NO CONCLUSIVE EVIDENCE OF A GALLBLADDER PROBLEM. I WAS TOLD THE BLOOD WORK INDICATED AN ELEVATED WHITE COUNT, INFLAMMATION OF THE LIVER AND GALLBLADDER AND ELEVATED LIVER AND KIDNEY FUNCTIONS. I WAS ALSO TOLD THE ULTRASOUND OF THE GALLBLADDER/ABDOMEN, THE X-RAYS OF THE ABDOMEN WITH CHEST AND THE CT OF THE ABDOMEN AND PELVIS WITH CONTRAST SHOWED NO CONCLUSIVE EVIDENCE THAT I HAD A GALLBLADDER ATTACK OR THAT I PASSED A GALLSTONE. AT THIS POINT I WAS ADMITTED TO THE HOSPITAL FOR OBSERVATION (ALTHOUGH, AT SOME LATER POINT, I WAS ADMITTED AS A PATIENT). MY BLOOD COUNTS WERE MONITORED DAILY (DR (B)(6)) DURING THE HOSPITALIZATION ((B)(6) 2021 THRU (B)(6) 2021). FOR THE NIGHT OF (B)(6) 20211 USED A CPAP PROVIDED BY THE HOSPITAL BUT ON (B)(6) 2021 MY HUSBAND BROUGHT MY DREAMSTATION CPAP (THAT I LATER LEARNED WAS BEING RECALLED) TO THE HOSPITAL FOR ME TO USE. ON (B)(6) 20211 WAS SENT FOR AN MRI OF THE ABDOMEN WITHOUT CONTRAST (DR (B)(6)). AGAIN, THE RESULTS WERE INCONCLUSIVE FOR GALLBLADDER. ON (B)(6) 2021 DR (B)(6) SENT ME FOR A NUCLEAR MEDICINE HEPATOBILIARY SCAN THAT AGAIN DID NOT SHOW CONCLUSIVE EVIDENCE OF AN ACUTE GALLBLADDER ATTACK. DURING THE TIME I WAS IN THE HOSPITAL, I WAS ALSO BEING FOLLOWED DAILY BY DR (B)(6) (PA GI CONSULTANTS, P.C.). DR (B)(6) TOLD ME ON (B)(6) 2021 THAT SHE FELT I SHOULD BE SEEN BY THE GENERAL SURGERY SERVICE. THEY RECOMMENDED THAT I CONSULT WITH PINNACLEHEALTH SURGICAL ASSOCIATES REGARDING THE POSSIBLE NEED FOR THE ELECTIVE REMOVAL OF MY GALLBLADDER. THERE WERE SOME CONFLICTING AND NON-CONCLUSIVE RESULTS FROM ALL OF THE TESTS BUT THE NUCLEAR MEDICINE HEPATOBILIARY SCAN (WHICH WAS GIVEN WITHOUT THE CCK HORMONE) GAVE THE IMPRESSION THAT THE TEST COULD REPRESENT CHRONIC CHOLECYSTITIS. I CHALLENGE THAT IMPRESSION SINCE I HAVE NEVER HAD ANY PREVIOUS SYMPTOMS. UPON RELEASE FROM THE HOSPITAL, I SCHEDULED A FOLLOW-UP WITH MY INTERNIST (DR (B)(6)) AND WITH A GENERAL SURGEON FROM PINNACLEHEALTH SURGICAL ASSOCIATES (DR (B)(6)). DR (B)(6) SAW ME ON (B)(6) 2021 AND DID BOTH BLOOD WORK AND AN EXAM. THE BLOOD WORK WAS NOT QUITE AS GOOD AS IT WAS IN THE HOSPITAL. PLEASE NOTE THAT I WAS USING THE DREAMSTATION CPAP AT THIS TIME; I STARTED USING MY OLDER RESMED S-9 CPAP THE NIGHT OF (B)(6) 2021 WHEN I LEARNED OF THE RECALL. DR (B)(6) REQUESTED THAT I COME BACK FOR MORE BLOOD WORK AND ANOTHER EXAM ON (B)(6) 2021. THE BLOOD WORK FROM THIS VISIT SHOWED IMPROVEMENT BUT WAS STILL NOT BACK TO NORMAL. I WAS AGAIN INSTRUCTED TO COME BACK ON (B)(6) 2021 TO BE SEEN AND REPEAT THE BLOOD WORK. IT HAS CONTINUED TO IMPROVE BUT I AM TO BE SEEN FOR MORE BLOOD WORK ON (B)(6) 2021. MY APPOINTMENT WITH THE GENERAL SURGEON, DR (B)(6), WAS ON (B)(6) 2021. HE REVIEWED ALL OF THE BLOOD WORK AND TESTS THAT HAD BEEN DONE TO DATE AND NOTED THAT THERE WERE CONFLICTING RESULTS. HE REQUESTED THAT I BE SEEN BY MY GI DOCTOR, DR (B)(6), ALTHOUGH I WAS ONLY ABLE TO GET AN APPOINTMENT WITH HIS PA, (B)(6) ON (B)(6) 2021. INITIALLY, MRS (B)(6) WANTED ME TO REPEAT THE HEPATOBILIARY SCAN BUT WITH THE CCK HORMONE. AFTER CONSULTING WITH DR (B)(6) LATER THAT MORNING, SHE CALLED TO TELL ME THAT THEY BOTH FELT I SHOULD HAVE MY GALLBLADDER REMOVED AND THAT I SHOULD SET UP AN APPOINTMENT WITH DR (B)(6). MY SECOND CONSULT WITH DR (B)(6) WAS (B)(6) 2021. I TOLD HIM I AM STILL NOT COMFORTABLE HAVING SURGERY FOR SOMETHING THAT HAD NEVER GIVEN ME PROBLEMS BEFORE AND FOR WHICH WE CAN GET NO CONCLUSIVE PROOF THAT THE PROBLEM REALLY EXISTS. THE EPIGASTRIC PAIN I EXPERIENCED WAS MORE SIMILAR TO THE SEVERE GASTRITIS I HAVE PREVIOUSLY EXPERIENCED FROM MEDICATIONS. I HAD BEEN USING HYDROXYCHLOROQUINE FOR ARTHRITIC PAIN AND FEEL THAT WAS THE CAUSE OF THE EPIGASTRIC PAIN. MY LAST DOSE TAKEN WAS THE EVENING OF (B)(6) 2021 AND THE SEVERE PAIN OCCURRED EARLY THE NEXT MORNING. SINCE AN EGD TEST WAS NOT DONE AT THE HOSPITAL, I HAVE NO WAY OF PROVING THAT BUT I HAVE NOT EXPERIENCED THE PAIN SINCE STOPPING THE HYDROXYCHLOROQUINE TABLETS. EVEN THOUGH DR (B)(6), THE GI DOCTOR, FELT I SHOULD HAVE THE GALLBLADDER SURGERY, DR (B)(6) SAID HE WOULD NOT DO SURGERY AFTER HE CAREFULLY REVIEWED ALL OF MY MEDICAL RECORDS FOR THESE PROBLEMS OVER THE LAST 18 MONTHS AND READ THE TYPE OF PROBLEMS THAT COULD OCCUR BY USING THE RECALLED CPAP. THE PHILIPS RECALL CENTERS AROUND DEGRADED FOAM. THE FOLLOWING IS EXCERPTED FROM THEIR WEBSITE "IN THE EVENT OF EXPOSURE TO DEGRADED FOAM: THE POTENTIAL RISKS OF DEGRADED FOAM EXPOSURE INCLUDE: IRRITATION (SKIN, EYE, AND RESPIRATORY TRACT), INFLAMMATORY RESPONSE, HEADACHE, ASTHMA, ADVERSE EFFECTS TO OTHER ORGANS (E.G, KIDNEYS AND LIVER) AND TOXIC CARCINOGENIC AFFECTS. TODAY, PHILIPS RESPIRONICS HAS RECEIVED SEVERAL COMPLAINTS REGARDING THE PRESENCE OF BLACK DEBRIS / PARTICLES WITHIN THE AIRPATH CIRCUIT (EXTENDING FROM THE DEVICE OUTLET, HUMIDIFIER, TUBING, AND MASK). PHILIPS ALSO HAS RECEIVED REPORTS OF HEADACHE, UPPER AIRWAY IRRITATION, COUGH, CHEST PRESSURE AND SINUS INFECTION. IN THE EVENT OF EXPOSURE TO CHEMICAL EMISSIONS: THE POTENTIAL RISK OF EXPOSURE DUE TO CHEMICAL EMISSIONS FROM AFFECTED FOAM INCLUDE: HEADACHE/DIZZINESS, IRRITATION (EYES, NOSE, RESPIRATORY TRACT, SKIN), HYPERSENSITIVITY, NAUSEA / VOMITING, TOXIC AND CARCINOGENIC EFFECTS. TO DATE, PHILIPS HAS NOT RECEIVED REPORTS OF PT IMPACT OR SERIOUS HARM AS A RESULT OF THIS ISSUE." I HAVE NO DEFINITE PROOF AT THIS TIME THAT MY PHILIPS DREAMSTATION CPAP CAUSE THE PROBLEM BUT IT IS VERY SUSPICIOUS THAT MY SOB STOPPED SHORTLY AFTER DISCONTINUING THE USE OF THIS CPAP. ALSO, UPON RELEASE FROM THE HOSP I WAS STILL USING THE CPAP BECAUSE I WAS NOT AWARE OF THE RECALL AND MY BLOODWORK ON (B)(6) 2021 WAS SLIGHTLY ELEVATED AGAIN. TWO WEEKS LATER MY BLOOD WORK WAS AGAIN TRENDING TOWARDS THE NORMAL VALUES. IT IS NOT BACK TO NORMAL YET BUT GETS BETTER WITH EACH TEST. I HAVE HAD NO RECURRENCE OF SYMPTOMS SINCE STOPPING USE OF THE DREAMSTATION CPAP. I ATTEMPTED TO CONTACT PHILIPS ABOUT MY SYMPTOMS THE DAY I LEARNED OF THE RECALL BUT WAS UNABLE TO REACH THEM BY PHONE. A FEW DAYS LATER, THE WEBSITE ALLOWED YOU TO ENTER YOUR SERIAL NUMBER TO DETERMINE IF YOUR DEVICE WAS AFFECTED. I ENTERED MY SERIAL NUMBER AND CONFIRMED THAT IT WAS ONE OF THE RECALLED DEVICES. ON (B)(6) 2021 I WAS FINALLY ABLE TO SPEAK WITH SOMEONE AT PHILIPS AND REPORT MY EXPERIENCE. OTHER THAN THE CONTACT I INITIATED, I HAVE HAD NO CONTACT FROM PHILIPS OR MY DME SUPPLIER ABOUT THE RECALL AND WOULD STILL BE USING MY DREAMSTATION CPAP IF I HAD NOT SEEN A NEWS ARTICLE ABOUT THE RECALL. I HAVE ONLY USED BLUE (B)(6) TO CLEAN MY CPAP AS I WAS INSTRUCTED WHEN I RECEIVED IT ON (B)(6) 2019. PHILIPS QUESTIONED ME ABOUT USING THE OZONE CLEANERS WHEN WE FINALLY SPOKE BUT I HAVE NEVER DONE SO. |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102845 | 2021-07-28 | 2021-01-21 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DREAM STN AUTO CPAP HUM HT CELLDOM | DSX500T11C |  | SYMPTOMS OVER 6 MONTHS USING THE PHILIPS CPAP I'VE NOTICED WHEEZING, COUGHING AT NIGHT. USED PHILIPS DREAMSTATION FOR 2 YEARS DEVELOPED WHEEZING AND DRY COUGH DURING LAST 6 MONTHS OF USE. INFORMED TO STOP USE BECAUSE PRODUCT WAS RECALLED. |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102844 | 2021-07-28 | 2018-05-02 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION BI-PAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DREAMSTATION |  |  | HEART, HEADACHES, CANCER, ET; I GOT A PHILIPS RESPIRONICS BI-PAP DREAM STATION FROM THE VA SOME TIME AGO IT HAS BEEN A PROBLEM SINCE DAY ONE, IT HAS CAUSED LEAKS, FILTER ISSUES, DID NOT WORK PROPERLY AND MADE ME ILL ETC. IT IS DREAMSTATION SERIAL NUMBER (B)(4). THIS IS THE THIRD OR 4TH ONE WHEN FINALLY I WAS INSTRUCTED TO CEASE USING IT AND GET A NEW ONE? FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102842 | 2021-07-28 | 2021-06-25 | I, n, j, u, r, y | Voluntary report | RESPIRONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | WENT TO PROVIDER TO PURCHASE DISPOSABLE SUPPLIES. WAS TOLD ABOUT MACHINE PROBLEMS AND RECALL. I HAVE BEEN HAVING PROBLEMS FOR ABOUT 3 MONTHS. HAVE BEEN HAVING TERRIBLE HEADACHES, COUGHING, SINUS ISSUES WITH PRESSURE IN MY EYES AND SNEEZING. I THOUGHT THAT THESE WERE AFFECTED BY POLLEN ALLERGIES. I DIDN'T SEE A LINK OR CORRELATION WITH THE CPAP MACHINE UNTIL I READ THE FDA REPORT ONLINE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102837 | 2021-07-28 |  | I, n, j, u, r, y | Voluntary report | PHILLIPS DREAM MACHINE, SLEEP APNEA MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I WAS USING MY PHILLIPS DREAMSTATION CPAP MACHINE AND WAS AWAKENED BY AN INTENSE AND HORRIBLE SMELL COMING FROM THE MACHINE I WAS INHALING. IT WAS INTENSE ENOUGH TO WAKE ME UP. EVEN THOUGH I IMMEDIATELY TOOK IT OFF, THE SMELL PERMEATED ME FOR MOST OF THE DAY IN MY NOSTRILS AND IN MY MOUTH. I STOPPED USING IT COMPLETELY WHEN THIS HAPPENED WHICH WAS ABOUT (B)(6). I THEN CALLED MY SLEEP APNEA DOCTOR LAST WEEK ABOUT THE PROBLEM AND ASKED HOW TO GET A NEW MACHINE SAYING I DIDN'T REALLY WANT TO GET THE SAME BRAND AFTER MY EXPERIENCE WHICH I TOLD HER ABOUT. IT WAS THEN SHE ASKED: YOU HAVE A PHILLIPS DREAM MACHINE? IT WAS THEN I LEARNED OF THE RECALL AND THE NEED TO REGISTER, WHICH I DID TODAY. UP UNTIL THEN, I KNEW NOTHING ABOUT THE RECALL AND HADN'T RECEIVED ANY NOTIFICATION FROM ANYONE, NOT FROM THE DOCTOR WHO PRESCRIBED A MACHINE FOR ME, NOT FROM THE MEDICAL SUPPLY VENDOR I PURCHASED IT FROM NOR FROM THE MANUFACTURER, PHILLIPS. THEY SOUND AS IF THEY HAVEN'T HAD MANY COMPLAINTS AND THEY INDICATE THEY'RE GREAT AT MAKING SURE THEIR MACHINES ARE SAFE. YET HERE WE ARE. HOW DO I LEARN WHAT I'VE INHALED? ARE YOU RESEARCHING THIS TO PROVIDE THE PUBLIC WITH THE COMPLETE INFORMATION? FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102857 | 2021-07-29 | 2019-01-01 | I, n, j, u, r, y | Voluntary report | PHILLIPS DREAMMACHINE CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I HAVE HAD A HEADACHE EVERYDAY FOR THE PAST 2 1/2 YEARS. THEY SOMETIMES LAST MINUTES TO ALL DAY. NOTHING SEAMS TO REALLY HELP. I HAVE HAD MRI'S CAT SCAN AND X-RAYS. I HAVE TRIED PREVENTIVE MEDICINE WITH NO HELP. I DON'T KNOW IF THE CPAP MACHINE HAS CAUSED THIS OR NOT. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-07
MW5102860 | 2021-07-29 | 2020-07-01 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAMSTATION AUTOCPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS / RESPIRONICS, INC. | REF DSX500H11C | PROPERTY OF AMERICAN HOMEPATIE |  | MID 2020 STARTED TO HAVE PERSISTENT SINUS AND UPPER RESPIRATORY CONGESTION WHICH WILL NOT CLEAR UP. PER MY HEART DOCTORS REQUEST I WAS REFERRED TO A PULMONARY DOCTOR FOR A SLEEP STUDY WHICH RECOMMENDED A BIPAP BECAUSE OF THE SEVERITY OF SLEEP APNEA. THE STUDY WAS COMPLETED AND RESULTS INDICATED I NEED THE DEVICE WHICH THE PULMONARY DOCTOR SENT A ORDER TO THE DEVICE PROVIDER TO FILL. DEVICE PROVIDER CONTACTED ME STATING MEDICARE WILL NOT ALLOW THE ORDER BECAUSE 5 YEARS HADN'T LAPSED FROM THE TIME I HAD MY EXISTING DEVICE AND ALSO WOULD NOT ALLOW BECAUSE I SHOULD HAVE BEEN TESTED WITH A CPAP ALSO IN THE SLEEP STUDY. SO, TODAY I HAVE A MACHINE WHICH IS HAZARDOUS TO MY HEALTH CAUSING ONGOING RESPIRATORY ISSUES WHICH IS NOT WHAT MY HEART DOCTOR AND PULMONARY DOCTOR SAID I NEEDED. HOWEVER, "(B)(6) FAILS TO PROVIDE THE DEVICE BECAUSE ?" SLEEP STUDY RESULTS FROM A SLEEP STUDY IN 2021 INDICATED I NEEDED TO HAVE MY PRESENT DEVICE WHICH I RECEIVED IN 2017 REPLACED WITH A BIPAP TO ADEQUATELY TREAT MY SEVERE SLEEP APNEA WHICH NEGATIVELY AFFECTS A HEART CONDITION I HAVE. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | I |  | 2021-07
MW5102862 | 2021-07-29 | 2021-07-26 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS BI-PAP DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. |  |  |  | TBI, STROKES, OTHER SERIOUS; UNABLE TO SLEEP, BREATHING ISSUES, HEADACHES, "OTHER MACHINE HEALTH ISSUES." HORRIBLE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102870 | 2021-07-29 | 2018-10-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS / RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | MY HUSBAND, (B)(6) WAS DIAGNOSED WITH ADENOCARCINOMA OF THE LUNG IN (B)(6) 2017. HE HAD A RIGHT UPPER LOBE LOBECTOMY ON (B)(6) 2017. HE HAS HAD A CHRONIC COUGH EVER SINCE. HE HAS HAD CLEAR LUNG IMAGING SINCE. TOMORROW WE ARE GOING IN FOR IMAGING ON HIS KIDNEY, SOMETHING FOUND WHEN HE HAD AN APPENDECTOMY. HE HAS USED A RESPIRONICS CPAP SINCE 1999. HIS LAST CPAP MACHINE WAS A DREAMSTATION, WHICH WAS USED AND MAY HAVE HAD SO CLEAN USED ON IT AS IT WAS PURCHASED USED FROM A MEDICAL PROVIDER THAT WAS A SLEEP MEDICAL PROVIDER.(MY EMPLOYER- (B)(6) NEVER USED SO CLEAN ON HIS DREAMSTATION, BUT CLEANED IT WITH SOAP, WATER, VINEGAR. I AM (B)(6) WIFE, (B)(6). I AM AFRAID TO ASK ABOUT THE DETAILS OF THE USED DEVICE...IE IF SO CLEAN WAS USED ON IT AS I WORK THERE AND WANT TO KEEP MY JOB). I WORK AS A RN AT A LUNG AND SLEEP CLINIC THAT HAS PRESCRIBED AND DISPENSED RESPIRONICS PAP DEVICES. IT IS POSSIBLE MY HUSBAND'S DEVICE WAS USED FOR CPAP SET UPS IN CLINIC AND CLEANED BETWEEN USES WITH SO CLEAN, BUT I DO NOT KNOW THIS FOR A FACT. HE USED THIS DREAM STATION PRIOR TO THE DIAGNOSIS OF LUNG CANCER, BUT HAS NOT USED CPAP SINCE THE RECALL. HIS LUNG CANCER WAS FOUND INCIDENTALLY ON CARDIAC IMAGING OR LIKELY IT WOULD HAVE NOT BEEN DIAGNOSED AT SO EARLY A STAGE. TEST RESULT CUT ME OFF SO WILL RESTART HERE. HISTOLOGIC TYPE: TUMOR #1 ADENOCARCINOMA, LEPIDIC PREDOMINANT (GRADE 1) (90% LEPIDIC, 10% ACINAR) TUMOR #2 ADENOCARCINOMA, ACINAR PREDOMINANT (GRADE 2) (90% ACINAR, 10% LEPIDIC) HISTOLOGIC GRADE TUMOR #1, (WELL DIFFERENTIATED) OF TUMOR #2,MODERATELY DIFFERENTIATED) OF, VISCERAL PLEURA INVASION NOT IDENTIFIED. TUMOR EXTENSION NOT IDENTIFIED, MARGINS UNINVOLVED BY INVASIVE CARCINOMA (BRONCHIAL, VASCULAR), DISTANCE OF INVASIVE TUMOR FROM CLOSEST MARGIN IS AT LEAST 1 CM. TREATMENT EFFECT: NOT APPLICABLE. LYMPH-VASCULAR INVASION NOT IDENTIFIED, PATHOLOGIC STAGING (PTNM) PRIMARY TUMOR (PT) PT1A. (TUMOR 2 CM OR LESS IN GREATEST DIMENSION, SURROUNDED BY LUNG OR VISCERAL PLEURA, MORE AVAILABLE IF NEEDED). REGIONAL LYMPH NODES (PN) NUMBER OF LYMPH NODES, EXAMINED: 21 NUMBER OF LYMPH NODES, INVOLVED: 0 N1 NODES: 13 N1 POSITIVE NODES: 0 N2 NODES: 8 N2 POSITIVE NODES: 0, MORE INFO IF NEEDED, PATHOLOGIC STAGE SUMMARY: FINAL PTN: PT1AN0 V2016 PATHOLOGIC STAGE: 1A. THE PATHOLOGIC STAGE PRESUMES NO DISTANT METASTASIS. THERE IS MORE TO THE REPORT IF NEEDED. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102876 | 2021-07-29 | 2017-03-13 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION ONE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500T11C |  |  | CHRONIC COUGH STARTED; CHEST CT SHOWED PULMONARY NODULES, NON MALIGNANT. MULTIPLE TESTS TO TRY TO DETERMINE CAUSE OF CHRONIC COUGH; NO DEFINITIVE CAUSE FOUND. MULTIPLE TREATMENTS WITHOUT ANY HELP WITH COUGH. USED PHILIPS DREAMSTATION FOR OVER TEN YEARS; USED SOCLEAN MACHINE FOR MULTIPLE YEARS. PHILIPS NOW HAS A RECALL ON THIS PRODUCT R/T REPORTED ADVERSE EFFECTS. SERIAL # (B)(4), HUMIDIFIER SERIAL # (B)(4). FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102881 | 2021-07-29 | 2021-05-27 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | REV00 | DSXHCP |  | IT WAS NOTED IN A CT SCAN, MY LOWER LOBES IN MY LUNGS HAVE ATELECTASIS AND SWOLLEN AXILLARY LYMPH NODES. RECENT LIVER FUNCTION LABS SHOW LOW ALKALINE PHOSPHATASE. I HAVE NEVER HAD EITHER ISSUE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02581 | 2021-07-29 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP A30-S SILVER SERIES | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS INC | 1111150 | 1111150 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN THE INITIAL REPORT, SECTION B5 WAS INCORRECT, THE CORRECT B5 SHOULD BE- THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DISCOMFORT IN THE OROPHARYNGEAL AREA. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. CLINICAL CODE IN H6 SECTION WAS INCORRECT IN THE PREVIOUS REPORT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND NO EVIDENCE OF CONTAMINATION. AN INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY MANUFACTURER. THE MANUFACTURER FOUND DUST AND DEBRIS CONTAMINATION ON THE AIR INLET FILTER AND IN SOME OTHER AREAS, WATER INGRESS ON THE INSIDE OF THE ENCLOSURE. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERROR LOGGED. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH WATER INGRESS, DUST AND DEBRIS CONTAMINATION. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE PASSED ALL FINAL TESTING. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02635 | 2021-07-29 | 2020-01-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | BLX500H15 | BLX500H15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CAUSED A PATIENT TO DEVELOP LUNG CANCER. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE REPORTED EVENT IS CURRENTLY UNKNOWN. THE DEVICE IS NOT RETURNING TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER REQUESTED THE DEVICE TO BE RETURNED ON 02/24/2022. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT REFUSES TO RETURN THE DEVICE. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102867 | 2021-07-29 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAMSTATIONAUTOCPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500H11C |  |  | BLACK DEBRIS IS IN THE WATER OF THE CPAP MACHINE. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02547 | 2021-07-29 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP AUTOSV H/HT/C, DS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX900T11C | DSX900T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING HEADACHE, WATER BECOMES CLOUDY AND WATER LOOKS FOAMY RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02599 | 2021-07-29 | 2021-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | ITX400H15W | ITX400H15W |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02633 | 2021-07-29 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2100X15B | IN2100X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO PATIENT HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, A MALFUNCTION OF THE VENTILATOR WAS OBSERVED. THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02428 | 2021-07-29 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AUX500T15 | AUX500T15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | SECTION B5 WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT SHOULD BE REPORTED AS FOLLOWS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED INFLAMED AIRWAYS LEADING TO LOW O2. THE PATIENT ALSO REPORTS SOMETHING ENTERED HIS MOUTH TWICE DURING USE AND WAS SWALLOWED. THERE WAS NO PATIENT HARM OR INJURY. NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. IN SECTION D4 ADDITIONAL UNIQUE DEVICE IDENTIFIER (UDIS) NUMBER WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT. IT HAS BEEN CORRECTED OR UPDATED IN THIS REPORT. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02374 | 2021-07-29 | 2017-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/HT/CELL, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | AUX400T15C | AUX400T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. . THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. SECTION H6 CODES CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT HAS BEEN CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02598 | 2021-07-29 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WAS 1 ERROR FOUND. THE MANUFACTURER CONCLUDED THAT, THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02597 | 2021-07-29 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02587 | 2021-07-29 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK PARTICLES IN WATER AND AIRWAY RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02586 | 2021-07-29 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING BLACK PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. AFTER THE FIRST ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, THE CUSTOMER STATED THAT THE DEVICE WAS RETURNED. THE MANUFACTURER NOT YET RECEIVED THE DEVICE IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION G2(COMPANY REPRESENTATIVE), G4 (COMBINATION PRODUCT) WAS INCORRECTLY CAPTURED IN FOLLOW UP 2 REPORT. SECTION A1 NOT CAPTURED IN PREVIOUS REPORTS WAS CAPTURED AND SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02555 | 2021-07-29 | 2021-06-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON OCT 9, 2023, AND SECTION H11 SHOULD BE REPORTED AS: CHRONIC BRONCHITIS WAS MISSED TO CAPTURE IN THE PREVIOUS REPORT. THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING, /CHAMBER, NASAL/THROAT IRRITATION OR SORENESS, CHEST CONGESTION, ASTHMA, SINUS INFECTIONS THAT ARE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGATION OF EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, AND/OR HEADACHE, HYPERSENSITIVITY, ASTHMA, (NEW OR WORSENING) INFLAMMATORY RESPONSE, REDUCED CARDIOPULMONARY RESERVE, CHRONIC BRONCHITIS AND OTHER. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02554 | 2021-07-29 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN DIFFICULTY BREATHING/SHORT OF BREATH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING DIFFICULTY BREATHING/SHORT OF BREATH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02463 | 2021-07-29 | 2021-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECOME DEGRADED AND THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT ALSO ALLEGED SEEING BLACK DEBRIS / PARTICLES IN THE CHAMBER OF THE DEVICE AND HAVING NASAL / THROAT IRRITATION OR SORENESS, DIFFICULTY IN BREATHING, / SHORTNESS OF BREATH, LIGHTHEADNESS, DIZZINESS, HEADACHES. THE DEVICE WAS RETURNED TO THE MANUFACTURER QUALITY PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. DURING THE INITIAL VISUAL INSPECTION THE MANUFACTURER, THE DEVICE WAS DISASSEMBLED FOR INTERNAL VISUAL INSPECTION. THE MANUFACTURER FOUND NO SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WITHIN THIS DEVICE. THE MANUFACTURER FOUND UNIDENTIFIED CONTAMINATION, DISCOLORATION OF BOTTOM ENCLOSURE AND REAR PANEL OF THE DEVICE AND RATTLE NOISE IN THE DEVICE. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND TO CONTAIN 1 ERROR CODE. THE MANUFACTURER CONCLUDES THERE IS NO EVIDENCE OF SOUND ABATEMENT FOAM BREAKDOWN WITHIN THE DEVICE AND UNIDENTIFIED DARK CONTAMINANT WAS FOUND THROUGHOUT THE DEVICE. SECTION H6 HAS BEEN UPDATED / CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02462 | 2021-07-29 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT IS UNABLE TO PROVIDE DETAILS ON THE SERIAL NUMBER OF THE DEVICE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING/CHAMBER,BLACK PARTICLES IN THE DEVICE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION D4 AND H4 WAS MISSING IN THE INITIAL MDR AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02435 | 2021-07-29 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02579 | 2021-07-29 | 2021-06-17 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP DIZZINESS AND SHORTNESS OF BREATH. THE PATIENT WAS HOSPITALIZED AND UNDERWENT A HEART CATHETERIZATION IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A PATIENT REPORTED DIZZINESS AND SHORTNESS OF BREATH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT WAS HOSPITALIZED AND UNDERWENT A HEART CATHETERIZATION IN RESPONSE TO THE REPORTED EVENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02461 | 2021-07-29 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02432 | 2021-07-29 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE 21 ERRORS FOUND. THE MANUFACTURER CONCLUDED THAT, THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. THE DEVICE WAS SCRAPPED DUE TO AGE. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02427 | 2021-07-29 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX700H11 | DSX700H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02426 | 2021-07-29 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400T11 | DSX400T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING MACHINE IS NOT WORKING PROPERLY. THE DEVICE IS MAKING LOUD NOISES AND IT SHOOTS WATER INTO HIS MASK AND BLACK PARTICLES IN WATER CHAMBER AND AIR TUBE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE - THE PATIENT ALLEGED MACHINE IS NOT WORKING PROPERLY. THE DEVICE IS MAKING LOUD NOISES AND IT SHOOTS WATER INTO HIS MASK AND BLACK PARTICLES IN WATER CHAMBER AND AIR TUBE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02636 | 2021-07-29 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY.THE PATIENT ALSO ALLEGED BLACK PARTICLES IN MACHINE. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL PART OF THE DEVICE AND FOUND MULTIPLE INDETERMINATE CONTAMINATES. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND FOUND DUST CONTAMINATION IN THE BLOWER ASSEMBLY. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION, BUT A SMALL AMOUNT OF LIQUID INGRESS,MINERAL DEPOSITS AND DUST CONTAMINATION WERE OBSERVED AND ARE CONSISTENT WITH BEING FROM AN EXTERNAL SOURCE. SECTION H6 WERE UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02631 | 2021-07-29 | 2021-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DEX400S13 | DEX400S13 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER WAS ALSO REPORTED AN ALLEGATION THAT A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE COUGH WITH SPUTUM AND HOARSENESS. THE PATIENT ALSO ALLEGED VISUALIZATION OF SMALL PARTICLES IN THE HOSE AND MASK. THE PATIENT ALSO ALLEGED COUGH WITH SPUTUM AND HOARSENESS. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, NO FOAM PARTICLES WERE OBSERVED IN THE BLOWER KIT. ADDITIONALLY, THE POWER CONNECTOR OF THE UNIT IS CORRODED AND THE UNIT CANNOT BE POWERED ON IN ORDER TO BE ABLE TO COLLECT THE ERROR LOG, MACHINE HOURS, ERROR CODE NUMBERS, AND SOFTWARE VERSION. THE DEVICE HAS BEEN SCRAPPED. SECTION D8, D9, H3, AND H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02373 | 2021-07-29 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED.THE PATIENT DID NOT RECEIVE MEDICAL INTERVENTION.THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. ADDITIONAL INFORMATION AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER RECEIVED INFORMATION ALLEGING NASAL /THROAT IRRITATION, NAUSEA, HEADACHES, DIZZINESS, HAD TWO EPISODES OF WEIRD FEELINGS IN LEGS AND ARMS, DEVICE HAD STRANGE ODOUR, MADE NOISES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM.THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL.THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION.THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME.IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02629 | 2021-07-29 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DEVICE NOISY AND BLACK AROUND MOUTH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02545 | 2021-07-29 | 2021-04-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP A COUGH, SORE THROAT, AND NASAL IRRITATION. THE PATIENT REQUIRED MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION FOR ANTIBIOTICS AND STEROIDS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE /RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM THAT BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE COUGH, SORE THROAT, AND NASAL IRRITATION. THE PATIENT REQUIRED MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION OF ANTIBIOTICS AND STEROIDS. THESE EVENTS WERE ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE INTERNAL AND EXTERNAL PARTS OF THE DEVICE AND FOUND MODEM SIDE DOOR WAS MISSING, UNKNOWN WHITE AND BLACK CONTAMINATION (DUST LIKE), INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM, AND SOME POTENTIAL VERY SMALL HAIRS OR FIBERS AT THE AIR INPUT OF THE BLOWER BOX. POTENTIAL LIGHT DUST ON TOP OF THE BLOWER MOTOR AND BLOWER MOTOR SEAL. TINY SPECS OF WHITE AND BLACK UNKNOWN CONTAMINATION ON THE BOTTOM OF THE BLOWER BOX. YELLOWISH PIECE OF UNKNOWN CONTAMINATION AT THE BOTTOM OF THE BLOWER BOX. BLACK CONTAMINATION, CONSISTENT WITH KERATIN, AT THE AIR OUTLET OF THE BLOWER BOX, CONSISTENT WITH ER 2243857 V01. MANUFACTURER ALSO OBSERVED OBSERVED MOIST SPOTS ON THE SOUND ABATEMENT FOAM IN THE BLOWER BOX AND CONFIRM THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM ON BOTH SIDES OF THE BLOWER BOX. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER VERIFIED FOR GOOD POWER SUPPLY AND POWER CORD, AIRFLOW WAS VERIFIED TO BE OPERATING CORRECTLY. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS OBSERVED IN THIS DEVICE. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02502 | 2021-07-29 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX500T12 | CAX500T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP A COUGH, HEADACHES, AND SWOLLEN EYELIDS. THE PATIENT REQUIRED MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION FOR STEROIDS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 08, 2024, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP A COUGH, HEADACHES, AND SWOLLEN EYELIDS. THE PATIENT REQUIRED MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION FOR STEROIDS. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTIONS B1, B2, H1 AND H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: B1 WAS CORRECTED FOR ONLY THE PRODUCT PROBLEM (BOTH ADVERSE EVENTS AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02470 | 2021-07-29 | 2021-07-20 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AUX400T15C | AUX400T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT EXPERIENCED DYSPNEA, UPPER CHEST PAIN, AND DIZZINESS. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | SECTION B5 WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT SHOULD BE REPORTED AS FOLLOWS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT EXPERIENCED DYSPNEA, UPPER CHEST PAIN,DIZZINESS, INABILITY TO STAND AND ANXIETY. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE EVENT. NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. IN SECTION D4 ADDITIONAL UNIQUE DEVICE IDENTIFIER (UDIS) NUMBER WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT. IT HAS BEEN CORRECTED OR UPDATED IN THIS REPORT. SECTION H6 HEALTH EFFECT CLINICAL CODE WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT. IT HAS BEEN CORRECTED OR UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02632 | 2021-07-29 | 2021-06-30 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/WIFI, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11W | DSX500H11W |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP LYMPHOMA OF THE SMALL INTESTINE. THE PATIENT REQUIRED MEDICAL INTERVENTION IN THE FORM CHEMOTHERAPY AND SURGERY FOR SMALL BOWEL OBSTRUCTION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02617 | 2021-07-29 | 2021-06-25 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION BIPAP PRO W/HUM/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX600H11C | DSX600H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP HYPOXIA THAT RESULTED IN A CEREBROVASCULAR ACCIDENT (CVA). THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED HYPOXIA THAT RESULTED IN A CEREBROVASCULAR ACCIDENT (CVA). PATIENT ALSO ALLEGED SEVERE HEADACHES. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS NOTIFICATION WAS REVIEWED BY THE PMS CLINICAL EXPERT DUE TO THE PATIENT REPORTING WENT TO EMERGENCY ROOM FOR LACK OF OXYGEN TO THE BRAIN AND WAS KEPT OVERNIGHT WHERE THEY DID WITH A LIGHT STROKE AFTER A MRI. HER LEFT SIDE OF BRAIN WAS AFFECTED WHICH CAUSED ISSUES TO RIGHT SIDE OF HER BODY. PATIENT IS EXPERIENCING SEVERE HEADACHES. THEN SHE WAS RUSHED TO THE HOSPITAL LAST NIGHT WITH A SEVERE HEADACHE. BASED ON INFORMATION AVAILABLE AT THIS TIME, THERE IS NOT ENOUGH EVIDENCE TO SUPPORT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE ALLEGED SERIOUS INJURY. THERE IS INSUFFICIENT INFORMATION RELATED TO: DEVICE USE AND MAINTENANCE, TIMELINE OF EVENTS, RELEVANT PAST MEDICAL HISTORY, AND MEDICAL INTERVENTION INFORMATION. HOWEVER, THE LIGHT STROKE EVENT IS ASSESSED AS NOT RELATED TO THE RECALLED DEVICE (I.E. A FOAM DEGRADATION CONCERN) IN THIS CASE. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED IN THE BASE UNIT. OPERATIONAL TESTING SHOWED THE UNIT PROVIDED AIRFLOW. SECONDARY FINDINGS: SLIGHT BLACK CONTAMINATION AT THE BLOWER SEAL OF THE THE BLOWER BOX AND IT IS CONSISTENT WITH THE KERATIN CONTAMINATION. EIGHT ERRORS WERE FOUND. PIL FOUND SLIGHT BLACK CONTAMINATION AT THE BLOWER SEAL OF THE THE BLOWER BOX AND IT IS CONSISTENT WITH THE KERATIN CONTAMINATION. PIL FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. SECTION E1 E-MAIL ADDRESS WAS CAPTURED AND SECTION D8, D9, H2, H3 AND H6 WERE UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02600 | 2021-07-29 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGES PARTICLES IN AIRWAY FROM DEVICE, SHE COMPLAINS SHE HAS A PERSISTING COUGH DURING AND AFTER USE. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGED PARTICLES IN TUBING/CHAMBER, SAYS SHE WAS HAVING TROUBLE BREATHING, WEAK AND NO ENERGY FLARE UP OF COPD/ HOSPITALIZED, THROAT IRRITATION. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION A1 WAS CORRECTED IN THIS REPORT. SECTION B1 WAS CORRECTED FOR BOTH ADVERSE EVENTS AND PRODUCT PROBLEM (ONLY THE PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO HOSPITALIZATION (INITIAL OR PROLONGED)- (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECTS - CLINICAL CODES AND HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102913 | 2021-07-30 | 2021-04-30 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS RESPIRONICS DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. |  |  |  | BOUGHT A NEW PHILIPS DREAMSTATION CPAP MACHINE AND WAS NEVER HAPPY WITH IT FROM DAY ONE, BOUGHT IT IN (B)(6) 2021. I WAS TIRED, SLEEPY ALL THE TIME, HAD A HEADACHE QUITE OFTEN AND DID VISIT THE SUPPLIER AROUND EARLY (B)(6) TO REPORT MY CONCERN. ALSO WENT TO DOCTOR ABOUT MY ABOUT THIS ON (B)(6). JUST FEEL TERRIBLE AS IF I WAS NOT USING A CPAP OR WORSE. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | I |  | 2021-07
2518422-2021-02397 | 2021-07-30 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO, GER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DEX400S13 | DEX400S13 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02642 | 2021-07-30 | 2021-06-03 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP BACTERIAL PNEUMONIA. THE PATIENT WAS HOSPITALIZED AND PRESCRIBED ANTIBIOTICS IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP BACTERIAL PNEUMONIA. THE PATIENT WAS HOSPITALIZED AND PRESCRIBED ANTIBIOTICS IN RESPONSE TO THE REPORTED EVENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02621 | 2021-07-30 | 2021-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTOSV ASSY 60 SRS 30CM,DOM | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | 960T | 960T |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | CORRECTED SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING BLACK PARTICLES IN HIS DEVICE, HEADACHES, DIFFCULTY TAKING A BREATH,OCASSIONAL DRY THROAT RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 CLINICAL CODING CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02602 | 2021-07-30 | 2021-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02603 | 2021-07-30 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX700S11 | DSX700S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DIZZY AND SEEN BLACK PARTICLES RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. AFTER THE FINAL ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, CUSTOMER DECLINED TO RESPOND TO THE GFE RELATED QUESTIONS AND TERMINATED THE CALL.THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02604 | 2021-07-30 | 2021-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02605 | 2021-07-30 | 2021-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS INITIALLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN INITIAL REPORT SECTION B5 WAS MENTIONED INCORRECTLY, CORRECT B5 SHOULD BE - THE PATIENT ALLEGED VISUALIZATION OF PARTICLES IN THE DEVICE, IRRITATION, COUGHING, SNEEZING. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND FLORAL SMELL IN THE BLOWER MOTOR. ALSO FOUND CONTAMINATION ON THE HEATING PLATE AND ON THE RUBBER SEAL FOR THE RESERVOIR. THE MANUFACTURER RECEIVED HUMIDIFIER AND FOUND CONTAMINATION IN THE HUMIDIFIER. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND 1 ERRORS. THE DEVICE WAS APPLIED POWER AND THE DEVICE OPERATED PROPERLY. THE MANUFACTURER CONCLUDES CONTAMINATION FOUND WERE CONSISTENT WITH CONTAMINATION ON THE HEATING PLATE AND ON THE RUBBER SEAL FOR THE RESERVOIR. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. IN THIS REPORT, SECTION D9, D10, G3, H3, H6 HAS BEEN UPDATED OR CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02606 | 2021-07-30 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700S11 | DSX700S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF BLACK PARTICLES IN TUBING/CHAMBER RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALSO ALLEGED HEADACHES, NASAL AND THROAT IRRITATION. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF BLACK PARTICLES IN TUBING/CHAMBER RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALSO ALLEGED HEADACHES, NASAL AND THROAT IRRITATION. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS UPDATED REPORT WILL BE FILED. SECTION B5 CORRECTED AND H6 (CLINICAL CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS) UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02607 | 2021-07-30 | 2021-12-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED HEADACHE, SHORTNESS OF BREATH, NASAL AND THROAT IRRITATION OR SORENESS, DIZZINESS. THERE WAS NO REPORT SERIOUS OR PERMANENT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND INSECTS IN DEVICE, WATER INGRESS IN THE BLOWER BOX.. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND ZERO ERRORS. THE DEVICE WAS APPLIED POWER AND THE DEVICE OPERATED PROPERLY. THE MANUFACTURER CONCLUDES CONTAMINATION FOUND WERE CONSISTENT WITH INSECT IN THE DEVICE AND WATER INGRESS CONSISTENT WITH BLOWER BOX. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. IN THIS REPORT, SECTION D9, G3, H3, H6 HAS BEEN UPDATED OR CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02608 | 2021-07-30 | 2021-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF (B)(4). THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02609 | 2021-07-30 | 2021-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02614 | 2021-07-30 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02644 | 2021-07-30 | 2021-06-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500S11 | DSX500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON MAR 25, 2025 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING SHORTNESS OF BREATH, COUGH, SINUS INFECTION, HEADACHES, LIGHTHEADED, ASTHMA, SNEEZING, ITCHING, MUCUS. THE PATIENT DID NOT REPORT TO RECEIVE MEDICAL INTERVENTION ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGED VISUALIZATION OF PARTICLES. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02634 | 2021-07-30 | 2021-06-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | RIN2110X15B | RIN2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SET POSITIVE END EXPIRATORY PRESSURE DECREASED WITHOUT MANUAL ADJUSTMENT. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR'S SET POSITIVE END EXPIRATORY PRESSURE DECREASED WITHOUT MANUAL ADJUSTMENT. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-07
2518422-2021-02640 | 2021-07-30 | 2021-06-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A SPOT IN THE PATIENT'S LUNG. THE PATIENT WAS HOSPITALIZED DUE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING SPOT IN THE PATIENT'S LUNG, BREATHING PROBLEM, HEADACHES AND VOMIT RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT WAS HOSPITALIZED DUE TO THE REPORTED EVENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02620 | 2021-07-30 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED ON THIS DEVICE IN MDR 2518422-2021-02099. THIS REPORT WAS SUBMITTED AS A DUPLICATE REPORT OF THE PREVIOUSLY SUBMITTED REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02520 | 2021-07-30 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE REPORTER ALSO ALLEGED VISUALIZATION OF DUST PARTICLES IN TUBING AND PATIENT HAVING HEADACHES, COUGH, TROUBLE TO SLEEP. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND FOUND EVIDENCE OF WATER INGRESS IN BLOWER BOX AND ON BLOWER. ALSO FOUND MULTIPLE CONTAMINANTS THROUGHOUT THE DEVICE. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION BUT PRESENCE OF MULTIPLE SOURCES OF CONTAMINATION AND WATER INGRESS IN THE AIR PATHWAY, WHICH IS CONSISTENT WITH THE SOURCE OF CONTAMINATION BEING EXTERNAL. SECTION D9, H3 AND H6 WERE UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02618 | 2021-07-30 | 2021-06-25 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX400H11C | DSX400H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON OCT 09, 2023, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. ADDITIONAL INFORMATION WAS RECEIVED ALLEGATION OF KIDNEY DISEASE/TOXICITY. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02344 | 2021-07-30 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX500T12 | CAX500T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED HEADACHES, RUNNY NOSE, SHORTNESS OF BREATH, NASAL CONGESTION, PARTICLES IN THE AIRWAY AND ELECTRICAL ODOR. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING HEADACHES, RUNNY NOSE, SHORTNESS OF BREATH, NASAL CONGESTION, PARTICLES IN THE AIRWAY AND ELECTRICAL ODOR RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02362 | 2021-07-30 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/HT/CELL, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX400T12C | CAX400T12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED BURNING/DRY THROAT SENSATION, SHORTNESS OF BREATH, PATIENTS VOICE COMES AND GOES AND BLACK PARTICLES IN THE WATER TANK. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND UNIT SCRAPPED. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02418 | 2021-07-30 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02643 | 2021-07-30 | 2021-06-17 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP BREAST CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING LIGHT HEADED, CHRONIC SORE THROAT, IRRITATION OF NASAL PASSAGES, HEADACHES AND FLARES UP ASTHMAS, AND SOME NAUSEA , AND IS A BREAST CANCER RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102915 | 2021-07-30 | 2021-06-25 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS, INC. |  |  |  | START WITH SORE THROAT ON (B)(6) 2021, GOT CONTINUALLY WORSE OVER NEXT 3 DAYS. WENT TO (B)(6) ON (B)(6) 2021 STARTED ANTIBIOTIC. THROAT TIGHTENING (B)(6) 2021. COULD NOT OPEN MOUTH OR SWALLOW ANTIBIOTICS (B)(6) 2021. WENT TO PRIMARY PHYSICIAN (B)(6) 2021 AND SHE SENT ME TO SPECIALIST TO LANCE ABSCESS ON MY TONSIL. HAD PICK LINE PUT IN AND 4 INFUSION TREATMENTS OVER NEXT 32 HOURS. CONTINUED ON LARGE DOSAGE OF ANTIBIOTICS FOR 2 WEEKS. LOSS OF WORK, DOCTOR BILLS, INFUSION BILLS, LOSS OF WEIGHT. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02641 | 2021-07-30 | 2021-06-24 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP BLACK MOLD IN THEIR LUNGS. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02420 | 2021-07-30 | 2021-07-19 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX500T12 | CAX500T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A COUGH, RUNNY EYES, AND A RUNNY NOSE. THE REPORTED INFORMATION STATES THE PATIENT WAS PRESCRIBED TYLENOL AND ANTIBIOTICS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT IS ALSO ALLEGING VISUALIZATION OF PARTICLES IN THE DEVICE AND THE FILTERS ARE DIRTY. THE MANUFACTURER CONTACT INFORMATION HAS ALSO BEEN UPDATED. THE PREVIOUS REPORT WAS ALSO FILED AS A FOLLOW UP/FINAL, HOWEVER THE MANUFACTURER'S INVESTIGATION IS NOW ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. PLEASE SEE SECTION G1 FOR UPDATED INFORMATION. | Y | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02611 | 2021-07-30 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | AUX500T15C | AUX500T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. . THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. SECTION H6 CODES CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT HAS BEEN CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02187 | 2021-07-30 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX700T11 | DSX700T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102894 | 2021-07-30 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. | DSX500T11C |  |  | I USE A PHILIPS DREAM STATION CPAP, MODEL DSX500T11C, SERIAL NO. (B)(4), WHICH HAS BEEN RECALLED BY PHILIPS. I HAVE REGISTERED THIS UNIT ON THE PHILIPS RECALL WEBSITE. I NOTICED THAT THE AIR FILTER HAD BLACK PARTICLES AND MY PHYSICIAN RECOMMENDED THAT I DISCONTINUE USE OF THE CPAP UNTIL I GET A REPLACEMENT. I HAVE SLEEP APNEA AND NON-USE OF A DEVICE IS AFFECTING MY HEALTH. I HAVE CALLED THE PHILIPS INFORMATION LINE BUT THEY COULD NOT TELL ME WHEN I WOULD BE CONTACTED BY PHILIPS ABOUT A REPLACEMENT OR WHAT I COULD DO IN THE INTERIM. THEY SUGGESTED THAT I CONTACT THE DISTRIBUTOR WHO SUPPLIED THE CPAP FOR HELP, BUT THEY REFERRED ME BACK TO PHILIPS. I AM VERY WORRIED ABOUT NOT USING THE DEVICE FOR SLEEP APNEA AS WELL AS WHAT MAY HAVE HAPPENED TO ME WITH PARTICLES IN MY BODY BY USING THE CPAP FOR THE PAST FEW YEARS. I NOTICE A HOARSENESS IN MY VOICE AS WELL AS JAW PAIN NEAR MY LEFT EAR. I DON'T KNOW IF THIS IS RELATED AND I HAVE ASKED MY DOCTOR FOR AN OPINION ON HOW TO RECTIFY THOSE ISSUES. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-07
2518422-2021-02601 | 2021-07-30 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP W/HUMID | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-07
MW5102912 | 2021-07-30 | 2021-02-26 | I, n, j, u, r, y | Voluntary report | RESPIRONICS DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. | DSX500 |  |  | DIAGNOSIS OF RECURRENCE OF LUNG CANCER FOLLOWING RADIATION TREATMENT. MALIGNANT NEOPLASM OF BRONCHUS AND LUNG; HISTORY OF TRACHEAL CARCINOMA INTERPRETATION: TRACHEAL NODULE, BIOPSY: INVASIVE CARCINOMA, MORPHOLOGICALLY CONSISTENT WITH PREVIOUSLY DIAGNOSED LOW GRADE MUCOEPIDERMOID CARCINOMA. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102897 | 2021-07-30 |  | I, n, j, u, r, y | Voluntary report | CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. | DREAMSTATION | DSX500T11C |  | NUMEROUS TIMES, WHILE USING MY PHILIPS DREAMWEAVER CPAP, I NOTICED SORES DEVELOPING INSIDE MY NOSTRILS. WHEN TRYING TO FIND OUT IF IT WAS COMMON FOR CPAP USERS, I COULDN'T FIND A LINK. ADDITIONALLY, WITH NO OTHER KNOWN CATALYST, I HAVE HAD FREQUENT CHEST CONGESTION AND IRRITATION. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102911 | 2021-07-30 | 2021-07-18 | I, n, j, u, r, y | Voluntary report | CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS INC. |  |  |  | BLADDER INFECTION BECAUSE OF KLEBSIELLA BACTERIA. HEADACHES SINCE (B)(6) 2021. BODY SWELLING. DISCONTINUED CPAP USE AFTER CALLING SUPPLIER AND LEARNED OF RECALL. DID NOT USE ANY CLEANING DEVISE ON MY CPAP. WAS GIVEN MACROBID ANTIBIOTIC FOR BLADDER TREATMENT WAS GIVEN ON (B)(6) 2021. RECEIVED A CALL ON (B)(6) 2021 WHILE I WAS AWAY FROM HOME STATING THAT I NEEDED A DIFFERENT ANTIBIOTIC BECAUSE OF DIFFERENT TYPE OF BACTERIA FOUND IN FURTHER TESTING. TESTING ON (B)(6) 2021 WAS DONE AT AN EXPRESS CARE FACILITY. I STOPPED USING THE CPAP AROUND (B)(6) 2021. RIGHT AWAY I NOTICED MORE CONGESTION AND COUGHING UP MUCOUS. I STARTED USING THE CPAP (B)(6) 2021. I IMMEDIATELY FELT BETTER RESTED AND THOUGHT IT WAS THE BEST THING EVER. DO NOT HAVE ANY FIGURES OF TESTS. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-07
MW5102899 | 2021-07-30 | 2020-01-07 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I HAVE SEVERE SLEEP APNEA. I WAS GIVEN A NEW PHILIPS RESPIRONICS DREAMSTATION FOR SLEEP APNEA ON (B)(6) 2020 AND I HAVE BEEN USING THIS MACHINE SINCE THEN. I HAD OTHER DREAMSTATIONS IN THE PAST AND WAS GIVEN THIS ONE WHEN THERE WAS AN UNSOLVABLE TECHNICAL PROBLEM WITH MY OLD ONE. THIS NEW ONE, RIGHT FROM THE START, NEEDED TO HAVE DISPOSABLE FILTERS REPLACED VERY OFTEN. THE WHITE FILTERS QUICKLY BECAME EXTREMELY DIRTY, UNLIKE FILTERS ON PAST MACHINES. I ALSO NOTICED SOME LUNG IRRITATION FROM THAT TIME, BUT BECAUSE WE WERE GOING INTO THE PANDEMIC AND I DO HAVE SOME ALLERGIES, I DIDN'T DO ANYTHING ABOUT IT. RECENTLY, I HAD BEEN NOTICING SOME DEBRIS IN THE BOTTOM OF THE WATER RESERVOIR AFTER EACH NIGHT'S USE. I AM RIGOROUS IN CLEANING THE MACHINE - WASHING THE HEADGEAR, NOSE PILLOWS, AND WATER RESERVOIR EVERY MORNING WITH DAWN DISHWASHING DETERGENT, THE TUBING EVERY 3 DAYS, AND USING ONLY DISTILLED WATER IN THE RESERVOIR. I NEVER RECEIVED A RECALL NOTICE FROM PHILIPS AND ONLY LEARNED ABOUT THE RECALL WHEN I ACTUALLY TALKED WITH A REP AT MY MEDICAL SUPPLY COMPANY TODAY AND MENTIONED THE CONTINUAL DIRTY FILTERS. MY CHECK OF THE SERIAL NUMBER SAYS THAT MY MACHINE IS AFFECTED BY THE RECALL. IT IS RECOMMENDED THAT I STOP USE IMMEDIATE BECAUSE OF POTENTIAL TOXIC PARTICLES FROM A SOUND ABATEMENT FOAM. SINCE LAST SPRING, I HAVE ALSO HAD CHEST PAIN ON THE LEFT, FOR WHICH I HAVE HAD 2 MAMMOGRAMS AND ULTRASOUND TESTS AS WELL AS PHYSICAL EXAMS OF THE AREA AND AN MRI OF THE THORACIC SPINE, WHICH HAVE NOT SHOWN ANY REASON FOR THE PAIN I HAVE BEEN HAVING SINCE SPRING 2020. I DON'T KNOW IF THIS COULD BE RELATED, BUT IT COULD BE. FDA SAFETY REPORT ID # (B)(\4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102909 | 2021-07-30 | 2021-07-21 | I, n, j, u, r, y | Voluntary report | DREAM STATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | SO-CLEAN AND PHILIPS RESPIRONICS. SO-CLEAN ADVERTISES IT IS SAFE WITH THE CPAP. RECEIVED LETTER FROM PHILIPS TODAY (7/28/2021) THAT A RECALL IS OCCURRING. I HAVE BEEN SUFFERING WITH SORE THROAT, HEADACHE, CONGESTION, COUGH, FATIGUE EXACERBATED OVER THE LAST WEEK. INTERMITTENTLY OVER THE PAST YEAR OR YEAR AND A HALF I HAVE HAD THESE SYMPTOMS EXACERBATE AND DISSIPATE. ALWAYS WAKING UP WITH AT LEAST SORE THROAT. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-07
MW5102902 | 2021-07-30 |  | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I HAVE SLEEP APNEA AND AM THE OWNER/USER OF A RECALLED PHILIPS DREAM STATION BREATHING THERAPY DEVICE. I DID NOT RECEIVE ANY NOTIFICATION FROM THE DEVICE PROVIDER, (B)(6), OR THE DEVICE MANUFACTURER, PHILIPS, THAT THE DEVICE WAS ON A RECALL LIST. I FIRST LEARNED OF THE PHILIPS RECALL IN (B)(6) 2021 UPON READING AN ONLINE ARTICLE. I IMMEDIATELY CONTACTED (B)(6)CUSTOMER SERVICE TO ASCERTAIN WHETHER MY DEVICE WAS ON THE RECALL LIST, AND TO INQUIRE AS TO WHAT THE APPROPRIATE FOLLOW UP WOULD BE. I WAS INITIALLY TOLD BY THEM THEY WERE UNAWARE OF ANY RECALL OF PHILIPS CPAP DEVICES. THEY SUBSEQUENTLY CONFIRMED THERE WAS A RECALL BUT COULDN'T TELL ME WHETHER MY DEVICE WAS ON THE LIST OR WHAT TO DO. THEY ADVISED ME TO CONTACT PHILIPS. ON (B)(6)THROUGH ADDITIONAL RESEARCH ON THE PHILIPS WEBSITE I WAS ABLE TO CONFIRM THAT MY DEVICE WAS LIKELY ON THE RECALL LIST. WHILE ON THEIR WEBSITE, I TRIED UNSUCCESSFULLY TO REGISTER MY DEVICE FOR THE RECALL THROUGH THE PHILIPS ONLINE RECALL REGISTRATION PORTAL. WHEN THAT FAILED I CALLED THE PHILIPS RECALL HELP NUMBER. SPOKE WITH (B)(6). SHE REGISTERED MY UNIT OVER TELEPHONE, AND PROVIDED ME A CONFIRMATION # (B)(4). SHE SAID THE UNIT WOULD BE REPLACED. SHE SAID IF I DIDN'T HEAR FROM THEM IN TWO WEEKS, CALL BACK TO DETERMINE STATUS. NOT HEARING FROM PHILIPS IN THE TWO WEEK TIME FRAME, I CALLED THEM TODAY, (B)(6) 2021. SPOKE WITH (B)(6). (B)(6) COULD OFFER ME NO ASSISTANCE, NO INFORMATION, OR PROVIDE ME A TIMELINE WHEN I MIGHT BE HEARING FROM PHILIPS AGAIN, OR WHEN MY DEVICE MIGHT BE REPLACED. HE SAID THEY WERE STILL AWAITING REGULATORY APPROVAL FROM SOME AGENCY, BUT DIDN'T SPECIFY THE AGENCY OR WHAT SPECIFICALLY THEY WERE AWAITING APPROVAL FOR. I'M FINDING ALL THIS TO PRETTY FRUSTRATING. THE VENDOR IS BEING LESS THAN HELPFUL. AS PATIENTS, WE ARE PRETTY MUCH STRANDED, LEFT TWISTING SLOWLY IN THE WIND. EVEN GETTING AN APPOINTMENT WITH A PULMONOLOGIST TO EXPLORE ALTERNATIVES HAS PROVED CHALLENGING. WHILE I UNDERSTAND THE FDA CAN'T OFFER MUCH HELP WITH THE PULMONOLOGIST APPOINTMENT, I AM ASKING FOR ASSISTANCE IN DEALING WITH PHILIPS, AND HOPEFULLY GETTING A RESOLUTION TO THE DEVICE REPLACEMENT ISSUE. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | I |  | 2021-07
MW5102927 | 2021-08-02 | 2020-10-05 | I, n, j, u, r, y | Voluntary report | PHILLIPS RESTMORE CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS INC. | U/K | U/K |  | ON (B)(6) 2020; I HAVE A PROBLEM WITH MY CPAP MACHINE. I WAS EITHER WET IN THE MORNING OR TANK RAN DRY. I SPOKE TO RESPIRATORY AND SHE SAID IT IS OLDER BUT LET'S REPLACE TANK AND ALL TUBING. I NOTICED I COULDN'T GET MY CARD OUT AND USED NEEDLE NOSE PLIERS. I DIDN'T NOTICE A DIFFERENCE AND EVER MORNING WOKE UP WITH A BARKING COUGH, WET AND TANK EMPTY. THEY SAID THERE WAS NOT ANYTHING THEY COULD DO BECAUSE I WAS NOT DUE UNTIL (B)(6) 2021. SO I USED IT, THROAT AND COUGH BECAME CHRONIC AND WAS WET. I HAD BEEN HOSPITALIZED SEVERAL TIMES DUE TO PULMONARY DISTRESS. I PUT TUBE SOCKS ON TUBING WHICH HELPED. I WAS NOT SLEEPING WELL. IN (B)(6), I GOT A NEW CARD AND WAS ABLE TO GET IT IN AND OUT. PRIOR TO THAT IT WAS STUCK IN MACHINE. THEN JUST AS I THOUGHT I COULD MANAGE 6 MONTHS, I GOT RECALL NOTICE. I WAS ADVISED NOT TO USE IT AS I HAVE HAD CANCER 3 YEARS AGO AND MY LUPUS HAD FLARED. AS I WAS PACKING IT UP, I WAS STUNNED TO SEE BLACK MOLD. I CLEAN MY MACHINE WITH VINEGAR AND CHANGE FILTER BUT NEVER LOOKED AT THIS PART OF MACHINE. IT SURE ANSWERED THE QUESTIONS OF WHY HAVE I DEVELOPED ASTHMA, REACTIVE AIRWAY DISEASE AND CHRONIC SORE THROAT. I COUGH FOR 30 MINUTES WHEN I RISE IN AM OFTEN STARTING DAY WITH NEBULIZER. THE CUSTOMER SERVICE AT PHILIPS STATED BETWEEN CARCINOGENS AND BLACK MOLD SHE UNDERSTOOD WHY BREATHING BECAME AN ISSUE. SHE STATED PUT GLOVES ON AND A MASK AND PUT MACHINE IN BAG. WHEN THEY KNEW WHAT THEY WERE GOING TO BE DOING WITH FIXING OR REPLACING THE RECALLED MACHINES WHICH MINE WAS ONE OF I WOULD GET A LETTER. I HAVE HAD AN EXTREMELY DIFFICULT TIME FOR THE LAST MONTH WITH SLEEP AND I ALSO HAVE CARDIAC ISSUES, HBP AND MODERATE SEVERE SLEEP APNEA AND I HAVE NOT BEEN ABLE TO SLEEP. THE ISSUE THAT I AM BRINGING FORWARD IS MY MACHINE HAS BLACK MOLD GROWING IN IT AND THAT HAS GOT NOTHING TO DO WITH THE RECALL AND BECAUSE OF BLACK MOLD I SHOULD HAVE IT REPLACED WHETHER IT'S 2 YEARS 3 YEARS OR 5 YEARS NOT WHAT THEY'RE TELLING ME. THERE IS NOTHING I CAN DO, BUT PHILLIP'S DISAGREES BECAUSE THE BLACK MOLD IS AS SERIOUS AS OR EVEN MORE SERIOUS THAN CARCINOGENS. PHILIPS INFORMED ME THAT MY MACHINE SHOULD BE REPLACED BECAUSE OF MOLD ISSUE. THEY DON'T EXPECT ANYTHING TO HAPPEN WITH REPLACEMENT OF THE MACHINE BECAUSE THEY PUT ME ON HIGH PRIORITY DUE TO THE BLACK MOLD UNTIL 6 TO 8 WEEKS FROM NOW. SO THERE ARE TWO ISSUES ONE IS IT IS RECALLED WHICH I UNDERSTAND YOU HAVE TO GO THROUGH THAT PROCESS. HOWEVER MY MACHINE HAS BLACK MOLD ON IT AND I SHOULD NOT HAVE TO GO THROUGH THE RECALL PROCESS AND SHOULD IMMEDIATELY BEEN PROVIDED WE WITH A NEW MACHINE BECAUSE SHE TOLD ME IT WAS VERY DANGEROUS TO BE BREATHING BLACK MOLD. DO MY LUNGS EVER HEAL? WHAT PERMANENT DAMAGE HAS THE BLACK MOLD CAUSED. I USE SYMBICORT AND 2 DIFFERENT NEBULIZERS. IT DESTROYED MY TEETH; 7 MONTHS AGO MY TEETH WERE GREAT. I PRACTICE SUPER DENTAL HYGIENE. THAT IS GONE FOR EVER. A TOOTH BROKE AND I HAD 16 CAVITIES. I AM (B)(6) AND NEVER HAD ONE, SO THAT SHOULD SAY SOMETHING. THE TOOTH THAT BROKE IS MY FRONT TOOTH. TALK ABOUT BEING GRATEFUL FOR MASKS.. BETWEEN BREATHING BLACK MOLD IN EVERY NIGHT, LOSING MY TEETH, TAKING 3 MEDICATIONS HAVE PULMONARY DISTRESS AND WAS TAKEN TO HOSPITAL AND BP UNCONTROLLED MY CPAP HAS CAUSED NUMEROUS MEDICAL AND PREVENTABLE CONDITIONS. IT ALL FALLS TOGETHER ON A TIME LINE. I MUST HAVE ACCIDENTALLY CLICKED IT. I AM SO SORRY. MIND YOU MY BREATHING IS SO POOR, I MIGHT AS WELL BE. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02672 | 2021-08-02 | 2021-07-29 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP NASOPHARYNGEAL CANCER CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02671 | 2021-08-02 | 2021-07-29 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, INTL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | INX500S15 | INX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP NODULES IN THEIR LUNGS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02652 | 2021-08-02 | 2021-07-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | FRX500S14 | FRX500S14 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02670 | 2021-08-02 | 2021-07-28 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO, FR | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | FRX500S14 | FRX500S14 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED THAT A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM ALLEGEDLY BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP ASTHMA. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE DEVICE WAS RETURNED TO THE MANUFACTURER QUALITY PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. DURING THE INITIAL VISUAL INSPECTION OF THE DEVICE FOUND WHITE COLORED TALC-LIKE CONTAMINANT ON THE NON-RESPIRONIC FILTER, IN THE AIR INLET PORT AND IN THE AIR-PATH; AND NO PROOF OF ODORS ON THE EXTERNAL ENCLOSURE OR DURING THE INTERNAL INSPECTION. THE FINDING OF CONTAMINANTS IN THE AIR INLET PORTS INDICATES THE SOURCE OF THE CONTAMINANTS IS EXTERNAL TO THE DEVICES AND IS NOT RELATED TO A FAILURE OF THE DEVICES. THE DEVICE WAS DISASSEMBLED FOR INTERNAL VISUAL INSPECTION. THE MANUFACTURER FOUND EVIDENCE OF WHITE COLORED TALC-LIKE CONTAMINANT ON THE INSIDE OF THE DEVICE AND IN THE DEVICES AIR-PATH. THE WHITE COLORED CONTAMINATE WAS ON THE OUTER SIDE OF THE BLOWER MOTOR/SOUND ABATEMENT FOAM HOUSING AND ON THE INSIDE PORTION OF THE UI PANEL AND ON THE INSIDE PORTION OF THE REAR PANEL AND INSIDE PORTION OF THE BOTTOM ENCLOSURE ALONG WITH THE WHITE POWDER IN THE BLOWER MOTOR AND BLOWER BOX. THERE IS NO EVIDENCE OF THE SOUND ABATEMENT FOAM DEGRADATION. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND FOUND THE POWER SUPPLY WOULD INTERMITTENTLY CUT OFF POWER TO THE DEVICE. A KNOWN GOOD POWER SUPPLY WAS APPLIED TO THE DEVICE AND THE DEVICE POWERED ON. THE MANUFACTURER REVIEWED THE DEVICE'S EVENT LOGS AND FOUND THERE WERE NO ERROR CODES. THE MANUFACTURER CONCLUDES THERE IS NO EVIDENCE OF FOAM DEGRADATION WITHIN THE DEVICE. THE CONTAMINATION FOUND WITHIN THE DEVICE IS CONSISTENT WITH BEING FROM AN EXTERNAL SOURCE. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02612 | 2021-08-02 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | CAX400T12 | CAX400T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02666 | 2021-08-02 | 2021-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERFACE BOARD WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR WAS RETURNED FOR ROUTINE PREVENTATIVE MAINTENANCE. DURING THE EVALUATION A SERVICE REQUIRED CODE WAS OBSERVED IN THE DOWNLOADED EVENT LOG AND THE INTERFACE BOARD WAS REPLACED TO ADDRESS THE ISSUE. DURING ADDITIONAL TESTING, A SERVICE REQUIRED CODE WAS OBSERVED IN THE EVENT LOG AND THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02654 | 2021-08-02 | 2021-01-01 | I, n, j, u, r, y | Manufacturer report | RENTAL DREAMSTATION EXPERT, FR | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | RFRX501S14 | RFRX501S14 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP PANCREATIC CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED OF PANCREATIC CANCER. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED AND FOUND EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED. THE EXTERNAL INSPECTION OF THE DEVICES FOUND NO EVIDENCE OF THERMAL OR PHYSICAL DAMAGE, THE INTERNAL INSPECTION OF THE DREAMSTATION FOUND EVIDENCE OF INSECTS IN BETWEEN THE EDGES OF THE TOP AND BOTTOM ENCLOSURES ALONG WITH THE DIRT/DUST IN THE BLOWER MOTOR. FINDING THE DIRT/DUST CONTAMINATES IN THE BLOWER MOTOR IS EVIDENCE THAT THE CONTAMINATES WERE IN THE AIR PATH OF THE DEVICE AND WAS A RESULT OF EXTERNAL CONTAMINATION. IT WAS OBSERVED NO EVIDENCE OF THE SOUND ABATEMENT FOAM DEGRADATION. THE LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02679 | 2021-08-02 | 2021-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02678 | 2021-08-02 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02669 | 2021-08-02 | 2021-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 02 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | I |  | 2021-08
MW5102900 | 2021-08-02 | 2021-05-17 | D, e, a, t, h | Voluntary report | PHILIPS RESPIRONICS TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | PHILIPS RESPIRONICS / RESPIRONICS, INC. |  |  |  | PATIENT WAS USING A PHILIPS RESPIRONICS TRILOGY 100 VENT CONTINUOUSLY. PATIENT HAD A SUDDEN CARDIAC ON-SET EVENT THAT RESULTED IN HOSPITALIZATION SEVERAL DAYS BEFORE PATIENTS DEATH. FAMILY LEARNED OF RECALL AND IS QUESTIONING A CORRELATION BETWEEN PATIENT PREVIOUS SYMPTOMS AND SYMPTOMS LISTED IN RELATION TO THE RECALL. SPECIFICALLY FROM OFF-GASSING AND NOT PARTICULATES. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5102931 | 2021-08-02 | 2014-01-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500 |  |  | I HAVE USED RESPIRONICS SYSTEM ONE AND THEN THE DREAMSTATION FROM 2013 UNTIL YESTERDAY. I HAVE BEEN COUGHING UP BLACK PARTICLES FOR YEARS. I EXHIBIT SYMPTOMS LISTED IN THE RECALL NOTIFICATION. TESTS ARE NEEDED. PLEASE ARRANGE. THANK YOU. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5102937 | 2021-08-02 | 2020-01-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | BREATHING PROBLEMS AT NIGHT; I HAVE SLEEP APNEA AND ASTHMA HAVE BEEN USING PHILIPS CPAP MACHINES SINCE 2009 IN 2019 GOT A NEW MACHINE AND AFTER A SOME TIME STARTED HAVING SEVERE HEADACHES DAILY, NOSE BLEEDS, NAUSEA AND NOTICED I WASN'T ABLE TO BREATHE AS WELL WITH THIS MACHINE, FREQUENTLY WAKING UP COUGHING. THERE WAS A PROBLEM WITH THE HUMIDIFIER AND SO I THOUGHT PERHAPS THAT WAS THE ISSUE. I WAS ABLE TO GET THE HUMIDIFIER REPLACED BUT THERE WASN'T MUCH DIFFERENCE IN FACT I HAD TO KEEP THE HUMIDITY SET LOW BECAUSE IT'S DIFFICULT TO BREATHE WHEN IT IS SET HIGHER THE LAST 7 MONTHS OR SO I'VE BEEN WAKING UP WITH HEADACHES LASTING THROUGHOUT THE DAY, NOSE BLEEDS, DIZZINESS AND STOMACH ISSUES ON AND OFF . IN (B)(6) 2020 I STARTED HAVING ABDOMINAL ISSUES, IN (B)(6) 2020 I HAD EMERGENCY GALL BLADDER SURGERY. I USED MY PERSONAL CPAP IN THE HOSPITAL AND THE HEADACHES, NASAL, THROAT IRRITATION, COUGHING DURING THE NIGHT CONTINUED THROUGH THE ENTIRE TIME INCLUDING WHILE I WAS IN THE HOSPITAL AND CONTINUE TILL TODAY. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5102938 | 2021-08-02 | 2021-07-23 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS BI-LEVEL CPAP DEVICE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | PHILIPS CPAP DEVICE HAS LEAD PARTICLES. RECEIVED LETTER FROM (B)(6) ADVISING ME TO STOP USING MY CPAP MACHINE. HE RECOMMEND ME ORDERING A NEW MACHINE FROM A DIFFERENT DISTRIBUTOR. THIS MACHINE IS CRITICAL FOR ME AND MY CURRENT HEALTH. UNFORTUNATELY MEDICARE WILL NOT ALLOW SINCE THEY ONLY BUY 1 MACHINE IN A LIFETIME. I AM TRYING TO GET A FILTER, BUT VERY SCARED TO USE SINCE THE ORIGINAL MACHINE IS UNSAFE AND FDA APPROVED. THE FILTERS ARE ON MASSIVE BACK ORDER AND NO ONE CAN TELL YOU WHEN YOU WILL RECEIVE THE FILTER. WILL THIS FILTER BE FDA APPROVED. WHY ARE WE HAVING RESTRICTIONS WITH (B)(6), THIS IS AND WAS OUT OF THE CONSUMERS CONTROL. I FIND THIS UNETHICAL FOR (B)(6) TO RISK OUR LIVES WITH SO LITTLE REGARD. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
MW5102940 | 2021-08-02 | 2021-03-01 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | PHILIPS / RESPIRONICS, INC. |  |  |  | I HAVE BEEN USING THIS PHILIPS RESPIRONICS TRILOGY 100 BIPAP SINCE (B)(6) 2020, AND HAVE EVERY NIGHT FOR ABOUT 5-8HRS. SOME TIME DURING THE FIRST QUARTER OF THIS YEAR 2021, I STARTED HAVING HEADACHES EVERY MORNING, WHICH I WOULD NEVER GET, BUT AS A RESULT STARTED TAKING TYLENOL IN THE MORNING QUITE OFTEN. ALSO, WHEN I WOKE UP, I FELT INEXPLICABLY NAUSEOUS TOO. MY THROAT WAS SORE AND IRRITATED, AND I HAD STUFFY SINUSES, WITH A DRY COUGH. FOR SEVERAL MONTHS I HAD TO USE MY ASTHMA PUFFERS MORE OFTEN TOO BECAUSE MY CHEST FELT TIGHT AND FOR SOME UNKNOWN REASON I HAD MORE ASTHMA. OVER THE LAST MONTH OR SO, THE UNIT SEEMED TO HAVE A SMELL EMANATING FROM IT WHEN I WOULD WEAR MY FACE MASK. I AM EXTREMELY UPSET TO HEAR THAT I HAVE BEEN EXPOSED TO THIS DEGRADED AND/OR OFF-GASSED PE-PUR FOAM. I HAVE BEEN WROUGHT WITH FEAR AND WORRY ABOUT THE OVERALL CARCINOGENIC EFFECTS THAT HAVE ALREADY OCCURRED. PHILIPS DID NOT NOTIFY ME AT ALL PERSONALLY. I AM NOT GOING TO USE THIS DEVICE ANY MORE. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5102941 | 2021-08-02 | 2021-06-23 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I HAVE HAD MULTIPLE SINUS INFECTIONS RECENTLY FROM UNKNOWN ORIGINS. I WAS BEGINNING TO THINK THAT MY CPAP MAY BE THE CULPRIT, BUT I. DEPEND ON IT TO SLEEP AT NIGHT SO I HAVE CONTINUED TO UTILIZE IT. UPON LEARNING OF THE RECALL ON MY MACHINE, I BELIEVE THIS MAY BE THE CAUSE. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | I |  | 2021-08
MW5102945 | 2021-08-02 | 2019-02-15 | I, n, j, u, r, y | Voluntary report | DREAMSTATION BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX700H11C |  |  | PHILLIPS RESPIRONICS RECALLED MY BIPAP MACHINE. I STARTED USING THE BIPAP IN HOSPITAL AFTER A RUPTURED BRAIN ANEURYSM AND SUBARACHNOID HEMORRHAGE STROKE. WHEN RELEASED I HAD A SLEEP APNEA TEST AND THE DOCTOR ORDERED A DREAMSTATION BIPAP I USE WITH SUPPLEMENTAL OXYGEN CONCENTRATOR. IN 2020 I WAS DIAGNOSED WITH LIVER PROBLEM, I HAD FOLLOW UP CT SCAN IN (B)(6) 2021 AND A .9MM NODULE WAS FOUND ON THE LEFT LOBE OF MY THYROID GLAND. A THYROID SONOGRAM IN (B)(6) 2021 SHOWED NODULE IS NOW 1.3MM AND THERE IS A 6MM MASS ON MY RIGHT LOBE OF THE THYROID GLAND. NODULE ISN'T LARGE ENOUGH TO DO A FINE NEEDLE BIOPSY YET, RECHECK IN 6 MONTHS. A CT OF A CLAVICLE LUMP WAS DONE AND THEY FOUND NOTHING ON THE LUMP BUT NOTED THE NODULE ON THE LEFT LOBE AS A COMPLEX CYSTIC NODULE, THE MASS ON MY RIGHT LOBE AS A CYST. THEY NOTED LEFT MAXILLARY SINUS DISEASE AND THEY SAW A GLIMPSE OF MEDIASTINAL LYMPH NODES. I HAVE BLURRY VISION. I WENT TO OPHTHALMOLOGIST AND GOT NEW PRESCRIPTION. I HAVE HAD WET MACULAR DEGENERATION IN MY LEFT EYE SINCE 2005 AND HAVE NO VISION IN LEFT EYE. I GO BACK TO A RETINA SPECIALIST IN 6 MONTHS TO SEE IF I HAVE A RETINA PROBLEM. CURRENTLY DOING PHYSICAL THERAPY FOR LOSS OF BALANCE, DIZZINESS AND FALL RISK. HAVE RECORDS OF ALL TESTS FROM (B)(6) 2019 THROUGH TODAY AVAILABLE. MORE TO COME. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02661 | 2021-08-02 | 2021-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVATION EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02476 | 2021-08-02 | 2021-06-28 | I, n, j, u, r, y | Manufacturer report | TRILOGY | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED NAUSEA, COUGH, DRY THROAT AND SEIZURES. THERE WAS NO REPORT OF THE PATIENT RECEIVING MEDICAL INTERVENTION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A VENTILATOR SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED NAUSEA, COUGH, DRY THROAT AND SEIZURES. THERE WAS NO REPORT OF THE PATIENT RECEIVING MEDICAL INTERVENTION. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y |  |  | 2021-08
2518422-2021-02648 | 2021-08-02 | 2021-07-29 | I, n, j, u, r, y | Manufacturer report | BIPAP AUTO, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | 760P | 760P |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG CANCER. THE PATIENT REQUIRED SURGICAL INTERVENTION IN RESPONSE TO THE REPORTED EVENT . THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02655 | 2021-08-02 | 2021-07-13 | I, n, j, u, r, y | Manufacturer report | BIPAP S/T C SERIES, INTERNATIONAL | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS INC | IN1061S | IN1061S |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP NASAL CONGESTION, FACE IRRITATION, AND NEOPLASMS. THE PATIENT WILL BE RECEIVING RADIOTHERAPY IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02610 | 2021-08-02 | 2020-04-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP DIFFICULTY BREATHING CAUSED BY LUNG INFLAMMATION. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION FOR ANTIBIOTICS, STEROIDS AND AN INHALER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON OCTOBER 09, 2023, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGED OF EYE IRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/ OR HEADACHE, ASTHMA (NEW OR WORSENING), INFLAMMATORY RESPONSE, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, OTHER, CONGESTIVE HEART FAILURE. THE SECTION E1 WAS UPDATED AND SECTIONS B1, B2, H1 AND H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR BOTH ADVERSE EVENTS AND PRODUCT PROBLEM (ONLY THE PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECTS - CLINICAL CODES AND HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02702 | 2021-08-03 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP HEADACHES, DIZZINESS, SHORTNESS OF BREATH AND BLOOD CLOTS IN THE KNEE AND ANKLE. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A DOCTOR'S ASSESSMENT AND COMPUTED TOMOGRAPHY (CT) TESTING. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE SOUND ABATEMENT FOAM. PATIENT ALLEGED TO TO DEVELOP HEADACHES, DIZZINESS, SHORTNESS OF BREATH AND BLOOD CLOTS IN THE KNEE AND ANKLE. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A DOCTOR'S ASSESSMENT AND COMPUTED TOMOGRAPHY (CT) TESTING. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURER CONCLUDES NO FURTHER ACTION IS NECESSARY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED IN THE BASE UNIT. THE EXTERNAL INVESTIGATION SHOWED THAT THERE WERE NO SIGNS OF CONTAMINATION ON THE OUTSIDE OF THE DEVICE. THE DEVICE WAS HOOKED UP TO A KNOWN GOOD POWER SUPPLY AND POWER CORD, AIRFLOW WAS VERIFIED TO BE OPERATING CORRECTLY. THE INTERNAL INVESTIGATION SHOWED THAT THERE WERE NO SIGNS OF CONTAMINATION OR SOUND ABATEMENT FOAM DEGRADATION INSIDE OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE INVESTIGATOR CONFIRMED THERE WAS NO SIGNS OF CONTAMINATION OR SOUND ABATEMENT FOAM DEGRADATION INSIDE OF THE DEVICE. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02688 | 2021-08-03 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON JULY 01, 2024, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICESTHE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED LUNG CANCER AND HAVING SHORTNESS OF BREATHE. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. SECTIONS B1, B2, H1 AND H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR ONLY THE PRODUCT PROBLEM (BOTH ADVERSE EVENTS AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02492 | 2021-08-03 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, GER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DEX500S13 | DEX500S13 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02689 | 2021-08-03 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/HT/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX400T11C | DSX400T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN PREVIOUS REPORT SECTION B5 WAS MENTIONED INCORRECTLY, CORRECT B5 SHOULD BE THE PATIENT ALLEGED VISUALIZATION OF PARTICLES, CHEST PAINS, HEADACHES, DIZZINESS, LIGHTHEADEDNESS, NAUSEA, DIFFICULTY BREATHING/SHORTNESS OF BREATH. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL AND INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND WHITE UNKNOWN DUST LIKE CONTAMINATION AND SOME VERY SMALL POTENTIAL HAIRS AT THE AIR INPUT OF THE BLOWER BOX, SMALL BLACK UNKNOWN CONTAMINATION ON THE BOTTOM OF THE BLOWER BOX. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS. THE MANUFACTURER CONCLUDES MULTIPLE CONTAMINATION FOUND WERE CONSISTENT WITH CONTAMINATION OF UNKNOWN WHITE DUST AND HAIRS AT THE AIR INPUT OF THE BLOWER BOX, BLACK UNKNOWN CONTAMINATION ON THE BOTTOM OF THE BLOWER BOX. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. IN THIS REPORT, SECTION D9, G3, H3, H6 HAS BEEN UPDATED OR CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02639 | 2021-08-03 | 2021-07-08 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION BIPAP AUTOSV H/C, DS | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS,INC | DSX900H11C | DSX900H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG NODULES. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THE PATIENT HAS HAD THREE DEVICES, BUT THERE IS NO ADDITIONAL INFORMATION OR SERIAL NUMBERS OF THE OTHER DEVICES. THE MANUFACTURER HAS REQUESTED FURTHER INFORMATION ON THE DEVICES; HOWEVER, THE CUSTOMER HAS REFUSED TO PROVIDE THAT INFORMATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGED THE PATIENT HAS BEEN IN AND OUT OF THE HOSPITAL FOR YEARS. PATIENT FEELS THEY MAY HAVE CANCER. THE PATIENT CANNOT BREATHE AND FEELS LIKE SOMETHING IS IN HIS LUNGS. THE PATIENT LOST THEIR HOME AND THINKS IT'S DUE TO MOLD. THE PATIENT HAS LUNG NODULES THAT ARE GETTING BIGGER. THE PATIENT ALSO REPORTS THAT "FOR MANY YEARS HE HAS HAD DOCTOR APPOINTMENTS, CT SCANS, X-RAYS, ETC. TRYING TO FIGURE OUT WHAT IS WRONG. NO OTHER MEDICAL INTERVENTION WAS REPORTED. THE REPORTED EVENT AND ITS SEVERITY WERE REVIEWED BY THE MANUFACTURER'S CLINICAL EXPERT. BASED ON AVAILABLE INFORMATION THE EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. ALSO, BASED ON THE CONCLUSIONS REACHED THE HARM DESCRIBED IN THIS COMPLAINT IS ASSESSED AS UNRELATED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02590 | 2021-08-03 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02591 | 2021-08-03 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, JP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | RJPX500S16 | RJPX500S16 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING BLACKISH SPUTUM RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. DURING THE EVALUATION, DIRT WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION. IN THIS REPORT, SECTION A1, A3, D8, D9, H2, H3 AND H6 HAS BEEN UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02696 | 2021-08-03 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11C | DSX400H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SARCOIDOSIS AND CHEST PAINS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE /RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM THAT BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE SARCOIDOSIS AND CHEST PAINS, BREATHING PROBLEMS, SORE THROAT, DRY SINUSES, CHOKING AND CHEST PAINS FOR MONTHS BEFORE. THE REPORTED EVENT THAT THE PATIENT WAS DIAGNOSED WITH SARCOIDOSIS; HAS BREATHING PROBLEMS AND CHEST PAIN IS ASSESSED AS A SERIOUS INJURY BUT NOT RELATED TO THE DEVICE. THEREFORE, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. HOWEVER, USE OF THE RECALLED DEVICE COULD POSSIBLY CAUSE OR CONTRIBUTE TO THE REPORTED NON-SERIOUS SYMPTOMS OF HEADACHES, SORE THROAT, DRY SINUS. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. H6 HEALTH EFFECT- CLINICAL CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02701 | 2021-08-03 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA. THE PATIENT WAS GIVEN NEBULIZER TREATMENTS IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING SINUS, SLEEP DYSFUNCTION, ASTHMA, CONGESTION, EYES ITCHING, HEADACHE, COUGH, NASAL/THROAT IRRITATION OR SORENESS, EAR ITCHING, NAUSEA, LIGHTHEADEDNESS, HYPERSENSITIVITY, DRY THROAT AND DIFFICULTY OF BREATHING/SHORTNESS OF BREATH. RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. SECTION H6 (CLINICAL CODE) CORRECTED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02709 | 2021-08-03 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING HEADACHE, NAUSE, VOMITING, PNEUMONIA, BLOOD PRESSURE FLUCTUATION, LOW HEART RATE, SHORTNESS OF BREATH, FATIGUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION A1, A2, A3 AND SECTION H6 HEALTH EFFECT- CLINICAL CODE (MISSED TO CAPTURE IN THE PREVIOUS REPORTS), TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02690 | 2021-08-03 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02677 | 2021-08-03 | 2021-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02659 | 2021-08-03 | 2021-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, A SERVICE REQUIRED CODE RELATED TO "HIGH INTERNAL OXYGEN" WAS OBSERVED. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
MW5102974 | 2021-08-03 |  | I, n, j, u, r, y | Voluntary report | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | CAX500H12C | CAX500H12C |  | BEGAN COUGHING AND SNEEZING AND MUCOUS FROM SINUSES, SLOWLY OVER THE PAST YEAR. I DIDN'T REPORT OR FUSS ABOUT BECAUSE I DIDN'T THINK IT WAS A BIG DEAL. BUT KNOWING THE RECALL REASONS I'M BEGINNING TO THINK THESE PROBLEMS IS CAUSING IT. I WILL DISCUSS WITH MY PHYSICIAN FOR TESTING. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02732 | 2021-08-03 | 2021-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN PREVIOUS REPORT THE B5 SECTION WAS INCORRECT. THE CORRECT B5 SECTION IS, THE PATIENT HAS ALLEGED OF HAVING NASAL/THROAT IRRITATION AND COUGH. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND FOUND EVIDENCE OF CONTAMINATION IN ISO PORT, IN BLOWER BOX AND BLOWER. THE MANUFACTURER FOUND NO EVIDENCE OF NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. THE HUMIDIFIER WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR EVALUATION. THE MANUFACTURER FOUND CONTAMINATION FOUND IN HUMIDIFIER AND HUMIDIFIER TANK SEAL, MINERAL DEPOSITS FOUND IN HUMIDIFIER TANK AS WELL AS ON THE HEATING PLATE. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND NO ERRORS LOGGED. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF CONTAMINATION ON CONTAMINATION IN ISO PORT, IN BLOWER BOX AND BLOWER, IN HUMIDIFIER AND HUMIDIFIER TANK SEAL, MINERAL DEPOSITS FOUND IN HUMIDIFIER TANK AS WELL AS ON THE HEATING PLATE. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. SECTION D9, G3 AND H6 HAS BEEN UPDATED/CORRECTED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02560 | 2021-08-03 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR WAS REPLACED TO ADDRESS. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02675 | 2021-08-03 | 2021-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02562 | 2021-08-03 | 2021-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02561 | 2021-08-03 | 2021-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | FR2110X14B | FR2110X14B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR HAD POWER SOURCE ISSUES. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S INTERNAL BATTERY NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02660 | 2021-08-03 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. THE DETACHABLE BATTERY WAS REPLACED TO ADDRESS THE VENTILATOR INOPERATIVE CONDITION, NOT THE SYSTEM BOARD. DURING ADDITIONAL TESTING, AN ISSUE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02674 | 2021-08-03 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR A LOW INSPIRATORY PRESSURE ALARM CONDITION. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS RETURNED TO THE MANUFACTURER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE FAILED TEST STEPS RELATED TO THE ISSUE. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WILL BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02673 | 2021-08-03 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR A BATTERY DEPLETED CONDITION. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS RETURNED TO THE MANUFACTURER AND FAILED TO CHARGE AND FAILED TO RUN ON AC POWER. THE DEVICE'S BATTERY CONNECTOR WILL BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02681 | 2021-08-03 | 2021-07-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2200X15B | IN2200X15B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE'S TIDAL VOLUME WAS NOT DISPLAYED. THE DEVICE'S EXTERNAL FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02680 | 2021-08-03 | 2021-07-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2200X11B | DS2200X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD, PROPORTIONAL VALVE AND SOLENOID VALVE WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02682 | 2021-08-03 | 2021-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02663 | 2021-08-03 | 2021-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02662 | 2021-08-03 | 2021-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02676 | 2021-08-03 | 2021-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02700 | 2021-08-03 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC. | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP PROSTATE CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP,AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AND CAUSED A PATIENT TO DEVELOP PROSTATE CANCER,LINGERING ELECTRICAL OR CHEMICAL SMELL IN HIS NOSE,SORE THROAT,COPD,SHORTNESS OF BREATH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 (HEALTH EFFECT - CLINICAL CODE) HAS BEEN CORRECTED AND (TYPE OF INVESTIGATION,INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS) HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02697 | 2021-08-03 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA, HEADACHE, SORE THROAT, AND A SPOT IN THEIR LUNG. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING EXACERBATION OF ASTHMA, HEADACHES, THROAT ISSUES, SOMETHING IS IN LUNG AND HOSPITLALIZATION. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HEALTH EFFECT - CLINICAL CODE HAS BEEN CORRECTED, AND TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02710 | 2021-08-03 | 2021-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
MW5102982 | 2021-08-03 | 2021-01-04 | I, n, j, u, r, y | Voluntary report | PHILLIPS RESPIRONICS DREAMSTATION BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | SINCE RECEIVING A NEW BIPAP, I HAVE HAD MULTIPLE ASTHMA EXACERBATIONS REQUIRING ORAL STEROIDS; TWO PNEUMONIAS, ONE REQUIRING HOSPITALIZATION; MULTIPLE CASES OF SINUSITIS; AND LUNG NODULES ATTRIBUTED TO INFLAMMATION. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5102983 | 2021-08-03 | 2020-10-01 | I, n, j, u, r, y | Voluntary report | PHILLIPS RESPIRONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS / RESPIRONICS, INC. |  |  |  | WHITE COATING NOTED IN THE MASK OF MY PHILLIPS RESPIRONICS DREAMSTATION CPAP( SERIAL NUMBER: (B)(4)). THE COATING STUCK TO THEN INSIDE OF THE MASK AND I NEEDED TO SCRUB IT OFF. I HAVE ONLY USED (B)(6) BABY SHAMPOO, AS INSTRUCTED BY (B)(6) HEALTHCARE, TO CLEAN THE MASK AND TUBING. I DID DEVELOP A RED RASH ON MY CHEEKS AFTER THIS OCCURRED. I DON'T KNOW IF I WILL HAVE ANY LASTING OR DELAYED EFFECTS FROM MY EXPOSURE TO THE BURNING/MELTING POLYURETHANE IN MY EQUIPMENT. THIS IS PART OF THE ONGOING NATIONWIDE RECALL. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
MW5102973 | 2021-08-03 | 2019-07-16 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I HAVE USED A CPAP SINCE 2015. ALSO USE SOCLEAN WITH THE SYSTEM. I HAVE HAD UNEXPLAINED RASHES, LUMPS UNDER SKIN AND HOSPITALIZE (B)(6) 2019 FOR SOMETHING IN MY LUNG. SINCE THEN I HAVE A LOT OF MUSCLE LOSS, A COUGH THAT WILL NOT GO AWAY, HEART FLUTTERING, DIZZINESS, EYE TWITCHING ON AND OFF. ALSO MY VISION SEAM TO BE GETTING WORSE. I HAVE HAD TO USE MY READING GLASSES NOT MY PRESCRIPTION GLASSES JUST TO GET COMPUTER SCREEN AND TV. HAVE SOME MEDICAL TEST THAT WILL BE HIGH LIKE FOR MY LIVER FUNCTION AND LOW FUNCTION ON MY ANION GAP FUNCTION. (WHICH I THINK HAS TO DO WITH ENERGY IN THE BODY OR SOMETHING LIKE THAT). I RECEIVED A NEW MACHINE IN (B)(6) 2020 AND WHEN I WAS PUTTING NEW WATER IN CHAMBER FOUND BLACK LIKE PARTICLES IN THE CHAMBER. I CONTACTED BOTH PHILLIPS AND THE SUPPLIES WITH NEITHER HAVING ANY RESPONSE. THE SUPPLIER TOLD ME MY WATER CHAMBER MOST LIKELY HAD MOLD IN IT AND SENT A NEW ONE AND CHARGED ME FOR IT. I TOLD THEM IT WAS NOT MOLD. I EVEN SHOWED THE NEW ONE AND OLD ONE TO SOMEONE AND THEY COULD NOT TELL DIFFERENCE. I ALSO MY COMPLIANCE RATE WAS 97% USAGE BEFORE I FOUND THE PARTICLES. MY RATE NOW IS 0%. IT HAS BEEN ALMOST A MONTH NOW WITH NO RESPONSE FROM PHILLIPS OR THE SUPPLIER. I WENT TO DOCTOR THE OTHER DAY HE WAS NOT SURE WHAT TO DO BUT DID HEAR WHIZZING IN MY BREATHING AND PUT ME ON ANTIBIOTICS TO SEE IF THAT MIGHT HELD BEEN ON THEM FOR 3 DAYS DOES NOT SEAM TO BE WORKING. I HAVE HAD LOT DIFFERENT TEST, BUT DO NOT KNOW WHAT I SHOULD BE TESTED FOR IF I HAVE PLASTIC OR FORM PARTICLES INSIDE OF ME. ANY SUGGESTIONS? I DO NOT HAVE FIRST MACHINE PHILLIPS SYSTEM ONE BUT DO HAVE THE PHILLIPS DREAM STATION. I DONATED THE PHILLIPS SYSTEM ONE AND I NOTIFIED THE HOSPITAL I DONATED IT TO BECAUSE I THINK THAT IS THE ONE THAT I MIGHT OF HAVE SOMETHING IN MY LUNG FROM. BOTH SYSTEMS ARE ON RECALL AND HAVE BEEN REGISTERED WITH PHILLIPS. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5102972 | 2021-08-03 |  | I, n, j, u, r, y | Voluntary report | DREAMSTATION (2 CPAP DEVICES | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS / RESPIRONICS, INC. |  |  |  | ALL THREE OF OUR FAMILY DEVICES BY PHILIPS RESPIRONICS ARE ON THE RECALL LIST -- TWO PHILIPS RESPIRONICS AUTOCPAP DREAMSTATION MACHINES AND OLDER REMSTAR PRO MACHINE THAT WAS USED A WHILE AGO. MY HUSBAND AND I HAVE AND HAD SERIOUS PROBLEMS DURING THE USE OF THE PHILIPS RESPIRONICS MACHINES, INCLUDING SINUS PROBLEMS, SEVERE HEADACHES, INCREASED HEART RATE, CHEST PRESSURE, COUGHING, AND PROBLEMS WITH OUR BREATHING (AND UPPER AIRWAY). MY HUSBAND HAS A HISTORY OF HAVING A STROKE OR TRANSIENT ISCHEMIC ATTACK (TIA) AND ALSO HAS HEART ATRIAL FIBRILLATION (A-FIB). THE TIA LIKELY OCCURRED DURING THE USE THE MACHINES (THE TIA DATE IS UNKNOWN) AND SEVERAL A-FIB PROBLEMS OCCURRED DURING THE USE OF THE DREAMSTATION AND REMSTAR PRO; WE DO NOT KNOW IF THE MACHINE USE HAD ANYTHING TO DO WITH THIS. WHEN WE REGISTERED OUR AUTOCPAP DREAMSTATIONS, WE COULD NOT FIND RECORD OF ANY CONFIRMATION NUMBER IN AN EMAIL, PHONE TEXT, OR PAPER MAIL FORM, BUT WE DID RECEIVE CONFIRMATION OF THE REMSTAR PRO. WE CALLED AND EMAILED PHILIPS RESPIRONICS TO ASK FOR THE CONFIRMATION NUMBERS FOR THE DREAMSTATIONS AND THEY DID NOT PROVIDE THEM. IN THE CALLS AND EMAILS, WE ALSO TRIED TO EXPLAIN OUR HEALTH PROBLEMS FURTHER (BOTH THE PROBLEMS MENTIONED ABOVE AND ALSO OUR HEALTH PROBLEMS THAT OCCUR WITH DISCONTINUATION OF THE DEVICES). DISCONTINUATION OF CPAP MACHINE USE INCREASES MY HUSBAND'S RISK FOR FUTURE A-FIB AND STROKE/TIA PROBLEMS. I HAVE SEVERE SLEEP APNEA, MENTAL/COGNITIVE PROBLEMS, AND PHYSICAL PROBLEMS THAT WORSEN WITHOUT THE USE OF A CPAP MACHINE, ALL TO THE POINT OF LIFE/WORK DISABILITY. WHEN WE CALL, THE REPRESENTATIVES WILL NOT ANSWER OUR QUESTIONS, GIVE US THE CONFIRMATION NUMBERS WE NEED, OR DISCUSS OUR HEALTH CONCERNS. THEY DROPPED CALLS, REQUESTS TO SPEAK TO SUPERVISORS DO NOT WORK (THEY GIVE US A SUPERVISOR PHONE NUMBER WHERE WE LEAVE A MESSAGE AND NOBODY CALLS BACK). EMAILS DO NOT LEAD TO ANY RESPONSE. WE STRUGGLE TO GET ANOTHER CPAP MACHINE AND HAVE TO EITHER DISCONTINUE USE OF THE PHILIPS RESPIRONICS CPAP MACHINES WITH RISK TO OUR HEALTH OR CONTINUE USE OF ONE OF THE RECALLED MACHINES, WHICH ALSO RISKS POSSIBLE CONTINUED OR WORSENED HEALTH PROBLEMS. WE UNDERSTAND THAT PHILIPS RESPIRONICS IS UNDER MUCH STRESS AND STRAIN RIGHT NOW; WE REACH OUT TO THEM ONLY OUT OF HOPE FOR HELP TO GET A UNIQUE DEVICE FROM THEM OR, PREFERABLY, FROM ANOTHER MANUFACTURER. OUR INSURANCE, SLEEP DOCTOR, AND MANY OTHER RESOURCES HAVE TRIED TO HELP US SECURE ANOTHER MACHINE FOR RENTAL; ALL COULD NOT HAPPEN BECAUSE OF THE SHORTAGE OF CPAP MACHINES. BECAUSE WE ARE UNDER MUCH FINANCIAL STAIN AND ARE HAVING HEALTH PROBLEMS, IT WOULD ASSIST IF WE COULD GET REIMBURSEMENT IF WE HAVE TO PURCHASE A CPAP MACHINE OUT OF POCKET. WE DO/DID NOT WANT TO DO OFFICIAL REPORTING BECAUSE WE JUST HOPE TO WORK WITH PHILIPS RESPIRONICS FOR HELP. WE HAVE MEDICAL DOCUMENTATION OF ALL MENTIONED HEALTH PROBLEMS WITH THE START AND END DATES. THE MEDICAL DOCUMENTS ARE MANY WITH A VARIETY OF DATES: TOO MANY TO PROVIDE IN THIS REPORT. IF NECESSARY, WE WILL PROVIDE ALL OF THIS DOCUMENTATION TO VERIFY MY REPORT AND ITS COMMENTS. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | I |  | 2021-08
MW5102971 | 2021-08-03 | 2021-07-27 | I, n, j, u, r, y | Voluntary report | CPAP MACHINE DREAM STATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS / RESPIRONICS, INC. | DSX500H11 |  |  | AFTER YEARS OF A CHRONIC COUGH (BEEN USING CPAP AND SOCLEAN FOR 11 YEARS) I WAS TESTED FOR LUNG CONCERNS. I WAS JUST DIAGNOSED WITH RESTRICTIVE LUNG DISEASE. I HAVE NO CLUE WHAT THE TEST UNIT, LOW RANGE OR HIGH RANGE ARE. THIS IS FROM PATIENT-RECEIVED NOTES FROM EVALUATION. RECORDS AVAILABLE. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5102967 | 2021-08-03 | 2021-07-13 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS / RESPIRONICS, INC. |  |  |  | I RECEIVED A RECALL LETTER FOR MY CPAP MACHINE FROM PHILIPS RESPIRONICS. IT WARNED OF "SERIOUS INJURY WHICH CAN BE LIFE-THREATENING" WITH CONTINUED USE OF MY MACHINE. I HAVE EXPERIENCED THE FOLLOWING SYMPTOMS: HEADACHE, COUGH AND INCREASED NASAL/SINUS CONGESTION. THE LETTER STATES THAT I SHOULD "DISCONTINUE USE OF MY DEVICE AND CONTACT MY DOCTOR OR DME PROVIDER TO DETERMINE OPTIONS FOR CONTINUED TREATMENT". I CANNOT STOP USING MY CPAP BECAUSE I HAVE SEVERE SLEEP APNEA. IF I DID I MY SYMPTOMS WOULD BE UNBEARABLE. I HAVE CONTACTED BOTH MY DOCTOR (WHO SAYS I SHOULD STOP USING MY MACHINE) AND MY DME PROVIDER WHO SAYS I AM CURRENTLY NOT ELIGIBLE FOR A NEW CPAP MACHINE. I HAVE CONTACTED PHILIPS THROUGH THEIR WEBSITE AND BY PHONE. I HAVE REGISTERED WITH THEM. I HAVE ALSO BEEN TOLD BY A CUSTOMER SERVICE REP. THAT I CANNOT GET A NEW MACHINE FROM PHILIPS AT THIS TIME. THUS MY ONLY OPTION IS TO KEEP USING MY RECALLED MACHINE. THIS IS NOT ACCEPTABLE. THERE IS NO MEDICAL TESTS FOR THIS ISSUE. PHILLIPS KNOWS IT HAS A SERIOUS PROBLEM WITH CERTAIN OF THEIR MACHINES AND IS RECALLING THEM. HOWEVER PHILIPS IS NOT PROVIDING THE APPROPRIATE CORRECTIVE ACTION FOR IT'S MORE SERIOUS CPAP USERS. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5102966 | 2021-08-03 | 2021-06-11 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS / RESPIRONICS, INC. |  |  |  | I WOKE UP WITH A SWOLLEN UVULA AFTER SLEEPING WITH RESPIRONICS DREAMSTATION AUTO CPAP. IT WAS SWOLLEN TO THE POINT THAT I COULD NOT SPEAK WITHOUT GAGGING UNCONTROLLABLY AND I HAD DIFFICULTY BREATHING. I WENT TO URGENT CARE AND RECEIVED A PRESCRIPTION FOR PREDNISONE WHICH RESOLVED THE SWELLING OVER THE NEXT SEVERAL DAYS. A RAPID STREP TEST WAS DONE AT URGENT CARE WHICH WAS NEGATIVE. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02749 | 2021-08-03 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
MW5102975 | 2021-08-03 | 2021-01-15 | I, n, j, u, r, y | Voluntary report | PHILLIPS DREAMSTATION RESPIRONICS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500T11C |  |  | PROGRESSIVELY FATIGUED AND SHORT OF BREATH. CONGESTED, COUGHING AND HEADACHES. ACHING JOINTS, SCRATCHY EYES AND THROAT. PROGRESSIVELY GETTING WORSE OVER TIME. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5102976 | 2021-08-03 | 2021-03-14 | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500H11C |  |  | A FEW WEEKS AFTER I STARTED USING MY PHILIPS DREAMSTATION CPAP MACHINE (STARTED USING IT (B)(6) 2021) I STARTED GETTING SINUS IRRITATION WHERE MY NOSE WOULD BE STUFFY ALL THE TIME. I WOULD ONLY HAVE NASAL STUFFINESS ON OCCASION PREVIOUSLY. ALSO MY THROAT AND AIRWAY IN GENERAL SEEMED TO BE IRRITATED AFTER USING THE CPAP. I HAVE FILED WITH PHILIPS TO GET THE MACHINE REPLACED - I ONLY BECAME AWARE OF THE RECALL ON THE CPAP ABOUT TWO WEEKS AGO, AROUND (B)(6), BUT I UNDERSTAND THE RECALL WAS THE END OF (B)(6) 2021. NO ONE FROM ROTECH OR PHILIPS REACHED OUT TO ME - I SAW THIS INFORMATION ONLINE. I HAVEN'T HEARD ANYTHING FROM PHILIPS SINCE AND I STOPPED USING MY CPAP AFTER I FOUND OUT. NOW I'M WITHOUT A CPAP AND INSURANCE ONLY PAYS FOR A NEW UNIT EVERY 3 YEARS. VERY FRUSTRATING. NOT SURE WHAT WILL HAPPEN WITH MY HEALTH LONG-TERM. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5102977 | 2021-08-03 |  | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500T11 |  |  | A BLACK DEPOSIT WAS NOTICEABLE ON MY PILLOW AND ON THE FILTERS. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
MW5102980 | 2021-08-03 | 2021-03-01 | I, n, j, u, r, y | Voluntary report | PHILIPS CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | FOR THE PAST SEVERAL MONTHS, I BEGAN HAVING CHEST PAINS, DIZZINESS, ESOPHAGEAL IRRITATION, NAUSEA, LOWER ABDOMINAL PAINS AND GENERALIZED MALAISE SHORTLY AFTER BEGINNING THE USE OF MY CPAP. NOW I HAVE DEVELOPED A SEVERE UTI THAT HAS IMPACTED MY KIDNEYS. I'VE BEEN TO SEVERAL DOCTORS AND HAD AN ER VISIT TRYING YO FIGURE OUT WHAT IS WRONG. I'VE BEEN GETTING WORSE OVER THE PAST SEVERAL MONTHS. I'VE HAD LAB TESTS, SCANS AND TAKING ANTIBIOTICS AND I FEEL LIKE I'M JUST GETTING WORSE. I JUST VISITED MY PRIMARY CARE PHYSICIAN ON AUGUST 1ST FOR THE RECENT UTI WAS GETTING WORSE AND MY OTHER SYMPTOMS ARE NOT RESOLVING. I WAS GETTING DOWN BECAUSE I WAS HAVING TO STAY ON BED A LOT AND RESTRICTING MY ACTIVITIES JUST SO I WOULDN'T HAVE PAIN OR THE OTHER SYMPTOMS. I RECEIVED INFORMATION ABOUT THE RECALL OF MY PHILLIPS CPAP A FEW DAYS AGO AND NOW I WONDER IF THIS MAY BE THE CAUSE OF MY PROBLEMS. MY DOCTOR IS CONFUSED ABOUT MY SYMPTOMS VS MY TEST RESULTS. NOW THAT I KNOW MORE ABOUT THE RECALL, IT MAY HELP MY DOCTOR FINALLY DIAGNOSE WHAT IS GOING ON WITH ME. I CAN PROVIDE RECENT TEST RESULTS ELECTRONICALLY FROM MY ELECTRONIC HEALTH RECORDS IF YOU CAN SEND ME AND WHERE TO SEND IT. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5102985 | 2021-08-03 | 2021-07-01 | I, n, j, u, r, y | Voluntary report | PHILLIPS DREAM STATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DREAM STATION |  |  | PHILLIPS DREAM STATION CPAP MACHINE, CAUSES SEVERE COUGHING, AND STUFFY NOSE. SOMETIMES WORSE SOMETIMES LESS, AND THIS IS WITH JUST STRAIT AIR NOT WITH THE HEAT FEATURE, AND KNOW I DO NOT USE AN OZONE 3 CLEANER. I HAVE GOTTEN TIRED OF (B)(6) REPS WHO THINK THEY KNOW EVERYTHING, AND THAT'S THE FIRST THING THEY JUMP TOO. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5102986 | 2021-08-03 | 2019-12-01 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I STARTED NOTICING SHORTNESS OF BREATH AND THE FEELING OF HEAVINESS ON MY CHEST. I LATER FOUND OUT ABOUT THE PHILIPS RESPIRONICS RECALL INVOLVING MANY OF THEIR DEVICES. I USE THEIR DREAM STATION C-PAP. I STOPPED USING THE DEVICE UPON HEARING ABOUT THE RECALL AND HAVEN'T RESUMED USING IT SINCE EVEN AFTER SPEAKING WITH MY MEDICAL PROVIDER. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5102989 | 2021-08-03 | 2021-06-12 | I, n, j, u, r, y | Voluntary report | CPAP DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS, INC. |  |  |  | I WAS NOT NOTIFIED BY MANUFACTURER THAT CPAP MACHINE HAD A PROBLEM AND THAT I SHOULD NOT USE IT. I DID RECEIVE CONTACT TO ORDER NEW SUPPLIES AS RECENTLY AS TODAY. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5102990 | 2021-08-03 | 2021-07-07 | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DXS500T11W | DSX500T11W |  | ACUTE ASTHMA; I HAD CHEST TIGHTNESS, SHORTNESS OF BREATH, CONGESTION, A COUGH, AND SINUS CONGESTION FOR ABOUT 10 DAYS. WENT TO EMERGENCY ROOM ON (B)(6) 2021. HAD AN EKG. WAS TREATED WITH A NEBULIZER, O2, AND ALBUTEROL. WAS PRESCRIBED PREDNISONE AND ALBUTEROL. IT'S NOW (B)(6) AND I AM STILL HAVING BREATHING PROBLEMS, ALTHOUGH THEY ARE BETTER THEN IN EARLY (B)(6). FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5102994 | 2021-08-03 | 2021-01-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO BIPAP HUMIDIFIER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | J2050169532D2 |  |  | I THOUGHT I HAD DEVELOPED A SINUS INFECTION BUT AFTER TRYING TO REMEDY THE SITUATION WITH OTC MEDS THERE WAS NO IMPROVEMENT. I TALKED TO DR AND WAS PRESCRIBED AMOXICILLIN AND AFTER ENTIRE TREATMENT I'M STILL DEALING WITH SINUS ISSUES/HEADACHES ETC. I USE A DREAMSTATION CPAP W/HUMIDIFIER AND STARTED WITH SINUS ISSUES A GOOD WHILE BACK, NOW JUST FINDING OUT ABOUT RECALL ISSUES WITH MACHINE. I PAID CASH FOR MY MACHINE, AS MY INSURANCE DIDN'T COVER. I CAN'T AFFORD TO GO OUT AND PURCHASE A NEW ONE. I JUST PAID THIS ONE OFF! I DRIVE TRUCK FOR A LIVING, HAULING WIDE LOADS, WITHOUT MY MACHINE I LITERALLY FALL ASLEEP NUMEROUS TIMES THROUGHOUT THE DAY. I DON'T HAVE ANOTHER ALTERNATIVE TO USING MY MACHINE. I CAN'T AFFORD TO REPLACE IT ON MY OWN, BUT I DEFINITELY CAN NOT GO WITHOUT USING IT IF I WANT TO KEEP MY JOB. MY DR. SAID HE COULD GIVE ME A NEW SCRIPT FOR A NEW MACHINE, OTHER THAN THAT HE ADVISED NO HELP. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5102956 | 2021-08-03 | 2021-07-30 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | PT USING PHILIPS DREAMSTATION DEVICE, INVOLVED IN RECALL. REPORTS CHRONIC COUGH, HEADACHE, DIZZINESS. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | I |  | 2021-08
MW5102996 | 2021-08-03 | 2018-07-02 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION AUTO CPAP HUMCELL DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500H11C |  |  | BEGAN USING THE PHILIPS RESPIRONICS CPAP DREAMSTATION MACHINE AND BEGAN HAVING ONGOING SINUS PROBLEMS, CHEST PRESSURE, HEAVY COUGH, WHEEZING AND RESPIRATORY PROBLEMS. ALL OF THESE ARE ONGOING PROBLEMS. I HAD A CHEST XRAY BEFORE COVID STARTED FOR THE COUGH. RESULTS DIDN'T SHOW ANYTHING AT THAT TIME. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5102997 | 2021-08-03 | 2021-06-01 | I, n, j, u, r, y | Voluntary report | DREAMSTNAUTOCPAPHUMHTCELL DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DREAMSTATION |  |  | USING A CPAP MACHINE FOR SLEEP APNEA. BEGAN HAVING BREATHING PROBLEMS SEVERAL MONTHS AGO BEGAN DIFFERENT BREATHING TREATMENTS UNTIL I WAS TOLD THAT THE PHILIPS REPIRONICS MACHINE HAD BEEN RECALLED. I CONFIRMED THAT MY MACHINE HAD BEEN RECALLED AND WAS INSTRUCTED BY PHILIPS TO DISCONTINUE USE. AFTER TWO WEEKS OF DISCONTINUED USE MY BREATHING IMPROVED. HOWEVER I DESPERATELY NEED A CPAP MACHINE TO ADDRESS MY SLEEP APNEA. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02716 | 2021-08-03 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02718 | 2021-08-03 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02550 | 2021-08-03 | 2021-06-18 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP INFLAMMATORY NODULES AND SHORTNESS OF BREATH. THE PATIENT REQUIRED MEDICAL INTERVENTION IN THE FORM OF A DOCTOR'S ASSESSMENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02717 | 2021-08-03 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02650 | 2021-08-03 | 2021-07-06 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION BIPAP AVAPS30 AAM H/HT, CAN | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | CAX1130T12 | CAX1130T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED CONGESTIVE HEART FAILURE. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING TO DEVELOPED CONGESTIVE HEART FAILURE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM.THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE EVENT IS CURRENTLY UNKNOWN. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02616 | 2021-08-03 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX500T12C | CAX500T12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION A1 (PATIENT IDENTIFIER) AND A3 (PATIENT SEX) WERE MISSED IN THE INITIAL REPORT, WHICH WERE CORRECTLY UPDATED IN THIS FOLLOW UP REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02549 | 2021-08-03 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP A COUGH. THE PATIENT REQUIRED MEDICAL INTERVENTION IN THE FORM OF INHALERS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP A COUGH. THE PATIENT REQUIRED MEDICAL INTERVENTION IN THE FORM OF INHALERS. THE REPORTED EVENT OF COUGH AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. DURING THE INVESTIGATION OF THE DEVICE, PIL OBSERVED AN UNKNOWN CONTAMINANT ON THE ANTI-SKID PAD ON THE BOTTOM ENCLOSURE, AND AROUND THE OUTSIDE OF THE P4 JACK INSERT. AN UNKNOWN CONTAMINANT WAS OBSERVED ON THE TOP ENCLOSURE. THE FRONT PANEL WAS OBSERVED TO HAVE A HAIR-LIKE PARTICLE ON IT. AN UNKNOWN DUST CONTAMINANT WAS OBSERVED ON THE BLOWER BOX, BLOWER SEAL, AND BOTTOM ENCLOSURE. EVIDENCE OF LIQUID INGRESS WITH MINERAL DEPOSITS WAS OBSERVED TO THE BLOWER AND BLOWER BOX. INV1023 ADDRESSES THE ISSUE OF LIQUID INGRESS. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERORRS. THE MANUFACTURER CONCLUDES THE DEVICE HAS DUST CONTAMINATION. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION D8, D9 AND H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02551 | 2021-08-03 | 2021-06-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP A ASTHMA, SORE THROAT, HEADACHES, AND SINUS INFECTIONS. THE PATIENT REQUIRED MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION FOR ANTIBIOTICS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP A ASTHMA, SORE THROAT, HEADACHES, AND SINUS INFECTIONS. THE PATIENT REQUIRED MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION FOR ANTIBIOTICS REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02615 | 2021-08-03 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02548 | 2021-08-03 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND RESULTED IN A CRACKED TOOTH AND DAMAGE TO THE UPPER RIGHT JAW. THE PATIENT REQUIRED MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION FOR ANTIBIOTICS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 13 , 2024 AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT IN A CRACKED TOOTH AND DAMAGE TO THE UPPER RIGHT JAW. THE PATIENT REQUIRED MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION FOR ANTIBIOTICS. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02553 | 2021-08-03 | 2021-06-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX400H11 | DSX400H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP BREAST CANCER. THE PATIENT REQUIRED MEDICAL INTERVENTION IN THE FORM OF A DOCTOR'S ASSESSMENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THE END USER INDICATED OZONE OR OTHER UNAPPROVED CLEANING AND DISINFECTION METHOD WAS USED. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02506 | 2021-08-03 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A RESPIRATORY INFECTION. THE PATIENT REQUIRED MEDICAL INTERVENTION IN THE FORM OF A DOCTOR'S ASSESSMENT AND PRESCRIPTION FOR ANTIBIOTICS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 18 , 2024 AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT ALSO ALLEGED HEADACHE FOR 3 WEEKS, SORE THROAT, AND HEAVY COUGH, FEELING BAD, RESPIRATORY INFECTION. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02505 | 2021-08-03 | 2021-06-24 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED SINUS INFECTIONS. THE PATIENT REQUIRED MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION FOR ANTIBIOTICS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM ALLEGEDLY BECAME DEGRADED AND CAUSED SINUS INFECTIONS. THE PATIENT REQUIRED MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION FOR ANTIBIOTICS. THE DEVICE WAS RETURNED TO THE MANUFACTURER QUALITY PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. DURING THE INITIAL VISUAL INSPECTION THE MANUFACTURER FOUND NO DISTINCT WEAR, TEAR, OR CONTAMINATION ON THE ENCLOSURE OF THE DEVICE. THE DEVICE WAS DISASSEMBLED FOR INTERNAL VISUAL INSPECTION. THE MANUFACTURER FOUND EVIDENCE OF DUST CONTAMINATION ON THE INTERIOR OF ALL ENCLOSURE PIECES, AS WELL AS DUST CONTAMINATION IN THE BLOWER. THE MANUFACTURER ALSO FOUND EVIDENCE OF WATER INGRESS AND MINERAL DEPOSIT CONTAMINATION ON THE BLOWER AND IN THE BLOWER BOX. NO EVIDENCE OF SOUND ABATEMENT FOAM BREAKDOWN WAS OBSERVED, AND THE FOAM APPEARS INTACT. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND CONFIRMED THE DEVICE PROVIDED FLOW. THE MANUFACTURER CONCLUDES THERE IS NO EVIDENCE OF SOUND ABATEMENT FOAM BREAKDOWN WITHIN THE DEVICE. THE CONTAMINATION WAS FOUND IN THE AS OBSERVED IN THE BLOWER AND IN THE BLOWER BOX; THE MANUFACTURER ALSO FOUND EVIDENCE OF WATER INGRESS IN THE DEVICE. SECTION H6 HAS BEEN UPDATED/CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02613 | 2021-08-03 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX500T12 | CAX500T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING FILTER WAS BLACK RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. SECTION A1 AND D4 WAS CORRECTED WHICH WAS WRONG IN INITIAL REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02711 | 2021-08-03 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02751 | 2021-08-03 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION EXPERT | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | CAX501T12 | CAX501T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. AFTER THE THIRD ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, THE CUSTOMER STATED THAT THE DEVICE WILL NOT BE RETURNED. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS UPDATED REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02720 | 2021-08-03 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. CORRECTION WERE MADE IN A1, A3 SECTIONS AND H6 SECTION WAS UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02683 | 2021-08-04 | 2021-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE REPORTER ALSO ALLEGED BLACK MATERIAL FOUND ON THE FILTER AND STOMA SITE OF THE PATIENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER PERFORMED VISUAL INSPECTED THE DEVICE. THE MANUFACTURER OPENED THE DEVICE TO CHECK FOR ANY CONTAMINATION/BLACK PARTICLES WITHIN THE FLOW PATHS. THERE ARE NO SIGNS OF CONTAMINATES ENTERING THE FLOW PATH FROM OUTSIDE THE DEVICE. THE MANUFACTURER FOUND NO SIGNS OF FOAM DEGRADATION. THE DEVICE POWERED ON AND PROVIDED THERAPY. THE ERROR LOG OF THE DEVICE WAS DOWNLOADED AND REVIEWED. THE DEVICE WAS TESTED AND THE MANUFACTURER CONFIRMED THE DEVICE OPERATES AS DESIGNED. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF FOAM DEGRADATION. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
MW5103014 | 2021-08-04 | 2020-05-24 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500T11C | DSX500T11 |  | THIS HAS BEEN ONGOING SINCE AROUND 3 MONTHS AFTER RECEIVING MY PHILIPS DREAMSTATION AUTO CPAP ON (B)(6) 2020. STARTING GETTING SINUS HEADACHES ON A REGULAR BASIS, USUALLY 5 DAYS A WEEK. OFTEN THESE HEADACHES WERE PRESENT UPON WAKING. I SOUGHT MEDICAL ASSISTANCE FROM MY DENTIST FOR TEETH GRINDING THINKING THAT COULD BE CAUSING THE ISSUE. NEW MOUTH GUARD WAS MADE AND IT DID NOT SOLVE THE PROBLEM. I SAW AN OPTOMETRIST AFTER THAT THINKING THAT MY EYES COULD BE STRAINING AND CAUSING THE HEADACHES. I RECEIVED A MILD PRESCRIPTION FOR READING GLASSES THAT DID NOT STOP THE HEADACHES. I THEN WENT TO AN OPHTHALMOLOGIST TO HAVE MY EYES FULLY CHECKED. THERE WAS NO ISSUE FOUND WITH MY EYES. THE OPHTHALMOLOGIST RECOMMENDED THAT I SEE A NEUROLOGIST. THE NEUROLOGIST RECOMMENDED A NERVE BLOCK AROUND MY EYE. WE DID THIS AND IT DID HELP FOR AROUND SIX WEEKS. DURING THIS TIME I ALSO NOTED TO MY DOCTOR THAT I HAD BECOME HYPERSENSITIVE ON MY SKIN. I FIRST NOTICED THIS DURING ACUPUNCTURE. I HAD BEEN DOING ACUPUNCTURE FOR SEVERAL MONTHS ON A BI-WEEKLY BASIS WITH NO ISSUES. SEEMINGLY OUT OF NOWHERE I HAD TO STOP GOING BECAUSE EVERY NEEDLE WOULD MAKE ME JUMP AND IT HURT. I HAVE NEVER HAD THIS ISSUE BEFORE. ALSO DURING THIS TIME I HAVE NOTED A FEW TIMES THAT A NASAL SECRETION WHICH DUE TO THE COLOR SEEMED LIKE BRAIN FLUID LEAKING THROUGH. THIS ONLY LASTED FOR A DAY EACH TIME AND I HAVE NOT YET MADE AN APPOINTMENT WITH THE ENT TO GET THIS EXAMINED. ALSO OF NOTE, THE PHILIPS CPAP MACHINE I RECEIVED ON (B)(6) 2020 HAS GOTTEN SIGNIFICANTLY LOUDER AND MY WIFE HAS CONFIRMED THIS WITH ME ALSO. AS I WRITE THIS, I HAVE A HEADACHE THAT I WOKE UP WITH TODAY. I STILL HAVE MY OLDER RESMED MACHINE TO USE AND I WILL BE DISCONTINUING USE OF THE PHILIPS MACHINE IMMEDIATELY. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103003 | 2021-08-04 |  | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO CPAP HUM DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500H11 |  |  | SINCE I BEGAN USING A NEW CONTINUOUS CPAP DEVICE, I HAVE BEEN SUFFERING FROM UNEXPLAINED DIZZINESS, HEADACHES AND HYPERSENSITIVITY. I HAVE REPORTED TO MEDICAL AT WORK AND I HAVE REPORTED THE PROBLEMS TO MY MEDICAL PROVIDER. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02583 | 2021-08-04 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WOULD NOT DELIVER 100% OXYGEN. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
MW5103018 | 2021-08-04 | 2021-01-17 | I, n, j, u, r, y | Voluntary report | CPAP PHILLIPS RESPIRONICS SYSTEM ONE HTD HUMID | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | PHILIPS / RESPIRONICS, INC. | SYSTEM ONE |  |  | FOUND BLACK FOAM PARTICULATES IN THE WATER TANK OF THE CPAP - RESPIRONICS REMSTAR PLUS, DOM S/N (B)(4); SYSTEM ONE HTD HUMID, DOM S/N (B)(4). FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103005 | 2021-08-04 |  | N, o, , a, n, s, w, e, r, , p, r, o, v, i, d, e, d | Voluntary report | PHILIPS RESPIRONICS DREAM STATION WITH HEATED HUMIDIFIER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS / RESPIRONICS, INC. | DSX500H11C |  |  | I WAS ADVISED BY MY MEDICAL PROVIDER TO USE A CPAP DEVICE AS MY OXYGEN COUNT WAS LOW DURING SLEEP. I CONTRACTED WITH MY LOCAL PHARMACY TO PURCHASE AND USE A PHILIPS RESPIRONICS DREAM STATION CPAP. I ENDED THE LAST PAYMENT (B)(6) 2021. DURING THE USE OF MY DREAM STATION, I EXPERIENCED CONSIDERABLE RESPIRATORY DISTRESS. IN THE MORNINGS AFTER A NIGHT OF USING THE CPAP, MY SINUS AND LUNGS WERE CODED WITH WHITE TO ANY MUCUS. UNDER GREAT DISTRESS, I WOULD CLEAR AND SPIT THIS MUCUS FOR 15-20 MINUTES. AT TIMES I THOUGHT I WAS HAVING A HEART ATTACK. WHEN I BECAME AWARE OF THE RECALL, I REALIZED WHAT CAUSED THE DISTRESS OF MY RESPIRATORY SYSTEM. |  | N | N | LAY USER/PATIENT |  | 2021-08
MW5103019 | 2021-08-04 |  | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS DREAM STATION CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | HEADACHES, CONSTANT COUGH, WINDED, SORE THROAT AND NOW POSSIBLE KIDNEY ISSUES AS I HAVE CONSTANT URINATION. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103027 | 2021-08-04 | 2021-07-26 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DREAMSTATION |  |  | CPAP MACHINE PHILIPS BRAND BLACK PARTICLES IN THE WATER TRAY. FDA SAFETY REPORT ID # (B)(6). |  | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02727 | 2021-08-04 | 2021-07-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | IN2200X15B | IN2200X15B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED CALIBRATION. THE DEVICE'S OXYGEN BLENDING MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02729 | 2021-08-04 | 2021-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S DISPLAY BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02730 | 2021-08-04 | 2021-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02726 | 2021-08-04 | 2021-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE BLOWER MOTOR WAS OBSERVED.. THE DEVICE'S BLOWER MOTOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02725 | 2021-08-04 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02728 | 2021-08-04 | 2021-07-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02744 | 2021-08-04 | 2021-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION RELATED TO THE DEVICE'S OXYGEN BLENDING MODULE BOARD OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED DUE TO PHYSICAL DAMAGE. THE DEVICE'S UNIVERSAL PORTING BLOCK WAS REPLACED FOR PHYSICAL DAMAGE. THE ACTIVE EXHALATION CONTROL MODULE WAS NOT REPLACED. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02685 | 2021-08-04 | 2021-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096B | 1054096B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE REPORTER ALSO ALLEGED BLACK PIECES IN THE CHILD'S VOMIT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER PERFORMED VISUAL INSPECTED THE DEVICE. THE MANUFACTURER OPENED THE DEVICE TO CHECK FOR ANY CONTAMINATION/BLACK PARTICLES WITHIN THE FLOW PATHS. THERE ARE NO SIGNS OF CONTAMINATES ENTERING THE FLOW PATH FROM OUTSIDE THE DEVICE. THE MANUFACTURER FOUND NO SIGNS OF FOAM DEGRADATION. THE DEVICE POWERED ON AND PROVIDED THERAPY. THE ERROR LOG OF THE DEVICE WAS DOWNLOADED AND REVIEWED. THE DEVICE WAS TESTED AND THE MANUFACTURER CONFIRMED THE DEVICE OPERATES AS DESIGNED. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF FOAM DEGRADATION. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02746 | 2021-08-04 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02684 | 2021-08-04 | 2021-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096B | 1054096B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE REPORTER ALSO ALLEGED BLACK MATERIAL FOUND ON THE FILTER AND STOMA SITE OF THE PATIENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER PERFORMED VISUAL INSPECTED THE DEVICE. THE MANUFACTURER OPENED THE DEVICE TO CHECK FOR ANY CONTAMINATION/BLACK PARTICLES WITHIN THE FLOW PATHS. THERE ARE NO SIGNS OF CONTAMINATES ENTERING THE FLOW PATH FROM OUTSIDE THE DEVICE. THE MANUFACTURER FOUND NO SIGNS OF FOAM DEGRADATION. THE DEVICE POWERED ON AND PROVIDED THERAPY. THE ERROR LOG OF THE DEVICE WAS DOWNLOADED AND REVIEWED. THE DEVICE WAS TESTED AND THE MANUFACTURER CONFIRMED THE DEVICE OPERATES AS DESIGNED. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF FOAM DEGRADATION. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02743 | 2021-08-04 | 2021-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLANK. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR'S LCD SCREEN WAS BLANK. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT COULD NOT BE CONFIRMED OR DUPLICATED. THERE WAS NO MALFUNCTION OF THE LCD SCREEN. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02739 | 2021-08-04 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR WAS ALARMING FOR A "VENTILATOR INOPERATIVE" ALARM CONDITION. THE DEVICE WAS EVALUATED AT THE MANUFACTURER'S SERVICE CENTER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02741 | 2021-08-04 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR WAS ALARMING FOR A "VENTILATOR INOPERATIVE" ALARM CONDITION. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND A "VENTILATOR INOPERATIVE" ALARM WAS OBSERVED IN THE DEVICE'S DOWNLOADED EVENT LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02740 | 2021-08-04 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR'S SYSTEM BOARD WAS REPLACED TO ADDRESS A VENTILATOR INOPERATIVE ALARM CONDITION. AFTER FURTHER EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE SYSTEM BOARD WAS NOT REPLACED. THE DEVICE'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02742 | 2021-08-04 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR WAS ALARMING FOR A "VENTILATOR INOPERATIVE" ALARM CONDITION. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT COULD NOT BE CONFIRMED OR DUPLICATED. THE DEVICE PERFORMED TO MANUFACTURER'S SPECIFICATIONS. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
MW5103028 | 2021-08-04 |  | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTARION C-PAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | EXPERIENCED SHALLOW BREATHING, SHORTNESS OF BREATH, UNABLE TO INHALE SUFFICIENTLY. REQUIRED AMBULANCE TRANSPORT TO ER AT 2AM. WAS USING PHILIPS DREAMSTATION C-PAP AND HAD RECENTLY BEGAN USING SO CLEAN OZONE GENERATOR FOR CLEANING THE C-PAP. RECEIVED OXYGEN AND OTHER TREATMENTS. RELEASED AFTER SEVERAL HOURS. (B)(6). FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02747 | 2021-08-04 | 2021-08-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02745 | 2021-08-04 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, A VENTILATOR'S LCD SCREEN WAS BLANK. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
MW5103020 | 2021-08-04 | 2021-07-22 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DREAMSTATION AUTO CPAP |  |  | WENT TO EMERGENCY ROOM FOR CHEST TIGHTNESS/SHORTNESS OF BREATH. FOLLOW UP LAB TEST REVEALED LUNG SCARRING, LUNG NODULES, ELEVATED DIAPHRAGM. USE PHILLIPS DREAM STATION FOR ALMOST 5 YEARS. MODEL WAS ON RECALL LIST. NON-SMOKER. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02686 | 2021-08-04 | 2021-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | PHILIPS RESPIRONICS INC. | 1054096B | 1054096B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02721 | 2021-08-04 | 2021-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02722 | 2021-08-04 | 2022-10-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 14, 2024. SECTION H11 SHOULD BE REPORTED AS THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO A CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES SOUND ABATEMENT FOAM BECOME DEGRADED AND ALLEGED VISUALIZATION OF PARTICLES. PATIENT ALSO ALLEGED OF SINUS IRRITATION, NASAL/THROAT IRRITATION, ALLERGY MEDICATION AND NASAL FLUSHES, SINUS SURGERY, SPHINCTER PALATOPLASTY, DIZZINESS RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM (ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR) SECTION B2 WAS CHANGED TO BLANK (PREVIOUSLY IT WAS OTHER SERIOUS) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION SECTION H6 HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02552 | 2021-08-04 | 2021-06-25 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 14, 2024, AND SECTION H11 SHOULD BE REPORTED AS THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED OF VISUALIZATION OF PARTICLES IN THE DEVICE, DIFFICULTY IN BREATHING, NASAL IRRITATION, SORE THROAT, DRYNESS IN THROAT, NAUSEA, VOMITING, EAR ITCHING, SINUS INFECTION AND IRRITATION, SPOT ON SKIN AND UNABLE TO SLEEP. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM (PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR). SECTION B2 WAS CHANGED TO OTHER SERIOUS (PREVIOUSLY IT WAS BLANK). SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02719 | 2021-08-04 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER INITIALLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN THE INITIAL REPORT SECTION B5 WAS NOT UPDATED, THE CORRECT B5 SHOULD BE - THE PATIENT HAS ALLEGED PARTICLES IN THE DEVICE, HEADACHES, DRY MOUTH. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO EVIDENCE OF CONTAMINATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. A HUMIDIFIER WAS ALSO RETURNED TO THE MANUFACTURER ALONG WITH DEVICE. AN EXTERNAL VISUAL INSPECTION OF THE HUMIDIFIER WAS COMPLETED BY THE MANUFACTURER AND FOUND NO EVIDENCE OF CONTAMINATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND SLIGHT CONTAMINATION ON THE DRY BOX SEAL. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERROR LOGGED. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH CONTAMINATION ON THE DRY BOX SEAL OF THE HUMIDIFIER. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02585 | 2021-08-04 | 2021-06-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP CANCER. THE DETAILS OF THE MEDICAL INTERVENTION THAT THE PATIENT REQUIRED ARE CURRENTLY UNKNOWN. THE PATIENT REFUSED TO PROVIDE DETAILS ON THE SERIAL NUMBER OF THE DEVICE OR CONTACT INFORMATION. IF FURTHER INFORMATION REGARDING THE DEVICE BECOME AVAILABLE, IT WILL BE SUBMITTED IN THE FOLLOW-UP REPORT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY SUBMITTED MDR 2518422-2021-02585-2 WITH INCORRECT SECTIONS B1, B2, H1, H6. CORRECTIONS TO PREVIOUS MDR ARE MADE IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY PRODUCT PROBLEM WAS CHECKED IN PREVIOUS MDR). SECTION B2 WAS CORRECTED TO HOSPITALIZATION (INITIAL OR PROLONGED). (PREVIOUSLY IT WAS BLANK). SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6- HEALTH EFFECT - IMPACT CODE WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02422 | 2021-08-04 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES IN DEVICE AND WATER CHAMBER. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02698 | 2021-08-04 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11 | DSX700T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP THROAT CANCER. THE PATIENT WAS HOSPITALIZED FOR TREATMENT IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON OCTOBER 02, 2023, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO BIPAP DEVICE'S SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP THROAT CANCER. THE PATIENT WAS HOSPITALIZED FOR TREATMENT IN RESPONSE TO THE REPORTED EVENT. THE END USER INDICATED OZONE OR OTHER UNAPPROVED CLEANING AND DISINFECTION METHOD WAS USED. THE REPORTED EVENT AND ITS REPORTED SEVERITY WAS REVIEWED BY THE PMS CLINICAL EXPERT AND EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGATION OF LARYNGEAL CANCER. SECTION E1 WAS UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02363 | 2021-08-04 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, INTL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | INX500S15 | INX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02695 | 2021-08-04 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED ON THIS DEVICE IN MDR 2518422-2021-02643. THIS REPORT WAS SUBMITTED AS A DUPLICATE REPORT OF THE PREVIOUSLY SUBMITTED REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02694 | 2021-08-04 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02653 | 2021-08-04 | 2021-07-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO, FR | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | FRX500S14 | FRX500S14 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | IN THE PREVIOUS REPORT, SECTION B5 WAS MISSED TO CAPTURE AND IT WAS CAPTURED IN THIS REPORT. THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED IRRITATION OF THE RESPIRATORY TRACT, DRY COUGH, FEELING OF PRESSURE IN THE CHEST, CLEAR NASAL SECRETION, MORNING HEADACHE, BLACK DUST IN THE FILTER AND MASK. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED. THERE WAS ONE ERROR FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02733 | 2021-08-04 | 2021-06-18 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SHORTNESS OF BREATH, KIDNEY ISSUES, AND A HEART ATTACK. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY SUBMITTED MDR 2518422-2021-02733-2 WITH INCORRECT SECTIONS B1, B2, H1, H6. CORRECTIONS TO PREVIOUS MDR ARE MADE IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY PRODUCT PROBLEM WAS CHECKED IN PREVIOUS MDR). SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY IT WAS BLANK). SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6- HEALTH EFFECT - CLINICAL CODE, HEALTH EFFECT - IMPACT CODE WAS CORRECTED IN THIS. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02424 | 2021-08-04 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/WIFI,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11W | DSX500T11W |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES, LIGHT BROWN GUNK FLOATING IN TANK EVERY TIME CHANGES WATER AND SMELLS BAD EVERY NOW AND THEN WAKES CUSTOMER UP RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. CORRECTED INFORMATION PROVIDED IN B5, THE INFORMATION (THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES, LIGHT BROWN GUNK FLOATING IN TANK EVERY TIME CHANGES WATER AND SMELLS BAD EVERY NOW AND THEN WAKES CUSTOMER UP RELATED TO A CPAP DEVICE) WAS NOT INCLUDED IN INITIAL REPORT. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02423 | 2021-08-04 | 2021-06-28 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING CANCER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02687 | 2021-08-04 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | PHILIPS RESPIRONICS, INC. | 1054655B | 1054655B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE DEVICE'S FILTERS ARE BLACKISH IN COLOR. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. THE MANUFACTURER VISUALLY INPSECTED THE DEVICE EXTERNALLY AND FOUND NO ISSUES. THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE WAS APPLIED POWER AND THE MANUFACTURER FOUND THE DEVICE TO OPERATE PROPERLY. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02638 | 2021-08-04 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/HT/CELL, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX400T12C | CAX400T12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DISCOVERED NODULES FORMING IN ONE LUNG. THE PATIENT REQUIRED MEDICAL INTERVENTION IN THE FORM OF A DOCTOR'S ASSESSMENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 01, 2024, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DISCOVERED NODULES FORMING IN ONE LUNG. THE PATIENT REQUIRED MEDICAL INTERVENTION IN THE FORM OF A DOCTOR'S ASSESSMENT. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02515 | 2021-08-04 | 2021-06-29 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP TONGUE CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP TONGUE CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE LIGHT BEIGE DUST OR DIRT CONTAMINATE, DARK FIBER CONTAMINATE, AND POTENTIAL KERATIN CONTAMINATE, A SEPARATE SMALL YELLOW BODY CONTAMINATE. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND NO ERRORS LOGGED. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF LIGHT BEIGE DUST OR DIRT CONTAMINATE, DARK FIBER CONTAMINATE, AND POTENTIAL KERATIN CONTAMINATE, A SEPARATE SMALL YELLOW BODY CONTAMINATE. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION D9, G3 AND H6 HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02449 | 2021-08-04 | 2022-01-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | PHILIPS RESPIRONICS, INC. | 1054655TP | 1054655TP |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THE DEVICE WAS NOT IN PATIENT USE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. THE MANUFACTURER FOUND THE BLOWER IMPELLER AND BLOWER CAP TO BE CONTAMINATED FROM AN EXTERNAL SOURCE. THERE WAS NO VISUAL OBSERVATION OF FOAM DEGRADATION IN THE DEVICE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02658 | 2021-08-04 | 2021-07-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DEX400S13 | DEX400S13 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. CORRECTED DATA: B3 DATE OF EVENT: PREVIOUSLY REPORTED AS 07/27/2021 ; UPDATED TO 07/30/2021 D8 WAS THIS DEVICE SERVICED BY A THIRD PARTY?: PREVIOUSLY REPORTED AS NO ; UPDATED TO UNKNOWN. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02656 | 2021-08-04 | 2021-07-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DEX500S13 | DEX500S13 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED NOSEBLEEDS AND SEVERE RESPIRATORY PROBLEMS MUCUS IN THE THROAT BLACK PARTICLES. THE DEVICE WAS RETURNED TO THE THIRD PARTY SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED AND DUST CONTAMINATION WAS FOUND. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED . THERE WERE NO ERROR FOUND. THE THIRD PARTY SERVICE CENTER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. SECTION H6, IN THE PREVIOUS REPORT; THEY MISSED TO CAPTURE CLINICAL CODE FOR NOSEBLEEDS AND SEVERE RESPIRATORY PROBLEMS MUCUS IN THE THROAT BLACK PARTICLES, THAT WAS CAPTURED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02416 | 2021-08-04 | 2021-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX500T12C | CAX500T12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT ALLEGED VISUALIZATION OF PARTICLES IN THE TUBING AND CHAMBER AND THAT THE AIR WAS WARMER THAN USUAL. THERE WAS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL INVESTIGATION WAS PERFORMED ON THE HUMIDIFIER AND THE DREAMSTATION. NO CONTAMINATION WAS FOUND ON THE OUTSIDE OF THE DEVICE. THE DEVICE WAS HOOKED UP TO A KNOWN GOOD POWER SUPPLY AND CORD. DEVICE POWERED ON PROPERLY AND ERROR LOGS WERE RETRIEVED. AN INTERNAL INVESTIGATION WAS PERFORMED ON THE HUMIDIFIER AND DREAMSTATION BY THE PRODUCT INVESTIGATION LAB. CONTAMINATION WAS FOUND IN THE HUMIDIFIER AND MINERAL DEPOSITS IN THE RESERVOIR. CONTAMINATION IN ISO HEATED TUBE PORT WERE ALSO OBSERVED. NO CONTAMINATION WAS FOUND IN THE DREAMSTATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND ZERO ERRORS ON THE DEVICE. THE MANUFACTURER WAS UNABLE TO CONFIRM THE CUSTOMER'S COMPLAINT. THE MANUFACTURER DOES CONCLUDE CONTAMINATION WAS OBSERVED IN THE HUMIDIFIER, RESERVOIR, AND ISO PORT. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. PLEASE SEE SECTIONS D9, D10, G1, AND H3 FOR THIS UPDATED CODING AND INFORMATION. | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103055 | 2021-08-05 | 2021-07-30 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DNX500H11C | DNX500H11C |  | ABNORMAL CHEST X-RAY; ON (B)(6) 2021 I WENT TO (B)(6) FOR A CHEST X-RAY AND BLOODWORK FOR UPCOMING SURGICAL CLEARANCE. LATER THAT DAY I WAS CONTACTED AND TOLD THAT MY CHEST X-RAY HAD CHANGED FROM A PRIOR CHEST X-RAY. THE REPORT STATES "NEW MILD MEDIAL LEFT LOWER LOBE INTERSTITIAL PROMINENCE NOTED" AND "MILD RETICULATION OF THE LEFT LOWER LOBE." I AM SCHEDULED FOR ANOTHER CHEST X-RAY ON (B)(6) 2021. IN ADDITION I HAD BLOODWORK WHICH REFLECTED A HIGH VALUE ON THE EOSINOPHILS OF 4.5% WHERE THE STANDARD RANGE IS 1.0-3.0%. THIS REPORT IS BEING DONE DUE TO CONCERNS OVER THE USE OF THE PHILIPS RESPIRONIC DREAMSTATION MACHINE WHICH IS REQUIRED DUE TO A STROKE AND SLEEP APNEA. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02664 | 2021-08-05 | 2021-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 02 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, SERVICE REQUIRED CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG AND THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD, INTERFACE BOARD, AND SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02665 | 2021-08-05 | 2021-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING AND A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED EVENT LOG. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02797 | 2021-08-05 | 2021-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2200X11B | DS2200X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02750 | 2021-08-05 | 2021-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS,INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RECEIVED BY THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED THIS DEVICE ON MDR 2518422-2021-02750-2. PLEASE DISREGARD MDR 2518422-2021-02750-2 AS IT WAS FILED IN ERROR. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02801 | 2021-08-05 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2200X11B | DS2200X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SCREEN WAS CRACKED AND NON FUNCTIONAL. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS EVALUATED BY A FIELD SERVICE ENGINEER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S DISPLAY BOARD AND DISPLAY WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02802 | 2021-08-05 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, WHILE TESTING THE DEVICE'S POWER FAIL ALARM IT DID NOT SOUND FOR THE SPECIFIED 2 MINUTES. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02800 | 2021-08-05 | 2021-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2200X11B | DS2200X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS FAILING A TEST STEP DURING TESTING. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR WAS FAILING A TEST STEP DURING TESTING. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS EVALUATED BY A THIRD PARTY FIELD SERVICE ENGINEER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD, BLOWER AND SOLENOID VALVE NEED REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION, A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02799 | 2021-08-05 | 2021-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2200X11B | DS2200X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02798 | 2021-08-05 | 2021-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2200X11B | DS2200X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02788 | 2021-08-05 | 2021-08-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VARIATION IN THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE DELIVERY WAS OBSERVED. THE PATIENT EXPERIENCED BLEEDING FROM THE LUNGS. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 13, 2024, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AND REPORTED A VENTILATOR HAD A VARIATION IN THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE DELIVERY WAS OBSERVED. THE PATIENT EXPERIENCED BLEEDING FROM THE LUNGS TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTIONS B1, B2, H1 AND H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR ONLY THE PRODUCT PROBLEM (BOTH ADVERSE EVENTS AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02796 | 2021-08-05 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A MALFUNCTION OF A VENTILATOR'S TOUCHSCREEN. THERE WAS NO PATIENT HARM OR INJURY REPORTED. THE VENTILATOR WAS EVALUATED BY A FIELD SERVICE ENGINEER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S DISPLAY CABLES WERE RESETED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
MW5103044 | 2021-08-05 | 2017-06-01 | I, n, j, u, r, y | Voluntary report | TRILOGY 100 VENTILATOR | CONTINUOUS, VENTILATOR, HOME USE | PHILIPS / RESPIRONICS, INC. | 1054260B | REV00 |  | MY SON USES THE TRILOGY 100 VENTILATOR. THEY HAVE BEEN RECALLED WHICH WE STILL HAVE NOT BEEN NOTIFIED EXCEPT I SAW IT IN AN ARTICLE. HE HAD A HORRIBLE EYE REACTION IN (B)(6) OF 2017. SEVERAL OPHTHALMOLOGIST HAVE SEEN HIM SAYING IT IS A ALLERGIC REACTION AND WE NEED TO FIX THE PROBLEM RIGHT WAY. HE IS ON AN EYE DROP SO HIS EYES DO NOT SWELL SHUT THAT WILL HURT HIS CORNEAS OVER TIME. THE CONDITION IS WORSENING. HE IS ALLERGIC TO ALL MATERIALS EXCEPT COTTON. THE INTERIOR FOAM IN THE TRILOGY IS MADE OF POLYURETHANE FOAM THAT IS DESECRATING WHICH IS CAUSING OFF GASSING. WE HAVE BEEN TOLD THAT THERE ARE NO OTHER VENTILATORS SO NO ONES CAN BE REPLACED EVEN THOUGH HE IS HAVING BAD PHYSICAL REACTIONS. IS THE A TIME LIMIT ON THE REPLACEMENT? WAS THE COMPANY OR OUR DME COMPANY SUPPOSE TO TELL US? WHAT CAN I DO IN THE MEAN TIME TO HELP HIM? THANK YOU VERY MUCH. (B)(6) |  | Y | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02791 | 2021-08-05 | 2021-08-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02809 | 2021-08-05 | 2021-03-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A (CPAP, BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING MAJOY PINK SLIME ISSUES WITH MACHINE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 WAS UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02850 | 2021-08-05 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBE/CHAMBER, PARTICLES IN AIRWAY FROM DEVICE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING PARTICLES IN TUBE/CHAMBER, PARTICLES IN AIRWAY FROM DEVICE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION G1 WAS CORRECTLY UPDATED. SECTION H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02753 | 2021-08-05 | 2021-06-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500S11 | DSX500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A SINUS INFECTION AND BRONCHITIS. THE REPORTED INFORMATION STATES THE PATIENT WAS PRESCRIBED ANTIBIOTICS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A SINUS INFECTION AND BRONCHITIS. THE REPORTED INFORMATION STATES THE PATIENT WAS PRESCRIBED ANTIBIOTICS. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02637 | 2021-08-05 | 2021-07-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/HT/CELL, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AUX400T15C | AUX400T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED MOUTH CANCER. THE PATIENT REQUIRED MEDICAL INTERVENTION IN THE FORM OF A DOCTOR'S ASSESSMENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE MOUTH CANCER. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02703 | 2021-08-05 | 2021-07-30 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO, FR | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | FRX500S14 | FRX500S14 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT WAS DIAGNOSED WITH NODULES IN THE THYROID. THE PATIENT DID REQUIRE MEDICAL INTERVENTION IN THE FORM OF THYROID ISTHMUSECTOMY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON OCT 12, 2023 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT WAS DIAGNOSED WITH NODULES IN THE THYROID. THE PATIENT DID REQUIRE MEDICAL INTERVENTION IN THE FORM OF THYROID ISTHMUSECTOMY. THE DEVICE WAS RETURNED TO THE THIRD PARTY SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. DURING THE EVALUATION, THE DUST/DIRT CONTAMINATION WAS OBSERVED. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED. THERE WERE 2 ERRORS FOUND. THE THIRD PARTY SERVICE CENTER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. DEVICE WAS SCRAPPED. SECTIONS D8, D9, H2, H3 AND H6 HAS BEEN UPDATED IN THIS REPORT. THE MANUFACTURER INADVERTENTLY CAPTURED B1 AND H1 IN PREVIOUS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02704 | 2021-08-05 | 2021-08-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED VERTIGO ISSUES AND ISCHEMIC OPTIC NEUROPATHY. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM A DOCTORS ASSESSMENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 14, 2024, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED VERTIGO ISSUES AND ISCHEMIC OPTIC NEUROPATHY. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM A DOCTORS ASSESSMENT.THE PATIENT DID NOT REPORT TO RECEIVE MEDICAL INTERVENTION. SECTION B1 WAS CORRECTED FOR BOTH ADVERSE EVENT AND PRODUCT PROBLEM. (PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION. TO SERIOUS INJURY. SECTION H6 WAS CORRECTED AND UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02769 | 2021-08-05 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02770 | 2021-08-05 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF BLACK PARTICLES, THE END USER INDICATED OZONE OR OTHER UNAPPROVED CLEANING AND DISINFECTION METHOD WAS USED RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM.THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL.THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION.THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME.IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. SECTION A1 AND G1 HAS BEEN CORRECTED ADN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02771 | 2021-08-05 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED DUE TO AGE. SECTION D8, D9 AND H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02772 | 2021-08-05 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02821 | 2021-08-05 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN CHAMBER/TUBING RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02822 | 2021-08-05 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/WIFI, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX400H11W | DSX400H11W |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DETAILS ON THE SERIAL NUMBER OF THE DEVICE REMAIN ARE CURRENTLY UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02833 | 2021-08-05 | 2021-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/ RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS VISIBLE FOAM DEGRADATION AND UNIT SCRABBED DUE TO AGE. | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103041 | 2021-08-05 | 2019-08-07 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILLIPS RESPIRONIC DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS RESPIRONIC / RESPIRONICS, INC. |  |  |  | INFLAMMATION HEADACHES; I HAVE A PHILLIPS RESPIRONIC FIRST GENERATION DREAMSTATION BIPAP SINCE 2018 AND WAS NEVER PREVIOUSLY WARNED ABOUT THE RISKS OF USING A SO CLEAN DEVICE TO CLEAN THE BIPAP. I AM CONCERNED THAT NEITHER THE RESPIRATORY THERAPIST, NOR DOCTOR THAT AUTHORIZED THE NEED FOR BIPAP THERAPY, GAVE A THOROUGH VIEWPOINT OF RISK USING THE CLEANING METHOD. I USED THE SO CLEAN IN GOOD FAITH THAT I WAS DOING ALL THAT I COULD TO KEEP ME AND LOVED ONES SAFE FROM CATCHING COVID THAT LIVED IN MY HOME. I DID USE THE CLEANING METHOD NO MORE THAN 5-6 TIMES A YEAR. I AM CONCERNED NOW I HAVE NOTICED SOME SWELLING AND INFLAMMATION IN AND AROUND CHEST AREA AND UNSURE WHY. I AM CONCERNED THAT PHILLIPS RESPIRONIC HAS NOTED THAT POSSIBLY THE WARRANTY MAY BE NULLIFIED IF CUSTOMERS HAVE USED THE OZONE CLEANING METHOD THAT THEY WERE NEVER FULLY AWARE OF THE RISKS TO BEGIN WITH. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103037 | 2021-08-05 | 2021-06-15 | I, n, j, u, r, y | Voluntary report | PHILLIPS DREAMSTATION CPAP DEVICE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | PERSISTENT COUGH, CHOKING FROM PHILLIPS CPAP MACHINE. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02782 | 2021-08-05 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES, HEADACHES, TROUBLE BREATHING THROUGH OUT THE NIGHT, DIFFICULTY BREATHING AND SHORT OF BREATH. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRABBED DUE TO (B)(6) 2017. CORRECTED INFORMATION PROVIDED IN B.5. (THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES, HEADACHES, TROUBLE BREATHING THROUGH OUT THE NIGHT, DIFFICULTY BREATHING AND SHORT OF BREATH. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY WAS NOT INCLUDED IN INITIAL MDR), CORRECTED INFORMATION PROVIDED IN H6 (HEAACHE AND DIFFICULTY BREATHING WAS NOT CAPTURED IN INITIAL MDR). | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02793 | 2021-08-05 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02780 | 2021-08-05 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. AFTER THE SECOND ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, CUSTOMER COULD NOT ABLE TO PROVIDE THE INFORMATION ABOUT DEVICE RETURN. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02699 | 2021-08-05 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX400T11 | DSX400T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP BREAST CANCER. THE PATIENT HAD SURGERY IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP BREAST CANCER. THE PATIENT HAD SURGERY IN RESPONSE TO THE REPORTED EVENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. DURING EVALUATION FOUND EVIDENCE OF DUST/DIRT CONTAMINATION INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM WAS OBSERVED THROUGHOUT DEVICE ENCLOSURE, AND AIRPATH, SUGGESTING A SOURCE EXTERNAL TO THE DEVICE. SLIGHT BLACK CONTAMINATION AT THE BLOWER SEAL OF THE BLOWER BOX IS CONSISTENT WITH THE KERATIN CONTAMINATION OBSERVED . A FIBROUS, HAIR-LIKE CONTAMINANT WAS FOUND INSIDE THE BLOWER BOX. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERORRS. THE MANUFACTURER CONCLUDES THE DEVICE HAS DUST/DIRT CONTAMINATION CONSISTENT WITH WATER INGRESS. THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION D8, D9 AND H6 HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02792 | 2021-08-05 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02787 | 2021-08-05 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SET PRESSURE DELIVERY WAS NOT ADEQUATE. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN EVALUATED BY THE THIRD PARTY SERVICE CENTER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE THIRD PARTY INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A VENTILATOR'S SET PRESSURE DELIVERY WAS NOT ADEQUATE. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. DEVICE HAD A SCRATCHED SCREEN. BUTTONS ARE WORN. CRACKED BY SD DOOR. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02790 | 2021-08-05 | 2021-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, A FAILURE OF THE EXTERNAL FLOW SENSOR WAS OBSERVED. THE DEVICE'S EXTERNAL FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02789 | 2021-08-05 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, A VENTILATOR'S PRESSURE DELIVERY WAS DECREASED, AND THE DEVICE FAILED TEST STEPS DURING TESTING. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE'S SENSOR BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02784 | 2021-08-05 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A MALFUNCTION OF THE VENTILATOR'S EXTERNAL FLOW SENSOR OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A MALFUNCTION OF THE VENTILATOR'S EXTERNAL FLOW SENSOR OCCURRED. THE DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S EXTERNAL FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02794 | 2021-08-05 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02786 | 2021-08-05 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S OXYGEN CELL WAS RECONNECTED AND THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02785 | 2021-08-05 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, THE DEVICE'S SET OXYGEN PERCENTAGE DID NOT MATCH THE MEASURED OXYGEN LEVEL. THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02761 | 2021-08-05 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A BENIGN MASS IN HIS LOWER THROAT WITH DIFFICULTY BREATHING, CHEST PAIN, HEADACHES, NAUSEA, AND SORE THROAT/SINUSES. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02762 | 2021-08-05 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02764 | 2021-08-05 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT ALSO STATES HAS BEEN EXPERIENCING NAUSEA, ACID REFLUX, VOMITING, MUCOUS IN THE THROAT, SICKNESS AFTER USING THE MACHINE, LOSS OF APPETITE, WEIGHTLOSS, NUTRIENT DEFICIENCY, CRAMPS, DEHYDRATION, NASAL BURNING WHILE USING MACHINE, HEADACHES, EXHAUSTION, PALPITATIONS, FATIGUE, AND HIGH BLOOD PRESSURE. PATIENT STATES THAT SHE ALSO NOTICED THAT THE FILTER GOT DIRTIER MUCH FASTER THAN NORMAL, AND A FEW BLACK THREADS IN THE RESERVOIR. SHE HAS ALSO NOTICED A CLICKING SOUND IN THE MACHINE FROM TIME TO TIME. PATIENT STATES THAT HER LIVER ENZYMES ARE HIGH. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT ALSO STATES HAS BEEN EXPERIENCING NAUSEA, ACID REFLUX, VOMITING, MUCOUS IN THE THROAT, SICKNESS AFTER USING THE MACHINE, LOSS OF APPETITE, WEIGHTLOSS, NUTRIENT DEFICIENCY, CRAMPS, DEHYDRATION, NASAL BURNING WHILE USING MACHINE, HEADACHES, EXHAUSTION, PALPITATIONS, FATIGUE, AND HIGH BLOOD PRESSURE. PATIENT STATES THAT SHE ALSO NOTICED THAT THE FILTER GOT DIRTIER MUCH FASTER THAN NORMAL, AND A FEW BLACK THREADS IN THE RESERVOIR. SHE HAS ALSO NOTICED A CLICKING SOUND IN THE MACHINE FROM TIME TO TIME. PATIENT STATES THAT HER LIVER ENZYMES ARE HIGH. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS UPDATED REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02767 | 2021-08-05 | 2021-06-22 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON OCT 25, 2024 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF RED SUBSTANCE AND BLACK PARTICLES. THE PATIENT ALSO ALLEGED OF HAVING SHORTNESS OF BREATH/DIFFICULTY BREATHING, INCREASED BLOOD PRESSURE, NASAL CONGESTION, SINUSES CONGESTION, AND ASTHMA RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY THE PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECTS: CLINICAL CODES AND HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02768 | 2021-08-05 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED AN ALLEGATION OF FILTERS TURNING BLACK AND THERE WERE BIG BLACK FLAKES FLOATING IN WATER. THE PATIENT ALSO EXPERIENCED HEADACHE, DIZZINESS AND NASAL SORE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS UPDATED REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. SECTION(S) A1(PATIENT INITIALS), E1 (INITIAL REPORTER-NAME AND ADDRESS), H6 (CLINICAL CODES) HAS BEEN CORRECTED IN THIS REPORT WHICH WAS MISSED TO CAPTURE IN INITIAL REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02619 | 2021-08-05 | 2021-07-12 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION BIPAP AVAPS30 H, DS | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | DSX1130H11 | DSX1130H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA, THYROID ISSUE AND BLOOD CLOTS IN THEIR LUNGS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS : THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED ASTHMA, THYROID ISSUE AND BLOOD CLOTS IN THEIR LUNGS. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENT OF ASTHMA, THYROID ISSUE AND BLOOD CLOTS IN THEIR LUNGS AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURER'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE INFORMATION AVAILABLE, THE MANUFACTURER CONCLUDES NO FURTHER ACTION IS NECESSARY. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. THE PATIENT HAS SWAPPED OUT THE DEVICE AT THE DISTRIBUTER AND AT THIS TIME NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILLED. SECTION B1, B2 HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT-IMPACT CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS AND SECTION H10 HAVE BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02783 | 2021-08-05 | 2021-03-08 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/HT/CELL, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | AUX400T15C | AUX400T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A RESPIRATORY INFECTION. THE PATIENT WAS PRESCRIBED AMOXICILLIN IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | SECTION G1 WAS MISSED TO CAPTURE IN THE FOLLOW-UP 2, WHICH WAS UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02779 | 2021-08-05 | 2021-06-26 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION GO AUTO W/BT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSG500S11 | DSG500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103047 | 2021-08-05 | 2021-07-01 | I, n, j, u, r, y | Voluntary report | RESPIRONICS DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500H11C |  |  | I HAVE BEEN USING THE PHILLIPS DREAMSTATION CPAP. I WAS RECENTLY DIAGNOSED WITH MULTIPLE MYELONOMA. I WAS ENCOURAGED BY MY DOCTOR TO FILL AN INCIDENT REPORT SINCE THERE IS A RECALL FOR THIS MACHINE RELATED TO THE OFF-GAS OF CERTAIN CHEMICALS THAT CAN BE LIFE THREATENING. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103091 | 2021-08-06 | 2020-10-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS / RESPIRONICS, INC. | DREAMSTNAUTOBIPAPHUMCELL DOM |  |  | COUGHING, SHORTNESS OF BREATH; MY MEDICAL DEVICE OF CONCERN HAS BEEN VOLUNTARILY RECALLED--MY PHILIPS BRAND BIPAP MACHINE FOR THE TREATMENT OF SLEEP APNEA. I KNOW THAT THE FDA MUST KNOW THAT PHILIPS IS SAYING A REPLACEMENT OR REPAIR OF THE MILLIONS OF DEFECTIVE MACHINES WILL TAKE A YEAR OR LONGER. IN THE MEANTIME, A PATIENT'S ONLY CHOICE IS TO PURCHASE A NEW (B)(6) MACHINE OR WAIT THE YEAR OR LONGER FOR A REPLACEMENT. MEDICAL INSURANCE, PRIVATE OR PUBLIC, FOLLOWS A FIVE-YR REPLACEMENT PLAN FOR CPAP AND BIPAP MACHINES. IF PHILIPS DOESN'T REPLACE, SPECIFICALLY, MY MACHINE BY OCTOBER 2023, THEN I CAN GO AHEAD WITH ORDERING A REPLACEMENT THEN. I WAS TOLD THIS, TODAY, BY THE MEDICAL SUPPLY COMPANY. I WAS TOLD BY A REPRESENTATIVE OF MY PULMONARY AND SLEEP MEDICINE CLINIC, TODAY, THAT 3.5 MILLION OTHER PEOPLE NEED REPLACEMENT MACHINES, TOO, SO I CAN BUY ONE LIKE MANY PATIENTS HAVE DONE. I CAN'T AFFORD TO BUY A REPLACEMENT MACHINE AS A MEDICAID PATIENT. I ALSO CANNOT AFFORD TO PAY FOR A CT SCAN OF MY CHEST, NECK, AND HEAD, BUT I WOULD LIKE TO BOTH VERIFY ANY DAMAGE FROM MY USE OF THE DEFECTIVE PHILIPS BIPAP AS WELL AS SET BASELINE CONDITIONS TO MARK ANY FUTURE CHANGES FROM THE CARCINOGENIC AFFECT OF MELTING INSULATION BEING FANNED INTO MY LUNGS. MY PRIMARY CARE PROVIDER AND THE SLEEP MEDICINE PROVIDERS HAVE BOTH ACKNOWLEDGED THE SERIOUSNESS OF THE SITUATION, YET NEITHER WILL ORDER SCANS TO FOR ME TO SCREEN THE EFFECTS OF THE MACHINE. NEITHER WILL ORDER THE TWO COMMON SEROLOGICAL TESTS GIVEN IN OCCUPATIONAL EXPOSURE EVENTS WITH SIMILARLY AEROSOLIZED AND INHALED INSULATION. I WOULD WAKE UP CHOKING, TURNING TO WRENCHING FOLLOWED BY HAVING TO WASH MY EYES WITH A WARM WASHCLOTH TO REMOVE THE VISCOUS MUCUS BUILDUP FROM THE TOXIC AIR OFF GASSING INTO MY EYES ('TOXINS AND CHEMICAL OFF-GASSING' ARE WORDS I TOOK VERBATIM FROM THE RECALL LETTER). IF MY EYES RESPONDED TO THE AIR THIS WAY, WHAT DID THE AIR DO TO MY THROAT, SINUSES, AND LUNGS? I WANT TO KNOW. I AM GOING TO SALE THINGS ONLINE AND AT THE PAWNSHOP SO I CAN PAY FOR THE PRIVILEGE OF USING A PHILIPS BRAND COMPUTER TOMOGRAPHY SCANNER--THE IRONY IS NOT FUNNY. I WILL DO THIS, ILLEGALLY, AS I WILL SELF-REFER AND LIE ABOUT NOT HAVING INSURANCE COVERAGE. PHILIPS CORP. HAD A 2019 GROSS REVENUE OF $1.58 BILLION AND I AM GOING TO HUSTLE TO GET TO USE, LIKELY, ONE OF THEIR CT SCANNERS TO DETERMINE IF THERE FAILURE TO ENGINEER SAFE MATERIALS HAS IRREPARABLY ENGINEERED MY FATE. I AM SO TIRED SINCE I HAVE NOT BEEN REGULARLY USING MY BIPAP SINCE (B)(6). NOW I KNOW WHY, THEY SAY NOT TO USE IT, BUT THEY WON'T REPLACE IT UNLESS I GIVE THEM ONE THOUSAND DOLLARS. ANY MONEY I HAVE IS GOING TO PAY FOR MY HEALTH SCREENINGS THAT MY DOCTORS WON'T ORDER. I WILL CONTINUE WITH MY 3-4 HRS OF SLEEP, IN HOURLY INCREMENTS, UNTIL I CAN FIGURE OUT WHAT TO DO TO HELP MYSELF. NO TESTS HAVE BEEN DONE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103080 | 2021-08-06 |  | I, n, j, u, r, y | Voluntary report | PHILLIPS DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS / RESPIRONICS INC. | DSX500H11C |  |  | I HAVE PHILLIPS CPAP MACHINE I HAD SINUS PROBLEMS AND EXTREME NAUSEA EVERY MORNING UNTIL I STOPPED USING THE MACHINE I THINK IT IS STILL AFFECTING MY BREATHING. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02724 | 2021-08-06 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "LOW OXYGEN FLOW" ALARM CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02836 | 2021-08-06 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. THE DEVICE HAD EVIDENCE OF WATER INGRESS. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02838 | 2021-08-06 | 2021-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02840 | 2021-08-06 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | FR2110X14B | FR2110X14B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S BLOWER MOTOR WAS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S MACHINE FLOW SENSOR NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR'S MACHINE FLOW SENSOR NEEDED TO BE REPLACED TO ADDRESS A "VENTILATOR INOPERATIVE" ALARM CONDITION. THE COMPONENT REPLACEMENT DID NOT CORRECT THE ISSUE. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02842 | 2021-08-06 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02835 | 2021-08-06 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1032804 | 1032804 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLANK. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE WAS EVALUATED AT THE MANUFACTURER'S SERVICE CENTER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02843 | 2021-08-06 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S POWER SUPPLY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02837 | 2021-08-06 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR KEYPAD FUNCTION FROZE. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S FRONT PANEL/LED PCA WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02845 | 2021-08-06 | 2021-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02839 | 2021-08-06 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02444 | 2021-08-06 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054097 | 1054097 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ON FIRE. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF A VENTILATOR THAT WAS ON FIRE. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION AND THERMAL DAMAGE WAS CONFIRMED. DURING THE INVESTIGATION, THE MANUFACTURER VISUALLY INSPECTED THE DEVICE AND FOUND PHYSICAL DAMAGE TO THE POWER CORD AT THE STRAIN RELIEF. PRODUCT LABELING STATES "PERIODICALLY INSPECT THE POWER CORD FOR DAMAGE OR SIGNS OF WEAR. DISCONTINUE USE AND REPLACE IF DAMAGED". ALL EVIDENCE OF FIRE DAMAGE WAS EXTERNAL TO THE DEVICE. THERE WAS NO INTERNAL DAMAGE FOUND. THE DEVICE WAS POWERED ON USING BATTERY POWER AND OPERATED NORMALLY. THE DEVICE'S EVENT LOG WAS REVIEWED AND THERE WERE NO ERROR CODES NOTED AT THE STATED TIME OF THE EVENT. THE MANUFACTURER CONCLUDES THAT THE THERMAL EVENT OCCURRED DUE TO ARCING FROM THE DAMAGED POWER CORD IN AN OXYGEN RICH ENVIRONMENT. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02734 | 2021-08-06 | 2021-07-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE REPORTER OF THE EVENT HAS CONFIRMED THE DEVICE WILL NOT BE RETURNED FOR EVALUATION OR REPAIR. | IT WAS INITIALLY REPORTED, THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE REPORTER OF THE EVENT HAS CONFIRMED THE DEVICE WILL NOT BE RETURNED FOR EVALUATION OR REPAIR. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02735 | 2021-08-06 | 2021-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, LEAKAGE INSIDE THE OXYGEN BLENDING MODULE HOUSING WAS DETECTED. THE OXYGEN BLENDING MODULE MANIFOLD ASSEMBLY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02646 | 2021-08-06 | 2021-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | GBX500H15 | GBX500H15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE CHILD'S MUCUS IN HIS NOSE WAS BLACK IN THE MORNING. THERE WAS NO REPORT OF SERIOUS HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. THE EXTERNAL PART OF THE DEVICE WAS VISUALLY INSPECTED BY THE MANUFACTURER AND FOUND NO SIGNS OF CONTAMINATION. THE INTERNAL PART OF THE DEVICE WAS VISUALLY INSPECTED BY THE MANUFACTURER AND FOUND EVIDENCE OF CONTAMINATION ON THE BOTTOM PANEL AND ON THE INLET FILTER OF THE BLOWER BOX. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE WAS APPLIED POWER AND AIRLFLOW WAS VERIFIED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS LOGGED. THE MANUFACTURER CONCLUDES THE DEVICE HAS EVIDENCE OF CONTAMINATION WITHIN THE DEVICE. THE MANUFACTURER CONFIRMS THERE IS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION WITHIN THE DEVICE. THE PREVIOUS REPORT DID NOT MENTION THE CHILD HAVING MUCUS IN THEIR NOSE. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02647 | 2021-08-06 | 2021-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | GBX500H15 | GBX500H15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103087 | 2021-08-06 |  | I, n, j, u, r, y | Voluntary report | DREAMSTNAUTOCPAPHUMCELL DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS INC / RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | BILATERAL CORNEAL INFLAMMATION; WITHIN ABOUT 1 YEAR OF BEING SWITCHED FROM RESMED S9 CPAP TO PHILIPS DREAMSTATION, MY EYES WOULD GET IRRITATED AND DRY OUT. I SOUGHT MEDICAL ATTENTION AND WAS TOLD I HAD INFLAMMATION OF THE CORNEA AND DRY EYE WHICH REQUIRED TREATMENT, WHICH IS ONGOING. COINCIDENTAL OR RELATED IS UNDETERMINED AT THIS TIME. ADDITIONALLY, BLACK PARTICULATES WOULD SPORADICALLY ACCUMULATE IN MY NOSTRILS; AND A SWEET PUNGENT ODOR WOULD OCCUR ESPECIALLY WHEN MY DISTILLED WATER IN THE HUMIDIFIER RESERVOIR GOT LOW. WITHIN A FEW MONTHS OF BEING ISSUED THE "NEW" PHILIPS MACHINE I TOOK IT BACK TO THE DME PROVIDER WITH THE COMPLAINT IT WAS EXCESSIVELY NOISY COMPARED TO THE RESMED UNIT AND HAD A LIKE-NEW ODOR PRESENT WHEN IN USE. THE RESPIRATORY TECH MADE AN AIRFLOW ADJUSTMENT AND REPLACED THE HUMIDIFIER GASKET. THEY ALSO ACKNOWLEDGED THERE WAS YELLOWING OF THE GASKET LIKELY CAUSED BY THE SOCLEAN MACHINE. THEY NEVER STATED I SHOULD STOP USING THE OZONE CLEANING METHODS/PROTOCOL I WAS USING UNTIL I CONTACTED THEM IN (B)(6) 2021 WITH COMPLAINTS ABOUT THE PARTICULATES AND ODORS. I WAS ADVISED TO CALL MY PHYSICIAN AND CONTACT PHILIPS TO FIND ALTERNATIVES TO CPAP THERAPY. THE WEBSITE FOR THE PHILIPS RECALL ACKNOWLEDGED MY UNIT WAS ONE OF THOSE AFFECTED BY THE RECALL SO I REGISTERED IT ON (B)(6) 2021. PHILIPS SAID THEY COULD NOT TELL ME TO STOP USING THE UNIT, ONLY MY PHYSICIAN COULD MAKE THAT DECISION, AND BASED ON MY COMPLAINTS I WAS ADVISED TO STOP. WITHOUT CPAP I SUFFER EXCESSIVE SNORING AND NIGHTLY APNEA EPISODES. I GET REFERRED AMONG THE PHYSICIAN, PHILIPS AND MEDLOGIC WITHOUT RESOLUTION. INSURANCE WILL NOT AUTHORIZE A REPLACEMENT UNIT, PHILIPS WILL NOT OFFER A REPLACEMENT AND MY PHYSICIAN IS PLACED IN A MEDICAL-LEGAL CONUNDRUM IN WHICH THEY CANNOT RECOMMEND CONTINUING USE OF THE AFFECTED CPAP MACHINE, WHILE AT THE SAME TIME ACKNOWLEDGING THE NEED FOR NIGHTLY THERAPY. SUBSEQUENTLY I WAS ADVISED TO SUBMIT THIS REPORT FOR DOCUMENTATION AND POSSIBLE SUPPORT FOR TIMELY RESOLUTION. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103088 | 2021-08-06 | 2020-10-09 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS CPAP RNTCPAP5 | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS / RESPIRONICS, INC. | RNTCPAP DREAMSTATION |  |  | BLOOD CANCER - MPN; THE PHILIPS RESPIRONICS CPAP DEVICE THAT I USED SINCE (B)(6) 2017 FOR SLEEP APNEA HAS ISSUED A RECALL, DUE TO POTENTIAL RISKS RELATED TO THE FOAM USED TO REDUCE THE SOUND AND VIBRATION OF THE MOTOR. ONE OF THE RISKS IS POTENTIAL CANCER CAUSED FROM THE TOXIC GASES. I WAS DIAGNOSED WITH MYELOPROLIFERATIVE NEOPLASM IN (B)(6) 2020. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103090 | 2021-08-06 | 2019-08-21 | I, n, j, u, r, y | Voluntary report | CPAP DEVICE DREAM STATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DREAM STATION | V1.18 |  | I HAVE BEEN TO A PULMONOLOGIST FOR HAVING COMPLAINTS OF PNEUMONIA THREE TIMES, SHORTNESS OF BREATH AND FEELING LIKE I AM CHOKING WITH MY PHILIPS CPAP DEVICE. I STARTED USING AN INHALER AND AZITHROMYCIN AS NEEDED WHEN I START HAVING SYMPTOMS THAT MAY LEAD TO PNEUMONIA. PULMONOLOGIST COULD NOT FIGURE OUT THE REASON FOR ALL MY SYMPTOMS AND CALLED IT "BAD LUCK". IT MAY BE RELATED TO MY CPAP DEVICE THAT HAS RECENTLY BEEN RECALLED AS I HAVE SLEEP APNEA DUE TO A DEVIATED SEPTUM. I NOW USE INHALER ANORO AND AZITHROMYCIN AS NEEDED. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-08
2518422-2021-02841 | 2021-08-06 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
MW5103095 | 2021-08-06 | 2019-02-20 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | RECEIVED A NEW PHILIPS DREAMSTATION IN (B)(6) 2019. ABOUT A MONTH LATER, WENT TO A HOTEL AND WAS IN THE HOT TUB. THAT EVENING, HAD SEVERE SINUS REACTION THAT LASTED 2-3 DAYS. EVER SINCE THEN, I'VE HAD A DRY COUGH THAT AFFECTS ME EVERY DAY. I WENT TO AN ENT TO LOOK INTO THE CAUSE. SUGGESTED IT MAY BE DUE TO REFLUX, SO OVER THE LAST COUPLE YEARS I'VE GONE IN FOR MULTIPLE TESTS TO DETERMINE THE EXTENT OF THE REFLUX, EVEN GOING SO FAR AS TO CONSIDER A RISKY PROCEDURE TO CONTROL IT- ALL TO TRY TO ALLEVIATE THE COUGH THAT I'VE BEEN DEALING WITH. I HAVE HAD SINUS ISSUES FOR YEARS, AND HAVE GOTTEN SINUS INFECTIONS OFF AND ON WHILE USING THE DEVICE. THE LAST FEW MONTHS, I'VE EXPERIENCED DAILY HEADACHES THAT I HAVE BEEN ATTRIBUTING TO MY SINUS ISSUES. I USE A SOCLEAN DEVICE WITH MY CPAP, AND EARLIER THIS WEEK WAS INFORMED OF A RECALL THAT PHILIPS HAS ISSUED FOR THE DEVICE THAT I USE INVOLVING SOCLEAN OZONE CLEANERS AND A FOAM THAT IS USED IN THE CPAP DEVICES. IT IS SAID TO CAUSE RESPIRATORY ISSUES, COUGHS, HEADACHES, ETC. I MAY HAVE NO WAY TO KNOW IF MY SYMPTOMS WERE RELATED TO A SINUS INFECTION AND REFLUX AS I'D ORIGINALLY THOUGHT, OR IF THEY WERE ALL ACTUALLY A RESULT OF THE USE OF MY CPAP DEVICE, BUT THE TIMING OF THE ISSUES WITH WHEN I STARTED TO USE THE DEVICE MAKES ME FEEL IT'S WORTH NOTING HERE JUST IN CASE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02844 | 2021-08-06 | 2021-07-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | EE2110X15B | EE2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF A VENTILATOR INOPERATIVE CONDITION OCCURRING. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02803 | 2021-08-06 | 2021-06-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY SUBMITTED MDR 2518422-2021-02803-3 WITH INCORRECT SECTIONS B1, B2, H1, H6. CORRECTIONS TO PREVIOUS MDR ARE MADE IN THIS REPORT AS FOLLOWS. SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY PRODUCT PROBLEM WAS CHECKED IN PREVIOUS MDR). SECTION B2 WAS CORRECTED TO LIFE-THREATENING (PREVIOUSLY IT WAS BLANK). SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6- HEALTH EFFECT - IMPACT CODE WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103078 | 2021-08-06 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS / RESPIRONICS INC. |  |  |  | PATIENT HAS BEEN USING A RECALLED CPAP MACHINE SINCE (B)(6) 2020. OVER THE LAST FEW MONTHS, HE HAS BEEN NOTICING BLACK PEPPER LIKE MATERIAL IN THE HUMIDIFIER BOX, AND TERRIBLE SMELLS COMING OUT OF THE MACHINE DURING USE. HE THOUGHT HE THREW UP IN THE MACHINE IT SMELLED THAT BAD. IN THE LAST COUPLE OF MONTHS HE HAS BEEN EXPERIENCING SEVERE HEADACHES, DIZZINESS, AND NOW VISION CHANGES. HE HAS BEEN REFERRED TO HIS PCP FOR FURTHER EVALUATION OF HIS SYMPTOMS OF CONCERN AND HAS REQUIRED DISCONTINUATION OF HIS CPAP THERAPY UNTIL HE CAN GET A REPLACEMENT MACHINE. WE HAVE NO TIMELINE ON REPLACEMENT. STRENGTH: 5-20 CWP, FREQUENCY:HS, DIAGNOSIS FOR USE: OSA, LEFT HEMIDIAPHRAGM PARALYSIS. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02877 | 2021-08-06 | 2021-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | ARX500S15 | ARX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. CORRECTED DATA: D8 WAS THIS DEVICE SERVICED BY A THIRD PARTY?: PREVIOUSLY REPORTED AS NO ; UPDATED TO UNKNOWN | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02856 | 2021-08-06 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500S11 | DSX500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING OR CHAMBER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED COUGH AND SEVERE SLEEP APNEA. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING OR CHAMBER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED COUGH AND SEVERE SLEEP APNEA. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. SECTION D8, D9, H7,H9 WAS CORRECTED IN THIS REPORT WHICH WAS MISSED IN INTIAL REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02855 | 2021-08-06 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | DREAMSTATION AUTO CPAP | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING NASAL/THROAT IRRITATION RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02866 | 2021-08-06 | 2020-05-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION EXPERT, FR | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | RFRX501S14 | RFRX501S14 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP THYMIC CARCINOMA. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. CORRECTED DATA: D1 UNIQUE IDENTIFIER (UDI) NUMBER: PREVIOUSLY REPORTED AS 00606959427815 ; UPDATED TO NOT AVAILABLE. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02517 | 2021-08-06 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF (B)(4). THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02863 | 2021-08-06 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN THE INITIAL REPORT SECTION B5 WAS NOT UPDATED, THE CORRECT B5 SHOULD BE - THE PATIENT HAS ALLEGED PARTICLES IN THE DEVICE AND IN FILTERS,HEADACHES, SINUS INFECTION, COUGH, TERRIBLE MUCUS, NASAL/THROAT IRRITATION OR SORENESS, CHEST PRESSURE, UPPER AIRWAY IRRITATION, RESPIRATORY TRACT IRRITATION. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF FINE DUST CONTAMINATION IN THE BLOWER AND THROUGHOUT THE AIRPATH, BLACK MARK ON THE FRONT PANEL ALIGNS WITH DAMAGE TO THE TOP ENCLOSURE, WHITE CONTAMINATION IN THE AIR INLET . THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. A HUMIDIFIER WAS ALSO RETURNED TO THE MANUFACTURER ALONG WITH DEVICE. AN INTERNAL VISUAL INSPECTION OF THE HUMIDIFIER WAS COMPLETED BY THE MANUFACTURER AND FOUND AN UNKNOWN BROWN CONTAMINATION IN THE CHAMBER INTERIOR NEAR THE WIRE COVER. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND 32 ERRORS LOGGED. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH FINE DUST CONTAMINATION IN THE BLOWER AND THROUGHOUT THE AIRPATH, BLACK MARK ON THE FRONT PANEL ALIGNS WITH DAMAGE TO THE TOP ENCLOSURE, WHITE CONTAMINATION IN THE AIR INLET AND AN UNKNOWN BROWN CONTAMINATION IN THE CHAMBER INTERIOR NEAR THE WIRE COVER OF THE HUMIDIFIER WAS INCONSISTENT WITH THE SOUND ABATEMENT FOAM. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION D9,G3,H6 HAS BEEN UPDATED/CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02862 | 2021-08-06 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103085 | 2021-08-06 | 2021-06-18 | I, n, j, u, r, y | Voluntary report | SERIES 1 BIPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS RESPIRONICS / RESPIRONICS, INC. | P120964114EC2E |  |  | EYES BEGAN EXCESSIVELY WATERING, TEARING UP. IS CURRENTLY ((B)(6) 2021) CONTINUING EVEN AFTER RECEIVING A PRESCRIPTION ANTI INFLAMMATORY EYE DROP FROM EYE DOCTOR. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02752 | 2021-08-06 | 2021-06-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP T CELL LYMPHOMA. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02804 | 2021-08-06 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SINUS INFECTIONS, STROKES, AND TRANSIENT ISCHEMIC EPISODES. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED ON THIS DEVICE IN MDR 2518422-2021-01891. THIS REPORT WAS SUBMITTED AS A DUPLICATE REPORT OF THE PREVIOUSLY SUBMITTED REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02813 | 2021-08-06 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02810 | 2021-08-06 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES AND NOT BEING ABLE TO TALK AFTER USE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02811 | 2021-08-06 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02812 | 2021-08-06 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING BLACK PARTICLES, NOSE BLEEDS, SORE THROAT AND COUGH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK PARTICLES, NOSE BLEEDS, SORE THROAT AND COUGH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02622 | 2021-08-06 | 2021-06-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG CANCER. THE REPORTED INFORMATION STATES THE PATIENT IS UNDERGOING CHEMOTHERAPY IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG CANCER. THE ABOVE REPORTED EVENT IS ASSESSED AS POSSIBLY RELATED TO THE DEVICE IN THIS CASE. HENCE, THE DEVICE MAY HAVE CAUSE OR CONTRIBUTE TO THE ALLEGED SERIOUS INJURY. THE REPORTED INFORMATION STATES THE PATIENT IS UNDERGOING CHEMOTHERAPY IN RESPONSE TO THE REPORTED EVENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AS PER THE INVESTIGATION OF THE DEVICE EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED IN THE BASE UNIT AND OPERATIONAL TESTING SHOWED THE UNIT PROVIDED AIRFLOW. AS PER SECONDARY FINDINGS OF THE BASE UNIT A FOREIGN WHITE POWDERY SUBSTANCE WAS FOUND THROUGHOUT THE BLOWER MOTOR BOX STEMMING FROM THE AIR INTAKE PORT TO THE AIR OUTLET PORT. THE FINS OF THE BLOWER MOTOR AND OUTLET PORT OF THE BLOWER MOTOR ARE WHITE, POTENTIAL DARK KERATIN PARTICLES WERE FOUND ON THE BLOWER BOX OUTLET. PRODUCT INVESTIGATION LABORATORY FOUND DARK KERATIN PARTICLES ON THE BLOWER BOX OUTLET PORT, FOREIGN WHITE SUBSTANCE THROUGHOUT THE BLOWER BOX ASSEMBLY OF THE DEVICE. PIL FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS. THE MANUFACTURER CONCLUDES UNABLE TO ADDRESS THE SYMPTOMS DESCRIBED AND CONFIRM THE PRESENCE OF CONTAMINATION IN THE AIRPATH. THE RESULTS OF THIS INVESTIGATION DO NOT IMPACT THE CALCULATED RISKS. IN THIS REPORT, SECTION D9, G3, H3, H6 HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02516 | 2021-08-06 | 2021-06-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP BASEL CELL PARSENOMA. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND B5 SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP BILATERAL BASEL CELL CARCINOMA DIAGNOSED 9-16-2020, SKIN REDDNESS, IRREGULAR SHAPED LESIONS THAT ARE DISCOLORED AND A BUMP ON THE LEFT SIDE OF THE NOSTRIL. THE PATIENT DID REPORT TO RECEIVE MEDICAL INTERVENTION AND SAW A DERMATOLOGIST, ONCOLOGIST AND HAD SURGERY AND SUPERFICIAL ELECTRON THERAPY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. SECTION G1 HAS BEEN CORRECTED WITH THE MOST CURRENT CONTACT OFFICE (AND MANUFACTURING SITE FOR DEVICES) OR COMPOUNDING OUTSOURCING FACILITY AND ADDRESS SECTION H6 WAS UPDATED WITH THE APPROPRIATE HEALTH EFFECT - CLINICAL CODE AND HEALTH EFFECT - IMPACT CODE. TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS WERE CORRECTED AS THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02778 | 2021-08-06 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED ON THIS DEVICE IN MDR 2518422-2021-02040. THIS REPORT WAS SUBMITTED AS A DUPLICATE REPORT OF THE PREVIOUSLY SUBMITTED REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02854 | 2021-08-06 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED BREATHING ISSUES, PARTICLES IN TUBING. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. PIL OBSERVED THE FOLLOWING FROM AN EXTERNAL AND INTERNAL INSPECTION: MISSING MODEM SIDE PANEL DOOR. VERY LITTLE UNKNOWN WHITE AND BLACK CONTAMINATION (DUST-LIKE), INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM, AT THE AIR INPUT OF THE BLOWER BOX. ALSO, A POTENTIAL TINY BLACK HAIR OR FIBER WAS AT THE INPUT OF THE BLOWER BOX. SLIGHT DUST-LIKE CONTAMINATION ON TOP OF THE BLOWER MOTOR. REAR PANEL O-RING WAS DEFORMED AND NOT FITTING PROPERLY. PIL USED A FITT (FOAM INTEGRITY TEST TOOL) AND THE ASSOCIATED PROCEDURE ER 2245460 V01 AND WAS NOT ABLE TO CONFIRM THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM. THERE IS NO VISIBLE DAMAGE OTHER THAN THE REAR PANEL O-RING, OR FUNCTIONALITY FAILURES OF THE DEVICE, WHICH SUGGEST THAT THE SOURCE OF CONTAMINATION WAS EXTERNAL TO THE DEVICE. PIL WAS NOT ABLE TO CONFIRM THE COMPLAINT OR ADDRESS THE SYMPTOMS SPECIFIED. PIL WAS NOT ABLE TO GET CARE ORCHESTRATOR INFORMATION FROM DEVICE. REVIEW OF THE DEVICE LOG SHOWED: -ERRORS LOGGED: 0 ERRORS WERE LOGGED. -DEVICE WAS LIKELY RESET PRIOR TO PIL RECEIPT AS INDICATED BY THE DEVICE HAVING 10.8 MACHINE HOURS AND NO BLOWER OR THERAPY HOURS. RISK TAG (ER 2219697 V20) ASSOCIATED WITH THIS MODE OF FAILURE: IRRIT02, INFECT01. THE RESULTS OF THIS INVESTIGATION DO NOT IMPACT THE CALCULATED RISKS. SECTION H6 HEALTH CLINICAL CODE HAS BEEN CORRECTED AND MEDICAL DEVICE PROBLEM CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103082 | 2021-08-06 |  | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS RESPIRONICS / RESPIRONICS, INC. |  |  |  | KEPT HAVING HEADACHES, THROAT PAIN AND PUS IN THROAT- REPORTED TO SEVERAL OF MY DOCTORS. WAS EVENTUALLY REFERRED TO ENT AND HAD A CHRONIC SINUS INFECTION/LARYNGOPHARYNGEAL REFLUX ACTING UP AT SAME TIME. NOTICED THAT EVERY MORNING FROM SLEEPING ON BIPAP I WAS HAVING THIS PROBLEM. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02982 | 2021-08-09 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DETAILS ON THE SERIAL NUMBER OF THE DEVICE REMAIN ARE CURRENTLY UNKNOWN. IF FURTHER INFORMATION REGARDING THE DEVICE BECOME AVAILABLE, IT WILL BE SUBMITTED IN THE FOLLOW-UP REPORT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES, CONGESTION RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02908 | 2021-08-09 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DEVICE STRANGE ODOR, FILTER ISSUE, PARTICLES IN TUBING/CHAMBER AND HEADACHE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. IN INITIAL REPORT, SECTION A1 AND H6 CLINICAL CODE WAS NOT CAPTURED AND SECTION G3 DATE WAS CAPTURED AS 07/14/2021, SHOULD BE 01/01/2021, CORRECTED AND UPDATED THESE FIELDS IN THIS REPORT. SECTION G1 AND H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02897 | 2021-08-09 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | RIN2110X15B | RIN2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VARIATION IN THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE DELIVERY WAS OBSERVED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN EVALUATED BY THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VARIATION IN THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE DELIVERY WAS OBSERVED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE VENTILATOR'S SOFTWARE WAS RE-LOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02834 | 2021-08-09 | 2021-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP A40, INTERNATIONAL | VENTILATOR, CONTINUOUS, MINIMAL VENTILATORY SUPPORT, FACILITY USE |  | 1111169 | 1111169 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02625 | 2021-08-09 | 2021-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO, INTL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | 761P | 761P |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02773 | 2021-08-09 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02774 | 2021-08-09 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX700T11 | DSX700T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN THE INITIAL REPORT SECTION B5 WAS NOT UPDATED, THE CORRECT B5 SHOULD BE - THE PATIENT HAS ALLEGED PARTICLES IN THE DEVICE, HEADACHE, UNDERWENT MRI AND CT SCAN. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER AND FOUND NO EVIDENCE OF CONTAMINATION. THE MANUFACTURER FOUND EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. . THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND ONE E-200 ON ERROR LOG. THE MANUFACTURER CONFIRMED THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION D9,G3,H6 HAS BEEN UPDATED/CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02775 | 2021-08-09 | 2021-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02754 | 2021-08-09 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING COUGH AND CHEST HURTS SHE HAS BEEN SEEN BY THREE DIFFERENT PULMONOGIST AND SEEN BLACK AND WHITE PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING COUGH AND CHEST HURTS SHE HAS BEEN SEEN BY THREE DIFFERENT PULMONOGIST AND SEEN BLACK AND WHITE PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, FOLLOWUP REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02777 | 2021-08-09 | 2021-07-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION ST30, GB | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS INC | GBX1030S20 | GBX1030S20 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02624 | 2021-08-09 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING BLACK PARTICLES, HEATING PLATE IS NOT WORKING PROPERLY AND SINUSES WERE VERY IRRITATING RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. SCRATCHES ON UPPER ENCLOSURE AND UI PANEL WAS OBSERVED. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE 5 ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. SECTIONS D8, D9, H2, H3 AND H6 WERE UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02518 | 2021-08-09 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL, AU | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | AUX500T15C | AUX500T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02776 | 2021-08-09 | 2021-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02859 | 2021-08-09 | 2021-07-13 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON AUGUST 05, 2024 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING CHAMBER, DIFFICULTY IN BREATHING, HEADACHE, SINUS INFECTION, CHEST PAIN, SORE THROAT, HEART PROBLEM, DRYNESS, HARD TO COUGH AND LUNG DISEASE RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02899 | 2021-08-09 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | RIN2110X15B | RIN2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VARIATION IN THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE DELIVERY WAS OBSERVED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN EVALUATED BY THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VARIATION IN THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE DELIVERY WAS OBSERVED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE VENTILATOR'S SOFTWARE WAS RE-LOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02933 | 2021-08-09 | 2021-07-13 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION EXPERT W/HUM/HT/CELL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | CAX501T12C | CAX501T12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT SWELLING AND REDNESS IN THE LEFT EYE. THE PATIENT HAS RECEIVED MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION FOR STEROID. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A PATIENT DEVELOPED SWELLING AND REDNESS IN THE LEFT EYE, BURSITIS IN LEFT ELBOW RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS RECEIVED MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION FOR STEROID. BASED ON INFORMATION AVAILABLE AT THIS TIME, THERE IS NOT ENOUGH EVIDENCE TO SUPPORT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE ALLEGED SERIOUS INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 HAS BEEN UPDATED IN THIS REPORT. SECTION D4 UDI NUMBER HAS BENN CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02861 | 2021-08-09 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX400H11C | DSX400H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02967 | 2021-08-09 | 2021-07-02 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID/HT, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | CAX400T12 | CAX400T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP DIFFICULTY BREATHING AND SHORTNESS OF BREATH. THE PATIENT WAS HOSPITALIZED DUE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. CORRECTED DATA: E1 ADDRESS (COUNTRY): PREVIOUSLY REPORTED AS USA ; UPDATED TO CANADA. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02848 | 2021-08-09 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02846 | 2021-08-09 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP AUTO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500S11 | DSX500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON MAY 08, 2023, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER PREVIOUSLY REPORTED THAT A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM ALLEGEDLY BECAME DEGRADED. TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGATION OF VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTION A1 AND B7 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02934 | 2021-08-09 | 2021-08-03 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION EXPERT W/HUM/HT/CELL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX501T12C | CAX501T12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE LUNG TUMOR. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE REPORTED EVENT IS CURRENTLY UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT ALLEGED TO HAVE LUNG TUMOR. NO MEDICAL INTERVENTION WAS REPORTED. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. NO ERRORS WERE FOUND. THE MANUFACTURER CONCLUDES THAT THE CUSTOMERS ALLEGATION OF SOUND ABATEMENT FOAM COULD NOT BE CONFIRMED AND THERE WAS NO VISIBLE FOAM DEGRADATION. SECTION D4 UNIQUE IDENTIFIER (UDI) # CORRECTED. SECTION G1 UPDATED. SECTION D8, D9, H2, H3 AND H6 HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02963 | 2021-08-09 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THE END USER INDICATED OZONE OR OTHER UNAPPROVED CLEANING AND DISINFECTION METHOD WAS USED. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02962 | 2021-08-09 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED DIFFICULTY BREATHING/SHORT OF BREATH. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02965 | 2021-08-09 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700S11 | DSX700S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES AIRWAY FROM DEVICE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02958 | 2021-08-09 | 2021-06-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED PATIENT STATED PARTICLES IN TUBING/CHAMBER, DIFFICULTY BREATHING/SHORT OF BREATH, STATED FILTERS COMES OUT MORE DARKER THEN BEFORE WHEN CLEANING AND COUGHING. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PATIENT STATED PARTICLES IN TUBING/CHAMBER, DIFFICULTY BREATHING/SHORT OF BREATH, STATED FILTERS COMES OUT MORE DARKER THEN BEFORE WHEN CLEANING AND COUGHING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. CORRECTIONS WERE MADE IN B3, B5 SECTIONS AND H6 SECTION WAS UPDATED IN THIS REPORT. CORRECTION TO B3: THE DATE OF EVENT WAS NOT REPORTED IN BOTH THE INITIAL AND FOLLOW UP REPORT IN B3 SECTION. THE DATE OF EVENT SHOULD BE REPORTED AS 06/08/2021. CORRECTION TO B5: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PATIENT STATED PARTICLES IN TUBING/CHAMBER, DIFFICULTY BREATHING/SHORT OF BREATH, STATED FILTERS COMES OUT MORE DARKER THEN BEFORE WHEN CLEANING AND COUGHING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02964 | 2021-08-09 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02906 | 2021-08-09 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS INITIALLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN THE INITIAL REPORT SECTION B5 WAS NOT UPDATED, THE CORRECT B5 SHOULD BE - THE PATIENT HAS ALLEGED FILTERS ARE GETTING BLACK, DIFFICULTY BREATHING/SHORT OF BREATH. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND BEIGE DUST OR DIRT CONTAMINATE IN FILTER INTAKE AREA AND BLOWER AREA, MODERATE BROWN CONTAMINATION IN CRACKS AND CREVASSES OF UPPER AND LOWER ENCLOSURE, DRIED BROWN LIQUID WITH TWO ADDITIONAL SMALLER SPOTS ON THE LOWER ENCLOSURE INSIDE, DARK SHEEN CONTAMINATE ON THE BACK OF THE UI PANEL, CORROSION, MINOR AMOUNT OF KERATIN ON BLOWER BOX OUTLET PORT .THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE CUSTOMER'S COMPLAINT OF THE FILTER TURNING BLACK WAS NOT CONFIRMED. THE FILTER WAS NOT RETURNED FOR EVALUATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND 3 PRE-TEST ERRORS WERE LOGGED WHILE SETTING UP THE UNIT FOR TESTING AND TWO INSTANCES OF E 53 PRIOR TO THAT. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH BEIGE DUST OR DIRT CONTAMINATE IN FILTER INTAKE AREA AND BLOWER AREA, MODERATE BROWN CONTAMINATION IN CRACKS AND CREVASSES OF UPPER AND LOWER ENCLOSURE, DRIED BROWN LIQUID WITH TWO ADDITIONAL SMALLER SPOTS ON THE LOWER ENCLOSURE INSIDE, DARK SHEEN CONTAMINATE ON THE BACK OF THE UI PANEL, CORROSION, MINOR AMOUNT OF KERATIN ON BLOWER BOX OUTLET PORT. THE CONTAMINATES WERE SOURCED FROM EXTERNAL ENVIRONMENTAL CONDITIONS. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION D9,G3,H6 HAS BEEN UPDATED/CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02907 | 2021-08-09 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02860 | 2021-08-09 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. A CORRECTION TO B5 WAS MADE AND SHOULD BE THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING BLACK MOLD AROUND, LIGHTHEADED, HEADACHE, PARTICLES IN TUBING/CHAMBER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02903 | 2021-08-09 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2100X15B | IN2100X15B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "VENTILATOR INOPERATIVE" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02952 | 2021-08-09 | 2021-06-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A SPOT ON THEIR LUNG. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02904 | 2021-08-09 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, A VENTILATOR'S NAVIGATIONAL BUTTONS FAILED TO RESPOND AS DESIGNED. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S FRONT PANEL/KEYPAD LED BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
MW5103116 | 2021-08-09 |  | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP WITH HUMIDIFIER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | WHITE HARD STUFF IN NOSE IN MOURNING, CONGESTION, STUFFY NOSE, SORE THROAT, HEADACHES, DIZZINESS, COUGHING, UNABLE TO EAT IN MORNING FOR HOURS, THROWING UP IF I TRIED TO EAT TOO SOON IN MORNING, LOSS OF SLEEP, WAKE UP EXHAUSTED IN MORNING, SO DIZZY HAVE TO USE CANE WHEN I GET OUT OF BED, LOSS OF QUALITY IF LIFE, SUPPOSE TO HELP ME SLEEP AND NOT STOP BREATHING, TOO MANY SIDE AFFECTS, MY DOCTOR SAYS KEEP USING I COULD STOP BREATHING IN MY SLEEP HAVE TO USE THE DREAMSTATION ONLY PER DOCTOR, SLEEP LAB AND INSURANCE COMPANY. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103114 | 2021-08-09 | 2021-06-22 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS REMSTAR PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS HEALTHCARE / RESPIRONICS, INC. | REMSTAR PRO | P08102699 03F8 |  | I HAVE BEEN USING A PHILIPS REMSTAR PRO SINCE (B)(6) 2014 AND EXPERIENCED BREATHING PROBLEMS, STOPPED USING THE ITEM ON THE ABOVE DATE AND MY BREATHING PROBLEMS ARE BETTER NOW. THERE IS BLACK MATERIAL INSIDE THE BREATHING TUBE WHERE IT CONNECTS TO THE MACHINE. I HAVE NOT BEEN ABLE TO FIND A LAB THAT WILL TEST FOR THE POLYESTER BASED POLYURETHANE FOAM USED IN THE PRODUCT THAT WAS THE REASON FOR THE RECALL ON THE DEVICE. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103113 | 2021-08-09 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | TRILOGY 100 VENTILATOR | VENTILATOR, CONTINUOUS, FACILITY USE | PHILIPS / RESPIRONICS, INC. | TRILOGY 100 VENTILATOR |  |  | BROTHER'S TRILOGY 100 BECAME INOPERABLE, STARTED SOUNDING A LOUD ALTERNATING ALARM (BEEP, BEEP...BEEP, BEEP...BEEP, BEEP), THE SCREEN WENT BLACK, AND IT STARTED BUZZING. BEEPING SUBSIDED AFTER PRESSING THE POWER BUTTON AND THEN THE RIGHT BUTTON, BUT IT WOULD NOT POWER BACK ON AND THE BUZZING CONTINUED UNTIL THE BATTERY DIED HOURS LATER. AFTER THE ISSUE BEGAN, HE BEGAN USING HIS BACKUP TRILOGY 100. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103106 | 2021-08-09 | 2021-07-27 | I, n, j, u, r, y | Voluntary report | PHILLIPS DREAMSTATION BIPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500H11C |  |  | BEGAN TO HAVE VERY BAD SORE THROAT, ASTHMA ATTACK, SINUS CONGESTION, ALLERGIC REACTION. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103060 | 2021-08-09 | 2021-07-25 | I, n, j, u, r, y | Voluntary report | PHILLIPS RESPIRONICS REMSTAR AUTO DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I NOTICED A BLACK SUBSTANCE WHEN I WOKE FROM A DISTURBED SLEEP PATTERN IN MY SALIVA, AFTER FORCEFULLY COUGHING IT OUT. THE MACHINE ALMOST STOPPED WORKING AND FORCED THE DEBRIS THROUGH THE TUBE AND INTO MY THROAT. I STOPPED USING THE DEVICE IMMEDIATELY, WHICH IN TURN HAS PROGRESSED INTO DANGEROUS HEALTH PROBLEMS. I HAVE NOT BEEN ABLE TO OBTAIN A NEW MACHINE RESULTING IN LUNG DAMAGE, WEIGHT GAIN, LOSS OF JOB, DEPRESSION, DAYTIME SLEEPINESS, AND SUICIDAL THOUGHTS. NONE TILL I WAS FORCED TO STOP USING THE DEVICE; NOW I AM HAVING LUNG PROBLEMS, DEPRESSION, LOSS OF JOB, SEXUAL DYS, HEART PROBLEMS, AND WEIGHT GAIN. SINCE STOPPING THE DEVICE, I HAVE BEEN PRESCRIBED TRELEGY, ELLIPTA AND PROAIR, ALSO ADDERALL FOR DAYTIME SLEEPINESS EXCESSIVE. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103118 | 2021-08-09 | 2021-03-15 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAM STATION CPAP MACHINE. | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS, INC. | DREAM STATION |  |  | IN (B)(6) 2021 MY WIFE HAD BEEN USING THE PHILIPS DREAM STATION CPAP BREATHING MACHINE FOR 1.5 YEARS UNTIL WE LEARNED FROM SOCIAL MEDIA THERE WAS A RECALL ON THE CANCER CAUSING MACHINES. WE WERE INFORMED VIA PHILIPS' WEBSITE THAT THEY WOULD BE REPLACED AND TO STOP USING IT IMMEDIATELY. MY WIFE HAS BEEN WITHOUT A BREATHING MACHINE AND CAN GET ZERO INFO FROM PHILIPS AS TO WHEN IT WILL BE REPLACED. THE DREAM STATION 2 MACHINE IS OK BUT CAN'T GET THEM TO SEND ONE NOR CAN WE PURCHASE ONE. THIS IS ABSOLUTELY DISGUSTING. HELP. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02902 | 2021-08-09 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE TO CHARGE ITS INTERNAL BATTERY WAS OBSERVED. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S POWER SUPPLY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02668 | 2021-08-09 | 2021-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 02 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEPS DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02667 | 2021-08-09 | 2021-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 02 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02970 | 2021-08-09 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, A FAILURE OF THE EXTERNAL FLOW SENSOR WAS OBSERVED. THE EXTERNAL FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02969 | 2021-08-09 | 2021-07-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, A FAILURE OF THE FLOW SENSOR WAS OBSERVED. THE COMPONENT WAS CONTAMINATED. THE EXTERNAL FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02868 | 2021-08-09 | 2021-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD, INTERFACE BOARD, SYSTEM BOARD, AND POWER MANAGEMENT BOARD NEED TO BE REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG AND THE DEVICE'S OXYGEN BLENDING MODULE BOARD, INTERFACE BOARD, SYSTEM BOARD, AND POWER MANAGEMENT BOARD NEED TO BE REPLACED TO ADDRESS THE ISSUE. THIS WAS INCORRECT. ONLY THE OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
MW5103119 | 2021-08-09 | 2021-08-01 | I, n, j, u, r, y | Voluntary report | DREAM MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/RESPIRONICS, INC. |  |  |  | DURING SLEEP WOKE UP CHOKING SEVERAL TIMES HEADACHES, DIZZINESS, CHEST TIGHTNESS, SEVERE DRY MOUTH. BLACK PARTICLES THROUGH OUT HOSE I WASHED IT BUT I KEPT WAKING AND CHOKING WHILE USING THE NASAL CPAP DREAM MACHINE. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02972 | 2021-08-10 | 2021-07-13 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SEIZURES, STROKES, AND CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02959 | 2021-08-10 |  | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO AUTO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSG500S11 | DSG500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02978 | 2021-08-10 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED ON THIS DEVICE IN MDR 2518422-2021-02809. THIS REPORT WAS SUBMITTED AS A DUPLICATE REPORT OF THE PREVIOUSLY SUBMITTED REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02975 | 2021-08-10 | 2021-03-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL, AU | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | AUX500T15C | AUX500T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SINUS INFECTION. THE PATIENT WAS PRESCRIBED ANTIBIOTICS FOR THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED AN INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. A CORRECTION TO THE B5 SHOULD BE : THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT REPORTED SINUS INFECTION AND SNEEZING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT WAS PRESCRIBED 2 COURSES OF ANTIBIOTICS TO MANAGE THE REPORTED EVENTS AND EYES WERE ALSO STINGING AND THE AIR DID NOT SMELL NICE WHEN USING DEVICE. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 HEALTH EFFECT - CLINICAL CODE HAS BEEN CORRECTED AND TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS HAS BEEN UPDATED IN THIS REPORT. SECTION B7 HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02973 | 2021-08-10 | 2021-02-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL, AU | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AUX500T15C | AUX500T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A SINUS INFECTION. THE PATIENT WAS PRESCRIBED ANTIBIOTICS FOR THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS : THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED UNABLE TO BREATH, HAVING HIS NOSE BLOCKED; WENT TO HIS GP AND WAS DIAGNOSED WITH SINUS INFECTION, WAS GIVEN ANTIBIOTICS AND THE ISSUE WAS SETTLED. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENT OF SINUS INFECTION AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURER'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE INFORMATION AVAILABLE, THE MANUFACTURER CONCLUDES NO FURTHER ACTION IS NECESSARY. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILLED. SECTION B1, B2 HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT-IMPACT CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02974 | 2021-08-10 | 2021-08-04 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL, AU | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AUX500T15C | AUX500T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA, PLEURISY, AND BRONCHIECTASIS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY SUBMITTED MDR 2518422-2021-02974-2 WITH INCORRECT SECTIONS B1, B2, H1, H6. CORRECTIONS TO PREVIOUS MDR ARE MADE IN THIS REPORT AS FOLLOWS. SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY PRODUCT PROBLEM WAS CHECKED IN PREVIOUS MDR) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY IT WAS BLANK) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02998 | 2021-08-10 | 2021-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN THE PREVIOUSLY SUBMITTED REPORT SECTION B5 WAS INCOMPLETE, B5 SECTION WILL BE- THE PATIENT ALLEGED OF HAVING COUGH. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL AND INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF CONTAMINATION IN THE INLET PORT, WHICH WAS LIKELY OF EXTERNAL ORIGIN AND CONTAMINATION INSIDE THE BLOWER BOX. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND ONE INSTANCE OF E-106. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF CONTAMINATION IN THE INLET PORT, WHICH WAS LIKELY OF EXTERNAL ORIGIN AND CONTAMINATION INSIDE THE BLOWER BOX. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. SECTION D9, G3 AND H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02988 | 2021-08-10 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN INITIAL REPORT SECTION B5 WAS MENTION INCORRECTLY, CORRECT B5 SHOULD BE - THE PATIENT ALLEGED PARTICLES IN AIR PATH, COUGHING, UNABLE TO SWALLOW, THROAT IRRITATION. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INPSECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN IN THIS DEVICE. AN INTERNAL VISUAL INPSECTION WAS COMPLETED BY THE MANUFACTURER AND FOUND POTENTIAL MINERAL DEPOSITS ON THE BLOWER MOTOR AND THE BLOWER BOX, INDICATING POTENTIAL WATER INGRESS. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS LOGGED. THE DEVICE WAS APPLIED POWER AND THE DEVICE OPERATED PROPERLY. THE MANUFACTURER CONCLUDES A MINERAL DEPOSTIE CONTAMINATION FOUND WERE CONSISTENT WITH WATER INGRESS. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. IN THIS REPORT, SECTION D9, G3, H3, H6 HAS BEEN UPDATED OR CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02989 | 2021-08-10 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THERE WAS NO PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING NASAL/THROAT IRRITATION, HEADACHE, DRY NOSE, PARTICLES IN THE AIRWAY RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HEALTH EFFECT - CLINICAL CODE HAS BEEN CORRECTED, AND TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02966 | 2021-08-10 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11 | DSX400H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. UNIT WAS SCRAPPED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02987 | 2021-08-10 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/WIFI,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11W | DSX500T11W |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING/CHAMBER,AIRWAY RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02865 | 2021-08-10 | 2021-05-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11C | DSX400H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03004 | 2021-08-10 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED ON THIS DEVICE IN MDR 2518422-2022-05175. THIS REPORT WAS SUBMITTED AS A DUPLICATE REPORT OF THE PREVIOUSLY SUBMITTED REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03002 | 2021-08-10 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP PRO W/HUM/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX600H11C | DSX600H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP,AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION BURNING, SMOKE, ELECTRICAL ODOR RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM.THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL.THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION.THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME.IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE,AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02986 | 2021-08-10 | 2021-06-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SINUS INFECTIONS. THE PATIENT HAD SINUPLASTY SURGERY IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED SEVERE SINUS INFECTIONS, HEADACHES, NASAL IRRITATION AND PAIN. THE PATIENT DID NOT REPORT TO RECEIVE MEDICAL INTERVENTION. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03001 | 2021-08-10 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11C | DSX400H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WAS ONE ERROR FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. UNIT SCRAPPED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03003 | 2021-08-10 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP PRO W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX600T11C | DSX600T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF BLACK PARTICLES, HEADACHES, SHORT OF BREATH, CONGESTION IN CHEST, COUGHING, RESPIRATORY DISTRESS RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. AFTER THE FINAL ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, THE CUSTOMER DECLINED TO RESPOND TO THE GFE RELATED QUESTIONS AND TERMINATED THE CALL. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02847 | 2021-08-10 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP AUTO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02852 | 2021-08-10 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING TO HAVE BLACK PARTICLES IN WATER TANK, ERRORS, AND A SORE THROAT WHILE USING THIS DEVICE. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED EXTERNALLY AND THE MANUFACTURER OBSERVED DUST/DIRT CONTAMINATION INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM WAS OBSERVED THROUGHOUT DEVICE ENCLOSURE, AND AIRPATH. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED AND THE MANUFACTURER OBSERVED MINERAL DEPOSIT SPOTS ON THE BOTTOM OF THE BLOWER MOTOR AND INSIDE THE BLOWER MOTOR CASING, ITS INDICATES POTENTIAL WATER INGRESS INSIDE THE DEVICE; UNKNOWN DUST-LIKE CONTAMINATION WAS OBSERVED ON TOP OF THE BLOWER MOTOR AND BLOWER MOTOR SEAL. THE MANUFACTURER OBSERVED BLACK CONTAMINATION, CONSISTENT WITH KERATIN, AT THE AIR OUTLET OF THE BLOWER BOX. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE TWO ERRORS (# IRRIT02, # INFECT01) FOUND. THE MANUFACTURER CONCLUDES THAT THEY CONFIRM THE CUSTOMER'S ALLEGATION AND THERE IS DUST/DIRT CONTAMINATION AND THE SOURCE OF CONTAMINATIONS WERE EXTERNAL TO THE DEVICE. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02997 | 2021-08-10 | 2021-08-09 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO, FR | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | FRX500S14 | FRX500S14 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER AND PITUITARY ADENOMA. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02991 | 2021-08-10 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LESIONS ON THEIR LUNGS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO DEVELOP LESIONS ON THEIR LUNGS APPROXIMATELY 2 YEARS AGO. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURER CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REPORTED BY THE PATIENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. NO CARE ORCHESTRATOR DATA WAS AVAILABLE FOR DOWNLOAD. DURING INVESTIGATION OF THE DEVICE,A DARK DUST CONTAMINANT ON THE TOP ENCLOSURE, FRONT ENCLOSURE, REAR PANEL, AND BOTTOM ENCLOSURE WAS OBSERVED. THE PCA APPEARED TO HAVE VERDIGRIS OF THE ANTENNA, HOWEVER, THERE WAS NO EVIDENCE THAT THIS AFFECTED THE PERFORMANCE OF THE PCA IN ANY WAY. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED A PIECE OF WHITE PLASTIC TYPE PARTICLE WAS OBSERVED ON THE BOTTOM ENCLOSURE UNDER THE BLOWER BOX. AN UNKNOWN DUST CONTAMINANT WAS OBSERVED ON THE BLOWER SEAL. EVIDENCE OF LIQUID INGRESS WAS PRESENT ON THE BLOWER AND BLOWER BOX. A KERATIN-LIKE SUBSTANCE WAS OBSERVED ON THE BLOWER BOX OUTLET. THE INVESTIGATOR CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION IN THE DEVICE. SECTION D8, D9, H2, H3 AND H6 WERE UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02992 | 2021-08-10 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A SINUS INFECTION. THE PATIENT WAS PRESCRIBED ANTIBIOTICS AND STEROIDS FOR THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 12, 2024. SECTION H11 SHOULD BE REPORTED AS THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A SINUS INFECTION. THE PATIENT WAS PRESCRIBED ANTIBIOTICS AND STEROIDS FOR THE REPORTED EVENT. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM (ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR). SECTION B2 WAS CHANGED TO BLANK (PREVIOUSLY IT WAS HOSPITALIZATION). SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02851 | 2021-08-10 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/ HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF A THERMAL EVENT OCCURRING TO THE DC POWER SUPPLY AND THE DC POWER INPUT OF THE CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE. THE USER REPORTED THE UNIT STOPPED WORKING, HE NOTICED A BURNING ELECTRICAL ODOR AND A THERMAL EVENT ON THE POWER SUPPLY AT THE POWER INLET PORT ON THE CPAP. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. UPON COMPLETION OF THE INVESTIGATION, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY BECAME AWARE OF AN ALLEGATION OF A THERMAL EVENT OCCURRING TO THE DC POWER SUPPLY AND THE DC POWER INPUT OF THE CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE. THE USER REPORTED THE UNIT STOPPED WORKING, HE NOTICED A BURNING ELECTRICAL ODOR AND A THERMAL EVENT ON THE POWER SUPPLY AT THE POWER INLET PORT ON THE CPAP. THERE WAS NO PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER AND NO ERROR WAS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT WAS SCRAPPED. SECTION A1, E1, H7, H9 WERE CORRECTED AND CAPTURED IN THIS REPORT. SECTION D8, D9, H2, H3 AND H6 WERE UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02971 | 2021-08-10 | 2021-07-12 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200S11 | DSX200S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SINONASAL UNDIFFERENTIATED CARCINOMA. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM AND PATIENT ALLEGED SINONASAL UNDIFFERENTIATED CARCINOMA. THE PATIENT DID NOT REPORT TO RECEIVE MEDICAL INTERVENTION. THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02912 | 2021-08-10 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED FILTERS TURNING BLACK AND PARTICLES IN CHAMBER. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WAS ONE ERROR FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02691 | 2021-08-10 | 2021-07-12 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SCAR TISSUE IN THEIR LUNGS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP,AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AND CAUSED A PATIENT TO DEVELOP SCAR TISSUE IN THEIR LUNGS RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02938 | 2021-08-10 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICES LCD DISPLAY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02936 | 2021-08-10 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 02 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF A VENTILATOR'S LCD SCREEN WAS OBSERVED. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S LCD DISPLAY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02940 | 2021-08-10 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 02 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG WAS OBSERVED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02941 | 2021-08-10 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02935 | 2021-08-10 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 02 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02937 | 2021-08-10 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 02 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02939 | 2021-08-10 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 02 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02737 | 2021-08-10 | 2021-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. | IT WAS INITIALLY REPORTED, A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S OUTLET SIDE PANEL WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02738 | 2021-08-10 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS NOT GENERATING PRESSURE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE BLOWER AIR PATH WAS BLOCKED WITH DIRT AND DUST. THE DEVICE'S BLOWER AIR PATH WAS CLEANED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02869 | 2021-08-10 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE LCD SCREEN WAS OBSERVED TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02968 | 2021-08-10 | 2021-07-14 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION EXPERT W/HUM/HT/CELL, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX501T12C | CAX501T12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP DIFFICULTY BREATHING, AND SEVERE SINUS INFECTIONS. THE PATIENT WAS PRESCRIBED ANTIBIOTICS FOR THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON MAR 19, 2025 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGED TO DEVELOP DIFFICULTY BREATHING, AND SEVERE SINUS INFECTIONS. THE PATIENT WAS PRESCRIBED ANTIBIOTICS FOR THE REPORTED EVENT RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM.THE REPORTED EVENT OF HAVING DEVELOP DIFFICULTY BREATHING, AND SEVERE SINUS INFECTIONS. THE PATIENT WAS PRESCRIBED ANTIBIOTICS FOR THE REPORTED EVENT WAS REVIEWED BY THE PMS CLINICAL EXPERT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. DURING THE INVESTIGATION, THE MANUFACTURER OBSERVED AN UNKNOWN DUST CONTAMINANT ON THE FLIP DOORS, PCA (PRINTED CIRCUIT ASSEMBLY), TOP ENCLOSURE, REAR PANEL, UI (USER INTERFACE) PANEL, BOTTOM ENCLOSURE, INSIDE ISO (INTERNATIONAL ORGANIZATION FOR STANDARDIZATION) PORT, BLOWER, AND BLOWER SEAL. THE MANUFACTURER ALSO OBSERVED BLACK HAIRLIKE CONTAMINANT ON THE FLIP DOORS, TOP ENCLOSURE, UI (USER INTERFACE) PANEL, INSIDE THE ISO (INTERNATIONAL ORGANIZATION FOR STANDARDIZATION) PORT, REAR PANEL, BOTTOM ENCLOSURE, AND BLOWER. EVIDENCE OF LIQUID INGRESS WITH AN UNKNOWN DUST CONTAMINANT CONSISTENT WITH MINERAL DEPOSITS WAS OBSERVED ON THE BLOWER. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED. THE MANUFACTURER USED A FITT (FOAM INTEGRITY TEST TOOL) AND WAS NOT ABLE TO CONFIRM THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THAT THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE MANUFACTURER CONFIRMED THE PRESENCE OF CONTAMINATION IN THE AIRPATH AND UNABLE TO ADDRESS THE PATIENT'S SYMPTOMS. SECTIONS D9, H2, H3 AND H6 HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02905 | 2021-08-10 | 2021-06-25 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED BRONCHITIS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF PRESCRIBED STEROIDS AND ANTIBIOTICS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 04, 2024 AND SECTION H11 SHOULD BE REPORTED AS THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT ALLEGED DEVELOPED BRONCHITIS, SORE THROAT AND COUGH. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF PRESCRIBED STEROIDS, ANTIBIOTICS AND MUCINEX. THE SECTIONS A2 (AGE), A3 (GENDER), B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM (ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR) SECTION B2 WAS CHANGED TO BLANK (PREVIOUSLY IT WAS OTHER SERIOUS) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION SECTION H6 HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02999 | 2021-08-10 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02870 | 2021-08-10 | 2021-07-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02867 | 2021-08-10 | 2021-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | DS2200X11B | DS2200X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONSTANT ALARM CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD AND MACHINE FLOW SENSOR WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02909 | 2021-08-10 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN INITIAL REPORT SECTION B5 WAS INCORRECT, CORRECT B5 SHOULD BE - THE PATIENT ALLEGED NASAL IRRITATION, SORE THROAT, SINUS ISSUE, HEADACHE, COUGHING. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND A BLUE AND BLACK CONTAMINATION AT THE INLET AIR FILTER HOLE. AN INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND CONTAMINATION INSIDE OF THE CASE ON THE BOTTOM BACK AND TOP PANELS OF THE CASE. TINY RESIDUE SPOTS ON SOUND ABATEMENT FOAM AND A DUSTY SMOKE (NON-TOBACCO) ODOR. THE MANUFACTURER FOUND EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS. THE MANUFACTURER CONCLUDES CONTAMINATION FOUND WERE CONSISTENT WITH BLUE AND BLACK CONTAMINATION AT THE INLET AIR FILTER HOLE, CONTAMINATION INSIDE OF THE CASE ON THE BOTTOM BACK AND TOP PANELS OF THE CASE, TINY RESIDUE SPOTS ON SOUND ABATEMENT FOAM AND A DUSTY SMOKE (NON-TOBACCO) ODOR. THE MANUFACTURER CONFIRMED THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. IN THIS REPORT, SECTION D9, G3, H3, H6 HAS BEEN UPDATED OR CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02910 | 2021-08-10 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11 | DSX700T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING FILTERS TURN BLACK, SHORTNESS OF BREATH ON AND OFF COUGH, LOTS OF HEADACHES AND LUNG DISCOMFORT RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION A1 CORRECTED AND H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02911 | 2021-08-10 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLE IN HUMIDIFIER, STARTED COUGHING AND LOSING HIS VOICE AND CHEST TIGHTNESS. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED. IN INITIAL REPORT, SECTION A1 AND H6 CLINICAL CODES WAS NOT CAPTURED, SECTION G3 DATE WAS WRONGLY CAPTURED AS 07/14/2021, SHOULD BE 07/01/2021, CORRECTED AND UPDATED THESE FEILDS IN THIS REPORT SECTION G1 AND H6 WERE UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02705 | 2021-08-10 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED ASTHMA AND EYE IRRITATION. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP,AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AND THE PATIENT DEVELOPED ASTHMA AND EYE IRRITATION RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM.THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN OBSERVED AN UNKNOWN DUST CONTAMINANT ON THE INSIDE OF THE BLOWER BOX AT THE LOCATION OF THE FILTER. AN UNKNOWN DARK CONTAMINANT WAS OBSERVED ON THE TOP ENCLOSURE.AN UNKNOWN DUST CONTAMINANT WAS OBSERVED ON THE BOTTOM ENCLOSURE, REAR PANEL, ON THE BLOWER, AND THE BLOWER SEAL. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED.IT IS CONFIRMED NO PRESENCE OF DEGRADED SOUND ABATEMENT FOAM. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERORRS. THE MANUFACTURER CONCLUDES THE DEVICE EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02706 | 2021-08-10 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500S11 | DSX500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED BREAST CANCER. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED BREAST CANCER. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL AND INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF 2 PIECES OF CONTAMINATION IN THE BLOWER BOX. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND NO ERRORS LOGGED. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF 2 PIECES OF CONTAMINATION WERE FOUND IN THE BLOWER BOX. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION D9, G3 AND H6 HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02707 | 2021-08-10 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP EYE IRRITATION AND DYSPNEA. THE PATIENT REQUIRED MEDICAL INTERVENTION IN THE FORM OF SEVERAL THROAT SURGERIES. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON 03/10/2025, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED SEVERAL THROAT SURGERIES AND DEVELOPED EYE IRRITATION, BREATHING DIFFICULTY AND SINUS. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGED DIZZINESS, HEADACHE, NAUSEA / VOMITING, LUNG DISEASE. ADDITIONAL INFORMATION WAS RECEIVED ABOUT ALLEGED VISUALIZATION OF PARTICLE. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02814 | 2021-08-10 | 2021-06-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED INFLAMMATION IN THE RIGHT UPPER LOBE OF THE LUNG WHICH WAS REVEALED ON A CT SCAN AS A GROUND GLASS OPACITY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED INFLAMMATION IN THE RIGHT UPPER LOBE OF THE LUNG WHICH WAS REVEALED ON A CT SCAN AS A GROUND GLASS OPACITY. IN THE PREVIOUS REPORT SECTION B5 WAS INCORRECT, IT SHOULD HAVE BEEN REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED HAVING COUGH AND THE PATIENT DEVELOPED INFLAMMATION IN THE RIGHT UPPER LOBE OF THE LUNG WHICH WAS REVEALED ON A CT SCAN AS A GROUND GLASS OPACITY. BASED ON INFORMATION AVAILABLE AT THE TIME OF THIS REVIEW, THERE IS INSUFFICIENT EVIDENCE TO PRONOUNCE A SERIOUS INJURY. NOR IS THERE ANY REPORT OF MEDICAL INTERVENTION REQUIRED OR RECEIVED TO PRECLUDE A LIFE-THREATENING INJURY OR PERMANENT IMPAIRMENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED / CORRECTED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02736 | 2021-08-10 | 2021-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DISPLAY WAS OBSERVED (PHYSICAL DAMAGE). THE DEVICE'S DISPLAY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02715 | 2021-08-10 | 2021-06-25 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/HT/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX400T11C | DSX400T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED BACTERIAL PNEUMONIA. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED BACTERIAL PNEUMONIA. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE EVENT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THEREFORE, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. SECTION G4 HAS BEEN CORRECTED / UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02876 | 2021-08-10 | 2021-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | AUX500T15 | AUX500T15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO THE DEVICE'S SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02714 | 2021-08-10 | 2021-06-28 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP/HUM/HT/WIFI, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11W | DSX700T11W |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT WAS DIAGNOSED WITH LUNG CANCER. THE PATIENT'S TUMOR HAS BEEN SURGICALLY REMOVED IN RESPONSE TO THE EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT WAS DIAGNOSED WITH LUNG CANCER. THE PATIENT'S TUMOR HAS BEEN SURGICALLY REMOVED IN RESPONSE TO THE EVENT. ADDITIONAL INFORMATION WAS RECEIVED AND ADDED TO THE REPORT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER ON 05/01/2023 FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED ON 05/10/2023. THE EXTERNAL INVESTIGATION SHOWED THAT THERE WERE NO SIGNS OF CONTAMINATION ON THE OUTSIDE OF THE DEVICE. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED IN THIS DEVICE. THE DEVICE WAS HOOKED UP TO A KNOWN GOOD POWER SUPPLY AND POWER CORD, AIRFLOW WAS VERIFIED TO BE OPERATING CORRECTLY. THE ERROR LOG SHOWED THAT THERE WERE ZERO INSTANCES OF ERRORS ON THE DEVICE. THE INTERNAL INVESTIGATION SHOWED THAT THERE WERE SIGNS OF AN UNKNOWN DUST CONTAMINATE ON THE FILTER PORT. THERE WERE NO SIGNS OF SOUND ABATEMENT FOAM DEGRADATION IN THE BLOWER BOX. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. SECTION D8, D9, H2, H3 AND H6 WERE UPDATED. SECTION E1 WAS CORRECTED BY CAPTURING EMAIL WHICH WAS MISSED IN INITIAL REPORT. THE MANUFACTURER PREVIOUSLY SUBMITTED MDR 2518422-2021-02714-2 WITH INCORRECT SECTIONS B1, B2, H1, H6. CORRECTIONS TO PREVIOUS MDR ARE MADE IN THIS REPORT AS FOLLOWS. SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY PRODUCT PROBLEM WAS CHECKED IN PREVIOUS MDR). SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY IT WAS BLANK). SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6- HEALTH IMPACT CODE WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103172 | 2021-08-10 | 2019-11-20 | I, n, j, u, r, y | Voluntary report | PHILLIPS DREAMSTATION CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | SEVERE SINUS INFECTION LED TO DISCOVERY OF FUNGAL BALL CLOSED OFF IN SINUS. SURGERY DEFERRED (PANDEMIC) UNTIL (B)(6). INFLAMMATION CONTINUES AND IS STILL BEING TREATED. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103169 | 2021-08-10 | 2019-03-20 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS REMSTAR PLUS C-FLEX | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS / RESPIRONICS, INC. | REMSTAR PLUS, DOM REF 260P | NA |  | DIAGNOSED WITH RENAL CELL CARCINOMA AND INVASIVE DCIS AFTER 4 YEARS USING THE PHILIPS REMSTAR PLUS C-FLEX DEVICE. NOTED 4 MM LUNG UPPER LOBE MODULES. JUST NOTIFIED IN 2021 OF POTENTIAL HARM FROM USE OF THIS CPAP MACHINE. MY CPAP WAS REPLACED IN 2020 WITH RESMED UNIT. DID THE USE OF THIS PRODUCT CONTRIBUTE TO MY DIAGNOSIS. IS THERE A CLASS ACTION SUIT RELATED TO THIS RECALL? FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103164 | 2021-08-10 | 2021-04-17 | I, n, j, u, r, y | Voluntary report | DREAMSTATION 2 | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS COMPANY/ RESPIRONICS, INC. | 122220-01278 | DSX500T11C |  | ON (B)(6) 2021 FOLLOWING 12 HOURS OF WHEEZING, CONTACTED MD ON CALL AND SENT TO (B)(6) RESPIRATORY URGENT CARE (B)(6). EXAM WITH CHEST XRAY AND DIAGNOSED WITH ASTHMA AND PLACED ON PREDNISONE 20 MG AND PROVENTIL HFA INHALANT FOR WHEEZING AND SHORTNESS OF BREATH.(THIS WAS A NEW DIAGNOSIS AND TREATMENT FOR ME ) I HAVE NEEDED THE INHALER ON A DAILY BASIS. I HAD BEEN USING THE DREAMSTATION 2 PRODUCT FROM PHILLIPS CO SINCE (B)(6) 2020 WHICH PRESENTLY HAS BEEN RECALLED. IT LISTED ASTHMA, EYE IRRITATION, HEADACHE AND NAUSEA AS ADVERSE PROBLEMS ASSOCIATED WITH USE OF THIS PRODUCT. COINCIDENTALLY HAD EYE IRRITATION AND INFLAMMATORY RESPONSE BILATERALLY AND SEEN AND TREATED BY OPTOMETRIST ON (B)(6) 2021 WHO DIAGNOSED INFLAMMATION AND CORNEAL EDEMA. TREATED WITH AN ANTIHISTAMINE EYE GTT AND EYE UNG AS WELL. A FOLLOW UP WAS DONE LATER WITH AN OPHTHALMOLOGIST WHO CONCURRED AND STATED THERE WAS GOOD INITIAL RESOLVE OF PROBLEM AND TO FINISH TREATMENT. INITIALLY, I BEGAN HAVING SEVERE HEADACHE AND NAUSEA AFTER INITIAL USE OF C-PAP MACHINE WHICH WAS TREATED BY INTERNIST. I HAVE REMAINED ON THE PROTONIX TWICE DAILY AND ALTHOUGH NAUSEA HAS SUBSIDED I STILL HAVE DAILY HEADACHES WHICH HAVE LESSENED NOW WITH CESSATION OF MACHINE. THEY HAVE NOT STOPPED AND MY CONCERN IS SIGNIFICANT OVER THE PARTICLES AND CHEMICALS THAT HAVE BEEN PUSHED THROUGH MY LUNGS SINCE (B)(6). I AM CONCERNED THAT ALTHOUGH PHILLIPS COMPANY WAS AWARE OF THIS PROBLEM IN APRIL THEY DID NOT SHARE THAT WITH THE CONSUMER, ME, UNTIL (B)(6) 2021. THERE IS A REAL PROBLEM THERE AS I HAD 3 SIGNIFICANT SIDE EFFECTS LISTED AND CONTINUE TO HAVE THEM. HAD I KNOWN IN (B)(6) THE ADVERSE EFFECTS MAY NOT HAVE BEEN AS SIGNIFICANT OR AS DETRIMENTAL TO MY HEALTH. I HAVE GRAVE CONCERNS ABOUT THE LONG TERM EFFECTS. THANK YOU FOR YOUR ASSIST IN THIS MATTER. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103160 | 2021-08-10 | 2021-05-26 | I, n, j, u, r, y | Voluntary report | RESPIRONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS / RESPIRONICS, INC. | DSX500H11C |  |  | SINCE (B)(6) 2021, I HAD HEADACHES ALMOST DAILY. I BEGAN TO NOTICE BY MID (B)(6) THAT DAYS WHEN I DIDN'T USE THE CPAP MACHINE, MY HEADACHES WERE NON EXISTENT. I NOTICED THE LENGTH OF USE EACH NIGHT WAS DIRECTED RELATED TO THE SEVERITY OF THE HEADACHE THE NEXT DAY. I ALSO EXPERIENCED SEVERAL ISSUES IF FUNNY FEELINGS IN MY CHEST. IT WASN'T REALLY PRESSURE AND IT WAS HARD TO DESCRIBE. SEVERAL HEART TEST WERE DONE AND NOTHING WAS FOUND. THE FIRST MONTH OFF THE MACHINE I HAD NO HEADACHES. I AM NOW EXPERIENCING THEM SOME DAYS EVEN WITHOUT USING THE MACHINE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103162 | 2021-08-10 | 2017-11-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION C-PAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILLIPS / RESPIRONICS, INC. |  |  |  | RE: PHILLIPS DREAM STATION C-PAP MACHINE. I BEGAN USING THIS DEVICE IN (B)(6) 2017. IN (B)(6) 2017 I HAD THE FIRST OF A SERIES OF REACTIVE AIRWAY DISEASE EPISODES, WHICH LASTED ABOUT 10 WEEKS BEFORE APPROPRIATE TREATMENT WAS IDENTIFIED (PREDNISONE, 10 DAYS). I HAD NOT HAD ANY RAD EPISODES BEFORE (B)(6) 2017. THE EPISODE BEGAN WITH A VIRAL RESPIRATORY INFECTION IN EARLY (B)(6). WHILE MY CO-WORKERS GOT SICK AND RECOVERED, MY SYMPTOMS WENT ON FOR WEEKS. SYMPTOMS INCLUDED SHORTNESS OF BREATH, NON-PRODUCTIVE COUGH, PRONOUNCED FATIGUE, FLU-LIKE SYMPTOMS, SEVERELY SWOLLEN SINUSES WITHOUT MUCH CONGESTION, MID-FACIAL HEADACHES, AND REDUCED CARDIO-VASCULAR FUNCTIONING. I AM A VERY ACTIVE RECREATIONAL CYCLIST AND AT THAT TIME HAD BEEN REGULARLY TAKING EXTENDED RIDES 3 OR 4 TIMES A WEEK OF 30 TO 70 MILES, SOMETIMES UP TO 100. DURING THESE EPISODES I BECAME NOTICEABLY OUT OF BREATH WALKING UP A SINGLE FLIGHT OF STAIRS. AFTER EXPERIMENTATION IT SEEMED THAT A 10-DAY COURSE OF PREDNISONE WAS USUALLY ENOUGH TO SETTLE THINGS. ABOUT A YEAR AND A HALF AGO I SWITCHED ALLERGISTS. ALLERGY TESTING HAD INDICATED THAT I WAS NOT RESPONDING TO POLLENS AND OTHER ALLERGENS FOR WHICH I HAD PREVIOUSLY RECEIVED IMMUNIZATION THERAPY. I HAD BEEN DIAGNOSED AS HAVING CHRONIC NON-ALLERGIC RHINITIS-SINUSITIS. AFTER HEARING ABOUT MY EPISODES THE ALLERGIST COMMENTED THAT IT SOUNDED AS IF I MIGHT HAVE SOMETHING LIKE N-ERD (NSAID EXACERBATED RESPIRATORY DISORDER). HE BEGAN TREATING ME ON THAT ASSUMPTION, AND I STARTED A TREATMENT REGIMEN THAT NOW INCLUDES BREO AND SPIRIVA INHALERS, ZILEUTON, AND BUDESONIDE RINSES 2 X DAILY. I LATER (2019) DEVELOPED SIGNIFICANT REACTIONS TO ALCOHOL, ESPECIALLY WINE AND BEER. AFTER AS LITTLE AS 2 OUNCES MY SINUSES SWELL UP ENOUGH TO RESTRICT THE AIR FLOW THROUGH MY NOSE SIGNIFICANTLY, I GET HEADACHES, AND I EXPERIENCE FATIGUE AND OTHER FLU-LIKE SYMPTOMS THROUGH THE NEXT DAY. I HAD NOT PREVIOUSLY HAD THIS ISSUE. THESE EPISODES HAVE REPEATED ON A REGULAR BASIS SINCE THEN. THEY APPEAR TO BE TRIGGERED BY A RESPIRATORY VIRUS, SOME FORM OF RESPIRATORY DISTRESS, ALCOHOL CONSUMPTION AND/OR POOR AIR QUALITY. I LIVE IN (B)(6) AND EXPERIENCED ANOTHER EPISODE IN (B)(6) 2020 FROM THE WILDFIRE SMOKE. AFTER THE VIRUS HAS SET IN I EXPERIENCE ESCALATING SYMPTOMS AS DESCRIBED ABOVE THAT CONTINUE FOR WEEKS (THE FIRST WAS 10 WEEKS, THE SECOND 8 WEEKS. ETC.). I'VE PROBABLY HAD AT LEAST 6 EPISODES ALL TOLD, INCLUDING THE PRESENT EPISODE. HOWEVER, BECAUSE THIS CONDITION IS UNUSUAL IT HAS BEEN DIFFICULT TO GET PROPER MEDICAL ATTENTION WHEN I PRESENT TO A PROVIDER WHO DOES NOT KNOW MY HISTORY. THIS PAST WEEKEND I WAS ABOUT 3 WEEKS INTO ANOTHER EPISODE THAT BEGAN WITH A RESPIRATORY VIRUS, MY WIFE HAD IT ABOUT 5 DAYS AFTER I STARTED, BUT RECOVERED QUICKLY. I BEGAN TREATMENT WITH A 5-DAY PREDNISONE COURSE AND A Z-PACK. AFTER I FINISHED THE MEDS I FELT BETTER, ALTHOUGH NOT 100% (MAYBE 60%). THE NEXT DAY, A SATURDAY, I BECAME VERY SHORT OF BREATH AND DIZZY AFTER MILD EXERTION, WHICH DID NOT ALLEVIATE AFTER AN HOUR. I COULD NOT REACH MY PRIMARY PROVIDER (ONLY A CONTRACT TRIAGE NURSE), MY ALLERGIST (NO WEEKEND COVERAGE), OR MY SINUS SPECIALIST (NO WEEKEND COVERAGE). HEARING MY SYMPTOMS THE TRIAGE NURSE TOLD ME TO GO TO THE ER, DESPITE SHARING MY MEDICAL HISTORY. I WENT BY AMBULANCE, SPENT 4 HOURS THERE, HAD AN EKG, AND BLOODWORK. ALL VITALS, THE EKG AND BLOODWORK WERE WNL. THE PROVIDERS AGREED THAT I NEEDED TO HAVE MY PREDNISONE TREATMENT EXTENDED, AND SENT ME HOME WITH A PRESCRIPTION. THIS HELPED, BUT I AM STILL RECOVERING. SO, TWO MONTHS AFTER I STARTED USING THE PHILLIPS CPAP I DEVELOPED AN UNUSUAL RAD DISORDER. IT HAS WORSENED OVER TIME, (WHILE I CONTINUED TO USE THE MACHINE), HAS BEEN VERY DISRUPTIVE IN MY LIFE, REQUIRES ME TO TAKE 4 DIFFERENT MEDS, WASTES TIME, MONEY, AND MEDICAL RESOURCES, COMPLICATES GETTING PROPER CARE AND PUTS ME AT HIGHER RISK FOR SERIOUS MEDICAL SEQUELAE (E.G., COVID COMPLICATIONS). I CHANGED MACHINES LAST WEEK AND HOPE THAT WILL HELP. I HAVE LOTS OF MEDICAL RECORDS AND LAB RESULTS DATING BACK TO (B)(6) 2017 WHICH SUPPORT THIS COMPLAINT. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103161 | 2021-08-10 | 2020-01-01 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DREAMSTATION AUTOCPAP | DSX500H11C |  | ONGOING RESPIRATORY ISSUES; I WAS USING A PHILIPS DREAMSTATION CPAP MACHINE FOR ABOUT 2 YEARS. I STARTED EXPERIENCING HEADACHES AND SPORADIC BREATHING ISSUES WHEN I WAS WAKING UP. I VISITED A DR AND THEY PRESCRIBED SOME MEDS FOR ME. MY SISTER AND I ENDED UP BEING THE PRIMARY CARE FOR MY MOTHER AS SHE CAME HOME INTO HOME HOSPICE. THE CPAP MACHINE HAS LEFT ME WITH A CONSTANT IRRITATION IN MY THROAT, SINUSES, AND LUNGS. I'M CONSTANTLY COUGHING AND STILL HAVE A LOW-GRADE HEADACHE. SINCE MY MOTHER HAS PASSED I HAVE GOTTEN BACK AND FOLLOWED UP WITH MY PHYSICIAN AND WE ARE TRYING TO TREAT THE SYMPTOMS NOW THAT WE KNOW THE CPAP WAS CAUSING THESE ISSUES. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103146 | 2021-08-10 | 2020-01-11 | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSXHCP |  |  | REPORTED TO MY DOCTOR AND CPAP SUPPLIER THAT USING THE MACHINE CAUSED ME TO HAVE RESPIRATORY ISSUES DURING USE. THEY TRIED A VARIETY OF MASK OPTIONS AND NONE WORKED. IT WAS PRESUMED THAT THE DEVICE WAS CAUSING MY DIFFICULTY BREATHING AND HEADACHES. I VOLUNTARILY STOPPED USING THE DEVICE BECAUSE OF STRUGGLING TO BREATH WHICH INTERN NOW CAUSING MY SLEEP APNEA TO GO UNTREATED. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103156 | 2021-08-10 | 2021-08-08 | I, n, j, u, r, y | Voluntary report | DREAM STATION AUTO PHILIPS RESPIRONICS CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. | DSX500T11C |  |  | SINUS INFECTIONS; LATELY HAVING MILD HEADACHES, BUT LAST NIGHT WORSEN AND WAS COMPLETELY CONGESTED. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103154 | 2021-08-10 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS DREAMSTATION CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS INC. |  |  |  | RECALL ON PHILIPS DREAMSTATION REQUIRES FDA APPROVAL TO REPLACE THE MACHINE. PLEASE APPROVE SO WE CAN GET BACK TO LIVING. THANK YOU. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103153 | 2021-08-10 |  | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | FOUND OUT MY UNIT IS RECALLED. REGISTERED MY PRODUCT OVER 2 MONTHS AGO AND HAVEN'T HEARD BACK FROM PHILIPS ABOUT A REPLACEMENT. IT'S MEDICALLY NECESSARY I HAVE A CPAP. I CAN'T AFFORD A REPLACEMENT AND NEED PHILIPS TO REPLACE MY AFFECTED DEVICE. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103149 | 2021-08-10 | 2021-08-02 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAM STATION 1 RESPIRONICS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | RECALL ON DME PRODUCT PHILIPS RESPIRONICS DREAM STATION 1. HAVE BAD SLEEP APNEA AND HAVE TO CONTINUE USE. DEVELOPED A SEVERE SINUS INFECTION WITH BLOOD SHOT EYES, EYE PAIN, AND NECK PAIN. REGISTERED IT ON THE PHILIPS WEBSITE, BUT HAVE NOT HEARD ANYTHING BACK FROM THEM. |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103129 | 2021-08-10 | 2021-04-17 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DRM STATION A P-FLEX HH & MODEM |  |  | I HAVE BEEN EXPERIENCING A DRY COUGH FOR ABOUT 4 MONTHS. FIRST I THOUGHT IT WAS RELATED TO THE PFIZER-BIONTECH COVID-19 VACCINE AS IT STARTED TWO DAYS AFTER MY FIRST SHOT BUT WITH THE RECENT PHILIPS RESPIRONICS RECALL I AM NOW WONDERING IF IT IS RELATED TO THAT INSTEAD. I HAVE TAKEN 2 MONTHS OF NASONEX AND ALMOST A MONTH OF SYMBICORT BUT THE COUGH IS STILL THERE. I HAVE ALSO TRIED VARIOUS HOME REMEDIES. I DID HAVE AN CHEST X-RAY DONE WHICH I WILL FIND MORE ABOUT IN THE COMING DAYS. CHEST X-RAY RESULTS TO COME. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103131 | 2021-08-10 | 2021-07-29 | I, n, j, u, r, y | Voluntary report | CPAP DREAM STATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. |  |  |  | COMPLAINANT REPORTS HIS CPAP IS EMITTING GASES CAUSED BY FOAM BREAKING DOWN. HE IS AWARE OF THE PHILIP RECALL FOR HIS CPAP, AND HE HAS INITIATED HIS COMPLAINT THROUGH THEIR PROCESS AS STATED IN THEIR RECALL. COMPLAINANT STATES HE IS NOT RECEIVING ANY GUIDANCE FROM THE COMPANY REGARDING THE NEXT STEP TO RETRIEVE AND REPLACE HIS CPAP. NOTIFICATION DATE: 07/30/2021, HEALTH PROFESSIONAL ATTENDED, HOSPITALIZATION NOT REQUIRED. DIFFICULTY BREATHING/RESPIRATORY; SLEEP PROBLEMS, NERVOUS. PROVIDER NAME: DR. (B)(6), DEFECTIVE: EMITTING GASES CAUSED BY FOAM. ON (B)(6) 2021 DAL-DO CCC: COMPLAINANT WAS CONTACTED AND PROVIDED THE FDA WEBSITE LANDING PAGE FOR THIS RECALL: AS THE MEDICAL DEVICE MANUFACTURER, PHILIPS RESPIRONICS HAS A RESPONSIBILITY NOT ONLY TO ENSURE THE MANUFACTURE OF SAFE AND EFFECTIVE DEVICES, BUT ALSO TO ESTABLISH AN APPROPRIATE MITIGATION STRATEGY TO REDUCE PUBLIC HARM IN THE EVENT OF DEVICE FAILURE OR DEFECT. THE FDA IS REVIEWING THE INFORMATION PHILIPS RESPIRONICS HAS PROVIDED REGARDING REPLACEMENT DEVICES WHILE CONTINUING TO WORK WITH PHILIPS RESPIRONICS ON THEIR CORRECTIVE ACTIONS FOR EXISTING DEVICES. AS THE FDA REVIEWS THIS INFO, THE FDA WILL DETERMINE WHETHER THE PROPOSED REPLACEMENT DEVICES POSE ANY ADD'L RISK TO PEOPLE WHO USE THESE DEVICES. COMPLAINT: (B)(4). |  | Y | N | I |  | 2021-08
MW5103141 | 2021-08-10 | 2020-11-28 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO CPAP HUMHT CELL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DREAMSTATION | NA |  | RUBBER SEALS BEGAN FALLING OUT OF MY DREAMSTATION BIPAP. I HAVE ADDED A BACTERIAL FILTER BASED ON MY DME PROVIDER RECOMMENDATION AND STOPPED USING MY HUMIDIFIER. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103155 | 2021-08-10 | 2020-09-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. | DSX500T11C |  |  | I DEVELOPED SIGNIFICANT SINUS INFECTION, SEVERE NASAL IRRITATION, HEADACHE. IT HAS REQUIRED REPEATED TREATMENT BY MY PHYSICIAN. I NEVER HAVE HAD ANY SINUS TROUBLE NOR NASAL IRRITATION BEFORE. I QUESTIONED IF IT WAS FROM MY CPAP DEVICE PRIOR TO BEING NOTIFIED OF RECALL. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02758 | 2021-08-10 | 2021-08-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, AHP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AHX500H11 | AHX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED LUNG CANCER. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY SUBMITTED MDR 2518422-2021-02758-1 WITH INCORRECT SECTIONS B1, B2, H1, H6. CORRECTIONS TO PREVIOUS MDR ARE MADE IN THIS REPORT AS FOLLOWS. SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY PRODUCT PROBLEM WAS CHECKED IN PREVIOUS MDR) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY IT WAS BLANK) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6- HEALTH IMPACT CODE WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103147 | 2021-08-10 | 2020-04-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500H11C |  |  | THIS IS AN ESTIMATED DATE THAT I STOPPED USING MY PHILIPS RESPIRONICS CPAP MACHINE ON A REGULAR BASIS DUE TO RESPIRATORY IRRITATION THAT CAUSED ME TO HAVE TROUBLE BREATHING USING THE MACHINE. I HAVE SOUGHT TREATMENT FROM ENT SPECIALIST TO TRY AND DETERMINE THE CAUSE OF MY REDUCED QUALITY OF BREATHING SINCE STARTING MY THERAPY TO TRY TO RESTART MY THERAPY. READING THE SYMPTOMS ASSOCIATED WITH ISSUES WITH THE MACHINE, I HAVE SUFFERED ALL THESE SYMPTOMS AND THEY STILL PERSIST EVEN AFTER DISCONTINUING USE. I HAVE CONTINUAL ISSUES WITH MY UPPER SINUSES (WHICH INCLUDES POST NASAL DRIP AND CONGESTION) FREQUENT HEADACHES AND A FEELING OF TIGHTNESS IN MY CHEST. I THOUGHT I HAD A DEVELOPING HEALTH CONDITION THAT WAS IMPACTING MY THERAPY BUT NOW BELIEVE AFTER READING THESE REPORTS THAT THE THERAPY CAUSED MY CONDITION. I BEGAN USE OF PHILIPS RESPIRONICS DREAMSTATION IN (B)(6) 2019 AND BEGAN EXPERIENCING PROBLEMS WITH MY RESPIRATORY SYSTEM BETWEEN (B)(6) 2019 AND (B)(6) 2020. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02878 | 2021-08-10 | 2021-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02881 | 2021-08-10 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/CELL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11C | DSX400H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED THAT HE HAS SEEN BLACK PARTICLES IN THE WATER CHAMBER ON HIS DEVICE AND IS ALSO COMPLAINING THAT THERE IS A WHISTLING SOUND THAT THE MACHINE MAKES WHILE IN USE.. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE 1 ERROR IS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED DUE TO 400 MODEL . INITIAL MDR SECTION A1,B3,B5,H7,H9 IS CORRECTED AND H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02756 | 2021-08-10 | 2021-08-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED NOSE CANCER. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON OCTOBER 09, 2023 AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED NOSE CANCER WHILE USING THIS DEVICE. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGED CANCER. AFTER SECOND ATTEMPT ON 07/19/2022 THE CUSTOMER HUNG UP THE CALL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION E1 WAS UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02882 | 2021-08-10 | 2021-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02713 | 2021-08-10 | 2021-06-24 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX700H11 | DSX700H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED DYSPNEA AND VOMITING. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED DYSPNEA AND VOMITING. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE EVENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02757 | 2021-08-10 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED KIDNEY CANCER. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED KIDNEY CANCER. THE DEVICE WAS RETURNED TO THE THIRD PARTY SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED . THERE WERE ONE ERROR FOUND. THE THIRD PARTY SERVICE CENTER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02712 | 2021-08-10 | 2021-06-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED A SEVERE ASTHMA AND CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD) EXACERBATION. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02755 | 2021-08-10 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT WAS DIAGNOSED WITH BRONCHITIS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF PRESCRIBED INHALERS, (B)(6), AND ANTIBIOTICS. THE DETAILS OF THE DEVICE'S SERIAL NUMBER ARE UNKNOWN. IF FURTHER INFORMATION REGARDING THE DEVICE BECOME AVAILABLE, IT WILL BE SUBMITTED IN THE FOLLOW-UP REPORT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT WAS DIAGNOSED WITH BRONCHITIS. THE REPORTED EVENT OF SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT WAS DIAGNOSED WITH BRONCHITIS AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. THE PATIENT HAS SWAPPED OUT THE DEICE AT THE DISTRIBUTOR AND AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02759 | 2021-08-10 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500S11 | DSX500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED PULMONARY HYPERTENSION. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT DID NOT REPORT TO RECEIVE MEDICAL INTERVENTION. THE PATIENT ALSO ALLEGED NASAL/THROAT IRRITATION OR SORENESS, SINUS INFECTIONS CONTINUED FOR SEVERAL WEEKS, NAUSEA AFTER WEARING MASK, PULMONARY HYPERTENSION, HEART IS ENLARGED ON ONE SIDE. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL INVESTIGATION SHOWED THAT THE LOCATED A LIGHT AMOUNT OF BEIGE DUST OR DIRT CONTAMINATE AND FIBER CONTAMINATE ON THE OUTER PANELS, SURFACE CRACKS AND CREVASSES, IN AND AROUND THE AIR INLET AREA, IN THE AIR INLET CANAL, AROUND THE SD CARD SLOT AREA, AROUND THE MODEM AREA, AND IN AND AROUND THE OUTLET PORT. THE AIR INLET CANAL HAS A SMALL CLEAR, LIGHT GREY, AND BLACK CONTAMINATE. THE MANUFACTURER VISUALLY INSPECTED THE INTERNAL INVESTIGATION SHOWED THAT THE UNIT FOUND A VERY SMALL AMOUNT OF BEIGE DUST OR DIRT CONTAMINATION ON THE TOP OF THE BLOWER BOX, NEAR THE AIR INTAKE AREA, AND IN THE BLOWER BOX CHIMNEY. THE MANUFACTURER REMOVED THE BLOWER BOX LID TO FIND THAT THE BOWER AND INTERNAL BLOWER BOX AREA CONTAINED A VERY SMALL AMOUNT OF BEIGE DUST OR DIRT CONTAMINATE. THE MANUFACTURER FOUND MINOR BEIGE DUST OR DIRT CONTAMINATE ON THE INSIDE AND OUTSIDE OF THE UNIT. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. THERE IS NO VISIBLE DAMAGE OR FUNCTIONALITY FAILURES OF THE DEVICE, WHICH SUGGEST THAT THE SOURCE OF CONTAMINATION WAS EXTERNAL TO THE DEVICE. SECTION H6 WERE UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02883 | 2021-08-10 | 2021-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO THE DEVICE'S SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02880 | 2021-08-10 | 2021-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO THE DEVICE'S SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02884 | 2021-08-10 | 2021-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUM/CELL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX700H11C | DSX700H11C |  | A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO THE DEVICE'S SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02886 | 2021-08-10 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES AND COUGH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE - THE PATIENT ALLEGED VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. CORRECTED INFORMATION WAS PROVIDED IN A1 (PREVIOUSLY PATIENT INITIALS WAS NOT INCLUDED) AND SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02879 | 2021-08-10 | 2021-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO THE DEVICE'S SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02885 | 2021-08-10 | 2021-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11 | DSX500T11 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO THE DEVICE'S SOUND ABATE FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CF1 PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103204 | 2021-08-11 | 2021-04-05 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | EXPERIENCED BRAIN FOG, SEVERE HEADACHES, NAUSEA/QUEASINESS, DIZZINESS. STOPPED USING THE RECALLED PHILIPS DREAMSTATION AUTO CPAP ON (B)(6) 2021 AND ALL SYMPTOMS HAVE GRADUALLY STOPPED. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103208 | 2021-08-11 | 2021-08-03 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS INC. | OLYMPIA RESPIRATORY CONFIRMED |  |  | WOKE UP FEELING AIR HUNGER, LIGHT HEADED, HEADACHE, DRY COUGH, WEAK VOICE, FEEL FEVERISH WITHOUT FEVER, SWEATING, DIZZY, FOGGY BRAIN, HEAVY CHEST, HEAVY BREATHING WITH MINIMAL EXERTION AND NAUSEA. I SOUGHT MEDICAL ATTENTION AS I FELT LIKE I WASN'T GETTING ENOUGH OXYGEN. I WAS GIVEN AN EKG AND COVID TEST. (B)(4) DIAGNOSED BRONCHITIS BUT TOLD ME TO CALL RIGHT AWAY IF SYMPTOMS BECOME WORSE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103180 | 2021-08-11 | 2021-03-08 | I, n, j, u, r, y | Voluntary report | CPAP DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. |  |  |  | I RECEIVED A PHILIPS RESPIRONICS CPAP MACHINE AROUND EARLY (B)(6) 2021. ABOUT A MONTH AFTER USE I NOTICED I WAS HAVING TROUBLE BREATHING, SPITTING OUT BLACK/GRAY STUFF THROUGHOUT THE NIGHT, THROWING UP BLACK TINTED PHLEGM EVERY MORNING AFTER USING CPAP, FEELING LIKE I WAS CHOKING, NOT FEELING LIKE I SLEPT AND HAVING TO INCREASE EMERGENCY ASTHMA INHALER USAGE. I TOLD MY DOCTOR I THOUGHT THE MACHINE WAS MAKING THINGS WORSE AND SHE DECREASED MY PRESSURE TO 4 (FROM 5) AND ENCOURAGED ME TO CLEAN IN MORE FREQUENTLY (I WAS ALREADY CLEANING IT DAILY). SO I WENT ON WITH THE NEW ADJUSTMENTS AND THE SAME PROBLEMS CONTINUED AND BEGAN TO GET WORSE OVER TIME. THEN I ASSUMED THAT MAYBE THE BLACK AND GRAY STUFF WERE FROM NOSEBLEEDS I GET BECAUSE IT TASTED METALLIC BUT I JUST NEVER RECALLED SPITTING UP GRAY/BLACK STUFF IN THE PAST FROM IT. THE BLACK DEBRIS STARTED TO GATHER AROUND MY MOUTH, IN MY GUMS, IN MY SALIVA IN ADDITION TO WHAT I WAS THROWING UP. I EVENTUALLY STOPPED USING MY MACHINE EVERY NIGHT AND BEGAN DOING IT EVERY OTHER NIGHT BECAUSE IT WAS MAKING ME FEEL SO SICK. I NOTICED THE NIGHTS I DIDN'T USE IT I FELT MORE RESTED AND WAS NOT THROWING UP AS MUCH BLACK STUFF. THE NIGHTS I DID USE IT THE ISSUES STATED ABOVE CONTINUED. I FOUND OUT THROUGH A PATIENT OF MINE (I'M A MENTAL HEALTH THERAPIST) THAT THEIR CPAP MACHINE WAS RECALLED. I ASKED THEM WHAT BRAND AND THEY SAID PHILIPS RESPIRONICS. I ASKED WHEN THEY HEARD ABOUT THAT AND THEY SAID ABOUT 2 MONTHS AGO. I FOUND OUT FROM MY PATIENT THAT IT WAS OFFICIALLY RECALLED SOME TIME IN JUNE OF 2021 TO THEIR KNOWLEDGE. NO ONE CONTACTED ME TO TELL ME. I REACHED OUT TO MY DOCTOR'S OFFICE ABOUT THE MATTER WHO DID NOT SAY ANYTHING ABOUT KNOWING ABOUT THIS. THERE WAS NO COMMUNICATION FROM PHILIPS RESPIRONICS OR WHERE I GET MY EQUIPMENT FROM. IT WAS DESCRIBED AS "BLACK DEBRIS FROM THE FOAM OR CERTAIN CHEMICALS RELEASED INTO THE DEVICE'S AIR PATHWAY MAY BE INHALED OR SWALLOWED BY THE PERSON USING THE DEVICE". I FOUND OUT ON (B)(6) 2021. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103189 | 2021-08-11 | 2021-01-31 | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | SUFFERING FROM SEVERE DRY EYES, SINUS ISSUES, AND WHEEZING. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-08
MW5103193 | 2021-08-11 | 2018-10-18 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | HAVING HEADACHES, SINUS ISSUES AND EYE IRRITATION AFTER USING CPAP MACHINE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103195 | 2021-08-11 | 2021-08-06 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS DREAMSTATION AUTO CPAP HUMCELL DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  | REF. # DSX500H11C |  | I HAVE A PHILIPS DREAMSTATION CPAP MACHINE SERIAL #: (B)(4). ON A COUPLE OF OCCASIONS OVER THE PAST COUPLE OF MONTHS, I HAVE FOUND A BLACK GRAPHITE-LIKE SUBSTANCE CLOGGING THE TEMPORARY AND PERMANENT AIR FILTERS. THE SUBSTANCE PERMANENTLY STAINS THE FILTERS AND CANNOT BE WASHED OUT. THE FILTERS MUST BE DESTROYED. DISTURBED ABOUT THE SUBSTANCE, I NOTIFIED (B)(6) WHO TOLD ME NOT TO WORRY ABOUT IT. LAST WEEK I RECEIVED A LETTER FROM PHILIPS STATING THAT THE SOUND ABATEMENT FOAM THEY USED WAS DEGRADING AND COULD BE INHALED. THEY INSTRUCTED ME TO COMPLETE A FORM AND NOTED THE COMPANY WAS "DEPLOYING A PERMANENT CORRECTIVE ACTION." PHILLIPS DID NOT STATE WHAT THE CORRECTIVE ACTION WOULD BE NOR DID THEY PROVIDE A TIME-FRAME FOR THIS CORRECTIVE ACTION. HOW DO I KNOW IF I INHALED ANY OF THE CHEMICALS? I AM APPRAISING THE FDA OF THIS SITUATION. I KNOW NOTHING ABOUT THE REMAINING TEST QUESTIONS ON THE FORM. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103197 | 2021-08-11 | 2021-08-03 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | J2991273810DF |  |  | BREATHING DIFFICULTY; HAD ASTHMA LIKE ATTACK. ALSO HAVE BEEN HAVING THESE TERRIBLE COUGHING FITS THROUGH OUT THE DAY, AND SINUS CONGESTION EVERY MORNING AFTER I DISCONNECT FROM MY DREAMSTATION CPAP. FOUND OUT LAST FRIDAY FROM A FB POST THAT THERE WAS A PROBLEM WITH MY CPAP. CALLED MY DME VENDOR AND THEY CONFIRMED. ASKED WHY I WAS NOT NOTIFIED FROM COMPANY OR DME VENDOR, THEY CLAIM THEY JUST FOUND OUT. WAS TOLD TO SEE MY LUNG SPECIALIST. WAS ASKED TO REGISTER MY MACHINE WITH PHILIPS RESPIRONICS AND I DID ON FRIDAY (B)(6) 2021. SAW DOCTOR THIS MORNING (B)(6) 2021. I WAS ADVISED NOT TO USE MACHINE. THEY SENT MY DME A NEW REQUEST BUT THE DME HAS NO SOLUTION OR MACHINES TO REPLACE MY RECALLED MACHINE. WAS TOLD TO BUY NEW ONE BECAUSE MANUFACTURE DOES NOT KNOW WHEN THEY WILL HAVE REPLACEMENTS. I RECEIVED AN E-MAIL FROM MANUFACTURER DOES NOT HELPFUL. I WILL HAVE NO OTHER CHOICE THAN TO BUY A DIFFERENT ONE. IN THE MEAN TIME I STILL HAVE COUGHING AND CHEST TIGHTNESS, AND I DO NOT KNOW WHAT PERMANENT DAMAGE I HAVE SUSTAINED. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-08
MW5103199 | 2021-08-11 |  | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | AFTER USING MY CPAP MACHINE FOR OVER A YEAR. NOTICE SHORTNESS OF BREATH AND LUNGS FEEL FULL AND CAN HAVE A HARD TIME TAKING DEEP BREATH. DID NOT HAVE THIS FEELING BEFORE BUYING THIS CPAP MODEL. HAVE A DR'S APPOINTMENT ON (B)(6) 2021 TO CHECK MY LUNGS. HAVE SHORTNESS OF BREATH. DOING SIMPLE OUT DOOR PROJECT, OR EVEN FAST PACE WALK. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103207 | 2021-08-11 | 2020-06-01 | I, n, j, u, r, y | Voluntary report | CPAP DREAM STATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS, INC. |  |  |  | I HAVE BEEN USING THE DREAM STATION CPAP FOR A COUPLE OF YEARS NOW AND THIS PRODUCT WAS JUST RECALLED. IN THE PAST YEAR OR SO I HAVE BEEN EXPERIENCING HEADACHES, SORE THROAT, TIGHTENING IN MY CHEST AND WORSENING ASTHMA CONDITIONS. I AM NERVOUS THAT IT MAY BE BECAUSE THE CPAP MACHINE DEFECTS I HAVE BEEN USING FOR A WHILE NOW. I AM NOT SURE WHAT NEXT STEPS WITH MY PHYSICIAN I SHOULD TAKE, BUT WANT YOU TO BE AWARE OF THIS. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103198 | 2021-08-11 | 2020-04-18 | I, n, j, u, r, y | Voluntary report | PHILIPS CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | J23986459703D OR H23984710C774 |  |  | RECEIVED A LETTER FROM PHILIPS RESPIRONICS - SLEEP AND RESPIRATORY CARE REGARDING MEDICAL DEVICE RECALL IN APRIL 2021. ON (B)(6) 2020 WENT TO (B)(6) FOR PNEUMONIA, WAS GIVEN LEVOFLOXACIN 750MG, 1 TABLET. ON (B)(6) 2020 WENT TO (B)(6) FOR PNEUMONIA AGAIN, ADMITTED FOR 5 DAYS, RIGHT MIDDLE LOBE PNEUMONIA/ SEPSIS/ MUCUS PLUGGING OF BRONCHI; COUGH WITH HEMOPTYSIS; PLEURITIC CHEST PAIN; LEUKOCYTOSIS/ PLEURISY, GIVEN AMOXICILLIN-CLAVULANATE ONE TABLET FOR 5 DAYS. DISCHARGED ON (B)(6) 2020. ON (B)(6) 2020 WENT TO EMERGENCY STILL NOT BREATHING GOOD AND CHEST PAIN AND WEAK, PLEURAL EFFUSION, RIGHT; PLEURISY ON (B)(6) 2020 WENT TO (B)(6) EMERGENCY IN (B)(6) TO SECOND OPINION. ADMITTED FOR 3 WEEKS, ACUTE HYPOXIC RESPIRATORY FAILURE; SEPSIS DUE TO PNEUMONIA; RIGHT SIDED RECURRENT LOCULATED PLEURAL EFFUSION; INDETERMINATE QUANTIFERON RESULT; NEEDED A THORACENTESIS. HAS RECURRENT PLEURAL EFFUSION AND AFTER MULTIPLE THORACENTESES HAD A VATS PROCEDURE DONE WITH COMPLETE DECORTICATION AFTER EVACUATION OF LARGE ABSCESS AND RELEASE OF TRAPPED LUNG ON (B)(6) 2020. PROCEDURES: (B)(6) 2020 - US R THORACENTESIS WITH 800 CC FLUID REMOVED (B)(6) 2020 - USR THORACENTESIS WITH 150 CC FLUID REMOVED (B)(6) 2020 - CT GUIDED DRAINAGE OF LOCULATED R PLEURAL FLUID POCKET (B)(6) 2020 - RIGHT VATS COMPLETE DECORTICATION AFTER BIG ABSCESS EVACUATION, RIGHT VATS RELEASE OF TRAPPED LUNGS DISCHARGED ON (B)(6) 2020. AS OF (B)(6) 2021 STILL NOT BACK TO NORMAL, SHORTNESS OF BREATH, FATIGUE, CONTENTIVE ISSUES, SEVERAL FALLS. SEEN A NEUROLOGIST ON (B)(6) 2021. THERE IS NERVE DAMAGE IN BOTH LOWER LEGS AND SOME MUSCLE DAMAGE, MORE IN THE LEFT LEG. ALSO NERVE DAMAGE IN LOWER BAG. HAD MULTIPLE BLOOD TEST DONE, ON THE FOLLOW UP APPT ON (B)(6) 2021. WE WERE INFORMED OF POSSIBLE LYME DISEASE, LOW B12, ANEMIC, AND DIABETES. LYME DISEASE FIRST POSITIVE THEN NEGATIVE. APPOINTMENT ON (B)(6) WITH PULMONOLOGIST IN (B)(6). STILL HAVING SHORTNESS OF BREATH AND VERY FATIGUED. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02900 | 2021-08-11 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VARIATION IN THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE DELIVERY WAS OBSERVED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE EVALUATION HAS NOT BEEN COMPLETED BY THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VARIATION IN THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE DELIVERY WAS OBSERVED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. THE VENTILATOR'S SOFTWARE WAS RE-LOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02858 | 2021-08-11 | 2021-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03108 | 2021-08-11 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG AND THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03104 | 2021-08-11 | 2021-06-16 | I, n, j, u, r, y | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260B | 1054260B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE BLOOD OXYGEN LEVEL OF A PATIENT TO DECREASE. THE PATIENT WAS ADMITTED TO THE HOSPITAL. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. ADDITIONAL INFORMATION WAS RECEIVED, SECTION B5 WAS CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING TROUBLE WAKING UP , WOKE UP VERY CONFUSED AND WAS ADMITTED TO THE HOSPITAL AND HAS LOW BLOOD OXYGEN RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. PATIENT HAD ANOXIA EVENT MOST LIKELY, ALSO HAD IRRITATION IN THE THROAT. THE REPORTED EVENT OF PATIENT HAD TROUBLE WAKING UP, WOKE UP CONFUSED AND WAS ADMITTED TO THE HOSPITAL OVERNIGHT FOR HYPOXEMIA AND ANOXIC EVENT, AND ALSO HAD THROAT IRRITATION AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS NOT ASSESSABLE DUE TO INSUFFICIENT INFORMATION. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED AND CORRECTED BY CAPTURING HEALTH EFFECT - CLINICAL CODE AND HEALTH EFFECT - IMPACT CODE WHICH WERE INCORRECT IN INITIAL REPORT. SECTION D4 WAS CORRECTED BY CAPTURING MODEL NUMBER AND CATALOG NUMBER WHICH WERE INCORRECT IN INITIAL REPORT. | Y | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03105 | 2021-08-11 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS FAILING CALIBRATION DURING TESTING. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE WAS EVALUATED BY A THIRD PARTY ENGINEER AND THE CUSTOMERS COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03109 | 2021-08-11 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03046 | 2021-08-11 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03045 | 2021-08-11 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUES. | THE MANUFACTURER FILED A DUPLICATE REPORT FOR THIS MALFUNCTION ON MDR 2518422-2021-03109. REPORT, MDR WAS A DUPLICATE AND WAS FILED IN ERROR. NO DUPLICATE MALFUNCTION OCCURRED AND ALL REPORTING IS CORRECT ON MDR 2518422-2021-03109. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02871 | 2021-08-11 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A "VENTILATOR INOPERATIVE" ALARM WAS OBSERVED BY THE TECHNICIAN. THE DEVICE'S BLOWER MOTOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02873 | 2021-08-11 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | DS2200X11B | DS2200X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. THE DEVICE WAS RETURNED UNREPAIRED AT THE REQUEST OF THE CUSTOMER. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02857 | 2021-08-11 | 2021-07-25 | I, n, j, u, r, y | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC |  |  |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW 5102233) ALLEGING A VENTILATOR'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT WAS HOSPITALIZED FOR SEIZURES AND INCREASED CARBON DIOXIDE LEVELS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING A VOLUNTARY MEDWATCH (MW 5102233) ALLEGING A VENTILATOR'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT WAS HOSPITALIZED FOR SEIZURES AND INCREASED CARBON DIOXIDE LEVELS. AFTER RECEIVING FURTHER INFORMATION FROM THE PATIENT IT IS DETERMINED THAT THE INITIAL REPORT SHOULD HAVE BEEN FILED AS ADVERSE EVENT ONLY. | Y | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02874 | 2021-08-11 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | DS2200X11B | DS2200X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD, PROPORTIONAL VALVE, AND THREE WAY SOLENOID VALVE WERE REPLACED TO ADDRESS THE ISSUES. | THE MANUFACTURER HAD PREVIOUSLY REPORTED THAT A DEVICE'S SYSTEM BOARD, PROPORTIONAL VALVE, AND THREE WAY SOLENOID VALVE WERE REPLACED TO ADDRESS SERVICE REQUIRED CODES FOUND IN THE ERROR LOG. THIS WAS INCORRECT. THE DEVICE'S SYSTEM BOARD WAS NOT REPLACED. THE DEVICE'S SYSTEM BOARD TUBING AND BLOWER TUBING WERE REPLACED ALONG WITH THE PROPORTIONAL VALVE AND THREE WAY SOLENOID VALVE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02898 | 2021-08-11 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040004 | 1040004 |  | THE MANUFACTURER RECEIVED INFORMAITON ALLEGING A VENTILATOR WAS ALARMING FOR A "VENTILATOR INOPERATIVE" ALARM CONDITION AND THE PATIENT REQUIRED ARTIFICIAL RESPIRATION VIA A RESUSCITATION BAG. THE PATIENT WAS PLACED ON ANOTHER DEVICE WITHOUT HARM. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. | SECTION B2 REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE CHECKED INCORRECTLY IN PREVIOUS REPORT. IT HAS BEEN UNCHECKED IN THIS REPORT. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02896 | 2021-08-11 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2110X15B | IN2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VARIATION IN THE VENTILATOR'S DELIVERY OF PRESSURE SUPPORT WAS OBSERVED. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VARIATION IN THE VENTILATOR'S DELIVERY OF PRESSURE SUPPORT WAS OBSERVED. ADDITIONAL INFORMATION RECEIVED INDICATES THE VENTILATOR'S SOFTWARE WAS RE-LOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03060 | 2021-08-11 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S PRESSURE INCREASED. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S SENSOR BOARD NEEDS REPLACED DUE TO CONTAMINATION. | IT WAS INITIALLY REPORTED, DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S SENSOR BOARD NEEDS REPLACED DUE TO CONTAMINATION. THE DEVICE'S INLET AIR PATH ASSEMBLY WAS ALSO REPLACED DUE TO A TEST STEP FAILURE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02901 | 2021-08-11 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VARIATION IN THE VENTILATOR'S PRESSURE DELIVERY WAS OBSERVED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE EVALUATION HAS NOT BEEN COMPLETED BY THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VARIATION IN THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE DELIVERY WAS OBSERVED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. THE VENTILATOR'S SOFTWARE WAS RE-LOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03005 | 2021-08-11 | 2021-08-16 | I, n, j, u, r, y | Manufacturer report | BIPAP PRO BIFLEX W/HUM SYS ONE 60SRS,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DS660HS | DS660HS |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON (B)(6) 2024 AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO BIPAP DEVICE SOUND ABENDENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGED KIDNEY DISEASE TOXICITY, ACUTE KIDNEY INJURY, ACUTE HYPOXIC RESPIRATORY FAILURE, SHORTNESS OF BREATH, TACHYCARDIA, NOSE BLEEDS, STROKE AND MUCUS. THE SECTIONS B1, B2, H1, H2 AND H6 HAVE BEEN UPDATED OR CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR ADVERSE EVENTS AND PRODUCT PROBLEM. (ONLY PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECTS - CLINICAL CODES AND HEALTH EFFECTS - IMPACT CODE WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02996 | 2021-08-11 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE WITH THE DEVICE'S LCD SCREEN. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS EVALUATED AT THE MANUFACTURER'S SERVICE CENTER AND THE LCD SCREEN WAS UNABLE TO BE VIEWED. THE DEVICE'S LCD SCREEN NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR'S LCD SCREEN WAS REPLACED DUE TO A BLANK SCREEN ISSUE. DURING FURTHER EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER A "SERVICE REQUIRED" ALARM CONDITION WAS OBSERVED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02995 | 2021-08-11 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02994 | 2021-08-11 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED BY A THIRD PARTY TO ADDRESS A MALFUNCTION. ADDITIONAL INFORMATION RECEIVED FROM THE THIRD PARTY SERVICE CENTER STATES THE DEVICE'S BLOWER MOTOR WAS ALSO REPLACED TO ADDRESS A MALFUNCTION. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02872 | 2021-08-11 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A "VENTILATOR INOPERATIVE" ALARM WAS OBSERVED BY THE TECHNICIAN. THE DEVICE'S BLOWER MOTOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03006 | 2021-08-11 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | 760P | 760P |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VOMITING, HEADACHE, NASAL/THROAT IRRITATION AND VISUALIZATION OF BLACK PARTICLES IN WATER TANK RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H8: HEALTH EFFECT - CLINICAL CODES WAS MISSED IN INITIAL REPORT, WHICH IS UPDATED IN THIS REPORT. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02825 | 2021-08-11 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING HEADACH AND SINUS INFECTION RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION A1 AND H6(HEALTH EFFECT - CLINICAL CODE) WAS MISSED IN THE INITIAL REPORT, WHICH IS UPDATED IN THIS REPORT. SECTION H6 TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS IS UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103216 | 2021-08-11 | 2021-08-07 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC | DSXHCP |  |  | INCREASE IN COUGHING, CHEST PRESSURE, SINUS ISSUES, NAUSEA, HEADACHES, UPPER AIRWAY IRRITATION, DIZZINESS, NON ALLERGIC RHINITIS, EAR PAIN AND PRESSURE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02853 | 2021-08-11 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. CORRECTED B5 SECTION SHOULD BE REPORTED AS : THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALSO ALLEGED TO HAVE HEADACHES. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. ( CLINICAL CODE IS UPDATED IN THIS REPORT). | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103214 | 2021-08-11 |  | I, n, j, u, r, y | Voluntary report | PHILIPS CPAP DREAM STATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / / RESPIRONICS, INC. |  |  |  | THIS HAPPENED OVER SEVERAL MONTHS, PLUS. I WAS HOSPITALIZED (B)(6) 2020, (B)(6) 2021, (B)(6) 2021, AGAIN (B)(6) 2021. SHORTNESS OF BREATH, COUGH, CHEST PRESSURE, RESPIRATORY TRACT. NAUSEA, HEADACHE, ETC. THE DOCTORS DID NOT RECOGNIZE THAT IT MIGHT BE MY CPAP MACHINE EVEN THOUGH IT HAD BEEN PUT ON RECALL. I DIDN'T KNOW IT WAS ON RECALL. I WOULD BE UNABLE TO GET MY BREATH SO BAD I WOULD FINALLY HAVE TO GO TO THE HOSPITAL. I USED MY PHILIPS CPAP FAITHFULLY AT HOME BUT DID NOT TAKE IT IN TO THE HOSPITAL WITH ME. I WOULD BE IN THE HOSPITAL 7 OR 8 DAYS AND BEGIN TO GET BETTER. THEN I CAME HOME AND STARTED BACK ON MY CPAP AND WITH IN A FEW DAYS I WAS RIGHT BACK IN THE HOSPITAL AGAIN. NOT A PLEASANT EXPERIENCE. THE LAST TIME I WAS IN I CAME HOME AND GOT TIGHTNESS AND BREATHLESS AGAIN AND WENT TO AN IMMEDIATE CARE UNIT. I ASK FOR SOME SOLUTION FOR A NEBULIZER AS MY DOCTORS COULD NOT PRESCRIBE IT WITH OUT SEEING ME WHICH WAS NOT POSSIBLE FOR SEVERAL WEEKS. I WAS IN TROUBLE AND ABOUT TO GO BACK IN THE HOSPITAL. THE NEBULIZER HELPED ME STAY OUT OF THE HOSPITAL SO FAR. I HAVE NOT USED MY PHILIPS CPAP SINCE I GOT THE RECALL NOTICE ON MY FRONT PORCH ON (B)(6) 2021. I AM SOME BETTER BUT FEEL LIKE I AM CLOGGED UP WITH MUCUS AND CAN'T GET IT UP VERY WELL. ALSO NOT SLEEPING VERY WELL. I HAVE TRIED TO REGISTER MY CPAP AND HAVE GONE ON THE COMPUTER SEVERAL TIMES AND CAN'T SEEM TO FIND WHERE TO REGISTER OR GET A FORM TO REGISTER. I WILL TRY TO CALL THE NUMBER I HAVE. I'VE HAD MY MACHINE SINCE 2018 AND DON'T THINK MY INSURANCE WILL COVER ANOTHER MACHINE. I'VE USED A CPAP FOR ABOUT 20 YEARS. ALL THE SUPPLIERS SAY THERE IS A SHORTAGE OF MACHINES NOW. I WILL NOT USE A PHILIPS BRAND AGAIN. I DO NOT WANT A PHILIPS AGAIN. I HAVE BEEN THROUGH HELL WITH THE HOSPITAL AND DOCTORS NOT UNDERSTANDING. SOME DO NOT TREAT PEOPLE OTHER THAN A NUMBER AND A DOLLAR SIGN. (B)(6) TO COME IN TO MY HOSPITAL ROOM AND HARDLY SLOW DOWN BEFORE GOING OUT AGAIN. THEY TOOK 3 BLOOD TEST SOME DAYS. ONCE THE LAB TECK BARELY GOT OUT OF MY ROOM AND ANOTHER CAME IN TO DRAW BLOOD. MINUTES APART. YOU WOULD THINK THE DOCTORS WOULD BE MORE ORGANIZED THAN THAT. GET YOUR ORDER IN EARLY SO ALL BLOOD WORK COULD BE DONE WITH ONE DRAW. SEEMS I WAS JUST A PUNCHING BAG. HOW COULD MY BLOOD CHANGE IN A DAY? ONE A DAY IS TO MUCH BUT THE DOCTORS TOLD ME THAT'S WHAT YOU GET WHEN YOU COME TO THE HOSPITAL. WELL, WHAT I GO FOR IS TO GET HELP. I WAS PURPLE ALL OVER MY HANDS AND ARMS. MY CPAP WAS KILLING ME AND I DIDN¿T KNOW IT AND GOT NO HELP IN THE HOSPITAL. IS THERE NO WHERE TO GO FOR HELP? IT IS NOT EASY TO FILE A COMPLAINT ABOUT THIS SITUATION. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03015 | 2021-08-11 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING COUGH, HEAVINESS IN CHEST, BREATHING ISSUES, FILTER WAS DARKER BLACK COLOR RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 HEALTH EFFECT CLINICAL CODE CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03012 | 2021-08-11 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT ALSO ALLEGED BLACK PARTICLES IN AIRWAY FROM DEVICE AND HAVING HEADACHE, DIZZINESS, LIGHTHEADED AND SHORTNESS OF BREATH. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL PART OF THE DEVICE AND FOUND NO EVIDENCE OF CONTAMINATION. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION, AND THE FOAM APPEARS INTACT. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03017 | 2021-08-11 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT ALSO ALLEGED PARTICLES IN TUBING AND HUMIDIFIER CHAMBER, WITH SYMPTOMS OF LOSS OF SENSE OF SMELL, HEADACHES AND ALLERGIES. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION.THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND FOUND MINIMAL FINE DUST CONTAMINATION. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED IN THIS DEVICE. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND 5 ERROR CODES. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION BUT FINE DUST CONTAMINATION IN THE DEVICE, WHICH IS LIKELY FROM EXTERNAL SOURCE. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03010 | 2021-08-11 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING COUGH ,DRY MOUTH AND THROAT RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. ADDITIONAL INFORMATION WAS RECEIVED AND ADDED TO THE REPORT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER ON 04/19/2023 FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED ON 06/14/2023. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL PART OF THE DEVICE AND FOUND MINIMAL SIGNS OF DUST CONTAMINATION. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED BY THE MANUFACTURER AND THEY OBSERVED UNKNOWN WHITE AND BLACK CONTAMINATION (DUST-LIKE), INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM, AT THE AIR INPUT OF THE BLOWER BOX. ALSO, A POTENTIAL TINY BLACK HAIR, CONSISTENT WITH KERATIN AT THE AIR OUTLET OF THE BLOWER BOX OR FIBER WAS AT THE INPUT OF THE BLOWER BOX AND LIGHT DUST-LIKE CONTAMINATION ON TOP OF THE BLOWER MOTOR, INDICATING POTENTIAL WATER INGRESS BLOWER MOTOR SEAL AND ON TOP OF THE BLOWER BOX. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION, THERE IS NO VISIBLE DAMAGE OR FUNCTIONALITY FAILURES OF THE DEVICE, WHICH SUGGEST THAT THE SOURCE OF CONTAMINATION WAS EXTERNAL TO THE DEVICE. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03011 | 2021-08-11 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED PARTICLES IN TUBING/CHAMBER, DEVICE/POWER CORD OR SUPPLY TOO HOT TO TOUCH, NASAL/THROAT IRRITATION OR SORENESS, SHORTNESS OF BREATH, EYE IRRITATION. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03022 | 2021-08-11 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED BY THE MANUFACTURER. THE MANUFACTURER FOUND NO SIGNS OF CONTAMINATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND VERY LIGHT BEIGE DUST OR DIRT CONTAMINATE. THE MANUFACTURER ALSO LOCATED LIGHT AND DARK FIBER OR HAIR CONTAMINATE. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS LOGGED. THE MANUFACTURER CONCLUDES NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. THE MANUFACTURER ALSO CONFIRMED CONTAMINATES WERE SOURCED FROM EXTERNAL ENVIRONMENTAL CONDITIONS. IN THE INITIAL REPORT SECTION B5 THE PATIENT ALLEGED VISUALIZATION OF PARTICLES, THROAT IRRITATION AND COUGHING THAT WAS NOT ADDED TO THE REPORT. IN THIS REPORT WHICH HAS BEEN UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03024 | 2021-08-11 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT ALSO ALLEGED COUGHING, PHLEGM, IRRITATION, SHARPNESS IN THROAT, FATIGUE, SHORTNESS OF BREATH, DARK SPOTS ABOVE THE LID OF THE WATER CHAMBER, HIGH BLOOD PRESSURE. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL INVESTIGATION SHOWED THAT THERE WERE NO SIGNS OF CONTAMINATION ON THE OUTSIDE OF THE DEVICE. THE INTERNAL INVESTIGATION SHOWED THAT THERE WERE NO SIGNS OF CONTAMINATION OR SOUND ABATEMENT FOAM DEGRADATION INSIDE OF THE DEVICE. THE ALLEGATIONS OF BLACK SPOTS ABOVE THE LID OF THE WATER TANK ARE UNABLE TO BE VERIFIED AS NO HUMIDIFIER WAS RETURNED WITH THE DEVICE. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS OBSERVED IN THIS DEVICE. SECTION H6 WERE UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03023 | 2021-08-11 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING BLACK PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION G3 DATE WAS INCORRECT IN THE FIRST FOLLOW UP REPORT AND THE CORRECT DATE WILL BE (01-JUL-2021). | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03020 | 2021-08-11 | 2021-07-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A LUNG INFECTION, AND NODULES ON THEIR LUNGS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03016 | 2021-08-11 | 2021-08-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED BLACK PARTICLES IN THEIR NOSE, AND EXPERIENCED HEADACHES, DIZZINESS, SORE THROAT, BLURRY VISION, PARTIAL BLINDNESS, UNABLE TO WALK AND/OR DRIVE SOMETIMES, VISUAL DISTURBANCE, AND HAD THEIR WALKING. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM (PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR) SECTION B2 WAS MADE OTHER SERIOUS (PREVIOUSLY IT WAS BLANK) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY SECTION H6 HEALTH EFFECT- IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03018 | 2021-08-11 | 2021-06-29 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SINUS INFECTIONS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. IN PREVIOUS REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SINUS INFECTIONS AND HOSPITALIZED FOR BREATHING. THE REPORTED EVENT AND ITS SEVERITY WERE REVIEWED BY THE MANUFACTURER'S CLINICAL EXPERT. BASED ON INFORMATION AVAILABLE AT THIS TIME, THERE IS NOT ENOUGH EVIDENCE TO SUPPORT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE ALLEGED SERIOUS INJURY. THERE IS INSUFFICIENT INFORMATION RELATED TO DEVICE USE AND MAINTENANCE, TIMELINE OF EVENTS, RELEVANT PAST MEDICAL HISTORY, AND MEDICAL INTERVENTION INFORMATION. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03019 | 2021-08-11 | 2021-07-15 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A COLLAPSED LUNG. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY SUBMITTED MDR 2518422-2021-03019 -1 WITH INCORRECT SECTIONS B1, B2, H1, H6. CORRECTIONS TO PREVIOUS MDR ARE MADE IN THIS REPORT AS FOLLOWS. SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY PRODUCT PROBLEM WAS CHECKED IN PREVIOUS MDR) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY IT WAS BLANK) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6- HEALTH IMPACT CODE WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03053 | 2021-08-11 | 2021-04-30 | D, e, a, t, h | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED CANCER. THE PATIENT REPORTEDLY PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM AND CAUSED CANCER. THE PATIENT REPORTEDLY PASSED AWAY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. THE MANUFACTURER WAS NOT CAPTURED DATA IN SECTION H9-CORRECTION/REMOVAL REPORTING NUMBER IN PREVIOUS REPORT AND IT HAS BEEN CORRECTED IN THIS REPORT. SECTION A1 WAS UPDATED IN THIS REPORT. SECTION H6 HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03008 | 2021-08-11 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED SMALL BLACK PARTICLES FLOATING IN THE WATER TANK. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03007 | 2021-08-11 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTRE FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED DUE TO AGE. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03013 | 2021-08-11 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP PRO W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX600T11C | DSX600T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03014 | 2021-08-11 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON OCT 12, 2023, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES IN WATER CHAMBER, NASAL AND THROAT IRRITATION, SHORTNESS OF BREATH, NAUSEA, LIGHT-HEADEDNESS, DIZZINESS, MILD HEADACHE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGED OF RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE OTHER. SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY. TO MALFUNCTION. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03036 | 2021-08-11 | 2021-02-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/HT/CELL, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AUX400T15C | AUX400T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED SINUS INFECTIONS AND VOMITING. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF PRESCRIBED ANTIBIOTICS AND NASAL SPRAYS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT IS ALSO ALLEGING INFLAMED NERVE ENDINGS IN TEETH, FEELING UNWELL, BLOOD POISONING AND NECK SURGERY WHILE USING THE DEVICE. PATIENT ALLEGED THE DEVICE IS POISONING HIM. THE ALLEGATIONS OF INFLAMED NERVE ENDINGS, NECK SURGERY, AND BLOOD POISONING WERE REVIEWED BY THE POST MARKET SURVEILLANCE CLINICAL EXPERTS AND THE EVENTS ARE ASSESSED AS NOT RELATED TO THE DEVICE. UNDERLYING HEALTH CONDITIONS AS WELL AS UPDATED MANUFACTURER CONTACT INFORMATION HAS ALSO BEEN PROVIDED. PLEASE SEE SECTIONS B7, G1 AND H6 FOR THIS UPDATED INFORMATION. THE PREVIOUS REPORT WAS ALSO FILED AS A FOLLOW UP/FINAL, HOWEVER THE MANUFACTURER'S INVESTIGATION IS NOW ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. THE UPDATED CODING IN H6 REFLECTS THIS CHANGE. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03009 | 2021-08-11 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING BLACK STUFF (BLACK PARTICLES) RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03069 | 2021-08-11 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES IN AIRWAY RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTRE FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF DEVICE WAS INSPECTED. THE DEVICE POWERED ON, AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03031 | 2021-08-11 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES (B)(4). THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02829 | 2021-08-11 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, INTL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | INX500S15 | INX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02824 | 2021-08-11 | 2021-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03038 | 2021-08-11 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/HT/CELL, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | AUX400T15C | AUX400T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED CONGESTION, COUGH, AND PLASTIC BREAKING INTO TWO PIECES. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF PRESCRIBED ANTIBIOTICS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED PLASTIC BREAKING FROM MASK, COUGH, BLOCKED NOSE, CONGESTION, HEAVY HEADACHES AND PAINFUL EARS SINCE THIS EVENT AND PATIENT HAD A CT SCAN AND WAS TAKING PAIN KILLERS AND ANTIBIOTICS FOR SOME DAYS. THE REPORTED EVENTS COUGH, BLOCKED NOSE, CONGESTION, HEAVY HEADACHES AND PAINFUL EARS DUE TO THE DEVICE AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENTS ARE NOT MEET THE DEFINITION OF A REPORTABLE SERIOUS INJURY IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURER CONCLUDES NO FURTHER ACTION IS NECESSARY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03059 | 2021-08-11 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN INITIAL REPORT SECTION B5 WAS INCOMPLETELY MENTIONED AND THE UPDATED SECTION B5 SHOULD BE- THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING OF SEEING PARTICLES IN TUBING/CHAMBER AND HAVING NASAL/THROAT IRRITATION OR SORENESS. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO EVIDENCE OF CONTAMINATION. THE MANUFACTURER COMPLETED AN INTERNAL VISUAL INSPECTION. THE MANUFACTURER FOUND EVIDENCE OF LOOKED LIKE CONDENSATION OR WATER DROPLETS IN THE SOUND ABATEMENT FOAM ON THE LEFT SIDE OF THE BLOWER BOX. THE MANUFACTURER FOUND EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. SECTION D4, D9, G3, H3 AND H6 HAS BEEN UPDATED / CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03029 | 2021-08-11 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX200T11 | DSX200T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTRE FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED DUE TO AGE. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03068 | 2021-08-11 | 2021-06-22 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX500S11 | DSX500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 13, 2024, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER PREVIOUSLY REPORTED A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECOME DEGRADED AND THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT ALSO ALLEGED HAVING NASAL / THROAT IRRITATION OR SORENESS, DIFFICULTY IN BREATHING, / SHORTNESS OF BREATH, LIGHTHEADNESS, DIZZINESS, VOMITING, NAUSEA, PNEUMONIA, ASTHMA. THE PATIENT DID NOT REPORT TO RECEIVE MEDICAL INTERVENTION. SECTION B1 WAS CORRECTED FOR BOTH ADVERSE EVENT AND PRODUCT PROBLEM. (PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION. TO SERIOUS INJURY. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03033 | 2021-08-11 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES BLACK DEBRIS INSIDE THE CHAMBER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DIFFICULTY BREATHING. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 CORRECTED AND UPDATED IN THIS REPORT | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03026 | 2021-08-11 |  | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03027 | 2021-08-11 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING CAUSED HER TO HAVE A SLIGHT COUGH AND FOR HER TO TRY AND CATCH HER BREATH, SHE HAS SEEN SOME DARK SPOTS IN THE TUBING ALSO WHERE ITS ON HER NOSTRALS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING CAUSED HER TO HAVE A SLIGHT COUGH AND FOR HER TO TRY AND CATCH HER BREATH, SHE HAS SEEN SOME DARK SPOTS IN THE TUBING ALSO WHERE ITS ON HER NOSTRALS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. CORRECTION AND UPDATED INFORMATIONS WERE CAPTURED IN SECTION H6 OF THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03034 | 2021-08-11 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. SECTION B5 SHOULD BE PREVIOUSLY REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE SOUND ABATEMENT FOAM. PATIENT ALLEGED TO DEVELOP HEADACHES, CHEST PAIN, BREATHING ISSUES. PATIENT WAS TAKING MEDICATIONS. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURER CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REPORTED BY THE PATIENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED IN THE BASE UNIT. SECONDARY FINDINGS SUCH AS CARE ORCHESTRATOR DATA WAS UNAVAILABLE. DUST/DIRT CONTAMINANT FOUND IN THE OUTSIDE OF THE FRONT AND REAR PANELS. DUST/ DIRT CONTAMINANT FOUND AROUND THE AIR INLET. FINE BLACK CONTAMINANT FOUND ON THE INSIDE OF THE UPPER ENCLOSURE, BOTTOM ENCLOSURE, FRONT PANEL, AND REAR PANEL. DUST/DIRT FOUND ON THE BLOWER BOX. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE INVESTIGATOR CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION IN THE DEVICE, BUT CAN CONFIRM THE PRESENCE OF A DUST/DIRT LIKE CONTAMINANT ON THE FRONT PANEL, REAR PANEL, AND BLOWER BOX AND CAN CONFIRM THE PRESENCE OF A FINE BLACK CONTAMINANT ON THE UPPER ENCLOSURE, LOWER ENCLOSURE, FRONT PANEL, AND REAR PANEL. SECTION H6 HEALTH EFFECT - CLINICAL CODE WHICH WAS MISSED IN PREVIOUS REPORT WAS CAPTURED. SECTION D8, D9, H2, H3 AND H6 WERE UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03030 | 2021-08-11 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03028 | 2021-08-11 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE 1 ERROR FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02849 | 2021-08-11 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WAS ONE ERROR FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03072 | 2021-08-11 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP AVAPS30 H/HT/C, DS | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | DSX1130T11C | DSX1130T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | IN PREVIOUS REPORT SECTION B5 SHOULD BE CORRECTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED TO SEEING THE BALCK PARTICLES IN WATER CHAMBER, THROAT SORENESS AND HEADACHE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. AFTER THE SECOND ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, CUSTOMER DECLINED TO RESPOND TO THE GFE RELATED QUESTIONS AND TERMINATED THE CALL. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND CONCLUSION HAS BEEN UPDATED IN THIS REPORT. SECTION H6 CLINICAL CODE HAS BEEN CORRECTED/UPDATED IN HIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03067 | 2021-08-11 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. CORRECTION: THIS WAS PREVIOUSLY REPORTED AS A CPAP DEVICE, BUT IT IS A BIPAP DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03025 | 2021-08-11 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700S11 | DSX700S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED ADDITIONAL INFORMATION ALLEGING PARTICLES AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING/CHAMBER, DIFFICULTY BREATHING/SHORT OF BREATH, WHEEZING, HEADACHES RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION B5 CORRECTED AND H6 CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103217 | 2021-08-11 | 2021-07-07 | I, n, j, u, r, y | Voluntary report | PHILLIPS RESPIRONICS DREAM STATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | STARTED HAVING ISSUES WHILE WEARING THE CPAP IN THE MIDDLE OF THE NIGHT, FELT IRRITATION IN MY RESPIRATORY TRACT. RUNNY NOSE AND COUGH BEGAN. STARTING ON JULY 7 SEVERE HEADACHES EVERY NIGHT WITH AND WITHOUT WEARING THE DEVICE. STILL CURRENTLY HAVING HEADACHES, RUNNY NOSE, AND COUGH AS A RESULT OF THE USE OF THE MACHINE. WAS HOSPITALIZED ON JULY 16 FOR SEVERE HEADACHES AND SINUS PAIN. WENT TO DENTIST TO RULE OUT DENTAL ISSUES ON JULY 20. HAD A CATSCAN ON AUGUST 3. ONLY DISCOVERED ISSUE ABOUT RECALL ON AUGUST 10 WHILE CALLING ABOUT A REPLACEMENT HOSE FOR THE CPAP MACHINE. VISITED VARIOUS DOCTORS WHO RULED OUT OTHER CAUSES FOR THE SEVERE HEADACHES. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-08
2518422-2021-02993 | 2021-08-12 | 2021-07-27 | D, e, a, t, h | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | RIN2110X15B | RIN2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT EXPIRED WHILE THE VENTILATOR WAS IN USE. A REQUEST WAS MADE TO VERIFY IF THE DEVICE SHUTDOWN WHILE IN USE. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED AND CONFIRMS THE DEVICE WAS MANUALLY POWERED OFF USING THE PROPER SEQUENCE OF KEY PRESSES. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND NO MALFUNCTION WAS OBSERVED. THE DEVICE WAS FOUND TO ALARM AND OPERATE AS DESIGNED. |  | Y | N | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03043 | 2021-08-12 | 2021-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE'S LCD SCREEN WAS BLANK. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03047 | 2021-08-12 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE'S LCD SCREEN WAS BLANK. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03040 | 2021-08-12 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, THE DEVICE'S FLOW SENSOR ASSEMBLY WAS DAMAGED AND FAILED TO PASS TESTING. THE DEVICE'S FLOW SENSOR ASSEMBLY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03044 | 2021-08-12 | 2021-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
MW5103221 | 2021-08-12 | 2020-01-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP DEVICE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | AFTER ABOUT A YEAR OF USING MY CPAP DEVICE I STARTED EXPERIENCING UPPER RESPIRATORY SYMPTOMS: CHRONIC POST NASAL DRIP, CHRONIC COUGHING, CHRONIC SINUS INFECTIONS, NASAL CRUSTING DUE TO INFECTION OF NOSE. IT WAS NOT UNTIL I LEARNED FROM THE DEVICE RECALL THAT I WAS QUITE SURE MY CPAP DEVICE MAY HAVE "...". NO LAB TESTS WERE DONE. MY ENT PHYSICIAN DIAGNOSED MY RHINITIS AND MY PCP MY CHRONIC SINUSITIS. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-08
2518422-2021-03114 | 2021-08-12 | 2021-06-30 | D, e, a, t, h | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED PARTICLES IN A PATIENT'S LUNGS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF HOSPITALIZATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON JAN 23, 2025 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A VENTILATOR'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED PARTICLES IN A PATIENT'S LUNGS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF HOSPITALIZATION AND PATIENT WAS EXPIRED. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED. THERE WERE NO ERRORS FOUND. THE MANUFACTURER'S SERVICE CENTER CONCLUDES THAT THEY COULD CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS A VISIBLE FOAM DEGRADATION. DEVICE WAS SCRAPPED. SECTIONS D8, D9, H2, H3 AND H6 HAS BEEN UPDATED IN THIS REPORT. | Y | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03041 | 2021-08-12 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, THE DEVICE'S FLOW SENSOR ASSEMBLY MALFUNCTIONED. THE DEVICE'S FLOW SENSOR ASSEMBLY NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03049 | 2021-08-12 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03102 | 2021-08-12 | 2021-08-04 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP, INTL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | INX700S15 | INX700S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT WAS DIAGNOSED WITH PULMONARY FIBROSIS. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03048 | 2021-08-12 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03042 | 2021-08-12 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, THE DEVICE'S FLOW SENSOR ASSEMBLY MALFUNCTIONED. THE DEVICE'S FLOW SENSOR ASSEMBLY NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03115 | 2021-08-12 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LITTLE FLAKES OF BLACK WHEN NOSE BLOWN, COUGHING, PAIN IN CHEST RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S INVESTIGATION LABORATORY FOR FURTHER EVALUATION. THE INTERNAL AND EXTERNAL ASPECTS OF THE DEVICE WERE INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT COULD NOT CONFIRM THE CUSTOMER'S COMPLAINT AND THERE WAS NO VISIBLE FOAM DEGRADATION AND CONTAMINATION. FAULTS OR ERRORS CONTAINED WITHIN THE LOGS SUPPORT SENSOR BOARD CONTAMINATION ALL OTHER DEVICE FUNCTIONALITY APPEARS AS NO PROBLEM FOUND. SENSOR BOARD FAILURE BECAUSE OF SENSOR CONTAMINATION MT-1 AND MT-4. SECTION B5 WAS INCORRECT IN INITIAL REPORT, IT SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING LITTLE FLAKES OF BLACK WHEN NOSE BLOWN, COUGHING, PAIN IN CHEST RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. SECTION A1, A3, E1 (E-MAIL ADDRESS) CORRECTED. SECTION H6 CORRECTED AND UPDATED IN THIS REPORT. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03058 | 2021-08-12 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED NODULES AND PARTICLES IN THE LUNGS. THE EXACT TREATMENT THE USER RECEIVED IS UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 13, 2024, AND SECTION H11 SHOULD BE REPORTED AS THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED NODULES AND PARTICLES IN THE LUNGS. THE PATIENT DID NOT REPORT TO RECEIVE MEDICAL INTERVENTION. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM (ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR). SECTION B2 WAS MADE BLANK (PREVIOUSLY IT WAS OTHER SERIOUS). SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECT - CLINICAL CODE AND HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03112 | 2021-08-12 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1032804 | 1032804 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02945 | 2021-08-12 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP AUTOSV | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | DSX900H11C | DSX900H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES AND RUNNING NOSE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION D8 AND H6 UPDATED AND CORRECTED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03084 | 2021-08-12 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03082 | 2021-08-12 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE COUGHING IN THE WINTER AND ITCHY . THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE EVENT IS CURRENTLY UNKNOWN. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HEALTH IMPACT CODE WAS MISSED TO CAPTURE,NOW IT IS CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03096 | 2021-08-12 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03095 | 2021-08-12 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED TO SEE SMALL BLACK PIECES IN THE WATER CHAMBER, MOTOR NOISE. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND AND UNIT SCRAPPED. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING TO SEE SMALL BLACK PIECES IN THE WATER CHAMBER, MOTOR NOISE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION H6 WERE UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02953 | 2021-08-12 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO EVIDENCE OF CONTAMINATION. THE MANUFACTURER COMPLETED AN INTERNAL VISUAL INSPECTION. THE MANUFACTURER FOUND EVIDENCE OF CONTAMINATION IN THE BOTTOM PANEL. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS LOGGED. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. SECTION D9, G3, H3 AND H6 HAS BEEN UPDATED / CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03090 | 2021-08-12 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. ADDITIONAL INFORMATION WAS RECEIVED AND ADDED TO THE REPORT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER ON 03/13/2023 FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED ON 05/10/2022. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03091 | 2021-08-12 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGING AN ISSUE.RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED SMALL PIECES IN THE WATER CHAMBER. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. ADDITIONAL INFORMATION WAS RECEIVED AND ADDED TO THE REPORT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER ON (MENTION D9 DATE (03/13/2023) FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED ON (MENTION G3 DATE (05/10/2023). THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE 0 ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED. SECTION D8, D9, H2, H3 AND H6 WERE UPDATED | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03103 | 2021-08-12 | 2021-08-04 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION GO AUTO W/BT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSG500S11 | DSG500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT WAS DIAGNOSED WITH PULMONARY FIBROSIS. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT WAS DIAGNOSED WITH PULMONARY FIBROSIS. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02944 | 2021-08-12 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02431 | 2021-08-12 | 2021-07-06 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/HT/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400T11C | DSX400T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED LESIONS IN BOTH KIDNEYS. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02430 | 2021-08-12 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED CHEST AND BACK PAIN. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 14, 2024. SECTION H11 SHOULD BE REPORTED AS THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED ALLEGATION OF AN ISSUE RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CHEST AND BACK PAIN. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM (ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR). SECTION B2 WAS CHANGED TO BLANK (PREVIOUSLY IT WAS HOSPITALIZATION). SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03055 | 2021-08-12 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED EXACERBATION OF ASTHMA AND COPD, SEVERE COUGH, DIZZINESS, AND HEADACHE. THE PATIENT RECEIVED ANTIBIOTICS AND REPORTS HOSPITALIZATION AS A RESULT OF THE SYMPTOMS EXPERIENCED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED EXACERBATION OF ASTHMA AND COPD, SEVERE COUGH, DIZZINESS, AND HEADACHE. THE PATIENT RECEIVED ANTIBIOTICS AND REPORTS HOSPITALIZATION AS A RESULT OF THE SYMPTOMS EXPERIENCED. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF AN UNIDENTIFIED CONTAMINATION AT THE AIR INTAKE UNDER THE FILTER AREA OF THE BLOWER BOX, KERATIN ACCUMULATION (BLACK CONTAMINATION) BY THE AIR OUTLET, SLIGHT WHITE DUST ON THE BLOWER, RUBBER RING OF BLOWER BOX, AND ON THE TOP OF THE BLOWER BOX. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND NO ERRORS LOGGED. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF MULTIPLE CONTAMINANTS THAT ARE NOT CONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02429 | 2021-08-12 | 2021-06-22 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, AHP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AHX500H11 | AHX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED PNEUMONIA. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING COVID, PNEUMONIA, SHORTNESS OF BREATH, LIGHTHEADEDNESS AND DIZZINESS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02893 | 2021-08-12 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02895 | 2021-08-12 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUMID/HT | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX700T11 | DSX700T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. SECTION H6 MEDICLA DEVICE PROBLE CODE WAS CODED WRONGLY AND INVESTIGATION FINDINGS WAS MISSED TO CAPTURE IN FOLLOW UP REPORT (2518422-2021-02895-2), NOW ITS CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02894 | 2021-08-12 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02892 | 2021-08-12 | 2021-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/WIFI | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11W | DSX500T11W |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT ALLEGED BLACK PARTICLES IN BETWEEN MACHINE AND HUMIDIFIER, COUGHING MORE AND HAD HEADACHES. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL AND INTERNAL INVESTIGATION SHOWED THAT AN UNKNOWN CONTAMINANT ON THE EXTERIOR OF THE TOP ENCLOSURE, INSIDE OF THE ISO PORT AROUND THE OUTSIDE OF THE ISO PORT, AND ON THE OUTSIDE OF THE REAR PANEL. A DISCOLORATION OF THE BOTTOM ENCLOSURE WAS OBSERVED WHERE THE MODEM WOULD BE PLACED. AN INTERNAL INSPECTION OF THE DEVICE WAS PERFORMED, AND PIL OBSERVED THE TOP ENCLOSURE HAD A YELLOW DISCOLORATION TO IT. THE O-RING ON THE REAR PANEL HAD AN UNKNOWN CONTAMINANT ON IT. THE BOTTOM ENCLOSURE WAS OBSERVED TO HAVE AN UNKNOWN CONTAMINANT, HAIR-LIKE PARTICLES, AND A DISCOLORATION ON IT. THE BLOWER BOX WAS OBSERVED TO HAVE AN UNKNOWN DUST CONTAMINANT ON IT, AS WELL AS ON THE BLOWER AND BLOWER SEAL. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. THERE IS NO VISIBLE DAMAGE OR FUNCTIONALITY FAILURES OF THE DEVICE, WHICH SUGGEST THAT AN UNKNOWN CONTAMINANT WAS OBSERVED ON THE BLOWER BOX. SECTION H6 WERE UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03056 | 2021-08-12 | 2021-06-29 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED TWO INCIDENCES OF SQUAMOUS CELL CANCER ON THE FACE, AS WELL AS HEADACHES. THE EXACT TREATMENT THE USER RECEIVED IS UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED TWO INCIDENCES OF SQUAMOUS CELL CANCER ON THE FACE, AS WELL AS HEADACHES AND DEVICE HAS A STRANGE ODOR. THE MEDICAL INTERVENTION RECEIVED BY THE PATIENT WAS UNKNOWN. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE DEVICE WAS EVALUATED AND THERE WAS NO SIGNS OF CONTAMINATION ON THE EXTERNAL PART OF THE DEVICE. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND FOUND THAT THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION IN THE BLOWER BOX, AND THE ERROR LOG SHOWED THAT THERE WERE TWO INSTANCES OF ERRORS, THERE WAS AN E-53 CONTINUE ERROR AND AN E-176 CONTINUE ERROR. THE DEVICE WAS LIKELY RESET PRIOR TO PIL RECEIPT DUE TO MACHINE HOURS BEING 7243.6 AND THE THERAPY BLOWER HOURS BEING 0. THE DEVICE RISK TAGS (ER 2219697 V20) ASSOCIATED WITH THIS FAILURE MODE INCLUDE: DEGRAD04, IRRIT01, IRRIT02, IRRIT05. THE RESULTS OF THIS INVESTIGATION DO NOT IMPACT THE CALCULATED RISKS. THE MANUFACTURER CONCLUDES NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION AND BREAKDOWN IN THE DEVICE. SECTION H6 WERE UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03057 | 2021-08-12 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED EXACERBATION OF ASTHMA AND HEADACHES. THE EXACT TREATMENT THE USER RECEIVED IS UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02980 | 2021-08-12 | 2021-08-03 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, JP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | RJPX500S16 | RJPX500S16 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED SORE THROAT, ASTHMA, AND DETERIORATION IN THE SENSE OF SMELL. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A DOCTOR'S ASSESSMENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED SORE THROAT, ASTHMA, AND DETERIORATION IN THE SENSE OF SMELL. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A DOCTOR'S ASSESSMENT. THE REPORTED EVENT AND ITS SEVERITY WERE REVIEWED BY THE MANUFACTURER'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURER CONCLUDES THAT NO FURTHER ACTION IS NECESSARY. THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03113 | 2021-08-12 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING NASAL/THROAT IRRITATION OR SORENESS, DIZZINESS, NAUSEA RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. SECTION A1, E1(TELEPHONE NUMBER), H6(CLINICAL CODE), H9 CORRECTED IN THIS REPORT. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03071 | 2021-08-12 | 2021-06-20 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 14, 2024, AND SECTION H11 SHOULD BE REPORTED AS THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO EXPERIENCE CANCER. ADDITIONAL INFORMATION WAS RECEIVED ABOUT ALLEGATION OF NASAL/THROAT IRRITATION AND SORENESS. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM (PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR). SECTION B2 WAS CHANGED TO OTHER SERIOUS (PREVIOUSLY IT WAS BLANK). SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECT - CLINICAL CODE WAS UPDATED AND HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03070 | 2021-08-12 | 2021-06-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE CANCER. THE REPORTED EVENT AND ITS REPORTED SEVERITY WAS REVIEWED BY THE PMS'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03054 | 2021-08-12 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS NOT DISPLAYING INFORMATION. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS NOT DISPLAYING INFORMATION. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS RETURNED TO THE MANUFACTURER SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED AND THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION OF NO INFORMATION SHOWING ON DISPLAY OR DAMAGED DISPLAY AND UNIT IS SCRAPPED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03111 | 2021-08-12 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03106 | 2021-08-12 | 2021-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE WITH A VENTILATOR'S VOLUME. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03119 | 2021-08-12 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION OCCURRED RELATED TO THE OXYGEN BLENDING MODULE BOARD. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION OCCURRED RELATED TO THE OXYGEN BLENDING MODULE BOARD. THERE WAS NO HARM OR INJURY REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03120 | 2021-08-12 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2200X11B | DS2200X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS NOT DELIVERING CORRECT VOLUME. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR WAS NOT DELIVERING CORRECT VOLUME. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE WAS EVALUATED BY A FIELD SERVICE ENGINEER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BLOWER MOTOR, MACHINE FLOW SENSOR, AND SYSTEM BOARD NEED REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03118 | 2021-08-12 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A MALFUNCTION OF A VENTILATOR'S KEYPAD. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A MALFUNCTION OF A VENTILATOR'S KEYPAD. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02946 | 2021-08-12 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PHILIPS BIPAP DEVICE | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS INC. | DSX1130H11C | DSX1130H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03165 | 2021-08-12 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03073 | 2021-08-12 | 2021-06-29 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION BIPAP PRO, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX600S11 | DSX600S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP OVARIAN CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | CORRECTIONS TO PREVIOUS MDR ARE MADE IN THIS REPORT AS FOLLOWS. SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY PRODUCT PROBLEM WAS CHECKED IN PREVIOUS MDR) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY IT WAS BLANK) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03077 | 2021-08-12 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS INITIALLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN THE INITIAL REPORT SECTION B5 WAS NOT UPDATED, THE CORRECT B5 SHOULD BE - THE PATIENT HAS ALLEGED BLACK PARTICLES IN THE DEVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND TINY RESIDUE SPOTS ON SOUND ABATEMENT FOAM, DUSTY SMOKE (NON-TOBACCO) ODOR, UNIDENTIFIED WHITE POWDER-LIKE CONTAMINATION AND SMALL PIECES OF LIGHT-COLORED HAIR AT THE AIR INTAKE UNDER THE FILTER AREA OF THE BLOWER BOX .THE MANUFACTURER FOUND EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. A HUMIDIFIER WAS ALSO RETURNED TO THE MANUFACTURER ALONG WITH DEVICE. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND WHITE DUST, AND A LIGHT-COLORED SHORT PIECE OF POTENTIAL HAIR ON THE HEATER PLATE. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND 7 ERRORS LOGGED. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH TINY RESIDUE SPOTS, DUSTY SMOKE (NON-TOBACCO) ODOR, WHITE DUST FROM HUMIDIFIER AND SMALL PIECES OF LIGHT-COLORED HAIR AT THE AIR INTAKE UNDER THE FILTER AREA OF THE BLOWER BOX, UNKNOWN WHITE POWDER, A LIGHT-COLORED SHORT PIECE OF POTENTIAL HAIR ON THE HEATER PLATE OF THE HUMIDIFIER WERE INCONSISTENT WITH THE SOUND ABATEMENT FOAM.THE MANUFACTURER CONFIRMED THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION D9,G3,H6 HAS BEEN UPDATED/CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03021 | 2021-08-12 | 2021-07-14 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 06, 2024, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTION B1 WAS CORRECTED FOR BOTH ADVERSE EVENT AND PRODUCT PROBLEM. ( PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION. TO SERIOUS INJURY. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03078 | 2021-08-12 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING BREATHING PARTICLES, THE DEVICE DID NOT PUT OUT ENOUGH AIR WHEN IT WAS WORKING AND MADE A LOT OF NOISE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. AFTER THE FINAL ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, CUSTOMER DECLINED TO RESPOND TO THE GFE RELATED QUESTIONS AND TERMINATED THE CALL. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03083 | 2021-08-12 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03088 | 2021-08-12 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGING AN ISSUE.RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED SMALL PIECES IN THE WATER CHAMBER AND MOTOR NOISE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. ADDITIONAL INFORMATION WAS RECEIVED AND ADDED TO THE REPORT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER ON (MENTION D9 DATE (03/13/2023) FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED ON (MENTION G3 DATE (05/10/2023). THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE 0 ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03089 | 2021-08-12 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF SMALL BLACK PIECES IN THE WATER CHAMBER AND THE DEVICE HAD A MOTOR NOISE. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE COMPLAINT IS NOT CONFIRMED AND THERE IS NO VISIBLE FOAM PARTICLES. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND UNIT IS SCRAPPED. THE MANUFACTURER PREVIOUSLY NOT CAPTURED, "THE VISUALIZATION OF SMALL BLACK PIECES IN THE WATER CHAMBER AND THE DEVICE HAD A MOTOR NOISE." IN SECTION B5 AND IT HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03087 | 2021-08-12 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING SMALL BLACK PEICES IN THE WATER CHAMBER RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDING TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED DUE TO AGE. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03086 | 2021-08-12 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED TO SEE SMALL BLACK PIECES IN THE WATER CHAMBER AND MOTOR NOISE. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING SMALL BLACK PIECES IN THE WATER CHAMBER AND MOTOR NOISE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND NO VISIBLE FOAM DEGRADATION AND UNIT WAS SCRAPPED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02942 | 2021-08-12 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN INITIAL REPORT SECTION B5 WAS INCOMPLETELY MENTIONED AND THE UPDATED SECTION B5 SHOULD BE- THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING OF SEEING BLACK PARTICLES AND HAVING COUGH. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL AND INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND EVIDENCE OF UNIDENTIFIED WHITE AND BLACK POWDER-LIKE CONTAMINATION AND SMALL PIECES OF WHITE AND BLACK COLORED HAIR AT THE AIR INTAKE UNDER THE FILTER AREA OF THE BLOWER BOX. THE MANUFACTURER FOUND EVIDENCE OF UNKNOWN BLACK CONTAMINATION INSIDE OF THE BLOWER MOTOR CASING WITH SOME WHITE SPECS OF DUST. THE MANUFACTURER FOUND EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. THE HUMIDIFIER WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR EVALUATION. THE MANUFACTURER FOUND BROWN AND BLACK CONTAMINATION AND A BLACK PIECE OF POTENTIAL HAIR WAS FOUND ON THE BOTTOM OF THE HUMIDIFIER ENCLOSURE. THE MANUFACTURER ALSO FOUND POTENTIAL WATER INGRESS AND MINERAL DEPOSITS ON THE AIR INLET TUBE AND HEATER PLATE SHIELD. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND THE ERROR LOG SHOWED 1 ERROR LOGGED. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH UNKNOWN WHITE AND BLACK POWDER-LIKE CONTAMINATION AND SOME WHITE SPECS OF DUST, INCONSISTENT WITH THE SOUND ABATEMENT FOAM. THE MANUFACTURER CONFIRMED THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION D9, D10, G3, H3 AND H6 HAS BEEN UPDATED / CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02943 | 2021-08-12 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING STRANGE SMELL, BLACK SPECS, BREATHING PROBLEMS, HEADACHES AND LIGHT HEADED RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS UPDATED REPORT WILL BE FILED. CORRECTED INFORMATION PROVIDED IN B5 AND H6, THE INFORMATION (THE MANUFACTURER RECEIVED INFORMATION ALLEGING STRANGE SMELL, BLACK SPECS, BREATHING PROBLEMS, HEADACHES AND LIGHT HEADED RELATED TO A CPAP DEVICE) WAS NOT INCLUDED IN INITIAL REPORT. CORRECTED INFORMATION PROVIDED IN D8, (WAS THIS DEVICE SERVICED BY A THIRD PARTY) WAS NOT UPDATED IN EARLIER REPORT. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03075 | 2021-08-12 | 2021-06-29 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A PULMONARY EMBOLYSM. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM AND PATIENT STATED THAT DIFFICULTY IN BREATHING. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. THE PATIENT HAS SWAPPED OUT THE DEVICE AT THE DISTRIBUTOR AND AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03081 | 2021-08-12 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03196 | 2021-08-13 | 2021-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | RJP2100X16B | RJP2100X16B |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, A VENTILATOR WAS ALARMING FOR A "VENTILATOR INOPERATIVE" ALARM CONDITION. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S SYSTEM BOARD AND BLOWER ASSEMBLY WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03199 | 2021-08-13 | 2021-08-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03197 | 2021-08-13 | 2021-08-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, AN OPEN CONDITION WAS FOUND WITHIN THE DEVICE'S POWER CORD. THE DEVICE'S POWER CORD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03198 | 2021-08-13 | 2021-08-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03172 | 2021-08-13 | 2021-06-28 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGED TO DEVELOP LUNG CANCER,LIGHT HEADEDNESS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE REPORTED EVENT OF HAVING LUNG CANCER,LIGHT HEADEDNESS WAS REVIEWED BY THE PMS CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 AND UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103266 | 2021-08-13 | 2021-08-11 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | J263914199295 |  |  | REPORTER STATED SHE OWN TWO CPAP MACHINES FROM PHILIPS. SHE DID NOT KNOW ABOUT THE RECALL UNTIL SHE WAS SURFING THE INTERNET AND CAME ACROSS THE RECALL INFORMATION, SHE IMMEDIATELY CALLED PHILIPS AND IT WAS CONFIRMED. CALLER SAID SINCE SHE'S BEEN ON THE MACHINE, SHE HAS DEVELOPED ASTHMA AND WAS DIAGNOSED WITH REACTIVE AIRWAY DISEASE (RAD). SHE STATED THAT SHE HAS NEVER SMOKED OR DRANK A DAY IN HER LIFE. CALLER SAID SHE ALSO HAS NAUSEA, HEADACHE, SHORTNESS OF BREADTH, VERTIGO, DIZZINESS, CHEST PAIN, MUSCLE ACHE AND BAD NERVE PAIN EVER SINCE SHE BEEN ON THE MACHINE. SHE SAID SHE NEVER KNEW THAT ALL THE SYMPTOMS SHE'S BEEN EXPERIENCING IS COMING FROM USING THE CPAP MACHINE. |  | N | Y | I |  | 2021-08
2518422-2021-02805 | 2021-08-13 | 2021-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING COUGHING MORE OFTEN AND SOME WHITE PARTICLES IN NASAL AREA AND MASK RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING COUGHING MORE OFTEN AND SOME WHITE PARTICLES IN NASAL AREA AND MASK RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION H6 WAS UPDATED AND CORRECTED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03097 | 2021-08-13 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION.UNIT WAS SCRAPPED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03098 | 2021-08-13 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02808 | 2021-08-13 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF BLACK SPECS IN WATER CHAMBER AND TUBING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02807 | 2021-08-13 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED PARTICLES COMING OUT IN WATER CHAMBER. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE 1 ERROR FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. SECTION H6 WAS UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02947 | 2021-08-13 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTOSV | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | DS960XTS | DS960XTS |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED PARTICLES IN TUBING/CHAMBER AND BURNING/SMOKE/ELECTRICAL ODOR. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. ADDITIONAL INFORMATION WAS RECEIVED AND ADDED TO THE REPORT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER ON DATE 03/28/2023 FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED ON DATE 06/29/2023. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. UPON EXTERNAL/INTERNAL EXAMINATION, PIL DID OBSERVE DUST/DIRT CONTAMINATION AS WELL AS A STRONG MUST/MILDEW ODOR WHILE EXAMINING THE BLOWER, HOWEVER PIL DID NOT OBSERVE ANY EVIDENCE OF BLOWER OR THERMAL FAILURE. PIL WAS NOT ABLE TO REPRODUCE THE ERRORS AND OBSERVED NO PARTICULATE IN THE AIRPATH THAT WOULD SUGGEST SMOKE BEING GENERATED IN THE DEVICE. PIL OBSERVED NO EVIDENCE OF DEGRADED SOUND ABATEMENT FOAM. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER 3 ERRORS WERE FOUND. THE MANUFACTURER CONCLUDES THAT THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION FOUND WITHIN THE DEVICE. SECTION D8, D9, H2, H3 AND H6 WERE UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03215 | 2021-08-13 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO SERIES ASSY, INT | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | IN751S | IN751S |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02823 | 2021-08-13 | 2021-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO SERIES ASSY, INT | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | 751P | 751P |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03094 | 2021-08-13 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP RESPIRATORY INFLAMMATION. THE PATIENT WAS PRESCRIBED AN INHALER IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING TO DEVELOP RESPIRATORY INFLAMMATION RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM.THE PATIENT WAS PRESCRIBED AN INHALER IN RESPONSE TO THE REPORTED EVENT.THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03175 | 2021-08-13 | 2021-06-28 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THE PATIENT DID NOT REPORT TO RECEIVE MEDICAL INTERVENTION. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02949 | 2021-08-13 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND NO EVIDENCE OF CONTAMINATION ON THE OUTSIDE OF THE DEVICE. AN INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE HUMIDIFIER WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR EVALUATION. THE MANUFACTURER FOUND EVIDENCE OF CONTAMINATION IN THE HUMIDIFIER, ON THE BOTTOM OF THE HUMIDIFIER, ON BOTH SIDE OF THE HEATING PLATE, ON DRY BOX, ON THE RESERVOIR SEAL AND INSIDE OF THE ISO PORT. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND NO ERRORS LOGGED. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF CONTAMINATION ON THE BOTTOM OF THE HUMIDIFIER, ON BOTH SIDE OF THE HEATING PLATE, ON DRY BOX, ON THE RESERVOIR SEAL AND INSIDE OF THE ISO PORT. THE MANUFACTURER CONFIRMED THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. SECTION D9, D10, G3 AND H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03167 | 2021-08-13 | 2021-06-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT ALSO ALLEGED BLACK STUFF IN MACHINE AND HAVING NASAL/THROAT BURNING AND IRRITATION. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL PART OF THE DEVICE AND FOUND FAULTY BLOWER, EVIDENCE OF BLOWER BOX FOAM AND WATER CONTAMINATION TO MOOG BLOWER. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND FOUND FAULTY BLOWER, EVIDENCE OF BLOWER BOX FOAM AND WATER CONTAMINATION TO MOOG BLOWER. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION, BUT FOUND FAULTY BLOWER, EVIDENCE OF BLOWER BOX FOAM AND WATER CONTAMINATION TO MOOG BLOWER WERE OBSERVED AND ARE CONSISTENT WITH BEING FROM AN EXTERNAL SOURCE. SECTION H6 WERE UPDATED IN THIS REPORT. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03168 | 2021-08-13 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES AND SINUS INFECTION CONGESTION, ALSO STATES THAT CLEANED DEVICE AND LIQUID CAME OUT BROWN WHILE CLEANING THE DEVICE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS UPDATED REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02948 | 2021-08-13 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSG500H11 | DSG500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING RESPIRATORY INFECTION AND PARTICLES IN AIRWAY RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. CORRECTIOM INFORMATION WAS PROVIDED IN D8, G1 AND G4. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03169 | 2021-08-13 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DEVICE IS NOISY AND PARTICLES IN THE AIR WAYS RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE - THE PATIENT ALLEGED DEVICE IS NOISY AND PARTICLES IN THE AIR WAYS RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03166 | 2021-08-13 | 2021-06-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES IN DEVICE AND SHOWN IN LUNGS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02918 | 2021-08-13 | 2021-06-30 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID/HT, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | CAX400T12 | CAX400T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE SECTIONS B1, B2, H1 AND H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO BOTH ADVERSE EVENTS AND PRODUCT PROBLEM. ( IN THE PREVIOUS MDR ONLY PRODUCT PROBLEM WAS CHECKED.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02826 | 2021-08-13 | 2021-05-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11W | DSX500T11W |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DETAILS OF THE DEVICE SERIAL NUMBER ARE UNKNOWN. IF FURTHER INFORMATION REGARDING THE DEVICE BECOME AVAILABLE, IT WILL BE SUBMITTED IN THE FOLLOW-UP REPORT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW- UP REPORT WILL BE FILED. UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02923 | 2021-08-13 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02922 | 2021-08-13 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02921 | 2021-08-13 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX700T11 | DSX700T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02920 | 2021-08-13 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03125 | 2021-08-13 | 2021-07-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES IN TUBING CHAMBER AND ALSO ALLEGED HEADACHE AND COUGH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02919 | 2021-08-13 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03174 | 2021-08-13 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11 | DSX700T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03173 | 2021-08-13 |  | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03177 | 2021-08-13 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING BLACK LIQUID ON BOTTOM RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03178 | 2021-08-13 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02832 | 2021-08-13 | 2021-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02831 | 2021-08-13 | 2021-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03176 | 2021-08-13 |  | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGED PARTICLES OR BLACK DEBRIS AROUND THE UNDER RUBBER LID OF THE REVISOR RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03170 | 2021-08-13 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03131 | 2021-08-13 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING BLACK FILTERS, RASPY VOICE, NASAL/THROAT IRRITATION OR SORENESS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK FILTERS, RASPY VOICE, NASAL/THROAT IRRITATION OR SORENESS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. AFTER THE SECOND ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, THE CUSTOMER DECLINED TO RESPOND TO THE GFE RELATED QUESTIONS AND TERMINATED THE CALL. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION G1-MANUFACTURER CONTACT OFFICE DETAILS WERE CORRECTED AND UPDATED. SECTION H6 HAS BEEN CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03133 | 2021-08-13 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON MAY 06, 2025, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PATIENT STATES WHILE YOU CLEANING YOU NOTICE BLACK FLECKS FLOATING IN WATER AND HAS SHORTNESS BREATH, FEVER, PRESSURE ON CHEST RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM.. THE PATIENT DID NOT REPORT TO RECEIVE MEDICAL INTERVENTION. SECTION B1 WAS CORRECTED FOR BOTH ADVERSE EVENT AND PRODUCT PROBLEM. ( PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.). SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.). SECTION H1 WAS CHANGED FROM MALFUNCTION. TO SERIOUS INJURY. SECTION H6 WAS CORRECTED AND UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03085 | 2021-08-13 | 2021-07-14 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION EXPERT W/HUM, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | CAX501H12 | CAX501H12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM ALLEGEDLY BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. AN EXTERNAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND THE GROSS PARTICLE FILTER CLOGGED WITH DUST AND DIRT. THE ULTRA FINE FILTER WAS ALSO HEAVILY CONTAMINATED WITH DUST. AN INTERNAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF MINOR DUST CONTAMINATION THROUGHOUT THE DEVICE'S BLOWER BOX AND BLOWER. THE MANUFACTURER ALSO FOUND EVIDENCE OF WATER INGRESS TO THE BLOWER BOX AND BLOWER. THE MANUFACTURER TESTED THE DEVICE FOR OVERHEATING USING THE RETURNED FILTERS. THE MANUFACTURER CONFIRMED THE DEVICE'S TEMPERATURES WERE BELOW THE MAXIMUM TEMPERATURE STANDARDS. THE DEVICE WAS TESTED FOR PARTICULATE BY PLACING A BACTERIAL FILTER AT THE PATIENT PORT DURING TESTING. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND CONFIRMED THE DEVICE PROVIDED AIRFLOW. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED AND FOUND ONE ERROR CODE FOR A REBOOT ERROR WAS LOGGED. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103233 | 2021-08-13 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAM STATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I NOTICED A FILM ON TOP OF THE WATER IN THE MORNING. WHEN I TOUCHED THE FILM, IT WAS BLACK ON MY FINGERTIP. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103234 | 2021-08-13 | 2021-02-04 | I, n, j, u, r, y | Voluntary report | DREAM STATION C-PAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | J296517804D6B |  |  | RECURRENT SINUSITIS/BRONCHITIS/PNEUMONIA. MULTIPLE COVID TESTS HAVE BEEN PERFORMED W/ EACH ONSET OF ILLNESS. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103235 | 2021-08-13 | 2020-08-20 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS DREAM STATION CPAP DEVICE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | HELLO, MY NAME IS (B)(6). MY BIRTH DATE IS (B)(6). I HAVE BEEN USING MY RESPIRONICS CPAP MACHINE SINCE 2019. I HAVE REGISTERED MY MACHINE ON THE RECALL SITE. MY MACHINE IS INCLUDED IN THE RECALL. I HAVE BEEN AWAITING INSTRUCTIONS AS TO REMEDIATION OR REPLACEMENT. I HAVE DISCUSSED WITH MY DR AND HAVE WEIGHED MY OPTIONS BETWEEN DISCONTINUING THE USE OF MY MACHINE OR TAKING A CHANCE ON THE CONSEQUENCES OF USING THE MACHINE. I HAVE NO CHOICE TO CONTINUE USING THE RESPIRONICS MACHINE BECAUSE OF THE DANGEROUS CONSEQUENCES IN DISCONTINUING USE. MY HUSBAND IS THOROUGHLY CLEANING ALL MACHINE PARTS DAILY INCLUDING THE WET FILTER TO KEEP AS MANY PARTICLES OUT AS POSSIBLE. I AM WRITING THIS LETTER IN RESPONSE TO A LETTER FROM PHILIPS CONCERNING THE MACHINE. I AM INCLUDING THE MULTIPLE ISSUES I HAVE HAD AND HAVE SEEN SPECIALISTS ABOUT WITH NO OBVIOUS REASONS FOR MY SYMPTOMS UNTIL NOW. THE ISSUES INCLUDE, BUT ARE NOT LIMITED TO THE FOLLOWING: DAILY HEADACHES AND DIZZINESS, DAILY EYE IRRITATION, DAILY NASAL CONGESTION WITH SEVERE THROAT PHLEGM, MULTIPLE SINUS INFECTIONS REQUIRING 4 REGIMENS OF ANTIBIOTICS; REOCCURRING COUGH, ELEVATED LIVER FUNCTION TESTS. I AM IN A SEVERE DILEMMA NOW. WITHOUT THE USE OF MY MACHINE, I SUFFER FROM SEVERE SLEEP DEPRIVATION. THE FIRST SINUS INFECTION I HAD WAS DURING COVID 19. THE INFECTION WAS SO BAD THAT I COULD NOT USE MY MACHINE EFFECTIVELY FOR 3 DAYS. AS A RESULT, I FELL DOWN MY GRANITE STEPS WHILE ON MY WAY TO GET A COVID TEST SO I COULD SEE MY DOCTOR. I SUFFERED A COMPLEX FRACTURE OF BOTH BONES IN MY WRIST AND FRACTURED MY PATELLA. AS YOU CAN IMAGINE, NOT USING THE MACHINE IS NOT AN OPTION. MY QUALITY OF LIFE HAS GOT TO IMPROVE. I AM IN DESPERATE NEED OF A NEW MACHINE. PLEASE LET ME KNOW IF YOU REQUIRE MORE INFORMATION. ANY ASSISTANCE IS APPRECIATED. SINCERELY,(B)(6). FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103238 | 2021-08-13 | 2018-08-14 | I, n, j, u, r, y | Voluntary report | PHILIPS CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | ON (B)(6) 2018 I WOKE UP WITH PRESSURE IN MY CHEST FOR 10 MINUTES, IT DIDN'T GO AWAY. MY HUSBAND DROVE ME TO THE HOSPITAL 4 MINUTES FROM MY HOME. THEY TOLD ME I WAS HAVING A HEART ATTACK AND THEY WERE CALLING IN LIFE FLIGHT BECAUSE I NEEDED TO BE TRANSFERRED TO (B)(6) HOSPITAL. NEXT I KNOW THE LIFE FLIGHT CREW PUT ME IN AN AMBULANCE AND SAID THEY WERE DRIVING ME THERE DUE TO THE WEATHER CONDITIONS. I DON'T REMEMBER ARRIVING THERE I ONLY REMEMBER WHEN I WOKE UP THERE WERE A TEAM OF DOCTORS STANDING IN MY ROOM AND THE ONE ASKED ME IF I REMEMBER ANYTHING THAT HAPPENED. I ONLY REMEMBERED BEING IN AN AMBULANCE IS WHAT I SAID. ONE DOCTOR SAID I WILL TELL YOU, YOU ARE ONE LUCKY GIRL TO BE ALIVE AND HERE WITH US TODAY. YOU SUFFERED A HEART ATTACK ALSO KNOWN BY V-FIB 'VENTRICLE FIBULATION'. THE ARTERY BLEW WHILE WE WERE PUTTING IN THE STENT IN YOUR HEART AND YOU CODED ACTUALLY YOU CODED TWICE ON US, ONE DURING THE PROCEDURE AND AFTER YOU WERE BROUGHT UP TO THE ROOM. NOT SURE BUT I WILL RELEASE RECORDS IF NEEDED. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103251 | 2021-08-13 |  | I, n, j, u, r, y | Voluntary report | PHILIPS BIPAP SYSTEM ONE | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | PHILIPS / RESPIRONICS, INC. |  |  |  | ADDITIONAL INFORMATION RECEIVED ON 10/14/2021 FOR MW5103251 FU1; I WOULD LIKE TO UPDATE MY MEDICAL DEVICE REPORT TO THE FOOD AND DRUG ADMINISTRATION'S MEDWATCH PROGRAM. AROUND (B)(6) 2021 MY DOCTOR FOUND, WITH THE HELP OF AN MRI, A LARGE MALIGNANT MASS IN MY COLON AND UNDISTINGUISHABLE SPOTS ON MY LIVER. I THEN WAS GIVEN AN ULTRA SOUND TO HELP IDENTIFY IF THE SPOTS WERE CANCEROUS OR NOT AND THEY FOUND THE SPOTS WERE CYSTS. I AM NOW SCHEDULED FOR A HEMI COLECTOMY THE (B)(6) 2021. SO FAR I HAVE BEEN VERY LUCKY NOT HAVING A CANCEROUS LIVER AND A RESECTION OF THE COLON LEAVING ME WITH NOT HAVING TO HAVE A COLOSTOMY BAG. I AM WONDERING IF THE DEFECTIVE DEVICE HAD ANY THING TO DO WITH MY DEVICE. I HAD AND USED THE SAME BIPAP MACHINE FOR APPROXIMATELY 12 YEARS. I APPRECIATE YOUR TIME AND CONSIDERATION OF THE INFORMATION I HAVE GIVEN YOU. |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103255 | 2021-08-13 | 2018-01-01 | I, n, j, u, r, y | Voluntary report | DREAM MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS, INC. |  |  |  | AFTER BEGINNING TO USE A CPAP, I WOULD AWAKEN SEVERAL TIMES DURING THE NIGHT COUGHING FOR SOMETIME. AFTER STOPPING USING THE CPAP, DUE TO IT'S RECALL, I FIND I STILL HAVE A COUGH, SOMETIMES A COUPLE TIMES A DAY. I HAVE SEEN AN ENT DOC, SOME IRRITATION. EXISTS. THE CPAP WAS RECALLED DUE TO DEGRADING PLASTIC INSULATION AND WAS SAID TO BE SENDING BLACK PARTICLES INTO USERS. RESULTS AVAILABLE THROUGH (B)(6). FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103256 | 2021-08-13 |  | I, n, j, u, r, y | Voluntary report | RESPIRONICS DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | MEDICAL ISSUES WITH CPAP USE; I HAVE BEEN USING A PHYSICIAN PRESCRIBED RESPIRONICS DREAMSTATION CPAP MACHINE FOR A COUPLE OF YEARS. I HAVE EXPERIENCED BOUTS OF SEVERE DIZZINESS, EYE AND NOSE IRRITATION, INEXPLICABLE BOUTS OF COUGHING, OCCASIONAL VOMITING WITH NO APPARENT CONNECTION TO FOOD OR BEVERAGE CONSUMPTION, SINUS INFECTIONS, RUNNY NOSE, DIFFICULTY BREATHING AND OCCASIONAL HEADACHES. WITH THE EXCEPTION OF THE RUNNY NOSE AND OCCASIONAL HEADACHES, I DID NOT EXPERIENCE THESE PRIOR TO THE USE OF THE CPAP. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103258 | 2021-08-13 | 2021-07-15 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS, INC. |  |  |  | I HAVE BEEN USING THE PHILIPS DREAMSTATION CPAP MACHINE FOR ABOUT 5 YEARS. I STARTED COUGHING VIOLENTLY IN (B)(6) 2021 AND BECAME AWARE OF THE RECALL AT THE END OF (B)(6) OF THIS DEVICE BECAUSE OF THE DEGRADATION OF THE FOAM USED. I AM STILL COUGHING AND HAVE BEEN SEEN BY THE URGENT CARE, MY PRIMARY MD AND A PULMONOLOGIST. I JUST HAD A CT SCAN OF MY LUNGS AND THE FINDINGS INCLUDED BIBASILAR ATELECTASIS. I HAVE STOPPED USING THE MACHINE AS I FEAR CONTINUED USE MAY CAUSE FURTHER DAMAGE. ON (B)(6) 2021 I WAS FLYING FROM (B)(6) AND HAD A DIFFICULT TIME BREATHING AND WAS COUGHING PROFUSELY. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103259 | 2021-08-13 | 2021-08-12 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | RESPIRONICS SYSTEM ONE CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. | P13173962AF7E |  |  | I HAVE JUST READ ABOUT SYMPTOMS FROM USING A PHILIPS RESPIRONICS CPAP MACHINE AND REALIZE THAT I AM EXPERIENCING SOME OF THEM. I DIDN'T REALIZE THERE WAS A PROBLEM UNTIL JUST READING ABOUT IT. I AM EXPERIENCING THROAT IRRITATION, NASAL IRRITATION AND CONTINUOUS DRIPPING, AN ANNOYING TICKLE IN THE THROAT CAUSING COUGHING, MILD HEADACHES AND EYE ITCHINESS. NO TESTS UNDERTAKEN. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103263 | 2021-08-13 | 2018-05-17 | I, n, j, u, r, y | Voluntary report | CPAP DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. |  |  |  | I HAD STARTED USING A PHILIPS CPAP MACHINE IN FEBRUARY AND NOT LONG AFTER STARTED HAVING SEVERE HEADACHES, NASAL IRRITATION, RAW AREAS, BURNING, RESPIRATORY IRRITATION, LARYNGITIS AND THEN BLOODY BEGAN WITH BLOODY NOSES. PRODUCT WAS RECALLED. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03074 | 2021-08-13 | 2021-07-06 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP PNEUMONIA AND A SPOT ON THEIR LUNG. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. SECTION B5 SHOULD BE PREVIOUSLY REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF BLACK PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT WAS ALLEGING WHILE SHUTING DOWN THE MACHINE IT AFFECTS HER BRAIN AND BREATHING DOWN AND CAUSED PATIENT TO DEVELOP PNEUMONIA AND A SPOT ON THEIR LUNG. IT WAS ALSO ALLEGED THAT THE PATIENT HAS CANCER WHILE USING THE DEVICE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED. THERE WAS ONE ERROR (ERR_PSENS_UNABLE_T O_OBTAIN_BUS) FOUND. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION WAS OBSERVED. IT WAS CONFIRMED THAT THE POSITIVE DEPOSIT OF FOAM PARTICULATE ON FIT TOOL. SECONDARY FINDINGS SHOWS THAT EVIDENCE OF LIQUID INGRESS. THE MANUFACTURER CONCLUDES THAT THEY COULD CONFIRM THAT THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION WAS OBSERVED AND SECONDARY FINDINGS SHOWS THAT EVIDENCE OF LIQUID INGRESS. SECTION H6 WAS CAPTURED INCORRECTLY AND SECTION H10 DID NOT CAPTURE ANY ERRORS IN PREVIOUS FOLLOW UP REPORT, NOW IT HAS BEEN CORRECTED AND UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03122 | 2021-08-13 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. IN INITIAL REPORT SECTION B5 WAS MENTION INCORRECTLY, CORRECT B5 SHOULD BE - THE PATIENT ALLEGED PARTICLES IN AIR PATH, NASAL IRRITATION, SORE THROAT, DIFFICULTY BREATHING. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL AND INTERNAL VISUAL INSPECTIONS OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND DARK DUST OR DIRT CONTAMINATION OBSERVED IN FILTER INTAKE AREA, BLOWER BOX AREA, OUTLET PORT AND ON THE INTERNAL PLASTIC SURFACES, DRIED MINERAL DEPOSITS IN THE OUTLET PORT WITH MINOR WHITE AND DARK FIBER CONTAMINATE, WHITE HAIR OR FIBER CONTAMINATE LAYING ON THE BLOWER MOTOR, MINOR AMOUNT OF KERATIN ON THE BLOWER OUTLET SEAL AND AROUND THE BLOWER BOX OUTLET PORT, DARK SHEEN ON THE BACK OF THE UI PANEL. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND 1 INSTANCE OF ERROR. THE MANUFACTURE CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH DARK DUST OR DIRT CONTAMINATION FILTER INTAKE AREA, BLOWER BOX AREA, OUTLET PORT AND ON THE INTERNAL PLASTIC SURFACES, DRIED MINERAL DEPOSITS IN THE OUTLET PORT WITH MINOR WHITE AND DARK FIBER CONTAMINATE, MINOR AMOUNT OF KERATIN ON THE BLOWER OUTLET SEAL AND AROUND THE BLOWER BOX OUTLET PORT, DARK SHEEN ON THE BACK OF THE UI PANEL AND WHITE HAIR OR FIBER CONTAMINATE LAYING ON THE BLOWER MOTOR INCONSISTENT WITH SOUND ABATEMENT FOAM. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03126 | 2021-08-13 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED DEVICE IS NOISY, NASAL/THROAT IRRITATION OR SORENESS,BREATHING ISSUES,DUST PARTICLES AND/OR FIBERS IN THE TANK,DRY COUGH, HEADACHE . THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. THE DEVICE WAS EVALUATED. THERE WAS MENTION OF EXTERNAL AND INTERNAL DAMAGE EVIDENT, REAR PANEL MISSING. SCRAPES/MARKING ON BOTTOM OF DEVICE. MISSING FOOT. 0 ERRORS, 0 BLOWER HOURS, AND 0 THERAPY HOURS WERE LOGGED, SUGGESTING THE DEVICE DATA WAS RESET PRIOR TO PIL RECEIPT. 1,680.8 MACHINE HOURS WERE LOGGED.CARE ORCHESTRATOR DATA WAS UNAVAILABLE,PLASTIC OF THE BLOWER BOX IS CRACKED AND BROKEN IN SEVERAL PLACES INCLUDING NEAR AIR INLET AND NEARBY SCREW BOSS, FERRITE BEAD CHANNEL, AS WELL AS PART OF THE AIR OUTLET.,LEADS OF P3 ON PCA BENT,VERY FINE DUST/DIRT CONTAMINATION INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM WAS OBSERVED ON AND IN THE BLOWER SEAL APPEARED TO NOT BE FULLY SEATED, LIKELY DUE TO THE PHYSICAL DAMAGE TO THE BLOWER BOX UPON EXTERNAL AND INTERNAL ASPECT OF THE DEVICE. THE MANUFACTURER CONCLUDES THAT PIL IS UNABLE TO ADDRESS THE SYMPTOMS SPECIFIED IN THE COMPLAINT. PIL IS UNABLE TO CONFIRM THE COMPLAINT REGARDING PARTICLES IN THE WATER TANK DUE TO NOT RECEIVING A HUMIDIFIER OR WATER TANK WITH THE DEVICE. PIL CAN CONFIRM SEVERE PHYSICAL DAMAGE TO THE DEVICE, HOWEVER IT IS UNDETERMINED WHEN THE DEVICE SUSTAINED THE DAMAGE. PIL IS ABLE TO CONFIRM THE PRESENCE OF CONTAMINATION IN THE AIRPATH. PIL FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. SECTION H6 WERE UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03129 | 2021-08-13 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX600H11 | DSX600H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING SINUS INFECTION,BLACK PARTICLES IN HUMIDIFIER TANK,DIFFICULTY BREATHING/SHORT OF BREATH,PARTICLES IN TUBING/CHAMBER,HEADACHES,EXCESS MUCUS RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03123 | 2021-08-13 |  | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTRE FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF DEVICE WAS INSPECTED. THE DEVICE POWERED ON, AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. SECTION H6 UPDATED IN THIS FOLLOW UP REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03127 | 2021-08-13 | 2021-07-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX700H11 | DSX700H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. IN PREVIOUS REPORT SECTION B5 MENTION INCORRECTLY, CORRECT B5 SHOULD BE THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PARTICLES IN AIR PATH. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTIONS OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND UNKNOWN WHITE AND BLACK DUST LIKE CONTAMINATION, AND SOME VERY SMALL POTENTIAL HAIRS OR FIBERS AT THE AIR INPUT OF THE BLOWER BOX. SOME OF THE PCA VIAS WERE DISCOLORED TO A COPPER-LIKE OR YELLOWISH COLOR. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER AND FOUND LIGHT DUST-LIKE CONTAMINATION ON TOP OF THE BLOWER MOTOR AND THE BLOWER MOTOR SEAL. ALSO FOUND LIQUID SPOTS ON THE BOTTOM OF THE BLOWER MOTOR, LIGHT BROWN LIQUID SPOTS ON THE BOTTOM OF THE BLOWER BOX, INDICATING POTENTIAL WATER INGRESS. A BLACK CONTAMINATION, CONSISTENT WITH KERATIN FOUND AT THE AIR OUTLET OF THE BLOWER BOX. TINY UNKNOWN SALT-LIKE TAN CONTAMINATION ON THE INSIDE OF THE BLOWER MOTOR SEAL AND INSIDE OF THE BLOWER MOTOR. TINY UNKNOWN BLACK, WHITE AND TAN SPECS OF CONTAMINATION FOUND ON THE BOTTOM THE BLOWER BOX. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND 0 ERRORS LOGGED WHILE WAS POWERING DEVICE ON. THE MANUFACTURER CONCLUDES MULTIPLE CONTAMINATION OF DUST, TAN, DISCOLORATION FOUND WERE EXTERNAL TO THE DEVICE AND WATER INGRESS AND LIQUID SPOTS WERE CONSISTENT TO THE BLOWER MOTOR AND BLOWER BOX. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. IN THIS REPORT, SECTION D9, G3, H3, H6 HAS BEEN UPDATED OR CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03128 | 2021-08-13 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX600H11C | DSX600H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN AIRWAY STANGE ODOR NOISY NASAL/THROAT IRRITATION DIFFICULTY BREATHING NAUSEA LIGHTHEADED DIZZY RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT DID NOT REPORT TO RECEIVE MEDICAL INTERVENTION. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN AIRWAY STANGE ODOR NOISY NASAL/THROAT IRRITATION DIFFICULTY BREATHING NAUSEA LIGHTHEADED DIZZY RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL FINDINGS WAS EXTERNAL INVESTIGATION SHOWED THAT THERE WERE NO SIGNS OF CONTAMINATION ON THE OUTSIDE OF THE DEVICE.AND INTERNAL PART OF THE DEVICE WAS INSPECTED CONTAMINATION IN FILTER PORT, E-83 ERROR ,THE EXTERNAL INVESTIGATION SHOWED THAT THERE WERE NO SIGNS OF CONTAMINATION ON THE OUTSIDE OF THE DEVICE. THE DEVICE WAS HOOKED UP TO A KNOWN GOOD POWER SUPPLY AND POWER CORD, AIRFLOW WAS VERIFIED TO BE OPERATING CORRECTLY. THE CARE ORCHESTRATOR DATA SHOWED THAT THE PATIENT HAD A COMPLIANCE RATING OF 0% AND USED THE HUMIDIFIER ON ADAPTIVE SETTING 4. THE ERROR LOG SHOWED THAT THERE WAS ONE INSTANCE OF AN E-83 (ERR_FSENS_UNABLE_T O_OBTAIN_BUS) REBOOT ERROR. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WAS ONE ERRORS FOUND (E-83 ERROR). THE INTERNAL INVESTIGATION SHOWED THAT THERE WERE SIGNS OF CONTAMINATION IN THE FILTER PORT ON THE BLOWER BOX CONTAMINATION IN FILTER PORT , E-83 ERROR .THERE WAS ALSO AN UNKNOWN CONTAMINATE ON THE OUTPUT PORT. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. SECTION G1, G3 INCORRECTLY CAPTURED IN PREVIOUS REPORT AND H6 WERE UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03171 | 2021-08-13 | 2021-06-29 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500S11 | DSX500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO REQUIRE CARDIAC SURGICAL INTERVENTION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | PATIENT ALLEGATION MISSED IN PREVIOUS REPORT AND B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM AND CAUSED NASAL/THROAT IRRITATION OR SORENESS. PATIENT STATED NOT GETTING A LOT OF AIR, MAKES NOISE, HEADACHES AND LIGHTHEADNESS. PATIENT REQUIRE CARDIAC SURGICAL INTERVENTION. THE REPORTED EVENTS OF NASAL/THROAT IRRITATION OR SORENESS, HEADACHE AND LIGHTHEADEDNESS ARE ASSESSED AS NON-SERIOUS INJURIES AND THE REPORTED HEART SURGERY IS NOT RELATED TO THE DEVICE. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HAS BEEN CORRECTED/UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03124 | 2021-08-13 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03138 | 2021-08-13 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING BLACK SPOTS IN THE WATER CHAMBER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. CORRECTED INFORMATION WAS PROVIDED IN E1. (PREVIOUSLY INITIAL REPORTER TELEPHONE NUMBER WAS NOT INCLUDED) AND G3 DATE INCORRECTLY CAPTURED IN INITIAL REPORT, IT SHOULD BE 30-JUN-2021. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE - THE PATIENT ALLEGED BLACK SPOTS IN THE WATER CHAMBER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03130 | 2021-08-13 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03140 | 2021-08-13 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11C | DSX400H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED DIFFICULTY BREATHING, CHEST PAIN, HEADACHE, DIZZINESS. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED DIFFICULTY BREATHING, CHEST PAIN, HEADACHE, DIZZINESS. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED AND CORRECTED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03143 | 2021-08-13 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS INITIALLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN INITIAL REPORT SECTION B5 WAS MENTIONED INCORRECTLY, CORRECT B5 SHOULD BE - THE PATIENT ALLEGED SINUS ISSUE. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND CONTAMINATION INSIDE OF THE CASE ON THE BOTTOM BACK AND TOP PANELS OF THE CASE. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS. THE DEVICE WAS APPLIED POWER AND THE DEVICE OPERATED PROPERLY. THE MANUFACTURER CONCLUDES CONTAMINATION FOUND WERE CONSISTENT WITH CONTAMINATION INSIDE OF THE CASE ON THE BOTTOM BACK AND TOP PANELS OF THE CASE. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03145 | 2021-08-13 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN PREVIOUSLY SUBMITTED REPORT SECTION B5 WAS INCOMPLETE, CORRECT B5 SECTION IS- THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED VISUALIZATION OF PARTICLES. THE PATIENT ALSO ALLEGED OF HAVING DRY CONSTANT COUGH, TICKLING IN HER THROAT, HEADACHES, DIZZINESS, AND NAUSEA. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL AND INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF STICKY CONGEALED CONTAMINATION BETWEEN BOTTOM ENCLOSURE AND REAR PANEL, SIMILAR CONTAMINATION IN CREVICES AROUND THERAPY BUTTON, DUST CONTAMINATION ON AND IN THE BLOWER, KERATIN-LIKE CONTAMINATION AROUND THE BLOWER OUTPUT SEAL, LIQUID INGRESS TO THE BLOWER BOX, MINERAL DEPOSITS. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND 1 ERROR LOGGED. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF STICKY CONGEALED CONTAMINATION BETWEEN BOTTOM ENCLOSURE AND REAR PANEL, SIMILAR CONTAMINATION IN CREVICES AROUND THERAPY BUTTON, DUST CONTAMINATION ON AND IN THE BLOWER, KERATIN-LIKE CONTAMINATION AROUND THE BLOWER OUTPUT SEAL, LIQUID INGRESS TO THE BLOWER BOX, MINERAL DEPOSITS. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. SECTION D9, G3 AND H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03136 | 2021-08-13 | 2020-03-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03146 | 2021-08-13 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORT THIS DEVICE (B)(6) ON MDR 2518422-2021-03146 . THE MDR 2518422-2021-03146 WAS INCORRECTLY FILED AND IS A DUPLICATE TO MDR 2518422-2022-23601. PLEASE DISREGARD MDR 2518422-2021-03146 AS IT WAS FILED IN ERROR. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03144 | 2021-08-13 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF BLACK PARTICLES FLOATING IN WATER TANK; NASAL/THROAT IRRITATION OR SORENESS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03147 | 2021-08-13 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX600T11C | DSX600T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING OF VISUALIZATION OF PARTICLES RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. AFTER THE FINAL ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, CUSTOMER DECLINED TO RESPOND TO THE GFE RELATED QUESTIONS AND TERMINATED THE CALL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. IN INTIAL REPORT, SECTION G3 WAS WRONGLY CAPTURED AS 08/03/2021, SHOULD BE 07/02/2021, CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03141 | 2021-08-13 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03142 | 2021-08-13 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES ALSO THE PATIENT ALLEGED TO HAVE DIFFICULTY BREATHING/SHORT OF BREATH. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED EXTERNALLY AND THE MANUFACTURER OBSERVED DUST/DIRT CONTAMINATION WITH DEGRADED SOUND ABATEMENT FOAM WAS OBSERVED THROUGHOUT THE DEVICE ENCLOSURE AND AIRPATH,SUGGESTING A SOURCE EXTERNAL TO THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED AND THE MANUFACTURER OBSERVED EVIDENCE OF WATER INGRESS ON THE BLOWER BOX ENCLOSURE. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY CONFIRMED THE CUSTOMER'S ALLEGATION AND THERE IS PRESENCE OF CONTAMINATION IN THE AIRPATH AND THE SOURCE OF CONTAMINATIONS WERE EXTERNAL TO THE DEVICE. SECTION H6 HEALTH IMPACT CODE WAS MISSED TO CAPTURE IN INITIAL MDR AND IT IS CORRECTED AND UPDATED IN THIS REPORT . | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03139 | 2021-08-13 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN THE PREVIOUS REPORT THE B5 WAS NOT UPDATED, THE CORRECT B5 SHOULD BE - THE PATIENT HAS ALLEGED VISUALIZATION OF PARTICLES,NASAL/THROAT IRRITATION, CLOGGED NOSE. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO EVIDENCE OF CONTAMINATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND CONTAMINATION ON THE INLET PORT AND INSIDE THE BLOWER BOX.THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND E-130 AND E-53 ON THE ERROR LOG. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH CONTAMINATION ON THE INLET PORT AND INSIDE THE BLOWER BOX. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION D9,G3,H6 HAS BEEN UPDATED/CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03134 | 2021-08-13 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING BLACK MATERIAL FLOATIG IN WATER RESERVOIR, PARTICLES IN TUBING/CHAMBER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE ONE ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02588 | 2021-08-14 | 2021-06-28 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT WAS DIAGNOSED WITH LUNG CANCER. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02589 | 2021-08-14 | 2017-05-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION BIPAP AUTOSV H, DS | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | DSX900H11 | DSX900H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED SINUS INFECTIONS AND CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD). THE PATIENT DID REQUIRE MEDICAL INTERVENTION IN THE FORM OF PRESCRIBED ANTIBIOTICS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON OCT 12, 2023, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF BLACK PARTICLES RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT ALLEGED TO DEVELOPED SINUS INFECTIONS, CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD), ENLARGED SPLEEN, HIGHER APNEA-HYPOPNEA INDEX(AHI) AND HIGHER PERIODIC BREATHING EVENTS. PATIENT USED OXYGEN BAR TO REDUCE AHI AND ALSO ALLEGED OF HAVING SEVERE CENTRAL APNEA AND CHEYNE-STOKES.THE PATIENT DID REQUIRE MEDICAL INTERVENTION IN THE FORM OF PRESCRIBED ANTIBIOTICS. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE PATIENT HAS ALLEGED SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, NAUSEA, VOMITING AND CANCER (SKIN CANCER). SKIN CANCER. SECTION H6 HEALTH EFFECTS: CLINICAL CODES AND HEALTH EFFECTS - IMPACT CODE WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02592 | 2021-08-14 | 2021-07-06 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT HAS AN UNEXPLAINED HEART ATTACK. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED AN UNEXPLAINED HEART ATTACK. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN.THE REPORTED EVENT OF UNEXPLAINED HEART ATTACK DUE TO THE DEVICE AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENTS ARE ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURER CONCLUDES NO FURTHER ACTION IS NECESSARY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02593 | 2021-08-14 | 2021-07-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT HAS HAD A STROKE. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS HAD A STROKE. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE DEVICE WAS EVALUATED. AS PER THE INVESTIGATION OF THE DEVICE EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED IN THE BASE UNIT. AS PER SECONDARY FINDINGS OF THE BASE UNIT OBSERVED AN UNKNOWN CONTAMINANT ON THE REAR PANEL AND A DUST CONTAMINANT ON THE BLOWER AND THE BLOWER BOX. EVIDENCE OF LIQUID INGRESS WAS OBSERVED ON THE BLOWER. INV1023 ADDRESSES THE ISSUE OF LIQUID INGRESS THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE 4 ERRORS FOUND. THE MANUFACTURER CONCLUDES UNABLE TO ADDRESS THE SYMPTOMS DESCRIBED AND CONFIRM THE PRESENCE OF CONTAMINATION IN THE AIRPATH. IN THIS REPORT, SECTION D9, G3, H3, H6 HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02594 | 2021-08-14 | 2021-07-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500S11 | DSX500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED RESPIRATORY INFECTIONS. THE PATIENT DID REQUIRE MEDICAL INTERVENTION IN THE FORM OF PRESCRIBED ANTIBIOTICS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED THE PATIENT RESPIRATORY INFECTIONS WITH MUCUS. THE REPORTED EVENT OF RESPIRATORY INFECTIONS WITH MUCUS AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS POSSIBLY RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THE PATIENT DID REQUIRE MEDICAL INTERVENTION IN THE FORM OF PRESCRIBED ANTIBIOTICS. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED THE PATIENT RESPIRATORY INFECTIONS WITH MUCUS. THE PATIENT DID REQUIRE MEDICAL INTERVENTION IN THE FORM OF PRESCRIBED ANTIBIOTICS. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H5 WAS PREVIOUSLY BLANK NOW CORRECTLY UPDATED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02595 | 2021-08-14 | 2021-07-06 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED SINUS INFECTIONS. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECOME DEGRADED AND CAUSED THE PATIENT TO DEVELOP SINUS INFECTIONS. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE EVENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER QUALITY PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. DURING THE INITIAL VISUAL INSPECTION THE MANUFACTURER, THE DEVICE WAS DISASSEMBLED FOR INTERNAL VISUAL INSPECTION. THE MANUFACTURER FOUND NO SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WITHIN THIS DEVICE. THE MANUFACTURER FOUND CONTAMINATION ON THE BOTTOM PANEL INSIDE THE DEVICE. THE MANUFACTURER ALSO FOUND SMALL PIECE OF BLACK RUBBER IN THE BLOWER BOX UNDERNEATH THE BLOWER MOTOR AND CONTAMINATION WAS FOUND INSIDE THE BLOWER. THE MANUFACTURER CONCLUDES THERE IS NO EVIDENCE OF SOUND ABATEMENT FOAM BREAKDOWN WITHIN THE DEVICE. SECTION H6 HAS BEEN UPDATED / CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02817 | 2021-08-15 | 2017-10-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT WAS DIAGNOSED WITH LUNG CANCER. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE DETAILS OF THE DEVICE SERIAL NUMBER ARE UNKNOWN. IF FURTHER INFORMATION REGARDING THE DEVICE BECOME AVAILABLE, IT WILL BE SUBMITTED IN THE FOLLOW-UP REPORT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT WAS DIAGNOSED WITH LUNG CANCER. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS UPDATED REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. SECTION(S) D4 (MODEL#, SERIAL#, UDI#, CATALOG#), G4 AND H4 HAS BEEN CORRECTED AND UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02820 | 2021-08-15 | 2017-01-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED ASTHMA, CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD), LUPUS, AND MIGRAINES. THE PATIENT DID REQUIRE MEDICAL INTERVENTION IN THE FORM OF PRESCRIBED INHALERS. THE DETAILS OF THE DEVICE SERIAL NUMBER ARE UNKNOWN. IF FURTHER INFORMATION REGARDING THE DEVICE BECOME AVAILABLE, IT WILL BE SUBMITTED IN THE FOLLOW-UP REPORT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING OF THE PATIENT DEVELOPED ASTHMA, CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD), LUPUS, AND MIGRAINES AND THE PATIENT DID REQUIRE MEDICAL INTERVENTION IN THE FORM OF PRESCRIBED INHALERS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02818 | 2021-08-15 | 2021-01-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT WAS DEVELOPED A CHRONIC COUGH. THE PATIENT DID REQUIRE MEDICAL INTERVENTION IN THE FORM OF PRESCRIBED INHALERS AND STEROIDS. THE DETAILS OF THE DEVICE SERIAL NUMBER ARE UNKNOWN. IF FURTHER INFORMATION REGARDING THE DEVICE BECOME AVAILABLE, IT WILL BE SUBMITTED IN THE FOLLOW-UP REPORT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 06 , 2024 AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGING DEVELOPED A CHRONIC COUGH.THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION SECTION H6 WAS CORRECTED AND UPDATED | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02983 | 2021-08-15 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP, GER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DEX700S13 | DEX700S13 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED A SKIN DISEASE. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH THE PRODUCT PROBLEM AND ADVERSE EVENT WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK . (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECTS: IMPACT CODE WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02916 | 2021-08-15 | 2021-06-29 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION BIPAP AUTOSV H/HT/C, DS | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | DSX900T11C | DSX900T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED LUNG CANCER. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE DETAILS OF THE DEVICE SERIAL NUMBER ARE UNKNOWN. IF FURTHER INFORMATION REGARDING THE DEVICE BECOME AVAILABLE, IT WILL BE SUBMITTED IN THE FOLLOW-UP REPORT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE /RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM THAT BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE EXTREME HEADACHES, RESPIRATORY AND INFLAMMATION ISSUES, HYPERSENSITIVITY, LUNG CANCER 1 NODULE RIGHT LUNG AND IRRITATION. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. SINCE THERE IS NO CUSTOMER INFORMATION, THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, FOLLOW-UP REPORT WILL BE FILED. H6 HEALTH EFFECT- CLINICAL CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02915 | 2021-08-15 | 2021-07-22 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED KIDNEY CANCER. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE DETAILS OF THE DEVICE SERIAL NUMBER ARE UNKNOWN. IF FURTHER INFORMATION REGARDING THE DEVICE BECOME AVAILABLE, IT WILL BE SUBMITTED IN THE FOLLOW-UP REPORT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED KIDNEY CANCER. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THERE IS NO CUSTOMER INFORMATION AND WE CANNOT REACH OUT TO THE CUSTOMER, HENCE, ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION CAN NOT BE MADE AT THIS TIME. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS UPDATED REPORT WILL BE FILED.. SECTION H6 WAS MISSED TO CAPTURE IN PREVIOUS MDR ,NOW IT IS CORRECTED AND UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02914 | 2021-08-15 | 2018-01-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED PULMONARY NODULES. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE DETAILS OF THE DEVICE SERIAL NUMBER ARE UNKNOWN. IF FURTHER INFORMATION REGARDING THE DEVICE BECOME AVAILABLE, IT WILL BE SUBMITTED IN THE FOLLOW-UP REPORT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PULMONARY NODULES (DIAGNOSED 2018) HARM TO RESPIRATORY SYSTEM, CELLULAR DAMAGE, DNA DAMAGE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02766 | 2021-08-15 | 2021-06-28 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT WAS DIAGNOSED WITH BLADDER CANCER. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP,AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PATIENT WAS DIAGNOSED WITH BLADDER CANCER,HEADACHE AND BREATHING PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02828 | 2021-08-15 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED A COUGH. THE PATIENT DID REQUIRE MEDICAL INTERVENTION IN THE FORM OF PRESCRIBED ANTIBIOTICS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER ON 03/23/2023 FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED ON 04/19/2023. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP COUGH. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE EVENT IS CURRENTLY UNKNOWN. THE REPORTED EVENT AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THE EVENT IS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL PART OF THE DEVICE. OUTSIDE OF THE UNIT HAS NORMAL WEAR AND TEAR. LIGHT AMOUNT OF BEIGE DUST OR DIRT CONTAMINATE AND FIBER CONTAMINATE ON THE OUTER PANELS, SURFACE CRACKS AND CREVASSES, IN AND AROUND THE AIR INLET AREA, IN THE AIR INLET CANAL, AROUND THE SD CARD SLOT AREA, AROUND THE MODEM AREA, IN AND AROUND THE OUTLET PORT WAS OBSERVED. LIGHT AMOUNT OF SMALL DARK FIBERS IN AND AROUND THE OUTLET PORT WAS ALSO OBSERVED. INTERNAL INVESTIGATION OF THE UNIT FOUND A VERY SMALL AMOUNT OF BEIGE DUST OR DIRT CONTAMINATION ON THE TOP OF THE BLOWER BOX, NEAR THE AIR INTAKE AREA, AND IN THE BLOWER BOX CHIMNEY. BLOWER AND INTERNAL BLOWER BOX AREA CONTAINED A VERY SMALL AMOUNT OF BEIGE DUST OR DIRT CONTAMINATE. ON THE INSIDE OF BOTTOM ENCLOSURE, SMALL BUNDLE OF FIBRE OR HAIR CONTAMINATE WERE FOUND. METAL HOUSING OF P4 CONNECTOR HAS SOME MATERIAL DISCOLORING. THE LOWER FRONT UI PANEL CHANNEL HAD A SMALL AREA OF DRIED ORANGE LIQUID CONTAMINATE WITH HAIR OR FIBERS EMBEDDED IN IT. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN, MINOR CONTAMINATES ON THE INSIDE AND OUTSIDE OF THE UNIT AND THESE CONTAMINATES FOUND TO BE SOURCED FROM EXTERNAL ENVIRONMENTAL CONDITIONS. SECTIONS B1, B2 HAVE CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE (2199), MEDICAL DEVICE CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02765 | 2021-08-15 | 2021-07-22 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED CHEST PAIN, COUGH, AND SHORTNESS OF BREATH. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED THE WATER BEING BLACK AND SOMETIMES GRAY AND THE PATIENT DEVELOPED CHEST PAIN, COUGH, AND SHORTNESS OF BREATH & COLD "SYMPTOMS" SUCH AS WEAKNESS FATIGUE AND SWELLING. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE EVENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE THIRTY TWO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02827 | 2021-08-15 | 2021-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO SERIES ASSY, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | 750P | 750P |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT FOUND SPOTS ON BOTH LUNGS. THE PATIENT DID REQUIRE MEDICAL INTERVENTION IN THE FORM OF DOCTOR'S ASSESSMENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER INCORRECTLY CAPTURED H6- HEALTH EFFECT-CLINICAL CODE IN THE PREVIOUS MDR 2518422-2021-02827 AND IT IS CORRECTED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02984 | 2021-08-15 | 2021-08-09 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP, GER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DEX700S13 | DEX700S13 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED DIZZINESS. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE SECTIONS B1, B2, D4, H1 AND H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM. ( IN THE PREVIOUS MDR BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED.) SECTION B2 WAS CORRECTED TO BLANK. ( PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION D4 PRIMARY UNIQUE DEVICE IDENTIFIER (UDI) WAS CORRECTED. SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION SECTION H6 HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03236 | 2021-08-16 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES IN AIRWAY OF DEVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. THE MANUFACTURER FOUND EVIDENCE OF CONTAMINANT ON BOTTOM PANEL INSIDE DEVICE AND ON IMPELLER OF BLOWER MOTOR. THE MANUFACTURER FOUND EVIDENCE OF WATER INGRESS TO THE BLOWER BOX. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND NO ERRORS LOGGED. THE MANUFACTURER CONCLUDES THE DEVICE HAS CONTAMINATION CONSISTENT WITH WATER INGRESS. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION D9, G3 AND H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02976 | 2021-08-16 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | AUX500T15 | AUX500T15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE CHEST TIGHTNESS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A NASAL SPRAY AND AN INHALER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE CHEST TIGHTNESS. PATIENT STATING, SHE HAS BECOME UNWELL WITH A SORE THROAT, SNEEZING AND RUNNY NOSE, MEDICAL INTERVENTION IN THE FORM OF A NASAL SPRAY AND AN INHALER. NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 CODES CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT HAS BEEN CORRECTED AND UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03224 | 2021-08-16 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03219 | 2021-08-16 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | INX500S15 | INX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03222 | 2021-08-16 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03225 | 2021-08-16 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03237 | 2021-08-16 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NASAL/THROAT IRRITATION OR SORENESS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING NASAL/THROAT IRRITATION OR SORENESS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, FOLLOWUP REPORT WILL BE FILED. SECTION G1(INCORRECTLY CAPTURED IN PREVIOUS REPORT), H6 WAS CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03246 | 2021-08-16 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE THE MANUFACTURER RECEIVED INFORMATION ALLEGED PARTICLES IN AIRWAY FROM DEVICE,BLACK PARTICLES FOUND IN DEVICE. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION. SECTION H6 UPDATED IN THIS REPORT. CORRECTED INFORMATION WAS PROVIDED IN SECTION G1. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03252 | 2021-08-16 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03243 | 2021-08-16 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION G1 & H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03240 | 2021-08-16 |  | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03239 | 2021-08-16 |  | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. UNIT SCRAPPED DUE TO AGE. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03249 | 2021-08-16 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM.THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03281 | 2021-08-16 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGATION ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED TO SEEING PARTICLES IN THE TUBING AND CHAMBER AND HAS EXPERIENCED HEADACHES, SINUS INFECTIONS, MALAISE AND NAUSEA. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03251 | 2021-08-16 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03223 | 2021-08-16 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11W | DSX700T11W |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02763 | 2021-08-16 | 2021-07-22 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED SINUS INFECTIONS. THE PATIENT DID REQUIRE MEDICAL INTERVENTION IN THE FORM OF PRESCRIBED ANTIBIOTICS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03188 | 2021-08-16 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03157 | 2021-08-16 | 2020-01-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP AN INFECTION AND NODULES IN THEIR LUNGS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION RECIVED ON NOV 12, 2024 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP AN INFECTION AND NODULES IN THEIR LUNGS. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03244 | 2021-08-16 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL PART OF THE DEVICE AND THERE WERE NO SIGNS OF CONTAMINATION ON THE OUTSIDE OF THE DEVICE. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND FOUND UNKNOWN DUST CONTAMINATE IN THE BLOWER BOX AND ON THE FILTER PORT. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION, BUT DUST CONTAMINATION WAS OBSERVED AND ARE CONSISTENT WITH BEING FROM AN EXTERNAL SOURCE. SECTION H6 WERE UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03250 | 2021-08-16 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | INX500S15 | INX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03248 | 2021-08-16 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11 | DSX700T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03245 | 2021-08-16 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03277 | 2021-08-16 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT ALSO ALLEGED PARTICLES IN TUBING/CHAMBER, NASAL/THROAT IRRITATION OR SORENESS, DRY MOUTH, SINUS DRAINAGE AND SINUS INFECTION. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL PART OF THE DEVICE AND FOUND NO DISTINCT WEAR, TEAR, OR CONTAMINATION ON THE ENCLOSURE. A SLIGHT YELLOW DISCOLORATION WAS OBSERVED ON THE ISO AIR OUTLET PORT. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND FOUND NO DISTINCT WEAR, TEAR, OR CONTAMINATION. NO EVIDENCE OF SOUND ABATEMENT FOAM BREAKDOWN WAS OBSERVED, AND THE FOAM APPEARS INTACT. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE DEVICE WAS POWERED ON AND WAS FOUND TO OPERATE PROPERLY. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION AND NO PROBLEM FOUND. SECTION D9, SECTION H3, CODES IN SECTION H6 WERE UPDATED/ CORRECTED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03155 | 2021-08-16 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, JP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | RJPX500S16 | RJPX500S16 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03156 | 2021-08-16 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, JP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | RJPX500S16 | RJPX500S16 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03192 | 2021-08-16 |  | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03190 | 2021-08-16 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03187 | 2021-08-16 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03280 | 2021-08-16 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED TO SEEING THE FILTER WAS BLACK, THROAT IS BURNING, LIPS WERE SWOLLEN. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02927 | 2021-08-16 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP NODULES IN THEIR LUNGS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON OCTOBER 09, 2023, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT THE PATIENT ALLEGED TO DEVELOP NODULES IN THEIR LUNGS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGED DIZZINESS AND HEADACHE. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. . THE SECTION E1 WAS UPDATED AND SECTIONS B1, B2, H1 AND H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED ONLY FOR PRODUCT PROBLEM (BOTH ADVERSE EVENTS AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK.(PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECTS - CLINICAL CODES AND HEALTH EFFECTS-IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03202 | 2021-08-16 | 2021-07-08 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION BIPAP AUTOSV H/HT/C, DS | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS,INC | DSX900T11C | DSX900T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP PULMONARY FIBROSIS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP PULMONARY FIBROSIS. ADDITIONAL INFORMATION WAS RECEIVED, AND THE PATIENT REPORTED THAT THEY HAD A HEART ATTACK AND DIFFICULTY BREATHING/SHORT OF BREATH AND AFTER ONE YEAR OF USE, HAS PULMONARY FIBROSIS AND IS PROGRESSING, WITH AREAS OF SUB-PLEURAL RETICULATION AT THE LUNG BASE. THE PATIENT ALSO REPORTED THAT THEY NOTICED A LONG HAIR SIMILAR TO A CAT HAIR, ABOUT 2 MM LONG, ON HIS PILLOW BUT WERE NOT SURE IF THIS WAS ANYTHING OR NOT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO THE ALLEGED SERIOUS INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HEALTH EFFECT - CLINICAL CODE HAS BEEN ADDED, AND TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03201 | 2021-08-16 | 2021-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400T11 | DSX400T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP HEADACHES AND SHORTNESS OF BREATH. THE PATIENT HAD HEART VALVE SURGERY IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. PATIENT ALLEGED TO DEVELOP HEADACHES AND SHORTNESS OF BREATH. THE PATIENT HAD HEART VALVE SURGERY IN RESPONSE TO THE REPORTED EVENT. THE REPORTED EVENT AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02926 | 2021-08-16 | 2021-07-06 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP PNEUMONIA. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP PNEUMONIA. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION B2 WAS CORRECTED AND H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03158 | 2021-08-16 | 2018-05-15 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP KIDNEY DISEASE, DIABETES, CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD), CONGESTIVE HEART FAILURE (CHF). THE PATIENT REQUIRED CARDIAC SURGERY AND MEDICATION TO TREAT THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PATIENT TO DEVELOP KIDNEY DISEASE, DIABETES, CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD), CONGESTIVE HEART FAILURE (CHF),SINUS ISSUES, RUNNY NOSE, HEADACHE, SWELLING, SORE THROAT, LOSS OF TASTE AND SMELL, STOMACH PROBLEMS, DIABETES, KIDNEY DISEASE, LUNG ISSUES, COPD, ANXIETY, WEAKNESS IN LEG, HEART STENT, CHF, AND AFIB. THE PATIENT REQUIRED CARDIAC SURGERY AND MEDICATION TO TREAT THE REPORTED EVENT RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY.THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE IS NO CUSTOMER INFORMATION AND WE CANNOT REACH OUT TO THE CUSTOMER, HENCE, ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION CAN NOT BE MADE AT THIS TIME. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 HEALTH IMPACT CODE WAS NOT CAPTURED IN INTIAL MDR,NOW IT IS CORRECTED AND UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03278 | 2021-08-16 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03110 | 2021-08-16 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE WAS RECALIBRATED TO ADDRESS THE ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03061 | 2021-08-16 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03092 | 2021-08-16 | 2021-07-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP RESPIRATORY INFLAMMATION. THE PATIENT WAS PRESCRIBED AN INHALER IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A PATIENT TO DEVELOP RESPIRATORY INFLAMMATION RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT WAS PRESCRIBED AN INHALER IN RESPONSE TO THE REPORTED EVENT. THE REPORTED EVENT IS ASSESSED AS SERIOUS INJURY BUT NOT RELATED TO THE DEVICE. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03076 | 2021-08-16 | 2021-06-24 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP BRONCHITIS. THE PATIENT REQUIRED A BREATHING TREATMENT IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03100 | 2021-08-16 | 2021-08-10 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION GO AUTO W/BT, BRAZIL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | BRG502S18 | BRG502S18 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED A SPOT ON HER LUNG. THE PATIENT DID REQUIRE MEDICAL INTERVENTION IN THE FORM OF PRESCRIBED MEDICATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED A SPOT ON HER LUNG. THE PATIENT DID REQUIRE MEDICAL INTERVENTION IN THE FORM OF PRESCRIBED MEDICATION. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03121 | 2021-08-16 | 2020-01-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID/HT, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | AUX400T15 | AUX400T15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED ASTHMA. THE PATIENT DID REQUIRE MEDICAL INTERVENTION IN THE FORM OF PRESCRIBED MEDICATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03035 | 2021-08-16 | 2021-08-09 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO, FR | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | FRX500S14 | FRX500S14 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED CANCER. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM AND THE PATIENT DEVELOPED CANCER. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND OBSERVED SMALL BLACK MATTER ON WORK SURFACE SHORTLY AFTER BEGINNING INVESTIGATION. A SMALL AMOUNT OF UNKNOWN CONTAMINATION WAS FOUND ON THE EDGE OF THE AIR INTAKE COVER. INTERNAL INVESTIGATION OF THE BASE UNIT NOTING MORE BUILDUP OF BLACK CONTAMINANTS, PARTICULARLY IN THE ENCLOSURE JOINTS. UPON EXAMINATION UNDER A MICROSCOPE, THE CONTAMINANTS PROVED TO BE SMALL INSECTS APPROXIMATELY 2-3MM LONG AND LESS THAN 1MM WIDE. PIL WAS ABLE TO CONFIRM THAT THERE ARE POTENTIAL IRRITANTS IN THE AIRPATH WHICH COULD BE A CAUSE OF THE COMPLAINT. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS. THE MANUFACTURER CONCLUDES THE DEVICE HAS CONTAMINATION WHICH HAS PROVIDED THE INSECTS. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03099 | 2021-08-16 | 2021-08-09 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, GER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DEX500S13 | DEX500S13 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED BILATERAL PULMONARY EMBOLISM. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03093 | 2021-08-16 | 2021-07-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP RESPIRATORY INFLAMMATION. THE PATIENT WAS PRESCRIBED AN INHALER IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT ALLEGED TO DEVELOP RESPIRATORY INFLAMMATION. THE PATIENT WAS PRESCRIBED AN INHALER IN RESPONSE TO THE REPORTED EVENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02985 | 2021-08-16 | 2021-08-09 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO, GER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DEX400S13 | DEX400S13 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOPED ASTHMA. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED THAT A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM ALLEGEDLY BECAME DEGRADED AND THE PATIENT DEVELOPED ASTHMA. THE DETAILS OF THE PATIENT'S REQUIRED MEDICAL INTERVENTION ARE UNKNOWN. THE PATIENT ALSO ALLEGED HAVING FATIGUE. THE DEVICE WAS RETURNED TO THE MANUFACTURER QUALITY PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. DURING THE INITIAL VISUAL INSPECTION OF THE DEVICE NO DISTINCT WEAR OR TEAR OF THE ENCLOSURE WAS FOUND. UNIDENTIFIED CONTAMINATION WAS OBSERVED ON THE EXTERIOR OF THE UPPER ENCLOSURE AND ON THE BACK PANEL NEAR THE AIR OUTLET PORT. THE DEVICE WAS DISASSEMBLED FOR INTERNAL VISUAL INSPECTION. THE MANUFACTURER FOUND VERY FINE DUST CONTAMINATION THROUGHOUT THE ENTIRE AIR PATH, AS WELL AS DUST CONTAMINATION IN THE BLOWER. A DARK CONTAMINANT AROUND THE BLOWER OUTLET SEAL APPEARS CONSISTENT WITH THE KERATIN CONTAMINATION. NO EVIDENCE OF SOUND ABATEMENT FOAM BREAKDOWN WAS OBSERVED, AND THE FOAM APPEARS INTACT. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND CONFIRMED THE DEVICE PROVIDED FLOW. THE MANUFACTURER CONCLUDES THERE IS NO EVIDENCE OF SOUND ABATEMENT FOAM BREAKDOWN WITHIN THE DEVICE. THE CONTAMINATION FOUND IN THE AS OBSERVED IN THE BLOWER AND AIR PATH. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03217 | 2021-08-16 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO, INTL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | 761P | 761P |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03231 | 2021-08-16 | 2021-07-22 | I, n, j, u, r, y | Manufacturer report | BIPAP AUTO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP UPPER RESPIRATORY INFECTIONS AND KIDNEY CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM AND CAUSED A PATIENT TO DEVELOP UPPER RESPIRATORY INFECTIONS AND KIDNEY CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. SINCE THERE IS NO CUSTOMER INFORMATION, THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103287 | 2021-08-16 | 2020-01-20 | I, n, j, u, r, y | Voluntary report | BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX600S11 |  |  | SWOLLEN THROAT/GLANDS HEADACHE; I THOUGHT MY BREATHING PROBLEMS AND SWOLLEN GLANDS WERE FROM A TOOTH THAT MAY NEED FILLING. BUT THEN MY THROAT STARTED SWELLING AND EARS RINGING. MY MOUTH WOULD GET SO DRY IF WAKE UP WITH MY TONGUE STUCK TO ROOF OF MOUTH. I HAD TO POOR IN WATER TO MOVE MY TONGUE AROUND. ALSO, DRY NOSE AND HEADACHES. THE SMELL OF CHEMICALS WERE OVERWHELMING, BUT DISSIPATED SOMEWHAT AFTER ALLOWING AIR TO BLOW BEFORE APPLYING TO MY FACE. NO TEST TAKEN YET. DOCTOR APPOINTMENT ON THE (B)(6) 2021. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103285 | 2021-08-16 | 2021-08-12 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS RESPIRONICS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | MY BIPAP IS ON THE RECALL LIST. THE S/N IS REGISTERED. SOMEONE REGISTERED THE S/N FOR ME. THERE ARE A FEW ISSUES WITH THAT. FIRST, WHEN SOMEONE WOULD CALL AND THEY WOULD NOT CALL ME. SECOND, I WOULD NOT RECEIVE ALL OF THE INFORMATION. I CALLED (B)(4) ON (B)(6) 2021 ASKING IF I WOULD RECEIVE A REPLACEMENT. SHE ASKED ONE QUESTION WHICH WAS IF THE S/N IS REGISTERED AND I SAID YES. INSTEAD OF ASSISTING ME AND SHE SAID SOMEONE WOULD CALL TO LET ME KNOW IF I NEED A REPAIR OR A REPLACEMENT. SHE DIDN'T LOOKUP ANYTHING. SHE DIDN'T ASK FOR THE S/N OR ANYTHING BUT SIMPLY DID NOT ASSIST. THE S/N IS (B)(4). FDA SAFETY REPORT ID# (B)(4). |  | N | Y | I |  | 2021-08
MW5103279 | 2021-08-16 | 2021-06-15 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | SHORTNESS OF BREATH; I HAVE BEEN USING MY PHILIPS DREAMSTATION FOR ABOUT 4 YEARS. IN THE LAST 6 MONTHS, I HAVE BEEN NOTICING THAT I WAKE UP GASPING FOR AIR AND A HORRIBLE TASTE IN MY MOUTH. MY HUMIDIFIER TRAY IS FILLED WITH SLIME THAT HAS TO BE CLEANED OUT EVERY DAY. I WAKE UP FEELING LIKE I DIDN'T SLEEP WELL. I WENT ONLINE TO SEE IF I COULD FIND ANYONE ELSE WITH THESE ISSUES AND FOUND THE RECALL NOTICE. I WENT TO PULMONOLOGIST AND WAS SENT FOR A CHEST XRAY AND A PULMONARY FUNCTION TEST. I WAS ABLE TO GET A RESMED DEVICE AS A REPLACEMENT AND MY SLEEP AND BREATHING HAVE BEEN MAKING GOOD IMPROVEMENTS BACK TOWARDS NORMAL. PHILIPS HAS HAD NO COMMUNICATION WITH ME SINCE I RECORDED MY DEVICE AS A RECALL ELIGIBLE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103277 | 2021-08-16 | 2020-09-02 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION AUTO CPAP HUMHT DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500T11 |  |  | I USED A PHILIPS CPAP FOR SEVERAL YEARS. I DEVELOPED DAILY DIZZINESS AND HEADACHES. MY NEUROLOGIST AT (B)(6) IN (B)(6) DIAGNOSED ME WITH INNER EAR DAMAGE. IN LIEU OF THE RECENT RECALL OF PHILIP'S CPAP MACHINES, I THINK THIS MAY HAVE BEEN CAUSED BY THE CPAP SINCE THESE SYMPTOMS WERE LISTED AS A POSSIBLE SIDE EFFECT OF THE GAS. MY NEUROLOGICAL WORKUP NEGATIVE, INCLUDING MRIS AND MY LAB WORK WAS NEGATIVE. |  | Y | N | LAY USER/PATIENT |  | 2021-08
2518422-2021-02748 | 2021-08-16 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP ST, C SERIES, INTERNATIONAL | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS INC | 1061421 | 1061421 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. CORRECTED DATA: D8 WAS THIS DEVICE SERVICED BY A THIRD PARTY?: PREVIOUSLY REPORTED AS NO ; UPDATED TO UNKNOWN | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03182 | 2021-08-16 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED.THE PATIENT DID NOT RECEIVE MEDICAL INTERVENTION.THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. ADDITIONAL INFORMATION AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES, DIFFICULTY IN BREATHING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM.THERE WAS NO REPORT OF PATIENT HARM OR INJURY. AFTER FIRST ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WAS UNSUCCESSFUL AND CUSTOMER DECLINED TO RESPOND TO THE GFE RELATED QUESTIONS AND TERMINATED THE CALL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03180 | 2021-08-16 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03181 | 2021-08-16 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03052 | 2021-08-16 | 2021-08-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S KEYPAD AND A MALFUNCTION OF THE LCD WAS OBSERVED. THE DEVICE'S FRONT PANEL BOARD AND LCD DISPLAY WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03051 | 2021-08-16 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG AND A MALFUNCTION OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICE'S POWER MANAGEMENT BOARD AND THE LCD WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03050 | 2021-08-16 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03247 | 2021-08-16 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT ALSO ALLEGED DEVICE BLINKS, SEEING LINE PARTICLES ARE IN CHAMBER AND HAS ODOR. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND FOUND EVIDENCE OF WATER INGRESS (BLOWER/FOAM/BLOWER BOX) CONFIRMED POSITIVE DEPOSIT OF FOAM PARTICULATE ON THE FIT TOOL. CONFIRMED EVIDENCE OF FOAM DEGREDATION ASSOCIATED TO THIS ALLEGATION. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION, BUT PRESENCE OF LIQUID INGRESS WERE OBSERVED AND ARE CONSISTENT WITH BEING FROM AN EXTERNAL SOURCE. IN INITIAL REPORT, SECTION G3 WAS WRONGLY CAPTURED AS 08/09/2021 SHOULD BE 07/14/2021, CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03227 | 2021-08-16 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A VENTILATOR INOPERATIVE CONDITION WAS OBSERVED. THE DEVICE'S FIO2 TUBING WAS FOUND TO BE OCCLUDED. THE DEVICE'S FIO2 TUBING WAS CLEANED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03226 | 2021-08-16 | 2021-07-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S AC PLUG CAUGHT FIRE WHEN PLUGGED INTO AN AC SOURCE. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. DURING THE INVESTIGATION BY THE MANUFACTURER, IT WAS CONCLUDED THE EVENT WAS CAUSED BY AN OUTSIDE SOURCE, DUST AND DIRT CONTAMINATION. THE DEVICE'S AC CONNECTOR AND BASE ENCLOSURE NEED REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03228 | 2021-08-16 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A SERVICE REQUIRED CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A VENTILATOR SERVICE REQUIRED CONDITION WAS OBSERVED. THE DEVICE'S FIO2 TUBING WAS FOUND TO BE OCCLUDED. THE DEVICE'S FIO2 TUBING WAS CLEANED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03191 | 2021-08-16 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03194 | 2021-08-16 | 2021-08-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED DREAMSTATION HAS BLACK SPECS IN THE WATER TANK. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. ADDITIONAL INFORMATION WAS RECEIVED AND ADDED TO THE REPORT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER ON 09/03/2022 FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED ON 09/03/2022. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. SECTION D8, D9, H2, H3 AND H6 WERE UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03186 | 2021-08-16 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03189 | 2021-08-16 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03183 | 2021-08-16 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED TO SEEING THE PARTICLES IN AIRPATH. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03184 | 2021-08-16 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM.THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03185 | 2021-08-16 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03179 | 2021-08-16 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03062 | 2021-08-16 | 2021-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, A MALFUNCTION OF A VENTILATOR'S LCD DISPLAY WAS OBSERVED. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICES LCD DISPLAY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03193 | 2021-08-16 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN THE AIRWAY COUGHING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION G3 DATE RECEIVED BY MANUFACTURER INCORRECTLY CAPTURED IN INITIAL REPORT IT SHOULD BE 14-JUL-2021. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03260 | 2021-08-17 | 2021-06-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED ON THIS DEVICE IN MDR 2518422-2021-01314. THIS REPORT WAS SUBMITTED AS A DUPLICATE REPORT OF THE PREVIOUSLY SUBMITTED REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03282 | 2021-08-17 | 2021-07-22 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD). THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. SECTION B5 SHOULD BE PREVIOUSLY REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF BLACK PARTICLES IN WATER CHAMBER AND TUBING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT ALLEGED TO DEVELOP CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD) AND LUNG CAPACITY DROPPED . THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03254 | 2021-08-17 | 2021-07-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03253 | 2021-08-17 | 2021-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WOULD NOT POWER ON. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03256 | 2021-08-17 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND ACTIVE EXHALATION CONTROL MODULE WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03255 | 2021-08-17 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03229 | 2021-08-17 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2100X15B | IN2100X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE ASSEMBLY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03329 | 2021-08-17 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, A VENTILATOR'S LCD SCREEN WAS BLANK. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03323 | 2021-08-17 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2100X15B | IN2100X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A "VENTILATOR INOPERATIVE" CONDITION WAS OBSERVED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03324 | 2021-08-17 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
MW5103311 | 2021-08-17 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAMSTATION 2 | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | NOTICED SOMETHING WHITE ON THE SIDE OF MY MOUTH AFTER WAKING UP AND TAKING THE CPAP MASK OFF. I WASN'T SURE WHAT IT WAS, SO I THREW IT AWAY. BUT I HAD MY CPAP MASK ON ALL NIGHT, SO I ASSUME IT HAD TO COME THROUGH OR FROM THE MACHINE. I'VE ALSO EXPERIENCED RECURRING HEADACHES. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103312 | 2021-08-17 | 2021-08-11 | I, n, j, u, r, y | Voluntary report | CPAP DREAMMACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I USED MY CPAP MACHINE AND HAVE SEEN BLACK SPECS IN THE WATER TANK. I HAVE A VERY SORE THROAT AND FEEL DIZZY. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-08
MW5103313 | 2021-08-17 | 2018-08-01 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS SYSTEM ONE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | 560P |  |  | WHEN PHILIPS RELEASED THEIR RECALL ON CPAP MACHINES AND THEN DESCRIBED THE SYMPTOMS ATTRIBUTABLE TO THIS RECALL, I NOTICED I'VE BEEN HAVING THESE FOR A WHILE NOW. I CAN'T PINPOINT THE EXACT DATE, SO I JUST PUT (B)(6) 2018 AS AN ESTIMATE. THE SYMPTOMS I HAVE ARE SINUS INFECTIONS, HEADACHES, MIGRAINES, EYE IRRITATION, WHEEZING, COUGHING. I'VE BEEN HAVING THESE FOR SEVERAL YEARS AND HAVE BEEN/CALLED DOCTORS ABOUT THEM, BUT THEY SAID I JUST NEED TO CLEAN MY CPAP MORE OFTEN. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-08
MW5103316 | 2021-08-17 | 2021-06-14 | I, n, j, u, r, y | Voluntary report | DREAM STATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS, INC. |  |  |  | TWO MONTHS AGO, PHILIPS RESPIRONICS ANNOUNCED A RECALL FOR CPAP MODELS, INCLUDING THE ONES THAT MY WIFE AND I USE. I REGISTERED ON THEIR WEB-SITE JUST FOUR DAYS LATER ON JUNE 18TH. HERE IT IS, AUGUST 14TH TWO MONTHS LATER AND WE ARE STILL WAITING FOR THEM TO REPAIR OR REPLACE OUR UNITS. I'VE CALLED THEM EVERY WEEK, AND THEY KEEP GIVING ME THE RUN-AROUND. WE ARE ENDANGERING OUR HEALTH BY CONTINUING TO USE THEIR PRODUCT, BUT THEY OBVIOUSLY DON'T CARE. SO WE ARE BUYING NEW MACHINES AT OUR OWN EXPENSE KNOWING IT IS UNLIKELY THAT WE WILL BE REIMBURSED. IS THIS ANY WAY TO TREAT THEIR CUSTOMERS? NO. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | I |  | 2021-08
MW5103319 | 2021-08-17 | 2021-08-14 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | BLACK SPECKS AND FLAKES FOUND IN TUBE AND HUMIDIFIER CHAMBER. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103320 | 2021-08-17 |  | I, n, j, u, r, y | Voluntary report | DREAM STATION AUTO CPAP WITH HUM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | DSX500H11 | IP22 |  | I HAVE CONTACTED THE CPAP PROVIDER AND MY PULMONOLOGIST ON A COUPLE OCCASIONS WITH MY ISSUES AS I HAVE HAD MY CPAP FOR YEARS AND THE ISSUES COME WITH USAGE OVER THE LAST COUPLE YEARS. I HAVE WOKE UP GASPING FOR AIR. I WAKE UP WITH HEADACHES AND HEAD PAIN ON THE LEFT SIDE OF MY HEAD WHICH I NEVER HAVE HAD THIS KIND OF PAIN. I WAKE UP FEELING LIKE I HAVE FLUID IN MY LUNGS AND IT TAKES A COUPLE HOURS AND MY INHALER(S) TO GET ENOUGH OUT WHERE I'M NOT COUGHING OR MY CHEST FEELING FULL. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103321 | 2021-08-17 |  | I, n, j, u, r, y | Voluntary report | DREAM MACHINE CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | ON RECALL LIST |  |  | MACHINE HAS HAD PARTICLES IN THE WATER CHAMBER WHEN IT HAS GONE DRY. EACH TIME THE MACHINE WENT DRY OF WATER IT MADE ME SICK WITH ALLERGIC SYMPTOMS THAT LASTED SEVERAL DAYS. THERE'S A RECALL. I FILLED OUT THE RECALL FORM 2 MONTHS AGO AND HAVE HEARD NOTHING. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103322 | 2021-08-17 | 2021-04-17 | I, n, j, u, r, y | Voluntary report | DREAM STATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500H11C |  |  | SEE BELOW EXPLANATION; I WAS SLEEPING WITH MY PHILIPS RESPIRONICS DREAMSTATION CPAP MACHINE ON. AT ABOUT 2 AM I WOKE UP COUGHING, HAVING DIFFICULTY BREATHING, AND WITH A HEADACHE. THE AIR THAT WAS COMING OUT OF MY CPAP MACHINE SMELLED LIKE TOXIC BURNING PLASTIC, AND THE BLOWER WAS MAKING A HIGH-PITCHED WHINING NOISE. I SHUT THE MACHINE OFF, AND GOT MY OLD CPAP AND USED IT TO GO BACK TO SLEEP. IN THE MORNING I OPENED THE MACHINE UP AND EXAMINED IT. I FOUND BLACK PARTICLES IN MY HUMIDIFIER TANK, AND BLACK DEBRIS IN THE AIR PASSAGES. I AM ENCLOSING PHOTOGRAPHS OF WHAT I FOUND. I HAVE NOT BEEN ABLE TO USE THE MACHINE SINCE THIS OCCURRED. I CONTACTED MY DME, (B)(4), AND TOOK IT IN TO BE SEEN BY THEIR TECHNICIAN. HE INFORMED ME THAT HE COULD SEND IT TO THE REPAIR FACILITY BUT THE COST OF REPAIRING IT WOULD BE ALMOST AS MUCH AS THE COST OF A NEW MACHINE. I ELECTED NOT TO HAVE IT SENT IN AND CONTACTED PHILIPS RESPIRONICS TO SEE WHAT THEY COULD DO. ON (B)(6) 2021 I SPOKE TO (B)(6) AND HE GAVE ME AN RGA NUMBER TO SEND THE MACHINE IN TO HAVE IT EXAMINED. I EXPLAINED WHAT HAD HAPPENED AND E-MAILED HIM PHOTOS OF THE PROBLEM WITH THE MACHINE. TWO DAYS LATER I WAS CALLED BY (B)(6) FROM PHILIPS RESPIRONICS AND I ALSO SENT HER PICTURES OF THE PROBLEM WITH THE MACHINE. I HAVE NOT SENT THE MACHINE INTO THEM AND I AM KEEPING IT FOR EVIDENCE. ABOUT A MONTH AFTER THIS ALL HAPPENED PHILIPS ISSUED A RECALL. I HAVE REGISTERED ON THEIR WEBSITE. AT THIS POINT I AM WITHOUT MY CPAP MACHINE AND PHILLIPS HAS NOT PROVIDED ANY REMEDY. I HAVE ATRIAL FIBRILLATION AND FAILURE TO TREAT FOR SLEEP APNEA CAN BE LIFE THREATENING. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103327 | 2021-08-17 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAMSTATION CPAP, AUTO CPAP, BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. | DSX400H11C | DSX400H11C |  | I WAS CLEANING THE HOSE TO MY DREAMSTATION CPAP DEVICE AND WHEN I FLUSHED IT WITH CLEAN WATER I NOTICED BLACK DEBRIS COMING OUT OF IT. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | I |  | 2021-08
MW5103330 | 2021-08-17 | 2021-07-15 | I, n, j, u, r, y | Voluntary report | CONTINUOUS VENTILATOR | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | SYSTEM ONE REMSTAR AUTO A-FLEX |  |  | I HAVE BEGUN WAKING UP WITH A HEADACHE AND SINUS PAIN EACH MORNING FOR ABOUT A MONTH. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03261 | 2021-08-17 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND NO ERROR LOGGED. THE MANUFACTURER CONFIRMED THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. IN THE PREVIOUSLY SUBMITTED REPORT SECTION B5 WAS INCOMPLETE. CORRECT SECTION B5 IS- THE PATIENT ALSO ALLEGED VISUALIZATION OF BLACK PARTICLES IN THE WATER CHAMBER. SECTION D9, G3 AND H6 HAS BEEN UPDATED/CORRECTED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03152 | 2021-08-17 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03276 | 2021-08-17 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02519 | 2021-08-17 | 2021-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03273 | 2021-08-17 | 2021-06-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP IDIOPATHIC PULMONARY FIBROSIS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP IDIOPATHIC PULMONARY FIBROSIS. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENT OF HEART ISSUES AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. THE PATIENT HAS SWAPPED OUT THE DEICE AT THE DISTRIBUTOR AND AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03275 | 2021-08-17 | 2021-07-22 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX400T11 | DSX400T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A TUMOR IN THEIR THROAT. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03274 | 2021-08-17 | 2021-06-28 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A LUNG INFECTION. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. CORRECTION: EVENT DATE IN B3 WAS INCORRECTLY REPORTED AS (B)(6) 2018. THE CORRECT EVENT DATE IS (B)(6) 2021. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03271 | 2021-08-17 | 2021-07-06 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP STAGE 2 KIDNEY FAILURE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP STAGE 2 KIDNEY FAILURE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND EVIDENCE OF A LIGHT COATING, BEIGE, DARK DUST OR DIRT CONTAMINATE AND DARK FIBER CONTAMINATE IN THE BLOWER HOUSING. THE MANUFACTURER FOUND EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION /BREAKDOWN. THE DEVICE'S EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND 1 ERROR WAS LOGGED. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH DUST OR DIRT, FIBERS, LIGHT COATING AND BEIGE. THE MANUFACTURER CONFIRMED THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. SECTION D9, G3, H3 AND H6 HAS BEEN UPDATED / CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03117 | 2021-08-17 | 2021-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TO POWER ON. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS NOT CONFIRMED. DURING THE EVALUATION AT THE MANUFACTURER'S SERVICE CENTER, VENTILATOR INOPERTATIVE CODES WERE OBSERVED IN THE DOWNLOADED EVENT LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
MW5103306 | 2021-08-17 |  | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DREAMSTATION |  |  | COUGH WITH INITIATION OF CPAP THERAPY. FDA SAFETY REPORT ID# (B)(4). |  | Y | N | LAY USER/PATIENT |  | 2021-08
MW5103299 | 2021-08-17 | 2021-07-01 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | USED CPAP SINCE 2001. CHANGED TO PHILLIPS RESPIRONICS 2018. BOUGHT SOCLEAN OZONE CLEANER 2019. DEVELOPED COUGH GRADUALLY OVER TIME. BECAME A PROBLEM SPRING 2021. NOTICED NEWS ITEM ABOUT RECALL OF MY CPAP MACHINE SUMMER 2021. CONSULTED WITH MY PULMONARY DOCTOR. CHEST X-RAY TAKEN AND WAS UNREMARKABLE. DISCONTINUED MY CPAP AND SOCLEAN. RETURNED TO EARLIER MODEL CPAP. COUGH WENT AWAY IN THREE WEEKS. FEEL BETTER NOW. |  | Y | N | I |  | 2021-08
2518422-2021-03285 | 2021-08-17 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED PARTICLES IN TUBING/CHAMBER, NOISY MACHINE, DRY MOUTH AND DRY EYE. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED. SECTION H6 WAS CORRECTED (ADDED DRY EYES, DRY MOUTH EVENTS) AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03289 | 2021-08-17 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON 10/09/2023, AND SECTION H10 SHOULD BE REPORTED AS: HE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISIBLE DEBRIS IN AIRPATH CIRCUITS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT SERIOUS HARM OR INJURY. ADDITIONAL INFORMATION WAS RECEIVED ABOUT ALLEGED EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, DIZZINESS OR HEADACHE, NAUSEA AND VOMITING. SECTION H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03288 | 2021-08-17 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX600T11C | DSX600T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03287 | 2021-08-17 | 2021-07-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS INITIALLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THERE WAS NO PATIENT HARM OR INJURY. IN INITIAL REPORT SECTION B5 WAS MENTION INCORRECTLY, CORRECT B5 SHOULD BE - THE PATIENT ALLEGED PARTICLES IN AIR PATH, HEADACHES, VOMITING. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND UNKNOWN WHITE AND BLACK DUST CONTAMINATION INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM, AND HAIRS AT THE AIR INPUT OF THE BLOWER BOX. FOUND BLACK CONTAMINATION, CONSISTENT WITH KERATIN, AT THE AIR OUTLET OF THE BLOWER BOX. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER AND FOUND UNKNOWN WHITE AND BROWN CONTAMINATION, ALONG WITH SOME BLACK HAIRS WERE ON THE BOTTOM OF THE BLOWER BOX. ALSO FOUND MINERAL DEPOSITS ON THE BOTTOM OF THE BLOWER MOTOR, INDICATING POTENTIAL WATER INGRESS. THE HUMIDIFIER WAS RETURNED FOR EVALUATION. THE MANUFACTURER FOUND UNKNOWN BROWN AND BLACK CONTAMINANTS AND DARK-COLORED HAIR STRANDS INSIDE OF HUMIDIFIER BOX. ALSO FOUND AIR INLET TUBE SEAL WAS OFF OF HUMIDIFIER AND LAYING IN THE BOTTOM OF THE ENCLOSURE. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS LOGGED. THE MANUFACTURER CONCLUDES MINERAL CONTAMINATION CONSISTENT WITH WATER INGRESS AND MULTIPLE CONTAMINANTS THAT ARE NOT CONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03305 | 2021-08-17 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. ADDITIONAL INFORMATION WAS RECEIVED, SECTION B5 WAS CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PARTICLES IN AIRWAY FROM DEVICE, DRY MOUTH FROM THE MACHINE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. SECTION H6 WAS CORRECTED BY CAPTURING HEALTH EFFECT - CLINICAL CODE WHICH WAS INCORRECT IN INITIAL REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03307 | 2021-08-17 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENR ALLEGED PARTICLES IN TUBING CHAMBER. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03309 | 2021-08-17 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON SEP 18, 2023, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES IN TUBING/CHAMBER, PARTICLES IN AIRWAY AND BLACK PARTICLES IN THE WATER AND TUBING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGATION OF ASTHMA (NEW OR WORSENING), KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, OTHER (UNSPECIFIED HEALTH CONDITION) AND LUNG SCARRING. THE SECTION E1 WAS UPDATED AND SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY THE PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECTS: CLINICAL CODES AND HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03303 | 2021-08-17 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DIZZY AND FOGGY , ALSO SEES A LOT OF PARTICLES IN DEVICE AND DEVICE HAS A STRANGE SMELL AND IS MAKING NOISE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS:THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DIZZY AND FOGGY , ALSO SEES A LOT OF PARTICLES IN DEVICE AND DEVICE HAS A STRANGE SMELL AND IS MAKING NOISE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 WAS CORRECTED (EVENT DIZZINESS WAS NOT UPDATED IN PREVIOUS REPORT) AND WAS UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03308 | 2021-08-17 | 2016-01-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING CHRONIC COUGH AND LUNG IRRITATION, WITH SINUS ISSUES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. CUSTOMER DECLINED TO RESPOND TO THE GFE RELATED QUESTIONS AND TERMINATED THE CALL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. CORRECTED B5SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE USER ALLEGED CHRONIC COUGH AND LUNG IRRITATION, WITH SINUS ISSUES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE PATIENT DID NOT REPORT TO RECEIVE MEDICAL INTERVENTION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. AN UPDATED REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03304 | 2021-08-17 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED SEEING BLACK PARTICLES IN THE DEVICE.. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03306 | 2021-08-17 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM.THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES.THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED.THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED.THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED.THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER.THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND UNIT SCRAPPED DUE TO AGE. THE MANUFACTURER MENTIONED INCORRECT SECTION H9-CORRECTION/REMOVAL REPORTING NUMBER IN PREVIOUS REPORT. IT HAS BEEN UPDATED IN THIS REPORT. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03301 | 2021-08-17 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT ALSO ALLEGED BLACK STUFF IN MACHINE AND HAVING NASAL/THROAT BURNING AND IRRITATION. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL PART OF THE DEVICE AND NOT FOUND ANYTHING. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND FOUND EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION H6 WERE UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03300 | 2021-08-17 |  | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03283 | 2021-08-17 | 2021-08-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, BL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | BLX500S15 | BLX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG CANCER. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A SURGERY FOR TUMOR REMOVAL. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. CORRECTED DATA: D8 WAS THIS DEVICE SERVICED BY A THIRD PARTY?: PREVIOUSLY REPORTED AS NO ; UPDATED TO UNKNOWN. E1 INITIAL REPORTER CONTACT / ADDRESS / PHONE NUMBERS. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03330 | 2021-08-17 | 2019-06-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP PARKINSON'S. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY MISSED INFORMATION IN SECTION B5 AND CAPTURED ALLEGATION ONLY AS PARKINSON'S INCREASED METABOLISM. AFTER REVIEW, IT WAS DETERMINED THAT SECTION B5 SHOULD BE FILED AS FOLLOWS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARKINSON'S INCREASED METABOLISM RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION NEED TO CAPTURE ABOUT THE ALLEGED ARTHRITIS. SECTION H6 HEALTH EFFECTS: CLINICAL CODE HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03294 | 2021-08-17 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK PARTICLES IN HER BED AND LUNG ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. AFTER THE FINAL ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, CUSTOMER DECLINED TO RESPOND TO THE GFE RELATED QUESTIONS AND TERMINATED THE CALL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. IN INITIAL REPORT, SECTTION H6 CLINICAL CODE WAS WROGLY CAPTURED, CORRECTED AND UPDATED IN THIS REPORT. SECTION G3 AND H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03291 | 2021-08-17 | 2023-03-24 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX900H11C | DSX900H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON AUG 13, 2024, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING WHEN WAKE UP IN THE MORNING BLACK PARTICLES ON IS EYELASHES/EYEBROWS AND INSIDE NOSTRILS AND HAVE TO PULL APART, WATER HASN'T BEEN GOING DOWN DOES NOT DRAIN NO MIST, FLUID IN LUNGS, PARTICLES ARE FLOATING IN WATER AND PARTICLES IN AIRWAY RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT. ADDITIONAL INFORMATION RECEIVED ABOUT INSOMNIA AND CHRONIC COUGH. SECTION H6 HEALTH EFFECTS: CLINICAL CODES WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03293 | 2021-08-17 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03292 | 2021-08-17 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DIFFICULTY BREATHING/SHORT OF BREATH, STRANGE ODOR AND DEVICE IS RELEASING PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. SECTION H6 HEALTH EFFECT - CLINICAL CODE IS UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03290 | 2021-08-17 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING IRRITATED THROAT, IRRITATED EYES, DRY EYES, HEADACHES, DIZZINESS, HYPERSENSITIVITY, COUGH, ECZEMA AND SHORTNESS OF BREATH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING IRRITATED THROAT, IRRITATED EYES, DRY EYES, HEADACHES, DIZZINESS, HYPERSENSITIVITY, COUGH, ECZEMA AND SHORTNESS OF BREATH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION G1-MANUFACTURER CONTACT OFFICE DETAILS WAS CORRECTED AND UPDATED. SECTION H6 HAS BEEN CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03341 | 2021-08-17 | 2021-06-14 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA. THE PATIENT REQUIRED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY MISSED INFORMATION IN SECTION B5 AND CAPTIRED ONLY AS ASTHMA ALLEGATION. AFTER REVIEW, IT WAS DETERMINED THAT SECTION B5 SHOULD BE FILED AS FOLLOWS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING ASTHMA RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION NEED TO CAPTURE ABOUT THE ALLEGED DYSPNEA, HEADACHE, COUGH AND FATIGUE. SECTION H6 HEALTH EFFECTS: CLINICAL CODES HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03342 | 2021-08-17 | 2021-07-26 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA. THE PATIENT REQUIRED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. CORRECTED DATA: D4 UNIQUE IDENTIFIER (UDI) NUMBER: PREVIOUSLY REPORTED AS (B)(4); UPDATED TO (B)(4) E1 INITIAL REPORTER CONTACT / ADDRESS / PHONE NUMBERS. E4 INITIAL REPORTER ALSO SENT REPORT TO FDA?: PREVIOUSLY REPORTED AS NO ; UPDATED TO UNKNOWN. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03302 | 2021-08-17 | 2021-08-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03264 | 2021-08-17 | 2021-07-02 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC. | DSX500S11 | DSX500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION B5 SHOULD BE CORRECTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP,AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING HAS SOMETHING IN HIS LUNGS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM.HE DID NOT PASS EXAM FROM HIS PHYSICIAN. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03296 | 2021-08-17 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THERE WAS NO PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF BLACK PARTICLES IN THE WATER TANK, PARTICLES IN TUBING/CHAMBER RELATED TO THE CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE TWO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03298 | 2021-08-17 | 2020-05-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED BEING TO THE LUNG DOCTOR,VISUALIZATION OF BLACK PARTICLES IN THE DEVICE AND AIRWAY. THE PATIENT HAD A CHEST X RAY, SPIT TEST AND WAS PRESCRIBED PREDNISONE AND BREATHING TREATMENTS. PATIENT ALLEGED THAT THE DEVICE LEAKED WATER, ALARMED AND FLASHED AND IS NO LONGER FUNCTIONING. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENT OF BEING TO THE LUNG DOCTOR,VISUALIZATION OF BLACK PARTICLES IN THE DEVICE AND AIRWAY. THE PATIENT HAD A CHEST X RAY, SPIT TEST AND WAS PRESCRIBED PREDNISONE AND BREATHING TREATMENTS. PATIENT ALLEGED THAT THE DEVICE LEAKED WATER, ALARMED AND FLASHED AND IS NO LONGER FUNCTIONING. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION.IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03266 | 2021-08-17 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. CORRECTED INFORMATION WAS PROVIDED IN H9. SECTION (PREVIOUSLY RES 88058 WAS NOT INCLUDED). H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03257 | 2021-08-17 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING/CHAMBER, THROAT SORENES AND HIS EAR IS HURTING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03265 | 2021-08-17 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03262 | 2021-08-17 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11C | DSX400H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03263 | 2021-08-17 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN PREVIOUS REPORT SECTION B5 WAS MENTIONED INCORRECTLY, CORRECT B5 SHOULD BE - THE PATIENT ALLEGED VISUALIZATION OF PARTICLES, HEADACHES, LIGHTHEADEDNESS, DIZZY, NASAL IRRITATION, SORE THROAT, COUGH. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL AND INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND DUST ON/IN THE BLOWER. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND 6 ERRORS. THE MANUFACTURER CONCLUDES CONTAMINATION FOUND WERE CONSISTENT WITH DUST ON/IN THE BLOWER. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. IN THIS REPORT, SECTION D9, G3, H3, H6 HAS BEEN UPDATED OR CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03259 | 2021-08-17 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. ADDITIONAL INFORMATION WAS RECEIVED, SECTION B5 WAS CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK DEBRIS PARTICLES IN THE TUBE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03149 | 2021-08-17 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03151 | 2021-08-17 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03150 | 2021-08-17 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03148 | 2021-08-17 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO AUTO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSG500S11 | DSG500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03297 | 2021-08-17 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | SECTION H6: HEALTH EFFECTS - CLINICAL CODE WAS MISSED IN THE PREVIOUS REPORT; IT HAS BEEN CORRECTED OR UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03318 | 2021-08-17 | 2021-07-19 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT DIFFICULTY IN BREATHING. THE PATIENT REPORTEDLY GOT HOSPITALIZED IN RESPONSE TO THE EVENT. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION FOR INHALER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A PATIENT REPORTED DIFFICULTY IN BREATHING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT REPORTEDLY GOT HOSPITALIZED IN RESPONSE TO THE EVENT. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION FOR INHALER. THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 HAS BEEN UPDATED. SECTION D4, UNIQUE IDENTIFIER (UDI) WAS INCORRECTLY CAPTURED IN PREVIOUS REPORT, IT HAS BEEN CORRECTED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03321 | 2021-08-17 | 2021-07-19 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA AND BRONCHITIS. THE PATIENT WAS PRESCRIBED PREDNISOLONE IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. CORRECTED DATA: E4 INITIAL REPORTER ALSO SENT REPORT TO FDA?: PREVIOUSLY REPORTED AS NO ; UPDATED TO UNKNOWN. G3 DATE RECEIVED BY MANUFACTURER: PREVIOUSLY REPORTED AS 07/22/2021 ; UPDATED TO 07/19/2021. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03322 | 2021-08-17 | 2021-07-19 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. B(2) WAS CORRECTED TO REFLECT 'OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS". THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03299 | 2021-08-17 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DEVICE WILL NOT TURN ON/DEVICE NOT FUNCTIONING, PARTICLES IN AIRWAY FROM DEVICE, BLACK PARTICLES INSIDE THE HOSE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. AFTER THE SECOND ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, CUSTOMER RESPONDED THAT I HAVE DISCARDED THE OLD DEVICE AND I AM UNABLE TO RETURN IT. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03286 | 2021-08-17 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX900T11C | DSX900T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTRE FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF DEVICE WAS INSPECTED. THE DEVICE POWERED ON, AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03295 | 2021-08-17 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500S11 | DSX500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED NASAL/THROAT IRRITATION. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE 1 ERROR FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03319 | 2021-08-17 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT SHORTNESS OF BREATH. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION FOR INHALER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON (B)(6) 2024 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT SHORTNESS OF BREATH. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION FOR INHALER SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTHE ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECTS: CLINICAL CODES AND HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103347 | 2021-08-18 |  | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500H11C |  |  | I HAVE DEVELOPED PERSISTENT, UNEXPLAINED COUGHING EPISODES EVER SINCE I STARTED USING MY PHILIPS RESPIRONICS CPAP MACHINE IN (B)(6) OF 2019 OR SHORTLY THEREAFTER. I LIVED WITH IT NOT KNOWING WHAT THE POSSIBLE CAUSE MIGHT BE. I FINALLY WENT TO A PRIMARY CARE PHYSICIAN, DR. (B)(6), ON (B)(6) 2021 TO COMPLAIN ABOUT MY COUGHING AND SHE THOUGHT I MIGHT HAVE DEVELOPED SOME SORT OF ASTHMA AND PRESCRIBED ADVAIR INHALER TO USE TWICE A DAY. I HAVE NEVER HAD ANY RESPIRATORY ISSUES MY ENTIRE LIFE. I FEEL AS THOUGH I HAVE BECOME DEPENDENT ON THIS INHALER. THIS WEEK I HAVE JUST RECEIVED A RECALL NOTICE FOR MY CPAP MACHINE. ONE OF THE EFFECTS OF THE DEFECTIVE MACHINE INCLUDE RISK OF CHEMICAL EXPOSURE DUE TO OFF-GASSING THAT CAN AFFECT THE RESPIRATORY TRACT AND CAUSE COUGHING AND SINUS ISSUES. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-08
MW5103343 | 2021-08-18 | 2019-08-28 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS BIPAP DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | J253475608700 | DSX700H11C |  | INITIALLY GOT A CPAP IN (B)(6) OF 2019. I EXPERIENCED A LOT OF LIGHTHEADEDNESS, DISORDERED BREATHING, AND ANXIETY IN THE DAYTIME AFTER USING IT. I HAD A REPEAT SLEEP STUDY AND WAS CHANGED TO BIPAP IN (B)(6) 2019. I EXPERIENCED SYMPTOMS SUCH AS HEADACHES, LIGHTHEADEDNESS, DISORDERED BREATHING, COUGH (WHILE USING THE DEVICE), NAUSEA ALMOST EVERY MORNING, FATIGUE, BRAIN FOG AND ANXIETY RELATED TO THESE SYMPTOMS. I TOLD DOCTORS THAT I FELT LIKE I WAS ON A DRUG ALTHOUGH HADN'T TAKEN ANY. I WENT THROUGH EXTENSIVE TESTING. I WAS OFF WORK IN 2020 FOR ABOUT 10 WEEKS AND EVENTUALLY ACCEPTED AN EARLY RETIREMENT (B)(6) OF 2021. SYMPTOMS HAVE BEEN SUCH THAT I HAD TO HAVE GROCERIES DELIVERED AT TIMES AND HAVE LIMITED THE DRIVING OF LONG DISTANCES. I SAW THE BLACK PARTICLES IN THE WATER CHAMBER AT TIMES, BUT HAD NO IDEA THESE WERE DANGEROUS. ADDITIONALLY, IN (B)(6) OF 2020 I PURCHASED A SO CLEAN (OZONE) DEVICE TO SANITIZE THE BIPAP. I WAS NEVER TOLD THAT THIS WAS DANGEROUS OR SHOULD NOT BE USED WITH THE BIPAP. WHILE SOME OF MY SYMPTOMS WERE PRE-EXISTING, THEY WERE NOT AS SIGNIFICANT. I BELIEVE MY SYMPTOMS WERE AGGRAVATED BY THE DEVICE. THERE IS CURRENTLY NOT A REPLACEMENT DEVICE AND MAY NOT BE ONE FOR A YEAR OR MORE, ACCORDING TO WHAT I HAVE BEEN TOLD. I UNDERWENT MULTIPLE TESTS THROUGHOUT 2019, 2020, AND 2021. I HAD MRI OF BRAIN WITH AND WITHOUT CONTRAST; EEG; HEARING EVALUATION AND VNG; MULTIPLE CONSULT WITH EXPERTS IN NEUROTOLOGY, NEURO-OPHTHALMOLOGY, HEADACHES, ETC. I HAD BLOODWORK AND OT/PT. I NEED TO REST A LOT DUE TO THE FATIGUE AND THE BRAIN FOG. I AM NOT GOING TO LIST EVERY SINGLE TEST HERE UNLESS YOU CALL ME AND WANT THE INFORMATION. AFTER EVERYTHING I'VE BEEN THROUGH I CAN'T SPEND SEVERAL MORE HOURS LOOKING UP DATES AND LISTING EVERYTHING. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103342 | 2021-08-18 | 2019-01-01 | I, n, j, u, r, y | Voluntary report | SLEEP APNEA MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. | DSX500T11C |  |  | I CAN NOT GIVE YOU AN EXACT DATE BUT FOR A FEW YEARS NOW I KEEP GETTING UPPER RESPIRATORY INFECTIONS WITH MY MACHINE AND SOLELY CAUSED FROM IT. I GET BAD ASTHMA WHICH I NEVER HAD UNTIL I USED THE MACHINE. I THOUGHT IT WAS ODD. NOW MY SHORT TERM MEMORY IS SHOT LIKE EMBARRASSINGLY SHOT. I HAVE BEEN TO EVERY DOCTOR POSSIBLE. I HAD ZERO IDEA UNTIL TODAY THAT THERE IS POSSIBLY A TOXIC CHEMICAL ISSUE WITH THE MACHINE. MY SLEEP DOCTOR TOLD ME ABOUT IT. I HAVE BEEN TO SO MANY DOCTORS TRYING TO FIGURE OUT WHAT MY ISSUE IS. I HAVE INFLAMMATION ISSUES TOO AND NO ONE CAN FIGURE ANYTHING OUT. I AM SO UPSET RIGHT NOW TO THINK THIS COULD BE MY SLEEP MACHINE. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103341 | 2021-08-18 | 2021-08-02 | I, n, j, u, r, y | Voluntary report | DREAMSTATION ASV | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500H11C |  |  | LONGER BRONCHITIS; BRONCHITIS SET IN AND STARTED MEDS. HOWEVER, FOUND BLACK DUSTLIKE PARTICLES BUILD UP IN THE HUMIDIFIER OF CPAP. BRONCHITIS IS GOING ON 3 WEEKS, WHERE IN PAST ONLY LAST 1-1 1/2 WEEK TOTAL. HAD TO GET EXTRA MEDICINES TO TRY TO FIGHT IT. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103332 | 2021-08-18 | 2020-01-14 | D, e, a, t, h | Voluntary report | PHILLIPS RESPIRONICS DREAM STATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DREAM STATION |  |  | MY HUSBAND, (B)(6), DIED (B)(6) 2020 AND HE WAS USING PHILLIPS RESPIRONICS DREAM STAGE C-PAP EQUIPMENT WHEN HE DIED IN HIS SLEEP. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103333 | 2021-08-18 | 2020-11-20 | I, n, j, u, r, y | Voluntary report | PHILLIPS CPAP DREAM STATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I WAS AND AM USING A PHILLIPS CPAP DREAM STATION FIRST I DISCOVERED A LUMP IN MY LEFT BREAST THEN I DISCOVERED A LUMP IN MY NECK AND HAD A CAT SCAN DONE. THEY COULD NOT UNDERSTAND WHAT WAS HAPPENING TO MY NECK. THEN MY EYES BECAME SWOLLEN SHUT I WENT TO A VISION CENTER AND WAS TOLD I HAD A DRY EYE INFECTION. PROBLEMS STILL EXIST. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-08
2518422-2021-03325 | 2021-08-18 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S FLOW SENSOR WAS REPLACED DUE TO DUST AND DIRT CONTAMINATION. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03063 | 2021-08-18 | 2021-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A SERVICE REQUIRED ALARM RELATED TO A LEAK. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SENSOR BOARD AND FLOW SENSOR ASSEMBLY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03331 | 2021-08-18 | 2021-01-26 | D, e, a, t, h | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040000 | 1040000 |  | THE MANUFACTURER WAS MADE AWARE OF AN INCIDENT THAT OCCURRED ON (B)(6) 2021 IN WHICH A PATIENT WAS PLACED ON A VENTILATOR AND THE OXYGEN SUPPLY TUBE WAS INCORRECTLY CONNECTED TO THE OXYGEN SOURCE. THE DEVICE ALARMED FOR INSUFFICIENT OXYGEN FLOW AND LOW OXYGEN INLET PRESSURE AS DESIGNED. IT IS REPORTED THAT THE STAFF DID NOT ADDRESS THE ALARMS AND THE PATIENT DETERIORATED AND DIED. |  | Y | N | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03313 | 2021-08-18 | 2020-09-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A SORE THROAT AND SWOLLEN TONSIL. THE PATIENT WAS PRESCRIBED ANTIBIOTICS IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. CORRECTION: EVENT DATE IN B3 WAS INCORRECTLY REPORTED AS (B)(6) 2021. THE CORRECT EVENT DATE IS (B)(6) 2020. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03315 | 2021-08-18 | 2021-05-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A COUGH, HEADACHES, AND FATIGUE. THE PATIENT WAS PRESCRIBED STEROIDS AND ANTIBIOTICS IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED ON THIS DEVICE IN MDR 2518422-2021-02075. THIS REPORT WAS SUBMITTED AS A DUPLICATE REPORT OF THE PREVIOUSLY SUBMITTED REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03314 | 2021-08-18 | 2021-05-24 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A COUGH AND SHORTNESS OF BREATH. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED TO DEVELOP COUGH AND SHORTNESS OF BREATH. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THE REPORTED EVENTS OF COUGH, SHORTNESS OF BREATH AND THEIR REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURER'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURER CONCLUDES NO FURTHER ACTION IS NECESSARY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE EXTERNAL INVESTIGATION SHOWED THAT THERE WAS AN UNKNOWN CONTAMINATION ON THE BOTTOM ENCLOSURE. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THERE WAS ALSO AN UNKNOWN CONTAMINATION IN THE BOTTOM OF THE BLOWER BOX. THE MANUFACTURER FOUND EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE WAS HOOKED UP TO A KNOWN GOOD POWER SUPPLY AND POWER CORD, AIRFLOW WAS VERIFIED TO BE OPERATING CORRECTLY. THE ERROR LOG SHOWED ZERO INSTANCES OF ERRORS ON THE DEVICE. THE MANUFACTURER CONCLUDES THAT THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION, UNKNOWN CONTAMINATION ON THE BOTTOM ENCLOSURE AND BOTTOM OF THE BLOWER BOX. THE MANUFACTURER CONFIRMED THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION B4 AND H4 HAD MISSING INFORMATION WHICH IS UPDATED IN THIS REPORT. SECTION H6 HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03311 | 2021-08-18 | 2021-06-22 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500S11 | DSX500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A KIDNEY INFECTION AND BLOOD IN URINE. THE PATIENT WAS PRESCRIBED ANTIBIOTICS IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03312 | 2021-08-18 | 2018-03-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A LUNG INFECTION AND SHORTNESS OF BREATH. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO DEVELOP A LUNG INFECTION AND SHORTNESS OF BREATH.THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION.AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03164 | 2021-08-18 | 2021-08-11 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX700T11 | DSX700T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A HEARING PROBLEM. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103356 | 2021-08-18 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. |  | REF NO DXX 800 HIIC |  | CALLER HAS HAD THE DREAMSTATION FOR APPROXIMATELY SIX MONTHS, SHE RECEIVED A CALL FROM HER MEDICAL ADVISOR, WHO INFORMED HER THE DEVICE WAS CAUSING SIDE EFFECTS FOR PEOPLE AND HAS BEEN RECALLED. CALLER HAS SINCE STOPPED USING THE DEVICE WHICH HAS CAUSED HER HEALTH TO DECLINE AND DETERIORATE RAPIDLY. SHE CALLED THE MANUFACTURER TO INQUIRE ABOUT A REPLACEMENT AND WAS INFORMED SEVERAL DEVICES HAVE BEEN RECALLED AND THEY ARE WAITING FOR APPROVAL FOR THE DEVICES TO BE FIXED. |  | N | Y | I |  | 2021-08
2518422-2021-03064 | 2021-08-18 | 2021-06-29 | I, n, j, u, r, y | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED HEMATEMESIS. THE PATIENT WAS ADMITTED TO THE HOSPITAL. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED HEMATEMESIS. THE PATIENT WAS ADMITTED TO THE HOSPITAL. DEVICE HAS STRANGE ODOR, BURNING/SMOKE/ELECTRICAL ODOR, AND GAS SMELL. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. DURING THE INVESTIGATION, WE COULD NOT OBSERVE ANY BLACK PARTICULATE OR EVIDENCE INDICATIVE OF FOAM DEGRADATION INSIDE THE UNIT. AFTER THE UNIT OPERATION, THE UNIT WAS DISASSEMBLED FOR VISUAL OBSERVATION. THE UNIT AIRPATH ASSEMBLY FOAM AND UNIT AIRPATH FOAM, FLOW PATH INCLUDING OUTLET, AND BACTERIA FILTER LOOKED CLEAN. OBSERVED A FEW FOREIGN DUST PARTICLES OF EXTERNAL ORIGIN INSIDE THE BLOWER. DURING THE INVESTIGATION, THEY FOUND THE MONITOR PRESSURE SENSOR AND CONTROL PRESSURE SENSOR SENSORS MOUNTED ON THE SENSOR BOARD TO BE FAULTY. THE SENSOR BOARD HAS BEEN REPLACED. SECTION H1 & H4 WAS CAPTURED INCORRECTLY IN THE PREVIOUS REPORT. IT HAS BEEN UPDATED AND CORRECTED IN THIS REPORT. IN THIS REPORT, SECTIONS D9, G3, H3, AND H6 HAVE BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03326 | 2021-08-18 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2100X15B | IN2100X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A FAILURE OF THE DEVICE'S EXTERNAL FLOW SENSOR CABLE WAS OBSERVED. THE DEVICE'S EXTERNAL FLOW SENSOR CABLE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03327 | 2021-08-18 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2200X15B | IN2200X15B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE AND OXYGEN SENSOR WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03367 | 2021-08-19 | 2021-08-12 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, BL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | BLX500H15 | BLX500H15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP BLACK SPOTS ON THE LUNG(S). THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE REPORTED EVENT IS CURRENTLY UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO THE FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER (B)(4).. CORRECTED DATA: D8 WAS THIS DEVICE SERVICED BY A THIRD PARTY?: PREVIOUSLY REPORTED AS NO ; UPDATED TO UNKNOWN. E1 INITIAL REPORTER ADDRESS. E4 INITIAL REPORTER ALSO SENT REPORT TO FDA?: PREVIOUSLY REPORTED AS NO ; UPDATED TO UNKNOWN. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03348 | 2021-08-19 | 2021-08-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE VENTILATOR'S LCD SCREEN OBSERVED. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03370 | 2021-08-19 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION EXPERT W/HUM/HT, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX501T12 | CAX501T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02890 | 2021-08-19 | 2021-07-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON OCT 09, 2023, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED VISUALIZATION OF PARTICLES. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGED CANCER. SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY THE PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECTS: CLINICAL CODES AND HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02889 | 2021-08-19 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY MISSED TO CAPTURE INFORMATION IN SECTION B5. AFTER REVIEW, IT WAS DETERMINED THAT SECTION B5 SHOULD BE FILED AS FOLLOWS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING CHAMBER AND HEADACHE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. HEALTH EFFECTS: CLINICAL CODES HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02888 | 2021-08-19 | 2021-06-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES AND PARTICLES IN TUBING/CHEMBER AND HEADACHE. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE 0 ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND SCRAPPED. A CORRECTION TO B5 WAS MADE AND SHOULD BE THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES AND PARTICLES IN TUBING/CHEMBER AND HEADACHE CUSTOMER HAVING THE LUNG ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTION A1 WAS UPDATED IN THIS REPORT WHICH WAS MISSED IN THE INITIAL REPORT. SECTION H6 WERE CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03345 | 2021-08-19 | 2021-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S NAVIGATIONAL BUTTON WAS OBSERVED. THE DEVICE WAS POWER CYCLED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03267 | 2021-08-19 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING COUGHING,RESPIRATORY TRACT ISSUES,PARTICLES IN AIRWAY AND TUBING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03270 | 2021-08-19 | 2021-07-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03268 | 2021-08-19 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION.A VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER.THE MANUFACTURER FOUND A SMALL PIECE OF WHITE PLASTIC OUTSIDE OF THE AIRPATH (LAYING ON BOTTOM OF THE UNIT ENCLOSURE).MATERIAL ORIGINATES EXTERNAL OF THE ORIGINAL DESIGN. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERROR LOGGED. THE MANUFACTURER CONCLUDES NO EVIDENCE OF DEGRADED SOUND ABATEMENT FOAM. IN THE INITIAL REPORT SECTION B5 WAS MENTIONED INCOMPLETELY WHICH SHOULD HAVE BEEN- THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM.THE PATIENT ALLEGED VISUALIZATION OF PARTICLES. THE PATIENT ALSO ALLEGED SYMPTOMS OF HEADACHES, DROUSYNESS AND CONGESTION. SECTION D9, G3, H6 HAS BEEN UPDATED / CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03269 | 2021-08-19 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX900T11C | DSX900T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02951 | 2021-08-19 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO SERIES ASSY, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | 750P | 750P |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03343 | 2021-08-19 | 2021-03-01 | I, n, j, u, r, y | Manufacturer report | BIPAP PRO, INTL | NON CONTINUOUS VENTILATOR | RESPIRONICS, INC. | 661P | 661P |  | MANUFACTURER BECAME AWARE OF AN ALLEGATION OF A USER WHILE ON A BI-LEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE HEARD CLICKING IN HIS TUBING AND A PIECE FROM THE MACHINE WENT INTO HIS THROAT AND INTO HIS LUNGS. USER DEVELOPED COUGH AND FEVER, WENT TO HIS DOCTOR WHERE A CT SCAN WAS PERFORMED. USER WAS HOSPITALIZED WITH PNEUMONIA, PLEURAL EFFUSION AND PERICARDIAL EFFUSION FOR 10 DAYS AND TREATED WITH ANTIBIOTICS. THE RADIOLOGICAL IMAGES DID NOT DETECT A FOREIGN BODY IN HIS LUNGS. USER WAS DISCHARGED HOME WITH EXISTING SYMPTOMS OF PERSISTENT COUGH, SHORTNESS OF BREATH AND TIREDNESS FROM STRENUOUS WORK. DEVICE IS NO LONGER IN USE AND HAS BEEN SENT FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED THE DEVICE WITH HUMIDIFIER FOR INVESTIGATION. THE MANUFACTURER WAS UNABLE TO CONFIRM THE COMPLAINT OF "PIECE FROM MACHINE WENT INTO THROAT". THE MANUFACTURER DOES CONFIRM THERE IS EVIDENCE OF PHYSICAL DAMAGE TO THE DEVICE EXTERIOR (DISPLAY SCREEN IS CRACKED AND THE POWER PUSH WHEEL IS BROKEN), DIRT/DEBRIS THROUGHOUT THE DEVICE, AND LIQUID/WATER INGRESS CAUSING CORROSION THROUGHOUT THE DEVICES, INCLUDING ON THE HUMIDIFIER LID SEAL, DRY BOX SEALS, AND ON AS WELL AS UNDER THE HEATER PLATE. THE DEVICES DO POWER UP AND PROVIDE FLOW. THERE WERE NO MISSING COMPONENTS INTERNALLY, NOR WERE THERE ANY SOLID OBJECTS FOUND IN THE BLOWER MOTOR ASSEMBLY. THE BUILD UP OF CORROSION, DIRT AND DEBRIS IN THE BLOWER MOTOR ASSEMBLY AND PHYSICAL DAMAGE TO THE DISPLAY SCREEN AND PUSH WHEEL SUGGEST POOR MAINTENANCE BY THE USER OF THE DEVICES AND NOT A FAILURE OF THE DEVICES UNDER NORMAL OPERATION. THE MANUFACTURER CONCLUDES THAT NO FURTHER ACTION IS NECESSARY AT THIS TIME. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03346 | 2021-08-19 | 2021-08-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03347 | 2021-08-19 | 2021-08-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE AND OXYGEN BLENDING MODULE BOARD WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02887 | 2021-08-19 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500S11 | DSX500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED RUBBER PARTICLES WERE IN THE CHAMBER AND HAS HEADACHES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION. CORRECTION INFORMATION WAS PROVIDED IN A1. SECTION H6 HEALTH CLINICAL CODE HAS BEEN CORRECTED AND TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103372 | 2021-08-19 | 2021-07-16 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I HAVE BEEN A USER OF THE PHILIPS DREAMSTATION CPAP MACHINE SINCE THE SPRING OF 2017. IN THE LAST COUPLE OF YEARS I HAD NOTICED SOME BLACK RESIDUE IN THE TUBING OF MY CPAP AND AROUND THE PART OF THE DEVICE WHERE I CONNECT THE TUBING TO THE ACTUAL MACHINE. AS A NOTE, I ALSO HAVE LIVED IN PRETTY HUMID ENVIRONMENTS, INCLUDING (B)(6). FOR THE LAST YEAR OR SO, I HAVE NOTICED INCREASED HEADACHES UPON WAKING UP IN THE MORNINGS. ADDITIONALLY, I STARTED TO FEEL TIGHTNESS IN MY CHEST IN THE MORNING, AND LATELY, I HAVE SEEN AN INCREASE IN THE NUMBER OF SINUS INFECTIONS AND RESPIRATORY ISSUES AFTER WAKING UP FROM SLEEP. AFTER COMING ACROSS THE RECALL NOTICE IN THE PAGES OF THE WALL STREET JOURNAL THIS WEEK, I HAVE NO DOUBT THAT MANY OF THESE PROBLEMS CAN BE DIRECTLY ATTRIBUTED TO THE RECALLED PHILIPS DREAMSTATION CPAP MACHINE WHICH I HAVE USED EVERY NIGHT FOR THE LAST THREE YEARS. THIS MACHINE THAT WAS MEANT TO PRESERVE MY HEALTH AND PREVENT SEVERE ILLNESS MAY NOW BE CONTRIBUTING TO HEALTH PROBLEMS AND POTENTIAL SEVERE ILLNESS. I ONLY HOPE NOTHING WORSE COMES OF THIS. IN THE MEANTIME, MY DOCTOR HAS RECOMMENDED I CONTINUE TO USE THE CPAP MACHINE SINCE SHE BELIEVES THE BENEFITS OUTWEIGH THE RISKS. THIS HAS TRULY BEEN A NIGHTMARE SCENARIO FOR ME WHERE I AM ASKED TO WEIGH THE COST OF POTENTIAL LONG-TERM HEALTH DAMAGE FOR USING THE MACHINE VERSUS NOT USING THE MACHINE. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103376 | 2021-08-19 | 2019-01-01 | I, n, j, u, r, y | Voluntary report | PHILIPS / RESPIRONICS, INC. | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500H11 |  |  | FROM (B)(6) 2019 I HAD SEVERE S.O.B., REQUIRING INHALERS, NEBULIZERS AND MULTIPLE MEDS INCLUDING ABXS AND STEROIDS. I HAD SEVERE LETHARGY, SKIN IRRITATION, SORE THROAT, AND HEAT IN MY FACE. I DISCOVERED THAT THERE IS A RECALL ON MY CPAP MACHINE. I USED CPAP MACHINE - THE DREAM STATION FROM PHILIPS RESPIRONICS FOR MORE THAN 5 YEARS. ON (B)(6) 2021, I STOPPED THE USE OF THE CPAP MACHINE WITH SIGNIFICANT IMPROVEMENT IN MY HEALTH IN THIS LAST WEEK. |  | Y | N | LAY USER/PATIENT |  | 2021-08
MW5103368 | 2021-08-19 | 2021-07-15 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I HAVE BEEN FINDING THE FLOATING MATERIAL IN MY PHILIP DREAMSTATION OFF AND ON, BUT WASN'T SURE WHAT IT WAS AND WOULD CLEAN IT OUT AND USE AGAIN, BUT NOW WITH THE RECALL I HAVE BEEN AFRAID TO USE IT AND WONDERING NOW IF IT IS THE CAUSE OF ALL MY HEADACHES/SINUS PROBLEMS. NONETHELESS I AM LEFT WITHOUT A SAFE MACHINE TO USE DURING THIS RECALL. I JUST PURCHASED MINE AGAIN THIS YEAR, MY THIRD ONE AND TWO OF THEM ON RECALL, MY FIRST WAS NOT BUT IT DOES NOT WORK. I HAVE CHRONIC SEVERE OBSTRUCTIVE SLEEP APNEA, MY DOCTOR HAS TOLD ME IN THE SLEEP STUDY THAT I TURN BLUE DURING THE NIGHT FROM OXYGEN DEPRIVATION AND STOP BREATHING MULTIPLE TIMES ON AND OFF, SO NEEDLESS TO SAY, I NEED A MACHINE. I TALKED WITH PHILIPS RESPIRONICS AGAIN YESTERDAY AND THEY RELAYED TO ME THEIR APOLOGIES AND HAD NO ANSWER AS TO WHEN, JUST THAT THEY CANNOT DO ANYTHING UNTIL FDA APPROVES IT. PLEASE HELP US, THERE ARE SO MANY PEOPLE HER WAITING IN LIMBO AND AFRAID TO USE THE MACHINE AND SUFFERING OR USING IT AND RISKING THE CANCER OR WHATEVER IT MAY CAUSE, THIS IS NOT FAIR TO US AND WE NEED HELP. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | I |  | 2021-08
MW5103366 | 2021-08-19 |  | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | REMSTAR PRO C-FLEX + | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | 460BPT |  |  | PHILIPS RESPIRONICS REMSTAR PRO C-FLEX + SYSTEM ONE BOUGHT THROUGH (B)(6). THIS IS REGARDING THE PHILIPS RECALL ON CPAP DEVICES. FDA SAFETY REPORT ID# (B)(4) |  | N | Y | I |  | 2021-08
MW5103375 | 2021-08-19 | 2020-12-15 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPERONICS DREAM STATION C-PAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I HAVE A PHILIPS RESPIRONICS DREAMSTATION CPAP MACHINE. IT BEGAN HAVING BLACK PARTICLES IN THE WATER CHAMBER, AND I BEGAN HAVING A COUGH, 'RUNNY EYES,' AND SORE THROAT. LEARNED IT WAS RECALLED ON 06/15/2021. I CANNOT SLEEP WITHOUT THE CPAP, SO I AM GOING TO MY SLEEP DR ON (B)(6) 2021 TO SEE IF I CAN GET ANOTHER MACHINE AND WHAT HARM THIS HAS DONE TO MY BODY. |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103365 | 2021-08-19 | 2021-06-15 | I, n, j, u, r, y | Voluntary report | DREAM STATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | PHILIPS DREAM CPAP INITIALLY I NOTICED SEVERE CONGESTION IN THE MORNING. I SPOKE WITH MY ALLERGIST WHO REMINDED ME TO USE MY COOL MIST HUMIDIFIER. I THEN NOTICED THAT I WAS BRINGING UP THICK WHITE MUCOUS AND EXPERIENCED HEADACHES UPON WAKING. SOME DAYS MY MUCOUS HAD BLACK DOTS IN IT. I NEVER ASSOCIATED THIS WITH MY CPAP. I NOW BELIEVE THAT THE SYMPTOMS WERE FROM THE CPAP BECAUSE UPON DISCONTINUED USE OF THE MACHINE THE HEADACHES AND MUCOUS DISCHARGE BEGAN TO LESSEN. IT IS MY UNDERSTANDING THAT A FIX WAS SUBMITTED TO THE FDA BY PHILLIPS IN JUNE YET NOTHING HAS BEEN DONE. PATIENTS NEED THEIR MACHINES. THIS IS A MATTER OF LIFE AND DEATH. PLEASE APPROVE A FIX IMMEDIATELY. MY HEALTH IS SUFFERING WITHOUT THE CPAP. I CAN BE REACHED AT (B)(6) PLEASE DO SOMETHING. YOU ARE SUPPOSED TO PROTECT US. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103364 | 2021-08-19 | 2018-03-16 | I, n, j, u, r, y | Voluntary report | DREAMSTATION S/T | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I BEGAN USING THE PHILIPS RESPIRONICS DREAM STATION S/T FOR MY DAUGHTER IN (B)(6) 2017. SHE HAD HER FIRST SEIZURE IN (B)(6) 2018 WHILE SLEEPING ON THE CONTINUOUS VENTILATORY SUPPORT MACHINE MENTIONED ABOVE. THE SEIZURES HAVE INCREASED IN NUMBER SINCE THEN AND SHE HAS AN AVERAGE OF 30 OR MORE SEIZURES PER MONTH, SOMETIMES IN CLUSTERS OF UP TO 12 OR 13 SEIZURES IN A 24 HOUR PERIOD. THIS HAS SEVERELY ALTERED HER QUALITY OF LIFE. SHE HAS BEEN PUT ON VARIOUS ANTI-SEIZURE MEDICATIONS, YET THE SEIZURES HAVE CONTINUED. MANY OF THE SEIZURES HAVE BEEN LIFE-THREATENING AS SHE HAS STOPPED BREATHING DURING SEIZURES WITH OXYGEN SATURATION LEVELS DROPPING AS LOW AS 12%. WE RECEIVED NOTIFICATION OF THE URGENT, LIFE-THREATENING, CLASS 1 RECALL YESTERDAY EVENING VIA (B)(6) ONLY TO LEARN THIS MORNING THAT PHILIPS RESPIRONICS KNEW ABOUT THIS PROBLEM MUCH EARLIER, AND BEGAN SENDING LETTERS OUT ON JUNE 12, 2021. WHEN I FOLLOWED ALL OF THE STEPS PHILIPS RESPIRONICS OUTLINES, I WAS GIVEN NO DATE WHEN CORRECTIONS WOULD BE MADE AND WAS NOT OFFERED A SAFE REPLACEMENT MACHINE. UNLESS A MIRACLE OCCURS IN THE NEXT FEW HOURS, I WILL BE FORCED TO HOOK MY DAUGHTER UP TO THIS BIPAP WHICH, RATHER THAN BEING LIFESAVING, NOW SOUNDS LIKE THE "MACHINE OF DEATH" BASED ON PHILIPS RESPIRONICS LIST OF THE DANGERS INVOLVED. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03344 | 2021-08-19 | 2021-08-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | RIN2110X15B | RIN2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VARIATION IN THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE DELIVERY AND POSITIVE INSPIRATORY PRESSURE WAS OBSERVED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VARIATION IN THE VENTILATOR'S POSITIVE END EXPIRATORY AND POSITIVE INSPIRATORY PRESSURE DELIVERY WAS OBSERVED. ADDITIONAL INFORMATION RECEIVED INDICATES THE VENTILATOR'S SOFTWARE WAS RE-LOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03349 | 2021-08-19 | 2021-08-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03335 | 2021-08-20 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLANK. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S LCD SCREEN NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER HAD PREVIOUSLY REPORTED THAT THE LCD SCREEN WAS REPLACED TO ADDRESS A BLANK LCD SCREEN. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE'S BUTTONS WOULD NOT RESPOND TO BUTTON PUSHES. THE DEVICE'S KEYPAD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
MW5103408 | 2021-08-20 | 2019-02-15 | I, n, j, u, r, y | Voluntary report | PHILIPS CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS, INC. |  |  |  | I WAS DIAGNOSED WITH GASTRIC LYMPHOMA IN (B)(6) 2019 AND UNDERWENT RADIATION THERAPY. MY ONCOLOGISTS PERIODICALLY EXAMINE ME FOR RECURRENCE AND WILL CONTINUE TO DO SO. IN (B)(6) 2021, I WAS DIAGNOSED WITH MUCOSIS FUNGEIDES WHICH IS A RARE LYMPHOMA OF THE CIRCULATORY SYSTEM WHICH EFFECTS THE SKIN AND I WILL MEET WITH AN EXPERT IN THAT FIELD ON (B)(6) TO DETERMINE TREATMENT. I AM UNAWARE OF THE LIKELIHOOD OF LINKING THE CANCER TO THE PHILIPS CPAP MACHINE, BUT I WANT THE FDA TO BE AWARE OF THE PROBLEM. THANKS. (B)(6). FDA SAFETY REPORT ID # (B)(4) |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03355 | 2021-08-20 | 2021-08-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO AUTO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSG500S11 | DSG500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03360 | 2021-08-20 | 2021-08-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03354 | 2021-08-20 | 2021-08-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | INX500S15 | INX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03350 | 2021-08-20 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED DIZZIY, DIFFICULTY BREATHING, LIGHTHEADED, HEADACHES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO CONTAMINATION. AN INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND SIGNS OF WATER INGRESS IN THE BLOWER BOX AND CONTAMINATION IN CASE. THE MANUFACTURER FOUND EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND ZERO ERRORS. THE DEVICE WAS APPLIED POWER AND THE DEVICE OPERATED PROPERLY. THE MANUFACTURER CONCLUDES CONTAMINATION FOUND WERE CONSISTENT WITH WATER INGRESS IN THE BLOWER BOX AND CONTAMINANTS IN CASE. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION IN THE DEVICE. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03339 | 2021-08-20 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S BLOWER MOTOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03338 | 2021-08-20 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03334 | 2021-08-20 | 2021-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SOFTWARE WAS RELOADED AND DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03332 | 2021-08-20 | 2021-07-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS DAMAGED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
MW5103407 | 2021-08-20 | 2021-08-12 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | LIVER AND KIDNEY IMPAIRMENT; AFTER LEARNING THE CPAP MACHINE I HAD USED FOR 5 YEARS WAS RECALLED, I WENT IN FOR A CHEST X-RAY, A LIVER AND KIDNEY FUNCTION TEST. I HAVE SCAR TISSUE ON MY LUNGS AND MY LIVER AND KIDNEYS ARE NOT WORKING AT FULL CAPACITY. I SEE MY PCM (B)(6) TO ADDRESS THESE ISSUES. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103383 | 2021-08-20 | 2016-06-21 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAM STATION CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | WITHIN A FEW MONTHS OF USING MY DREAM STATION CPAP MACHINE, I BECAME SHORT OF BREATH, MY CHEST HURT AND I HAD HEART PALPITATIONS. NO ONE COULD FIGURE OUT WHY MY OXYGEN LEVEL DROPPED WHENEVER I EXERTED MYSELF. I HAVE MILD ASTHMA, BUT NEVER HAD BREATHING ISSUES LIKE THIS. MY CHEST HURT SO BAD, I WAS HOSPITALIZED. I FOLLOWED UP WITH 4 CARDIOLOGISTS AND A PULMONOLOGIST FOR THE PAST 5 YEARS, BUT CONTINUE TO DECLINE. I HAVE MADE 6 TRIPS TO THE EMERGENCY ROOM BETWEEN 2016 AND 2021 FOR THIS ISSUE AS WELL. AT THIS POINT, I AM HOMEBOUND. I BEGAN COUGHING UP GREY/BLACK MUCUS ABOUT 4-6 MONTHS AGO AND MY DOCTOR FEELS THE COUGHING MAY HAVE CAUSED COSTOCHONDRITIS. I'M IN CONSTANT PAIN AND DON'T KNOW WHAT TO DO. I HAVE SUCH A HARD TIME BREATHING. NONE OF MY DOCTORS SEEM TO HAVE ANY INFORMATION ABOUT WHETHER THIS MAY HAVE BEEN RELATED TO THE CPAP MACHINE. THANK YOU FOR YOUR HELP. MY LAB RESULTS ARE ALL NORMAL WITH THE EXCEPTION OF ENLARGED LYMPH NODES A FEW YEARS AGO. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103381 | 2021-08-20 | 2021-07-01 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS DREAMSTATION ASV CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX900T11C |  |  | PHILIPS RESPIRONICS HAS RECALLED SEVERAL MILLION CPAP DEVICES IN AN ANNOUNCEMENT DATED JUNE 14, 2021. APPARENTLY THE FDA IS AWARE OF THIS RECALL. I HAVE BEEN A USER OF ONE OF THESE RECALLED DEVICES FOR OVER FOUR YEARS. THAT SPECIFIC DEVICE IS A DREAMSTATION ASV CPAP. I BECAME AWARE OF THIS RECALL AND THE ASSOCIATED WARNING "STOP USING THIS DEVICE IMMEDIATELY" THROUGH MY OWN READING. I HAVE YET TO RECEIVE ANY NOTICE FROM PHILIPS IN SPITE OF THE FACTS THAT CONTINUED USE OF THE DEVICE IS LIFE THREATENING, AND PHILIPS KNOWS THAT I AM A REGISTERED USER AND A DAILY USER OF THIS DEVICE. I MUST ASSUME THAT PHILIPS HAS NOT NOTIFIED OTHER USERS ALSO. I AM REGISTERED ON THE PHILIPS REGISTRATION WEBSITE. I AM REGISTERED AS A SPECIFIC USER OF THIS DEVICE AND MY MACHINE REPORTS USAGE TO PHILIPS DAILY. PHILIPS HAS YET TO CONTACT ME AND HAS NOT SUBMITTED ANY PLAN TO REPAIR OR REPLACE THESE DEVICES. FURTHER, I HAVE NOW DISASSEMBLED MY SPECIFIC DEVICE AND FOUND EVIDENCE OF THE SPECIFIC CONTAMINATION THAT CAUSED THE RECALL. THIS IS IN SPITE OF THE FACT THAT NONE OF THE FACTORS THAT PHILIPS SAYS MAY CONTRIBUTE TO THE CONTAMINATION ARE PRESENT IN MY SITUATION. I BELIEVE THAT PHILIPS IS SEEKING TO MISLEAD THE FDA AND END-USERS ABOUT THE SERIOUSNESS OF THIS RECALL. I HAVE BEEN BREATHING CARCINOGENIC PARTICLES (AND PROBABLY VAPORS) FROM THIS MACHINE FOR YEARS. I BELIEVE THE FDA SHOULD CAUSE PHILIPS TO IMMEDIATELY NOTIFY END-USERS AND APPROPRIATELY REPLACE THESE MACHINES OR COMPENSATE USERS FOR SUCH REPLACEMENT. IT IS CLEARLY UNNECESSARY TO ALLOW SUCH A SITUATION TO CONTINUE FOR OVER 60 DAYS WITH NO NOTICE AND NO SOLUTIONS. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03353 | 2021-08-20 | 2020-08-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03065 | 2021-08-20 | 2021-07-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S PRESSURE AND RESPIRATORY RATE WERE FLUCTUATING AND THE PATIENT EXPERIENCED DIFFICULTY BREATHING AND HIS OXYGEN DROPPED. THE PATIENT WAS ADMITTED TO THE HOSPITAL. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT COULD NOT BE DUPLICATED. THE MANUFACTURER CONCLUDES THE VENTILATOR OPERATED AND ALARMED AS DESIGNED AND DID NOT CAUSE OR CONTRIBUTE TO THE REPORTED EVENT. | THE MANUFACUTRER PREVIOUALY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR'S PRESSURE AND RESPIRATORY RATE WERE FLUCTUATING AND THE PATIENT EXPERIENCED DIFFICULTY BREATHING AND HIS OXYGEN DROPPED. THE PATIENT WAS ADMITTED TO THE HOSPITAL. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUMSTOMER'S COMPLAINT COULD NOT BE DUPLICATED. THE MANUFACTURER CONCLUDES THE VENTILATOR OPERATED AND ALARMED AS DESIGNED AND DID NOT CAUSE OR CONTRIBUTE TO THE REPORTED EVENT. DURING ADDITIONAL TESTING IN THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY, THE DEVICE FAILED A TEST STEP DURING TESTING RELATED TO THE NURSE CALL CONNECTOR. THE DEVICE'S INTERFACE BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. THIS FAILURE WOULD NOT HAVE CONTRIBUTED TO THE REPORTED EVENT. | Y | Y | HEALTH PROFESSIONAL |  | 2021-08
MW5103409 | 2021-08-20 | 2021-02-17 | I, n, j, u, r, y | Voluntary report | DREAM STATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500H11 |  |  | ON (B)(6) 2021 I WAS DIAGNOSED WITH CHONDROSARCOMA OF THE CRICOID. SURGERY WAS DONE TO REDUCE IT ON (B)(6) 2021. I WAS ADVISED THAT IT WOULD CONTINUE TO GROW AND MAY LEAD TO THE LOSS OF MY VOICE IF REMOVAL OF MY VOICE BOX BECOMES NECESSARY. IT'S GROWTH COULD ALSO IMPEDE MY BREATHING AND LEAD TO A TRACHEOTOMY. I AM CURRENTLY UNDER THE CARE OF DR. (B)(6) AT THE (B)(6) MEDICAL CENTER. THE MRI RESULTS I HAVE LISTED ARE FROM MY LAST MRI DONE AFTER SURGERY. I HAVE REGISTERED BOTH MY OLD AND NEW CPAP MACHINES WITH THE PHILIPS.COM/SRC-UPDATE AS THEIR SERIAL NUMBERS INDICATED THEY WERE BOTH INCLUDED IN THE RECENT RECALL. I DISCONTINUED USING THE MACHINE BUT WOULD LIKE TO BE CONTACTED CONCERNING MY RECENT UNEXPLAINED HEALTH PROBLEM AND IT'S POTENTIAL CONNECTION THE RECALL OF MY MACHINE. I USED MY OLD MACHINE FOR APPROXIMATELY 6 YEARS AND MY CURRENT MACHINE FOR ABOUT A YEAR. I HAVE BOTH MACHINES IF THEY NEED TO BE EXAMINED. I AM A NONSMOKER AND OTHERWISE IN GOOD HEALTH. I USED MY FIRST PHILIPS DREAMSTATION FROM APPROXIMATELY 2016 UNTIL 2021 WHEN I GOT A NEW MODEL I STOPPED USING MY CPAP MACHINE ON MONDAY, (B)(6) 2021 WHEN I HEARD OF THE RECALL. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-08
2518422-2021-03380 | 2021-08-20 | 2021-08-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND POWER SUPPLY WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03378 | 2021-08-20 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | RIN2110X15B | RIN2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VARIATION IN THE VENTILATOR'S PEAK INSPIRATORY PRESSURE DELIVERY AND PEAK END EXPIRATORY PRESSURE WAS OBSERVED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN EVALUATED BY THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VARIATION IN THE VENTILATOR'S PEAK INSPIRATORY PRESSURE DELIVERY AND PEAK END EXPIRATORY PRESSURE WAS OBSERVED. ADDITIONAL INFORMATION RECEIVED INDICATES THE VENTILATOR'S SOFTWARE WAS RE-LOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03377 | 2021-08-20 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S NAVIGATIONAL BUTTON FAILED TO RESPOND ACCURATELY. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S FRONT PANEL/KEYPAD LED BOARD WAS REPLACED TO ADDRESS THE ISSUE |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03374 | 2021-08-20 | 2021-07-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VARIATION IN THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE DELIVERY WAS OBSERVED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VARIATION IN THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE DELIVERY WAS OBSERVED. ADDITIONAL INFORMATION RECEIVED INDICATES THE VENTILATOR'S SOFTWARE WAS RE-LOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03340 | 2021-08-20 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03337 | 2021-08-20 | 2021-07-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WOULD NOT PASS TESTING. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03336 | 2021-08-20 | 2021-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WOULD NOT PASS TESTING. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN BLENDING MODULE ASSEMBLY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03333 | 2021-08-20 | 2021-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND SENSOR BOARD WERE REPLACED TO ADDRESS THE ISSUES. AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03375 | 2021-08-20 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLANK. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE WAS EVALUATED AT THE MANUFACTURER'S SERVICE CENTER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S LCD SCREEN NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03373 | 2021-08-20 | 2021-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S MODE SPONTANEOUSLY CHANGED FROM ASSIST CONTROL TO SPONTANEOUS TIMED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR ALLEGEDLY CHANGED MODES SPONTANEOUSLY FROM ASSIST CONTROL TO SPONTANEOUS TIMED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03379 | 2021-08-20 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VARIATION IN THE VENTILATOR'S PEAK INSPIRATORY PRESSURE AND POSITIVE END EXPIRATORY PRESSURE DELIVERY WAS OBSERVED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VARIATION IN THE VENTILATOR'S PEAK INSPIRATORY PRESSURE DELIVERY AND POSITIVE END EXPIRATORY PRESSURE WAS OBSERVED. ADDITIONAL INFORMATION RECEIVED INDICATES THE VENTILATOR'S SOFTWARE WAS RE-LOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
MW5103410 | 2021-08-20 |  | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILLIPS RESPIRONICS DREAMSTATION BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX700H11C |  |  | SINCE THE RECALL OF BY PHILLIPS RESPIRONICS OF THEIR DREAMSTATION MACHINE AND NO FURTHER INFORMATION FROM THEM, I'VE BEEN CONCERNED ABOUT FINDING A BLACK ALMOST LIKE MELTED PLASTIC FILM ON MY HUMIDIFIER SEAL. IT DIDN'T SEEM LIKE MOLD AND I REPLACED THE SEAL BUT NOW WONDER IF I COULD HAVE INHALED THAT BLACK JUNK FOR AWHILE WITHOUT KNOWING IT. I NEVER USED AN OZONE CLEANER. IS THIS THE TYPE OF PROBLEM THAT CAUSED THE RECALL? HOW CONCERNED SHOULD I BE AND I'D REALLY LIKE TO KNOW WHEN THE COMPANY KNEW ABOUT THE ISSUES. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03376 | 2021-08-20 | 2021-07-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | RIN2110X15B | RIN2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VARIATION IN THE VENTILATOR'S PEAK INSPIRATORY PRESSURE DELIVERY WAS OBSERVED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VARIATION IN THE VENTILATOR'S PEAK INSPIRATORY PRESSURE DELIVERY WAS OBSERVED. ADDITIONAL INFORMATION RECEIVED INDICATES THE VENTILATOR'S SOFTWARE WAS RE-LOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03381 | 2021-08-20 | 2021-08-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03359 | 2021-08-23 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN AIRWAY RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING PARTICLES IN AIRWAY RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION G1-MANUFACTURING CONTACT OFFICE INFO WAS PREVIOUSLY INCORRECT AND IS CORRECTLY UPDATED. SECTION G3 DATE WAS PREVIOUSLY INCORRECT (SHOULD HAVE BEEN 7/19/2021). | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03362 | 2021-08-23 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | SECTION B5 SHOULD BE RETURN AS THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING SINUS ISSUES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDING TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND UNIT SCRAPPED DUE TO AGE. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03416 | 2021-08-23 | 2021-07-27 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102506) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP TONGUE CANCER. THE PATIENT HAD RADIATION AND CHEMOTHERAPY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 07, 2024. SECTION H11 SHOULD BE REPORTED AS THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO DEVELOP SQUAMOUS CELL CARCINOMA ON THE TONGUE. THE PATIENT HAD RADIATION AND CHEMOTHERAPY IN RESPONSE TO THE REPORTED EVENT. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM (PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR). SECTION B2 WAS CHANGED TO OTHER SERIOUS (PREVIOUSLY IT WAS BLANK). SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03363 | 2021-08-23 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03364 | 2021-08-23 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS MISSED TO CAPTURE IN SECTION B5, IT SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED OF EXPERIENCING HEADACHES, CHEST PRESSURE AND PAIN, DRY MOUTH, LITTLE BIT OF DIZINESS, COUGHING AND FEELING EXTREMELY TIRED. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. AFTER THE THIRD ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, THE CUSTOMER STATED THAT THE DEVICE WILL NOT BE RETURNED. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 (CLINICAL CODE) HAS BEEN CORRECTED IN THIS REPORT. SECTION H6 (TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS) HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03409 | 2021-08-23 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. |  | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03410 | 2021-08-23 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX600T11C | DSX600T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING HAVING COUGH AND VISUALIZATION OF BLACK PARTICLES IN THE HUMIDIFIER RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. THE MANUFACTURER PREVIOUSLY NOT CAPTURED, HAVING COUGH AND VISUALIZATION OF BLACK PARTICLES IN THE HUMIDIFIER IN SECTION B5 AND IT HAS BEEN UPDATED IN THIS REPORT AND SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03365 | 2021-08-23 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11 | DSX400H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. IN THE INITIAL REPORT SECTION B5 WAS NOT UPDATED, THE CORRECT B5 SHOULD BE - THE PATIENT HAS ALLEGED PARTICLES IN THE DEVICE, SINUS PROBLEM, NOSE PROBLEM, EYE PROBLEM, SKIN WAS COMING OF NOSE, VISITED A DOCTOR FOR EYE PROBLEM.THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INPSECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO EVIDENCE OF CONTAMINATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF CONTAMINATION ON THE RESERVOIR SEAL, CONTAMINATION UNDER THE DRY BOX, CONTAMINATION IN THE BLOWER BOX. THE MANUFACTURER FOUND EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. A HUMIDIFIER WAS ALSO RETURNED TO THE MANUFACTURER ALONG WITH DEVICE. AN INTERNAL VISUAL INSPECTION OF THE HUMIDIFIER WAS COMPLETED BY THE MANUFACTURER AND FOUND A CONTAMINATION ON THE BOTTOM OF THE HUMIDIFIER ON EITHER SIDE OF THE HEATING PLATE, CONTAMINATION ON THE HEATING PLATE COVER ON THE BOTTOM OF THE HUMIDIFIER. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND E-101 AND E-130 ON THE ERROR LOG. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH CONTAMINATION ON THE RESERVOIR SEAL, CONTAMINATION UNDER THE DRY BOX, CONTAMINATION IN THE BLOWER BOX, CONTAMINATION ON THE BOTTOM OF THE HUMIDIFIER ON EITHER SIDE OF THE HEATING PLATE, CONTAMINATION ON THE HEATING PLATE COVER ON THE BOTTOM OF THE HUMIDIFIER.. THE MANUFACTURER CONFIRMED THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION D9,D10,G3,H6 HAS BEEN UPDATED/CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03361 | 2021-08-23 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED PARTICLE IN AIRWAY FELT LIKE PAINT CHIPS FLOATING IN MASK. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. CORRECTED INFORMATION PROVIDED IN B.5. (THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED PARTICLE IN AIRWAY FELT LIKE PAINT CHIPS FLOATING IN MASK.). SECTION D8, D9, H3 AND H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03392 | 2021-08-23 | 2021-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102243) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CHEST PRESSURE. THE PATIENT WAS ADMITTED TO THE HOSPITAL IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED, CAUSED THE PATIENT TO DEVELOP CHEST PRESSURE AND NAUSEA. THE PATIENT WAS ADMITTED TO THE HOSPITAL AND DISCHARGED IN RESPONSE TO THE REPORTED EVENT. PATIENT TOOK THREE CELEBREX OVER NEXT TWO DAYS AFTER HOSPITAL DISCHARGE. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FOLLOW-UP REPORT WILL BE FILED. SECTION(S) B1 AND B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- CLINICAL CODE HAS BEEN CORRECTED (MISSED TO CAPTURE ALL RELEVANT CODES) SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03393 | 2021-08-23 | 2018-03-28 | D, e, a, t, h | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102217) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PER ADDITIONAL INFORMATION RECEIVED THE PATIENT PASSED AWAY. PRIOR TO HIS DEATH, PATIENT REPORTED THAT HE WAS DIAGNOSED WITH NON-SMALL CELL LUNG CANCER/LARGE CELL. PATIENT ALSO REPORTED THAT HE WOULD WAKE UP IN THE MORNING WITH THROAT AND NASAL IRRITATION VERY OFTEN. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION B2 UPDATED TO DEATH. SECTION D1 CORRECTED. SECTION D4 CORRECTED/UPDATED. SECTION G1 CORRECTED. SECTION H1 UPDATED. SECTION H6 IMPACT CODE CORRECTED. SECTION H6 TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03390 | 2021-08-23 | 2020-10-08 | I, n, j, u, r, y | Manufacturer report | DREAM STATION BIPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX700 | DSX700 |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102336) ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO SEVERE SINUS INFECTIONS. THE PATIENT REQUIRED SURGERY AND PRESCRIPTION MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03406 | 2021-08-23 | 2021-07-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO BIFLEX | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DS750S | DS750S |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS UPDATED REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03398 | 2021-08-23 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO SERIES ASSY | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | 751P | 751P |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03404 | 2021-08-23 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | 761P | 761P |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103421 | 2021-08-23 | 2021-07-29 | I, n, j, u, r, y | Voluntary report | DREAMSTATION HUM CARE PACK DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C |  |  | BREATHING DIFFICULTY; BREATHING DIFFICULTY AND LUNG CONGESTION. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03395 | 2021-08-23 | 2021-05-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT ALSO ALLEGED BLACK SPECS FLOATING IN WATER TANK. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL PART OF THE DEVICE. UNKNOWN DUST/DIRT CONTAMINATION WAS OBSERVED ON SURFACE OF BASE UNIT. MODERATE UNKNOWN CONTAMINATION PRESENT ON SIDE PANEL OUTLET. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND FOUND UNKNOWN WHITE DUST CONTAMINATE PRESENT THROUGHOUT BLOWER BOX HOUSING, MOTOR CASING, MOTOR IMPELLER, AND SIDE PANEL O-RING. UNKNOWN DUST/DIRT CONTAMINATION AND FIBERS FOUND ON THE BLOWER CASING/IMPELLER AND BLOWER BOX WAS INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM CONTAMINATION. THE PRESENCE OF CONTAMINATION THROUGHOUT THE AIR PATH SUGGESTS A SOURCE EXTERNAL TO THE DEVICE. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND 1 ERROR CODE. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION, BUT CONFIRMS PRESENCE OF CONTAMINATION IN THE AIR PATH. SECTION H6 (MEDICAL DEVICE PROBLEM CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONSLUSIONS) WERE UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03394 | 2021-08-23 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DIRT AND DEBRIS IN AIR MASK AND TUBING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING DIRT AND DEBRIS IN AIR MASK AND TUBING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION D8 WAS CORRECTED AND SECTION H6 WAS UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03391 | 2021-08-23 | 2021-07-06 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC |  |  |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102316) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP KIDNEY DISEASE, PNEUMONIA, AND SPOTS ON THEIR LUNGS. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP KIDNEY DISEASE, PNEUMONIA, AND SPOTS ON THEIR LUNGS. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. AFTER FIRST ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WAS UNSUCCESSFUL AND CUSTOMER DECLINED TO RESPOND TO THE GFE RELATED QUESTIONS. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03455 | 2021-08-23 | 2021-07-02 | D, e, a, t, h | Manufacturer report | DREAMSTATION BIPAP AVAPS30 AAM H/HT/C,CA | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS INC | CAX1130T12C | CAX1130T12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT DEATH. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVED AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND ALLEGEDLY CAUSED A PATIENT DEATH. DESPITE MULTIPLE ATTEMPTS, THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03460 | 2021-08-23 | 2021-07-19 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG CANCER. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | THE MANUFACTURER PREVIOUSLY MISSED TO CAPTURE INFORMATION IN SECTION B5. AFTER REVIEW, IT WAS DETERMINED THAT SECTION B5 SHOULD BE FILED AS FOLLOWS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION ALLEGING LUNG CANCER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION NEEDS TO BE CAPTURED WAS ABOUT THE ALLEGED HEADACHE AND DYSPNEA. SECTION H6 HEALTH EFFECTS: CLINICAL CODES HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03459 | 2021-08-23 | 2021-07-27 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | THE MANUFACTURER PREVIOUSLY SUBMITTED MDR 2518422-2021-03459-1 WITH INCORRECT SECTIONS B1, B2, H1, H6. CORRECTIONS TO PREVIOUS MDR ARE MADE IN THIS REPORT AS FOLLOWS. SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY PRODUCT PROBLEM WAS CHECKED IN PREVIOUS MDR) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY IT WAS BLANK) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03366 | 2021-08-23 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING DIFFICULTY BREATHING AND VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03358 | 2021-08-23 | 2021-07-13 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED TO SEEING A SLIMY TEXTURE IN THE WATER RESERVOIR/TUBING AND PARTICLES IN THE TUBING/CHAMBER AND THE DEVICE GETS VERY HOT. THE PATIENT ALLEGES TO EXPERIENCING NAUSEA, INCREASED BLOOD SUGAR, SORENESS IN THE NOSE, DIFFICULTY BREATHING, LOW OXYGEN LEVELS AND COUGHING UP PARTICLES. THE FAMILY MEMBER HAS REPORTED THE PATIENT TO RECEIVE MEDICAL INTERVENTION THAT CPR WAS PERFORMED 4 TIMES SINCE JUNE OF 2021 AND THE PHYSICIAN HAS INSTRUCTED PATIENT TO CONTINUE TO USE NEBULIZER. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE EXTERNAL AND INTERNAL ASPECT OF THE DEVICE WAS INSPECTED AND UNKNOWN BROWN CONTAMINATION AND A POTENTIAL HAIR THAT IS APPROXIMATELY 1 INCH LONG, ON THE MODEM SIDE OF THE DEVICE, UNDER THE PANEL DOOR WERE OBSERVED. UNKNOWN WHITE AND BLACK CONTAMINATION (DUST-LIKE), INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM, AT THE AIR INPUT OF THE BLOWER BOX WAS OBSERVED. THE MANUAFCTIRER ALSO OBSERVED UNKNOWN BROWN CONTAMINATION AROUND THE SD CARD SLOT AND UNKNOWN POTENTIAL WHITE HAIRS OR FIBERS AT THE AIR INPUT OF THE BLOWER BOX. UNKNOWN LIGHT FILM OF GRAY CONTAMINATION ALL THROUGHOUT THE INSIDE OF THE DEVICE ENCLOSURE, UNKNOWN LIGHT BROWN DUST-LIKE CONTAMINATION ON THE TOP OF THE BLOWER BOX, AND ON THE TOP AND BOTTOM OF THE BLOWER MOTOR, AND THE BLOWER MOTOR SEAL AND THROUGHOUT THE INSIDE OF THE BLOWER BOX WAS OBSERVED BY THE MANUFACTURER DURING THE DEVICE EVALUATION. BLACK CONTAMINATION, CONSISTENT WITH KERATIN WAS OBSERVED ON THE AIR OUTLET OF THE BLOWER BOX, TINY UNKNOWN BLACK (INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM) AND LIGHT BROWN SPECS OF CONTAMINATION ON THE BOTTOM THE BLOWER BOX AND IN ISO PORT WAS ALSO OBSERVED. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION, BUT AN AMOUNT OF DUST / DIRT CONTAMINATION, KERATIN CONTAMINATION WERE OBSERVED AND ARE CONSISTENT WITH BEING FROM AN EXTERNAL SOURCE. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03356 | 2021-08-23 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING/CHAMBER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03419 | 2021-08-23 | 2021-07-27 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LEUKEMIA. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03420 | 2021-08-23 | 2021-07-27 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SEIZURES. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING CAUSED A PATIENT TO DEVELOP SEIZURES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03458 | 2021-08-23 | 2021-07-27 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP BREAST CANCER. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED BREAST CANCER. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED BREAST CANCER,AND CHEMO AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION,THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION.AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03396 | 2021-08-23 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO AUTO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSG500S11 | DSG500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | IN PREVIOUS MDR B5 VERBIAGE MENTIONED INCORRECTLY CLINICAL CODES ARE MISSING. IN THIS REPORT CORRECTED AND UPDATED AS BELOW. THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN VISUALIZATION OF PARTICLES AND PATIENT STATEES THAT I HAVE A RUNNY NOSE, CONGESTION, SORE THROAT RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03389 | 2021-08-23 | 2019-06-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102302) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER PREVIOUSLY RECEIVED A VOLUNTARY MEDWATCH (MW5102302) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA,RESPIRATORY ISUUES,DIFFICULTY BREATHING. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE.THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THERE IS NO CUSTOMER INFORMATION AND WE CANNOT REACH OUT TO THE CUSTOMER, HENCE, ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION CAN NOT BE MADE AT THIS TIME. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS UPDATED REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03405 | 2021-08-23 | 2021-12-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE ALLEGEDLY RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND THE DEVICE INFORMATION RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03439 | 2021-08-23 | 2021-08-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03438 | 2021-08-23 | 2021-08-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03437 | 2021-08-23 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03436 | 2021-08-23 | 2021-08-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03195 | 2021-08-23 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1032804 | 1032804 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" ALARM CONDITION RELATED TO CHECK CIRCUIT WAS OBSERVED. THE DEVICE WAS NOT IN PATIENT USE. THE REPORTED ISSUE WAS CONFIRMED. THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03328 | 2021-08-23 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO A VENTILATOR'S LEAK WAS OBSERVED. THE DEVICE WAS NOT IN PATIENT USE. THE REPORTED ISSUE WAS CONFIRMED. THE DEVICE FAILED TEST STEPS DURING TESTING. THE SENSOR BOARD AND THE FLOW SENSOR ASSEMBLY WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03107 | 2021-08-23 | 2021-08-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION RELATED TO PRESSURE. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03440 | 2021-08-23 | 2021-08-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICE'S LCD DISPLAY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03284 | 2021-08-23 | 2021-07-29 | I, n, j, u, r, y | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED HEADACHES AND A STROKE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED HEADACHES AND A STROKE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED. IN THE PREVIOUS REPORT B2, H7, AND H9 WERE OMITTED. THEY ARE REFLECTED ON THIS REPORT. H1 WAS INCORRECTLY REPORTED AS A MALFUNCTION AND SHOULD BE REPORTED AS A SERIOUS INJURY. IT IS CORRECTED ON THIS REPORT. | Y | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03385 | 2021-08-23 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING "SERVICE REQUIRED" ALARM CONDITIONS OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03386 | 2021-08-23 | 2021-08-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03387 | 2021-08-23 | 2021-08-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03388 | 2021-08-23 | 2021-08-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR'S LCD SCREEN WAS REPLACED TO ADDRESS A BLANK SCREEN ISSUE. THE CODING FOR THIS ISSUE IN SECTION H6 WAS INCORRECTLY ENTERED. THE CODING HAS BEEN CORRECTED TO REFLECT THE CORRECT ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
MW5103463 | 2021-08-23 | 2021-08-16 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/RESPIRONICS INC. |  |  |  | MY PHILIPS DREAM STATION CPAP WAS RECALLED IN JUNE AND I DIDN'T FIND OUT ABOUT IT UNTIL MY BOSS SAW AN ARTICLE ABOUT IT IN (B)(6)--TWO MONTHS LATER! MEANWHILE, I HAVE BEEN SUFFERING FROM SINUS ISSUES OF LATE THAT I COULD NOT EXPLAIN. (B)(6) HAS SOLD ME CPAP SUPPLIES IN THE PAST TWO MONTHS & NEVER INFORMED ME OF THIS RECALL. THEY ARE ALSO SLOW TO RESPOND TO GET ME A REPLACEMENT CPAP MACHINE AND AS OF TODAY, DO NOT EVEN ANSWER THEY TELEPHONE. MY PHYSICIAN ADVISED TO NOT USE THE MACHINE, BUT I AM STRUGGLING TO SLEEP. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103461 | 2021-08-23 | 2017-05-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | NEVER HAD KIDNEY ISSUES BEFORE USING UNIT. ABOUT A YEAR AFTER USE I DEVELOPED KIDNEY ISSUES, WORSENING ALLERGIES AND MORNING HEADACHES, FEELING WORSE THAN WHEN I WENT TO SLEEP. DOCTORS TOLD ME TO KEEP USING DESPITE ISSUES. WHEN I HEARD ABOUT RECALL I BROKE DOWN MY MACHINE BEST I COULD AND DID FIND SOME BLACK SPECKS IN SOME CREVICES. I DO MONITOR MY FILTERS AND CLEAN REGULARLY. I HAVE HAD MANY DOCTORS VISITS AND TESTS FOR KIDNEY AND ABDOMEN ISSUES BECAUSE OF PERSISTENT PAIN. THE DOCTORS HAVE BEEN UNABLE TO DIAGNOSE A SPECIFIC ISSUE FOR SEVERAL YEARS SO I LIVE WITH THE PAIN. I THINK IT'S BECAUSE I'M (B)(6) AND DON'T EXPECT ME TO HAVE ISSUES GENERALLY AFFECTING OLDER PATIENTS. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103454 | 2021-08-23 | 2021-07-22 | I, n, j, u, r, y | Voluntary report | DREAM STATION 1 CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | WITHIN TWO DAYS OF FEELING LIKE MY SINUSES WERE FLARED UP AND THINKING I MAY HAVE A SINUS INFECTION, I WENT TO AN URGENT CARE. I HAD FEVER, COUGHING (BROWN SPUTUM, BODY ACHES, HEADACHES, PAIN IN CHEST, LOW BACK PAIN. THE CHEST X-RAY CONFIRMED PNEUMONIA. I WENT BACK TO THE URGENT CARE TWICE AND TO HOSP ER. CAT SCAN SHOWED NO IMPROVEMENT OF PNEUMONIA ALTHOUGH SYMPTOMS, EXCEPT CHEST PAIN, WERE GETTING BETTER. I HAVE FOLLOW UP APPTS WITH PCP AND LUNG DR. NOT USING CPAP AND HAVING PROBLEMS WITH FUNCTIONING. 67 PAUSES PER HOUR AND OXY DROPPING TO 85 PERCENT. I AM FOLLOWING UP WITH PCP ON (B)(6) AND I AM TRYING TO LOCATE A LUNG DR. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | I |  | 2021-08
2518422-2021-03417 | 2021-08-23 | 2021-07-27 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION BIPAP AUTOSV H/C, DS | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS INC | DSX900H11C | DSX900H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102506) ALLEGING A BILEVEL CONTINUOUS POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP TONGUE CANCER. THE PATIENT HAD RADIATION AND CHEMOTHERAPY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03441 | 2021-08-23 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03415 | 2021-08-23 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102481) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A NODULE ON THEIR LUNG, AND DEAD BRAIN CELLS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 07, 2024, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102481) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A NODULE ON THEIR LUNG, AND DEAD BRAIN CELLS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM (ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR). SECTION B2 WAS CHANGED TO BLANK (PREVIOUSLY IT WAS OTHER SERIOUS). SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03442 | 2021-08-23 | 2021-08-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03444 | 2021-08-23 | 2021-08-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03454 | 2021-08-23 | 2021-08-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES RELATED TO THE INTERNAL BATTERY WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03066 | 2021-08-23 | 2021-08-10 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, INTL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | INX500T15 | INX500T15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD). THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | THE MANUFACTURER PREVIOUSLY SUBMITTED MDR 2518422-2021-03066-1 WITH INCORRECT SECTIONS B1, B2, H1, H6. CORRECTIONS TO PREVIOUS MDR ARE MADE IN THIS REPORT AS FOLLOWS. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03443 | 2021-08-23 | 2021-08-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03449 | 2021-08-24 | 2021-05-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX400H11C | DSX400H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03453 | 2021-08-24 | 2021-08-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO AUTO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | GBG502S15 | GBG502S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03430 | 2021-08-24 | 2021-08-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. | THIS MDR IS BEING SUBMITTED AS PART OF A BATCH SUBMISSION OF PREVIOUSLY CLOSED COMPLAINTS THAT WERE REASSESSED AS REPORTABLE FOAM DEGRADATION COMPLAINTS; DISCOVERED AS PART OF A RETROSPECTIVE REMEDIATION REVIEW. THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION, THE POLLEN FILTER, INLET AIR PATH ASSEMBLY AND REMOVABLE AIR PATH FOAM WERE REPLACED AS PER FIELD COMMUNICATION. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03429 | 2021-08-24 | 2021-08-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03463 | 2021-08-24 | 2024-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | RIN2110X15B | RIN2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR POWERED OFF AND FAILED TO POWER BACK ON. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUES. | THE MANUFACTURER PREVIOUSLY REPORTED THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR POWERED OFF AND FAILED TO POWER BACK ON. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUES. THE DEVICE EVALUATION WAS COMPLETED AND THE SYSTEM BOARD WAS NOT REPLACED TO ADDRESS THE ISSUE. THE DEVICE ERROR LOG WAS REVIEWED AND BATTERY DEPLETED ALARMS WERE OBSERVED. A LOSS OF POWER ALARM OCCURRED IMMEDIATELY FOLLOWING THE DEPLETED ALARMS. THIS ISSUE WAS FOUND TO BE CAUSED BY A SOFTWARE ALGORITHM CALCULATION ERROR. THE SOFTWARE WAS UPGRADED TO ADDRESS THE ISSUE. THE DEVICE WAS CLEANED AND PASSED ALL FINAL TESTING. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03474 | 2021-08-24 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
MW5103492 | 2021-08-24 | 2021-03-30 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | AUTO CPAP | DSX500T11C |  | I WAKE UP IN THE MORNING WITH HEADACHES WHICH GO AWAY AS THE DAY MOVES ALONG, I HAVE ASTHMA SO IT IRRITATES THAT. I HAVE BEEN WAKING UP WITH CHEST PRESSURE THAT GOES AWAY THROUGHOUT THE DAY. I ASK ABOUT A REPLACEMENT OR NEW MACHINE BUT THEY ONLY TELL ME THEY HAVE FILTERS BUT I HAVE A PHILIPS DREAM STATION THAT IS A HUMIDIFIED ONE AND HAVE READ IT WILL NOT WORK WITH THAT TYPE AND HUMIDIFIER BECAUSE THE MOISTURE WILL CAN'T GET THROUGH THE FILTER SO IT WILL CAUSE BACTERIA. I HAVE ASKED FOR A NEW DREAM STATION 2 MACHINE AS A REPLACEMENT AS MINE IS ON THE DO NOT USE LIST BUT CAN'T GET ANYWHERE WITH ANYONE. PHILIPS JUST BLAMES THE FDA BECAUSE THEY SAY THEY HAVEN'T GIVEN THE APPROVAL YET. SINUS. I HAVE BEEN CALLING MY DOCTORS AND PHILIPS BUT THEY WON'T SEE ME OR BE HEARD. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103484 | 2021-08-24 | 2021-06-14 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION CPAP DEVICE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS, INC. |  |  |  | PHILIPS RECALLED MY CPAP DREAMSTATION. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | I |  | 2021-08
MW5103479 | 2021-08-24 |  | I, n, j, u, r, y | Voluntary report | SYSTEM ONE REMSTAR AUTO A-FLEX | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | MY MOTHER BEGAN USING HER CPAP IN EARLY 2014. AT THAT TIME WE LIVED IN (B)(6). WE MOVED TO (B)(6) AT THE END OF 2014. WITHIN THE FIRST YEAR OF BEING IN (B)(6), MY MOTHER BEGAN DEVELOPING NASAL IRRITATION AND A COUGH. WE THOUGHT SHE HAD DEVELOPED AN ALLERGY. WE HAD HER TESTED FOR ALLERGIES, BUT SHE HAD NONE. MY MOTHER'S NASAL PASSAGES WERE QUITE INFLAMED SO THE ALLERGIST, IN LIEU OF ANY ALLERGIES, DIAGNOSED HER WITH CHRONIC RHINITIS, WHICH SHE'D PREVIOUSLY NEVER HAD. SHE ALSO DEVELOPED A COUGH. IT WAS MILD AT FIRST, BUT HAS BECOME CHRONICALLY PERSISTENT IN THE LAST 3-4 YEARS. SHE'S SEEN ENT, GI, PULMONARY, CARDIOLOGY, AND ALLERGY SPECIALISTS. THEY'VE RULED OUT COPD, PNEUMONIA, CHF, AND ASTHMA. NOT ONE OF THEM HAS BEEN ABLE TO DETERMINE WHY SHE HAS THIS PERSISTENT, CHRONIC COUGH AND CHRONIC RHINITIS. NOTHING THE DOCTORS HAVE TRIED HELP TO MITIGATE THE SYMPTOMS. HER BODY IS PRODUCING MORE PHLEGM THAN EVER BEFORE. THE PHLEGM COLLECTS IN HER THROAT AND SHE SOUNDS LIKE SHE'S DROWNING. SHE STOPPED USING PHILIPS RESPIRONICS CPAP (B)(6) 2021. HER SYMPTOMS HAVEN'T IMPROVED. I BELIEVE THERE'S PERMANENT DAMAGE TO HER RESPIRATORY SYSTEM, BOTH THE NOSE, VOCAL CORDS, AND LUNGS. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03426 | 2021-08-24 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLANK. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
MW5103477 | 2021-08-24 | 2020-10-08 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500T11 |  |  | I HAVE BEEN USING A PHILIPS DREAMSTATION AUTO CPAP SINCE OCTOBER 15 OF 2016. (PHILIPS RECALL CONFIRMATION CODE: (B)(4)) IN (B)(6) 2020, I BEGAN TO HAVE SIGNIFICANT SHORTNESS OF BREATH AND CHEST PAINS. ON (B)(6) 2020, MY CARDIOLOGIST (DR (B)(6)) PERFORMED A NUCLEAR STRESS TEST ON ME, DURING WHICH I COLLAPSED AND WAS RUSHED TO LONG BEACH MEMORIAL EMERGENCY DEPARTMENT WHERE I WAS SUBJECTED TO A BATTERY OF TESTS. NONE OF THE TESTS DEFINITIVELY EXPLAINED THE SHORTNESS OF BREATH AND CHEST PAINS. I CONTINUE, TO THIS DAY, TO HAVE SIGNIFICANT SHORTNESS OF BREATH AND OCCASIONAL CHEST PAIN. ON (B)(6) 2021, WHEN I INFORMED DR (B)(6) OF THE PHILIPS RECALL, HE INDICATED THAT PHILIPS WAS LIKELY "LIABLE" FOR MY SHORTNESS OF BREATH AND CHEST PAINS. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | I |  | 2021-08
2518422-2021-03450 | 2021-08-24 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO AUTO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSG500S11 | DSG500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02979 | 2021-08-24 | 2021-08-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, GB | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | GBX500H15 | GBX500H15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103475 | 2021-08-24 |  | I, n, j, u, r, y | Voluntary report | REMSTAR AUTO A-FLEX | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | TO WHOM THIS MAY CONCERN, MY MOTHER IS A HEALTHY (B)(6) YEAR OLD NON-SMOKER WHO DEVELOPED METASTATIC LUNG CANCER THAT TRAVELED TO HER BRAIN. SHE WAS DIAGNOSED AS OF (B)(6) 2021. MY MOM HAS BEEN USING THE REMSTAR AUTO SLEEP APNEA MACHINE MADE BY PHILLIPS. COINCIDENTALLY AS SHE WAS DIAGNOSED YESTERDAY WE HAPPENED TO SEE ON THE NEWS THAT PHILLIPS WAS HAVING A RECALL ON THEIR CPAP MACHINES. OF COURSE MY FAMILY LOOKED INTO IT MORE, AND CAME ACROSS OF THE FDA'S ARTICLE ON THE RECALL. WE HAVE ALREADY WENT TO THE PHILLIPS WEBSITE TO REGISTER HER DEVICE. WE USED THE SERIAL CODE ON THE DEVICE TO VERIFY THAT IT WAS ONE ON THE RECALL AND IT WAS. AND WE NOW HAVE A CONFIRMATION CODE FROM PHILLIPS. CANCER DOES NOT RUN IN OUR FAMILY, SO WHEN WE HEARD ABOUT THIS RECALL WE BECAME CONCERNED AND NOW WONDER COULD THIS HAVE BEEN A CAUSE FOR HER DIAGNOSIS. SHE HAS BEEN USING THIS SAME DEVICE FOR OVER 5 YEARS. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | I |  | 2021-08
2518422-2021-03473 | 2021-08-24 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BLOWER MOTOR WAS RECALIBRATED TO ADDRESS THE ISSUE. | IT WAS INITIALLY REPORTED, THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BLOWER MOTOR WAS RECALIBRATED TO ADDRESS THE ISSUE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S OUTLET SIDE PANEL WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03472 | 2021-08-24 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03480 | 2021-08-24 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE WOULD NOT CONNECT TO THE TEST STATION. THE SYSTEM BOARD RIBBON CABLES WERE DISCONNECTED. THE CABLES WERE RECONNECTED TO ADDRESS THE ISSUE.THE DEVICE HAD EVIDENCE OF TAMPERING. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S BLOWER MOTOR NEEDS REPLACED TO ADDRESS THE ISSUE. | IT WAS INITIALLY REPORTED, A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE WOULD NOT CONNECT TO THE TEST STATION. THE SYSTEM BOARD RIBBON CABLES WERE DISCONNECTED. THE CABLES WERE RECONNECTED TO ADDRESS THE ISSUE. THE DEVICE HAD EVIDENCE OF TAMPERING. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S BLOWER MOTOR NEEDS REPLACED TO ADDRESS THE ISSUE. DURING FINAL TESTING, THE DEVICE FAILED A TEST STEP. THE OUTLET SIDE PANEL WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
MW5103483 | 2021-08-24 | 2021-08-05 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I HAVE BEEN USING THE PHILIPS DREAM STATION BIPAP MACHINE SINCE 2018. I HAVE USED OZONE PERIODICALLY TO CLEAN THE MACHINE. I USED OZONE SINCE (B)(6) 2021. OZONE WAS USED EVERY 7-10 DAYS. I STARTED HAVING ABNORMAL SKIN IRRITATION AND PEELING ON (B)(6) 2021. I STARTED TAKING PICTURES OF THIS ABNORMAL SKIN PEALING ON (B)(6) 2021. ON (B)(6) 2021, I WAS REVIEWING MY EMAILS. I WAS A CERTIFIED DOT EXAMINER, AND I STILL RECEIVE DOT INFORMATION. I READ THE ARTICLE, AND IT PERTAIN TO THE PHILIPS RECALL OF THE DREAM STATION BIPAP MACHINE. I WAS NOT NOTIFIED OF THIS RECALL BY EITHER PHILIPS OR MY DME COMPANY. I COULD NOT PUT A PUZZLE TOGETHER WITH THE SKIN PEALING AND THE MULTIPLE RASHES ON MY FACE. IT WAS COINCIDENTAL OR BY GOD'S GRACE THAT I RECEIVED THIS ARTICLE FROM THE DOT. IF IT WAS NOT FOR THE DOT ARTICLE, I WOULD STILL BE USING MY BIPAP MACHINE. I AM UNSURE OF ALL THE CONSEQUENCES OF USING THE BIPAP MACHINE AT THIS TIME FRAME. I WILL BE SEEING MY PULMONOLOGIST ON (B)(6) 2021. I KNOW MY LUNG FUNCTIONS HAVE DECREASED. THIS WILL BE JUSTIFIED BY MY REPEAT SPIROMETRY THAT I WOULD BE HAVING DONE. I SAW THE DERMATOLOGIST AND I AM ON MORE MEDICATIONS NOW. I WAS DIAGNOSED WITH CONTACT DERMATITIS PLUS OTHER SKIN DIAGNOSES. MY NOSE HAS BEEN FEELING FOR THE LAST 16 DAYS. AS A PREVIOUS HEALTHCARE PROVIDER, I KNOW WHAT A SECOND-DEGREE BURN LOOKS LIKE. MY NOSE IS RAW AND PAINFUL. I WILL IN ENCLOSE PHOTOS OF THE CONTINUAL PROCESS OF MY FACIAL CHANGES. I AM MOST CONCERNED ABOUT MY RESPIRATORY STATUS. I HAVE PRIMARY PROGRESSIVE MS. MY WEAKNESS HAS INCREASED OVER THE LAST 2 TO 3 MONTHS. I HAVE HAD INCREASED PARESTHESIAS LEFT SIDE OF MY FACE AND LEFT ARM. I HAVE NOT BEEN AFFECTED BY THIS BY THE PRIMARY PROGRESSIVE MS (PPMS). I HAVE HAD INCREASE MUSCLE CONTRACTIONS WHICH IS USUALLY IN THE LEFT LEG BUT NOW IT IS IN THE RIGHT LEG. MY MUSCLES IN BILATERAL HAM, QUADS AND CALVES HAVE BEEN AS FIRM AS BRICKS EVEN WITH 2 MUSCLE RELAXERS, STRETCHING, TENS UNIT AND MASSAGE VIBRATOR. I HAVE A LIST THAT I PROVIDED TO MY PROVIDERS THAT EXPLAINS ALL THAT I HAVE HAD OVER 2 MONTHS. THESE WERE WRITTEN 3 WEEKS PRIOR TO THE NOTIFICATION EMAIL FROM THE DOT AGENCY. I WILL OBTAIN A PROOF OF THE INFORMATION FROM THE PROVIDER ON (B)(6) WHEN I SEE ANOTHER NEW PROVIDER, A NEPHROLOGIST, THIS TIME. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02932 | 2021-08-24 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | RGBX500S15 | RGBX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL AND INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND EVIDENCE OF DUST AND DIRT CONTAMINATE AND DARK FIBER CONTAMINATE THROUGHOUT THE BASE UNIT INCLUDING THE AIR INLET AREA, INLET CANAL, IN AND AROUND THE OUTLET PORT, UI PANEL CHANNELS, THE BLOWER BOX, THE BLOWER AND ISOLATORS, BLOWER IMPELLERS, AND BLOWER SEAL. THE HUMIDIFIER WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR EVALUATION. THE MANUFACTURER FOUND EVIDENCE OF DUST AND DIRT CONTAMINATE AND DARK FIBER CONTAMINATE THROUGHOUT THE HUMIDIFIER INCLUDING THE UPPER AND LOWER SURFACES. THE MANUFACTURER ALSO FOUND EVIDENCE OF AN EXCESSIVE AMOUNT OF WHITE POWDER OR POTENTIAL MINERAL DEPOSITS IN THE HUMIDIFIER AND IN THE TANK. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS LOGGED. THE DEVICE WAS APPLIED POWER AND THE DEVICE OPERATED PROPERLY. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH DARK FIBER, DUST AND DIRT CONTAMINATE AND ALSO CONSIST WITH WHITE POWDER OR POTENTIAL MINERAL CONTAMINATION, INCONSISTENT WITH THE SOUND ABATEMENT FOAM. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION D9, D10, G3, H3 AND H6 HAS BEEN UPDATED / CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103481 | 2021-08-24 | 2014-06-06 | I, n, j, u, r, y | Voluntary report | RESPIRONICS REMSTAR CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | 560P |  |  | RECEIVED A REPLACEMENT PHILLIPS CPAP MACHINE (USING CPAP FOR 30 YEARS) AND MACHINE NEVER SEEMED RIGHT. TOOK IT BACK 3 TIMES. MY BREATHING GRADUALLY GOT WORSE. WENT TO PULMONARY, EAR NOSE AND THROAT AND ASTHMA AND IMMUNOLOGY WITH NO RESOLVE. PROBLEM GETTING WORSE. I REPLACED WITH RES MED MACHINE (B)(6) 2021. MY PHILIPS MACHINE IS ON RECALL LIST. I REPORTED IT AND HAVE HEARD NOTHING. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02954 | 2021-08-24 | 2021-07-06 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP PRECANCEROUS CELLS IN THEIR THROAT AND NOSE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING TO DEVELOP PRECANCEROUS CELLS IN THEIR THROAT AND NOSE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03422 | 2021-08-24 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER , A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03425 | 2021-08-24 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITIONS OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S POWER SUPPLY AND INTERNAL BATTERY WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03424 | 2021-08-24 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLANK. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. AN ISSUE RELATED TO THE FLOW SENSOR ASSEMBLY WAS OBSERVED. THE DEVICE'S FLOW SENSOR ASSEMBLY WAS REPLACED TO ADDRESS THE ISSUE. THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02955 | 2021-08-24 | 2021-07-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500S11 | DSX500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SEIZURES. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03423 | 2021-08-24 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03159 | 2021-08-24 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. EXTERIOR INVESTIGATION STATES UNKNOWN DUST/DIRT CONTAMINATION OBSERVED. THE GROOVE BETWEEN WHERE THE TOP AND BOTTOM CLOSURE MEET LOOKS AS IF IT WAS ATTEMPTED TO HAVE THE TOP PRIED OFF WITH TOOLS. PIL OBSERVED THAT THE TOP ENCLOSURE WAS SCRATCHED, AND UI BUTTON KEYPAD WAS CRACKED AND DAMAGED. PIL OBSERVED A SLASH THROUGH BOTTOM ENCLOSURE AND LABELING, POSSIBLY LIKELY FROM A BOX CUTTER. INTERIOR INVESTIGATION STATES PIL OBSERVED A WHITE RESIDUE THROUGHOUT. THERE WAS A SLIGHT KERATIN-LIKE SUBSTANCE BUILDUP AT THE BLOWER BOX HOUSING OUTLET. PIL OBSERVED UNKNOWN DUST/DEBRIS CONTAMINATION THROUGHOUT INTERIOR AND MINERAL DEPOSITS ON MOTOR HOUSING. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED IN THE BASE UNIT. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERORRS. THE MANUFACTURER CONCLUDES THE DEVICE HAS DUST/DEBRIS CONTAMINATION. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION(S) D8, D9 AND H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02930 | 2021-08-24 | 2021-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING COUGH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE - THE PATIENT ALLEGED COUGH. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03475 | 2021-08-24 | 2021-04-14 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102141) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP HIVES. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER PREVIOUSLY SUBMITTED MDR 2518422-2021-03475-3 WITH INCORRECT SECTIONS B1, B2, H1, H6. CORRECTIONS TO PREVIOUS MDR ARE MADE IN THIS REPORT AS FOLLOWS. SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY PRODUCT PROBLEM WAS CHECKED IN PREVIOUS MDR) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY IT WAS BLANK) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6- HEALTH IMPACT CODE WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02957 | 2021-08-24 | 2021-07-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER INITIALLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. IN THE INITIAL REPORT SECTION B5 WAS NOT UPDATED, THE CORRECT B5 SHOULD BE - THE PATIENT HAS ALLEGED PARTICLES IN THE DEVICE, WHEEZING, COUGH, SHORTNESS OF BREATHE. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF UNIDENTIFIED CONTAMINATION AT THE AIR INTAKE UNDER THE FILTER AREA OF THE BLOWER BOX INCLUDES WHITE AND LIGHT BROWN CONTAMINANTS AT THE AIR INLET, SMALL PIECES OF LIGHT-COLORED HAIR. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. A HUMIDIFIER WAS ALSO RETURNED TO THE MANUFACTURER ALONG WITH DEVICE. AN INTERNAL VISUAL INSPECTION OF THE HUMIDIFIER WAS COMPLETED BY THE MANUFACTURER AND FOUND EVIDENCE OF UNIDENTIFIED CONTAMINATION INCLUDES LIGHT BROWN STAINS,WHITE CONTAMINANTS, BROWN CONTAMINATION ON THE OUTSIDE OF THE BLOWER BOX SEAL ALONG WITH HAIRS AT THE BOTTOM OF THE HUMIDIFIER BOX AND ON THE HEATER PLATE. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERROR LOGGED. THE MANUFACTURER CONCLUDES THAT THE UNIDENTIFIED CONTAMINATION AT THE AIR INTAKE UNDER THE FILTER AREA OF THE BLOWER BOX INCLUDES WHITE AND LIGHT BROWN CONTAMINANTS AT THE AIR INLET, SMALL PIECES OF LIGHT-COLORED HAIR, AND UNIDENTIFIED CONTAMINATION INCLUDES LIGHT BROWN STAINS,WHITE CONTAMINANTS, BROWN CONTAMINATION ON THE OUTSIDE OF THE BLOWER BOX SEAL OF THE HUMIDIFIER ALONG WITH HAIRS AT THE BOTTOM OF THE HUMIDIFIER BOX, AND ON THE HEATER PLATE WERE INCONSISTENT WITH THE SOUND ABATEMENT FOAM. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION D9,G3,H6 HAS BEEN UPDATED/CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03452 | 2021-08-24 | 2021-08-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO AUTO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | GBG502S15 | GBG502S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03384 | 2021-08-24 | 2021-08-17 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AUX500T15 | AUX500T15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP HEART ATTACK LIKE SYMPTOMS. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION B5 WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT SHOULD BE REPORTED AS FOLLOWS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO BE HOSPITALIZED FOR HEART ATTACK AND ASTHMA-LIKE SYMPTOMS. NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. SECTION H6 HEALTH EFFECT - CLINICAL CODE WAS MISSED TO CAPTURE IN PREVIOUS REPORT. IT HAS BEEN CORRECTED OR UPDATED IN THIS REPORT. UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03435 | 2021-08-24 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING/CHAMBER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING/CHAMBER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTION G1-MANUFACTURER CONTACT OFFICE INFO WAS PREVIOUSLY INCORRECT AND HAS NOW BEEN CORRECTLY UPDATED. SECTION G3 WAS CORRECTED AND IT SHOULD HAVE BEEN 7/16/2021 INSTEAD OF 8/17/2021. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03434 | 2021-08-24 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03433 | 2021-08-24 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03411 | 2021-08-24 |  | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03428 | 2021-08-24 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02956 | 2021-08-24 | 2021-07-28 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX700H11 | DSX700H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03408 | 2021-08-24 | 2021-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION EXPERT W/HUM, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | CAX501H12 | CAX501H12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN THE INITIAL REPORT, SECTION B5 WAS INCORRECT, THE CORRECT B5 SHOULD BE- THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER RECEIVED VOLUNTARY MEDWATCH (MW5102229 AND MW5102223). THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM.THE PATIENT HAS ALLEGED SEEING BLACK PARTICLES, CHEST PAIN, SINUS PRESSURE AND DIZZINESS. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL AND INTERNAL VISUAL INSPECTION OF DEVICE WERE COMPLETED BY THE MANUFACTURER AND FOUND EVIDENCE OF UNKNOWN DUST CONTAMINANT ON THE BOTTOM ENCLOSURE, UNKNOWN DUST CONTAMINANT ON THE BLOWER BOX, LIQUID INGRESS ON THE BLOWER AND BLOWER BOX, KERATIN-LIKE SUBSTANCE ON THE BLOWER BOX OUTLET. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERROR LOGGED. THE DEVICE WAS APPLIED POWER AND THE DEVICE OPERATED PROPERLY. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH LIQUID INGRESS ON THE BLOWER AND BLOWER BOX, KERATIN-LIKE SUBSTANCE ON THE BLOWER BOX OUTLET AND UNKNOWN DUST CONTAMINANT ON THE BOTTOM ENCLOSURE, UNKNOWN DUST CONTAMINANT ON THE BLOWER BOX WERE INCONSISTENT WITH SOUND ABATEMENT FOAM. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. IN THIS REPORT, SECTION A1, D9, G3, H6 HAVE BEEN UPDATED/CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03413 | 2021-08-24 | 2018-07-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP,AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM.THERE WAS NO REPORT OF PATIENT HARM OR INJURY. AFTER THE FINAL ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, THE CUSTOMER STATED THAT THE DEVICE WILL NOT BE RETURNED, HE NEED ONLY BIPAP,HE IS NOT USING CPAP.THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03412 | 2021-08-24 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING SEEING SMALL PARTICLES OF A DARK GREY OR BLACK SUBSTANCE THAT APPEARED TO BE URETHANE FOAM WHICH HAD GOTTON LOOSE AND WERE LOCATED ALMOST AT THE END OF THE TUBES IN THE NOSE MASK RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03485 | 2021-08-25 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING OF VISUALIZATION OF BLACK PARTICLES IN HUMIDIFIER CHAMBER, WATER CHAMBER RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION B5 IN THE INITIAL REPORT SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF BLACK PARTICLES IN HUMIDIFIER CHAMBER, WATER CHAMBER RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. SECTIONS B3 AND E1 (E-MAIL ADDRESS) CORRECTED. SECTION G3 WAS INCORRECTLY CAPTURED IN INITIAL REPORT (IT HAS TO BE 05-AUG-2021). | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03482 | 2021-08-25 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER PREVIOULSY REPORTED A CPAP DEVICE'S SOUND ABATEMENT FOAM ALLEGEDLY BECAME DEGRADED. THERE WAS NO PATIENT HARM OR INJURY. THE MANUFACTURER ALSO ALLEGED SEEING BLACK PARTICLES AND HAVING HEADACHES. THE MANUFACTURER RECEIVED NEW INFORMATION OF VOLUNTARY MEDWATCH (MW5106095) THAT SEEING BLACK / BROWN DUST IN THE TUBE AND WATER CHAMBER. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03483 | 2021-08-25 | 2021-07-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11 | DSX700T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. NEW INFORMATION WAS PROVIDED IN SECTION D4. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03481 | 2021-08-25 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING OF PARTICLES IN TUBING, SORE THROAT, CHEST PAINS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING, SORE THROAT, CHEST PAINS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTIONS D1 AND H6 HEALTH EFFECT-CLINICAL CODE CORRECTED. SECTIONS D4, H4, H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03160 | 2021-08-25 | 2021-07-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 WAS UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02929 | 2021-08-25 | 2021-07-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03484 | 2021-08-25 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING STRANGE ODOR AND VISUALIZATION OF PARTICLES AND AWFUL TASTE IN MOUTH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02931 | 2021-08-25 | 2021-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | A CPAP DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS VISIBLE FOAM DEGRADATION AND DEVICE SCRAPPED DUE TO PRE-DESIGNATION AND DATE (B)(6) 2017. SECTION G1,G2 WASS CORRECTED AND H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03161 | 2021-08-25 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03499 | 2021-08-25 | 2021-07-06 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A TUMOR ON THEIR BACK. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. SECTION B5 SHOULD BE PREVIOUSLY REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT ALLEGED TO DEVELOP A TUMOR ON THEIR BACK ,HEART ISSUES. SHE WAS GETTING ADDITIONAL TESTS ON THE TUMOR SOON. PATIENT ALSO HAS COPD AS PREEXISITING CONDITION. NO OTHER CLINICAL INFORMATION WAS REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03487 | 2021-08-25 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGED EXPERIENCING HEADACHE, BLACK PARTICLES, DIZZINESS, RESPIRATORY PROBLEMS AND SINUS INFECTIONS. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03490 | 2021-08-25 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING HEAVILY COUGH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. SECTION H6(CLINICAL CODE) CORRECTED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03489 | 2021-08-25 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION B5 SHOULD BE REPORTED AS THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES, PERSISTENT SINUS INFECTIONS, DIFFICULTY BREATHING, SORE THROAT, DIFFICULTY SWALLOWING, HEADACHES, SOME NAUSEA, KIDNEY STONES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDING TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERROR FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03486 | 2021-08-25 | 2011-07-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION B5 SHOULD BE REPORTED AS THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDING TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERROR FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03491 | 2021-08-25 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. SECTION B5 SHOULD BE PREVIOUSLY REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED PARTICLES IN TUBING/CHAMBER. DIFFICULTY BREATHING/SHORT OF BREATH. BURNING/SMOKE/ELECTRICAL ODOR. THE SMELL WOKE HIM UP. THE RESERVOIR WAS FULL OF BLACK PARTICLES. PATIENT STATED IT TOOK HIM TWO DAYS TO CLEAR HIS LUNGS. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. TTHE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND OBSERVED MODERATE UNKNOWN CONTAMINATION ON ALL SURFACES OF DEVICE. AT POWER INLET, THE RUST RESIDUE AND CORROSION INDICATE UNKNOWN LIQUID INGRESS TO THE DEVICE. IN INTERIOR INVESTIGATION OBSERVED SEVERE UNKNOWN CONTAMINATION THROUGHOUT THE INTERNALS OF THE DEVICE. OBSERVED UNKNOWN LIQUID INGRESS AT THE P4 CONNECTOR CABLE. WATER INGRESS HAS BEEN PREVIOUSLY ADDRESSED IN INV1023. AN UNKNOWN DUST/DIRT AND PARTICLE CONTAMINANT IS PRESENT THROUGHOUT BLOWER BOX HOUSING, BLOWER, AND BLOWER IMPELLER. COMPONENT J3 ON PCA SEEMS TO HAVE LIQUID DAMAGE. THE UNKNOWN CONTAMINANT SEEMS TO HAVE BEEN DRAWN INTO THE BLOWER AND THERE IS NO LONGER MOVEMENT OF THE IMPELLER. MOST LIKELY THIS CONTAMINANT CAUSED THE BLOWER TO SEIZE WHICH COULD ACCOUNT FOR THE ODOR. ALSO SEEN IN THE BLOWER HOUSING LOOKS TO BE AN UNKNOWN METAL CONTAMINANT. OBSERVED THE SOUND ABATEMENT FOAM WAS PRESENT WITH THE SAME UNKNOWN DUST/DIRT CONTAMINATION THAT IS PRESENT WITHIN THE REST OF THE DEVICE. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED IN THE BASE UNIT. THE UNKNOWN DUST/DIRT CONTAMINATION AND FIBERS FOUND ON THE BLOWER CASING/IMPELLER AND BLOWER BOX WAS INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM CONTAMINATION. THE MANUFACTURER CONCLUDES UNABLE DIRECTLY ADDRESS THE SYMPTOMS OUTLINED IN THE COMPLAINT AND CAN CONFIRM THAT THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN OBSERVED IN THE BASE UNIT. ALSO CAN CONFIRM THE PRESENCE OF CONTAMINATION IN THE AIRPATH AND UNKNOWN LIQUID INGRESS TO DEVICE. IN THIS REPORT, SECTION D9, G3, H3 HAS BEEN UPDATED OR CORRECTED. SECTION H6 (HEALTH EFFECT - CLINICAL CODE) HAS BEEN CORRECTED AND (TYPE OF INVESTIGATION,INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS) HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03492 | 2021-08-25 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS INITIALLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE - THE PATIENT ALLEGED PARTICLES IN TUBING/CHAMBER. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND FOREIGN WHITE POWDERY SUBSTANCE WAS FOUND ON THE BLOWER MOTOR FINS, POTENTIAL DARK KERATIN PARTICLES WERE FOUND ON THE BLOWER BOX OUTLET AND INLET PORT THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGREDATION/BREAKDOWN. THE DEVICE DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERROR. THE MANUFACTURER CONCLUDES NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN FOUND. PIL FOUND POTENTIAL KERATIN PARTICLES ON THE BLOWER BOX OUTLET AND INLET PORTS. PIL FOUND A FOREIGN WHITE SUBSTANCE ON THE BLOWER MOTOR FINS. SECTION D8, D9 AND H3, H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03502 | 2021-08-25 | 2021-07-02 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP BLOOD IN THEIR STOOL. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. SECTION B5 SHOULD BE PREVIOUSLY REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT ALLEGED TO DEVELOP HEADACHES AND BLOOD IN THEIR STOOL. NO MEDICAL INTERVENTION WAS REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HEALTH EFFECT - CLINICAL CODE WHICH WAS MISSED IN PREVIOUS REPORT WAS CAPTURED. SECTION H6 TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03503 | 2021-08-25 | 2021-06-30 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SEIZURES. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP,AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING CAUSED A PATIENT TO DEVELOP SEIZURES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM.THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03451 | 2021-08-25 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500H11C | DSX500H11C |  | A CPAP DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03371 | 2021-08-25 | 2021-07-28 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE INFLAMED JAWS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION FOR ANTIBIOTICS AND A NASAL SPRAY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION H6 CODING HAS BEEN CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03235 | 2021-08-25 | 2021-07-06 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG CANCER. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICE. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO DEVELOP LUNG CANCER. PATIENT ALSO REPORTED HAVING SINUS PROBLEMS, HEADACHES, COUGH AND INTERMITTENT DIZZINESS. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT. THIS NOTIFICATION WAS REVIEWED BY THE PMS CLINICAL EXPERT DUE TO THE PATIENT REPORTING HAS LUNG CANCER LUNG, WAS REMOVED AND BELIEVES DEVICE TO BE CAUSE OF THIS. HAD BEEN ON MACHINE FOR TWO YEARS BEFORE. PATIENT ALSO REPORTED HAVING SINUS PROBLEMS, HEADACHES, COUGH, AND INTERMITTENT DIZZINESS. BASED ON INFORMATION AVAILABLE AT THIS TIME, THERE IS NOT ENOUGH EVIDENCE TO SUPPORT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE ALLEGED SERIOUS INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HEALTH EFFECT CLINICAL CODES WERE MISSED TO CAPTURE IN INITIAL REPORT AND WERE CAPTURED IN THIS REPORT AND OTHER SECTION H6 FIELDS WERE UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03204 | 2021-08-25 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP PARTICLES IN THEIR LUNGS AND DIFFICULTY BREATHING. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED PARTICLES IN THEIR LUNGS AND DIFFICULTY BREATHING. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED PARTICLES IN THEIR LUNGS AND DIFFICULTY BREATHING AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION,THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED AND THERE WAS MENTION OF VISUAL FINDINGS ODUR EMANATING FROM BLOWER BOX, EVIDENCE OF WATER INGRESS TO BLOWER AND PARTICLE CONTAMINATION IN BLOWER DIRT TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03234 | 2021-08-25 | 2021-07-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11 | DSX700T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP THROAT CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 04, 2024, AND SECTION H11 SHOULD BE REPORTED AS THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED TO EXPERIENCE THROAT CANCER. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM (PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR). SECTION B2 WAS MADE OTHER SERIOUS (PREVIOUSLY IT WAS BLANK). SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03205 | 2021-08-25 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A SINUS INFECTION. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 25, 2024, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A SINUS INFECTION. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT.THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03206 | 2021-08-25 | 2021-07-29 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION BIPAP AUTOSV H/C, DS | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS INC | DSX900H11C | DSX900H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A TUMOR IN HIS NOSE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03316 | 2021-08-25 | 2021-07-15 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/HT/CELL, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | CAX400T12C | CAX400T12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP LUNG NODULES AND TUMOR IN TRACHEA. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE REPORTED EVENT IS CURRENTLY UNKNOWN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. CORRECTED DATA: D8 WAS THIS DEVICE SERVICED BY A THIRD PARTY?: PREVIOUSLY REPORTED AS NO ; UPDATED TO UNKNOWN. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03488 | 2021-08-25 | 2020-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03397 | 2021-08-25 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | A CPAP DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS VISIBLE FOAM DEGRADATION AND THE UNIT SCRAPPED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03203 | 2021-08-25 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP AN UPPER RESPIRATORY INFECTION. THE PATIENT WAS PRESCRIBED AN ORAL STEROID IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO EXPERIENCE UPPER RESPIRATORY INFECTION AND SPOT ON THE TOP OF MOUTH. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT.THE REPORTED EVENT OF UPPER RESPIRATORY INFECTION AND SPOT ON THE TOP OF MOUTH AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION,THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION.AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03504 | 2021-08-25 | 2021-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT COULD NOT BE DUPLICATED. DURING THE EVALUATION, SERVICE REQUIRED CODES WERE OBSERVED IN THE DOWNLOADED EVENT LOGS. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND SYSTEM BOARD NEED REPLACED TO ADDRESS THE ISSUE. THE DEVICE WAS SCRAPPED PER MANUFACTURER PROTOCOL. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02928 | 2021-08-25 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO BIFLEX | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DS750HS | DS750HS |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING HEADACHES AND SINUS INFECTIONS RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE - THE PATIENT ALLEGED HEADACHES AND SINUS INFECTIONS RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03495 | 2021-08-25 | 2021-07-14 | I, n, j, u, r, y | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED CARBON DIOXIDE AND CONFUSION. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF HOSPITALIZATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED CARBON DIOXIDE AND CONFUSION. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF HOSPITALIZATION. AFTER RECEIVING FURTHER INFORMATION FROM THE PATIENT IT IS DETERMINED THAT THE INITIAL REPORT SHOULD HAVE BEEN FILED AS ADVERSE EVENT ONLY. SECTION B1 AND B2 HAS BEEN UPDATED/CORRECTED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03493 | 2021-08-25 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103504 | 2021-08-25 | 2021-08-03 | I, n, j, u, r, y | Voluntary report | DREAMSTATIONGO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | ON (B)(6) 2021 ER VISIT. DX PNEUMONIA. CHEST X-RAY. ON (B)(6) 2021 ICU AND HOSPITAL STAY. DX PNEUMONIA AND BLOOD CLOTS ON MY LUNGS. CHEST X-RAY/CAT-SCAN. ALMOST VENTED. DISCHARGED ON HOME OXYGEN. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-02157 | 2021-08-25 | 2021-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS FAILING A TEST STEP. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOULSY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR WAS FAILING A TEST STEP. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE DID NOT FAIL A TEST STEP, THE DEVICE WAS ALARMING FOR AN UNAMED "SERVICE REQUIRED CODE". THE DEVICE WAS RETURNED TO THE MANUFACTURE FOR EVALUATION AND A SERVICE REQUIRED CODE WAS OBSERVED IN THE DOWNLOADED EVENT LOG, HOWEVER, THIS CODE WOULD NOT CAUSE THE DEVICE TO MALFUNCTION OR DELIVER THERAPY AS DESIGNED. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03505 | 2021-08-25 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS NOT WORKING. THERE WAS NO PATIENT HARM OR INJURY REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING ALLEGING A VENTILATOR WAS NOT WORKING. THERE WAS NO PATIENT HARM OR INJURY REPORTED. THE DEVICE WAS EVALUATED BY A THIRD PARTY SERVICE CENTER AND THE CUSTOMER'S COMPLAINT WAS NOT CONFIRMED. THE DEVICE WAS FOUND TO OPERATE AND ALARM AS DESIGNED. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03506 | 2021-08-25 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS NOT WORKING. THERE WAS NO PATIENT HARM OR INJURY REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING ALLEGING A VENTILATOR WAS NOT WORKING. THERE WAS NO PATIENT HARM OR INJURY REPORTED. THE DEVICE WAS EVALUATED BY A THIRD PARTY SERVICE CENTER AND THE CUSTOMER'S COMPLAINT WAS NOT CONFIRMED. THE DEVICE WAS FOUND TO OPERATE AND ALARM AS DESIGNED. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03507 | 2021-08-25 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A MALFUNCTION OF A VENTILATOR'S TOUCHSCREEN. THE DEVICE WAS NOT IN PATIENT USE. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03116 | 2021-08-25 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A MALFUNCTION OF A VENTILATOR'S TOUCHSCREEN. THERE WAS NO PATIENT HARM OR INJURY REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03372 | 2021-08-26 | 2021-08-18 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP SEVERE ASTHMA. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE REPORTED EVENT IS CURRENTLY UNKNOWN. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON OCT 25, 2024 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES, HEADACHE AND ASTHMA RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY THE PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03402 | 2021-08-26 | 2021-07-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. REPEATED ATTEMPTS TO HAVE THE DEVICE RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103537 | 2021-08-26 | 2021-06-04 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONIC CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DREAM STATION | DSX500H11C |  | I BEGAN HAVING AN ISSUE WITH COUGHING WHEN TRYING TO LAY FLAT ON MY BACK. I HAD A NEW DRY COUGH. I WAS HAVING HEADACHES ALMOST EVERY MORNING. MY NAIL BEDS STARTED ACHING, TURNED YELLOW AND STARTED PEELING. MY HANDS AND FINGERS SWELLED. I DEVELOPED A RASH ON MY ABDOMEN. I WAS USING A PHILIPS RESPIRONICS CPAP DEVICE AND THE FIRST PART OF JULY WAS NOTIFIED OF A RECALL. I WAS ADVISED TO STOP USE OF THE CPAP IMMEDIATELY. I HAVE CONTACTED PHILIPS AND REGISTERED MY ACKNOWLEDGMENT OF THE RECALL. THEY HAVE OFFERED NO TIMELINE FOR REPLACEMENT OR REPAIR OF THIS ISSUE. I AM FRANKLY AFRAID TO CONTINUE USE OF THIS DEVICE. WITHOUT THE CPAP, I AM WAKING UP WITH A COUGH, OCCASIONAL HEADACHE AND CHEST PRESSURE. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103541 | 2021-08-26 | 2021-08-02 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500T11C |  |  | I USE A PHILIPS CPAP MACHINE, MODEL #DSX500T11C. I WAKE UP IN THE MIDDLE OF THE NIGHT WITH A HEADACHE SEVERAL TIMES A MONTH. THIS HAS ONLY HAPPENED IN RECENT MONTHS (25 YEARS USING CPAP MACHINES). I SEE THAT THIS AND MANY OTHER PHILIPS MACHINES HAVE BEEN RECALLED (I CHECKED SERIAL #, MINE HAS). I AM CONCERNED THAT THE HEADACHES STEM FROM THE TOXIC AND POTENTIALLY CARCINOGENIC PARTICLES FROM THE DISINTEGRATING FOAM THAT LED TO THE RECALL. NOBODY NOTIFIED ME OF THE RECALL: NOT PHILIPS, NOT MY ENT, NOT THE (B)(6), NOT THE MEDICAL EQUIPMENT SUPPLIER (B)(4). THIS IS OUTRAGEOUS WHEN THERE IS A POTENTIALLY DEADLY PROBLEM WITH THE MACHINES. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03510 | 2021-08-26 | 2021-07-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED REPEATED ATTEMPTS WERE UNSUCCESSFULLY MADE TO OBTAIN INFORMATION FOR A VENTILATOR THAT WAS ALARMING FOR A VENTILATOR INOPERATIVE ALARM CONDITION. THE DEVICE HAS SINCE BEEN RETURNED TO A THIRD PARTY SERVICE CENTER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. AN OXYGEN CONCENTRATION FAILURE WAS ALSO OBSERVED. THE DEVICE'S SYSTEM BOARD, OXYGEN BLENDING MODULE BOARD AND CAPACITOR WERE REPLACED TO ADDRESS THE ISSUES. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03511 | 2021-08-26 | 2021-08-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2110X15B | IN2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VARIATION IN THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE DELIVERY AND POSITIVE INSPIRATORY PRESSURE WAS OBSERVED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VARIATION IN THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE DELIVERY AND POSITIVE INSPIRATORY PRESSURE WAS OBSERVED. ADDITIONAL INFORMATION RECEIVED INDICATES THE VENTILATOR'S SOFTWARE WAS RE-LOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03512 | 2021-08-26 | 2021-08-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2110X15B | IN2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VARIATION IN THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE DELIVERY AND POSITIVE INSPIRATORY PRESSURE WAS OBSERVED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VARIATION IN THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE DELIVERY AND POSITIVE INSPIRATORY PRESSURE WAS OBSERVED. ADDITIONAL INFORMATION RECEIVED INDICATES THE VENTILATOR'S SOFTWARE WAS RE-LOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03211 | 2021-08-26 | 2021-07-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/WIFI,CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX500T12W | CAX500T12W |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. ADDITIONAL INFORMATION WAS RECEIVED AND ADDED TO THE REPORT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER ON 02/22/2023 FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED ON 03/20/2023. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND THE UNIT WAS SCRAPPED. SECTION D8, D9, H2, H3 AND H6 WERE UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03518 | 2021-08-26 | 2021-08-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03516 | 2021-08-26 | 2021-08-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, A VENTILATOR'S LCD SCREEN WAS BLANK. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03515 | 2021-08-26 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLANK. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE WAS EVALUATED AT THE MANUFACTURER'S SERVICE CENTER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03514 | 2021-08-26 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1032804 | 1032804 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03517 | 2021-08-26 | 2021-08-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-02246 | 2021-08-26 | 2021-04-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. |  | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03508 | 2021-08-26 | 2021-07-13 | I, n, j, u, r, y | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS OVERHEATING AT NIGHT AND SHUT DOWN. THE PATIENT EXPERIENCED DIFFICULTY BREATHING AND WAS HOSPITALIZED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON MAR 18, 2025 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A VENTILATOR WAS OVERHEATING AT NIGHT AND SHUT DOWN. THE PATIENT EXPERIENCED DIFFICULTY BREATHING AND WAS HOSPITALIZED. ADDITIONAL INFORMATION WAS RECEIVED REGARDING THE ALLEGATION OF THE PATIENT'S DEATH DUE TO COVID AND HEART FAILURE IN THE HOSPITAL. THE SECTION D9, E1 WAS UPDATED AND SECTIONS B2, H1,H2, H3 AND H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B2 WAS CORRECTED TO DEATH AND HOSPITALIZATION. (PREVIOUSLY, IT WAS BLANK.). SECTION H1 WAS CHANGED FROM MALFUNCTION TO DEATH. SECTION H6 HEALTH EFFECTS: CLINICAL CODES AND HEALTH EFFECTS - IMPACT CODE WAS UPDATED. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. DURING THE INVESTIGATION, THE MANUFACTURER VISUALLY INSPECTED THE VENTILATOR AND DID NOT DETECT ANY HARDWARE DEFECTS ON THE DEVICE. THE DEVICE WAS RETURNED WITHOUT A DETACHABLE BATTERY AND THE INLET FILTER. THE VENTILATOR PASSED THE POSTTEST'S APPLICABLE TEST STEPS ON THE TRILOGY MULTIFUNCTION TEST STATION. THE DATE OF THE INCIDENT WAS REPORTED ON (B)(6) 2021, THERE WERE NOT ANY CRITICAL ALARMS OR ERROR CODES. ON THE DATE OF INCIDENT ((B)(6) 2021) THERE WERE ONLY 2 LOW PIP (PEAK INSPIRATORY PRESSURE) AND 2 LOW PEP (POSITIVE EXPIRATORY PRESSURE) PATIENT CIRCUIT ALARMS. AFTER THESE PATIENT CIRCUIT ALARMS, THE RESET KEY WAS PRESSED TO SILENCE THE ALARMS. THE DEVICE WAS STILL IN USE UNTIL (B)(6) 2022 WITHOUT ANY CRITICAL ERROR CODES. THE TRILOGY DEVICE WAS ALLOWED TO WARM UP ON THE BENCH TOP. DURING THIS TIME THERE WERE NOT ALARMS OR ERROR CODES. THE DEVICE WAS TESTED ON THE TRILOGY MULTI-FUNCTION TEST STATION AND PASSED THE APPLICABLE TEST STEPS. TEST STEPS WERE OBSERVED AS FAILURES BUT ARE RELATED TO PRODUCTION ENVIRONMENT TESTING PARAMETERS. THE DEVICE WAS RUN OVERNIGHT ON THE BENCH TOP WITHOUT ANY ALARMS, FAILURES OR OVERHEATING. THE DEVICE WAS OPENED AND DID NOT DETECT ANY SIGNS OF MELTED PLASTIC INSIDE OF THE DEVICE. TECHNICIAN INSPECTED THE AIRPATH FOAM AND FOUND IT TO BE FIRMLY SECURED, WITHOUT SIGNS OF DELAMINATION OR DEGRADATION. THE MANUFACTURER WAS NOT ABLE TO CONFIRM THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO CRITICAL ERROR CODES. THE MANUFACTURER CONCLUDES THAT THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE MANUFACTURER UNABLE TO DUPLICATE ANY FAILURE WITH THE VENTILATOR AND HAS DETERMINED THAT THE DEVICE FUNCTIONS AS DESIGNED. | Y | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03513 | 2021-08-26 | 2021-08-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2110X15B | IN2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VARIATION IN THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE DELIVERY AND POSITIVE INSPIRATORY PRESSURE WAS OBSERVED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VARIATION IN THE VENTILATOR'S POSITIVE END EXPIRATORY PRESSURE DELIVERY AND POSITIVE INSPIRATORY PRESSURE WAS OBSERVED. ADDITIONAL INFORMATION RECEIVED INDICATES THE VENTILATOR'S SOFTWARE WAS RE-LOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03530 | 2021-08-26 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP AUTOSV H/HT/C, DS | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS INC | DSX900T11C | DSX900T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE(BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03400 | 2021-08-26 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP/HUM/HT/CELL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES, DIFFICULTY BREATHING/SHORT OF BREATH AND HEADACHES RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED SEEING PARTICLES IN THE DEVICE, DIFFICULTY BREATHING/SHORT OF BREATH AND HEADACHES. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION(S) B5, D1, D4 AND H4 (PREVIOUSLY WAS NOT INCLUDED) WAS CORRECTED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03232 | 2021-08-26 | 2021-08-12 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, JP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | RJPX500S16 | RJPX500S16 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA. THERE WAS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THE TIME. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER COMPLETED AN INTERNAL VISUAL INSPECTION OF THE DEVICE. THE MANUFACTURER FOUND INDETERMINATE BLACK CONTAMINATION ON AIR OUTLET AND BLOWER BOX INLET, HAIR-LIKE CONTAMINANTS IN BOTTOM ENCLOSURE, WHITE DUST-LIKE CONTAMINATION AROUND FLOW SENSOR SEAL, ON BLOWER, AND WITHIN BLOWER BOX, BLACK HAIR-LIKE CONTAMINANT ON BLOWER OUTLET. THE MANUFACTURER ALSO FOUND BLACK RESIDUE CONSISTENT WITH FINDINGS OF KERATIN ON BLOWER OUTLET AND BLOWER BOX OUTLET. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND ERROR CODE E-4. THE MANUFACTURER CONCLUDE THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. THE MANUFACTURER ALSO CONFIRMS THE PRESENCE OF MULTIPLE CONTAMINANTS (CONSISTENT WITH HAIR, DUST, KERATIN) WITHIN THE AIR PATH, WHICH IS LIKELY OF EXTERNAL ORIGINAL AND NOT CONSISTENT WITH ORIGINATING FROM A DEVICE FAILURE. IN THE INITIAL REPORT CODE FOR RHINITIS WAS NOT MENTIONED WHICH HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103515 | 2021-08-26 | 2019-12-29 | I, n, j, u, r, y | Voluntary report | BIPAPAUTOSVADSYS0NE60SRSDOM | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | PHILIPS / RESPIRONICS, INC. |  |  |  | MY FATHER WHO IS (B)(6), NEVER SMOKED, NEVER DRANK, HAS USED ASV PAP FOR 4 YEARS AND ABOUT 9 MONTHS. HE USED A SO CLEAN MACHINE AND HEATED HUMIDITY DAILY. HE WAS HOSPITALIZED WITH EMESIS, WEAKNESS AND FOUND TO HAVE AN ABCESS IN HIS ABDOMEN. HE HAD DRAINS PLACED. HOSPITALIZED (B)(6) 2019 TO (B)(6) 2020 AND 19 DAYS TCU DISCHARGED HOME WITH DRAINS IN ABDOMEN. FELL DUE TO WEAKNESS AGAIN, HOSPITALIZED/(B)(6) 2021, ONE DRAIN REMOVED. TOLD TO FLUSH DRAIN TWICE DAILY. ON (B)(6) 2021 SEEN BY (B)(6) AFTER BIOPSY DONE; ON (B)(6) 2021 TOLD RESULTS OF BIOPSY, MUCINOUS ADENOCARCINOMA OF EITHER APPENDIX OR COLON. SEEN BY (B)(6) GASTROENTEROLOGY (B)(6) 2020 ASSESSMENT: MUCINOUS ADENOCARCINOMA (C80.1), INTRA-ABDOMINAL ABSCESS (K65.1), CHRONIC CONSTIPATION (K59.09) DESCRIPTION: IMPRESSION: MUCINOUS ADENOCARCINOMA COMPLICATED BY INTRA-ABDOMINAL ABSCESS. UNCLEAR PRIMARY SOURCE. SEEN BY (B)(6) ONCOLOGY SURGEON, DIDN'T RECOMMEND SURGERY, TOO RISKY, COULD BE FATAL. CANCER IS STAGE 4 DUE TO RUPTURE OF ABSCESS. (B)(6) ONCOLOGIST, DIDN'T RECOMMEND CHEMO, DUE TO POOR RESPONSE FOR THIS KIND OF CANCER AND WOULD BE LOW QUALITY OF LIFE AND IS SLOW GROWING. GAVE MY FATHER 1-3 YEAR LIFE EXPECTANCY. SECOND OPINION AT (B)(6), GAVE MY FATHER 3 MONTHS TO ONE YEAR. ON (B)(6) 2020 3RD HOSPITALIZATION (FATHER STILL USING SO CLEAN AND HEATED HUMIDITY.) FATHER WOKE UP THROWING UP. FEVER. ADMITTED TO (B)(6) WITH PNEUMONIA. INFLUENZA A AND B NEGATIVE. ON (B)(6) 2020 PRESENTED AT TUMOR CONFERENCE AT (B)(6). AT DISCHARGE TOLD MAY BE 1 YEAR LIFE EXPECTANCY, TUMOR BIGGER AND NOW IN OTHER AREA OF ABDOMEN. DUE TO PNEUMONIA, AND MYSELF BEING A PULMONARY NURSE. AN OVERNIGHT OXIMETRY WAS DONE TO SEE IF HE NEEDED OXYGEN. THE RESULTS WERE VERY ABNORMAL AND OXYGEN WAS ORDERED. DUE TO THE FACT IT WAS EXTREMELY DIFFICULT TO PULL APART THE OXYGEN ADAPTOR TO USE THE SO CLEAN MACHINE, WE QUIT USING THE SO CLEAN MACHINE AND JUST USED ASV AND OXYGEN SINCE THEN (WASHED MASK, TUBING AND HUMIDIFIER BY HAND). HE USED HEATED HUMIDITY SET AT 2 OR 3 OUT OF 5. MY FATHER IS DOING FAIRLY WELL WITHOUT ABDOMINAL PAIN OR RECURRENCE OF EMESIS. HIS APPETITE IS GOOD. DUE TO HIS NEED FOR AN ADVANCED ASV MACHINE HE HAS CONTINUED TO USE HIS ASV MACHINE SINCE THE RECALL WITHOUT HEATED HUMIDITY AND WITH A INLINE FILTER. HE CONTINUES TO USE OXYGEN WITH HIS ASV PAP MACHINE. SINCE HIS INITIAL DIAGNOSIS, HE HAS SURVIVED A COVID HOSPITALIZATION (B)(6) 2020 AND A NEW DIAGNOSIS OF CONGESTIVE HEART FAILURE WINTER OF 2021 AND HAS CONTINUED TO BECOME STRONGER. NOW NEEDING CARES 45 MINUTES DAILY(DOWN FROM 4 TIMES DAILY CARE) AND 2 HOURS AN ADDITIONAL 3 TIMES PER WEEK RESIDING IN HIS OWN APARTMENT CONDO. WE HAVE AN ORDER FOR A NEW ASV MACHINE BUT IT IS NOT CURRENTLY AVAILABLE(THOUGH NOT TECHNICALLY COVERED UNTIL (B)(6) 2021 AS THAT IS WHEN HIS MACHINE WILL BE 5 YEARS OLD) AT (B)(6). NOR AT (B)(6) TO MY KNOWLEDGE AS THEY DIDN'T RESPOND TO MY REQUEST. THE REASON I DON'T WANT THIS SUBMITTED WITH MY NAME OR MY FATHER'S ATTACHED TO PHILLIPS RESPIRONICS, IS BECAUSE I HAVE WORKED AT A LUNG AND SLEEP CLINIC FOR THE LAST 15 YEARS THAT DISPENSES PHILLIPS RESPIRONICS MACHINES. I AM AFRAID FOR MY JOB IF THEY LEARN THAT I HAVE REPORTED THIS TO THE FDA. THEY HAVE PHONE MEETINGS WITH STAFF FROM PHILLIPS RESPIRONICS. MY COMPANY HAS BEEN A MAJOR DISTRIBUTOR, UNTIL THE RECALL OF ONLY PHILLIPS RESPIRONICS PAP MACHINES FOR (B)(6). THEY STILL ARE DISTRIBUTING MACHINES BUT ONLY THE NEW NONAFFECTED DEVICES. I HAVE A PHOTO OF THE SERIAL NUMBER, BUT AM HAVING DIFFICULTY UPLOADING IT; (B)(6) 2020 WAS DATE RECEIVED THE DIAGNOSIS. BIOPSY WAS A DAY OR TWO BEFORE. MUCINOUS ADENOCARCINOMA STAGE 4, COMPLEX OBSTRUCTIVE SLEEP APNEA, MACULAR DEGENERATION, NEW DIAGNOSIS CHF WINTER 2021. MY FATHER HAS BEEN VERY HEALTHY HIS WHOLE LIFE, HE WAS A NONSMOKER AND DIDN'T DRINK ALCOHOL. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103516 | 2021-08-26 | 2021-08-23 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | MY MACHINE, WHICH UNFORTUNATELY DIDN'T WORK PROPERLY FROM THE DAY I BROUGHT IT HOME, FAILS ENTIRELY NOW. IT DOESN'T RECEIVE THE POWER IT NEEDS TO RUN. IT HAS NEVER ACCURATELY REPORTED THE HOURS I USE IT. NEVER. I TOOK IT TO THE SUPPLIER TWICE IN (B)(6) AND THEY SAID IT DIDN'T MATTER THAT IT WOULDN'T TRACK MY HOURS OF USAGE. NOW, IT WILL NOT TURN ON. MY EMAIL ADDRESS IS (B)(6). FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103518 | 2021-08-26 | 2021-06-28 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAM STATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | NOT SURE HOW IT HAPPENED BUT CHECKED FILTER AND IT MELTED TO ITSELF. I HAVE PICTURES AND SAVED IT. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103523 | 2021-08-26 | 2021-08-09 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500T11C | DREAMSTNAUTOCPAPHUMHTCELL DOM |  | OVERHEATING OF CPAP MACHINE; I HAVE ONE OF THE PHILIPS RESPIRONICS DREAMSTATION MACHINES THAT IS AFFECTED BY THE 2 VENTILATOR RECALLS. ON AUGUST 9, AT APPROXIMATELY 4:00 PM, MY APARTMENT'S AIR CONDITIONING SYSTEM FAILED. JUST 10 MINUTES LATER, I SMELLED A BURNING ODOR IN MY CPAP MASK. I CHECKED THE BACTERIA FILTER I HAD INSTALLED THAT MORNING AND NOTICED A FEW BLACK PARTICLES IN IT AND IN THE EMPTY HUMIDIFIER CHAMBER. IN ADDITION, I CHECKED MY THERMOMETER THAT I KEEP AT MY BEDSIDE AND NOTICED MY BEDROOM'S TEMPERATURE WAS 93 DEGREES FAHRENHEIT. THE DREAMSTATION MACHINE WAS WARM, THOUGH NOT HOT, TO THE TOUCH. I IMMEDIATELY UNPLUGGED THE MACHINE AND STARTED USING A CPAP FROM ANOTHER MANUFACTURER, WHICH I HAD JUST ORDERED THAT WEEK I REACHED OUT TO PHILIPS RESPIRONICS MULTIPLE TIMES REGARDING THE RECALL AS WELL AS THIS INCIDENT, AND HAD REGISTERED MY MACHINE BACK IN JUNE ON THE DAY THE RECALL WAS ANNOUNCED. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103532 | 2021-08-26 | 2020-10-19 | I, n, j, u, r, y | Voluntary report | RESPIRONICS DREAM STATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500T11C |  |  | I WAS USING MY CPAP AND I HAD HEADACHES, DIZZINESS, TROUBLE BREATHING (CONTINUES EVEN AFTER) AND OTHER HEALTH CHALLENGES. I AM ONLY USING THIS TIME PERIOD AS AN APPROXIMATION. I WAS DIAGNOSED WITH METHEMOGLOBINEMIA. I HAD METHEMOGLOBINEMIA UNTIL I STOPPED MY CPAP. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03519 | 2021-08-26 | 2021-06-30 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION BIPAP AVAPS30 H, DS | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | DSX1130H11 | DSX1130H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP RESTRICTIVE LUNG DISEASE. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03456 | 2021-08-26 | 2021-07-30 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP AN ISSUE WITH THEIR PARATHYROID GLAND. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP AN ISSUE WITH THEIR PARATHYROID GLAND. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03162 | 2021-08-26 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP NODULES IN THEIR LUNGS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE 15 ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. SECTIONS B1, H1 AND H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM (BOTH THE PRODUCT PROBLEM AND ADVERSE EVENTS AND WAS CHECKED IN THE PREVIOUS MDR.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECTS - CLINICAL CODES AND HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03163 | 2021-08-26 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A COUGH AND IRRITATED EYES. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A COUGH AND IRRITATED EYES. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. IN THE INITIAL REPORT SECTION B5 WAS NOT UPDATED, THE CORRECT B5 SHOULD BE - THE PATIENT HAS ALLEGED PARTICLES IN THE DEVICE, CHEST COMPRESSION, IRRITATION IN EYE, COUGH, TAKING ALLERGY MEDICATION AND USING INHALER. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO EVIDENCE OF CONTAMINATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF CONTAMINATION IN BLOWER BOX, CONTAMINATION ON BLOWER MOTOR, WATER INGRESS IN THE BLOWER BOX. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND 3 INSTANCES OF E-4 ON THE ERROR LOG. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH CONTAMINATION IN BLOWER BOX, CONTAMINATION ON BLOWER MOTOR, WATER INGRESS IN THE BLOWER BOX. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION B1,B2,H1,D9,G3,H6 HAS BEEN UPDATED/CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03498 | 2021-08-26 | 2021-07-29 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | 999998 | 999998 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO REQUIRE LASER SURGERY TO THEIR TONGUE DUE TO MODERATE CELL CHANGES. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03523 | 2021-08-26 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED DEVICE HAS STRANGE ODOUR, NASAL/THROAT IRRITATION OR SORENESS, VISUALIZATION OF BLACK PARTICLES IN HIS FLUID. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. SECTION H6 UPDATED IN THIS REPORT. CORRECTED INFORMATION WAS PROVIDED IN H.6. SECTION (IN PREVIOUS MDR SORE THROAT WAS NOT INCLUDED). | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03212 | 2021-08-26 | 2021-07-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUMID, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | CAX700H12 | CAX700H12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03478 | 2021-08-26 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11C | DSX400H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | IN SECTION H6 HEALTH EFFECT - CLINICAL CODE MISSED TO CAPTURE IN PREVIOUS REPORTS, IT HAS BEEN UPDATED OR CORRECTED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03479 | 2021-08-26 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN AIRWAY FROM DEVICE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PARTICLES IN AIRWAY FROM DEVICE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03477 | 2021-08-26 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11C | DSX400H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING BLACK PARTICLES ARE IN TUBING AND DEVICE HAS STRANGE ODOR AND RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03208 | 2021-08-26 | 2021-07-02 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11C | DSX400H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP FLUID IN THEIR LUNGS. THE PATIENT REQUIRED REMOVAL OF THE FLUID AS A RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT HAS ALLEGED TO DEVELOP FLUID IN LUNGS. THE PATIENT REQUIRED REMOVAL OF THE FLUID AS A RESPONSE TO THE REPORTED EVENT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURER CONCLUDES NO FURTHER ACTION IS NECESSARY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. VISUAL INSPECTION CONFIRMED POSITIVE DEPOSIT OF FOAM PARTICULATE ON THE FIT TOOL. PERFORMANCE TESTING/VERIFICATION OF COMPLAINT CONFIRMED THE EVIDENCE OF FOAM DEGRADATION ASSOCIATED TO THIS ALLEGATION. REVIEW OF SECONDARY FINDINGS SHOWED THE EVIDENCE OF LIQUID INGRESS. PIL WAS ABLE TO CONFIRM THE COMPLAINT OR ADDRESS THE SYMPTOMS SPECIFIED. THE DEVICE WAS HOOKED UP TO A KNOWN GOOD POWER SUPPLY AND POWER CORD, AIRFLOW WAS VERIFIED TO BE OPERATING CORRECTLY. THE ERROR LOG SHOWED THAT THERE WERE NO INSTANCE OF ERRORS ON THE DEVICE. PIL WAS ABLE TO CONFIRM THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM IN THE DEVICE. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03531 | 2021-08-26 | 2021-01-01 | I, n, j, u, r, y | Manufacturer report | RENTAL-DREAMSTATION AUTO CPAP, GB | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | RGBX500S15 | RGBX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. DESPITE MULTIPLE ATTEMPTS, THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. CODES IN SECTION H6, SECTION E1 WERE UPDATED OR CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03529 | 2021-08-26 | 2021-07-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | IN PREVIOUS REPORT SECTION D9 WAS INCORRECTLY CAPTURED AS : 02/23/2023 BUT IT SHOULD BE REPORTED AS: 02/21/2023 SECTION D9 HAS BEEN CORRECTED/UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03476 | 2021-08-26 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11C | DSX400H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK PARTICLES IN THE MACHINE, DOES NOT FEEL GOOD RESPIRATORY RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. CORRECTED INFORMATION WAS PROVIDED IN SECTIONS G1, H6. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03209 | 2021-08-26 | 2021-07-10 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP BREATHING PROBLEMS. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP BREATHING PROBLEMS. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. SECTION B2 HAS BEEN CORRECTED / UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03210 | 2021-08-26 | 2021-07-06 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP NASAL AND THROAT INFECTIONS. THE PATIENT WAS PRESCRIBED ANTIBIOTICS IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP NASAL AND THROAT INFECTIONS. THE PATIENT WAS PRESCRIBED ANTIBIOTICS IN RESPONSE TO THE REPORTED EVENT. IN THE INITIAL REPORT, H6 SECTION WAS NOT UPDATED. WHICH IS UPDATED IN THIS REPORT. THE DEVICE ALONG WITH A HUMIDIFIER WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE AND HUMIDIFER WERE COMPLETED BY THE MANUFACTURER. THE MANUFACTURE FOUND EVIDENCE OF CONTAMINATION ON THE BOTTOM OF BASE AND HUMIDIFIER, BETWEEN THE UNITS, ALONG THE FLIP-LID HINGE, AND IN THE BOTTOM OF THE HUMIDIFIER CHAMBER. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER.THE MANUFACTURER FOUND EVIDENCE OF DUST ON AND IN THE BLOWER, A DUSTY SMOKE (NON-TOBACCO) ODOR, UNIDENTIFIED CONTAMINATION ON THE INTERIOR OF THE HUMIDIFIER BOTTOM COVER. THE MANUFACTURER FOUND EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND 1 INSTANCE OF E-153 ON THE ERROR LOG. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT CONTAMINATION ON THE BOTTOM OF BASE AND HUMIDIFIER, BETWEEN THE UNITS, ALONG THE FLIP-LID HINGE, AND IN THE BOTTOM OF THE HUMIDIFIER CHAMBER, DUST ON AND IN THE BLOWER, A DUSTY SMOKE (NON-TOBACCO) ODOR, UNIDENTIFIED CONTAMINATION ON THE INTERIOR OF THE HUMIDIFIER BOTTOM COVER. THE MANUFACTURER CONFIRMED THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION D9,D10,G3,H6 HAS BEEN UPDATED/CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03207 | 2021-08-26 | 2021-01-14 | D, e, a, t, h | Manufacturer report | DREAMSTATION BIPAP AUTOSV H/C, DS | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS INC | DSX600T11 | DSX600T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT DEATH. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED A BIPAP DEVICE'S SOUND ABATEMENT FOAM ALLEGEDLY BECAME DEGRADED AND CAUSED A PATIENT DEATH. THE REPORTER OF THE EVENT REPORTED THE PATIENT HAD KIDNEY FAILURE, LUNG BLEEDING AND WAS HOSPITALIZED. DESPITE MULTIPLE ATTEMPTS, THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103565 | 2021-08-27 |  | I, n, j, u, r, y | Voluntary report | RESPIRONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | USE OF THE PHILIPS RESPIRONICS DREAMSTATION CPAP MACHINE WAS CAUSING CHEST CONGESTION AND TIGHTNESS AND I WAS WAKING UP WITH HEADACHES. I WAS NOT AWARE OF THE RECALL UNTIL A FRIEND LET ME KNOW A COUPLE OF DAYS AGO. I HAVE STOPPED USING MY CPAP MACHINE AND WILL FOLLOW UP WITH MY PRIMARY PHYSICIAN FOR AN APPOINTMENT FOR AN EVALUATION. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103566 | 2021-08-27 | 2021-07-07 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | RTDSX500 |  |  | BREATHING AND LUNG PROBLEMS. EXPERIENCE SHORTNESS OF BREATH AND FATIGUE. DR SENT TO GET CT FOUND OUT I HAVE BLOOD CLOTS. FOUND OUT TODAY THAT MY CPAP HAS BEEN RECALLED. PHILIPS. THINKING THE CAP COULD BE THE CAUSE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103569 | 2021-08-27 | 2021-03-27 | I, n, j, u, r, y | Voluntary report | CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | H27170233A9DC |  |  | I REPORTED TO MY DOCTOR AND THE COMPANY THAT GAVE ME THE DEVICE THAT I KEPT HAVING THIS ELECTRIC SHOCKING IN MY MOUTH THRUSTING ME OUT OF MY SLEEP, SCARING THE HELL OUT OF ME AND MY GRANDDAUGHTER. I WAS TOLD THAT THE MACHINE SHOULDN'T DO THAT, IT WASN'T EVERY SINGLE DAY, BUT WHEN IT DID HAPPEN IT HURT AND VERY SCARY. I HAVE A BAD COUGH, NASAL DRIP GOTTEN MUCH WORSE. CALLING MY DOCTOR TO GET A FULL WORK UP, NOW I'M REALLY SCARED. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03550 | 2021-08-27 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED DUE TO PHYSICAL DAMAGE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
MW5103574 | 2021-08-27 | 2021-06-30 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | RECENTLY I DEVELOPED A BAD, PRODUCTIVE UNCONTROLLABLE COUGH. AFTER READING ABOUT THE RECALL NOTICE IN THE NY TIMES (I NEVER RECEIVED A NOTICE OF THE RECALL FROM MANUFACTURER OR MEDICAL SUPPLY COMPANY) I STOPPED USING THE CPAP MACHINE AND THE COUGH IS NOW IMPROVING. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103580 | 2021-08-27 | 2021-05-10 | I, n, j, u, r, y | Voluntary report | DREAMSTATION HUM CORE PACK DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | FOR THE ABOUT THE LAST 3 MONTHS, I KEEP FEELING LIKE MY NOSE WAS BURNING, WOULD WAKE UP WITH A HEADACHE AND BOTH NOSTRILS STOPPED UP. FEW TIMES, I PULLED IT OFF IN THE MIDDLE OF THE NIGHT BECAUSE I FELT LIKE I COULDN'T BREATH AND MY NOSE WAS BURNING BADLY. WAS FEELING A SHORTNESS OF BREATH DURING THE DAY, SO I WOULD USE MY INHALER. I KEPT FEELING A FEELING OF SHORTNESS OR SOMETHING. NOW THAT I STOPPED WEARING IT SINCE I HEARD OF THE RECALL, I HAVEN'T HAD ANY OF THE SYMPTOMS ANYMORE. PHILIPS DREAMSTATION SLEEP APNEA. |  | Y | N | I |  | 2021-08
2518422-2021-03548 | 2021-08-27 | 2021-08-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03431 | 2021-08-27 | 2021-08-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03427 | 2021-08-27 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE TO POWER ON WAS OBSERVED. THE DEVICE'S POWER SUPPLY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03432 | 2021-08-27 | 2021-08-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03549 | 2021-08-27 | 2021-08-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. | IT WAS INITIALLY REPORTED, THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. THE DEVICE'S SYSTEM BOARD WAS ALSO REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
MW5103549 | 2021-08-27 | 2021-01-22 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500T11W |  |  | I WOKE UP BECAUSE MY CPAP MACHINE WAS GIVING ME A STRANGE TASTE/FEELING. I CHECKED THE HOSE AND WATER CHAMBER, THEY SEEMED FINE. THEN I CHECKED THE FILTER WHICH HAD BEEN CHANGED THE PREVIOUS FRIDAY (I CHANGE THE FILTER EVERY OTHER WEEK) AND IT WAS BLACK, LIKE SOOT. THE PERMANENT FILTER WAS ALSO COMPLETELY BLACK AND IT COULD NOT BE CLEANED. I HAD JUST RELIED THE FILTER, SO I DIDN'T REQUEST A NEW ONE, I BOUGHT THEM FROM (B)(6) AND PAID FOR THEM, THINKING I HAD DONE SOMETHING WRONG. I JUST FOUND OUT FROM MY PULMONOLOGIST THAT THERE IS A RECALL ON MY PHILIP'S DREAMSTATION CPAP MACHINE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03551 | 2021-08-27 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03532 | 2021-08-27 | 2021-07-13 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP BREAST CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03466 | 2021-08-27 | 2021-07-30 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP IRRITATION IN THE LUNG AIRWAYS. THE PATIENT HAS RECEIVED MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTION FOR NASAL SPRAY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03470 | 2021-08-27 | 2021-07-30 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE NASAL INFLAMMATION. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTION FOR SALINE, NASAL DROPS AND NASAL CREAMS. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03471 | 2021-08-27 | 2021-07-30 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP A POLYP AND SHORTNESS OF BREATH. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTION FOR PAIN KILLING PILLS. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCING HEADACHES, HIGH BLOOD PRESSURE, PAIN IN THE TEETH, SHORTNESS OF BREATH, AGGRESSIVE,LACK OF MOOD,TO DEVELOP A POLLYP AND SHORTNESS OF BREATH. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTION PILLS FOR PAIN KILLING AND BLOOD PRESSURE. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FOLLOW UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03468 | 2021-08-27 | 2021-07-30 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE SHORTNESS OF BREATH. THE MEDICAL INTERVENTION RECEIVED IN RESPONSE TO THIS EVENT IS CURRENTLY UNKNOWN. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03467 | 2021-08-27 | 2021-07-30 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP UPPER RESPIRATORY TRACT INFECTION. THE PATIENT HAS RECEIVED MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTION FOR INHALERS. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03465 | 2021-08-27 | 2020-02-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE SHORTNESS OF BREATH. THE PATIENT HAS RECEIVED MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTION FOR ANTIBIOTICS AND STEROIDS. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03464 | 2021-08-27 | 2021-07-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE DIFFICULTY IN BREATHING. THE PATIENT WAS REPORTEDLY HOSPITALIZED IN RESPONSE TO THE MEDICAL EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE DIFFICULTY IN BREATHING. THE PATIENT WAS REPORTEDLY HOSPITALIZED IN RESPONSE TO THE MEDICAL EVENT. THE REPORTED EVENT AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE.BASED ON INFORMATION AVAILABLE AT THE TIME OF THIS REVIEW, THERE IS INSUFFICIENT EVIDENCE TO PRONOUNCE A SERIOUS INJURY. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03448 | 2021-08-27 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE TINNITUS AND A SINUS INFECTION. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTIONS FOR ANTIBIOTICS. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 04 , 2024 AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE TINNITUS AND A SINUS INFECTION. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTIONS FOR ANTIBIOTICS.THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03447 | 2021-08-27 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUMID, AHP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AHX700H11 | AHX700H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE KIDNEY CANCER. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE REPORTED EVENT IS CURRENTLY UNKNOWN. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO HAVE KIDNEY CANCER, EXPERIENCING NASAL CONGESTION, LIGHTHEADED. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENT OF KIDNEY CANCER AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2 HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03542 | 2021-08-27 | 2020-01-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03565 | 2021-08-27 | 2021-07-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03494 | 2021-08-27 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DNX500H11C | DNX500H11C |  | THE MANUFACTURER WAS MADE AWARE OF AN ALLEGATION FROM A USER THAT HER CPAP DEVICE HAD A BURNING ODOR WHICH PROMPTED HER TO GO TO A HOSPITAL AND RECEIVED OXYGEN AND BREATHING TREATMENTS. THE USER WAS LATER DISCHARGED HOME. THERE WAS NO REPORT OF MELTING, WARPING OR VOIDS TO THE CPAP DEVICE. THE DEVICE WAS DISCARDED BY THE SUPPLIER. THE MANUFACTURER IS UNABLE TO CONFIRM THE ALLEGATION THE DEVICE WAS EMITTING A BURNING ODOR. THERE WAS NO REPORT OF SERIOUS OR PERMANENT INJURY. BASED ON THE AVAILABLE INFORMATION, NO FURTHER ACTION IS NECESSARY. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER WAS MADE AWARE OF AN ALLEGATION FROM A USER THAT HER CPAP DEVICE HAD A BURNING ODOR WHICH PROMPTED HER TO GO TO A HOSPITAL AND RECEIVED OXYGEN AND BREATHING TREATMENTS. THE USER WAS LATER DISCHARGED HOME. THERE WAS NO REPORT OF MELTING, WARPING OR VOIDS TO THE CPAP DEVICE. THE DEVICE WAS DISCARDED BY THE SUPPLIER. THE MANUFACTURER IS UNABLE TO CONFIRM THE ALLEGATION THE DEVICE WAS EMITTING A BURNING ODOR. THERE WAS NO REPORT OF SERIOUS OR PERMANENT INJURY. BASED ON THE AVAILABLE INFORMATION, NO FURTHER ACTION IS NECESSARY. THE REPORTED EVENT AND SEVERITY WERE REVIEWED BY CLINICAL EXPERT DUE TO PATIENT REPORTING OF HOSPITALIZATION BECAUSE OF FOUL SMELL FROM THE MACHINE AND TUBES WERE BURNED AND MELTED. PATIENT RECEIVED OXYGEN AND BREATHING TREATMENT. THESE EVENT MAY BE RELATED TO A PRODUCT PROBLEM AND/OR USER ERROR. BASE ON THE AVAILABLE INFORMATION AT THE TIME OF THE REVIEW, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03539 | 2021-08-27 | 2018-08-15 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC |  |  |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102027) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP PULMONARY HYPERTENSION AND CONGESTIVE HEART FAILURE. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP PULMONARY HYPERTENSION AND CONGESTIVE HEART FAILURE. THESE EVENTS WERE REVIEWED BY THE PMS CLINICAL EXPERT AND WERE ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03564 | 2021-08-27 | 2021-07-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03528 | 2021-08-27 | 2021-08-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03526 | 2021-08-27 | 2021-08-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03540 | 2021-08-27 | 2021-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 04, 2024, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM.THIS RECORD WAS REVIEWED BY THE PMS CLINICAL EXPERT DUE TO THE PATIENT'S ATTORNEY REPORTING ALLEGATIONS OF MUSCULOSKELETAL ARTHRITIS BEING TREATED WITH PREDNISOLONE AND METHOTREXATE. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. IN ADDITION, THE PATIENT'S ATTORNEY REPORTED ALLEGATIONS OF NAUSEA AND INFECTION. THE PATIENT DID NOT REPORT TO RECEIVE MEDICAL INTERVENTION.. SECTION B1 WAS CORRECTED FOR BOTH ADVERSE EVENT AND PRODUCT PROBLEM. ( PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION. TO SERIOUS INJURY. SECTION H6 WAS CORRECTED AND UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03543 | 2021-08-27 | 2021-07-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION B5: DESCRIBE AN EVENT OR PROBLEM CAPTURED INCORRECTLY IN THE PREVIOUS REPORT THAT SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED BEING LIGHT-HEADED IN THE MORNING. DIZZINESS, DRY COUGH, SHORTNESS OF BREATH, CHEST TIGHTNESS, AND FEELING LUMPS IN THE THROAT. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTIONS B5, B7, D4 AND H6 HAVE BEEN UPDATED AND CORRECTED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03525 | 2021-08-27 | 2021-08-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03541 | 2021-08-27 | 2021-07-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED MILD HEADACHES, SPECIAL BREATHING DIFFICULTIES, AND SHORTNESS OF BREATH. HE HAS BEEN TO A DOCTOR, AND HE HAS BEEN PRESCRIBED ALLERGY PILLS AS WELL AS MEANS OF INHALATION, BUT NOTHING HAS WORKED. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTION D4 UNIQUE IDENTIFIER NUMBER WAS CAPTURED INCORRECTLY IN THE PREVIOUS REPORT. IT HAS BEEN UPDATED AND CORRECTED IN THIS REPORT. SECTION H6 HEALTH EFFECTS - CLINICAL CODE HAS BEEN UPDATED OR CORRECTED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03527 | 2021-08-27 | 2021-08-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03524 | 2021-08-27 | 2021-08-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03522 | 2021-08-27 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES IN TUBING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. CORRECTED INFORMATION PROVIDED IN B5, THE INFORMATION (THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES IN TUBING RELATED TO A CPAP DEVICE) WAS NOT INCLUDED IN INITIAL REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03520 | 2021-08-27 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED ON THIS DEVICE IN MDR 2518422-2021-02458. THIS REPORT WAS SUBMITTED AS A DUPLICATE REPORT OF THE PREVIOUSLY SUBMITTED REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03457 | 2021-08-27 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/HT/CELL, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX400T12C | CAX400T12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING NOSE IRRITATION, ITCHING THROAT RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. SECTION H6(CLINICAL CODE), D4(UDI NUMBER) ARE CORRECTED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03572 | 2021-08-30 | 2021-08-03 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CHRONIC OBSTRUCTIVE PULMONARY DISEASE AND CONGESTIVE HEART FAILURE. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CHRONIC OBSTRUCTIVE PULMONARY DISEASE AND CONGESTIVE HEART FAILURE. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO SIGNS OF CONTAMINATION ON THE OUTSIDE OF THE DEVICE. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER CONFIRMED THERE WAS AN UNKNOWN CONTAMINATE ON THE INLET PORT, BLOWER MOTOR BROKEN AND SCREW JOINT BROKEN. THERE WAS ALSO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION INSIDE OF THE BLOWER BOX. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS LOGGED. THE DEVICE WAS HOOKED UP TO POWER SUPPLY, AIRFLOW WAS VERIFIED AND THE DEVICE OPERATED PROPERLY. THE MANUFACTURER FOUND ZERO ERRORS ON THE DEVICE. THE MANUFACTURER CONCLUDES EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION INSIDE OF THE BLOWER BOX . THE MANUFACTURER ALSO FOUND AN UNKNOWN CONTAMINATES ON THE INLET PORT. IN THE INITIAL REPORT CLINICAL CODE FOR SYMPTOMS NASAL AND THROAT IRRITATION WERE NOT MENTIONED WHICH HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03369 | 2021-08-30 | 2021-08-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, ES | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | ESX500S15 | ESX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03533 | 2021-08-30 | 2019-12-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID/HT, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AUX400T15 | AUX400T15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE SHORTNESS OF BREATH. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION FOR ACID REFLUX PILLS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS : THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED TO EXPERIENCE SORE EYES SHORTNESS OF BREATH, WHEEZE, STOMACH PAIN, PALPITATION IN THE EARS, SORE NOSE AND HEADACHES. THE REPORTED EVENTS WERE REVIEWED BY THE MANUFACTURER'S CLINICAL EXPERT. BASED ON THE INFORMATION AVAILABLE, THE MANUFACTURER CONCLUDES NO FURTHER ACTION IS NECESSARY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION B1, B2 HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT-IMPACT CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03562 | 2021-08-30 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX400H11C | DSX400H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A SINUS INFECTION. THE PATIENT WAS PRESCRIBED ANTIBIOTICS IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON JAN 01 , 2023 AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A SINUS INFECTION. THE PATIENT WAS PRESCRIBED ANTIBIOTICS IN RESPONSE TO THE REPORTED EVENT. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03547 | 2021-08-30 | 2021-08-02 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500S11 | DSX500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SHORTNESS OF BREATH. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGATION ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SHORTNESS OF BREATH. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THE ABOVE REPORTED EVENT IS ASSESSED AS POSSIBLY RELATED TO THE DEVICE IN THIS CASE. HENCE, THE DEVICE MAY HAVE CAUSE OR CONTRIBUTE TO THE ALLEGED SERIOUS INJURY. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03563 | 2021-08-30 | 2021-02-28 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102489) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA. THE PATIENT WAS PRESCRIBED STEROIDS IN RESPONSE TO THE REPORTED EVENT. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102489) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA. THE PATIENT WAS PRESCRIBED STEROIDS IN RESPONSE TO THE REPORTED EVENT. IN INITIAL REPORT SECTION B5 WAS INCOMPLETELY MENTIONED AND THE UPDATED SECTION B5 SHOULD BE- THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING OF HAVING COUGH AND DIFFICULTY BREATHING OR SHORT OF BREATH. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO EVIDENCE OF CONTAMINATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. THE DEVICE'S EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS LOGGED. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. SECTION D9, G3, H3 AND H6 HAS BEEN UPDATED / CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03538 | 2021-08-30 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, JP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | RJPX500S16 | RJPX500S16 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03566 | 2021-08-30 | 2021-07-13 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG NODULES, AND A MASS IN ONE LUNG. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PATIENT TO DEVELOP LUNG NODULES, AND A MASS IN ONE LUNG, HEADACHE, DIFFICULTY IN BREATHING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03586 | 2021-08-30 | 2021-08-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE DEVICE'S LCD SCREEN WAS OBSERVED TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03559 | 2021-08-30 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO, GER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DEX400S13 | DEX400S13 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN PREVIOUS REPORT, SECTION B5 WAS MENTIONED INCORRECTLY. CORRECT B5 SHOULD BE - THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES, HEADACHE, COUGHING, DIZZINESS AND EYE IRRITATION. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO AN AUTHORIZED SERVICE CENTER FOR EVALUATION. THE DEVICE WAS VISUALLY INSPECTED. THE SERVICE CENTER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE PASSED ALL FINAL TESTING. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03545 | 2021-08-30 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID/HT, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | CAX400T12 | CAX400T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03546 | 2021-08-30 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX500T12 | CAX500T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT ALLEGED SINUS PROBLEMS AND VISUALIZATION OF BLACK PARTICLES IN THE TUBING AND CHAMBER. THERE WAS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THE MANUFACTURER'S CONTACT INFORMATION HAS ALSO BEEN UPDATED. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. EXTERNAL INVESTIGATION OBSERVED A PARTIALLY UNSEATED FLIP-LID SEAL AND CONTAMINATION IN THE HUMIDIFIER. INTERNAL INVESTIGATION OBSERVED DUST ON AND IN THE BLOWER. PRODUCT INVESTIGATION LAB OBSERVED A DUSTY SMOKE (NON-TOBACCO) ODOR. SOUND ABATEMENT FOAM APPEARED INTACT, HOWEVER DEGRADED SOUND ABATEMENT FOAM RESIDUE ADHERED TO FITT WHEN COMPRESSED. ACCORDING TO THE LONG TEXT OF THE COMPLAINT, THE PATIENT ACKNOWLEDGED USING A SOCLEAN DEVICE PURPORTEDLY SOLD BY THE DME. PRODUCT INVESTIGATION LAB APPLIED POWER TO THE DREAMSTATION. VERIFIED DEVICE POWERS ON, PROVIDES AIRFLOW. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND SHOWED 1 E-130 ERROR LOGGED. THE MANUFACTURER IS UNABLE TO DIRECTLY ADDRESS THE COMPLAINT AS NO WATER TANK WAS RETURNED WITH THE DEVICE. THE MANUFACTURER WAS ABLE TO CONFIRM THE PRESENCE FOR DEGRADED SOUND ABATEMENT FOAM RESIDUE IN THE DEVICE. THE MANUFACTURER ALSO OBSERVED DUST CONTAMINATION IN BOTH THE BLOWER AND THE HUMIDIFIER. PLEASE SEE SECTIONS D9, D10, G1, H3 AND H6 FOR THIS UPDATED CODING AND INFORMATION. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03544 | 2021-08-30 | 2021-07-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL, AU | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | AUX500T15C | AUX500T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03588 | 2021-08-30 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S OXYGEN PERCENTAGE COULD NOT BE SET. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER HAD PREVIOUSLY REPORTED THAT A DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS A VENTILATORS INABILITY TO SET THE DELIVERED OXYGEN PERCENTAGE. THE INTERFACE BOARD WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. DURING THE INVESTIGATION, A "SERVICE REQUIRED" CODE WAS CONFIRMED. THE ROOT CAUSE OF THE INTERFACE BOARD FAILING WAS THE C27 4.85 VDC DETERIORATED DURING USE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03589 | 2021-08-30 | 2021-08-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | K181166 | K181166 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER HAD PREVIOUSLY REPORTED THAT THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS A SERVICE REQUIRED ERROR. A FAILURE OF THE DEVICE TO CHARGE ITS INTERNAL BATTERY WAS ALSO OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03585 | 2021-08-30 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | RJP2100X16B | RJP2100X16B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BLOWER MOTOR WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF A VENTILATOR INOPERATIVE CONDITION OCCURRING AND THE DEVICE'S BLOWER MOTOR WAS REPLACED TO ADDRESS THE ISSUE. THE SYSTEM BOARD WAS ALSO REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03567 | 2021-08-30 | 2021-07-02 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP HEADACHES AND DIZZINESS. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. IN PREVIOUS REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING THE PATIENT WAS IN THE HOSPITAL FOR OVER TWO WEEKS BECAUSE OF BLOOD FROM HIS HEAD AND ALSO HAS BEEN EXPERIENCING DIZZY AND LIGHT HEADED SPELLS FROM USING THE DEVICE. PATIENT ALSO REORTED PARTICLES IN TUBING/CHAMBER. THE REPORTED EVENT AND ITS SEVERITY WERE REVIEWED BY THE MANUFACTURER'S CLINICAL EXPERT. BASED ON INFORMATION AVAILABLE AT THIS TIME, THERE IS NOT ENOUGH EVIDENCE TO SUPPORT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE ALLEGED SERIOUS INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03537 | 2021-08-30 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AVAPS C SERIES INTL 30, FLOW GEN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | 1161T | 1161T |  | A BIPAP DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03587 | 2021-08-30 | 2021-08-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040000 | 1040000 |  | AA VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03445 | 2021-08-30 | 2021-08-19 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, GER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | DEX500S13 | DEX500S13 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE SINUS PROBLEMS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A PARANASAL SURGERY FOR SINUS ISSUES. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE NASAL CONGESTION, COUGHING, MUCUS, EYE IRRITATION AND SINUS PROBLEMS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A PARANASAL SURGERY FOR SINUS ISSUES. THE DEVICE IS NOT RETURNING TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER RECEIVED INFORMATION ON MAY 24, 2022 STATING THE PATIENT'S LAWYER REFUSES TO RELEASE THE DEVICE TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03607 | 2021-08-30 | 2021-06-01 | D, e, a, t, h | Manufacturer report | BIPAP AVAPS CSERIES | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | PHILIPS RESPIRONICS INC. | 1167P | 1167P |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED. THE PATIENT EXPIRED IN (B)(6) 2021. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03606 | 2021-08-30 | 2021-08-23 | I, n, j, u, r, y | Manufacturer report | BIPAP S/T CSERIES | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC | 1061457 | 1061457 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED BREAST CANCER. THE PATIENT'S MEDICAL INTERVENTION IS UNKNOWN AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME ALLEGEDLY DEGRADED AND CAUSED BREAST CANCER. THE PATIENT'S MEDICAL INTERVENTION IS UNKNOWN AT THIS TIME. UPON FURTHER REVIEW, THE REPORTED EVENT OF PRIMARY DIAGNOSIS OF PULMONARY FIBROSIS, HISTORY OF DIABETES AND BREAST CANCER AND THE PATIENT USING THE DEVICE 24/7 - CANNOT BE WITHOUT THE DEVICE WAS ASSESSED AS SERIOUS INJURY. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03535 | 2021-08-30 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO, INTL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | 761P | 761P |  | A BIPAP DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03534 | 2021-08-30 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO, INTL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | 761P | 761P |  | A BIPAP DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103612 | 2021-08-30 | 2021-06-25 | I, n, j, u, r, y | Voluntary report | PHILIPS REMSTAR AUTO A-FLEX CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. |  |  |  | PATIENT CALLED TO REPORT ADVERSE SYMPTOMS HE BELIEVES IS RELATED TO HIS RECALLED PHILIPS CPAP MACHINE. PATIENT STATED THAT HE HAS BEEN EXPERIENCING SYMPTOMS INCLUDING: BREATHING PROBLEMS, NASAL PASSAGE IRRITATION, VISION IMPAIRMENT, BLURRED VISION, CONGESTION, SEVERE FOOT PAIN IN RIGHT FOOT WHICH TURNED OUT TO BE GOUT, INCREASE IN ARTHRITIS SYMPTOMS INCLUDING SHOULDER, THUMB AND WRIST AREA, MORE THAN NORMAL DISCHARGE IN NASAL PASSAGE, AND EYE IRRITATION. PATIENT STATED THAT HE'S DISCOVERED IF HE SLEEPS ON HIS LEFT SIDE AND AT A 30 DEGREE ANGLE WITHOUT THE CPAP, HIS SYMPTOMS ARE LESSENED. PATIENT STATED HE LEARNED ABOUT THE RECALL BY MAILED LETTER WHICH HE BELIEVES HE RECEIVED IN EARLY JULY. |  | Y | Y | I |  | 2021-08
2518422-2021-03536 | 2021-08-30 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AVAPS C SERIES INTL 30, FLOW GEN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | 1161T | 1161T |  | A BIPAP DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103594 | 2021-08-30 | 2019-01-01 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAM STATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500T11C |  |  | EVERY TIME I USE MY PHILLIPS DREAMSTATION CPAP MY SINUS BLOCK UP AND BECOME IRRITATED AND I CANNOT BREATHE AT NIGHT. I HAVE LINGERING AND LONG TERM SINUS ISSUES. I HAVE BEEN SEEN BY ALLERGISTS AND ENT'S. THEY HAVE CONCLUDED THAT IT IS NOT RELATED TO ALLERGIES OR INFECTION. I HAVE HAD NUMEROUS TESTS FOR THIS ISSUE. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103593 | 2021-08-30 | 2021-07-30 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS CPAP MACHINES | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | THE AIR PRESSURE DIPS REALLY LOW. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103591 | 2021-08-30 | 2020-06-15 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTOBIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS | DreamStation Auto BiPAP w/Humid w/H Tube, w/Cell, US | DSX700T11C |  | SEVERE HEADACHES, LEFT EYE PAIN, CHEST PRESSURE, AND NASAL CHAPPING. COVID HAS RESULTED IN HER PRIMARY CONTACT WITH HER FAMILY VIA (B)(6) MEETING WITH MINIMAL CONTACT STILL ALLOWED AS OF (B)(6) 2021. I AM HER POWER OF ATTORNEY (SON), (B)(6) AND SHE IS UNABLE TO READ AND COMMUNICATE WELL DUE TO 2 STROKES IN 2018 AND SHE NOW RESIDES IN AN ASSISTED LIVING FACILITY IN (B)(6). I AM HER PRIMARY TRANSLATOR WHEN SHE IS ALLOWED TO LEAVE THE FACILITY. NO SPECIFIC TESTS HAVE BEEN RUN AS HER COMPLAINTS COME AND GO MOSTLY WITH NO WARNINGS AND LASTING FOR A FEW MINUTES TO SEVERAL HOURS. THIS IS A RECALL PRODUCT THAT I WAS MADE AWARE OF ONLY YESTERDAY (B)(6) 2021 VIA EMAIL FROM (B)(6). FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103588 | 2021-08-30 | 2018-01-01 | I, n, j, u, r, y | Voluntary report | RESPIRONICS SYSTEM ONE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | SERIES 60 |  |  | HELLO, I NOTICED BLACK FOAM PIECES COMING OUT OF MY PHILIPS RESPIRONICS CPAP DEVICE STARTING SEVERAL YEARS AGO - I ESTIMATE ~2017, AFTER PURCHASING THE DEVICE IN 2014. WHEN I TRAVELED, I WOULD PLACE MY CPAP IN A PLASTIC (B)(6) BAG. WHEN I WOULD GET TO MY DESTINATION AND GO TO SETUP MY CPAP, THAT'S WHEN I WOULD NOTICE A SMALL BLACK PIECE HAD FALLEN OUT INTO THE PLASTIC BAG. THIS HAPPENED MULTIPLE TIMES OVER THE YEARS WHEN I TOOK THE DEVICE APART FOR TRAVEL. I MENTIONED IT TO MY HUSBAND AND WAS CURIOUS ABOUT WHAT THESE BLACK PIECES WERE, BUT I DO NOT KNOW IF I MENTIONED IT TO MY CPAP PROVIDER. I DID WEB SEARCHES AND DID NOT SEE ANYTHING AT THAT TIME TO EXPLAIN WHAT THESE PIECES WERE. DUE TO INSURANCE I DID NOT QUALIFY FOR A NEW CPAP MACHINE FOR AT LEAST A FEW YEARS AFTER THIS STARTED (NOT UNTIL 2019 OR 2020). I AM (B)(6), I HAVE CHRONIC SINUS ISSUES AND LUNG TIGHTNESS AND CFS. I DO NOT KNOW IF THE USE OF THIS CPAP AFTER THE FOAM WAS FALLING APART HAS MADE MY CONDITION WORSE OR IF NATURALLY IT HAS WORSENED OVER THE YEARS. I AM EXTREMELY DISHEARTENED THAT MY CPAP COULD HAVE BEEN HARMING ME. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103607 | 2021-08-30 | 2021-06-17 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / REPISRONICS, INC. | DSX500T11C |  |  | DEVELOPED SEVERE FACIAL RASH THAT PERSISTED; SEVERE, CHRONIC HEADACHE; DIZZINESS; FEELING MALAISE; FATIGUE; INSOMNIA. CONSULTED WITH DOCTOR AND DME PROVIDER. ADJUSTMENTS WERE MADE TO EQUIPMENT (I.E. MASK TYPE WAS CHANGED A COUPLE OF TIMES; RAMP SPEED CHANGED FROM 15 TO 20. HEAT CHANGED FROM 3 TO 4. RESULTED IN NO LASTING IMPROVEMENT. NO ONE CONTACTED ME TO INFORM ME OF A RECALL OF THE PHILIPS DREAMSTATION CPAP MACHINE. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103584 | 2021-08-30 | 2021-05-21 | I, n, j, u, r, y | Voluntary report | DREAM STATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | DSX400T11 |  |  | ON (B)(6) 2021 A CHEST XRAY SHOWED WIDESPREAD PULMONARY NODULES IN BOTH MY LUNGS THAT WERE NOT PRESENT IN A CHEST XRAY TAKEN IN (B)(6) 2019. A CT SCAN (B)(6) 2021 AND ANOTHER ON (B)(6) 2021 SHOWED NUMEROUS SPICULATED PULMONARY NODULES THAT HAD INCREASED IN SIZE AND NUMBER IN THE 10 DAYS BETWEEN THE SCANS. THE SUSPECTED CAUSE FOR THE FORMATION OF THESE PULMONARY NODULES WAS DEFECTIVE COMPONENTS IN MY CPAP MACHINE. ON (B)(6) 2021, PHILIPS RESPIRONICS, THE MAKER OF MY CPAP MACHINE, A DREAMSTATION SERIAL NUMBER (B)(4) ISSUED A RECALL. PHILIPS ADVISED USERS TO DISCONTINUE USE OF THEIR CPAP MACHINES IMMEDIATELY STATING, EXAMPLES OF POTENTIAL RISKS INCLUDE EXPOSURE TO DEGRADED SOUND ABATEMENT FOAM PARTICLES OR EXPOSURE TO CHEMICAL EMISSIONS FROM THE SOUND ABATEMENT FOAM MATERIALS. I STOPPED USING MY DREAMSTATION CPAP MACHINE ON (B)(6) 2021. CURRENT DIAGNOSIS IS CRYPTOGENIC ORGANIZING PNEUMONIA. I'VE BEEN ON OXYGEN SINCE THE END OF (B)(6). I STARTED ON 30 MGS OF PREDNISONE AND 250 MGS OF AZITHROMYCIN ON (B)(6) 2021. MY IMMUNE SYSTEM IS COMPROMISED BECAUSE OF THE STEROIDS AND THE CDC HAS AUTHORIZED PERSONS TAKING 20 MGS OR MORE OF PREDNISONE TO GET A BOOSTER SHOT AGAINST COVID-19 WHICH I RECEIVED ON (B)(6) 2021. SINCE MY IMMUNE SYSTEM HAS BEEN COMPROMISED, I HAD TO CANCEL MY APPOINTMENT AT (B)(6) PULMONARY DEPARTMENT ON (B)(6) 2021. THE PURPOSE OF MY TRIP TO (B)(6) WAS TO DETERMINE ANY FUTURE DISABILITIES OR PERMANENT DAMAGE TO MY LUNGS. |  | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03595 | 2021-08-31 | 2021-08-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A "VENTILATOR INOPERATIVE" CONDITION WAS OBSERVED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03592 | 2021-08-31 | 2021-08-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03611 | 2021-08-31 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR "LOW CIRCUIT LEAK". THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SENSOR BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03594 | 2021-08-31 | 2021-08-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER HAD PREVIOUSLY REPORTED THAT THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS AN OBSERVED ISSUE. THIS IS INCORRECT. THE DEVICE WAS RETURNED UNREPAIRED AT THE REQUEST OF THE CUSTOMER. | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03609 | 2021-08-31 | 2021-08-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03610 | 2021-08-31 | 2021-08-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WOULD NOT POWER ON. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S INTERNAL POWER SUPPLY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03561 | 2021-08-31 | 2021-08-25 | I, n, j, u, r, y | Manufacturer report | BIPAP AUTO, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | 760P | 760P |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP THYROID CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03605 | 2021-08-31 | 2021-07-14 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP FLUID IN THEIR LUNGS. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP FLUID IN THEIR LUNGS, COUGHING AND CHOKING. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURE INSPECTED EXTERNALLY OBSERVED THAT DUST/DIRT CONTAMINATION ON TOP AND BOTTOM ENCLOSURE, PANEL, REAR PANEL, KEYPAD, BLOWER, BLOWER BOX, P4 POWER CONNECTOR, DC JACK COLOR INSERT. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE MANUFACTURER OBSERVED DIRT/DUST CONTAMINATION IN AIRPATH AND ARE CONSISTENT WITH BEING FROM AN EXTERNAL SOURCE. SECTION H6 TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03560 | 2021-08-31 | 2021-08-25 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO, INTL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | INX400S15 | INX400S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG AND BRAIN CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG AND BRAIN CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03570 | 2021-08-31 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP W/HUMID, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX200H12 | CAX200H12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY MISSED TO CAPTURE SOME INFORMATION. AFTER REVIEW, SHOULD BE FILED AS FOLLOWS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES, COUGH, NASAL AND THROAT IRRITATION RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTION H6 HEALTH EFFECTS: CLINICAL CODES HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03602 | 2021-08-31 | 2021-06-30 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A SINUS INFECTION. THE PATIENT WAS PRESCRIBED ANTIBIOTICS IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP A SINUS INFECTION,NAUSEA. THE PATIENT WAS PRESCRIBED ANTIBIOTICS IN RESPONSE TO THE REPORTED EVENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED EXTERNALLY AND INTERNALLY ,UNKNOWN WHITE AND BLACK CONTAMINATION (DUST LIKE) INCONSISTENT WITH DEGRADED SOUND ABTEMENT FOAM ,SOME POTENTIAL VERY SMALL HAIRS OR FIBERS AT THE AIR INPUT OF THE BLOWER BOX ,TOP AND BOTTOM OF THE BLOWER MOTOR AND THE BLOWER MOTOR SEAL,BLACK CONTAMINATION CONSISTENT WITH KERATIN AT THE AIT OUTLET OF THE BLOWER BOX.UNKNOWN WHITE DUST LIKE CONTAMINATION AT THE INSIDE OF THE BLOWER BOX ALONG WITH SOME UNKNOWN PIECES OF BLACK INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM,ONE TINY BLACK PIECE POTENTIAL SOUND ABATEMENT FOAM ON THE BOTTOM OF THE BLOWER BOX. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NINE ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THEY CONFIRMED THE PRESENCE OF MULTIPLE CONTAMINANTS THAT ARE NOT CONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM AND 1 PIECE OF ONTAMINATION THAT IS POTENTIAL DEGRADED SOUND ABATEMENT FOAM AND THERE IS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS OBSERVED IN THIS DEVICE. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03603 | 2021-08-31 | 2021-07-12 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11 | DSX400H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP KIDNEY CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO KIDNEY CANCER. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENT OF KIDNEY CANCER AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. DURING REVIEW OF THE ERROR LOGS, PIL DETERMINED THAT THE DEVICE WAS LIKELY RESET PRIOR TO PIL RECEIPT DUE TO THE MACHINE HAVING 7079.2 HOURS AND 0 BLOWER AND THERAPY HOURS. NO CARE ORCHESTRATOR DATA WAS AVAILABLE FOR DOWNLOAD. DURING THE INTERNAL INVESTIGATION OF THE DEVICE, PIL OBSERVED AN UNKNOWN DUST CONTAMINANT ON THE BOTTOM ENCLOSURE, BLOWER, AND BLOWER SEAL. EVIDENCE OF LIQUID INGRESS WAS OBSERVED TO THE BLOWER AND BLOWER BOX. INV1023 ADDRESSES THE ISSUE OF LIQUID INGRESS. A KERATIN-LIKE SUBSTANCE WAS OBSERVED ON THE BLOWER BOX OUTLET. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED. PIL CONFIRMED NO PRESENCE OF DEGRADED SOUND ABATEMENT FOAM USING A FITT AND THE ASSOCIATED PROCEDURE ER 2245460 (V01). SECTION(S) B1, B2 HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. CORRECTED INFORMATION WAS PROVIDED IN G1.SECTION. H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03604 | 2021-08-31 | 2021-06-30 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG EFFUSION WHICH REQUIRED DRAINING. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DEVELOP LUNG EFFUSION AND SHORT OF BREATH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND DUST/DIRT CONTAMINATION INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM WAS OBSERVED THROUGHOUT DEVICE ENCLOSURE, AND AIRPATH, SUGGESTING A SOURCE EXTERNAL TO THE DEVICE. EVIDENCE OF WATER INGRESS ON THE BLOWER AND BLOWER BOX. CIRCULAR MARKINGS FOUND ON BLOWER BOX LID. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGREDATION. THE DEVICE DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERROR. THE MANUFACTURER CONCLUDES THE DEVICE HAS DUST/DIRT CONTAMINATION INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM WAS OBSERVED THROUGHOUT DEVICE ENCLOSURE, AND AIRPATH, SUGGESTING A SOURCE EXTERNAL TO THE DEVICE. EVIDENCE OF WATER INGRESS ON THE BLOWER AND BLOWER BOX. CIRCULAR MARKINGS FOUND ON BLOWER BOX LID. THERE WAS NO EVIDANCE OF SOUND ABATEMENT FOAM DEGREDATION. SECTION D8, D9 AND H3, H6 UPDATED IN THIS REPORT. SECTION H6(CLINICAL CODE) CORRECTED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03496 | 2021-08-31 | 2021-05-17 | D, e, a, t, h | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS,INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A CARDIAC EVENT. THE PATIENT WAS HOSPITALIZED AND SUBSEQUENTLY EXPIRED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A VENTILATOR'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A CARDIAC EVENT. THE PATIENT WAS HOSPITALIZED AND SUBSEQUENTLY EXPIRED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03590 | 2021-08-31 | 2021-08-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | DS2200X11B | DS2200X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S MACHINE FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03593 | 2021-08-31 | 2021-08-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03461 | 2021-08-31 | 2021-08-07 | D, e, a, t, h | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040004 | 1040004 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT'S BLOOD OXYGEN LEVELS DECREASED WHILE USING A VENTILATOR. THE PATIENT WAS PLACED ON ANOTHER DEVICE AND EXPIRED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE LOGGED SERVICE REQUIRED CODE INDICATES LOW OXYGEN FLOW. AT THE TIME OF THE REPORTED EVENT, THE DEVICE WAS SET TO DELIVER 100% OXYGEN WITH A RESPIRATORY RATE THAT EXCEEDED 30 BREATHS PER MINUTE. WITH THE DEVICE SETTINGS, THE OXYGEN SOURCE WAS UNABLE TO DELIVER THE REQUIRED FLOW. THE DEVICE PASSED ALL TESTING AND WAS FOUND TO OPERATE AND ALARM AS DESIGNED. IT APPEARS THE EVENT WAS RELATED TO THE PATIENT'S DETERIORATING CONDITION. |  | Y | N | HEALTH PROFESSIONAL |  | 2021-08
MW5103622 | 2021-08-31 | 2021-08-22 | I, n, j, u, r, y | Voluntary report | TRILOGY MACHINE | CONTINUOUS, VENTILATOR, HOME USE | PHILIPS / RESPIRONICS, INC. |  |  |  | MY MOTHER EXPERIENCED A CRITICAL DECREASE IN O2 LEVELS WAS RUSHED TO HOSPITAL. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103623 | 2021-08-31 | 2019-06-03 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS DREAM MACHINE BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX600H11C |  |  | HEAVY COUGH BREATHING PROBLEM; RECEIVED PHILLIPS DREAM STATION, DSX600H11C BIPAP SN (B)(4) FROM PHILLIPS AND AROCARE WITH HELP OF MEDICARE. PHILLIPS NOW HAS RECALLED THIS MACHINE AND WE WERE NOTIFIED NOT TO USE THIS BIPAP IN (B)(6) 2021. NOW WITHOUT A MACHINE, I AM LOOSING SLEEP AND HAVING PROBLEMS RESTING, COUGHING AND LUNG PROBLEMS. HAVE AN APPOINTMENT WITH A PULMONARY DR. (B)(6) 2021. OTHER TEST RESULTS WILL BE AVAILABLE AFTER THIS APPT. MEDICARE HAS REFUSED TO REPLACE THIS MACHINE SINCE IT IS NOT OVER 5 YEARS OLD? NOW MEDICARE AND PHILLIPS IS STALLING AND WE ARE STILL WITHOUT A USABLE BIPAP MACHINE. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103632 | 2021-08-31 | 2020-05-08 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | RESPIRONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | RESPIRATORY TRACT IRRITATION; WHEN I WOKE UP IN THE MORNING ((B)(6) 2020) I NOTICED THAT THERE WAS BLACK POWDER AROUND MY NOSE AND ON MY MASK. I HAD A SORE THROAT FOR DAYS AFTER AND HAVE HAD AN INCREASE IN PHLEGM PRODUCTION, ESPECIALLY IN THE MORNINGS AFTER WAKING, FOR A WHILE NOW, WHICH STARTED AFTER THIS OCCURRED. ALSO, THE BEGINNING OF THIS YEAR I WAS DIAGNOSED WITH GASTRITIS AND HAD BIOPSIES OF SALMON MUCOSA IN MY ESOPHAGUS DURING AN ENDOSCOPY IN (B)(6) OF 2021. I HAVE ALSO BEEN HAVING CONTINUOUS UNEXPLAINED WEIGHT LOSS AND ABDOMINAL PAINS. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103635 | 2021-08-31 | 2021-06-04 | I, n, j, u, r, y | Voluntary report | DREAM STATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I STARTED TO HAVE DIFFICULTY BREATHING. AFTER MULTIPLE DOCTOR VISITS AND A TRIP TO THE ER I WAS DIAGNOSING INFLAMING MY LUNGS FROM AN UNKNOWN ENVIRONMENTAL CAUSE. FURTHER TESTING REVEALED I HAD SCARING ON MY LUNGS AND ONLY 47% LUNG CAPACITY. I AM STILL UNDERGOING TREATMENT AND TESTING. PULMONARY FUNCTION TESTING, LAB WORK, AND AN X-RAY HAVE ALSO BEEN COMPLETED. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103636 | 2021-08-31 | 2021-08-20 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | BIPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS MEDICAL / RESPIRONICS, INC. | B-FLEX 650 P |  |  | RECALLS HAVE BEEN MADE TO CONSUMERS REGARDING PHILIPS MEDICAL, CPAP/ BIPAP MACHINES. TO THIS DATE I CAN NOT CONFIRM IF MY BIPAP MACHINE FALLS UNDER THIS RECALL. MY MACHINE IS: BIPAP PRO, B-FLEX, SYSTEM ONE, REF # 650P, S/N (B)(4), FE 9D, WITH HUMIDIFIER, HAVING OVER 29000 HOURS OF USE. I DO NOT HAVE CANCER OR OTHER PROBLEMS THAT ARE NOTED ON RECALLS ON THE WWW. HOW LONG DO YOU NEED TO HOLD A GUN AT YOUR HEAD BEFORE IT GOES OFF? I DON'T WANT CANCER. PLEASE HELP ME OUT. THANK YOU. (B)(6). FDA SAFETY REPORT ID# (B)(4). |  | N | Y | I |  | 2021-08
MW5103641 | 2021-08-31 | 2021-01-01 | I, n, j, u, r, y | Voluntary report | PHILIP'S RESPIRONICS DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DreamStation Auto CPAP w/Humidifier, w/Cell, US | DSX500H11C |  | I HAD BEEN COUGHING UP PHLEGM FOR A LONG TIME. THOUGHT IT WAS ACID REFLUX AND WAS PRESCRIBED PROTONIX 2X DAILY. THEN GROWTH IN THYROID KEPT GETTING LARGER, INHIBITING SWALLOWING. ULTRASOUND SHOWED 2 MASSES, BOTH BENIGN AFTER BIOPSIED. LEG NUMBNESS CONTINUED. X-RAYS SHOWED HEMANGIOMAS ON LUMBAR SPINE AND KIDNEY LESIONS. WILL HAVE 3RD MRI TOMORROW. FIRST EMG WAS NORMAL BUT ONE THIS WEEK SHOWED NERVE DAMAGE. A1C WAS 5.7 SO NOT DIABETES. I'VE HAD TO GO FOR PHYSICAL THERAPY, CHIROPRACTIC IN CASE OF PINCHED NERVE (NO), TO A UROLOGIST, 2 NEUROLOGISTS AND A NEUROSURGEON. IT CONTINUES TO WORSEN AND AFFECTS MY ABILITY TO SLEEP, SIT, OR LIE DOWN. JUST THIS WEEK, (B)(6) TOLD ME ABOUT THE RECALL. I REGISTERED MY DEVICE AND DISCONTINUED USE OF THAT CPAP. MY PROVIDER WENT OUT OF BUSINESS SO PERHAPS THAT IS WHY I DIDN'T KNOW ABOUT THE RECALL. I'VE HAD MANY TESTS WITH MORE COMING. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103642 | 2021-08-31 | 2021-08-01 | I, n, j, u, r, y | Voluntary report | PHILIPS CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS, INC. | DSXHCP |  |  | I HAVE BEEN USING A CPAP FOR 10 YEARS BUT RECENTLY, IN THE PAST TWO YEARS, I'VE NOTICED MY LUNG FUNCTION WORSENING AND IT WASN'T UNTIL WE HEARD OF THE RECALL OF OUR CPAP MACHINES THAT I REALIZED, AFTER INSPECTING THE MACHINES, THAT THEY WERE THE REASONS WHY MY LUNG FUNCTION WERE WORSENING POSSIBLE EXACERBATING MY BRAIN FUNCTION AND HEART PROBLEMS. I WAS HAVING COUGHING SPELLS AT NIGHT INTERRUPTING MY SLEEP, FEELING LIKE THERE WAS SOMETHING RASPY IN MY THROAT, LIKE SAND, HEADACHES AND DISORIENTATION WHEN I WOKE UP THE NEXT MORNING. I WAS GETTING BLACK STUFF OUT OF MY THROAT AND NOSE AND I HAD NO IDEA WHAT IT WAS UNTIL, UPON INSPECTION, I SAW ALL THE BLACK STUFF IN MY MACHINE. MY MACHINE IS A PHILLIPS DREAM STATION, FOUR YEARS OLD. MY WIFE'S IS SAME DREAM STATION BUT ONLY 2 YEARS OLD. WE REGISTERED THEM BOTH ON THE RECALL REGISTRATION LINE ABOUT A MONTH AGO. I AM A (B)(6) MALE, I AM DEAF DUE TO INJURY AND A BRAIN TUMOR. MY WIFE HAS TO DO ALL THE PHONE CALLS. SHE HAS BEEN CALLING ALL OUR CARE PROVIDERS AND DOCTORS, MY HEART DOCTOR, MY LUNG DOCTOR, MY CPAP DOCTOR AND ALSO OF COURSE OUR FAMILY DOCTOR. I CANNOT SLEEP WITHOUT MY CPAP MACHINE SO I AM USING THE MACHINE WITH THE IN-LINE FILTER BUT I'M STILL CONCERNED ABOUT THE POSSIBILITY OF TOXINS COMING OFF THE BLACK DEBRIS THAT OUR CPAP DR. SAYS IS POSSIBLY CARCINOGENIC. MY WIFE IS VERY UNCOMFORTABLE WITHOUT HER MACHINE BUT SHE KNEW THAT I HAD TO HAVE A MACHINE AND SINCE HER MACHINE WAS NEWER, AND HAD LESS BLACK DEBRIS, WE HAD OUR DOCTORS CHANGE THE SETTINGS AND I AM USING HER MACHINE. PLEASE, WE NEED HELP!!! SINCERELY, (B)(6). NO TEST RELATED TO THE CPAP PROBLEM. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103643 | 2021-08-31 | 2019-11-20 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | FREQUENT SINUS INFECTIONS (7 TIMES IN 2020 AND 2-3 TIMES SO FAR IN 2021), SEVERE HEAD, EYE, NOSE PAIN. EYE IRRITATION TEARING AND CRUSTING. CONTINUOUS COUGH THAT HAS WORSENED. DIFFICULTY BREATHING. BRAIN FOG. WHILE USING MY CPAP HAVE HAD DIFFICULTY BREATHING. MY FILTERS HAVE REQUIRED BEING CLEANED AND CHANGED MORE FREQUENTLY. THE HUMIDITY HAS CHANGED WITH MY WATER CHAMBER BECOMING EMPTY WHEN USUALLY WOULD NOT. ALSO THE WATER IN MY WATER CHAMBER BECOMES HOT AS WELL AS THE METAL PLATE ON THE BOTTOM OF THE CHAMBER. I DO NOT HAVE ANY HEATING FUNCTION. I NEED A NEW CPAP MACHINE AND EQUIPMENT BUT BECAUSE MY MACHINE IS NOT 5 YEARS OLD I HAVE BEEN TOLD I AM NOT ELIGIBLE FOR ONE NOW AND HAVE TO WAIT UNTIL PHILIPS SENDS ME ONE. PHILIPS HAS TOLD ME THEY ARE WAITING ON THE FDA APPROVAL BEFORE THEY CAN SEND ME ONE. MY PULMONARY OFFICE HAS TOLD ME THAT I WAS HIGH RISK DUE TO HOW I ANSWERED SOME QUESTIONS AND NEED TO USE MY MACHINE. I HAVE TRIED NOT USING MY MACHINE BUT COULD NOT TOLERATE IT. SO I HAVE CONTINUED TO USE MY MACHINE. I WOULD REALLY APPRECIATE A NEW CPAP MACHINE AND EQUIPMENT. I HAVE COMPLETED THE REPORT. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103644 | 2021-08-31 | 2017-05-09 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DREAMSTATION |  |  | I RECENTLY LEARNED OF THE PHILIPS DREAMSTATION RECALL AND IT MAY EXPLAIN SOME OF THE SYMPTOMS I'VE BEEN EXPERIENCING FOR SEVERAL YEARS. I WAS DIAGNOSED THIS SPRING AS HAVING EXERCISED INDUCED ASTHMA, WHICH IS A NEW THING FOR ME. ALSO THIS YEAR I WAS EVALUATED FOR LEFT EYE IRRITATION AND PRESCRIBED ANTIBIOTIC EYE DROPS, BUT IT HAS NOT RESOLVED. FOR THE LAST SEVERAL YEARS I'VE BEEN EVALUATED FOR LIGHT-HEADEDNESS AND DIZZINESS, WITH NO CAUSE FOUND, OR RESOLUTION. I FREQUENTLY WAKE UP IN THE MORNING WITH A DULL HEADACHE EVEN THOUGH THE NUMBERS FROM MY CPAP MACHINE SAY THAT I AM GETTING EXCELLENT RESULTS FROM IT. I ALSO FREQUENTLY EXPERIENCE A MILD SORE THROAT. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103645 | 2021-08-31 |  | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS DREAM STATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I HAVE BEEN USING MY PHILIPS RESPIRONICS CPAP DEVICE FOR APPROXIMATELY A YEAR. DURING THAT TIME I HAVE HAD HEADACHES, COUGHING, SORE THROAT, AND CONGESTION AFTER USE. I NEVER USED ANY CLEANER OTHER THAN WHAT I WAS ADVISED TO USE BY MY DME VENDOR WHICH IS EITHER VINEGAR OR (B)(6). I HAVE ONLY EVER USED SMALL AMOUNTS OF (B)(6) ALONG WITH WATER FOR CLEANING. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | I |  | 2021-08
2518422-2021-03591 | 2021-08-31 | 2021-08-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-08
2518422-2021-03568 | 2021-08-31 | 2021-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL, AU | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | AUX500T15C | AUX500T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-08
MW5103648 | 2021-08-31 | 2021-02-12 | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DREAM STATION, DSX500H11C |  |  | I AM USING THE PHILIPS CPAP DREAM STATION PRESCRIBED BY MY SLEEP APNEA DOCTOR. APPROXIMATELY 6 MONTHS AGO I REPORTED TO MY DOCTOR THAT I HAVE A PERSISTENT COUGH AND FREQUENT HICCUP TYPE SYMPTOMS. AS OF TODAY THESE CONDITIONS REMAIN, AND THE COUGH HAS BECOME EVEN MORE FREQUENT. THE HICCUP SYMPTOM SEEMS TO COME FROM AN EXCESS OF AIR IN MY STOMACH, WHICH DOES CAUSE SOME CHEST AND STOMACH PRESSURE. DUE TO THESE ISSUES, I HAD AN UPPER AND LOWER GI PERFORMED IN THE PAST 6 MONTHS WHICH DID NOT REVEAL ANY CLEAR PROBLEMS. AT THIS TIME I'M NOT SURE WHAT TO DO. I PLAN TO CALL MY SLEEP APNEA PHYSICIAN FOR HER ADVICE. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103646 | 2021-08-31 | 2020-06-08 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS, INC. | 3194661 |  |  | SORE THROAT; DOCTOR SAYS COUGH COULD BE RELATED TO CPAP. TRIED 3 BOXES OF MUCINEX BUT DID NO GOOD FOR COUGH. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103656 | 2021-08-31 | 2021-07-27 | I, n, j, u, r, y | Voluntary report | DREAM STATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | REF-DSX500T11 |  |  | DIFFICULTY BREATHING AT NIGHT; WHEN CALLING OUR SUPPLIER ABSOLUTE RESPIRATORY CARE FOR A FACE MASK, I WAS NOTIFIED OF PROBLEMS WITH CPAP MACHINES AND TOLD TO CALL PHILLIPS. PHILLIPS TOOK THE SERIAL NUMBER AND INFORMED ME TO CALL MY DOCTOR. THE DOCTOR INFORMED ME I SHOULD NOT USE MY CPAP MACHINE IT IS IN RECALL. I THEN CALLED PHILLIPS AGAIN AND CONTACTED A REPRESENTATIVE THAT EXPLAINED THEY WERE IN DISCUSSION WITH REGULATORY AGENCY. I THEN CALLED MY INSURANCE COMPANY THAT TELLS ME MY MACHINE NEEDS TO BE 5 YRS OLD TO OBTAIN A REPLACEMENT. AS OF JULY 27TH I HAVE NO MACHINE TO INSURE MY OXYGENATION AT NIGHT. I AM SCHEDULED TO SEE MY NEUROLOGIST BECAUSE THERE IS QUESTIONS OF SEIZURE ACTIVITY DUE TO LACK OF OXYGEN AT NIGHT. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | I |  | 2021-08
MW5103655 | 2021-08-31 |  | I, n, j, u, r, y | Voluntary report | CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS, INC. | DSXHCP | DSX500H11C |  | I NOTICED A CHANGE IN MY SLEEP QUALITY SEVERAL MONTHS AGO. THEN BEGAN TO HAVE HEADACHES AND COUGHING, AND ASTHMA IS WORSENING. MY EYES ARE ALSO DRY AND ITCHY. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103653 | 2021-08-31 | 2021-06-14 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTOCPAPHUMCELL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | BREATHING PROBLEMS, SLEEP APNEA, COPD. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103651 | 2021-08-31 | 2007-01-01 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS REMSTAR PRO C-FLEX+ | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | WHILE USING THE CPAP MACHINE I HAVE HAD SEVERAL SEVERE SINUS INFECTIONS AND A CONSTANT NASAL DRIP. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
MW5103649 | 2021-08-31 | 2021-07-15 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  | DSX500T11 |  | SUFFERED FROM THE FOLLOWING SIDE EFFECTS: HEADACHE, UPPER AIRWAY IRRITATION, COUGH, CHEST PRESSURE, RESPIRATORY TRACT IRRITATION, HEADACHE/DIZZINESS, HYPERSENSITIVITY, AND NAUSEA. ALSO SUFFERED HEART ATTACK AFTER USE. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-08
2518422-2021-03662 | 2021-09-01 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | PHILIPS/RESPIRONICS, INC | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. |  | N | Y | LAY USER/PATIENT |  | 2021-09
MW5103681 | 2021-09-01 | 2021-03-15 | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP V1.0.2690 | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | MULTIPLE SYMPTOMS; HAVE BEEN USING THE DREAM STATION AUTO CPAP MACHINE SINCE 2016. HAVING INCREASING HEADACHES, EYE IRRITATION SHORTNESS OF BREATH WITH WHEEZING REQUIRING AN INHALER (DOCTORS STATE CHEST XRAY NORMAL, PULSE OXIMETRY NORMAL) BUT HARD TO BREATH. UPPER AIRWAY IRRITATION AND COUGH. ALL THESE SYMPTOMS ARE INCLUDED ON THE LIST OF REACTIONS FROM USING THIS MACHINE. POSSIBILITY OF CARCINOGENIC EFFECTS ARE LISTED, MAY NOT KNOW THAT UNTIL LATER. THE STATED EXPOSURE FROM THE DEVICE IS CHEMICAL AND PARTICLE. HARD TO KNOW FOR SURE WHEN EXACTLY THESE STARTED AS THEY HAVE COME ON SLOWLY BUT STARTED THE INHALER IN (B)(6) 2021. THE MACHINE HAS NOW BEEN RECALLED BUT SEEMS LIKE THE DAMAGE HAS BEEN DONE. HOW DO WE PURSUE COMPENSATION FOR THE LIFESTYLE CHANGE DUE TO SHORTNESS OF BREATH AND IF A CANCER OCCURS IN THE FUTURE?. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5103670 | 2021-09-01 |  | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | DREAMSTATION |  |  | ONGOING SINUS AND THROAT PROBLEMS. SMALL BLACK FLECKS HAVE BEEN IN MY MUCUS FOR 2 YEARS, REPORTED TO MY DOCTOR WITH NO SOURCE FOUND. ONE KNOWN ALLERGY, CEDAR. FOR SPECIFIC DATES, INFO MY DR. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5103673 | 2021-09-01 | 2021-06-17 | I, n, j, u, r, y | Voluntary report | PHILLIPS DREAMSTATION CPAP CPAP PRO AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. | IP22 |  |  | I'M NOT SURE IF THIS IS RELATED BUT I'VE BEEN A HEALTHY (B)(6) WOMAN WHO WAS STILL EMPLOYED ENJOYING MY WORK. I USE A CPAP DREAMSTATION PHILIPS RESPIRONICS FOE SLEEP APNEA. I'VE HAD A PAROTID BENIGN TUMOR FOR 10 YEARS. SINCE I STARTED USING MY DREAMSTATION IN 2018 I'VE HAD MORE SYMPTOMS WITH THE TUMOR. I DECIDED TO HAVE THE TUMOR REMOVED AND THEY FOUND CANCER. IT IS A RARE CANCER AND HAD TRAVELED ALONG THE NERVES IN MY FACE SO THE NERVES IN THE LEFT SIDE OF MY FACE ALONG WITH THE PAROTID GLAND AND SOME LYMPH GLANDS WERE TAKEN. THE LEFT SIDE OF MY FACE IS PARALYZED. I'VE HAD TWO OTHER SURGERIES SINCE THEN, ONE TO REMOVE MORE CANCER AND ONE TO PUT A GOLD LEAF IN MY EYE TO ALLOW ME TO OPEN AND CLOSE IT. I HAVE REGISTERED MY CPAP BUT NOTHING HAS BEEN DONE TO REPLACE IT YET. I'VE BEEN TRYING TO GET A DIFFERENT CPAP AFTER MY DOCTOR SENDING MY PRESCRIPTION TO ANOTHER COMPANY BUT IT IS TAKING FOREVER. I'M STILL HAVING TO USE THE DREAMSTATION BECAUSE I HAVE NOTHING ELSE. IF THESE ARE DIRECTLY RELATED TO MY ISSUE WITH CANCER I HAVE HUGE RESERVATIONS WITH USING IT. I JUST WANTED TO PUT THIS OUT THERE IN CASE IT RELATES DIRECTLY TO THE PHILIPS DREAMSTATION. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5103676 | 2021-09-01 | 2020-06-14 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS REMSTAR SYSTEM ONE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. | 20091231 |  |  | I EXPERIENCED DIFFICULTY HOLDING MY HEAD UP AND FEELING WEAK. I WAS ALSO HAVING TROUBLE WALKING AND DOING ANYTHING. I WENT TO THE EMERGENCY ROOM. I WAS DIAGNOSED WITH HEMOPTYSIS AND PNEUMONIA DUE TO ORGANISM. THEY CONDUCTED SEVERAL LAB TEST, AN ELG 12 LEAD, A CHEST XR. I WAS GIVEN ANTIBIOTICS AND RELEASED HOWEVER OTHERWISE ISSUES SURFACED. DURING THIS TIME I WAS CONTINUING TO USE MY CPAP MACHINE UNAWARE OF ANY ISSUES. I RETURNED TO THE EMERGENCY ROOM JUNE 30 DUE TO WEAKNESS AND DIFFICULTY BREATHING. THEY RAN NUMEROUS TEST AND ENDED UP SENDING ME TO (B)(6) CLINIC'S CARDIAC WARD. I WAS THERE TWO WEEKS TWO DAYS OF THOSE TWO WEEKS I WAS IN THE ICU. THEY CONCLUDED MY HEART HAD GONE INTO CONSTANT AFIB AND I HAVE BLOOD CLOTS IN MY LUNGS. I AM STILL DEALING WITH THE TIREDNESS AND I AM ON MEDICATION FOR THE BLOOD CLOTS. I BELIEVE WHAT I WENT THROUGH WAS DUE TO MY CPAP MACHINE ESPECIALLY SINCE THE DOCTORS ARE NOT SURE HOW THE CLOTS GOT IN MY LUNGS. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03648 | 2021-09-01 | 2021-07-12 | I, n, j, u, r, y | Manufacturer report | BIPAP AUTOSV ADV W/HUM,SYS ONE,60SRS,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC. | DS960HS | DS960HS |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE BLADDER CANCER. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A TREATMENT FOR CHEMOTHERAPY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP,AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AND CAUSED THE PATIENT TO HAVE BLADDER CANCER RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A TREATMENT FOR CHEMOTHERAPY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03657 | 2021-09-01 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP S/T, C-SERIES, CHINA | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | 1061457 | 1061457 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
MW5103684 | 2021-09-01 | 2021-08-21 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS /RESPIRONICS INC. |  |  |  | PHILIPS RESPIRONICS RECALLED MY DEVICE WHICH I HAVE SLEEP APNEA AND CAN DIE IN MY SLEEP. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-09
MW5103677 | 2021-09-01 |  | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAM MAKER CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | MY PHILIPS RESPRONIC CPAP DREAMMACHINE HAS BEEN RECALLED. I HAVE NOT TESTED MACHINE AND AM STILL USING IT. MACHINE IS DUE FOR NORMAL REPLACEMENT IN (B)(6) 2022. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | I |  | 2021-09
MW5103683 | 2021-09-01 | 2021-08-30 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAM STATION WITH HUMIDIFIER AND WIFI | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. |  |  |  | RECEIVED LETTER IN MAIL FROM PHILIPS ITS CPAP MACHINES HAD PIECES OF FOAM NOISE FILTER COMING LOOSE AND TRAVELING THROUGH AIR TUBE AND CAUSING PATIENT TO BREATHE IN THESE PARTICLES, RESULTING IN THROAT IRRITATION AND COUGH. I CALLED TO REGISTER MY SON'S CPAP MACHINE FOR RECALL, WAS TOLD THE ISSUE HAD BEEN REPORTED TO THE FDA ON JUNE 14, 2021 AND APPROPRIATE REGULATORY AUTHORITIES, AND AFTER I ASKED WAS TOLD RECALL DATES WERE NOT YET DETERMINED AND NEITHER WAS LENGTH OF TIME TO REPAIR OR REPLACE AND RETURN MACHINES TO PATIENTS. SLEEP APNEA IS A SERIOUS AILMENT, AND PHILIPS BEING THE ONE THAT MANUFACTURES THE CPAP MACHINES APPEARS TO BE COMPLETELY CLUELESS TO THE FACT NOT HAVING A HEAVILY DEPENDED UPON PIECE OF EQUIPMENT OR HAVING TO USE ONE THAT IS DEFECTIVE AND COULD CAUSE ILLNESS, LUNG DAMAGE AND POSSIBLY DEATH IS UNACCEPTABLE. PHILIPS AND THE FDA ARE BLITHELY AND BLINDLY TAKING THEIR TIME WHILE PATIENTS SUFFER AGAIN PATIENTS ARE BREATHING THESE PARTICLES INTO THEIR LUNGS AND NOT HAVING THEIR MACHINE TO USE NIGHTLY COULD CAUSE DEATH YOUNG PEOPLE HAVE DIED FROM SLEEP APNEA. WHAT ARE YOU ALL THINKING? I TOLD THE REP ON THE PHONE TO GET GOING AND REACH UP TO UPPER MANAGEMENT AND TELL THEM TO PUT SOME TRACTION UNDER THIS ISSUE AND THE FDA WHAT ARE YOU ALL WAITING FOR SOMEONE TO DIE BECAUSE THEY'VE INGESTED TOO MANY OF THESE FOAM PARTICLES? PHILIPS IS LOOKING FOR A CHEAP WAY OUT OF THIS WHAT WILL IT COST THEM IF SOMEONE DIES DUE TO PHILIPS' NEGLIGENCE? MERELY SENDING A LETTER AND REPORTING OR PUBLICIZING IT TO PEOPLE DOES NOT CUT IT THIS ISSUE REQUIRES SERIOUS ACTION MANDATE THAT PHILIPS GET THE RECALL STARTED. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03627 | 2021-09-01 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL, AU | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | AUX500T15C | AUX500T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03599 | 2021-09-01 | 2021-07-02 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A COLLAPSED LUNG. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | ADDITIONAL INFORMATION WAS RECEIVED ON OCT 24, 2024 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING CHEST PAIN, BRAIN FOG, COLLASPED LUNG, URAINE PROBLEM RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE SECTION E1 WAS UPDATED AND SECTIONS B1, B2, H1 AND H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR BOTH ADVERSE EVENTS AND PRODUCT PROBLEM (ONLY THE PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECTS - CLINICAL CODES AND HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03446 | 2021-09-01 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID/HT, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | AUX400T15 | AUX400T15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03601 | 2021-09-01 | 2021-07-15 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION B7 AND H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03600 | 2021-09-01 | 2021-07-28 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A CHEMICAL BURN. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A CHEMICAL BURN. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. IN PREVIOUS REPORT, SECTION B3, B5 AND H10 WAS INCORRECT. THE CORRECT B5 SHOULD BE- THE PATIENT TO ALLEGED CHEMICAL BURN AND VISUALIZATION OF BLACK PARTICLES IN THE DEVICE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL AND INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND DIRT/CONTAMINATION UNDER THE SD CARD, AIR INTAKE FLIP COVER ON THE BASE UNIT. THE MANUFACTURER ALSO RECEIVED HUMIDIFIER. A VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND DUST, DIRT, HAIR, ALGAE/MOLD WERE FOUND ON THE FLIP LID HINGE BARS, ON THE HEATED TUBE PORT/CONNECTOR, UNDER THE TOP ENCLOSURE LATCH, ON THE BOTTOM OF THE HUMIDIFIER CHAMBER, ON THE DRY BOX NEAR THE ISO SEAL AND ON THE HEATER PLATE .THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS. THE DEVICE WAS APPLIED POWER AND THE DEVICE OPERATED PROPERLY. THE MANUFACTURER CONCLUDES MULTIPLE CONTAMINATION FOUND WERE CONSISTENT WITH CONTAMINATION OF DUST, DIRT, HAIR, ALGAE/MOLD IN THE DEVICE AND HUMIDIFIER. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. IN THIS REPORT, SECTION B3, D9, G3, H3, H6 HAS BEEN UPDATED OR CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03660 | 2021-09-01 | 2021-08-02 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE USER ALLEGES HEALTH COMPLICATIONS INCLUDING DIABETES, INFLAMMATION, LEUKOCYTOSIS, AND HYPERTENSION. NO SERIAL NUMBER OR MODEL NUMBER WERE PROVIDED. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DIABETES, NAUSEA AND VOMITING, DYSPHAGIA GERDS, ACUTE INHALATION, HEADACHES, COUGHS, ESOPHAGITIS, HTN, TOXIC AND CARCINOGENIC EFFECT, INFLAMMATION, AND FLUID BUILD UP LUNG LEUKOCYTOSIS, PAIN SUFFERING, EMOTIONAL DISTRESS, LOSS OF ENJOYMENT, AND LOSS OF CONSORTIUM. NEEDING TACHEOMETRY SURGERY RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 IS UPDATED IN THIS REPORT. SECTION H6(CLINICAL CODE) IS CORRECTED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03643 | 2021-09-01 | 2021-07-02 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE LUNG CANCER. THE MEDICAL INTERVENTION RECEIVED IN RESPONSE TO THIS EVENT ARE CURRENTLY UNKNOWN. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE LUNG CANCER. IT IS UNKNOWN WHETHER THE PATIENT RECEIVED MEDICAL INTERVENTION. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03626 | 2021-09-01 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO AUTO W/BT, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | AUG500S15 | AUG500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03574 | 2021-09-01 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING OF BLACK PARTICLES ON HIS DEVICE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. AFTER THE THIRD ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, THE CUSTOMER DECLINED TO RESPOND TO THE GFE RELATED QUESTIONS AND TERMINATED THE CALL. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION B5 WAS INCORRECT IN INITIAL REPORT, IT SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK PARTICLES ON HIS DEVICE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03583 | 2021-09-01 | 2021-08-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO AUTO W/BT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSG500S11 | DSG500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03581 | 2021-09-01 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED ON THIS DEVICE IN MDR 2518422-2021-02461. THIS REPORT WAS SUBMITTED AS A DUPLICATE REPORT OF THE PREVIOUSLY SUBMITTED REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03575 | 2021-09-01 | 2021-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03584 | 2021-09-01 | 2021-08-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO AUTO W/BT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSG500S11 | DSG500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. . |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03573 | 2021-09-01 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11 | DSX700T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGATION ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED NASAL/THROAT IRRITATION OR SORENESS, DEVICE PUT BLACK PARTICLES IN HIS MOUTH, HUMIDIFIER IS NOT WORKING, NASAL CONGESTION AND NOSE BLEEDS. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 (HEALTH EFFECT - CLINICAL CODE) (MISSED TO CAPTURE PATIENT ALLEGATION IN PREVIOUS REPORT) HAS BEEN CORRECTED/UPDATED AND (TYPE OF INVESTIGATION,INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS) HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03598 | 2021-09-01 | 2021-08-10 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP BASAL CELL CARCINOMA. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING THE PATIENT DEVELOPED BASAL CELL CARCINOMA RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT. THIS EVENT IS ASSESSED AS SERIOUS BUT NOT RELATED TO THE DEVICE. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LAB FOR EVALUATION. THE EXTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE MANUFACTURER VISUALLY OBSERVED VERY LITTLE CONTAMINATION ON THE OUTSIDE OF THE DEVICE. AIR INLET PORT HAD SOME DIRT/DEBRIS AND IT WAS VISIBLE FROM OUTSIDE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED AND THE MANUFACTURED OBSERVED EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION AND BREAKDOWN IN THE BLOWER BOX, ON THE INSIDE SURFACE OF THE BLOWER UPPER CAP, AND IN THE BLOWER BOX. A BLACK, TAR-LIKE SUBSTANCE CONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM WAS ALSO FOUND ON AND IN THE BLOWER ASSEMBLY. THE MANUFACTURER ALSO OBSERVED THE TAR-LIKE SUBSTANCE FORMED A TEXTURED SURFACE IN THE UPPER BLOWER ASSEMBLY, THAT APPEARS SIMILAR TO PITTING, BUT NO PITTING WAS FOUND IN THE BLOWER ASSEMBLY MATERIAL. THE MANUFACTURER POWERED ON THE DEVICE USING A KNOWN GOOD POWER SUPPLY AND AC POWER CORD WHICH SHOWED THE BASE UNIT PROVIDED AIRFLOW. A HIGH-PITCHED WHINING NOISE WAS OBSERVED WHEN THERAPY WAS INITIATED AND THE MANUFACTURER ALSO OBSERVED THREE INSTANCES OF ERRORS (E-20 ERR_MOTOR_SPINUP_FLUX_LOW ). THE MANUFACTURER CONFIRMED EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION IN THE DEVICE. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03569 | 2021-09-01 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP W/HUMID, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | CAX200H12 | CAX200H12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03596 | 2021-09-01 | 2021-07-14 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THE PATIENT IS UNDERGOING CHEMOTHERAPY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03597 | 2021-09-01 | 2021-07-14 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/HT/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX400T11C | DSX400T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP NODULES IN THEIR LUNGS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP NODULES IN THEIR LUNGS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03617 | 2021-09-01 | 2021-06-30 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP COLON CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | CORRECTION TO PREVIOUS REPORT AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP COLON CANCER; HEADACHE; TROUBLE BREATHING; SPITS UP PHLEGM. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. SECTION H6 HEALTH EFFECTS: CLINICAL COD CORRECTED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03616 | 2021-09-01 | 2021-07-02 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LIVER AND PANCREATIC CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NASAI/THROAT SORENESS DIFFICULT BREATHING NAUSEA LIGHTHEADED DIZZY HEADACHE BLOODY NOSE STAGE 4 LIVER AND PANCREATIC CANCER AND PATIENT DEVICE HAS STRANGE ODOR, DEVICE/POWER CORD OR SUPPLY TOO HOT TO TOUCH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING NASAI/THROAT SORENESS DIFFICULT BREATHING NAUSEA LIGHTHEADED DIZZY HEADACHE BLOODY NOSE STAGE 4 LIVER AND PANCREATIC CANCER AND PATIENT DEVICE HAS STRANGE ODOR, DEVICE/POWER CORD OR SUPPLY TOO HOT TO TOUCH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, FOLLOWUP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03613 | 2021-09-01 | 2021-08-03 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03614 | 2021-09-01 | 2021-07-02 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP INTENSE SHAKING. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED PATIENT TO DEVELOP INTENSE SHAKING AND HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT USING THE DEVICE. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENT OF INTENSE SHAKING WHILE USING THE DEVICE AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THESE EVENTS ARE ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURER CONCLUDES NO FURTHER ACTION IS NECESSARY. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION AND AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTIONS B1, B2 HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03642 | 2021-09-01 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP NODULES IN THE LUNGS AND SHORTNESS OF BREATH. THE MEDICAL INTERVENTION RECEIVED IN RESPONSE TO THIS EVENT ARE CURRENTLY UNKNOWN. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP NODULES IN THE LUNGS AND SHORTNESS OF BREATH. THE MEDICAL INTERVENTION RECEIVED IN RESPONSE TO THIS EVENT ARE CURRENTLY UNKNOWN. THE REPORTED EVENT CAUSED THE PATIENT TO DEVELOP NODULES IN THE LUNGS AND SHORTNESS OF BREATH AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. THE PATIENT HAS SWAPPED OUT THE DEICE AT THE DISTRIBUTOR AND AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03641 | 2021-09-01 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE ORAL INFLAMMATION, SWELLING OF THE LIPS AND SORE IN THE MOUTH. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF DRUGS. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE ORAL INFLAMATION, SWELLING OF THE LIPS AND SORE IN THE MOUTH. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF DRUGS. THERE IS NO PATIENT HARM OR INJURY. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM (ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR) SECTION B2 WAS CHANGED TO BLANK (PREVIOUSLY IT WAS OTHER SERIOUS) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION SECTION H6 HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03646 | 2021-09-01 | 2021-08-04 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE A STROKE AND HEART STOPPAGE. THE MEDICAL INTERVENTION RECEIVED IN RESPONSE TO THIS EVENT ARE CURRENTLY UNKNOWN. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON MAY 23, 2024, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING THAT THE PATIENT HAD A STROKE AND HEART STOPPAGE RELATED TO THE CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGATION OF EYE IRRITATION, INTRAOCULAR HEMORRHAGE AND ASTHMA (NEW OR WORSENING). SECTION H6 HEALTH EFFECTS: CLINICAL CODES HAS BEEN UPDATED IN THE REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03640 | 2021-09-01 | 2021-07-07 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | CAX500T12 | CAX500T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE LUNG CANCER. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A SURGERY FOR RIGHT LOBECTOMY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED LUNG CANCER. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A SURGERY FOR RIGHT LOBECTOMY. THE REPORTED EVENT OF LUNG CANCER AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURER'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURER CONCLUDES NO FURTHER ACTION IS NECESSARY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. UNIT SCRAPPED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03659 | 2021-09-01 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2200X11B | DS2200X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, OBSTRUCTION CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S BLOWER MOTOR, FLOW SENSOR, PROPORTIONAL VALVE, SOLENOID VALVE AND TUBING WERE REPLACED TO ADDRESS THE ISSUE. THE DEVICE WAS CONTAMINATED WITH A GREASY SUBSTANCE THROUGHOUT THE AIR FLOW PATH. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03663 | 2021-09-01 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1032800 | 1032800 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS FAILING A CALIBRATION TEST. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSL REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR WAS FAILING A CALIBRATION TEST. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT COULD NOT BE CONFIRMED OR DUPLICATED. THE DEVICE OPERATED AND ALARMED AS DESIGNED. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03644 | 2021-09-01 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TOOTH LOSS AND UNSPECIFIED LUNG DAMAGE. THE MEDICAL INTERVENTION RECEIVED IN RESPONSE TO THIS EVENT ARE CURRENTLY UNKNOWN. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON OCT 24, 2024, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DEGRADED AND CAUSED THE PATIENT TOOTH LOSS AND UNSPECIFIED LUNG DAMAGE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03664 | 2021-09-02 | 2021-08-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR "HIGH INTERNAL OXYGEN". THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03679 | 2021-09-02 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | IT WAS INITIALLY REPORTED, THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS FOR ADDITIONAL INFORMATION OR TO HAVE THE DEVICE RETURNED FOR EVALUATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03694 | 2021-09-02 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400S11 | DSX400S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. IN THE PREVIOUSLY SUBMITTED REPORT SECTION B5 WAS INCOMPLETE, SECTION B5 IS- THE PATIENT ALSO ALLEGED VISUALIZATION OF PARTICLES AND HAVING COUGH. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL AND INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF FINE WHITE DUST LIKE AND HAIR LIKE PARTICULATES ON BLOWER, FINE BLACK DUST LIKE PARTICULATES CONSISTENT WITH KERATIN AROUND BLOWER SEAL, FLUID INGRESS IN THE BLOWER BOX AND ON THE BLOWER. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. THE DEVICE WAS HOOKED UP TO POWER SUPPLY, AIRFLOW WAS VERIFIED AND THE DEVICE OPERATED PROPERLY. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND NO ERRORS LOGGED. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF FINE WHITE DUST LIKE AND HAIR LIKE PARTICULATES ON BLOWER, FINE BLACK DUST LIKE PARTICULATES CONSISTENT WITH KERATIN AROUND BLOWER SEAL, FLUID INGRESS IN THE BLOWER BOX AND ON THE BLOWER. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. SECTION D9, G3 AND H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03647 | 2021-09-02 | 2021-08-04 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE SINUSITIS AND SHORTNESS OF BREATH. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTIONS FOR ANTIBIOTICS. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT ALSO ALLEGED COUGH AND THROAT IRRITATION. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL AND INTERNALLY PART OF THE DEVICE AND FOUND A SLIGHT INDISCRIMINATE ODOR DURING THERAPY. MISSING MODEM SIDE PANEL DOOR. SLIGHT UNKNOWN WHITE AND BLACK CONTAMINATION (DUST-LIKE), INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM, AT THE AIR INPUT OF THE BLOWER BOX. ALSO, SOME POTENTIAL TINY BLACK HAIR OR FIBERS WERE AT THE INPUT OF THE BLOWER BOX. SLIGHT DUST-LIKE CONTAMINATION ON TOP OF THE BLOWER MOTOR. TINY UNKNOWN WHITE SPECS OF CONTAMINATION ON THE BOTTOM THE BLOWER BOX. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE IS NO VISIBLE DAMAGE OR FUNCTIONALITY FAILURES OF THE DEVICE, WHICH SUGGEST THAT THE SOURCE OF CONTAMINATION WAS EXTERNAL TO THE DEVICE. THERE IS NO VISIBLE DAMAGE OR FUNCTIONALITY FAILURES OF THE DEVICE, WHICH SUGGEST THAT THE SOURCE OF CONTAMINATION WAS EXTERNAL TO THE DEVICE. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03693 | 2021-09-02 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/WIFI,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11W | DSX500T11W |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALSO ALLEGED TO COUGH. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. AFTER THE THIRD ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, RECIEVED AN AUTOMATED VOICE BUT COULDNOT CONNECT THE REPORTER THROUGH CALL OR VOICE MAIL. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03656 | 2021-09-02 | 2021-08-04 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE BLOOD CANCER. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THIS EVENT ARE CURRENTLY UNKNOWN. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION IN A VOLUNTARY MEDWATCH (MW5102535), ALLEGING THE PATIENT REPORTED BLOOD CANCER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THIS EVENT ARE CURRENTLY UNKNOWN. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6, HEALTH EFFECT - IMPACT CODE HAS BEEN CORRECTED AND TYPE OF INVESTIGATION, INVESTIGATION FINDINGS INVESTIGATION CONCLUSIONS HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03692 | 2021-09-02 | 2021-08-05 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX700T11W | DSX700T11W |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP INCREASING CO2 LEVELS. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE WITH A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP INCREASING CO2 LEVELS. THE PATIENT ALLEGEDLY WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. CORRECTION TO SECTIONS B2, D1 AND H1 WERE UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03576 | 2021-09-02 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS INITIALLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE - THE PATIENT ALLEGED PARTICLES IN TUBING. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO CONTAMINATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER AND FOUND CONTAMINATION IN THE ISO PORT. THERE WERE ALSO SIGNS OF WATER INGRESS IN THE BLOWER BOX AND ON THE BLOWER MOTOR. THERE WAS A SLIGHT FLORAL SMELL IN THE BLOWER BOX AS WELL. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERORRS. THE MANUFACTURER CONCLUDES THE DEVICE HAS CONTAMINATION CONSISTENT WITH WATER INGRESS. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. IN THIS REPORT, SECTION D9, G3, H3, H6 HAS BEEN UPDATED OR CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03655 | 2021-09-02 | 2021-07-15 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE SINUSITIS AND RESPIRATORY ISSUES. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A SURGERY FOR SINUS ISSUES AND PRESCRIPTIONS FOR ANTIBIOTICS AND STEROIDS. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE SINUSITIS AND RESPIRATORY ISSUES. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A SURGERY FOR SINUS ISSUES AND PRESCRIPTIONS FOR ANTIBIOTICS AND STEROIDS. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03578 | 2021-09-02 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN CHAMBER/TUBING, LIGHTHEADEDNESS, RESPIRATORY PROBLEMS, DRY COUGH, ASTHMA, DIFFICULTY BREATHING AND BRONCHITIS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING PARTICLES IN CHAMBER/TUBING, LIGHTHEADEDNESS, RESPIRATORY PROBLEMS, DRY COUGH, ASTHMA, DIFFICULTY BREATHING AND BRONCHITIS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTION G1-MANUFACTURER CONTACT OFFICE INFO WAS PREVIOUSLY INCORRECT AND HAS NOW BEEN CORRECTLY UPDATED. SECTION G3 WAS PREVIOUSLY INCORRECT. SECTION H6 WAS CORRECTED UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03612 | 2021-09-02 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID/HT, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | AUX400T15 | AUX400T15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM AND PATIENT HAVEING SORE THROAT. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. SECTION H6 CODES CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT HAS BEEN CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03628 | 2021-09-02 | 2020-02-20 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA, KIDNEY, AND LIVER ISSUES. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION STATING THAT THE DEVICE WAS SCRAPPED AND WILL NOT BE RETURNING TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. PATIENT ALSO ALLEGED COUGH. SECTION D1, G4 AND H6 ARE UPDATED WITH CORRECTED DATA. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03666 | 2021-09-02 | 2021-08-20 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP, GER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DEX700S13 | DEX700S13 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A HEADACHE. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED THAT A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM ALLEGEDLY BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A HEADACHE. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER QUALITY PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. DURING THE INITIAL VISUAL INSPECTION OF THE DEVICE MINIMAL WEAR AND TEAR WERE FOUND. THE MANUFACTURER FOUND EVIDENCE OF CONTAMINANTS PRESENT ON THE OUTSIDE OF THE UNIT, ON THE INLET FILTER, IN THE FILTER INLET AREA, AND ON THE HUMIDIFIER LATCH LEDGE. THE DEVICE WAS DISASSEMBLED FOR INTERNAL VISUAL INSPECTION. THE MANUFACTURER FOUND EVIDENCE OF CONTAMINATES PRESENT INSIDE THE UNIT. THE MANUFACTURER ALSO FOUND EVIDENCE OF 7 TYPES OF CONTAMINATES OF WHICH 3 WERE FOUND IN THE AIRPATH. THE MANUFACTURER FINDS THAT THESE CONTAMINATES WERE LIKELY SOURCED FROM EXTERNAL ENVIRONMENTAL CONDITIONS. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND CONFIRMED THE DEVICE PROVIDED FLOW AND ACCESSED THE ERROR LOG. THE MANUFACTURER CONCLUDES THERE IS NO EVIDENCE OF SOUND ABATEMENT FOAM BREAKDOWN WITHIN THE DEVICE. THE CONTAMINATION FOUND IN THE AS OBSERVED IN THE BLOWER AND AIR PATH AND ARE CONSISTENT WITH BEING FROM AN EXTERNAL SOURCE. SECTION D4 HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03577 | 2021-09-02 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11 | DSX700H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING V PARTICLESIN TUBING AND WATER, LIPS BEEN REAL IRRITATED AND SEEN THE PARTICLES, DEVICE IS NOISY RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03579 | 2021-09-02 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING HEADACHE, SINUS ISSUES AND TIREDNESS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 HEALTH EFFECT CLINICAL CODE (MISSED TO UPDATE IN THE INITIAL MDR), TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
MW5103715 | 2021-09-02 | 2021-08-31 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | I FILED FOR A REPLACEMENT OF MY PHILLIPS DREAMSTATION CPAP. PHILLIPS IS PROCESSING REPLACEMENT. I HAD NO IDEA THIS DEVICE HAD A RECALL BACK IN JUNE. I USED THIS DEVICE AND I NOTICED ONE NIGHT (B)(6) 2021. I SLEPT TERRIBLE AND FELT LIKE I COULDN'T BREATHE. I WENT TO WORK WITH A HEADACHE. WHEN I GOT HOME I INSPECTED MY CPAP AND CHECK THE FILTERS AND TO MY SURPRISE THEY BOTH WERE PITCH BLACK IN COLOR. I ATTEMPTED THE WASH THE FILTERS AND NOTICE THE BLACK SUBSTANCE WAS NOT COMING OUT. I ENDED UP REPLACING THEM AND THOUGHT NOTHING OF IT. I THEN DISCOVERED ON 8/31/2021 THAT MY DEVICE HAD A RECALL. I FEEL LIKE PHILIPS AND MY PROVIDER FAILED TO SEND OUT PROPER NOTIFICATIONS. I MEAN JUNE AND I DON'T FIND OUT UNTIL 8/31/2021. THIS DEVICE IS CELLULAR BASED. WHY DIDN'T THEY SEND OUT A ERROR CODE OR AT LEAST SEND ME A LETTER OR EMAIL. SOMETHING. I PLAN ON FOLLOWING UP WITH MY DOCTOR ON (B)(6)2021 TO DISCUSS MY OPTIONS SINCE MY SLEEP WILL BE HORRIBLE WITHOUT MY DEVICE. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-09
MW5103709 | 2021-09-02 | 2018-12-01 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DREAMSTATION AUTO CPAP | DSX500H11C |  | I HAVE A PHILIPS DREAMSTATION AUTO CPAP MACHINE (SERIAL NUMBER (B)(4)) AND HAPPENED TO FIND THE PRODUCT RECALL NOTICE THIS AFTERNOON AND AM VERY CONCERNED ABOUT HAVING HAD USED IT FOR THE PAST 3+ YEARS. I HAVE NOTICED PARTICULATE MATTER IN THE AIRWAY AS DESCRIBED IN THE RECALL NOTICE, AND I HAVE HAD AN ONGOING CHRONIC COUGH FOR YEARS THAT HAS YET TO BE RESOLVED DESPITE HAVING BEEN TO MULTIPLE SPECIALISTS INCLUDING PULMONOLOGISTS, ALLERGIST, ENT, GASTROENTEROLOGIST, ETC. I AM ESPECIALLY ANGERED THAT THE RECALL WAS DATED FROM JUNE OF THIS YEAR, BUT I HAVEN'T BEEN CONTACTED BY PHILIPS, OR MY INSURANCE COMPANY ((B)(6)), OR THE DME EQUIPMENT PROVIDER WHO MY INSURANCE COMPANY CONTRACTED THE PURCHASE OF THE UNIT THROUGH (B)(6), OR MY DME EQUIPMENT REPLACEMENT MAIL ORDER COMPANY (B)(6). I CAN UNDERSTAND ONE OF THE INVOLVED COMPANIES DROPPING THE BALL AND NOT NOTIFYING ME, BUT TO HAVE NONE OF THEM CONTACT ME ABOUT AN FDA RECALL IS SIMPLY RIDICULOUS. AS FOR NOW, I'M FOLLOWING THE INSTRUCTIONS IN THE RECALL (TO CEASE USING THE MACHINE UNTIL CONTACTED) BUT I'M HONESTLY NOT HOPEFUL THAT I'LL HEAR ANYTHING FROM PHILIPS GIVEN HOW THEY'VE HANDLED COMMUNICATION THUS FAR. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5103717 | 2021-09-02 | 2021-06-01 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | J275963004A87 |  |  | FOR THE PAST TWO YEARS FIGHTING A SINUS INFECTION THAT WOULD NOT GO AWAY WITH ANTIBIOTICS. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03580 | 2021-09-02 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING PNEUMONIA RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL PART AND INTERNAL PART OF THE DEVICE AND OBSERVED UNKNOWN WHITE AND BLACK CONTAMINATION (DUST-LIKE), INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM, AT THE AIR INPUT OF THE BLOWER BOX. UNKNOWN SILVER HAIRS OR FIBRES AND SOME WATER OR LIQUID SPOTS AT THE AIR INPUT OF THE BLOWER BOX. LIGHT GRAY FILM OF UNKNOWN CONTAMINATION THROUGHOUT THE INSIDE OF THE ENCLOSURE AND ON TOP OF THE BLOWER BOX. UNKNOWN DUST-LIKE CONTAMINATION ON TOP OF THE BLOWER MOTOR AND THE BLOWER MOTOR SEAL. POTENTIAL MINERAL DEPOSITS OR DRIED LIQUID SPOTS ON THE BOTTOM OF THE BLOWER MOTOR AND ON THE BOTTOM OF THE BLOWER BOX, INDICATING POTENTIAL WATER INGRESS. BLACK CONTAMINATION, CONSISTENT WITH KERATIN, AT THE AIR OUTLET OF THE BLOWER BOX, TINY UNKNOWN BLACK (INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM) AND WHITE SPECS OF CONTAMINATION ON THE BOTTOM THE BLOWER BOX. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION, BUT DUST OR DIRT CONTAMINATION WERE OBSERVED AND ARE CONSISTENT WITH BEING FROM AN EXTERNAL SOURCE. SECTION H6 HEALTH EFFECT CLINICAL CODE WAS UPDATED AND CORRECTED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03668 | 2021-09-02 | 2021-08-20 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, GER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DEX500S13 | DEX500S13 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO NEED A LIVER TRANSPLANT. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03691 | 2021-09-02 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING FILTER GETS BLACK RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
MW5103694 | 2021-09-02 | 2017-10-01 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500H11 | 3194661 |  | I BEGAN USING A PHILIPS RESPIRONICS DREAMSTATION AUTO CPSP HUM DOM (SERIAL NUMBER (B)(4), REF DSX500H11) IN (B)(6) 2017. WITHIN THREE MONTHS I BEGAN GAINING A WEIGHT, WHICH WAS THE OPPOSITE OF WHAT I HAD HOPE FOR NOW THAT I WAS GAINING MORE SLEEP. WITHIN 9 MONTHS I BEGAN COMPLAINING TO BOTH MY ALLERGY CLINIC AND CPAP DOCTORS MY BREATHING WAS BECOMING MORE LABORED AND I WAS HAVING AN INORDINATE AMOUNT OF NON-STOP NASAL DRAINAGE. IN 2018-2019 MY EOSINOPHILS INCREASED TO EXTREMELY HIGH LEVELS AND MY ALLERGY DOCTOR DIAGNOSED ME WITH ASTHMA. I HAD ONLY HAD SEASONAL POLLEN ALLERGIES PRIOR TO USING THE CPAP. ALTHOUGH I'VE NEVER BEEN HOSPITALIZED, MY ALLERGY DOCTOR HAS WARNED ME ON SEVERAL VISITS I SHOULD HAVE GONE TO THE EMERGENCY ROOM WITH MY LOW LEVELS OF OXYGEN. BETWEEN 2018-2021 MY LUNG HEALTH HAS WORSENED. I HAVE BEEN PRESCRIBED FASENRA IN 2019 AND DUPIXENT IN 2020. I TAKE ZYRTEC AND MONTELUKAST DAILY, USE A DULERA INHALER DAILY AND ALBUTEROL IN MY NEBULIZER. IN (B)(6) 2019 I TOOK ALLERGY BLOOD TEST FOR FOOD ALLERGIES AND ALTHOUGH I ONLY HAD MILD AND MODERATE, I REFRAIN FROM ANY FOODS THAT CAUSE ANY INFLAMMATION FOR ME (DAIRY, EGG WHITES AND WHEAT ARE A FEW). WHILE I DID SEE A MARKED IMPROVEMENT OF INFLAMMATION IN MY JOINTS, MY BREATHING REMAINS LABORED. AFTER A YEAR OF STRICT DIET, MY EOSINOPHILS LEVELS CAME DOWN SLIGHTLY SHOULD BE. IN (B)(6) 2021 I HAD A CALCIUM SCORE TEST IN WHICH THE REPORT NOTED THEY HAS A PARTIAL SCAN OF THE LUNGS AND THERE WAS 'ATELECTASIS' NOTED. I HAVE NEVER HAD THIS NOTED IN PREVIOUS LUNGS SCANS. AS I TYPE THIS MY LUNGS ARE COMPROMISED, I WILL BE AIMING TO DO FOUR ALBUTEROL TREATMENTS TODAY TO GET MY OXYGEN LEVELS UP. I AM A WHITE FEMALE, (B)(6) YEARS OF AGE, DO NOT SMOKE OR DRINK ALCOHOL. I WEIGH (B)(6). AND AM 5'0". I HAD AN UNEXPLAINED 32 LB. WEIGHT GAIN WITHIN THE FIRST 4 MONTHS OF USING THE CPAP. I EAT A HEALTHY ORGANIC BALANCED DIET, RARELY EATING JUNK FOOD SINCE IT COMMONLY CONTAINS DAIRY, EGGS OR WHEAT. WE USE SPECIFIC ALLERGY FREE CLEANING PRODUCTS AND USE ORGANIC ALLERGY FREE BEDDING AS SUGGESTED BY THE ALLERGY CLINIC. I USED TO EXERCISE DAILY, AND ALTHOUGH I AM STILL ACTIVE, AEROBIC ACTIVITY IS LIMITED DUE TO MY LUNG CAPACITY. I JUST KEEP MY PACE A BIT SLOWER. I HAVE CONSISTENTLY USED MY CPAP FROM DAY ONE. I REMEMBER THE FIRST MONTH OF SLEEP AS THE BEST I EVER HAD. THAT CHANGED AS TIME WENT ON. I OFTEN WAKE UP IN THE MIDDLE OF THE NIGHT COUGHING, SOMETIMES ACCOMPANIED BY A HEADACHE. I WAKE UP OFTEN AND JUST ASSUMED THAT WAS DUE TO DREAMING, BUT SOMETIMES WONDERED IF IT WAS AN ALLERGIC REACTION TO THE DEVICE ITSELF. I CAN PROVIDE MEDICAL RECORDS AND BLOODS TEST RESULTS FOR THIS TIME PERIOD AND BEFORE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03712 | 2021-09-03 | 2021-08-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03714 | 2021-09-03 | 2021-08-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2200X11B | DS2200X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S DISPLAY SCREEN WAS DAMAGED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S DISPLAY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03713 | 2021-09-03 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2200X11B | DS2200X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS DELIVERING BREATHS TOO QUICKLY. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S INLET AND PARTICULATE FILTERS WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03711 | 2021-09-03 | 2021-08-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03708 | 2021-09-03 | 2021-08-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03709 | 2021-09-03 | 2021-08-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03618 | 2021-09-03 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SENSOR BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03707 | 2021-09-03 | 2021-08-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03706 | 2021-09-03 | 2021-08-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03705 | 2021-09-03 | 2021-08-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03704 | 2021-09-03 | 2021-08-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03703 | 2021-09-03 | 2021-08-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S POWER MANAGEMENT BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03702 | 2021-09-03 | 2021-08-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICES LCD DISPLAY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03700 | 2021-09-03 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION OCCURRED RELATED LOW INSPIRATORY PRESSURE. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE VENTILATOR'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03619 | 2021-09-03 | 2021-08-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | FR2110X14B | FR2110X14B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S 3-WAY SOLENOID VALVE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03710 | 2021-09-03 | 2021-08-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03676 | 2021-09-03 | 2021-08-09 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102653) ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP BREATHING ISSUES. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION B5 SHOULD BE PREVIOUSLY REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT WAS FACING BREATHING ISSUE, BUILD UP OF MUCUS IN THE THROAT OVER SEVEN YEARS AND PROBLEM CONTINUED TO WORSEN. PATIENT BEGAN TO HAVE HEART ISSUES AND ENDED UP IN HOSPITAL WITH PNEUMONIA ALSO HAS COVID SYMPTOMS. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 HEALTH EFFECT - CLINICAL CODE ADDED IN THIS REPORT. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03658 | 2021-09-03 | 2021-08-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | NDX500H15 | NDX500H15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVING INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. ON THE PREVIOUSLY SUBMITTED REPORT, SECTION D1 HAD INCORRECT DATA. SECTION D1 AND SECTION D4 WERE UPDATED/CORRECTED IN THIS REPORT. THE MANUFACTURER RECEIVED INFORMATION STATING THAT THE DEVICE WAS SCRAPPED AND WILL NOT BE RETURNING TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03673 | 2021-09-03 | 2021-07-02 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/WIFI, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11W | DSX500H11W |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER OF THE TONGUE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER OF THE TONGUE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE REPORTED EVENT OF SQUAMOUS CELL CANCER OF THE TONGUE AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03677 | 2021-09-03 | 2021-07-02 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP PNEUMONIA. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO DEVELOP PNEUMONIA COMPLICATION OF ASTHMA, DIFFICULTY OF BREATHING/SHORTNESS OF BREATH AND WHEEZING NOISE COMING FROM HIS CHEST. THE REPORTED EVENT OF PNEUMONIA COMPLICATION OF ASTHMA, DIFFICULTY OF BREATHING/SHORTNESS OF BREATH AND WHEEZING NOISE COMING FROM HIS CHEST AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03728 | 2021-09-03 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN AIR CHAMBER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03649 | 2021-09-03 | 2020-11-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE DIFFICULTY IN BREATHING. THE PATIENT WAS REPORTEDLY HOSPITALIZED IN RESPONSE TO THE EVENT. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE DIFFICULTY IN BREATHING AND LUNG CANCER. THE PATIENT WAS REPORTEDLY HOSPITALIZED IN RESPONSE TO THE EVENT. AFTER FIRST ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION, THE CUSTOMER HAS RESPONDED ON (B)(6) 2022 STATED THE DEFECTIVE DEVICE WAS DISCARDED. ADDITION TO THAT, PATIENT HAS GOT THE REPLACEMENT. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03651 | 2021-09-03 | 2021-08-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO AUTO CPAP W/HUMID, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAG500H12 | CAG500H12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03652 | 2021-09-03 | 2021-08-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP/HUM/HT/CELL,CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | CAX700T12C | CAX700T12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN THE INITIAL REPORT, SECTION B5 WAS INCORRECT, THE CORRECT B5 SHOULD BE- THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED SEEING BLACK PARTICLES IN THE DEVICE AND HAVING A COUGH. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL AND INTERNAL VISUAL INSPECTION OF DEVICE WERE COMPLETED BY THE MANUFACTURER AND FOUND EVIDENCE OF WATER INGRESS ON PCA. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERROR LOGGED. THE MANUFACTURER CONCLUDES THE CONTAMINATE FOUND WAS CONSISTENT WITH WATER INGRESS ON PCA. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03653 | 2021-09-03 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID/HT, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | CAX400T12 | CAX400T12 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THIS FOLLOW-UP REPORT IS BEING FILED TO INCLUDE A CORRECTION FOR PREVIOUS REPORT 2518422-2021-03653. CONCOMINANT MEDICATION HAS BEEN ADDED TO SECTION H10. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03674 | 2021-09-03 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SINUS AND UPPER RESPIRATORY INFECTIONS. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 05, 2024, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING HAVING SINUS AND UPPER RESPIRATORY INFECTIONS,SHORTNESS OF BREATH AND DIFFICULTY BREATHING WHILE USING THE DEVICER RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTIONS B1, B2, H1 AND H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR ONLY THE PRODUCT PROBLEM (BOTH ADVERSE EVENTS AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03675 | 2021-09-03 | 2021-07-13 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG CANCER. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. EVIDENCE OF SOUND ABATEMENT WAS NOT OBSERVED ON THIS DEVICE. PIL IS UNABLE TO ADDRESS THE COMPLAINT OR SYMPTOMS. PIL FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. THE EXTERNAL INVESTIGATION SHOWED NO SIGNS OF CONTAMINATION ON THE OUTSIDE OF THE DEVICE. THE DEVICE WAS HOOKED UPTO A KNOWN GOOD POWER SUPPLY AND POWER CORDS, AIRFLOW WAS VERIFIED TO BE OPERATING CORRECTLY. THE INTERNAL INVESTIGATION OF THE HUMIDIFIER SHOWED THAT THE HUMIDIFIER TANK LATCH SPRING WAS BROKEN AND THERE WAS ALSO AN UNKNOWN CONTAMINATE ON THE RESERVOIR SEAL. THERE WERE SIGNS OF AN UNKNOWN CONTAMINATE ON THE INSIDE OF THE ENLOSURE. THERE WAS ALSO AN UNKNOWN DUST CONTAMINATE IN THE FILTER INLET PORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03697 | 2021-09-03 | 2021-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP A40, INTERNATIONAL | VENTILATOR, CONTINUOUS, MINIMAL VENTILATORY SUPPORT, FACILITY USE |  | 1111169 | 1111169 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03625 | 2021-09-03 | 2017-06-20 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102568) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP NODULES ON THEIR LUNG. THE PATIENT HAD SURGERY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03670 | 2021-09-03 | 2021-07-13 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11 | DSX400H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG CANCER, DIFFICULTY BREATHING AND COUGHING. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL PART OF THE DEVICE AND FOUND THERE WAS AN UNKNOWN STICKER RESIDUE ON THE DEVICE BAR CODE STICKER. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND FOUND CONTAMINATION IN BLOWER BOX, CONTAMINATION IN FILTER PORT, CONTAMINATION INSIDE OF BLOWER FAN. MINIMAL DUST CONTAMINATION IN THE BLOWER ASSEMBLY. THE HUMIDIFIER SHOWED A SMALL AMOUNT OF LIQUID INGRESS UNDER THE BOTTOM PLATE. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER VERIFIED FOR POWER SUPPLY AND POWER CORD, AIRFLOW WAS VERIFIED TO BE OPERATING CORRECTLY. THE ERROR LOG SHOWED THAT THERE WERE ZERO INSTANCES OF ERRORS ON THE DEVICE THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED IN THIS DEVICE. | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5103731 | 2021-09-03 | 2021-08-11 | I, n, j, u, r, y | Voluntary report | DREAM STATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | USING THE PHILIPS RESPIRONICS DREAM STATION CPAP DEVICE DEVELOPED A LACK OF SENSATION INSIDE MY UPPER AND LOWER LIP. THIS OCCURRED ROUGHLY A MONTH AFTER INSTALLING AN IN-LINE FILTER TO PROTECT AGAINST POSSIBLE DEGRADATION OF THE SOUND ABATEMENT FOAM WHICH WAS THE REASON FOR THE RECALL OF THIS DEVICE. I DISCONTINUED USING THE CPAP MACHINE A WEEK LATER ON (B)(6) 2021. FOR POINT OF INFORMATION, I BEGAN USING THIS DEVICE ON (B)(6) 2021 AND MY SLEEP THERAPY DOCTOR RECOMMENDED CONTINUING USING THE DEVICE AFTER THE RECALL OCCURRED ON JUNE 14, 2021 DUE TO MY SEVERE SLEEP APNEA. ALSO, PRIOR TO (B)(6), I HAD BEEN USING ANOTHER RECALLED PHILIPS CPAP MACHINE, THE SYSTEMONE SINCE 2014. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03672 | 2021-09-03 | 2016-02-02 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102658) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP PNEUMONIA. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. IN PREVIOUS REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DIFFICULTY OF BREATHING, SINUS INFECTION, GENERAL UPPER RESPIRATORY INFECTION, BRONCHITIS, ALLERGIES AND WAS HOSPITALIZED WITH PNEUMONIA. THE REPORTED EVENT AND ITS SEVERITY WERE REVIEWED BY THE MANUFACTURER'S CLINICAL EXPERT. BASE ON THE AVAILABLE INFORMATION AT THE TIME OF THE REVIEW, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENTS. UPON REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION, THE CUSTOMER RESPONDED THAT THEY DON'T HAVE A DEVICE ANYMORE. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HEALTH EFFECT - CLINICAL CODE HAS BEEN CORRECTED, AND TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED. SECTION B2 HAS BEEN CORRECTED / UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03720 | 2021-09-07 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/HT/CELL, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AUX400T15C | AUX400T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION B5 WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT SHOULD BE REPORTED AS FOLLOWS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED EYE WEEPING, HEADACHE AND OBSERVED BACK SPECS. THERE WAS NO PATIENT HARM OR INJURY. NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03695 | 2021-09-07 | 2020-11-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A BLOOD CLOT. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A BLOOD CLOT IN BRAIN. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURE INSPECTED EXTERNALLY OBSERVED THAT UNKNOWN DUST CONTAMINATION INSIDE THE BLOWER BOX AT THE AIR INLET, TOP ENCLOSURE, FRONT PANEL, BOTTOM ENCLOSURE, REAR PANEL. A KERATIN LIKE SUBSTANCE WAS OBSERVED ON THE BLOWER BOX. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND ONE ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE MANUFACTURER OBSERVED DUST CONTAMINATION AND KERATIN SUBSTANCES ARE CONSISTENT WITH BEING FROM AN EXTERNAL SOURCE. SECTION H6 TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03633 | 2021-09-07 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/HT/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400T11C | DSX400T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE - THE PATIENT ALLEGED PARTICLES IN AIRWAY. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL.THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION.THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 HAS BEEN UPDATED IN THIS REPORT . | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03768 | 2021-09-07 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03769 | 2021-09-07 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING SINUS INFECTIONS, VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03632 | 2021-09-07 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP W/HUM/HT/CELL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200T11C | DSX200T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03698 | 2021-09-07 | 2021-08-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO, GER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DEX400S13 | DEX400S13 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 08 2024, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. SECTION B1 WAS CORRECTED FOR BOTH ADVERSE EVENT AND PRODUCT PROBLEM. ( PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION. TO SERIOUS INJURY. SECTION H6 WAS CORRECTED AND UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03629 | 2021-09-07 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT ALLEGED OF PARTICLES IN TUBING/CHAMBER. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTRE FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF DEVICE WAS INSPECTED. THE DEVICE POWERED ON, AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. SECTION H6 (TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS) UPDATED IN THIS FOLLOW UP REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03721 | 2021-09-07 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL, AU | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | AUX500T15C | AUX500T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03761 | 2021-09-07 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION B5 SHOULD BE REPORTED AS THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES, NASAL/THROAT IRRITATION OR SORENESS RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03557 | 2021-09-07 | 2021-08-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03631 | 2021-09-07 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200T11 | DSX200T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN WATER TANK RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. AFTER THE FIRST ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, THE CUSTOMER STATED THAT THE DEVICE WILL RETURNING IT SOON. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING PARTICLES IN WATER TANK RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
MW5103765 | 2021-09-07 |  | I, n, j, u, r, y | Voluntary report | PHILIPS DREAM STATION C-PAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | INCREASED, MORE FREQUENT UPPER RESP INFECTIONS; INCREASED USE OF RESCUE INHALER; INCREASED DIFFICULTY TREATING INFECTIONS. LONGER USE OF ANTIBIOTICS. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5103762 | 2021-09-07 | 2021-03-02 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. |  |  |  | COUGH CONG; I HAD EXCESSIVE MUCUS AND SINUS TYPE ISSUES. MACHINE BECAME UNCOMFORTABLE TO USE HARD TO BREATHE WITH IT ON. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5103761 | 2021-09-07 | 2021-06-15 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS INC. |  |  |  | IN THE MIDDLE OF THE NIGHT I HAD TO REMOVE THE MASK BECAUSE THERE WAS A TERRIBLE SMELL COMING FROM THE DEVICE AND I STARTED COUGHING UNCONTROLLABLY. I USED IT SEVERAL MORE TIMES. THIS ONLY HAPPENED AFTER THE UNIT WAS ON FOR A COUPLE OF HOURS. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5103753 | 2021-09-07 | 2019-09-11 | I, n, j, u, r, y | Voluntary report | DREAMSTAAUTOCPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500T11C |  |  | I USE A DREAMSTATION CPAP MACHINE, THE MORNING OF (B)(6) 2019. I WOKE UP WITH A VERY SORE, DRY THROAT. I FELT LIKE I HAD INHALED A LOT OF SMOKE. THERE WERE BLACK PARTICLES OR SOME DUSTY SOOT VISIBLE ON MY PILLOWCASE, THE CPAP TUBING, FACE MASK AND THE FILTERS WERE BLACK. I CALLED MY SUPPLY COMPANY AND THE REP SAID SHE HAD NOT HEARD OF THIS HAPPENING BEFORE. I WAS TO GO TO THEIR STORE TO PICK UP NEW SUPPLIES. TWO DAYS LATER MY SYMPTOMS WERE MUCH WORSE. I HAD A VERY SORE THROAT, DEEP CHEST COUGH, FEVER, AND HEADACHE. DAY 2 MY SYMPTOMS WERE EVEN MORE SEVERE AND THE THIRD DAY MY FEVER SPIKED TO AROUND 102 DEGREES. I WENT TO URGENT CARE, HAD AN XRAY TAKEN AND WAS TOLD I HAD PNEUMONIA. I WAS PRESCRIBED MEDICATION AND SENT HOME. A COUPLE OF DAYS LATER I FELT BETTER. I WAS TO FOLLOW UP WITH ANOTHER XRAY WITH MY PCP IN A WEEK. I CONTACTED MY SLEEP APNEA DOCTOR'S OFFICE REGARDING THE INCIDENT AND WAS TOLD BY THE PERSON WHO ANSWERED THE PHONE TO IMMEDIATELY STOP USING THE CPAP MACHINE. SHE ARRANGED FOR ME TO GET A REPLACEMENT MACHINE. I KEPT THE DEFECTIVE MACHINE AND STILL HAVE IT. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5103749 | 2021-09-07 |  | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP DEVICE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | DSX200H11 | DSX200H11 |  | I HAVE BEEN USING ONE OF THE RECALLED PHILIPS CPAP DREAM STATION DEVICES FOR YEARS. YESTERDAY I ACCIDENTALLY FOUND OUT ABOUT THE RECALL WHILE ON THE INTERNET. I HAD NOT HAD ANY NOTIFICATION FROM PHILIPS ABOUT THE RECALL. I HAVE NOW REGISTERED MY DEVICE WITH PHILIPS AS INSTRUCTED. I CHECKED THE AREAS WHERE THE BLACK SPECKS FROM THE DEFECTIVE FOAM HAVE BEEN REPORTED, AND THEY ARE PRESENT IN MY DEVICE. FOR THE LAST SEVERAL MONTHS I HAVE NOTICED SOME RESPIRATORY SYMPTOMS AS DESCRIBED IN THE RECALL NOTICE. THESE INCLUDE A FEELING OF DRYNESS AND IRRITATION IN MY THROAT AND UPPER CHEST AND A FREQUENT FEELING OF NEEDING TO CLEAR MY THROAT. MY EYES HAVE ALSO BEEN IRRITATED. AT FIRST I ATTRIBUTED THIS TO SEASONAL ALLERGIES. HOWEVER, I HAD NOTICED IN THE LAST FEW MONTHS THAT THE USE OF OTC ANTIHISTAMINES AND EYE DROPS DO NOT SEEM TO BE VERY EFFECTIVE IN PREVENTING THE OCCURRENCE OF THESE SYMPTOMS. THIS IS UNUSUAL BECAUSE THEY HAVE BEEN VERY EFFECTIVE IN THE PAST FOR TREATING ANY ALLERGY SYMPTOMS. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5103744 | 2021-09-07 | 2019-12-17 | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | REDSXH |  |  | I EXPERIENCED TROUBLE BREATHING, FEVER, HEADACHES, DIZZINESS, NASAL CONGESTION, POSTNASAL DRIP, AND FATIGUE. THIS SENT ME TO URGENT CARE, WHERE I WAS GIVEN A BREATHING TREATMENT AND PRESCRIBED ANTIBIOTICS, AN INHALER, AND FLONASE. WHEN MY SYMPTOMS CAME BACK, I WAS GIVEN AN MRI WHERE THEY FOUND INFECTIOUS FLUID IN MY SKULL CAVITY. I WAS TREATED WITH ANTIBIOTICS TWICE FOR THIS SAME CONDITION BECAUSE IT CAME BACK AND DOCTORS DO NOT KNOW WHY. IT IS COMING BACK AGAIN. I HAD ALSO EXPERIENCED A POPPING AND PULLING SENSATION IN MY LEFT EAR THAT WAS CAUSED BY A MUCOSAL THICKENING / INFLAMMATORY REACTION OF MY LEFT MAXILLARY SINUS. THIS WAS ALSO FOUND IN MY MRI RESULTS. THIS STEMS FROM EITHER TRAUMA, INFECTION, CHEMICAL AGENTS, ALLERGIES OR NEOPLASM. |  | Y | N | LAY USER/PATIENT |  | 2021-09
2518422-2021-03555 | 2021-09-07 | 2021-08-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03630 | 2021-09-07 | 2021-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX200H11C | DSX200H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE:THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DIRTY FILTER, PARTICLES, FILTERS ARE TURNING BLACK RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03558 | 2021-09-07 | 2021-08-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03556 | 2021-09-07 | 2021-08-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03552 | 2021-09-07 | 2021-08-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03553 | 2021-09-07 | 2021-08-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03678 | 2021-09-07 | 2021-07-16 | I, n, j, u, r, y | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED BREATHING ISSUES. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF HOSPITALIZATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVED INFORMATION ALLEGING A VENTILATOR'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED BREATHING ISSUES. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF HOSPITALIZATION. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03681 | 2021-09-07 | 2021-03-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03680 | 2021-09-07 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A MALFUNCTION OF A VENTILATOR'S EXTERNAL FLOW SENSOR. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS EVALUATED BY A FIELD SERVICE ENGINEER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S EXTERNAL FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03554 | 2021-09-07 | 2021-08-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S NAVIGATIONAL BUTTON WAS OBSERVED. THE DEVICE'S FRONT PANEL/KEYPAD LED BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03684 | 2021-09-07 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A MALFUNCTION OF A VENTILATOR'S EXTERNAL FLOW SENSOR. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS EVALUATED BY A FIELD SERVICE ENGINEER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S FLOW SENSOR ASSEMBLY WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A MALFUNCTION OF A VENTILATOR'S EXTERNAL FLOW SENSOR AND WAS EVALUATED BY A FIELD SERVICE ENGINEER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S FLOW SENSOR ASSEMBLY WAS REPLACED TO ADDRESS THE ISSUE. ADDITIONAL INFORMATION RECEIVED INDICATES THAT THE EXTERNAL FLOW SENSOR CABLE WAS ALSO REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03685 | 2021-09-07 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS NOT WORKING. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. . | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR WAS NOT WORKING. THE DEVICE WAS NOT IN PATIENT USE. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03683 | 2021-09-07 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A FAILURE OF A VENTILATOR TO POWER ON. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A FAILURE OF A VENTILATOR TO POWER ON. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS EVALUATED BY A THIRD PARTY FIELD SERVICE ENGINEER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03818 | 2021-09-08 | 2021-09-01 | D, e, a, t, h | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID/HT, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX400T12 | CAX400T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT DEATH. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. IT WAS ALLEGED THAT THE PATIENT DIED DUE TO ESOPHAGEAL CANCER. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION D4 UNIQUE IDENTIFIER (UDI) # CORRECTED IN THIS REPORT. SECTION G1 CORRECTED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03819 | 2021-09-08 | 2021-08-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | AUX500T15 | AUX500T15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER 88058. | SECTION B5 WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT SHOULD BE REPORTED AS FOLLOWS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED BLACK PARTICLES IN WATER CHAMBER. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. SECTION H6 HEALTH EFFECT - CODE CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT HAS BEEN CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03785 | 2021-09-08 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A COUGH AND SINUS INFECTION. THE PATIENT WAS PRESCRIBED ANTIBIOTICS AND STEROIDS IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 05, 2024 AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A COUGH AND SINUS INFECTION. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03814 | 2021-09-08 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200H11C | DSX200H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DEVICE HAS STRANGE ODOR, HEADACHE, FEELS LIKE NOT ENOUGH AIR COMING THROUGH SINUS INFECTION, STATED CAN NOT REALLY BREATHE WHILE USING THE MACHINE AND SEEING PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS UPDATED REPORT WILL BE FILED. CORRECTED INFORMATION PROVIDED IN B5 AND H6, THE INFORMATION (THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEVICE HAS STRANGE ODOR, HEADACHE, FEELS LIKE NOT ENOUGH AIR COMING THROUGH SINUS INFECTION, STATED CAN NOT REALLY BREATHE WHILE USING THE MACHINE AND SEEING PARTICLES RELATED TO A CPAP DEVICE) WAS NOT INCLUDED IN INITIAL REPORT. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03813 | 2021-09-08 | 2021-08-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. IN THE INITIAL REPORT SECTION B5 WAS NOT UPDATED, THE CORRECT B5 SHOULD BE - THE PATIENT HAS ALLEGED BLACK PARTICLES IN THE DEVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN THE INITIAL REPORT H10 SECTION WAS MARKED INCORRECTLY. THE DEVICE ALONG WITH A HUMIDIFIER WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF LIGHT BEIGE DUST OR DIRT IN THE AIR INTAKE CANAL, ON THE INSIDE OF THE BLOWER BOX, ON THE INSIDE SURFACE OF THE BOTTOM ENCLOSURE, TWO SMALL BROWN DECEASED INSECTS ON THE INSIDE OF THE BOTTOM ENCLOSURE SURFACE SURROUNDED BY POTENTIAL BROWN FIBERS, TWO VERY SMALL DARK SPEC CONTAMINATES (NOT SOUND ABATEMENT FOAM) WERE LOCATED IN THE AIRPATH AT THE BOTTOM OF THE BLOWER BOX AND THE OUTLET PORT, A MINIMAL AMOUNT OF POTENTIAL KERATIN AROUND THE BLOWER BOX INNER RING OF THE OUTLET PORT. THE MANUFACTURER COULD NOT CONFIRM SOUND ABATEMENT FOAM DEGRADATION. . THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERROR LOGGED. THE DEVICE WAS APPLIED POWER AND THE DEVICE OPERATED PROPERLY. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH LIGHT BEIGE DUST OR DIRT IN THE AIR INTAKE CANAL, ON THE INSIDE OF THE BLOWER BOX, ON THE INSIDE SURFACE OF THE BOTTOM ENCLOSURE, TWO SMALL BROWN DECEASED INSECTS ON THE INSIDE OF THE BOTTOM ENCLOSURE SURFACE SURROUNDED BY POTENTIAL BROWN FIBERS, TWO VERY SMALL DARK SPEC CONTAMINATES (NOT SOUND ABATEMENT FOAM) WERE LOCATED IN THE AIRPATH AT THE BOTTOM OF THE BLOWER BOX AND THE OUTLET PORT, A MINIMAL AMOUNT OF POTENTIAL KERATIN AROUND THE BLOWER BOX INNER RING OF THE OUTLET PORT. THE MANUFACTURER DID NOT CONFIRMED THE PRESENCE OF SOUND ABATEMENT FOAM DEGRADATION DUE TO INSECT CONTAMINATION. SECTION D9,D10,G3,H6 HAS BEEN UPDATED/CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03639 | 2021-09-08 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. ADDITIONAL INFORMATION WAS RECEIVED, SECTION B5 WAS CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING OR CHAMBER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. AFTER THE FINAL ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, THE CUSTOMER LIFT THE CALL AND HUNG UP THE CALL. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03634 | 2021-09-08 | 2021-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03792 | 2021-09-08 | 2021-07-19 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 06, 2024, AND SECTION H11 SHOULD BE REPORTED AS THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING CHAMBER, NASAL THROAT IRRITATION AND SORENESS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT WAS RUSHED TO HOSPITAL. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM (PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR). SECTION B2 WAS CHANGED TO HOSPITALIZATION (INITIAL OR PROLONGED) (PREVIOUSLY IT WAS BLANK). SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03636 | 2021-09-08 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | PATIENT ALLEGATIONS WAS MISSED IN INTIAL REPORT AND B5 IS REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAD BEEN EXPERIENCING PARTICLES IN HIS THROAT THAT HE HAS BEEN COUGHING UP. IT CAUSES IRRITATIONS AND HAS A HACKING COUGH. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HAS BEEN CORRECTED/UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03637 | 2021-09-08 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE 0 ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03635 | 2021-09-08 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03638 | 2021-09-08 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER INITIALLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED VISUALIZATION OF BLACK PARTICLES, STUFFINESS. THERE WAS NO PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE INTERNAL INVESTIGATION SHOWED THAT THERE WERE SIGNS OF SOUND ABATEMENT FOAM DEGRADATION IN THE BLOWER BOX AND ON THE BLOWER MOTOR. THE EXTERNAL INVESTIGATION SHOWED THAT THERE WERE NO SIGNS OF CONTAMINATION ON THE OUTSIDE OF THE DEVICE. THE MANUFACTURER FOUND EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN IN THIS DEVICE. THE DEVICE WAS HOOKED UP TO A KNOWN GOOD POWER SUPPLY AND POWER CORD, AIRFLOW WAS VERIFIED TO BE OPERATING CORRECTLY. THE ERROR LOG SHOWED THAT THERE WERE ZERO INSTANCES OF ERRORS ON THE DEVICE. THE MANUFACTURER CONCLUDES THAT THERE WERE SIGNS OF SOUND ABATEMENT FOAM DEGRADATION IN THE BLOWER BOX AND ON THE BLOWER MOTOR. THE EXTERNAL INVESTIGATION SHOWED THAT THERE WERE NO SIGNS OF CONTAMINATION ON THE OUTSIDE OF THE DEVICE. THE MANUFACTURER FOUND EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN IN THIS DEVICE. SECTION D8, D9, G3, H6 HAS BEEN UPDATED/CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03790 | 2021-09-08 | 2020-06-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAM STATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102632) ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102632) RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
MW5103795 | 2021-09-08 | 2021-07-01 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPERONICS DREAMSTATION 1 CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS, INC. | DREAMSTATION |  |  | INCREASED ASTHMA SYMPTOMS. CONTINUING DRY COUGH, PAIN IN THROAT AND ESOPHAGUS, TROUBLE BREATHING AT TIMES. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5103794 | 2021-09-08 | 2021-09-04 | I, n, j, u, r, y | Voluntary report | DREAMSTATION S/T | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | PHILIPS / RESPIRONICS, INC. | DREAMST ST30 H/C, DS | REF: DSX1030H11C |  | I AM GETTING SURGERY ON (B)(6) 2021 TO GET MY TONSILS REMOVED DUE TO ISSUES WITH MY ALLERGIES AND ASTHMA. I AM ABLE TO RECOGNIZE A CLEAR, DISTINCT INCREASE IN ISSUES WITH THESE PROBLEMS AND THE START OF USING MY BIPAP ST (PHILIPS DREAMSTATION ST). I DIDN'T NOTICE IT FOR THE FIRST FEW MONTHS AND THOUGHT IT WAS MOSTLY DUE TO STRESS BROUGHT ON BY THE PANDEMIC, BUT NOW I FOUND OUT IT IS MOST LIKELY CAUSING THESE ISSUES AND MIGHT BE GIVING ME CANCER? ARE YOU KIDDING? I DEMAND EVERY SINGLE RED CENT I EVER PAID TO (B)(6) BACK. I DEMAND COMPENSATION FOR MY ALLERGY APPOINTMENTS AND CONSULTS, MY SUPPLIES FOR MY MACHINE, THE COST OF MY UPCOMING SURGERY, ALL OF IT. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5103793 | 2021-09-08 | 2021-08-25 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAMSTATION AUTO CPAP HUM DON | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500H11 |  |  | AFTER RECEIVING THE RECALL NOTICE, I REGISTERED MY UNIT AND STOPPED USING THE PHILIPS CPAP FOR TWO NIGHTS. DURING THOSE 2 NIGHTS I VOMITED. AFTER CONSULTING WITH (B)(6) HOSPITAL SLEEP CENTER - MY SLEEP STUDY SOURCE AND CPAP AUTHORIZER - I HAVE BEEN USING THE CPAP. CLEARLY, I DON'T KNOW IF THE CPAP STOPPAGE WAS "CAUSAL" BUT, AS A MINIMUM IT WAS "COINCIDENTAL" WITH THE VOMITING STARTING AND THEN STOPPING. MY CPAP IS OVER 5 YEARS OLD AND MY DR RECOMMENDS IMMEDIATE REPLACEMENT. HOW DO I MAKE THAT HAPPEN AND HOW CAN I INCREASE THE PRIORITY OF ITS REPLACEMENT? I HAVE NOT HAD A SLEEP STUDY RE-ACCOMPLISHED, AS I HAVE PERIODIC SLEEP CENTER UPDATES THAT SUBSTANTIATE THE ONGOING NEED FOR THE CPAP. A 186-AROUSAL EVENT. AROUSAL INDEX 31.6 PER HOUR OF SLEEP. A 4 OBSTRUCTIVE SLEEP APNEAS - MIXED APNEAS AND 5 CENTRAL APNEAS. TOTAL NUMBER OF APNEAS WAS 9 WITH AN APNEA INDEOF 1.5 PER HOUR. THERE WERE 27 HYPOPNEAS RESULTING IN A HYPOPNEA INDEX OF 4.6 AND A TOTAL AHI OF 6.1. A TOTAL OF 147 RERAS OCCURRED WITH A RERA INDEX OF 25.0 THE AHI WAS 15.2 SUPINE AND 11.2 NON-SUPINE AND 1.1 IN REM AND 7.0 NREM SLEEP. THE RERA INDEX WAS 52.2 SUPINE AND 11.2 NON-SUPINE AND 6.6 IN REM AND 28.1 IN NREM SLEEP. A TOTAL OF 23 OXYGEN ESATURATRATIONS OCCURRED DURING SLEEP WITH AN OXYGEN DESATURATION INDEX (ODI) OF 3.9 PER HOUR OF SLEEP. THE MAXIMUM DURATION OF OXYGEN DESATURATION WAS 142.0 SEC. THE MEAN OXYGEN DESATURATION DURING THE RECODING WAS 94.4% AND THE LOWEST OXYGEN SATURATION RECORDED DURING SLEEP WAS 89.0%. ASSESSMENT: OBSTRUCTIVE SLEEP APNEA-SEVERE (ICD 10-G47.33)] '. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-09
MW5103792 | 2021-09-08 | 2021-06-15 | I, n, j, u, r, y | Voluntary report | RESPIRONICS DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  | DSX500H11C |  | I STARTED FEELING A NOSE IRRITATION SINCE LATE 2020 AND REPORTED IT TO MY DOCTOR. WE TRIED DIFFERENT MEDICATIONS BUT IT NEVER COMPLETELY SUCCEEDED. ON (B)(6) 2021 I NOTICED A LITTLE GROWTH ON BOTH SIDES OF MY NECK AND THEY HAVE NOW GROWN TO BIG TUMORS ON BOTH SIDES OF MY NECK THAT HAVE BEEN DIAGNOSED AS UNDETERMINED CARCINOMA WITHOUT A PRIMARY CANCER SOURCE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5103790 | 2021-09-08 | 2021-07-01 | I, n, j, u, r, y | Voluntary report | REMSTAR PRO, -RESPIRONICS INC.- SYSTEM ONE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS INC. | 460P--DS6HFLG-- |  |  | CHRONIC SINUS PROBLEMS, UNABLE TO WEAR ALL NIGHT, BREATHING PROBLEMS REQUIRED C-PAP TO BREATHE, BREATHING PROBLEMS INCREASING FOR PAST 5 YEARS, NOW REQUIRING MEDICAL HELP, INSTRUCTED TO STOP USING. ON STEROIDS FOR BREATHING, SCARRING OF LUNGS, MEDICAL DR'S ARE JUST BECOMING AWARE OF PROBLEM AND DON'T KNOW WHAT TO DO. LUNG DR'S HAS ONLY GAVE STEROIDS, NO KNOWLEDGE OF PROBLEM AT TEST DATE. MEDICAL DR. IS AWARE OF RECALL, ONLY HAS INFO THAT I HAVE, BUT SAID CPAP¿S VERY HARD TO FIND AT THIS TIME. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | HEALTH PROFESSIONAL |  | 2021-09
MW5103789 | 2021-09-08 | 2019-01-01 | I, n, j, u, r, y | Voluntary report | PHILIPS SYSTEM ONE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I AM USING THE PHILIPS CPAP WHICH HAS JUST BEEN RECALLED. I HAD HUGE EYE REACTIONS AND HAVE SEEN SEVERAL DOCTORS REGARDING THE EXTREME PUFFINESS AND FINALLY DECIDED TO HAVE SURGERY. TWO MONTHS AFTER MY SURGERY I AM NOW WONDERING IF IT HAD TO DO WITH THIS MACHINE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03723 | 2021-09-08 | 2021-08-10 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SCAR TISSUE IN THEIR LUNGS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DUE TO SLEEP APNEA PATIENT HAD A BLOOD CLOT IN THE HEART AND GOT A SMALL VESSEL DISEASE. PATIENT WAS DIAGNOSED WITH FIBROUS SCAR TISSUE IN THE APEX OF THE LUNG. PATIENT ALSO REPORTED INFLAMMATORY CONDITIONS OF THE THROAT, SHORTNESS OF BREATH, AND COUGHING FITS THAT IN SOME CASES CAUSED VOMITING. THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 (HEALTH EFFECT - CLINICAL CODE) HAS BEEN CORRECTED AND (TYPE OF INVESTIGATION,INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS) HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03724 | 2021-09-08 | 2021-01-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03722 | 2021-09-08 | 2021-08-10 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03788 | 2021-09-08 | 2021-05-24 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102660) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO HAVE SINUS ISSUES. THE PATIENT REQUIRED SURGERY AND PRESCRIPTION MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED A VOLUNTARY MEDWATCH (MW5102660) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO HAVE SINUS ISSUES. THE PATIENT REQUIRED SURGERY AND PRESCRIPTION MEDICATION IN RESPONSE TO THE REPORTED EVENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03789 | 2021-09-08 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THERE WAS NO PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DIZZINESS, NAUSEA, HEADACHES AND COUGHING UP BLACK PHLEGM WITHIN THE LAST FOUR - FIVE MONTHS. PATIENT REPORTED DEVICE HAS BEEN SERVICED TWICE CHANGED BACK FROM HEATING PAD ON THE BOTTOM. THERE WAS NO PATIENT HARM OR INJURY. THE REPORTED EVENT AND ITS SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS A NON-SERIOUS INJURY. THEREFORE, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HEALTH EFFECT - CLINICAL CODE HAS BEEN CORRECTED AND TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-09
MW5103798 | 2021-09-08 | 2021-01-28 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSXHCP |  |  | RESPIRATORY ISSUES - SAW DOCTOR - ONGOING COUGH AND BREATHING AFFECTED. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03754 | 2021-09-08 | 2021-07-15 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | UDSX500S11 | UDSX500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. SECTION B5 SHOULD BE PREVIOUSLY REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT WAS ALLEGING TO DEVELOP CANCER, HEADACHE, NASAL/THROAT IRRITATION/SORENESS, DIFFICULTY BREATHING/SHORTNESS OF BREATH, HEART FAILURE, OXYGEN ON 24-7, HAS BEEN NAUSEATED AND VOMITING IN THE PAST. DEVICE WAS NOISY AND HAS A STRANGE ODOR. NO MEDICAL INTERVENTION WAS REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION G1 UPDATED. SECTION H6 HEALTH EFFECT - CLINICAL CODES WHICH WAS MISSED IN PREVIOUS REPORT WAS CAPTURED. SECTION H6 TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03682 | 2021-09-08 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2200X11B | DS2200X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVIC'ES MACHIN FLOW SENSOR NEEDS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03753 | 2021-09-08 | 2021-07-15 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP HEARING LOSS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO EXPERIENCE HEARING LOSS. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENT OF HEARING LOSS AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. IN INITIAL REPORT, SECTION G3 WAS CAPTURED AS 08/10/2021, SHOULD HAVE BEEN 07/15/2021, UPDATED G3 SECTION IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03805 | 2021-09-08 | 2021-08-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S TOUCHSCREEN WAS UNABLE TO BE ACCESSED. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR'S TOUCHSCREEN WAS UNABLE TO BE ACCESSED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS NOT CONFIRMED. NO MALFUNCTION OF THE TOUCHSCREEN WAS NOTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03734 | 2021-09-08 | 2021-08-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03806 | 2021-09-08 | 2021-08-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD AND MACHINE FLOW SENSOR WERE REPLACED TO ADDRESS THE ISSUE. THE DEVICE HAD EVIDENCE OF WATER INGRESS. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03804 | 2021-09-08 | 2021-08-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | IN2110X15B | IN2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S TOUCHSCREEN WAS UNABLE TO BE ACCESSED. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR TRILOGY EVO TOUCHSCREEN WAS UNABLE TO BE ACCESSED. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE WAS RETURNED TO THE MANUFACTURING SERVICE CENTER FOR EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THERE WAS NO ADDITIONAL FAILURE WERE OBSERVED. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION. THE DEVICE WAS SCRAPPED. SECTION H6 TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03807 | 2021-09-08 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S MACHINE FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03803 | 2021-09-08 | 2021-08-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT'S BLOOD OXYGEN SATURATION DECREASED, THE PATIENT COMPLAINED OF DIFFICULTY BREATHING AND A RESUSCITATOR BAG WAS USED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S OVERHEAD BLOWER FILTER WAS FOUND TO BE CLOGGED WHICH AFFECTED THE ACTIVE EXHALATION CONTROL MODULE FUNCTION. THE OVERHEAD BLOWER FILTER AND ACTIVE EXHALATION CONTROL MODULE NEED TO BE REPLACED TO ADDRESS THE ISSUES. | ADDITIONAL INFORMATION WAS RECEIVED ON 04 NOV, 2024 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT'S BLOOD OXYGEN SATURATION DECREASED, THE PATIENT COMPLAINED OF DIFFICULTY BREATHING AND A RESUSCITATOR BAG WAS USED. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. | Y | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03797 | 2021-09-08 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A MALFUNCTION OF A VENTILATOR'S TOUCHSCREEN. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING MALFUNCTION OF A VENTILATOR'S TOUCHSCREEN. THERE WAS NO HARM OR INJURY REPORTED AFTER THE FIRST ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, CUSTOMER DECLINED TO RESPOND TO THE GFE RELATED QUESTIONS AND TERMINATED THE CALL. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03795 | 2021-09-08 | 2021-09-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OVERHEAD BLOWER FILTER NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03796 | 2021-09-08 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DE2110X13B | DE2110X13B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. TUBING FROM THE SYSTEM BOARD WAS OBSERVED TO BE DISCONNECTED. THE TUBING WAS RECONNECTED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03733 | 2021-09-08 | 2021-07-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2200X11B | DS2200X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS EVALUATED BY A FIELD SERVICE ENGINEER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD NEEDS REPLACED AND THE SOFTWARE RELOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03732 | 2021-09-08 | 2021-07-16 | I, n, j, u, r, y | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR BECAME DISCONNECTED FROM AC POWER, THE DEVICE STOPPED WORKING AND THE PATIENT EXPERIENCED SHORTNESS OF BREATH AND DIFFICULTY BREATHING. THE PATIENT WENT TO THE HOSPITAL FOR MEDICAL INTERVENTION. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR BECAME DISCONNECTED FROM AC POWER, THE DEVICE STOPPED WORKING AND THE PATIENT EXPERIENCED SHORTNESS OF BREATH AND DIFFICULTY BREATHING. THE PATIENT WENT TO THE HOSPITAL FOR MEDICAL INTERVENTION. IN THE PREVIOUS REPORT B2 WAS OMITTED. THEY ARE REFLECTED ON THIS REPORT. H1 WAS INCORRECTLY REPORTED AS A MALFUNCTION AND SHOULD BE REPORTED AS A SERIOUS INJURY. IT IS CORRECTED ON THIS REPORT. | Y | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03725 | 2021-09-08 | 2021-08-10 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03701 | 2021-09-08 | 2021-08-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BLOWER MOTOR NEEDS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, "SERVICE REQUIRED' CODES RELATED TO THE DEVICE'S SPEAKERS WERE OBSERVED. THE EVALUATION IS STILL ONGOING FOR THIS MALFUNCTION AND A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE THIRD PARTY SERVICE CENTER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S BLOWER MOTOR NEEDED REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, "SERVICE REQUIRED' CODES RELATED TO THE DEVICE'S SPEAKERS WERE OBSERVED. THE SPEAKERS NEED REPLACED TO ADDRESS THE ISSUE. DURING ADDITIONAL TESTING, THE DEVICE FAILED A TEST STEP DURING TESTING. THE ACTIVE EXHALATION CONTROL MODULE NEEDS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
MW5103786 | 2021-09-08 | 2021-08-01 | I, n, j, u, r, y | Voluntary report | CONTINUOUS VENTILATOR | VENTILATOR, CONTINUOUS, FACILITY USE | PHILIPS ELECTRONICS LTD. | TRILOGY |  |  | SINUS, HEADACHES, RUNNY NOSE, AIRWAY IRRITATION COUGH. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5103788 | 2021-09-08 | 2021-09-03 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS RESPIRONICS REMSTAR AUTO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | SYSTEMONE |  |  | ERRONEOUS INFORMATION AT FDA. I WAS TOLD TO USE THIS FORM BUT THE SITUATION IS THAT THERE IS RECALL INFORMATION TO FILE FOR A DEVICE REPLACEMENT OR FIX OF PHILIPS RESPIRONICS CPAP/APAP MACHINES AND OTHER RESPIRATORY DEVICES AND THE FDA WEBSITE REFLECTS THE WRONG URL TO FILE A CLAIM. THIS SAME WRONG INFORMATION (URL FOR CLAIM SUBMISSION) IS ON THE PAPERWORK PROVIDED TO THE CLAIMANTS AS WELL AS THE FDA WEBSITE. IT NEEDS TO BE CHANGED FROM WWW.PHILLIPS.COM/SRC-UPDATE TO HTTPS://WWW.USA.PHILIPS.COM/HEALTHCARE/E/SLEEP/COMMUNICATIONS/SRC-UPDATE. MY SON FOUND THE CORRECT WEBSITE. YOU CAN FIND THE FDA PAGE WITH THE WRONG URL HERE: HTTPS://WWW.FDA.GOV/SAFETY/RECALLS-MARKET-WITHDRAWALS-SAFETYALERTS/ PHILIPS-ISSUES-RECALL-NOTIFICATION-MITIGATE-POTENTIAL-HEALTH-RISKS-RELATED-SOUND-ABATEMENT-FOAM. I'M BEING FORCED TO ANSWER A SERIES OF QUESTIONS WITH FALSE RESPONSES JUST, SO I CAN PROVIDE YOUR IT DEPARTMENT THE CORRECT INFORMATION TO UPDATE THE LANDING PAGE FOR THE PHILIPS RECALL OF THESE DEVICES. ADDITIONAL COMMENTS NOT APPLICABLE. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03661 | 2021-09-08 | 2021-08-09 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC. |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A (CPAP, BIPAP) DEVICE'S SOUND ABATEMENT FOAM. A USER ALLEGES UNSPECIFIED SERIOUS LUNG INJURY, RECURRENT RESPIRATORY INFECTIONS, ACUTE RESPIRATORY FAILURE REQUIRING INTENSIVE CARE AND OTHER RESPIRATORY ISSUES. NO SERIAL NUMBER OR MATERIAL NUMBER WERE PROVIDED. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP,AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING UNSPECIFIED SERIOUS LUNG INJURY, RECURRENT RESPIRATORY INFECTIONS, ACUTE RESPIRATORY FAILURE REQUIRING INTENSIVE CARE AND OTHER RESPIRATORY ISSUES RELATED TO A DEVICE'S SOUND ABATEMENT FOAM.NO SERIAL NUMBER OR MATERIAL NUMBER WERE PROVIDED. THERE IS NO CUSTOMER INFORMATION AND WE CANNOT REACH OUT TO THE CUSTOMER, HENCE, ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION CAN NOT BE MADE AT THIS TIME. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03798 | 2021-09-08 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED PARTICLES IN THE TUBING, SINUS INFECTIONS, AND LUNGS HAVE BEEN HURTING FOR OVER A YEAR. ALSO OBSERVED THAT BLACK STUFF CAME OUT WHEN HE DID THE TUBE CLEANING. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. DURING THE INVESTIGATION, CANNOT CONFIRM ANY EVIDENCE OF FOAM DEGENERATION OR FOAM CONTAMINATION. FAULTS OR ERRORS CONTAINED WITHIN THE LOGS SUPPORT SENSOR BOARD CONTAMINATION, WITH ALL OTHER DEVICE FUNCTIONALITY APPEARING AS HAVING NO PROBLEM. SENSOR BOARD FAILURE BECAUSE OF SENSOR CONTAMINATION AND POSSIBLY CONTAMINATION IN ALL SENSORS. SECTION H6 HEALTH EFFECTS-CLINICAL CODE WAS INCORRECTLY CAPTURED IN THE PREVIOUS REPORT. IT HAS BEEN CORRECTED IN THIS REPORT. IN THIS REPORT, SECTION D9, G3, H3, H6 HAS BEEN UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03766 | 2021-09-08 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES IN TUBING/CHAMBER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03767 | 2021-09-08 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING / CHAMBER, NASAL / THROAT IRRITATION OR SORENESS, NAUSEA, STRANGE ODOR RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. DURING LAB INVESTIGATION OBSERVED CONFIRM THAT THE DEVICE POWERS UP AND PROVIDES AIRFLOW USING A KNOWN GOOD POWER SUPPLY. PIL OBSERVED DIRT/DUST CONTAMINATION ON TOP AND BOTTOM ENCLOSURES, UI PANEL, REAR PANEL, SD COVER FLIP DOOR, ACCESSORY MODULE FLIP DOOR, BLOWER BOX UPPER CAP, BLOWER, BLOWER OUTLET SEAL, REAR PANEL O-RING, P4 POWER CONNECTOR. A CONTAMINATE CONSISTENT WITH KERATIN WAS OBSERVED ON THE BLOWER BOX. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF CONTAMINATION IN THE DEVICE AND CONFIRMED THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION D8, D9 AND H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03765 | 2021-09-08 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING BLACK STUFF WAS COMING OUT OF THE WATER TANK RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS UPDATED REPORT WILL BE FILED. CORRECTED INFORMATION PROVIDED IN B5, THE INFORMATION (THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK STUFF IS COMING OUT OF THE WATER TANK RELATED TO A CPAP DEVICE) WAS NOT INCLUDED IN INITIAL REPORT. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03770 | 2021-09-08 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03762 | 2021-09-08 | 2021-08-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN THE PREVIOUSLY SUBMITTED REPORT SECTION B5 WAS INCOMPLETE, SECTION B5 IS- THE PATIENT ALSO ALLEGED VISUALIZATION OF BLACK PARTICLES EXPELLED FROM MOUTH. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF MODERATE AMOUNT OF BEIGE AND DARK DUST OR DIRT CONTAMINATE ALONG WITH DARK FIBER CONTAMINATE THROUGHOUT THE BASE UNIT INCLUDING THE AIR INLET AREA, INLET CANAL, IN AND AROUND THE OUTLET PORT, UI PANEL CHANNELS, THE BLOWER BOX, THE BLOWER AND ISOLATORS, BLOWER IMPELLERS, AND BLOWER SEAL. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. THE HUMIDIFIER WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR EVALUATION. THE MANUFACTURER FOUND EVIDENCE OF MODERATE DARK FIBER CONTAMINATION IN AND AROUND THE OUTLET PORT, WHITE FIBER OR HAIR CONTAMINATE LOCATED ON THE ENTRANCE TO THE HUMIDIFIER INLET PORT, POTENTIAL MINERAL DEPOSITS ON THE INSIDE WATER TANK METAL PLATE, UNDETERMINED CONTAMINATE ON HUMIDIFIER HEATER PLATE, A MINOR BROWN CONTAMINATE WITH DARK FIBER CONTAMINATE ON THE INTERNAL UPPER HUMIDIFIER LID SEAL AND SURFACES. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND 3 INSTANCES OF ERROR 53, 10 INSTANCES OF ERROR 190, 2 INSTANCES OF ERROR 200. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF MULTIPLE CONTAMINANTS IN THE DEVICE. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. SECTION B3, D9, D10, G3 AND H6 HAS BEEN UPDATED/CORRECTED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03763 | 2021-09-08 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON 08/24/2023, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF BLACK PARTICLES IN THE WATER CHAMBER AND TUBING RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. ADDITIONAL INFORMATION WAS RECEIVED ABOUT ALLEGED DRY COUGH AND NASAL/THROAT IRRITATION OR SORENESS. SECTION H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03764 | 2021-09-08 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING NASAL AND THROAT IRRITATION RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL AND INTERNAL PART OF PART OF THE DEVICE SECONDARY FINDINGS ARE CONTAMINATION ON BOTTOM ENCLOUSER AND DUST CONTAMINATES ON THE BLOWER MOTOR. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER CONCLUDES THAT THERE WAS NO UNKNOWN PIECE OF CONTAMINATION ON THE INSIDE OF THE BOTTOM ENCLOUSER AND UNKOWN CONTAMINATION ON BOTTOM ENCLOUSER AND DUST CONTAMINATES ON THE BLOWER MOTOR. OBSERVED NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION AND BREAKDOWN. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03771 | 2021-09-08 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK PARTICLES WERE OBSERVED IN THE HUMIDIFIER CHAMBER. THE PATIENT ALSO ALLEGED TO HAVE A BURNING SENSATION IN THE MOUTH, TONGUE AND THROAT. THERE WAS NO REPORT OF SERIOUS HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR EVALUATION. THE DEVICE WAS VISUALLY INSPECTED. THE MANUFACTURER CONFIRMED WHITE AND BLACK CONTAMINATION CONSISTENT WITH DUST AND HAIR. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE MANUFACTURER CONCLUDES THE DEVICE HAD EVIDENCE OF CONTAMINATION FROM AN EXTERNAL SOURCE, CONSISTENT WITH DUST. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03772 | 2021-09-08 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03729 | 2021-09-09 | 2021-09-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL, AU | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AUX500T15C | AUX500T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE COUGH. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 04, 2024, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PATIENT TO HAVE COUGH AND HEADACHES RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTIONS B1, B2, H1 AND H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR ONLY THE PRODUCT PROBLEM (BOTH ADVERSE EVENTS AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03758 | 2021-09-09 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SHORTNESS OF BREATH. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO DEVELOP SHORTNESS OF BREATH. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENT OF SHORTNESS OF BREATH AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03800 | 2021-09-09 | 2021-07-15 | I, n, j, u, r, y | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A BLOOD CLOT IN THE LUNG. THE PATIENT IS CURRENTLY IN A REHABILITATION CENTER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED BLOOD CLOT IN THE LUNG. THE PATIENT IS CURRENTLY IN A REHABALITION CENTER. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURE INSPECTED EXTERNALLY OBSERVED THAT DUST, DIRT, HAIR AND OTHER PARTICLES CONTAMINATION AND FAILURE OF INTERNAL BATTERY. THERE WAS NO EVIDENCE OF BLACK FOAM DEGARDATION. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION AND THERE IS NO VISIBLE DAMAGE OR FUNCTIONALITY FAILURES OF THE DEVICE. THE MANUFACTURER CAN CONFIRM THE SENSOR FAILURES WHILE TESTING DUE TO CONTAMINATION SUCH AS DUST/DIRT, HAIR, PARTICLES THROUGHT OUT THE AIR/FLOW PATH , INSIDE THE TRANSISITION TUBING AND THAT ARE CONSISTENT FROM EXTERNAL SOURCE. SECTION H6 TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03715 | 2021-09-09 | 2021-08-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03849 | 2021-09-09 | 2021-09-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERFACE BOARD AND OXYGEN BLENDING MODULE BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03743 | 2021-09-09 | 2021-09-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03756 | 2021-09-09 | 2021-07-15 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION BIPAP PRO W/HUM/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX600H11C | DSX600H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A SPOT ON THEIR LUNG. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | IN PREVIOUS REPORT, THE MANUFACTURER DID NOT CAPTURE FEW ALLEGATIONS IN B5 VERBIAGE. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED THAT THE DEVICE AS SPOT ON THEIR LUNG, LUNG CANCER, COUGH, CANT BREATH WELL DUE TO CONGESTION, SHORTNESS OF BREATHE, RID OF FLEM IN HIS THROAT. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE EVENT IS CURRENTLY UNKNOWN. SECTION H6 CLINICAL CODE WAS UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03852 | 2021-09-09 | 2021-09-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
MW5103834 | 2021-09-09 |  | I, n, j, u, r, y | Voluntary report | CPAP DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC |  |  |  | CALLER REPORTING ON THE DREAMSTATION DEVICE SHE HAS BEEN USING FOR A LITTLE OVER THREE YEARS. SHE ALLEGES SOMETIME THIS YEAR SHE STARTED WAKING UP WITH SEVERE HEADACHES AND TIGHTNESS AROUND HER CHEST, SHE FELT PRESSURE TOWARDS HER BACK AS WELL AND IT WAS DIFFICULT TO BREATH. THESE SYMPTOMS ACCORDING TO CALLER KEPT GETTING WORSE AND WORSE SHE EVENTUALLY HAD TO STOP USING THE DEVICE. SHE HAD TO RELY ON INHALERS TO MANAGE HER ASTHMA AND BREATHING. SHE WENT BACK TO USING THE CPAP DEVICE ABOUT A MONTH AGO AND SHE EXPERIENCED THE SAME ADVERSE EVENTS. SHE HEARD ON THE NEWS WHAT OTHER PEOPLE EXPERIENCED AND SHE IMMEDIATELY STOPPED USING THE DEVICE. ON SEPTEMBER 8TH SHE RECEIVED A RECALL NOTICE FROM PHILIPS RESPIRONICS. |  | Y | Y | I |  | 2021-09
MW5103828 | 2021-09-09 | 2021-09-07 | I, n, j, u, r, y | Voluntary report | REMSTAR SE AUTO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS, INC. |  |  |  | TOXIC/CARCINOGENIC EFFECTS; RECEIVED A RECALL NOTICE FROM PHILIPS REGARDING THE CPAP MACHINE I'VE BEEN USING FOR YEARS. PE-PUR FOAM ISSUES. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | I |  | 2021-09
MW5103827 | 2021-09-09 | 2020-03-29 | I, n, j, u, r, y | Voluntary report | CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I'M EXPERIENCING ON AND OFF HEADACHE FOR A YEAR NOW. RECENTLY, THE PAIN BECOMES PERSISTENT AND WORSENING (SOMETIMES 3-5 DAYS ,NON STOP), TYLENOL NOT EFFECTIVE IN RELIEVING PAIN. I WENT TO THE DOCTOR DUE TO PERSISTENT HEADACHE, 9 OUT OF 10 PAIN LEVEL. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5103820 | 2021-09-09 | 2019-10-21 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAM MACHINE BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | RECALL OF PHILIPS RESPIRONICS DREAM MACHINE BIPAP IS ANNOUNCED IN FEB 2020. DR PRESCRIBED MACHINE IN (B)(6) 2019. LATER RECOMMENDED SOCLEAN OZONE CLEANER IN (B)(6) 2021. DR TAKES NO RESPONSIBILITY FOR THESE ERRORS. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03884 | 2021-09-09 | 2021-08-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | REMSTAR PRO C-FLEX | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS INC. | LA451SBR | LA451SBR |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03883 | 2021-09-09 | 2021-07-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO SERIES | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS INC. | 750P | 750P |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE REPORTER ALSO ALLEGED BLACK PARTICLES COMING OUT OF THE DEVICE WHILE RUNNING. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL PART OF THE DEVICE AND FOUND NO EVIDENCE OF THERMAL DAMAGE OR FAILURE. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND FOUND NO EVIDENCE OF THERMAL DAMAGE. THE MANUFACTURER DID OBSERVE EVIDENCE OF DIRT/DUST CONTAMINATION AND FOAM MATERIAL PRESENT IN THE REMSTAR. THE MANUFACTURER FOUND THERE IS DIRT/DUST ON THE INSIDE SURFACES (TOP SIDE OF THE PCA AND OUTER TOP PORTION OF THE BLOWER CAP). ALSO FOUND THE SOUND ABATEMENT FOAM MATERIAL IS MISSING IN THE BOTTOM INSIDE OF THE ENCLOSURE AND WAS IN THE BLOWER MOTOR/BOX. THE AIR PATH OF THE DEVICE WAS COMPROMISED AND THE CONTAMINATE WOULD HAVE BEEN THROUGHOUT THE DEVICES AIR PATH. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND TO CONTAIN 0 ERROR CODES. THE MANUFACTURER VERIFIED THE UNITS DO POWER-UP, PROVIDE FLOW. ALSO OBSERVE BLOWER MOTOR NOISE. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. ALSO FOUND CONTAMINANTS THROUGHOUT THE DEVICES AIR PATH WHICH INDICATES THE SOURCE OF THE CONTAMINANTS IS EXTERNAL TO THE DEVICE. ON THE PREVIOUSLY SUBMITTED REPORT, SECTION D9 HAD INCORRECT DATE, WHICH WAS UPDATED/CORRECTED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03620 | 2021-09-09 | 2021-08-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TESTING. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TESTING. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. DURING FINAL TESTING, THE DEVICE FAILED A TEST STEP. THE DEVICE'S MACHINE FLOW SENSOR WAS REPLACED DUE TO CONTAMINATION. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03621 | 2021-09-09 | 2021-08-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2200X11B | DS2200X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE TOUCH SCREEN WAS OBSERVED. THE DEVICE'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03736 | 2021-09-09 | 2021-08-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03737 | 2021-09-09 | 2021-08-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03735 | 2021-09-09 | 2021-08-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03740 | 2021-09-09 | 2021-08-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S POWER MANAGEMENT BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03738 | 2021-09-09 | 2021-08-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE MANIFOLD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03823 | 2021-09-09 | 2021-09-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S SENSOR BOARD AND POWER MANAGEMENT BOARD WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03744 | 2021-09-09 | 2021-09-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03741 | 2021-09-09 | 2021-08-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1032804 | 1032804 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03742 | 2021-09-09 | 2021-08-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03739 | 2021-09-09 | 2021-08-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S NAVIGATIONAL BUTTON WAS OBSERVED. THE DEVICE'S FRONT PANEL/KEYPAD LED BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03808 | 2021-09-09 | 2021-08-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BLOWER MOTOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03850 | 2021-09-09 | 2021-09-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERFACE BOARD AND OXYGEN BLENDING MODULE BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03745 | 2021-09-09 | 2021-09-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03801 | 2021-09-09 | 2021-08-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE FILTER NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03851 | 2021-09-09 | 2021-09-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERFACE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03848 | 2021-09-09 | 2021-08-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03856 | 2021-09-09 | 2021-08-11 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03787 | 2021-09-09 | 2021-07-14 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CONGESTIVE HEART FAILURE. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03822 | 2021-09-09 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID/HT, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | CAX400T12 | CAX400T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED THROAT HAS BEEN BURNING FOR THE PAST 3 MONTHS, ALSO NOTICED BLACK DEBRIS WHEN COUGHING. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. A FOREIGN WHITE POWDERY SUBSTANCE WAS FOUND ON THE BLOWER MOTOR'S FAN BLADES. POTENTIAL DARK KERATIN PARTICLES WERE FOUND ON THE BLOWER BOX OUTLET PORT. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WAS 1 ERROR FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. SECTION H6 WAS UPDATED AND CORRECTED BY ADDING HEALTH EFFECT - CLINICAL CODES WHICH WAS MISSED IN INITIAL REPORT. SECTION D4 WAS CORRECTED BY CAPTURING UNIQUE IDENTIFIER WHICH WAS INCORRECT IN INITIAL REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03821 | 2021-09-09 | 2021-08-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX500H12C | CAX500H12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING LOTS OF ISSUES WITH BLACK PARTICLES IN THE HOSE, ALSO ISSUES WITH MAKING AIR VERY HOT AND NOT ABLE TO BREATH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION D4 WAS CORRECTED AND H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03760 | 2021-09-09 | 2021-07-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | CAX500T12C | CAX500T12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A PULMONARY EMBOLISM. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03759 | 2021-09-09 | 2021-09-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID/HT, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | AUX400T15 | AUX400T15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO HAVE A CODE BLUE. THE PATIENT WAS HOSPITALIZED AT THE TIME OF THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 12, 2024 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED THE EVENT OF HOSPITALIZATION AND USING DEVICE FOR FIVE HOURS - BLUE CODE. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY THE PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03882 | 2021-09-09 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP RESPIRATORY ISSUES. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT DEVELOP RESPIRATORY ISSUES AND MUCOUS. THE REPORTED EVENT OF RESPIRATORY ISSUES, MUCOUS AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THE MEDICAL INTERVENTION RECEIVED BY THE PATIENT IN THE FORM OF PRESCRIBED MEDICATION. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. HENCE, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03755 | 2021-09-09 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SPOTS ON THEIR LUNG. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 05, 2024, AND SECTION H11 SHOULD BE REPORTED AS THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED PATIENT TO DEVELOP SPOTS ON THEIR LUNG. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM (ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR), SECTION B2 WAS MADE BLANK (PREVIOUSLY IT WAS OTHER SERIOUS), SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION, SECTION H6 HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03731 | 2021-09-09 | 2021-09-02 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO, FR | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | FRX500S14 | FRX500S14 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP ASTHMA. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT WAS ALLEGING TO DEVELOP ASTHMA. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION D4 UNIQUE IDENTIFIER (UDI) # CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03854 | 2021-09-09 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03757 | 2021-09-09 | 2021-07-15 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SHORTNESS OF BREATH. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03858 | 2021-09-09 | 2021-08-11 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03857 | 2021-09-09 | 2021-08-11 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO BECOME UNCONSCIOUS. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO BECOME UNCONSCIOUS. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03726 | 2021-09-09 | 2021-09-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO AUTO W/BT, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AUG500S15 | AUG500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE SINUSITIS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION FOR NASAL SPRAY AND SINUS DECONGESTANT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON 12 NOV, 2024 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED SINUSITIS. THE MEDICAL INTERVENTION THAT THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION FOR NASAL SPRAY AND SINUS DECONGESTANT. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03810 | 2021-09-09 | 2021-07-19 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT ALSO ALLEGED ASTHMA, COUGHS ALL NIGHT AND CANNOT BREATHE. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL AND INTERNAL INVESTIGATION SHOWED NO EVIDENCE OF FOAM DEGRADATION ASSOCIATED TO THIS ALLEGATION. THE VISUAL INSPECTION WAS UNREMARKABLE. NO EVIDENCE OF FOAM PARTICULATE. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN IN THE RETURNED MATERIALS. THERE IS NO VISIBLE DAMAGE OR FUNCTIONALITY FAILURES OF THE DEVICE, WHICH SUGGEST THAT THE SOURCE OF CONTAMINATION WAS EXTERNAL TO THE DEVICE. SECTION G1 CORRECTED AND H10 WRONGLY CHECKED IN PREVIOUS MDR. SECTION H6 WERE UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03860 | 2021-09-09 | 2019-09-23 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP DECREASED LUNG FUNCTION. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03861 | 2021-09-09 | 2021-06-14 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A POSSIBLE STROKE. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03817 | 2021-09-09 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID/HT, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX400T12 | CAX400T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN CHAMBER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03727 | 2021-09-09 | 2021-09-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AUX500T15 | AUX500T15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE SINUSITIS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION FOR NASAL SPRAY AND SINUS DECONGESTANT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON AUG 23, 2023 AND SECTION H11 SHOULD BE REPORTED AS THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE SINUSITIS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF A PRESCRIPTION FOR NASAL SPRAY AND SINUS DECONGESTANT. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM (ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR). SECTION B2 WAS CHANGED TO BLANK (PREVIOUSLY IT WAS OTHER SERIOUS). SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION . SECTION H6 HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03815 | 2021-09-09 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL, AU | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AUX500T15C | AUX500T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION B5 WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT SHOULD BE REPORTED AS FOLLOWS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED GETTING BLACK PARTICLES OUT OF HER NOSE WHEN SNEEZING OR BLOWING AND BLACK PARTICLES ARE NOW BEING FOUND IN THE MACHINES TUBES, MASK AND TANK. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. SECTION H6 HEALTH EFFECT - CLINICAL CODE WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT. IT HAS BEEN CORRECTED OR UPDATED IN THIS REPORT. UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03824 | 2021-09-09 | 2021-07-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE COPD (CHRONIC OBSTRUCTIVE PULMONARY DISEASE). THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE REPORTED EVENT IS CURRENTLY UNKNOWN. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON OCTOBER 12, 2023, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING COPD, NOSE BLEEDS AND SCRATCHY VOICE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE REPORTED EVENT IS CURRENTLY UNKNOWN. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGED SHORTNESS OF BREATH, HEADACHE, DRY THROAT, RINGING IN EARS, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, NAUSEA, VOMITING, LUNG DISEASE. SECTION E1, H6 WAS UPDATED. SECTION H6 CLINICAL CODE HAS BEEN CORRECTED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03825 | 2021-09-09 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE BREATHING ISSUES. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE REPORTED EVENT IS CURRENTLY UNKNOWN. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO EXPERIENCE BREATHING PROBLEMS, COVID. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENTS OF BREATHING PROBLEMS, COVID AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION AND AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTIONS B1, B2 HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. CORRECTED INFORMATION WAS PROVIDED IN SECTION G1. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03826 | 2021-09-09 | 2021-07-13 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION BIPAP AVAPS30 H/C, DS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX1130H11C | DSX1130H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE ASTHMA AND PNEUMONIA. THE PATIENT WAS REPORTEDLY HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT REPORTED ASTHMA, HEADACHES, DIZZINESS, PNEUMONIA RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT ALSO ALLEGED SEEING BLACK PARTICLES IN CHAMBER AND TUBING.THE PATIENT WAS REPORTEDLY HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. DURING EXTERIOR INVESTIGATION: THERE IS UNKNOWN DUST/DIRT CONTAMINATION PRESENT ON ALL SURFACES OF THE BASE UNIT. DURING INTERIOR INVESTIGATION: THERE IS UNKNOWN DEBRIS IN THE TRACKS OF THE UI PANEL. THERE ARE MINERAL DEPOSITS PRESENT ON THE MOTOR CASING SUGGESTING UNKNOWN LIQUID INGRESS. THERE IS AN UNKNOWN WHITE DUST CONTAMINATION FOUND ON THE BOTTOM ENCLOSURE, BLOWER BOX HOUSING, BLOWER MOTOR, BLOWER IMPELLER, AND REAR PANEL O-RING. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED IN THE BASE UNIT. A KERATIN-LIKE SUBSTANCE WAS OBSERVED AROUND THE BLOWER BOX OUTLET. THE UNKNOWN DUST/DIRT CONTAMINATION AND FIBERS FOUND ON THE BLOWER CASING/IMPELLER AND BLOWER BOX WAS INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM CONTAMINATION. THE PRESENCE OF CONTAMINATION THROUGHOUT THE AIRPATH SUGGESTS A SOURCE EXTERNAL TO THE DEVICE. MINERAL SPOTS CONSISTENT WITH WATER INGRESS WERE FOUND ON MOTOR CASING. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES UNABLE TO DIRECTLY ADDRESS THE SYMPTOMS OUTLINED IN THE COMPLAINT. CAN CONFIRM THAT THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN OBSERVED IN THE BASE UNIT. CONFIRMED THE PRESENCE OF CONTAMINATION IN THE AIRPATH. IN THIS REPORT, SECTION D9, G3, H3, H6 HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03827 | 2021-09-09 | 2021-08-11 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP ALZHEIMER'S DISEASE. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THIS EVENT IS CURRENTLY UNKNOWN. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03786 | 2021-09-09 | 2021-07-13 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP BLOOD CLOTS IN THEIR BRAIN. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03820 | 2021-09-09 | 2021-07-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | CAX500T12 | CAX500T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03896 | 2021-09-10 | 2021-08-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR A CIRCUIT DISCONNECT ALARM CONDITION. THE PATIENT'S BLOOD OXYGEN SATURATION DECREASED AND THE PATIENT NEEDED TO BE MANUALLY VENTILATED WITH AN AMBU BAG. THE PATIENT RECOVERED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR WAS ALARMING FOR A CIRCUIT DISCONNECTION ALARM CONDITION. THE PATIENT'S BLOOD OXYGEN SATURATION DECREASED AND THE PATIENT NEEDED TO BE MANUALLY VENTILATED WITH AN AMBU BAG. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS NOT DUPLICATED. THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. | Y | N | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03897 | 2021-09-10 | 2021-08-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VARIATION IN THE VENTILATOR'S PRESSURE DELIVERY WAS OBSERVED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR HAD A VARIATION IN THE VENTILATOR'S PRESSURE DELIVERY. THE DEVICE WAS EVALUATED BY A THIRD PARTY SERVICE CENTER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03802 | 2021-09-10 | 2021-06-25 | I, n, j, u, r, y | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260B | 1054260B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED SHORTNESS OF BREATH, HEADACHES AND EAR INFECTIONS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF ANTIBIOTICS AND HOSPITALIZATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED SHORTNESS OF BREATH, CONGESTION, HEADACHES, FLEM STUCK IN NOSE AND THROAT, EAR INFECTIONS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF ANTIBIOTICS AND HOSPITALIZATION. AFTER THE FIRST ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, THE CUSTOMER STATED THAT THE DEVICE IS ALREADY RETURNED. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FOLLOW UP REPORT WILL BE FILED. SECTION D4, E1 AND H4 CORRECTED AND UPDATED IN THIS REPORT. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03895 | 2021-09-10 | 2021-08-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2110X15B | IN2110X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S MACHINE FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
MW5103840 | 2021-09-10 | 2021-09-08 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | DSX500H11C |  |  | RECALL FAKE SITE; PHILIPS RESPIRONICS HAS EXISTING RECALL EFFORT FOR CPAP MACHINES AND SENT LETTER TO USERS WITH "PLEASE VISIT WEBSITE WWW.PHILIPS.COM/SRC-UPDATE" AND PHONE NUMBER MESSAGE ALSO DIRECTS TO THIS FAKE SITE. HOWEVER, THERE IS NO SUCH SITE FUNCTIONING. THIS IS A FAKE URL. ON MULTIPLE DAYS, THE SITE PROVIDES A "PAGE DOES NOT EXIST" ERROR. PHILIPS HAS NO FUNCTIONING SITE AVAILABLE. THEN THEIR PHONE MESSAGING SAYS "HIGHER THAN USUAL CALL WAITING". PERHAPS BECAUSE WHEN EVERY CONSUMER AFFECTED IS UNABLE TO GET TO THE RECALL WEBSITE THEY MUST CALL INSTEAD. THIS SEEMS AN INTENTIONAL BALK OR FRICTION-ADDING MOVE BY PHILIPS TO SAY THEY ARE RESPONDING BUT MAKING IT HARD FOR CONSUMERS IMPACTED BY THIS HELPIMPACTING CPAP PRODUCT TO GET A SOLUTION. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-09
MW5103854 | 2021-09-10 | 2020-08-15 | I, n, j, u, r, y | Voluntary report | DREAMSTATION HUMCORC PACK DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS, INC. | IP 22 |  |  | SINCE I HAVE HEART ISSUES AND UNABLE TO USE MACHINE. I SLEEP AT MOST 4-5 HOURS A NIGHT PUTTING ME AT RISK BECAUSE OF HEART ISSUES. COULD NOT BREATHE WHILE USING PHILIPS DREAMSTATION FOR SLEEP APNEA. I WOULD WAKE UP AFTER 1-2 HOURS USE NOT BEING ABLE TO BREATHE. I DISCONTINUED USE AND REPORTED TO DR. (B)(6) SHORTLY AFTER AS WELL AS PRIMARY DOCTOR (B)(6). DR. (B)(6) REFERRED ME TO (B)(6) PULMONARY SPECIALIST FOR EVALUATION. NOT ABLE TO USE AND ARE AT RISK WITH APNEA. MY BREATHING WAS BETTER THAN PREVIOUS YEAR. I KNEW THAT BECAUSE I HAD DISCONTINUED USE OF DREAMSTATION DEVICE. I NOTICED DIFFICULTY BREATHING APPROXIMATELY 2 YEARS BEFORE STOPPING USING DEVICE E. MY NOSE WAS STOPPING UP AND IT WOULD WAKE ME IN 2-3 HOURS. HAD PULMONARY TEST AND DR. (B)(6) RECEIVED NOTICE THAT RESULTS OF TEST INDICATED LUNG ISSUES. DR. (B)(6) DID NOT WANT ME TO STOP TAKING AMIODORONE WHICH WAS USED TO TREAT MY AF IB. HE CALLED THE PULMONARY DOCTOR AND HE RE-ISSUED A REPORT SAYING I COULD CONTINUE USE OF AMIODORONE. MY BREATHING CONTINUED TO GET WORSE USING CPAP AND I STOPPED AND USED ADHESIVE STRIPS ON MY NOSE TO BREATHE BETTER AT NIGHT. MY BREATHING TEST WAS BETTER THIS YEAR AFTER STOPPING USING CPAP BUT MY SLEEP APNEA IS WORSE AND WAS TOLD THAT IT WOULD TAKE AN UNKNOWN PERIOD OF TIME TO FIX PROBLEM WITH MY MACHINE. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5103867 | 2021-09-10 | 2020-02-12 | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS, INC. |  |  |  | I HAVE A PHILIPS DREAMSTATION CPAP MACHINE THAT I HAVE BEEN USING FOR 5 YEARS. I STARTED HAVING PROBLEMS WITH CONTINUOUS COUGH, SHORT OR BREATH, HEAD ACHES ABOUT A 1 1/2 YEARS AGO. I SAW MY PULMONARY DR. ON (B)(6) 2020 TO SEE WHAT MIGHT BE WRONG. WAS PRESCRIBED SINGULAIR AND FLUTICASONE NASAL SPRAY AND WAS GIVEN FOLLOW UP APPOINTMENT IN 1 MONTH TO DETERMINE IF MEDS HELPED. NO RELIEF FROM MEDS SO WAS GIVEN SYMBICORT INHALER STILL WITH NO RELIEF. DR. SENT ME TO CARDIOLOGIST THINKING IT MAY BE RELATED TO MY HEART, BUT THE STRESS TEST WAS NEGATIVE. I HAVE CONTINUED TO SEE MY PULMONOLOGIST AT LEAST EVERY MONTH. I ALSO SAW MY THYROID DR. FOR HYPOTHYROIDISM AND HAD SEVERAL THYROID SCANS THAT WERE FINE. ON (B)(6) 2021 I WENT FOR ANOTHER THYROID SCAN AND THEY FOUND A BINO DUAL GOITER. I QUIT USING MY CPAP MACHINE WHEN I SAW A NOTICE OF THE RECALL ON MY CPAP MACHINE WHICH SAID TO STOP USING IT DUE TO PROBLEMS RELATED TO SOME FOAM USED IN THE MACHINE. NOW I DO NOT SLEEP WELL BECAUSE I DO HAVE SLEEP APNEA. I HAVE CALLED PHILIPS AND WAS TOLD IT MAY BE 12 MONTHS BEFORE THEY CAN FIX OR REPLACE MY MACHINE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03912 | 2021-09-10 | 2021-08-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | A VENTILATOR WAS SERVICED BY A FIELD SERVICE ENGINEER. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE BY THE FIELD SERVICE ENGINEER, AN ISSUE RELATED TO THE EXTERNAL FLOW SENSOR CABLE WAS OBSERVED. THE EXTERNAL FLOW SENSOR CABLE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03913 | 2021-09-10 | 2021-08-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S NAVIGATIONAL BUTTON WAS OBSERVED. THE DEVICE'S FRONT PANEL BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03922 | 2021-09-10 | 2017-06-01 | I, n, j, u, r, y | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED AN EYE REACTION. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTION MEDICATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED AN EYE REACTION. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTION MEDICATION. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03622 | 2021-09-10 | 2021-08-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S POWER BUTTON DID NOT FUNCTION AS DESIGNED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S FRONT PANEL BOARD AND KEYPAD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03923 | 2021-09-10 | 2021-08-09 | I, n, j, u, r, y | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED HEADACHE, VOMITING, DEHYDRATION, NAUSEA AND BRADYCARDIA. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF INTRAVENOUS FLUIDS AND HOSPITALIZATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED HEADACHE, VOMITING, DEHYDRATION, NAUSEA AND BRADYCARDIA. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF INTRAVENOUS FLUIDS AND HOSPITALIZATION. AFTER RECEIVING FURTHER INFORMATION FROM THE PATIENT IT IS DETERMINED THAT THE INITIAL REPORT SHOULD HAVE BEEN FILED AS ADVERSE EVENT ONLY. SECTION B1 AND B2 HAS BEEN UPDATED/CORRECTED IN THIS REPORT. | Y | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03931 | 2021-09-10 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2200X11B | DS2200X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TESTING. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03932 | 2021-09-10 | 2021-08-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | FR2110X14B | FR2110X14B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S BLOWER AND MACHINE FLOW SENSOR WERE REPLACED TO ADDRESS THE ISSUE. THE DEVICE HAD EVIDENCE OF WATER INGRESS. | IT WAS INITIALLY REPORTED, DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S BLOWER AND MACHINE FLOW SENSOR WERE REPLACED TO ADDRESS THE ISSUE.THE DEVICE HAD EVIDENCE OF WATER INGRESS. THE DEVICE'S SYATEM BOARD WAS ALSO REPLACED DUE TO A TEST FAILURE. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
MW5103836 | 2021-09-10 | 2021-04-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DNX500H11C |  |  | THIS PAST LATE SPRING I BECAME AWARE THAT I WAS EXPERIENCING A SHORT DRY COUGH SEVERAL TIMES A DAY. THIS IS A NEW SYMPTOM. I AM NOT CONGESTED, DO NOT HAVE COVID, HAVE BEEN VACCINATED, AND AM OTHERWISE HEALTHY. THE COUGH SEEMS TO COME FROM A PERCEIVED IRRITATION IN MY CHEST; NOT MY THROAT OR NASAL AREA. I HAD BEEN USING A CPAP MACHINE FOR ABOUT 10 YEARS WHEN I SWITCHED TO THEY PHILLIPS "DREAMSTATION" CPAP MACHINE IN EARLY (B)(6) 2020. I HAVE NOW STOPPED USING THE RECALLED CPAP MACHINE, (B)(6) 2021. I HAVE NOT REPORTED THIS TO MY DOCTORS. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03665 | 2021-09-10 | 2021-08-27 | D, e, a, t, h | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT BECAME DISCONNECTED FROM A VENTILATOR AND EXPIRED. THE MANUFACTURER IS CURRENTLY INVESTIGATING THIS EVENT AND A FOLLOW-UP REPORT WILL BE FILED WHEN THE INVESTIGATION IS COMPLETE. | IT WAS INITIALLY REPORTED, THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT BECAME DISCONNECTED FROM A VENTILATOR AND EXPIRED. THE DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE PASSED ALL TESTING AND WAS FOUND TO OPERATE AND ALARM AS DESIGNED. | Y | N | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03799 | 2021-09-10 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR SHUTS DOWN RANDOMLY. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR SHUTS DOWN RANDOMLY. THERE WAS NO HARM OR INJURY REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03908 | 2021-09-10 | 2021-08-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03746 | 2021-09-10 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE 3 ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED DUE RO MULTIPLE E53 ERROR CODES. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03793 | 2021-09-10 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED THAT A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM ALLEGEDLY BECAME DEGRADED AND THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT ALSO ALLEGED SEEING PARTICLES IN THE DEVICE AND HAVING COUGH,CHEST PRESSURE AND HEADACHE. THE DEVICE WAS RETURNED TO THE MANUFACTURER QUALITY PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. DURING THE INITIAL VISUAL INSPECTION OF THE DEVICE FOUND NO EVIDENCE OF EVIDENCE OF DAMAGE, MALFUNCTION, OR SIGNIFICANT CONTAMINATION. THE DEVICE WAS DISASSEMBLED FOR INTERNAL VISUAL INSPECTION. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM. A SMALL AMOUNT OF DUST ACCUMULATION WAS OBSERVED IN THE BLOWER AND ON THE IMPELLER FINS. THE MANUFACTURER POWERED UP THE DEVICE AND VERIFIED AIRFLOW.THE DEVICE LOG SHOWED NO ERRORS WERE LOGGED. THE MANUFACTURER CONCLUDES THAT THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE CONTAMINATION FOUND WITHIN THE DEVICE IS CONSISTENT WITH BEING FROM AN EXTERNAL SOURCE. SECTION H6 HAS BEEN UPDATED/CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03747 | 2021-09-10 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGED BLACK PARTICLES IN CHAMBER AND ALSO STATED COUGHING WHILE ON THE DEVICE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03750 | 2021-09-10 | 2021-08-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX700H11C | DSX700H11C |  | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING/CHAMBER RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. SECTION B5 SHOULD BE PREVIOUSLY REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING/CHAMBER RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION A1 WAS CORRECTED (NOT CAPTURED IN INITIAL MDR). SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03749 | 2021-09-10 | 2021-08-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11 | DSX700H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03794 | 2021-09-10 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION B5 SHOULD BE REPORTED AS THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03751 | 2021-09-10 | 2021-08-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11 | DSX700T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03748 | 2021-09-10 | 2021-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED AND KIDNEY, SINUS INFECTION AND ISSUES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION G1, H6 WAS CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03836 | 2021-09-10 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION EXPERT W/HUM/CELL, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | CAX501H12C | CAX501H12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN THE INITIAL REPORT, SECTION B5 WAS NOT UPDATED, THE CORRECT B5 SHOULD BE- THE PATIENT HAS ALLEGED SEEING PARTICLES IN THE DEVICE, EYE IRRITATION, NASAL/THROAT IRRITATION. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL AND INTERNAL VISUAL INSPECTION OF DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO EVIDENCE OF CONTAMINATION. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERROR LOGGED. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION D9, G3, H6 HAVE BEEN UPDATED/CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03752 | 2021-09-10 | 2021-08-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | PREVIOUS REPORT HAS INCORRECT H3 SECTION. H3 SECTION HAS BEEN CORRECTED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03886 | 2021-09-10 | 2021-07-13 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE /RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM THAT BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE NASAL/THROAT IRRITATION AND ASTHMA. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. AFTER THE FIRST ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HEALTH EFFECT- CLINICAL CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03888 | 2021-09-10 | 2021-07-19 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP PNEUMONIA. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT DEVELOPED PNEUMONIA. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT.THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE AND FOUND EVIDENCE OF CONTAMINATION CONSISTENT WITH KERATIN AROUND THE AIRINLET. THE MANUFACTURER FOUND DUST CONTAMINATION ON AND IN THE BLOWER.THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND ZERO ERRORS LOGGED.THE MANUFACTURER CONCLUDES NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. THE MANUFACTURER CONFIRMED THE PRESENCE OF DUST AND KERATIN FOUND WITHIN THE DEVICE.SECTION D9, G3, H6 HAS BEEN UPDATED / CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03907 | 2021-09-10 | 2021-08-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03865 | 2021-09-10 | 2021-08-13 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03862 | 2021-09-10 | 2021-08-13 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP BREAST CANCER. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03867 | 2021-09-10 | 2021-07-13 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP EYE IRRITATION AND LUNG INFLAMMATION. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON DEC 26, 2023 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP EYE IRRITATION AND LUNG INFLAMMATION. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY THE PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03864 | 2021-09-10 | 2021-08-13 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03866 | 2021-09-10 | 2021-07-19 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THE PATIENT IS UNDERGOING TREATMENT IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03868 | 2021-09-10 | 2021-07-13 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP KIDNEY CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 08, 2024, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING INTERNAL ORGAN DAMAGE AND KIDNEY CANCER THAT ARE RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE SECTION B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY THE PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03911 | 2021-09-10 | 2021-08-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03887 | 2021-09-10 | 2021-07-13 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO HAVE A MYOCARDIAL INFARCTION. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER INITIALLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO HAVE A MYOCARDIAL INFARCTION. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. IN THE INITIAL REPORT SECTION B5 WAS NOT UPDATED, THE CORRECT B5 SHOULD BE - THE PATIENT HAS ALLEGED HEADACHES, SINUS INFECTION, COUGH, FAITGUE, THROAT IRRITATION, MYOCARDIAL INFARCTION AND UNDERWENT OPEN HEART SURGERY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO EVIDENCE OF CONTAMINATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF WATER INGRESS IN THE BLOWER BOX AND ON THE BLOWER MOTOR, CONTAMINATION ON INLET PORT. THE MANUFACTURER FOUND EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. A HUMIDIFIER WAS ALSO RETURNED TO THE MANUFACTURER ALONG WITH DEVICE. AN INTERNAL VISUAL INSPECTION OF THE HUMIDIFIER WAS COMPLETED BY THE MANUFACTURER AND FOUND NO EVIDENCE OF CONTAMINATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERROR LOGGED. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH WATER INGRESS IN THE BLOWER BOX AND ON THE BLOWER MOTOR, CONTAMINATION ON INLET PORT. THE MANUFACTURER CONFIRMED THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION D9,G3,H6 HAS BEEN UPDATED/CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03811 | 2021-09-10 | 2021-08-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, W/HUMID/HT, KR | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | KRX500T15 | KRX500T15 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03863 | 2021-09-10 | 2021-08-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC. | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP DECREASED LUNG FUNCTION. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP,AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP DECREASED LUNG FUNCTION,SEVERE PERIODIC PAIN JUST BEHIND THE STERNUM,ASSOCIATED MILD DIFFICULTY BREATHING,HEADACHE,ATTACK. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THERE IS NO CUSTOMER INFORMATION AND WE CANNOT REACH OUT TO THE CUSTOMER, HENCE, ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION CAN NOT BE MADE AT THIS TIME. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 (HEALTH EFFECT - CLINICAL CODE) HAS BEEN CORRECTED AND (TYPE OF INVESTIGATION,INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS) HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03828 | 2021-09-10 | 2021-07-15 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE COUGH AND WHEEZING. THE PATIENT WAS REPORTEDLY HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 07, 2024, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO EXPERIENCE PERSISTENT COUGH, PRODUCING PHLEGM, WHEEZING FROM THE DEVICE. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM (ONLY PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR). SECTION B2 HOSPITALIZATION WAS CHECKED. SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03909 | 2021-09-10 | 2021-08-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03910 | 2021-09-10 | 2021-08-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03812 | 2021-09-10 | 2021-08-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, KR | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC. | KRX500S15 | KRX500S15 |  | A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM WAS OBSERVED.THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03920 | 2021-09-10 | 2021-09-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03830 | 2021-09-10 | 2021-08-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE SORE THROAT. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTION FOR ANTIBIOTICS. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 09 2024, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE SORE THROAT. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTION FOR ANTIBIOTICS. SECTION B1 WAS CORRECTED FOR BOTH ADVERSE EVENT AND PRODUCT PROBLEM. ( PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION. TO SERIOUS INJURY. SECTION H6 WAS CORRECTED AND UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03891 | 2021-09-10 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11 | DSX700H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER INITIALLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF POSITIVE DEPOSIT OF FOAM PARTICULATE ON THE FIT TOOL. EVIDENCE OF FOAM DEGRADATION ASSOCIATED TO THIS ALLEGATION. EVIDENCE OF LIQUID INGRESS WAS OBSERVED. THE MANUFACTURER FOUND EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERROR LOGGED. THE MANUFACTURER CONCLUDES THAT POSITIVE DEPOSIT OF FOAM PARTICULATE ON THE FIT TOOL. EVIDENCE OF FOAM DEGRADATION ASSOCIATED TO THIS ALLEGATION. EVIDENCE OF LIQUID INGRESS WAS OBSERVED. THE MANUFACTURER FOUND EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03829 | 2021-09-10 | 2021-07-19 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE DIFFICULTY IN BREATHING. THE PATIENT WAS PRESCRIBED AN INHALER IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03831 | 2021-09-10 | 2021-07-19 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP ASTHMA. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE REPORTED EVENT IS CURRENTLY UNKNOWN. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY SUBMITTED MDR 2518422-2021-03831-1 WITH INCORRECT SECTIONS B1, B2, H1, H6. CORRECTIONS TO PREVIOUS MDR ARE MADE IN THIS REPORT AS FOLLOWS. SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY PRODUCT PROBLEM WAS CHECKED IN PREVIOUS MDR). SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY IT WAS BLANK). SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6- HEALTH EFFECT - IMPACT CODE WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03833 | 2021-09-10 | 2021-08-12 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE LUNG AND HEART ISSUES. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE REPORTED EVENT IS CURRENTLY UNKNOWN. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON AUG 15, 2023, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED LUNG AND HEART ISSUES. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGED VISUALIZATION OF PARTICLES, DIZZINESS, HEADACHE, NAUSEA / VOMITING, INFLAMMATORY RESPONSE, OTHER. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03832 | 2021-09-10 | 2021-08-12 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP COPD (CHRONIC OBSTRUCTIVE PULMONARY DISEASE). THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE REPORTED EVENT IS CURRENTLY UNKNOWN. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. PATIENT ALLEGED THAT IT CAUSED COPD (CHRONIC OBSTRUCTIVE PULMONARY DISEASE). THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENT OF COPD WITH ASSOCIATED DIFFICULTY BREATHING AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. THE PATIENT HAS SWAPPED OUT THE DEVICE AT THE DISTRIBUTOR AND AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03779 | 2021-09-13 | 2021-08-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03774 | 2021-09-13 | 2021-08-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO LCD SCREEN WAS OBSERVED. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03839 | 2021-09-13 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AHX500H11 | AHX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED SHORTNESS OF BREATH.THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS UPDATED REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03842 | 2021-09-13 | 2021-08-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF BLACK PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY.THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE.THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT IS SCRAPPED DUE TO AGE. A CORRECTION TO B5 WAS MADE AND SHOULD BE:THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF BLACK PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM AND IN THE PREVIOUS MDR D8 WAS MISSING. SECTION H6 WERE UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03838 | 2021-09-13 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING NAUSEA, FELT LIKE HAD VERTIGO, FLUTTERING IN HEART, FILTER SEEN BLACK IN IT RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03930 | 2021-09-13 | 2019-12-30 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102748) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP MIGRAINE HEADACHES. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING MIGRAIN, CONGESTION, FLUSHING WITH SWEATING, RAPID HEART RATE, NAUSEA, DIARRHEA, SEVERE STOMACH PAIN, KERATITIS, ACHES, FATIGUE, SINUS, BRATHING ISSUE, FACIAL SWELLING, ALLERGY ISSUE, INFLAMMATION RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. SECTION H6 (CLINICAL CODE) CORRECTED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03928 | 2021-09-13 | 2021-05-27 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102881) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ATELECTASIS AND SWOLLEN LYMPH NODES. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY SUBMITTED MDR 2518422-2021-03928-1 WITH INCORRECT SECTIONS B1, B2, H1, H6. CORRECTIONS TOPREVIOUS MDR ARE MADE IN THIS REPORT AS FOLLOWS. SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM (ONLY PRODUCT PROBLEM WAS CHECKED IN PREVIOUS MDR). SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS (PREVIOUSLY IT WAS BLANK). SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6- HEALTH IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03780 | 2021-09-13 | 2021-08-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A BLANK LCD SCREEN WAS OBSERVED BY THE TECHNICIAN. THE DEVICE'S LCS SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03777 | 2021-09-13 | 2021-08-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUES. AN ERROR CODE RELATED TO THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03773 | 2021-09-13 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03781 | 2021-09-13 | 2021-08-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03778 | 2021-09-13 | 2021-08-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A FAILURE OF THE DEVICE TO POWER ON WAS OBSERVED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03939 | 2021-09-13 | 2021-08-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2200X11B | DS2200X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03893 | 2021-09-13 | 2021-07-15 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A STROKE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE /RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM THAT BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE HIGH BLOOD PRESSURE AND STROKE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. NO MEDICAL INTERVENTION/TREATMENT OR ADDITIONAL INFORMATION WAS REPORTED. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION. SECTION H6 HEALTH EFFECT- CLINICAL CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03927 | 2021-09-13 | 2018-07-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102372) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A BRAIN TUMOR. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON OCTOBER 09, 2023, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING TO DEVELOP A BRAIN TUMOR, DIZZY, HEARING LOST ON THE LEFT SIDE, GALL BLADDER ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGED EYE IRRITATION, DIZZINESS AND/ OR HEADACHE, KIDNEY DISEASE/ TOXICITY, LIVER DISEASE/ TOXICITY, LUNG DISEASE, BRAIN TUMOR (ACOUSTIC NEUROMA). SECTION E1 WAS UPDATED IN THIS REPORT. SECTION H6 HEALTH EFFECTS - CLINICAL CODES WAS UPDATED IN THE REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03890 | 2021-09-13 | 2020-04-22 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102684) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP BREAST CANCER. THE PATIENT REQUIRED CHEMOTHERAPY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION IN A VOLUNTARY MEDWATCH (MW5102684), ALLEGING A PATIENT DEVELOPED BREAST CANCER AND OBSERVED BLACK PARTICLES IN HUMIDIFER AND TUBING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT REQUIRED CHEMOTHERAPY IN RESPONSE TO THE REPORTED EVENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 HAS BEEN UPDATED IN THIS REPORT. SECTION D1, D4, G4 AND H4 HAS BEEN CORRECTED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03892 | 2021-09-13 | 2021-07-15 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED THE PATIENT HAVING ASTHMA. THE PATIENT DID RECEIVE MEDICAL INTERVENTION AND REPORTED TO BE IN THE NURSING HOME.THE REPORTED EVENT OF HAVING ASTHMA WAS REVIEWED BY THE PMS CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE.NO OTHER MEDICAL INTERVENTION/TREATMENT OR ADDITIONAL INFORMATION WAS REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03929 | 2021-09-13 | 2021-04-19 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102751) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA. THE PATIENT WAS PRESCRIBED ALBUTEROL IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED A VOLUNTARY MEDWATCH (MW5102751) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA. THE PATIENT WAS PRESCRIBED ALBUTEROL IN RESPONSE TO THE REPORTED EVENT. THE REPORTED EVENT IS ASSESED AS NOT RELATED TO THE DEVICE. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03889 | 2021-09-13 | 2021-03-26 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/WIFI,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11W | DSX500T11W |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102704) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A COUGH AND CHEST PRESSURE. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE /RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM THAT BECAME DEGRADED AND CAUSED THE PATIENT TO HOSPITALIZE DUE TO COUGH, ATYPICAL CHEST PRESSURE, NAUSEA. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REPORTED BY THE PATIENT. AFTER THE FIRST ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03776 | 2021-09-13 | 2021-08-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BLOWER MOTOR AND INTERNAL TUBING WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03894 | 2021-09-13 | 2020-12-30 | I, n, j, u, r, y | Manufacturer report | PHILLIPS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSXSOOHI1C | DSXSOOHI1C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102700) ALLEGING A DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG CANCER. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED A VOLUNTARY MEDWATCH (MW5102700) ALLEGING A DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG CANCER. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT. SINCE THERE IS NO CUSTOMER INFORMATION, THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5103884 | 2021-09-13 | 2019-09-01 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | RESPIRONICS REMSTAR SE DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | REMSTAR SE DOM |  |  | MY CPAP MACHINE APPARENTLY HAS BEEN HAVING INTERMITTENT FOAM DETERIORATION FOR SEVERAL YEARS. AFTER BEING NOTIFIED BY MEDICARE NOT TO USE MAY MACHINE THIS MAY HAVE BEEN MY PROBLEM. I HAD BEEN CHANGING FILTERS AND GEAR, TO GET THE PARTICLES TO STOP, OR LETTING THE MACHINE RUN FOR A WHILE TO CLEAR UP THE AIRWAY. THE CURRENT RECALL DOES NOT INCLUDE MY MACHINE A REMSTAR SE, SO THAT IS WHY I AM REPORTING THIS. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-09
MW5103891 | 2021-09-13 | 2021-02-15 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | RECALL OF PHILIPS DREAMSTATION CPAP, HAVE SINUS INFECTIONS, SKIN AND EYE IRRITATION. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03941 | 2021-09-13 | 2021-08-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03942 | 2021-09-13 | 2021-08-20 | I, n, j, u, r, y | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2200X11B | DS2200X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT'S BLOOD OXYGEN SATURATION LEVEL DECREASED WHILE ON A VENTILATOR. THERE WAS NO PERMANENT HARM OR INJURY. THE DEVICE WAS TESTED AT THE REPORTING FACILITY. THE VENTILATOR PASSED ALL TESTING AND WAS PLACED BACK INTO USE. THE EVENT APPEARS TO HAVE BEEN CAUSED BY THE PATIENT CONDITION. |  | Y | N | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03940 | 2021-09-13 | 2021-08-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03938 | 2021-09-13 | 2021-08-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2200X11B | DS2200X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. THE SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03960 | 2021-09-13 | 2021-09-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, GB | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | GBX500S15 | GBX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE HUMIDIFIER HAD BLACK WATER. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. THE EXTERNAL PART OF THE DEVICE WAS VISUALLY INSPECTED AND THE MANUFACTURER CONFIRMED THERE WERE NO SIGNS OF CONTAMINATION ON THE DEVICE. THE INTERNAL PART OF THE DEVICE WAS VISUALLY INSPECTED BY THE MANUFACTURER AND THE MANUFACTURER FOUND EVIDENCE OF CONTAMINATION ON THE BOTTOM PANEL OF THE DEVICE AND IN THE BLOWER BOX. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND 3 ERRORS LOGGED. THE DEVICE WAS PROVIDED POWER AND THE DEVICE PROVIDED AIRFLOW. THE MANUFACTURER CONCLUDES THE DEVICE HAD EVIDENCE OF CONTAMINATION. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION WITHIN THE DEVICE. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03775 | 2021-09-13 | 2021-08-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
MW5103965 | 2021-09-14 |  | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS, INC. |  | DSX500T11C |  | PT STATES THAT SHE BEGAN HAVING PERIODIC HEALTH ISSUES WHILE USING HER PHILIPS DREAMSTATION CPAP. PT REPORTS THAT WHILE USING HER CPAP, HER LEFT NOSTRIL WAS STOPPED UP AND THAT SHE HAD TO CHANGE POSITIONS OR GET UP TO WALK AROUND TO UNCLOG HER NASAL PASSAGE. ADDITIONALLY PT STATES THAT SHE STARTED HAVING FREQUENT LEFT NOSTRIL NASAL DRIP ABOUT 1 YEAR AGO. PT ALSO REPORTS THAT WHILE USING HER CPAP SHE WOULD GET A LIGHT HEADACHE WHILE SLEEPING AND WOULD NEED TO TAKE MEDICATION AND GET UP TO WALK AROUND SO THAT IT WOULD NOT WORSEN. AT PRESENT PT STATES SHE HAS STOPPED USING THE CPAP AND THAT SINCE HER HEADACHES, NASAL DRIP AND CLOGGED NOSE SENSATION HAVE IMPROVED. PT NOTES THAT DESPITE IMPROVEMENT OF SYMPTOMS THAT SHE HAS NOT HAD A FULL NIGHTS SLEEP SINCE STOPPING THE CPAP AND THAT SHE HAS NOTICED HER DEPENDENCE ON THE DEVICE BECAUSE WITHOUT IT SHE CANNOT SLEEP WELL. PT REPORTS THAT IF SHE TRIES TO SLEEP ON HER BACK LIKE SHE DID WITH THE CPAP SHE CANNOT BREATHE. PT REPORTED CONCERN ABOUT HER SYMPTOMS AND STATED THAT INITIALLY SHE DID NOT REALIZE THAT HER SYMPTOMS WERE A REACTION TO HER CPAP. PT STATES SHE HEARD ABOUT THE CPAP RECALL ON THE RADIO AND RECEIVED A RECALL LETTER WHICH PROMPTED HER TO REGISTER HER DEVICE AND STOP USING THE CPAP. OF NOTE, PT REPORTS THAT SHE HAS AN UPCOMING APPOINTMENT WITH HER DOCTOR FOR EVALUATION ON (B)(6) 2021. |  | Y | N | LAY USER/PATIENT |  | 2021-09
MW5103942 | 2021-09-14 | 2019-05-15 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO CPAP WITH HEATED HUMIDIFIER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500H11 |  |  | ON (B)(6) 2018 STARTED USING PHILIPS DREAMSTATION CPAP MACHINE. ON (B)(6) 2019 WENT TO EMERGENCY ROOM, DIAGNOSED WITH AFIB. ON (B)(6) 2020 CHEST DISCOMFORT AND BREATHING NOT RIGHT, SHORTNESS OF BREATH. STOPPED USING TREADMILL. ON (B)(6) 2020 WENT TO THE EMERGENCY ROOM - EXTREME AFIB. ON (B)(6) 2020 HEADACHES, HEAD AND CHEST CONGESTION. ON (B)(6) 2020 WENT TO EMERGENCY ROOM - DIAGNOSED WITH PREMATURE VENTRICULAR CONTRACTIONS ALL SYMPTOMS ABOVE CONTINUED. MANY TESTS WERE DONE BUT FOUND NOTHING WRONG. ON (B)(6) 2021 WENT TO EMERGENCY ROOM - MILD CHEST PAIN, SHORTNESS OF BREATH, HOT AND SWEATING. NO DIAGNOSIS. ON (B)(6) 2021 CARDIOLOGIST DID HEART CATH - NOTHING WRONG, ADVISED I GET PULMONARY EVALUATION. ON (B)(6) 2021 I FOUND OUT THE PHILIPS DREAMSTATION CPAP'S ARE BEING RECALLED BECAUSE THEY MAY CAUSE AIRWAY IRRITATION AND INFLAMMATION SKIN, EYE AND RESPIRATORY TRACT IRRITATION; HEADACHE, ASTHMA, ORGAN DAMAGE (LIVER, KIDNEY AND OTHERS) INFLAMMATORY RESPONSE, COUGH, SINUS INFECTION, UPPER AIRWAY IRRITATION, CHEST PRESSURE, CARCINOGENIC EFFECTS, (B)(6) 2021; I MESSAGED MY SLEEP DOCTOR AND ASKED FOR A NEW DIFFERENT CPAP MACHINE (B)(6) 2021. MY DOCTOR MESSAGED ME THAT SHE SENT A PRESCRIPTION TO MY DME COMPANY ON (B)(6) 2021. STARTED USING A NEW RES-MED CPAP MACHINE (B)(6) 2021. HAVE NOT DETECTED AFIB OR PVC'S TO DATE SINCE (B)(6). ALL EKG'S ARE NORMAL. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5103936 | 2021-09-14 |  | I, n, j, u, r, y | Voluntary report | DREAMSTATION 1 | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | SINCE I STARTED USING DREAMSTATION 1 ABOUT 5 YRS AGO AFTER BEING DIAGNOSED WITH SLEEP APNEA, I HAVE HAD WORSENING SHORTNESS OF BREATH THAT HAS PROGRESSED TO DYSPNEA WITH MINIMAL EXERTION AND HAVE BEEN DIAGNOSED WITH RESTRICTIVE LUNG DISEASE AND ASTHMA. I HAVE SOME WHEEZING AND COUGHING. THIS HAS AFFECTED MY QUALITY OF LIFE. MY PULMONOLOGIST HAS NOW PRESCRIBED OXYGEN. I AM A NON-SMOKER. I WAS NEVER NOTIFIED OF THE RECALL. I HEARD ABOUT IT ON THE NEWS AND FILED A CLAIM WITH PHILIPS AND THEY SENT A REPLACEMENT DREAMSTATION 2. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5103931 | 2021-09-14 | 2021-08-03 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DREAMSTATION |  |  | DIAGNOSED WITH BRONCHIAL PNEUMONIA FOR NO APPARENT REASON. FOUND THAT CPCP MACHINE IS BEING RECALLED FOR RESPIRATORY PROBLEMS. CONTINUOUS PROBLEMS WITH PHLEGM AND RESPIRATORY DISTRESS. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5103930 | 2021-09-14 | 2021-02-13 | I, n, j, u, r, y | Voluntary report | CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DREAM STATION - DSX20H11C |  |  | CHEST PAINS SENT ME TO EMERGENCY HOSPITAL. ALL TEST NEGATIVE EXCEPT LACTIC ACID WAS HIGH. THROUGH FURTHER BLOOD TESTS FOUND KIDNEY CANCER (CLEAR CELL CARCINOMA). HAD KIDNEY REMOVED (B)(6) 2021. I HAD CT SCANS, PET SCANS, BLOOD TESTS AND MULTIPLE DOCTOR APPOINTMENT (PCP - ONCOLOGY - CARDIOLOGY - ETC. SURGERY. FDA SAFETY REPORTED ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5103929 | 2021-09-14 | 2021-04-23 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500T11 |  |  | I WAS DIAGNOSED WITH SLEEP APNEA AND HAVE USED THE PHILIPS DREAMSTATION NONCONTINUOUS VENTILATOR AT NIGHT (SERIAL NUMBER (B)(4)) FOR THE PAST 4 YEARS. DURING THE SECOND YEAR OF USE, I NOTICED BLACK MOLD GROWING IN THE HUMIDIFIER IN AN AREA THAT COULD NOT BE CLEANED BECAUSE OF THE WAY THE PLASTIC OVERLAPS. I COMPLAINED TO THE COMPANY THAT PROVIDED THE VENTILATOR TO ME, (B)(6), AND ASKED FOR A REPLACEMENT HUMIDIFIER. THEY ADVISED THAT I PULL THE PLASTIC APART AND DRIP RUBBING ALCOHOL INTO THAT AREA. IT SEEMED TO WORK OK. I CONTINUED TO USE MY BIPAP BUT HAVE NOTICED THROAT IRRITATION, AND RED EYES THAT I BLAMED ON OTHER HEALTH SITUATIONS LIKE TIREDNESS AND WEARING AN N95 DURING A PANDEMIC. IN (B)(6) 2021, I HAD AN EYE EXAM WITH MY EYE DOCTOR. HE EXPRESSED CONCERN ABOUT THE LINE OF IRRITATION THAT FOLLOWED WHERE MY EYELIDS CLOSED. HE SUGGESTED I WEAR AN EYE MASK TO TRY TO PROTECT MY EYES FROM THE AIR DISCHARGE VENT THAT BLEW ON MY EYES AT NIGHT. THIS MONTH I NOTICED A LOT MORE MOLD GROWING IN THE HUMIDIFIER. I AM A NURSE AND SO I CLEAN ALMOST ALL MY BIPAP EQUIPMENT DAILY AND I WORKED HARD TO CLEAN IT AGAIN. I USED VINEGAR, MILD SOAP WHICH SEEMED TO HELP, HOWEVER THE MOLD QUICKLY GREW AGAIN. TODAY I RECEIVED A RECALL NOTICE FOR THE BIPAP MACHINE THAT I USE AND IT ALL SEEMED TO MAKE SENSE. I'M WRITING TO LET YOU KNOW THAT I WAS DEFINITELY AFFECTED AND REPORTED THE MOLD PROBLEM SOON AFTER STARTING TO USE THIS MACHINE, BUT DID NOT GET AN ADEQUATE RESPONSE FROM MY VENDOR. I UPLOADED PICTURES OF THE VENTILATOR AND THE MOLD THAT COULD NOT BE REMOVED THROUGH CLEANING. THANK YOU FOR THE WORK YOU ARE DOING AND I'M GLAD I FOUND OUT ABOUT THIS SITE. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5103927 | 2021-09-14 | 2019-06-01 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION AUTO CPAP HUM DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. | REF DSX500H11 |  |  | I USE A RECALLED PHILIPS DREAMSTATION AUTO CPAP SN# (B)(4), WHICH PHILIPS CONFIRMS IS AFFECTED BY THE RECALL ISSUED IN JUNE 2021. THE SYMPTOM I NOW REALIZE WAS PROBABLY CAUSED BY THE DEVICE WAS BURNING, WATERING EYES, SEVERE ENOUGH THAT I HAD TO GET OUT OF BED AND RINSE THEM WITH COLD WATER, AFTER WHICH I FELT RELIEF ABOUT 20 MINUTES LATER. THIS HAPPENED INTERMITTENTLY DURING THE PAST TWO YEARS OR SO, BUT IT NEVER OCCURRED TO ME THAT IT MIGHT BE BEING CAUSED BY THE CPAP MACHINE. I HAVE NO ALLERGIES AND LIVE IN A RELATIVELY HYPOALLERGENIC ENVIRONMENT, AND NEVER HAD THESE SYMPTOMS PRIOR TO USING A CPAP MACHINE. THIS WAS MY THIRD ONE (PREVIOUS MODELS DEVELOPED MECHANICAL ISSUES) BUT IN RETROSPECT I REALIZE THAT THE BURNING EYES STARTED AFTER I RECEIVED THE DREAMSTATION CPAP MACHINE. SINCE I SUFFER FROM SEVERE SLEEP APNEA, IT'S NOT CLEAR WHAT I CAN DO, GIVEN THE SHORTAGE OF THIS MEDICAL DEVICE. GIVEN HOW LONG IT TAKES TO DEVELOP CANCER, I SUPPOSE I WILL SUFFER ANXIETY ABOUT THAT FOR THE FORESEEABLE FUTURE. IT IS BECAUSE OF THAT RISK THAT I CHECKED "LIFE-THREATENING" BOX ABOVE. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03996 | 2021-09-14 | 2021-08-18 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION BIPAP AVAPS/ST | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | PHILIPS RESPIRONICS, INC. |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A ,BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED CANCER. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF SURGERY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03998 | 2021-09-14 | 2021-08-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report |  | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE'S INFORMATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03992 | 2021-09-14 | 2021-08-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102752) ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED (SINUS INFECTION, UPPER RESPIRATORY ISSUES, FEVER,HEADACHES, COUGH AND BRONCHITIS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION OF PRESCRIBED STEROIDS AND ANTIBIOTICS. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED VOLUNTARY MEDWATCH (MW5102752) ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED (SINUS INFECTION, UPPER RESPIRATORY ISSUES, FEVER,HEADACHES, COUGH AND BRONCHITIS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION OF PRESCRIBED STEROIDS AND ANTIBIOTIC. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL AND EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN BUT DID OBSERVED EVIDENCE OF THERMAL DAMAGE TO Q8 ON THE PCA. IT APPEARS THE THERMAL DAMAGE WAS ISOLATED TO THE COMPONENT. THERE WAS NO REPORT OF THE ENCLOSURE OF THE DEVICE BEING COMPROMISED. SECTION D9, G3 AND H6 HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5103958 | 2021-09-14 | 2021-09-02 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | PHILIPS RESPIRONICS | DSX 500T11C |  |  | ON (B)(6) 2021, I BEGAN HAVING HEADACHES, NAUSEA, AND DIZZINESS. HEADACHE SO BAD I TOOK AN ASPIRIN ONE DAY, EVEN THOUGH I AM NOT "SUPPOSED" TO, DUE TO KIDNEY DISEASE. AFTER A FEW DAYS, NAUSEA WENT AWAY, HOWEVER MY EYES WERE VERY IRRITATED AND CONTINUED TO HAVE HEADACHES AND DIZZINESS OCCURRING DAILY. AROUND THE (B)(6). SYMPTOMS CONTINUED, MY THROAT WAS SCRATCHY, AND I BEGAN FEELING CHEST PRESSURE AND HAD DIFFICULTY BREATHING, I HAVE COPD, ASTHMA AND HEART DISEASE SO CALLED MY DOCTOR AND SHE PRESCRIBED PROAIR AND USE NEBULIZER AS NEEDED TO SEE IF THOSE HELPED. TODAY WHILE CLEANING DREAMSTATION CPAP MACHINE NOTICED BLACK SPECKS. I KNEW THERE WAS A RECALL ON DREAMSTATION MACHINES, SO I LOOKED AT IT MORE CLOSELY, CHECKED TO SEE IF MY MACHINE WAS PART OF THE RECALL, IT IS, AND NOW STRONGLY SUSPECT MY DREAMSTATION CPAP MACHINE IS CAUSING THE HEALTH PROBLEMS I AM HAVING. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5103939 | 2021-09-14 | 2021-09-12 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | WOKE UP AFTER USING CPAP WITH NAUSEA,VOMITED MULTIPLE TIMES, ALSO DIZZY. USED CPAP FOR NAP LATER IN THE DAY AND GOT DIZZY SHORTLY AFTER APPLYING CPAP. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03969 | 2021-09-14 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S BOARDS WERE MELTED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD AND POWER MANAGEMENT WERE REPLACED BY THE CUSTOMER PRIOR TO RETURNING FOR EVALUATION. DURING THE EVALUATION A COMPONENT ON THE FRONT PANEL WAS FOUND TO HAVE THERMAL DAMAGE. THE DEVICE'S FRONT PANEL KEYPAD/LED BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03971 | 2021-09-14 | 2021-09-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, GB | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | RGBX500S15 | RGBX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND EVIDENCE OF DUST OR DIRT CONTAMINATE, AND DARK FIBER CONTAMINATE IN THE INLET PORT, INLET CANAL, AND OUTLET PORT. THE MANUFACTURER FOUND A LARGE AMOUNT OF A YELLOW STICKY CONTAMINATE ON THE UPPER OUTSIDE SURFACES AND IN THE CREVASSES OF THE BASE UNIT. THE MANUFACTURER FOUND THE P4 CONNECTOR AND OUTER HOUSING ALONG WITH THE ASSEMBLY SCREWS ON THE BOTTOM OF THE UNIT COMPLETELY CORRODED AND RUSTED. THE MANUFACTURER FOUND A LARGE AMOUNT OF ORANGE CORROSION RUST DEPOSITS ON THE BOTTOM OF THE BASE AND HUMIDIFIER UNITS. THIS ORANGE CORROSION IS POTENTIALLY DUE TO THE P4 CONNECTOR CORRODING FROM WATER OVERFLOW. THE MANUFACTURER COMPLETED AN INTERNAL VISUAL INSPECTION. THE MANUFACTURER FOUND EVIDENCE OF FOUND A MODERATE AMOUNT OF LIGHT BEIGE DUST AND DIRT CONTAMINATE, AND DARK FIBER CONTAMINATE, AND BROWN CONTAMINATE THROUGHOUT THE BASE UNIT, AROUND THE OUTLET PORT, AND UI PANEL CHANNELS. THE MANUFACTURER FOUND A LARGE AMOUNT OF WATER SPOTS IN THE BLOWER, BLOWER IMPELLERS, BLOWER SEAL AND BLOWER BOX. THE MANUFACTURER FOUND LARGE AMOUNT OF FLAKEY RUST CORROSION ON THE P4 CONNECTOR AND UI PANEL CHANNELS. THE MANUFACTURER ALSO FOUND WATER SPOT CORROSION AND POTENTIAL MINERAL DEPOSITS WERE LOCATED IN THE INLET PORT TO THE HUMIDIFIER, THE HUMIDIFIER CHAMBER, OUTLET PORT, AND ON THE HEATER PLATE. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND THE ERROR LOG SHOWED DIFFERENT INSTANCES OF E30 AND E193. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH DUST OR DIRT CONTAMINATE, YELLOW STICKY CONTAMINATE, ORANGE CORROSION RUST, DARK FIBER CONTAMINATE AND POTENTIAL MINERAL, INCONSISTENT WITH THE SOUND ABATEMENT FOAM. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION D9, D10, G3, H3 AND H6 HAS BEEN UPDATED / CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03983 | 2021-09-14 | 2021-08-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03970 | 2021-09-14 | 2021-08-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03933 | 2021-09-14 | 2021-08-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03899 | 2021-09-14 | 2021-08-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S NAVIGATIONAL BUTTON WAS OBSERVED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S FRONT PANEL/KEYPAD LED BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03900 | 2021-09-14 | 2021-08-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03935 | 2021-09-14 | 2021-08-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03934 | 2021-09-14 | 2021-08-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | RIN2110X15B | RIN2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04022 | 2021-09-14 |  | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING WORSENING ASTHMA, ODD DRY COUGH, SHORTNESS OF BREATH, CHEST TIGHTNESS, MIGRAINE AND CLUSTER HEADACHES RELATED TO BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM (PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR) SECTION B2 WAS MADE OTHER SERIOUS (PREVIOUSLY IT WAS BLANK) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY SECTION H6 HEALTH EFFECT- IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03968 | 2021-09-14 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BLOWER MOTOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03902 | 2021-09-14 | 2021-09-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03901 | 2021-09-14 | 2021-08-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03903 | 2021-09-14 | 2021-09-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND INTERFACE BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03847 | 2021-09-14 | 2021-08-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2200X11B | DS2200X11B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S BLOWER CHASSIS AND SYSTEM BOARD TUBING, SYSTEM BOARD AND OXYGEN BLENDING MODULE BOARD AND HARNESS NEED TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03898 | 2021-09-14 | 2021-08-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED A TEST STEP DURING TESTING. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03985 | 2021-09-14 | 2021-08-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03984 | 2021-09-14 | 2021-08-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03904 | 2021-09-14 | 2021-09-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04021 | 2021-09-14 | 2021-05-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03982 | 2021-09-14 | 2021-08-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BLOWER MOTOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03994 | 2021-09-14 | 2021-08-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING HARD TO BREATH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING HARD TO BREATH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. . | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04016 | 2021-09-14 |  | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING FOAM FILTER IS DISSIPATING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. IN THE PREVIOUS REPORT, SECTION E1 (EMAIL ADDRESS) WAS CAPTURED WRONGLY, CORRECTED IN THIS REPORT. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04017 | 2021-09-14 | 2021-07-07 | D, e, a, t, h | Manufacturer report | DREAMSTATION BIPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX600S11 | DSX600S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A FUNGAL INFECTION IN THE LUNGS. THE PATIENT REPORTEDLY RECEIVED MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTION ANTIFUNGAL AND WAS HOSPITALIZED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON OCT 09, 2023 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED A VOLUNTARY MEDWATCH (MW5102626) ALLEGING A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING FUNGAL INFECTION IN THE LUNGS RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT REPORTEDLY RECEIVED MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTION ANTIFUNGAL AND WAS HOSPITALIZED. ADDITIONAL INFORMATION WAS RECEIVED ABOUT ALLEGATION OF CANCER AND DEATH. THE SECTIONS E1, B2, H1 AND H6 HAVE BEEN UPDATED IN THIS REPORT AS FOLLOWS: SECTION B2 WAS UPDATED TO DEATH. (PREVIOUSLY, IT WAS ONLY OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO DEATH. SECTION H6 HEALTH EFFECTS: CLINICAL CODES AND HEALTH EFFECTS- IMPACT CODE WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04018 | 2021-09-14 | 2021-08-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A SORE THROAT AND TONSIL ABSCESS. THE PATIENT REPORTEDLY RECEIVED MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTION ANTIBIOTICS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED A VOLUNTARY MEDWATCH REPORT (MW5102915) FROM A USER ALLEGING THE DREAMSTATION AUTO BIPAP CAUSED SORE THROAT AND TONSIL ABSCESS. THE PATIENT REPORTEDLY RECEIVED MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTION ANTIBIOTICS. THE REPORTED EVENT OF SORE THROAT AND TONSIL ABSCESS AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS MAY RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURE INSPECTED EXTERNALLY OBSERVED THAT DUST/DIRT CONTAMINATION INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM WAS OBSERVED IN DEVICE ENCLOSURE AND AIRPATH. SLIGHT BLACK CONTAMINATION AT THE BLOWER BOX IS CONSISTENT WITH THE KERATIN CONTAMINATION. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE MANUFACTURER OBSERVED PRESENCE OF CONTAMINATION IN THE DEVICE AND ARE CONSISTENT WITH BEING FROM AN EXTERNAL SOURCE. SECTION H6 TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04015 | 2021-09-14 | 2021-08-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED FREQUENT UNSPECIFIED RESPIRATORY ISSUES. THE PATIENT REPORTEDLY RECEIVED MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTION STEROIDS. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED FREQUENT UNSPECIFIED RESPIRATORY ISSUES. THE PATIENT REPORTEDLY RECEIVED MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTION STEROIDS. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03921 | 2021-09-14 | 2021-07-20 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX500T12 | CAX500T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE ASTHMA. THE PATIENT WAS REPORTEDLY HOSPITALIZED IN RESPONSE TO THE EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PATIENT TO HAVE ASTHMA. THE PATIENT WAS REPORTEDLY HOSPITALIZED IN RESPONSE TO THE EVENT RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE EVENT IS CURRENTLY UNKNOWN. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. DURING THE EXTERIOR INVESTIGATION, UNKNOWN DUST OR DIRT CONTAMINATION WAS PRESENT ON ALL SURFACES OF THE BASE UNIT. THERE IS AN UNKNOWN DUST CONTAMINATE PRESENT AT THE AIR INLET WHERE THE FILTER WOULD BE AND AT THE ISO PORT ENTRANCE. DURING THE INTERIOR INVESTIGATION, THERE WAS UNKNOWN DEBRIS IN THE TRACKS OF THE UI PANEL. THERE IS AN UNKNOWN WHITE DUST CONTAMINATION FOUND ON THE BOTTOM ENCLOSURE, BLOWER BOX HOUSING, BLOWER MOTOR, BLOWER IMPELLER, AND REAR PANEL O-RING. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN WAS NOT OBSERVED IN THE BASE UNIT. A KERATIN-LIKE SUBSTANCE WAS OBSERVED AROUND THE BLOWER BOX OUTLET. THE UNKNOWN DUST/DIRT CONTAMINATION AND FIBERS FOUND ON THE BLOWER CASING/IMPELLER AND BLOWER BOX WERE INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM CONTAMINATION. THE PRESENCE OF CONTAMINATION THROUGHOUT THE AIRWAY SUGGESTS A SOURCE EXTERNAL TO THE DEVICE. MINERAL SPOTS CONSISTENT WITH WATER INGRESS WERE FOUND WITHIN THE BLOWER BOX AND ON THE MOTOR CASING. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS. THE MANUFACTURER CONCLUDES UNABLE TO DIRECTLY ADDRESS THE SYMPTOMS OUTLINED IN THE COMPLAINT. CONFIRMED THAT THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN OBSERVED IN THE BASE UNIT. CONFIRMED THE PRESENCE OF CONTAMINATION IN THE AIRPATH. SECTION D4 WAS CAPTURED INCORRECTLY IN THE PREVIOUS REPORT. IT HAS BEEN CORRECTED IN THIS REPORT. IN THIS REPORT, SECTION D9, G3, H3, H6 HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03987 | 2021-09-14 | 2021-09-09 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, BR | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | BRX500S18 | BRX500S18 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP SPOT ON THE LUNGS. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THIS EVENT IS CURRENTLY UNKNOWN. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04019 | 2021-09-14 | 2020-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A (CPAP, BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PERSISTENT SINUS AND UPPER RESPIRATORY CONGESTION RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FOLLOW UP REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03974 | 2021-09-14 | 2021-08-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102695) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CHEST CONGESTION. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED A VOLUNTARY MEDWATCH (MW5102695) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CHEST CONGESTION. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. IN THE INITIAL REPORT SECTION B5 WAS NOT UPDATED, THE CORRECT B5 SHOULD BE - THE MANUFACTURER RECEIVED VOLUNTARY MEDWATCH (MW5102695).THE PATIENT HAS ALLEGED HARD COUGH, SHORTNESS OF BREATH, CONGESTION IN CHEST, FEVER, THROAT PROBLEM, LUNG INFECTION, SLEEP PROBLEM, PHELGM. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. IN THE INITIAL REPORT, SECTION B3 DATE WAS INCORRECTLY AND H10 SECTION WAS MARKED INCORRECTLY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF CONTAMINATION THROUGH-OUT; WHITE POWDERY SUBSTANCE INDICATIVE OF TALC/DUST .THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERROR LOGGED. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH CONTAMINATION THROUGH OUT THE DEVICE; WHITE POWDERY SUBSTANCE INDICATIVE OF TALC/DUST. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION B1,B2,B7,B3,D9,G3,H1,H6,H10 HAS BEEN UPDATED/CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03991 | 2021-09-14 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102712) ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | IN INTIAL REPORT, THE MANUFACTURER REPORTED INCORRECT INFORMATION IN SECTION H1 MENTIONED AS SERIOUS INJURY. AFTER REVIEW, IT WAS DETERMINED AS MALFUNCTION. THERE WAS NO PATIENT HARM OR INJURY OCCURED. SECTION H1 HAS BEEN CORRECTED IN THIS REPORT TO REFLECT THE CORRECT INFORMATION. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03963 | 2021-09-14 | 2021-07-20 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION BIPAP AUTOSV | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS,INC | D5X900T11 C | D5X900T11 C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102810) ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP BLOOD CLOTS IN THEIR KIDNEYS THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED A VOLUNTARY MEDWATCH (MW5102810) ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP BLOOD CLOTS IN THEIR KIDNEYS THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. SECTIONS D1, D2, D4, G4 AND H4 HAS BEEN CORRECTED / UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03965 | 2021-09-14 | 2021-07-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102834) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A STROKE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO EXPERIENCE A STROKE, COUGH, CHEST PAIN, BODY PAIN AND SHORTNESS OF BREATH. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENT OF STROKE AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. IN SECTION H6 HEALTH EFFECT - CLINICAL CODE, HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03966 | 2021-09-14 | 2014-02-04 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102716) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A COUGH AND CHRONIC BRONCHITIS. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON OCT 09, 2023, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102716) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A COUGH AND CHRONIC BRONCHITIS. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. ADDITIONAL INFORMATION WAS RECEIVED ABOUT ALLEGATION OF KIDNEY DISEASE/TOXICITY. SECTION H6 HEALTH EFFECTS: CLINICAL CODES WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03967 | 2021-09-14 | 2021-02-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102912) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS : THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED TO EXPERIENCE LUNG CANCER. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENT OF LUNG CANCER AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURER'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE INFORMATION AVAILABLE, THE MANUFACTURER CONCLUDES NO FURTHER ACTION IS NECESSARY. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILLED. SECTION B1, B2 HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION B7, SECTION H6 HEALTH EFFECT-IMPACT CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03943 | 2021-09-14 | 2019-01-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/HT/CELL, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | AUX400T15C | AUX400T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON JAN 11, 2024, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OCCURRENCE LUNG DISEASE SINCE COMMENCING TREATMENT IN JAN 2019. RESPIRATORY PROBLEMS STARTED 3 MONTHS LATER. HAVE INCREASED IN SEVERITY IN PAST YEAR. UNDERGO MANY RESPIRATORY TESTS UNDER A PHYSICIAN WHILE USING THE DEVICE. MEDICAL INTERVENTION WAS NOT SPECIFIED. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTION B1 WAS CORRECTED FOR BOTH ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) . (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION. TO SERIOUS INJURY. SECTION H6 HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03975 | 2021-09-14 | 2021-06-28 | I, n, j, u, r, y | Manufacturer report | PHILIPS DREAMSTATION BIPAP | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS INC |  |  |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102759) ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A RESPIRATORY INFECTION. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03993 | 2021-09-14 | 2021-07-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED KIDNEY CANCER. THE PATIENT HAD HIS KIDNEY SURGICALLY REMOVED. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE KIDNEY CANCER. THE PATIENT HAD HIS KIDNEY SURGICALLY REMOVED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H4 WAS MISSED TO CAPTURE, WHICH WAS UPDATED AND SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03972 | 2021-09-14 | 2019-01-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102835) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP AN ASTHMA ATTACK. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS INITIALLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE - THE PATIENT ALLEGED AND CAUSED A PATIENT TO DEVELOP AN ASTHMA ATTACK,COUGH, SINUS INFECTION RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN SECONDARY FINDINGS POTENTIAL DARK KERATIN PARTICLES WERE FOUND ON THE BLOWER BOX OUTLET PORT. THERE IS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. POTENTIAL DARK KERATIN PARTICLES WERE FOUND ON THE BLOWER BOX OUTLET PORT. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERORRS. THE MANUFACTURER CONCLUDES THE DEVICE HAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED IN THE BASE UNIT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03986 | 2021-09-14 | 2021-08-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S BLOWER MOTOR AND ACTIVE EXHALATION CONTROL MODULE WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03937 | 2021-09-15 | 2021-05-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | RIN2110X15B | RIN2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VARIATION IN THE VENTILATOR'S PRESSURE WAS OBSERVED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04054 | 2021-09-15 | 2021-08-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR STOPPED WORKING WHILE IN PATIENT USE. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR STOPPED WORKING WHILE IN PATIENT USE. THERE WAS NO HARM OR INJURY REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04039 | 2021-09-15 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED (B)(4). | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04052 | 2021-09-15 | 2021-08-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | CA2110X12B | CA2110X12B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION INDICATING THE DEVICE'S BATTERY WAS NOT CHARGING HAD OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR WAS ALARMING FOR A SERVICE REQUIRED ALARM CONDITION, INDICATING THE INTERNAL BATTERY WAS NOT CHARGING. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04053 | 2021-09-15 | 2021-08-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2200X11B | DS2200X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VARIATION IN THE VENTILATOR'S PRESSURE DELIVERY WAS OBSERVED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VARIATION IN THE VENTILATOR'S PRESSURE DELIVERY WAS OBSERVED. THE VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT COULD NOT BE CONFIRMED OR DUPLICATED. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04008 | 2021-09-15 | 2021-08-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | DS2200X11B | DS2200X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03853 | 2021-09-15 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS NOT WORKING. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF A VENTILATOR NOT WORK. THE DURABLE MEDICAL EQUIPMENT PROVIDER STATES THE DEVUCE WAS CHECKED AT THE PATIENTS HOUSE AND IS STIL IN USE. THE DEVICE WAS FOUND TO BE OPERATING AND POVIDING THERAPY AS DESIGNED. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04078 | 2021-09-15 | 2021-08-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE MANUFACTURER IS CURRENTLY INVESTIGATING THIS EVENT AND A FOLLOW-UP REPORT WILL BE FILED WHEN THE INVESTIGATION IS COMPLETE. | IT WAS INITIALLY REPORTED, THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04063 | 2021-09-15 | 2021-08-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040004 | 1040004 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR A LOW OXYGEN FLOW CONDITION. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
MW5104032 | 2021-09-15 |  | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | HAD DAILY HEADACHES. WENT TO SEVERAL DOCTORS COMPLAINING OF IT. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03989 | 2021-09-15 | 2021-08-16 | I, n, j, u, r, y | Manufacturer report |  | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED ASTHMA. IT IS UNKNOWN IF THE PATIENT REQUIRED MEDICAL INTERVENTION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORT A CPAP DEVICE'S SOUND ABATEMENT FOAM ALLEGEDLY BECAME DEGRADED AND CAUSED ASTHMA. IT IS UNKNOWN IF THE PATIENT REQUIRED MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THIS DEVICE WILL BE CORRECTED PER RES (B)(4). | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03462 | 2021-09-15 | 2021-08-20 | I, n, j, u, r, y | Manufacturer report | BIPAP A40 PRO | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | PHILIPS RESPIRONICS, INC. | AUX3100S19 | AUX3100S19 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BIPAP A40 PRO WAS INADEQUATELY VENTILATING THE PATIENT AND THE PATIENT WAS ADMITTED TO THE HOSPITAL FOR HYPERCAPNIA. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A BIPAP A40 PRO WAS ALLEGEDLY INADEQUATELY VENTILATING THE PATIENT AND THE PATIENT WAS ADMITTED TO THE HOSPITAL FOR HYPERCAPNIA. DESPITE OF THREE ATTEMPTS 07/18/2022, 08/15/2022 AND 09/08/2022; THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. THERE WAS NO RESPONSE FROM THE CUSTOMER TO HAVE THE DEVICE RETURNED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | Y | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03936 | 2021-09-15 | 2021-08-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | RIN2110X15B | RIN2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VARIATION IN THE VENTILATOR'S PRESSURE WAS OBSERVED. THERE WAS NO HARM OR INJURY REPORTED. THE PATIENT WAS INSTRUCTED TO REBOOT THE DEVICE TO RESET THE PRESSURE ISSUE. THE DEVICE WILL NOT BE RETURNING FOR EVALUATION AT THIS TIME. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
MW5103976 | 2021-09-15 |  | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS - DREAM STATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. |  | REFERENCE # DNX500H11C |  | MEDICAL DEVICE RECALL PHILIPS RESPIRONICS CPAP AND BI-LEVEL PAP DEVICES I CONTACTED PHILIPS TO A RECALL LETTER FOR MY CPAP. UPON CONTACTING PHILIPS VIA A TELEPHONE HELP/SERVICE LINE, I WAS INFORMED THAT PHILIPS DOES NOT KNOW WHEN THEY WILL BE ABLE TO REPAIR/REPLACE MY PHILIPS CPAP THAT IS PART OF THE RECALL. I ASKED IF THEY WOULD REIMBURSE FOR A TEMPORARY REPLACEMENT UNIT. I WAS TOLD WE DO NOT KNOW. I ASKED FOR A SUPERVISOR FOR ADDITIONAL INFORMATION. I WAS TOLD THERE IS NO ADDITIONAL INFORMATION. I ASKED WHAT DEPARTMENT I COULD SPEAK TO ABOUT TEMPORARY CPAP REPLACEMENT UNTIL PHILIPS HAS AN ANSWER TO MY HEALTH CONCERNS RELATED TO NOT HAVING ACCESS TO A CPAP. THEY GAVE ME AN OPTION TO ESCALATE MY CONCERNS RELATED TO THIS RECALL. I CHOSE TO ESCALATE MY ISSUE. I RECEIVED THE RECALL NOTICE LAST WEEK. I CONTACTED PHILIPS THIS MORNING 10:33 AM CT. IT WAS CONFIRMED THAT MY CPAP DEVICE IS PART OF THE RECALL. THERE WAS NO TIMEFRAME GIVEN FOR THE NEXT ACTION TO BE TAKEN BY PHILIPS INCLUDING THE ESCALATION OF MY HEALTH CONCERNS RELATED TO THE CPAP RECALL. MY DEVICE WAS REGISTERED IN THE RECALL CONFIRMATION NUMBER -(B)(4) MY NAME - (B)(6). I AM PROVIDING THIS INFORMATION TO THE FDA DUE TO THE POOR PREPARATION AND RESPONSE TO CONSUMERS BY PHILIPS. MY DOCTOR PRESCRIBED THIS CPAP AFTER SLEEP STUDY TESTING SUMMER 2020. I HAVE BEEN TESTED 4 TIMES SINCE THE 1990'S EACH TIME PRESCRIBED A CPAP WHICH I HAVE BEEN USING SINCE THE 90'S. PHILIPS SHOULD BE OFFERING AN OPTION FOR CONSUMERS TO TEMPORARILY REPLACE THE CPAP FOR EACH CONSUMER AFFECTED BY THE RECALL EITHER WITH A CPAP OR WITH A PROGRAM TO RENT A CPAP OR OBTAIN A LOANER. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | I |  | 2021-09
MW5103980 | 2021-09-15 | 2019-01-01 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS - SYSTEM ONE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | SYSTEM ONE |  |  | I BOUGHT A SOCLEAN MACHINE. I THEN NOTICED I WAS COUGHING WITH MY CPAP AT NIGHT. I ALSO NOTICED THAT THE MASK HAD BLACK MARKS THAT I COULD NOT WASH OFF. I NOW HAVE A NEW CPAP MACHINE THAT DOESN'T USE THE BLACK FOAM. I DON'T COUGH AS MUCH, BUT STILL DO ON OCCASION AND HAVE TO SLEEP WITH A COUGH DROP. I DO NOT USE MY SOCLEAN MACHINE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5103981 | 2021-09-15 | 2021-09-12 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS / RESPIRONICS, INC. |  |  |  | I HAVE A PHILIPS RESPIRONICS DREAMSTATION CPAP MACHINE FOR MY SLEEP APNEA. FIRST, IT HAS BEEN RECALLED AND THE MANUFACTURER WILL NOT PROVIDE AN ESTIMATE FOR REPLACEMENT. SECOND, THE MACHINE RANDOMLY STARTS BEEPING AND THEN SHUTS DOWN MULTIPLE TIMES PER NIGHT WHICH ADVERSELY AFFECTS MY BREATHING, CAUSES ME TO SNORE AND START CHOKING, AND IS DANGEROUS TO MY HEALTH. I OPENED A TICKET WITH PHILIPS AND THEY ESCALATED IT. THEY SAID SOMEONE SHOULD CALL WITHIN 2 WEEKS BUT IF THEY DON'T THEN CALL BACK. I CAN'T WAIT 2 WEEKS, I CANNOT BREATHE WHILE I AM SLEEPING. I CALLED THE DME SUPPLIER, (B)(4). THEY SAID BECAUSE THE MACHINE IS BEING RECALLED, THEIR TECHNICAL SUPPORT/REPAIR DEPARTMENT WILL NOT REPAIR IT. THEY HAVE NO LOANERS OR REPLACEMENT MACHINES IN STOCK. I HAVE CALLED 2 OTHER CPAP SUPPLIERS IN (B)(4) AND NEITHER OF THEM HAVE ANY CPAP MACHINES IN STOCK THAT I CAN PURCHASE; THEY ARE ALL BACK ORDERED AND THEIR IS NO ESTIMATE. THIS IS ABSURD AND RIDICULOUS. THIS IS A LIFE-SAVING MEDICAL DEVICE THAT I AM REQUIRED TO USE BECAUSE I CANNOT BREATHE ON MY OWN AT NIGHT AND ALL I AM GETTING IS A RUNAROUND, OBSTACLES, AND DEAD-ENDS. PHILIPS HAS SAID THAT THEY WOULD BEGIN REPLACING THE UNITS IN SEPTEMBER 2021 AND IT COULD TAKE UP TO A YEAR. A YEAR FOR CPAP MACHINE REPLACEMENTS FOR PEOPLE WHO CANNOT BREATH? I DON'T WANT TO EXAGGERATE BUT THIS IS REALLY A MATTER OF LIFE OR DEATH. A SLEEP STUDY RESULT INDICATED THAT I STOP BREATHING MULTIPLE TIMES PER MINUTE DURING SLEEP. THE MACHINE IS A NECESSITY FOR MY HEALTH, SLEEP, OXYGEN, AND BREATHING. IT IS UNACCEPTABLE THAT A MEDICAL DEVICE COMPANY DOES NOT CARE ABOUT THE HEALTH AND WELL-BEING OF ITS CUSTOMERS OR THE QUALITY OF ITS PRODUCTS AND DOES NOT STAND BY THEM. NOR DOES THE COMPANY PROVIDE ALTERNATIVES OR SOLUTIONS TO DEFECTS, AT LEAST NOT IN ANY TIMELY MANNER. HOW CAN SOMEONE WHO CAN'T BREATHE BE EXPECTED TO WAIT 2 WEEKS, 2 MONTHS, 6 MONTHS, A YEAR FOR A REPLACEMENT PRODUCT? HOW CAN SOMEONE WHO CAN'T BREATH BE TOLD THAT NO ONE WILL COME OUT TO REPAIR OR TROUBLESHOOT THE DEVICE JUST BECAUSE IT'S ON A RECALL LIST? I HAVE RUN OUT OF OPTIONS AND SUGGESTIONS. SOMETHING NEEDS TO BE DONE ABOUT THIS. I NEED A NEW OR REPLACEMENT SLEEP APNEA/CPAP MACHINE ASAP AND NO ONE CAN HELP ME AND NO ONE REALLY SEEMS TO CARE, ESPECIALLY THE MANUFACTURER, PHILIPS RESPIRONICS. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-09
MW5103982 | 2021-09-15 |  | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO CPAP. | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | UPPER RESPIRATORY/ CHEMICAL; PHILLIPS CPAP / NEVER NOTIFIED OF RECALL/ FOAM BROKE DOWN AND INGESTED. THIS IS RELATED TO THE PHILLIPS DREAMSTATION CPAP RECALL. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5103984 | 2021-09-15 |  | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DREAMSTATION 1 |  |  | WAS SUPPLIED WITH A PHILIPS CPAP DEVICE IN (B)(6) 2021 JUST PRIOR TO ITS RECALL. NO INFORMATION WAS PROVIDED ABOUT ANY DANGER ON THE USE OF THE MACHINE ALTHOUGH PHILIPS WAS AWARE OF THE PROBLEM. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-09
MW5103985 | 2021-09-15 | 2021-09-09 | I, n, j, u, r, y | Voluntary report | CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500H11C |  |  | WAS ADMITTED TO THE HOSPITAL AFTER GETTING VERY DIZZY. THE PHILIPS CPAP MACHINE HAS BEEN RECALLED AND I DON'T KNOW IF MY CONDITION WAS CAUSED BY THE RECALLED PRODUCT. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5103986 | 2021-09-15 | 2021-06-16 | I, n, j, u, r, y | Voluntary report | RESPIRONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I WAS DIAGNOSED WITH ASTHMA AND PRESCRIBED FLOVENT AND ALBUTEROL INHALERS. I HAVE NEVER HAD RESPIRATORY PROBLEMS IN MY LIFE. I WAS NOTIFIED BY RESPIRONICS THAT THE CPAP MACHINE I WAS USING, DREAMSTATION, WAS RECALLED DUE TO A FOAM WITHIN THE MACHINE FOR SOUND ABATEMENT DEGRADING AND BEING RELEASED INTO THE AIR PATHWAY AND CAN BE INHALED. IT MAY BE EXACERBATED BY USE OF OZONE CLEANERS WHICH I USED FOR ABOUT 3 MONTHS. ONE OF THE SIDE EFFECTS IS IRRITATION OF THE RESPIRATORY TRACT, COUGH, AND ASTHMA. I BEGAN HAVING SHORTNESS OF BREATH WITH EXERCISE AND DIFFICULTY EXHALING. I DO HAVE A CONGENITAL HEART DEFECT AND HAVE HAD SURGERY. BUT MY CARDIOLOGIST HAS PERFORMED TESTS AND DETERMINED THAT THESE SYMPTOMS WERE NOT HEART RELATED. I HAD PULMONARY FUNCTION TESTS DONE AND IT WAS ABNORMAL. MY EXPIRATORY VOLUME/PRESSURE WAS ABNORMAL. I AM FOLLOWED BY A PULMONOLOGIST. THE CPAP MACHINE HAS HELPED VERY MUCH WITH MY SLEEP AS I USED TO WAKE UP WITH HEADACHES EVERY DAY. HOWEVER I AM AFRAID TO CONTINUE USING IT. I FEEL IT COULD BE CONTRIBUTING TO MY RESPIRATORY PROBLEMS. THERE IS MODERATE DECREASE IN FEV1 AND FVC WITH NORMAL TLC SUGGESTING A MODERATE OBSTRUCTIVE VENTILATORY DEFECT. NORMAL DIFFUSION CAPACITY AND NON SPECIFIC AIRWAY RESISTANCE PATTERN IS NOTES. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5103998 | 2021-09-15 | 2021-07-14 | I, n, j, u, r, y | Voluntary report | DREAMSTATION 1 | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | MY CPAP MACHINE HAS BEEN RECALLED. AT THIS POINT, THERE ARE NO OTHER ALTERNATIVES FOR A REPAIR OR REPLACEMENT FROM PHILIPS DREAMSTATION CPAP DEVICE. I NEED MY MACHINE IN ORDER TO PREVENT SERIOUS HEALTH CONDITIONS WHICH INCLUDES MY HEART HEALTH. ACCORDING TO THE RECALL, THE MACHINE CAN EMIT FOAM PARTICLES AND GASES FROM THE BREAKDOWN OF THE FOAM. THIS GASES CAN CAUSE LUNG DAMAGE IN ADDITION TO THE HEART CONDITION FROM THE USE THE MACHINE. AT THIS POINT I FEEL THAT I AM IN A NO WIN SITUATION. THERE ARE RISKS TO USING MY MACHINE-THERE ARE RISKS IF I STOP USING IT. THIS IS ONE HOT MESS. PEOPLES LIVES HAVE BEEN PLACED IN JEOPARDY FROM THE USE OF THIS MACHINE. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5104025 | 2021-09-15 | 2021-09-12 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | RESPIRONICS DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | I FOUND OUT THROUGH A FRIEND THAT SOME CPAP'S WERE BEING RECALLED. I CHECKED INTO IT RIGHT AWAY AND REGISTERED MY DEVICE AS PART OF THE RECALL 7/21/2021 THROUGH THE PHILLIPS WEBSITE. NO ONE SAID TO STOP USING THE MACHINE AT THIS TIME. I CALLED (B)(6) TO ASK ABOUT REPLACEMENT OR IF I COULD RENT ANOTHER MACHINE- I WAS TOLD INSURANCE WOULDN'T COVER AND MACHINES ARE IN SHORT SUPPLY. I CALLED MY DR TO ASK ABOUT IT AND SAID I CAN CONTINUE TO USE MACHINE. AND SAID DON'T USE A OZONE CLEANER IF I HAVE BEEN- I GOT ONE FOR CHRISTMAS BUT ONLY USED 2 TIMES SINCE I DON'T LIKE THE SLIGHT SMELL IT LEAVES AND SPACE IT TAKES UP. ON (B)(6) 2021- CALLED (B)(6) TO ASK ABOUT THE VERY TINY BLACK DUST LIKE STUFF I NOTICED IN SINK WHEN CLEANING. THEY SAID THIS IS RELATED TO THE RECALL AND TO ASK DR IF I CAN STILL USE THE MACHINE. I ASKED IF I COULD GET A NEW OR DIFFERENT BRAND MACHINE- OR EVEN RENT. HE SAID I WOULD NEED MY DOCTOR TO WRITE RX FOR NEW MACHINE. CALLED DR (B)(6) OFFICE TO TELL ABOUT THE BLACK STUFF AND REQUEST SCRIPT. THEY WILL HAVE RX EARLY NEXT WEEK. DIDN'T SAY TO STOP USING MACHINE. CALLED PHILLIPS- THEY WILL PUT IN AN QUALITY ESCALATION FORM TO EXPEDITE REPLACEMENT OF MY MACHINE. THEY ASKED IF I HAVE NOTICED ANYTHING ELSE AND I TOLD ABOUT HOW DIRTY THE FILTERS WERE SO I REPLACED FILTERS ANYWAY. THEY REPORTED THAT ALSO IN MY INFO. ON (B)(6) 2021- CALLED (B)(6) TO ASK ABOUT USE OF MY MACHINE/AND REPLACING FILTERS MORE OFTEN- THEY RECOMMEND AN ADDITIONAL FILTER. SO WILL SEND TO ME. ON (B)(6) 2021- AFTER USING NEW ADDITIONAL FILTER FOR LESS THAN A WEEK, NOTICED IT IS LOOKING GREY SO STOPPED USING CPAP. ON (B)(6) 2021 CALLED DOCTOR (B)(6) OFFICE- TOLD ABOUT HOW OFTEN MY FILTERS ARE NEEDING CHANGED. AND NEW ADDITIONAL FILTER ON BACK GETTING DIRTY /GREY IN LESS THAN A WEEK. SHE WILL CHECK WITH (B)(6) IF I CAN STILL USE SINCE I HAVE SEVERE SLEEP APNEA AND CALL (B)(6) TO SEE IF THEY HAVE REFURB MACHINE. SAID THEY HAVEN'T HAD ANY PATIENTS WHO HAVE AN AFFECTED MACHINE, SO NOT SURE WHAT HE WILL RECOMMEND. ON (B)(6) - CALLED PHILLIPS 10:50AM- I AM JUST SICK WITH WORRY, IN TEARS THAT NO ONE TOLD ME TO STOP USING THIS MACHINE ONCE I NOTICED SIGNS IT IS AFFECTED. THE CALL TO PHILLIPS WAS NOT HELPFUL AT ALL. - SHE SAID THAT THE LAST TIME I CALLED THE REPRESENTATIVE SHOULD HAVE TOLD ME TO STOP USING THE AFFECTED MACHINE. AND THEY DID NOT SAY THAT. I ASKED ABOUT IF THEY HAVE INFORMATION ABOUT WHAT THIS COULD HAVE DONE TO MY LUNGS AND SHE SAID TO CHECK FOR INFORMATION ON THE WEBSITE. I ASKED IF THEY HAVE A MEDICAL TEAM / CONSULTANT THAT CAN ANSWER MEDICAL RELATED QUESTIONS DUE TO THE RECALL, SHE SAID NO THEY DO NOT HAVE A MEDICAL TEAM. WHEN I ASKED HOW LONG TILL I MAY GET A NEW MACHINE SINCE I EXPEDITED MY REPLACEMENT AND SHE SAID IT COULD BE 12 MONTHS. I AM JUST SICK. PHYSICALLY MAKES ME SICK THAT I COULD HAVE BREATHED IN THE PARTICLES AND GASSES AND COULD GET SICK FROM IT. NOW I REALIZE THAT MY EYE IRRITATION, ITCHY THROAT, THROAT AND SINUS CONGESTION RECENTLY IS LIKELY CAUSED BY MY CPAP! SO I'M WORRIED- WHAT OTHER PROBLEMS MIGHT THIS CAUSE FOR MY HEALTH?! I STOPPED USING THE MACHINE (B)(6) 2021. BUT EVEN TODAY THE REP AT KUBAT SAID THAT PHILLIPS HAD RECOMMENDED PATIENTS STOP USING THE MACHINES. WHEN?! NO ONE HAS TOLD ME. NEVER HEARD DIRECTLY FROM PHILLIPS, (B)(6) WHO I BOUGHT MY CPAP FROM, OR MY DOCTOR WHO I JUST SAW ON (B)(6) 2021 AFTER THE RECALL STARTED. MAYBE THEY HADN'T HEARD YET? I HAVE A HARD TIME FUNCTIONING WITHOUT THE CPAP. I EVEN STOPPED USING FOR A COUPLE OF WEEKS AFTER THE RECALL BUT JUST CAN'T FUNCTION WITHOUT IT. IT AFFECTS MY ABILITY TO FUNCTION WELL AT WORK, I GET EXHAUSTED, HEADACHES AND EVEN FORGETFUL. DON'T KNOW WHAT I CAN DO, CAN'T EVEN BUY A MACHINE BECAUSE NONE AVAILABLE. WHAT CAN I DO FROM HERE? SHOULDN'T PHILLIPS BE REQUIRED TO TELL PATIENTS? OR (B)(6) WHO SUPPLIES? EVEN THE PERSON WHO CALLS REGULARLY TO ORDER SUPPLIES DIDN'T KNOW OF RECALL. THIS SEEMS NEGLIGENT ALL THE WAY AROUND. (B)(6). THERE HAS TO BE OTHER PATIENTS WHO HAVE BEEN USING THEIR CPAP NOT KNOWING THE MACHINE WAS RECALLED. AND EVEN WORSE OTHER PATIENTS LIKE ME HAVE BEEN MISINFORMED ABOUT THE RECALL PROTOCOL AND DON'T KNOW THAT THEY SHOULD STOP USING THEIR MACHINES. I REALIZE THAT IN THIS BUSINESS EVERYONE POINTS TO EVERYONE ELSE TO MAKE THAT CALL TO STOP USING A MEDICAL DEVICE DUE TO LIABILITY, BUT AT WHAT COST TO THEIR PATIENTS? AS FAR AS I KNOW, I STILL WOULDN'T KNOW ABOUT THE RECALL IF MY FRIEND HADN'T MENTIONED IT KNOWING THAT I USE A CPAP. AND I TOLD MY SISTER WHO USES THE SAME MACHINE I DO, WHO DIDN'T KNOW ABOUT THE RECALL EITHER. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-09
MW5104028 | 2021-09-15 | 2021-09-10 | I, n, j, u, r, y | Voluntary report | DREAM MACHINE. C-PAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | J25564406C785 |  |  | HAVE CONTINUALLY FELT SCRATCHY THROAT. WAKE WITH HEADACHE SOB DAILY. USING CPAP EVERYDAY. BP STAYING IN 120/. AND 130/ S. BLOOD PRESSURE GOING UP 196/96 NOW THAT CPAP NOT AVAILABLE. FILTER I HAD IN MACHINE IS BLACK. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5104033 | 2021-09-15 | 2021-03-01 | I, n, j, u, r, y | Voluntary report | RESPIRONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | DNX500H11C |  |  | HAVE BEEN HAVING A SORE THROAT FOR MONTHS, I'VE SEEN MY PRIMARY CARD DOCTOR TWICE TRYING DIFFERENT THINGS TO FIND PROBLEM. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | I |  | 2021-09
MW5104034 | 2021-09-15 | 2021-09-12 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILLIPS RESPIRONICS DREAMSTATION CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | AUTOCPAP |  |  | USING A PHILIPS RESPIRONICS DREAMSTATION CPAP MACHINE, I NOTICED A HIGHLY UNUSUAL BUILDUP OF BLACK MATERIAL ON THE FILTERS. THIS UNIT IS CURRENTLY UNDER A RECALL. I HAVE PREVIOUSLY REGISTERED THIS MACHINE DUE TO THE RECALL. I HAVE NOW CHOSEN NOT TO CONTINUE USE OF THIS CPAP MACHINE AT THIS TIME. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-09
MW5104035 | 2021-09-15 | 2021-07-27 | I, n, j, u, r, y | Voluntary report | CPAP MACHINE DREAM STATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500H11C |  |  | I'VE BEEN USING PHILIPS RESPIRONICS CPAP MACHINE (DREAM STATION) SINCE 2018. HOWEVER AROUND (B)(6) 2020, I NOTICED THROAT IRRITATION, CHEST PRESSURE AND SHORTNESS OF BREATH. I TOOK COUGH MEDICATION AND USED ALBUTEROL TO CLEAR MY AIRWAY. ON (B)(6) OF THIS YEAR I WENT TO MY DOCTOR TO DISCUSS MY SHORTNESS OF BREATH. THE NURSE CHECKED MY OXYGEN SATURATION AND IT WAS ONLY 87%. IT WAS DROPPING WHILE THE DOCTOR WAS EXAMINING ME SO HE CALLED THE ATTENDING DOCTOR TO SEE ME AND THEY DECIDED TO SEND ME TO THE EMERGENCY ROOM. I WAS IMMEDIATELY GIVEN OXYGEN. THE ER DOCTOR ORDERED TESTS, CBC, CHEST X-RAY, 2D ECHO. THEY FOUND MY LUNGS WAS FULL OF FLUIDS AND INFLAMED. I STAYED IN THE HOSPITAL FOR ONE NIGHT AND THE NEXT MORNING I DID A SIX MINUTES WALK AND THE DOCTOR SAID FROM NOW ON, I WILL NEEDING SUPPLEMENTAL OXYGEN. PHILIPS IS NOW RECALLING THE CPAP MACHINE I WAS USING BECAUSE THE FOAM USED IN THE MACHINE MAY HAVE DEGRADED INTO PARTICLES WHICH COULD ENTERED IN THE DEVICE'S AIR PATHWAY INTO MY LUNGS. MY LIFE HAS BEEN ABRUPTLY CHANGED AND I'M NO LONGER ABLE TO DO A LOT OF THINGS BECAUSE I'M USING OXYGEN MACHINE. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | I |  | 2021-09
MW5103977 | 2021-09-15 | 2021-07-27 | I, n, j, u, r, y | Voluntary report | CPAP BRAND PHILLIPS RESPIRONIC DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I HAD COUGH AND CONGESTION WITH NO FEVER, HEADACHE, SORE THROAT OR SHORTNESS OF BREATH. THE COUGH AND CONGESTION HAVE BEEN PERSISTED FOR MORE THAN 3 MONTHS. I HAD BEEN TAKING MULTIPLE COURSE OF ANTIBIOTIC TO NO AVAIL. I HAD OCCASIONALLY SAW SOMETHING BLACK IN MY MOUTH. THE RESULT HAS NOT BEEN PROVIDED. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-09
2518422-2021-04004 | 2021-09-15 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED TO DEVELOP INFLAMMATORY LYMPH NODES IN THE LUNGS, COUGH, CANCER. THE REPORTED EVENT OF CANCER AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURER'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURER CONCLUDES NO FURTHER ACTION IS NECESSARY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. PIL/RIQL FINDS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN IN THE RETURNED UNIT. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE WAS HOOKED UP TO A KNOWN GOOD POWER SUPPLY AND POWER CORD, AIRFLOW WAS VERIFIED TO BE OPERATING CORRECTLY. THE ERROR LOG SHOWED ZERO INSTANCES OF ERRORS ON THE DEVICE. THE MANUFACTURER CONCLUDES THAT THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04040 | 2021-09-15 | 2021-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION B5 SHOULD BE PREVIOUSLY REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT ALLEGED DIFFICULT IN BREATHING/SHORT OF BREATH, LIGHT HEADACHES. IT WAS ALSO ALLEGED STRANGE ODOR AND BLACK SPOTS IN THE DEVICE. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03977 | 2021-09-15 | 2021-07-15 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX400T11 | DSX400T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A SPOT ON THEIR LUNG. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A SPOT ON THEIR LUNG. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. IN THE INITIAL REPORT, SECTION H10 WAS MARKED INCORRECTLY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO EVIDENCE OF CONTAMINATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF DIRT/DUST/DANDER CONTAMINATION. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND 12 INSTANCES OF .E-7, 5 INSTANCES OF E-105, 5 INSTANCES OF E-250, 1 INSTANCE OF E-22. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH DIRT/DUST/DANDER CONTAMINATION. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION D9,E1,G3,H6 HAS BEEN UPDATED/CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04043 | 2021-09-15 | 2021-07-15 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP WHEEZING. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04045 | 2021-09-15 | 2021-07-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX700H11 | DSX700H11 |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102862) ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A STROKE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED ON THIS DEVICE IN MDR 2518422-2021-02956. THIS REPORT WAS SUBMITTED AS A DUPLICATE REPORT OF THE PREVIOUSLY SUBMITTED REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04044 | 2021-09-15 | 2021-04-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103129) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A COUGH. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING COUGH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION B1 AND B2 WERE INCORRECTLY CAPTURED IN PREVIOUS REPORT. SECTION B1 IS CORRECTLY MARKED IN THIS REPORT AS PRODUCT PROBLEM ONLY. SECTION H1, H6(HEALTH EFFECT- IMPACT CODE) IS CORRECTED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03980 | 2021-09-15 | 2021-06-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03978 | 2021-09-15 | 2021-07-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/WIFI,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11W | DSX500T11W |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03981 | 2021-09-15 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED WAKES UP COUGHING UP MUCUS AND FEELING SICK TO STOMACH, SAW PARTICLES FLOATING IN THE WATER, MAKES NOISES, WATER CAKES UP AND FLOATS IN THE WATER, VOMITS. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WAS THREE ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED. CORRECTED INFORMATION PROVIDED IN B5 AND H6, THE INFORMATION (THE MANUFACTURER RECEIVED INFORMATION ALLEGED WAKES UP COUGHING UP MUCUS AND FEELING SICK TO STOMACH, SAW PARTICLES FLOATING IN THE WATER, MAKES NOISES, WATER CAKES UP AND FLOATS IN THE WATER, VOMITS) WAS NOT INCLUDED IN INITIAL REPORT. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03690 | 2021-09-15 | 2017-01-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH REPORT (MW5102804) FROM A USER ALLEGING THE DREAMSTATION AUTO BIPAP CAUSED AIR TRAPPING IN HER ABDOMEN WITH BLOATING AFTER A HIATAL HERNIA REPAIR SURGERY. AFTER REDUCTION OF AIR PRESSURE ON THE MACHINE WITH CONTINUED BLOATING, AN EMERGENCY SURGERY WAS PERFORMED DUE TO A NEW HIATAL HERNIA WITH BOWEL OBSTRUCTION. THE SURGEON FELT THAT THE PRESSURE FROM THE BIPAP CONTRIBUTED TO THE BOWEL BECOMING STUCK IN HER NEW HERNIA. THE PULMONOLOGIST PRESCRIBED ANOTHER PRESSURE REDUCTION TO A NON-THERAPEUTIC LEVEL WHICH GREATLY REDUCED THE BLOATING. THE USER DISCONTINUED THE USE OF THE DREAMSTATION AUTO BIPAP DUE TO THE RECALL AND IS USING AN OLDER NON-PHILIPS MACHINE AT A HIGHER PRESSURE AND IS NO LONGER EXPERIENCING BLOATING. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED A VOLUNTARY MEDWATCH REPORT (MW5102804) FROM A USER ALLEGING THE DREAMSTATION AUTO BIPAP CAUSED AIR TRAPPING IN HER ABDOMEN WITH BLOATING AFTER A HIATAL HERNIA REPAIR SURGERY. AFTER REDUCTION OF AIR PRESSURE ON THE MACHINE WITH CONTINUED BLOATING, AN EMERGENCY SURGERY WAS PERFORMED DUE TO A NEW HIATAL HERNIA WITH BOWEL OBSTRUCTION. THE SURGEON FELT THAT THE PRESSURE FROM THE BIPAP CONTRIBUTED TO THE BOWEL BECOMING STUCK IN HER NEW HERNIA. THE PULMONOLOGIST PRESCRIBED ANOTHER PRESSURE REDUCTION TO A NON-THERAPEUTIC LEVEL WHICH GREATLY REDUCED THE BLOATING. THE USER DISCONTINUED THE USE OF THE DREAMSTATION AUTO BIPAP DUE TO THE RECALL AND IS USING AN OLDER NON-PHILIPS MACHINE AT A HIGHER PRESSURE AND IS NO LONGER EXPERIENCING BLOATING. THE REPORTED EVENT OF LARGE HIATAL HERNIA THAT HAD PULLED OVER HALF OF PATIENT'S STOMACH THROUGH THE DIAPHRAGM AND ABDOMINAL BLOATED AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURE INSPECTED EXTERNALLY OBSERVED THAT SCRATCHES ON TOP, UNKNOWN BROWN RESIDUE ON KNOB, WHITE DUST DEPOSITS IN BLOWER OUTLET, WHITE DEPOSITS IN FILTER HOLE, DEBRIS CONSISTENT WITH DUST IN BOTTOM OF INSIDE OF DEVICE, WHITE DUST WAS OBSERVED O RING AND TOP OF BLOWER AND WATER MARKS ON BOTTOM OF BLOWER. THE MANUFACTURER USED A FITT (FOAM INTEGRITY TEST TOOL) AND WAS NOT ABLE TO CONFIRM THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION AND THERE IS NO VISIBLE DAMAGE OR FUNCTIONALITY FAILURES OF THE DEVICE. THE MA NUFACTURER OBSERVED DUST CONTAMINATION AND WATER MARKS IN THE DEVICE AND ARE CONSISTENT WITH BEING FROM AN EXTERNAL SOURCE. SECTION H6 TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04024 | 2021-09-15 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON OCT 09, 2023, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING BLACK PARTICLES IN LUNGS AND SHORTNESS OF BREATH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED ALLEGATION OF DIZZINESS AND/OR HEADACHE, BLACK DEPOSITS IN LUNGS / CYSTS IN RIGHT AND LEFT KIDNEYS. SECTION E1 WAS UPDATED SECTION H6 HEALTH EFFECTS - CLINICAL CODES. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04025 | 2021-09-15 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04062 | 2021-09-15 | 2021-06-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H1 IC | DSX500H1 IC |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04031 | 2021-09-15 | 2020-01-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11 | DSX700T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED A VOLUNTARY MEDWATCH REPORT (MW5102899) FROM A USER ALLEGING THE DREAMSTATION AUTO BIPAP CAUSED LUNG IRRITATION, CHEST PAIN AND SEEING SOME DEBRIS IN THE BOTTOM OF THE WATER RESERVOIR. THIS EVENT IS ASSESSED AS POSSIBLY RELATED TO THE DEVICE IN THIS CASE DUE TO USER ERROR ASSOCIATED WITH ANY OF THE FOLLOWING POSSIBILITIES: IMPROPER HUMIDIFICATION AND/OR TUBE TEMPERATURE SETTINGS, IMPROPER MASK FIT/USE, IMPROPER OR INADEQUATE TREATMENT PRESSURES, CERTAIN CLEANING CHEMICALS THE PATIENT MIGHT HAVE USED IN THE DEVICE AND/OR INADEQUATE CLEANING PRACTICES/FREQUENCY. ADDITIONALLY, THE REPORTED COMPLAINT COULD ALSO BE EXPLAINED BY DISEASE OR UNDERLYING CONDITION/PAST-MEDICAL HISTORY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURE INSPECTED EXTERNALLY OBSERVED THAT DUST/DIRT CONTAMINATION IN THE DEVICE ENCLOSURE AND AIRPATH, AND EVIDENCE WATER INGRESS ON THE BLOWER BOX. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION .THE MANUFACTURER OBSERVED DUST CONTAMINATION AND WATER INGRESS IN THE DEVICE AND ARE CONSISTENT WITH BEING FROM AN EXTERNAL SOURCE. SECTION H6 TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04003 | 2021-09-15 | 2021-07-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP PNEUMONIA AND LUNG DISEASE. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP,AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING CAUSED A PATIENT TO DEVELOP PNEUMONIA AND LUNG DISEASE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN SECONDARY FINDINGS PRIOR TO TAKING THE UNIT APART, THE DEVICE WAS OPERATED FOR 30 MINUTES IN WHICH IT WORKED AS DESIGNED. A FOREIGN WHITE POWDERY SUBSTANCE WAS FOUND THROUGHOUT THE BLOWER MOTOR AND BLOWER BOX. THE FINS OF THE BLOWER MOTOR AND OUTLET PORT OF THE BLOWER MOTOR ARE WHITE. ALSO EVIDENCE OF LIQUID INGRESS TO THE BLOWER BOX WAS OBSERVED. POTENTIAL DARK KERATIN PARTICLES WERE FOUND ON THE BLOWER BOX OUTLET PORT. IT WAS OBSERVED FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN,FOUND POTENTIAL DARK KERATIN PARTICLES ON THE BLOWER BOX OUTLET PORT,FOUND A WHITE POWDERY SUBSTANCE THROUGHOUT THE BLOWER MOTOR AND BLOWER BOX,OBSERVED LIQUID INGRESS MARKS ON THE INSIDE OF THE BLOWER BOX. IT WAS UNABLE TO ADDRESS THE ALLEGATION OF " PHENOMENA". THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO "ERRORS". THE MANUFACTURER CONCLUDES THE DEVICE HAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED IN THE BASE UNIT. IN THIS REPORT, | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03988 | 2021-09-15 | 2021-09-07 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP ASTHMA. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE REPORTED EVENT IS CURRENTLY UNKNOWN. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP ASTHMA. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE REPORTED EVENT IS CURRENTLY UNKNOWN. SECTION G3 WAS CORRECTED TO REFLECT THE BAD WHEN NEW INFORMATION WAS RECEIVED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04036 | 2021-09-15 | 2021-08-20 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP, AR | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | ARX200S15 | ARX200S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LIVER AND DIGESTIVE DISEASE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT WAS ALLEGED TO DEVELOP LIVER AND DIGESTIVE DISEASE.THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. DURING THE INVESTIGATION, OBSERVED AN UNKNOWN DUST CONTAMINANT ON THE FLIP DOORS, PCA, TOP ENCLOSURE, REAR PANEL, UI PANEL, BOTTOM ENCLOSURE, INSIDE ISO PORT, BLOWER, AND BLOWER BOX. OBSERVED BLACK HAIRLIKE CONTAMINANT ON THE FLIP DOORS, UI PANEL, AND BOTTOM ENCLOSURE. A KERATIN-LIKE SUBSTANCE WAS OBSERVED ON THE BLOWER BOX OUTLET. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND ERORRS. THE MANUFACTURER CONFIRMED NO PRESENCE OF DEGRADED SOUND ABATEMENT FOAM. IN THIS REPORT, SECTION D9, G3, H3, H6 HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04030 | 2021-09-15 | 2021-07-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP PRO BIFLEX | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DS650HS | DS650HS |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04057 | 2021-09-15 | 2021-08-04 | D, e, a, t, h | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUMID/HT, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | CAX700T12 | CAX700T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP CANCER. THE PATIENT HAS SINCE DIED. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO DEVELOP CANCER. THE PATIENT HAS SINCE DIED. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION OF DEVICE WERE COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND DUST/DIRT CONTAMINATION INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM THROUGHOUT DEVICE ENCLOSURE, AND AIR PATH, SUGGESTING A SOURCE EXTERNAL TO THE DEVICE. THE MANUFACTURER ALSO FOUND EVIDENCE OF WATER INGRESS ON THE BLOWER AND BLOWER BOX, STICKY CONTAMINATE ON BLOWER ISOLATORS. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERROR LOGGED. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE MANUFACTURER ALSO CONFIRMED THE PRESENCE OF CONTAMINATION IN THE AIR PATH. SECTION D9, G3, H6 HAS BEEN UPDATED / CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03946 | 2021-09-15 | 2021-07-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP PNEUMONIA. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE /RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM THAT BECAME DEGRADED AND CAUSED THE PATIENT TO HOSPITALIZE DUE TO PNEUMINIA. THE CUSTOMER ALSO ALLEGED NOT ABLE TO SLEEP AND DIFFICULTY BREATHIG. THESE EVENTS ARE ASSESSED SERIOUS BUT AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. AFTER THE FIRST ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WAS UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HEALTH EFFECT- CLINICAL CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03999 | 2021-09-15 | 2021-08-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX400T12C | CAX400T12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NUMEROUS BLACK FOAM PARTICLES WITHIN DEVICE. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING NUMEROUS BLACK FOAM PARTICLES WITHIN DEVICE. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED. IN THE INITIAL REPORT, SECTION A1, D4, G4 WAS MISSED TO CAPTURE, CORRECTED IN THIS REPORT. IN THE INITIAL REPORT, SECTION D1 WAS CAPTURED WRONGLY, CORRECTED IN THIS REPORT. SECTION H6 WERE UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04006 | 2021-09-15 | 2021-07-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SCAR TISSUE IN THEIR LUNG. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SCAR TISSUE IN THEIR LUNG. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. AS PER THE INVESTIGATION OF THE DEVICE EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED IN THE BASE UNIT. AS PER SECONDARY FINDINGS OF THE BASE UNIT OBSERVED UNKNOWN DUST/DIRT CONTAMINATION FOUND ON THE BLOWER CASING/IMPELLER AND BLOWER BOX WAS INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM CONTAMINATION. MINERAL SPOTS CONSISTENT WITH WATER INGRESS WERE FOUND ON THE MOTOR CASING. WATER INGRESS HAS BEEN PREVIOUSLY ADDRESSED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERORRS. THE MANUFACTURER CONCLUDES UNABLE TO ADDRESS THE SYMPTOMS DESCRIBED AND CONFIRM THE PRESENCE OF CONTAMINATION IN THE AIRPATH. IN THIS REPORT, SECTION D9, G3, H3, H6 HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04038 | 2021-09-15 | 2021-08-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION EXPERT W/HUM/HT, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | CAX501T12 | CAX501T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04005 | 2021-09-15 | 2021-07-15 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A SPOT ON THEIR LUNG. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03945 | 2021-09-15 | 2021-07-19 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP KIDNEY ISSUES. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING TO DEVELOP KIDNEY ISSUES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE.THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03944 | 2021-09-15 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP RESPIRATORY ISSUES. THE PATIENT WAS PRESCRIBED STEROIDS IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO EXPERIENCE UPPER RESPIRATORY ISSUES FOR WHICH THEY RECEIVED A STEROID INJECTION, ANTIBIOTICS, ORAL STEROIDS, AND NEBULIZER TREATMENTS. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENT OF UPPER RESPIRATORY ISSUES AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04055 | 2021-09-15 | 2021-08-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S TOUCHSCREEN LOCKED AND THE PARAMETERS WERE UNABLE TO BE CHANGED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR'S TOUCHSCREEN WAS LOCKED AND THE PARAMETERS WERE UNABLE TO BE CHANGED. THE VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03950 | 2021-09-15 | 2021-07-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG INFLAMMATION. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE /RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM THAT BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE LUNG INFLAMMATION, COUGH, HEADACHE AND SHE "RECEIVED" TWO SHOTS IN ARM FROM USING CPAP. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03948 | 2021-09-15 | 2021-07-19 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP NODULES IN THEIR LUNGS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED TO DEVELOP NODES IN THE LUNGS, LACK OF SLEEP, SHORT OF BREATHE, CHEST BURNS. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. PIL/RIQL FINDS UNKNOWN WHITE AND BLACK CONTAMINATION (DUST-LIKE), INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM, AT THE AIR INPUT OF THE BLOWER BOX IN THE RETURNED UNIT, SLIGHT DUST-LIKE CONTAMINATION ON TOP OF THE BLOWER MOTOR, SLIGHT BLACK CONTAMINATION CONSISTENT WITH KERATIN AT THE AIR OUTLET OF THE BLOWER BOX, TINY UNKNOWN WHITE SPECS OF CONTAMINATION ON THE BOTTOM THE BLOWER BOX. ALSO, A FEW TINY PIECES OF UNKNOWN BLACK (INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM) AND WHITE CONTAMINATION ON THE BOTTOM OF THE BLOWER BOX. THERE IS NO VISIBLE DAMAGE OR FUNCTIONALITY FAILURES OF THE DEVICE, WHICH SUGGEST THAT THE SOURCE OF CONTAMINATION WAS EXTERNAL TO THE DEVICE. THE DEVICE WAS HOOKED UP TO A KNOWN GOOD POWER SUPPLY AND POWER CORD, AIRFLOW WAS VERIFIED TO BE OPERATING CORRECTLY. THE ERROR LOG SHOWED ONE INSTANCE OF ERRORS ON THE DEVICE. THE MANUFACTURER CONCLUDES THAT THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION H6 HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03949 | 2021-09-15 | 2021-07-19 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP THYROID ISSUES. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP THYROID ISSUES. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE EVENT IS CURRENTLY UNKNOWN. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04037 | 2021-09-15 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP/HUM/HT/CELL,CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | CAX700T12C | CAX700T12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN THE INITIAL REPORT, SECTION H10 WAS MARKED INCORRECTLY AND B5 WAS INCORRECT, THE CORRECT B5 SHOULD BE- THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED SEEING PARTICLES IN THE DEVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL AND INTERNAL VISUAL INSPECTION OF DEVICE WERE COMPLETED BY THE MANUFACTURER AND FOUND EVIDENCE OF DUST/DIRT CONTAMINATION INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM THROUGHOUT THE DEVICE ENCLOSURE, AND AIRPATH, SLIGHT BLACK CONTAMINATION AT THE BLOWER SEAL OF THE BLOWER BOX, KERATIN CONTAMINATION. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERROR LOGGED. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH BLACK CONTAMINATION AT THE BLOWER SEAL OF THE BLOWER BOX, KERATIN CONTAMINATION AND DUST/DIRT CONTAMINATION WAS INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03947 | 2021-09-15 | 2021-07-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER AND HEADACHE. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURER CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAVE RECEIVED AT THAT TIME. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. PIL PERFORMED EXTERNAL VISUAL INSPECTION OF DREAMSTATION AND OBSERVED SMALL AMOUNTS OF CONTAMINATION ON THE EXTERIOR OF THE DEVICE. DURING INTERNAL INSPECTION OF THE DREAMSTATION, SOME CONTAMINATION AND DUST/DIRT INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM WAS OBSERVED INSIDE THE BLOWER BOX. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS OBSERVED IN THE BASE UNIT. THE BLOWER ALSO SHOWED SIGNS OF DUST CONTAMINATION. THE DEVICE WAS HOOKED UP TO A KNOWN GOOD POWER SUPPLY AND POWER CORD, AIRFLOW WAS VERIFIED TO BE OPERATING CORRECTLY. THE ERROR LOG SHOWED THAT THERE WERE NO INSTANCE OF ERRORS ON THE DEVICE. PIL WAS ABLE TO CONFIRM THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM IN THE DEVICE. EVIDENCE OF DUST CONTAMINATION ALSO OBSERVED. SECTION D8, D9, G3 HAS BEEN UPDATED. SECTION H6 HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04066 | 2021-09-16 | 2021-08-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE VENTILATOR WAS ALARMING. THE DEVICE'S SYSTEM BOARD, PROPORTIONAL VALVE AND SOLENOID VALVE WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04065 | 2021-09-16 | 2021-08-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1032800 | 1032800 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR HAD A LEAK ISSUE. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
MW5104039 | 2021-09-16 |  | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION AUTO CPAP W/HUMIDIFER/HEATER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500T11C |  |  | USING A PHILIPS DREAMSTATION CPAP WITH HUMIDIFIER FOR OVER 2 YEARS; BECAME SICK WITH COVID IN (B)(6) 2021 AND POST COVID HAD LINGERING PROBLEMS WITH BREATHING AND LUNG AND KIDNEY SPOTS DISCOVERED ON CT TEST. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5104041 | 2021-09-16 |  | I, n, j, u, r, y | Voluntary report | PHLLIPS DREAM STATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DREAMSTNAUTOCPAPHUMCELL DOM | DSX500H11C |  | I RECEIVED A RECALL OF MY PHILIPS RESPIRONICS CPAP MACHINE THAT WHEN INSULATION MATERIAL DEGRADES AND INFILTRATES THE AIR DUCTS CAN GIVE OFF TOXIC GAS WHICH IS FATAL. MOST ALL SYMPTOMS THAT WERE LISTED I HAVE BEEN EXPERIENCING. PROBLEM IS THEY PROVIDED A WEB ADDRESS TO REPORT YOUR DEVICE MODEL AND SN BUT THE ADDRESS COMES BACK 'DOESN'T EXIST'. I'VE TRIED SEVERAL TIMES WITH SAME NEGATIVE RESULTS. IT'S IMPORTANT THAT I USE THIS DEVICE WHEN SLEEPING BUT THE RECALL SAID TO IMMEDIATELY STOP FURTHER USE AND CONTACT THEM BUT THE ADDRESS DOESN'T EXIST. THE SYMPTOMS I HAVE GET WORSE BY THE DAY AND I DON'T KNOW WHAT TO DO. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-09
MW5104049 | 2021-09-16 | 2021-06-18 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I HAVE THE PHILIPS DREAMSTATION CPAP. SINCE BEFORE THE RECALL WAS SENT, I HAVE BEEN EXPERIENCING HEADACHES, SINUS PRESSURE/CONGESTION, CHEST PAIN, SHORTNESS OF BREATH, FATIGUE, COUGH, AND A CHANGE IN MY VOICE. SINCE THE RECALL, IT HAS STILL BEEN GETTING WORSE. TO THE POINT THAT I STRUGGLE TO GET WORDS OUT. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5104051 | 2021-09-16 | 2021-01-01 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DREAMSTATION |  |  | EXTREME TIREDNESS, WEAKNESS FOR PAST 6 MONTHS. PRODUCT RECALL USERS NOT NOTIFIED. |  | Y | Y | I |  | 2021-09
2518422-2021-04061 | 2021-09-16 | 2021-09-14 | D, e, a, t, h | Manufacturer report | BIPAP AVAPS C SERIES , CHINA, FLOW GEN | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS INC | 1167P | 1167P |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP FUNGAL INFECTION. THE PATIENT HAS REPORTEDLY HOSPITALIZED IN RESPONSE TO THE EVENT AND HAS SINCE DIED. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04048 | 2021-09-16 | 2021-07-05 | I, n, j, u, r, y | Manufacturer report | BIPAP AUTO SERIES ASSY, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | 750P | 750P |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED CHEST PAIN, DIFFICULTY BREATHING AND SHORTNESS OF BREATH. THE PATIENT DID RECEIVE MEDICAL INTERVENTION THAT REQUIRED HOSPITALIZATION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED CHEST PAIN, DIFFICULTY BREATHING AND SHORTNESS OF BREATH. THE PATIENT DID RECEIVE MEDICAL INTERVENTION THAT REQUIRED HOSPITALIZATION. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03952 | 2021-09-16 | 2021-06-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTOSV ADV W/HUM,SYS ONE,60SRS,DOM | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | DS960HS | DS960HS |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103020) ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP EYE IRRITATION. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON 11 NOV, 2024 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103020) ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP EYE IRRITATION. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECTS: CLINICAL CODES AND HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04064 | 2021-09-16 | 2021-08-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | EU2100X15B | EU2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S PARAMETERS COULD NOT BE ADJUSTED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04041 | 2021-09-16 | 2021-06-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A (CPAP, BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING HOARSENESS IN VOICE AND JAW PAIN NEAR LEFT EAR,HAD BLACK PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FOLLOW UP REPORT WILL BE FILED. CORRECTED INFORMATION WAS PROVIDED IN SECTION D 10 G1 AND H6. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03871 | 2021-09-16 | 2021-08-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04028 | 2021-09-16 | 2021-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04026 | 2021-09-16 | 2021-07-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. ADDITIONAL INFORMATION HAS BEEN ADDED AND THE B5 SHOULD BE REPORTED AS : THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING BREAST CANCER, HEADCHE, RESPIRATORY PROBLEM, WATERY EYES ON DAILY BASIS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT WAS HOSPITALIZED AND PERFORMED A BIOPSY FOR THE REPORTED EVENT. THE REPORTED EVENT THAT THE PATIENT HAS CANCER AND RESPIRATORY PROBLEMS IS ASSESSED AS SERIOUS INJURY BUT NOT RELATED TO THE DEVICE AND THE REPORTED PATIENT HAS WATERY EYES ON A DAILY BASIS, USES NOSE DROPS, AND HAS HEADACHES IS ASSESSED AS NON-SERIOUS INJURY. HENCE, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO THE ALLEGED SERIOUS INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HEALTH EFFECT - CLINICAL CODE, IMPACT CODE,TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS HAS BEEN UPDATED IN THIS REPORT. SECTION B1, B2, H1 CORRECTED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03881 | 2021-09-16 | 2021-09-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03875 | 2021-09-16 | 2021-08-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03880 | 2021-09-16 | 2021-09-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03873 | 2021-09-16 | 2021-08-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03879 | 2021-09-16 | 2021-08-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03874 | 2021-09-16 | 2021-08-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03878 | 2021-09-16 | 2021-08-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03872 | 2021-09-16 | 2021-08-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03876 | 2021-09-16 | 2021-08-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03877 | 2021-09-16 | 2021-08-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04029 | 2021-09-16 | 2021-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN PREVIOUS REPORT SECTION B5 WAS MENTIONED INCORRECTLY, CORRECT B5 SHOULD BE - THE PATIENT ALLEGED VISUALIZATION OF PARTICLES, CONGESTED CHEST. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL AND INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND BROWN AND DUST-LIKE CONTAMINATION AND TINY HAIRS ALL OVER THE EXTERIOR OF THE ENCLOSURE, WHITE AND BLACK UNKNOWN DUST LIKE CONTAMINATION AND VERY SMALL POTENTIAL HAIRS AT THE AIR INPUT OF THE BLOWER BOX, BLACK CONTAMINATION WITH KERATIN AT THE AIR OUTLET OF THE BLOWER BOX, WHITE DUST CONTAMINATION ON THE TOP OF THE BLOWER MOTOR. ALSO FOUND POTENTIAL MINERAL DEPOSITS ON THE BOTTOM OF THE BLOWER MOTOR, INDICATING POTENTIAL WATER INGRESS. THE MANUFACTURER RECEIVED HUMIDIFIER AND FOUND BROWN AND DUST-LIKE CONTAMINATION AND TINY HAIRS ALL OVER THE EXTERIOR OF THE ENCLOSURE AND ON THE HEATER PLATE, UNKNOWN BROWN AND BLACK CONTAMINANTS AND DARK-COLORED HAIR STRANDS INSIDE OF HUMIDIFIER BOX AND UNKNOWN WHITE CONTAMINATION ON THE AIR INLET TUBE SEAL, UNKNOWN BROWN CONTAMINATION IN THE ISO PORT SEAL AREA, WATER TANK HAD BROWN, BLACK AND MINERAL DEPOSITS ALL THROUGHOUT THE TANK. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS. THE DEVICE WAS APPLIED POWER AND THE DEVICE OPERATED PROPERLY. THE MANUFACTURER CONCLUDES MULTIPLE CONTAMINATION FOUND WERE CONSISTENT WITH CONTAMINATION OF DUST, KERATIN, HAIRS IN THE DEVICE, MINERAL DEPOSITS CONSISTENT WITH WATER INGRESS. ALSO CONCLUDES MULTIPLE CONTAMINATION FOUND WERE CONSISTENT WITH CONTAMINATION OF DUST, HAIRS, MINERAL DEPOSIT IN THE HUMIDIFIER. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. IN THIS REPORT, SECTION D9, D10, G3, H3, H6 HAS BEEN UPDATED OR CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03869 | 2021-09-16 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING NASAL/THROAT IRRITATION OR SORENESS, BLACK FLAKES COMING OUT OF MACHINE AND SHORT OF BREATH RELATED TO A CPAP DEVICES SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED AND THE MANUFACTURER OBSERVED WATER INGRESS AND RUST CONTAMINATION IN P4 POWER CONNECTOR. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND ONE ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION, BUT WATER INGRESS AND RUST CONTAMINATION WERE OBSERVED. SECTION D8 AND D9 HAS BEEN UPDATED IN THIS REPORT. IN SECTION H6 HEALTH EFFECT - CLINICAL CODE, MEDICAL DEVICE PROBLEM CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS HAS BEEN CORRETCED / UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04095 | 2021-09-16 | 2021-07-07 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC. | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED DIFFICULTY BREATHING AND SHORTNESS OF BREATH. THE PATIENT DID RECEIVE MEDICAL INTERVENTION WHICH RESULTED IN HOSPITALIZATION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP,AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DIFFICULTY BREATHING AND SHORTNESS OF BREATH.THE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM.THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL.THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION.THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME.IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE,AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03955 | 2021-09-16 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS:THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUNDABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT DEVELOP ASTHMA. THE REPORTED EVENT DEVELOP ASTHMA AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. PIL OBSERVED THE FOLLOWING FROM AN EXTERNAL AND INTERNAL INSPECTION: UNKNOWN WHITE AND BLACK CONTAMINATION (DUST-LIKE), INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM, AT THE AIR INPUT OF THE BLOWER BOX. SLIGHT DUST-LIKE CONTAMINATION ON TOP OF THE BLOWER MOTOR AND THE BLOWER MOTOR SEAL. UNKNOWN FILM OF CONTAMINATION ON THE TOP AND BOTTOM OF THE BLOWER MOTOR. TINY UNKNOWN BLACK (INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM) AND WHITE SPECS OF CONTAMINATION ON THE BOTTOM THE BLOWER BOX. PIL USED A FIT (FOAM INTEGRITY TEST TOOL) AND THE ASSOCIATED PROCEDURE ER (B)(4) V01 AND WAS NOT ABLE TO CONFIRM THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM. THERE IS NO VISIBLE DAMAGE OR FUNCTIONALITY FAILURES OF THE DEVICE, WHICH SUGGEST THAT THE SOURCE OF CONTAMINATION WAS EXTERNAL TO THE DEVICE. PIL WAS NOT ABLE TO CONFIRM THE COMPLAINT OR ADDRESS THE SYMPTOMS SPECIFIED. PIL WAS NOT ABLE TO GET CARE ORCHESTRATOR INFORMATION FROM DEVICE. REVIEW OF THE DEVICE LOG SHOWED: -ERRORS LOGGED: 0 ERRORS WERE LOGGED. -DEVICE WAS LIKELY RESET PRIOR TO PIL RECEIPT AS INDICATED BY THE DEVICE HAVING 1100.8 MACHINE HOURS AND NO BLOWER OR THERAPY HOURS. RISK TAG (ER (B)(4) V20) ASSOCIATED WITH THIS MODE OF FAILURE: IRRIT02, INFLECT. THE RESULTS OF THIS INVESTIGATION DO NOT IMPACT THE CALCULATED RISKS. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03957 | 2021-09-16 | 2021-07-22 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX700H11 | DSX700H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP THROAT NODULES. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP THROAT NODULES. A CORRECTION TO THE B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT DEVELOPED NODULES IN THROAT, DIFFICULTY OF BREATHING/SHORTNESS OF BREATH, RESPIRATORY ISSUES, LUNG AND THROAT IRRITATION, CHEST IS BURNING, EYES WERE RED, COUGH, HEADACHES, CAN'T WALK, BAD OXYGEN LEVELS AND LIGHTHEADEDNESS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. BASED ON THE AVAILABLE INFORMATION AT THE TIME OF THE REVIEW, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENTS. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03956 | 2021-09-16 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SPOTS ON THEIR LUNGS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON OCT 23, 2024. SECTION H11 SHOULD BE REPORTED AS THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SPOTS ON THEIR LUNGS AND SHORTNESS OF BREATH. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THERE IS NO PATIENT HARM OR INJURY. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM (ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR). SECTION B2 WAS CHANGED TO BLANK (PREVIOUSLY IT WAS HOSPITALIZATION). SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04081 | 2021-09-16 | 2021-08-23 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION BIPAP AUTOSV H/HT/C, DS | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS INC | DSX900T11C | DSX900T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LEUKEMIA. THE PATIENT IS UNDERGOING CHEMOTHERAPY IN RESPONSE TO THE REPORTED EVENT. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LEUKEMIA. THE PATIENT IS UNDERGOING CHEMOTHERAPY IN RESPONSE TO THE REPORTED EVENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL AND INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO EVIDENCE OF CONTAMINATION. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS LOGGED. THE DEVICE WAS APPLIED POWER AND THE DEVICE OPERATED PROPERLY. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. SECTION D9, G3, H3 AND H6 HAS BEEN UPDATED / CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04077 | 2021-09-16 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM AND BECAME DEGRADED AND THE PATIENT ALLEGED OF PARTICLES IN TUBING/CHAMBER, NASAL/THROAT IRRITATION OR SORENESS, SHORT OF BREATH, NAUSEAS, HEADACHES, WATERY EYES, BLACK PARTICLES IN DEVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. EXTERNAL EXAMINATION OBSERVED DUST OR DIRT CONTAMINATION ON THE OUTER SURFACES, CRACKS, CREVICES, AROUND THE OUTLET PORT AND MODEM AREA. BEIGE CONTAMINATION WAS LOCATED AROUND THE AIR INLET CANAL. INTERNAL INVESTIGATION WAS PERFORMED AND LOCATED A LIGHT GREY FILM OF CONTAMINATION ON THE INSIDE SURFACE OF THE FRONT PANEL AND SMALL AMOUNT OF DARK CONTAMINANT AROUND WHERE THE BLOWER OUTLET SEAL WOULD SIT ON THE BLOWER BOX. BEIGE CONTAMINANT ON THE INSIDE SURFACE OF THE BLOWER BOX BELOW THE BLOWER WAS ALSO OBSERVED. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN, AND OTHER CONTAMINATIONS FOUND TO BE SOURCED FROM EXTERNAL ENVIRONMENTAL CONDITIONS. SECTION H6 MEDICAL DEVICE CODE, COMPONENT CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04082 | 2021-09-16 | 2020-01-14 | D, e, a, t, h | Manufacturer report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC |  |  |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103332) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT DEATH. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED A VOLUNTARY MEDWATCH (B)(4) REPORTED A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND ALLEGEDLY CAUSED A PATIENT DEATH. DESPITE MULTIPLE ATTEMPTS, THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04094 | 2021-09-16 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION EXPERT W/HUM/CELL, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX501H12C | CAX501H12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED BLACK SPOTS ON DEVICE. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE 1 ERROR FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED DUE TO AGE. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04073 | 2021-09-16 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION G1 WAS CORRECTED AND H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04049 | 2021-09-16 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED RESPIRATORY ISSUES AND MUCOUS . THE PATIENT DID RECEIVE MEDICAL INTERVENTION REQUIRING MEDICATION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED ON THIS DEVICE IN MDR 2518422-2021-03882. THIS REPORT WAS SUBMITTED AS A DUPLICATE REPORT OF THE PREVIOUSLY SUBMITTED REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04047 | 2021-09-16 | 2021-06-28 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED FLUID IN HIS LEFT LUNG. THE PATIENT DID RECEIVE MEDICAL INTERVENTION WITH PRESCRIBED MEDICATIONS. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03953 | 2021-09-16 | 2021-08-23 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103088) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP MYELOPROLIFERATIVE NEOPLASM. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED AND CAUSED A PATIENT TO DEVELOP MYELOPROLIFERATIVE NEOPLASM. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03870 | 2021-09-16 | 2021-08-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03954 | 2021-09-16 | 2020-05-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102964) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA ATTACKS. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY RECEIVED A VOLUNTARY MEDWATCH (MW5102964) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA ATTACKS. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. IN THE PREVIOUSLY SUBMITTED REPORT SECTION B5 WAS INCOMPLETE, CORRECT SECTION B5 IS- THE PATIENT ALSO ALLEGED VISUALIZATION OF PARTICLES FROM THE DEVICE AND HAVING NOSE SORENESS & IRRITATION, NASAL DRIP, COUGH. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF AN UNKNOWN DUST CONTAMINATION WAS OBSERVED INSIDE THE UI PANEL DISPLAY SCREEN. AN INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF AN UNKNOWN DUST CONTAMINATION WAS OBSERVED ON THE BLOWER CAP, ON THE BLOWER, ON THE BLOWER OUTPUT SEAL, BLOWER ISOLATOR, ISO PORT O-RING, AND IN THE BLOWER BOX, DARK CONTAMINATION CONSISTENT WITH THE KERATIN CONTAMINATION WAS AROUND THE BLOWER SEAL IN THE BLOWER BOX. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. THE DEVICE WAS HOOKED UP TO POWER SUPPLY, AIRFLOW WAS VERIFIED AND THE DEVICE OPERATED PROPERLY. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND 2 INSTANCES OF E-4 AND E-200. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF MULTIPLE CONTAMINANTS. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. SECTION D9, G3 AND H6 HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03951 | 2021-09-16 | 2021-07-22 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103020) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP NODULES AND SCARRING OF THEIR LUNGS. THE PATIENT WAS PRESCRIBED PRESCRIPTION MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04033 | 2021-09-16 | 2021-08-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200H11C | DSX200H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DRY MOUTH,CHEST TIGHTNESS,HEADACHE,DIZZINESS,BLACK PARTICLES THROUGH OUT HOSE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HEALTH EFFECT- CLINICAL CODE WAS MISSED TO CAPTURE IN PREVIOUS MDR,NOW IT IS CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04074 | 2021-09-16 | 2021-08-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING MASK TURNING BROWN, BLACK PARTICLES IN FILTER RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04109 | 2021-09-16 | 2021-08-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | INX700S15 | INX700S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04108 | 2021-09-16 | 2021-08-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX700T11 | DSX700T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A (CPAP, BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04051 | 2021-09-16 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP A CYST ON THE RIGHT KIDNEY. THE PATIENT DID NOT REPORT RECEIVING MEDICAL INTERVENTION THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO DEVELOP A CYST ON THE RIGHT KIDNEY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENT OF CYST ON THE RIGHT KIDNEY AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04050 | 2021-09-16 | 2021-07-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE HEADACHES AND WAS DIAGNOSED WITH A-SID. THE REPORTED INFORMATION STATES THE PATIENT UNDERWENT A CARDIO-VERSION PROCEDURE TWICE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO EXPERIENCE A HEADACHES AND WAS DIAGNOSED WITH A-SID. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE CUSTOMER. THE REPORTED EVENT OF CARDIOVERSION PROCEDURE TWICE AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5104065 | 2021-09-17 | 2019-01-31 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | RESPIRONICS PHILIPS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I HAVE A PHILIPS CPAP DREAMSTATION AND FROM THE FIRST USE I SMELLED A CHEMICAL SMELL THAT THE MACHINE WAS OMITTING. BOTH THE DEVICE SUPPLIER AND THE PULMONARY PRACTICE DID NOT ADDRESS MY CONCERNS. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04060 | 2021-09-17 | 2021-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04136 | 2021-09-17 | 2021-07-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURYTHIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING TO HAVE HEADACHE,NAUSEA,CHEST PAIN,LIGHT HEADINESS,DIZZYNESS,THROAT IRRITATION AND PARTICLES IN TUBING/CHAMBER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE EVENT IS CURRENTLY UNKNOWN. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HEALTH EFFECT-CLINICAL CODE WAS MISSED TO CAPTURE IN PREVIOUS MDR,NOW IT IS CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04132 | 2021-09-17 | 2021-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX400S12 | CAX400S12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED BLACK PARTICLES IN THE FILTER AFTER USE. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUNDN AND SCRAPPED. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK PARTICLES IN THE FILTER AFTER USE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION H6 WERE UPDATED IN THIS REPORT. PATIENT USED A FILTER AND BLACK PARTICLES IN THE FILTER AFTER USE. WOULD LIKE SOMEONE TO RETURN CALL AS SOON AS POSSIBLE . | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04123 | 2021-09-17 | 2021-08-26 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, CN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | CNX500T17 | CNX500T17 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP PULMONARY NODULES. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04111 | 2021-09-17 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING SEEING WHITE FLAKES IN TUBING AND WATER CHAMBER RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION D1,D4,H7 WAS CORRECTED AND SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04119 | 2021-09-17 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. SECTION B5 SHOULD BE PREVIOUSLY REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DEVICE IS NOISY AND CLIENT HAD PARTICLE IN THE "TUBING" WHEN SHE RINSE IT OUT RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04118 | 2021-09-17 | 2021-08-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04075 | 2021-09-17 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | ADDITIONAL INFORMATION WAS RECEIVED ON SEPT 28, 2023 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING NOTICED BLACK PARTICLES IN AIR DEVICE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURE INSPECTED EXTERNALLY AND INTERNALLY, OBSERVED MINOR DUST/DIRT CONTAMINATION ON TOP AND BOTTOM ENCLOSURES, KEYPAD, CONTROL DIAL, SD COVER FLIP DOOR, UI PANEL, REAR PANEL, BLOWER, BLOWER BOX, P4 POWER CONNECTOR, DC JACK COLOR INSERT. HAIRLIKE SUBSTANCE ON TOP ENCLOSURE AND CONTROL DIAL. DARK UNKNOWN CONTAMINATE ON INSIDE OF THE UI PANEL. POSSIBLE LIQUID INGRESS WAS OBSERVED ON THE P4 POWER CONNECTOR. A CONTAMINATE CONSISTENT WITH KERATIN WAS OBSERVED ON THE BLOWER BOX. THE MANUFACTURER USED A FITT (FOAM INTEGRITY TEST TOOL) AND WAS NOT ABLE TO CONFIRM THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER WAS UNABLE TO DIRECTLY ADDRESS THE SYMPTOMS OUTLINED IN COMPLAINT. THEY DID OBSERVE DUST/ DIRT CONTAMINATION IN THE LIKELY FROM A SOURCE EXTERNAL TO THE DEVICE. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04117 | 2021-09-17 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING/CHAMBER, PATIENT IS HAVING LUNG ISSUES AND HAS BEEN PLACED ON OXYGEN AND PATIENT HAS SEEN RUBBER IN CHAMBER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL.THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION G1 CORRECTED, H6 CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04110 | 2021-09-17 | 2021-07-15 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11C | DSX400H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CHEST PAINS. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CHEST PAINS. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. PATIENT ALSO ALLEGED EXPERIENCING COUGHING, INFLAMMATION OF LUNGS, PARTICLES IN FILTER, GREASY WATER, HEADACHES, LIGHT HEADEDNESS, DIZZINESS, SORENESS OF THROAT, RUNNING NOSE, RUNNING EYES, WAKING UP SCREAMING IN PAIN, ODOR FROM DEVICE, AND HAS BEEN IN THE HOSPITAL A LOT SINCE USAGE OF DEVICE. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. DURING THE EXTERIOR INVESTIGATION, OBSERVED UNKNOWN DUST/DIRT CONTAMINATION PRESENT ON ALL SURFACES OF THE BASE UNIT. FOUND AN UNKNOWN DUST CONTAMINATE PRESENT AT THE AIR INLET WHERE THE FILTER WOULD BE AND AT THE ISO PORT ENTRANCE. DURING THE INTERNAL INVESTIGATION OF THE BASE UNIT, OBSERVED UNKNOWN DEBRIS IN THE TRACKS OF THE UI PANEL. THERE IS AN UNKNOWN WHITE DUST CONTAMINATION FOUND ON THE BOTTOM ENCLOSURE, BLOWER BOX HOUSING, BLOWER MOTOR, BLOWER IMPELLER, AND REAR PANEL O-RING. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED IN THE BASE UNIT. SECONDARY FINDINGS: THE UNKNOWN DUST/DIRT CONTAMINATION AND FIBERS FOUND ON THE BLOWER CASING/IMPELLER AND BLOWER BOX WAS INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM CONTAMINATION. THE PRESENCE OF CONTAMINATION THROUGHOUT THE AIRPATH SUGGESTS A SOURCE EXTERNAL TO THE DEVICE. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS. THE MANUFACTURER CONCLUDES UNABLE TO DIRECTLY ADDRESS THE SYMPTOMS OUTLINED IN THE COMPLAINT. CONFIRMED THAT THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN OBSERVED IN THE BASE UNIT. CONFIRMED THE PRESENCE OF CONTAMINATION IN THE AIRPATH. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04083 | 2021-09-17 | 2021-06-11 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102966) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A SWOLLEN UVULA. THE PATIENT WAS PRESCRIBED STEROIDS IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04087 | 2021-09-17 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103087) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CORNEAL INFLAMMATION. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 14, 2024. SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED A VOLUNTARY MEDWATCH (MW5103087) ALLEGING BILATERAL CORNEAL INFLAMMATION, EYES WOULD GET IRRITATED, DRY OUT, BLACK PARTICULATES WOULD SPORADICALLY ACCUMULATE IN NOSTRILS; AND A SWEET PUNGENT ODOR WOULD OCCUR ESPECIALLY WHEN DISTILLED WATER IN THE HUMIDIFIER RESERVOIR GOT LOW AND DEVICE EXCESSIVELY NOISY RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM (ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR). SECTION B2 WAS CHANGED TO BLANK (PREVIOUSLY IT WAS OTHER SERIOUS). SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04105 | 2021-09-17 | 2020-04-18 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103198) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP PNEUMONIA. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON OCT 12, 2023, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED A VOLUNTARY MEDWATCH (MW5103198) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP PNEUMONIA. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. ADDITIONAL INFORMATION WAS RECEIVED ALLEGED OF RESPIRATORY TRACT IRRITATION OTHER AND PNEUMONIA, BRONCHITIS MUCUS PLUG, PLEURISY, RESPIRATORY FAILURE, CHEST PAIN, SEPSIS, LARGE ABSCESS OF RIGHT LUNG. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL.THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER LATER, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 HEALTH EFFECTS - CLINICAL CODES WAS UPDATED IN THE REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04042 | 2021-09-17 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. SECTION B5 SHOULD BE PREVIOUSLY REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES IN THE WATER LOOKS LIKE WORMY, DIRT RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT WAS ALLEGING HEADACHES, DIZZINESS, (LUNGS) HURTS TO BREATH, DRAINAGE, WATERY EYES, NAUSEA. IT WAS ALSO ALLEGED STRANGE ODOR. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION G1 UPDATED IN THIS REPORT. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04122 | 2021-09-17 | 2021-09-10 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, GER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DEX500S13 | DEX500S13 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CARCINOMA. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5104061 | 2021-09-17 | 2021-09-01 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | REMSTAR AUTO A-FLEX CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | REMSTAR SYSTEM ONE SERIES 50 | 550P (MAYBE) |  | DEAR FDA DEVICE RECALLS DIVISION: I AM WRITING AS A CONCERNED PATIENT IN POSSESSION OF A RECALLED RESPIRATORY CARE DEVICE THAT IS NOT CLEARLY LISTED ON PHILIPS WEB SITE AS RECALLED. I HAVE AN ALTERNATIVE DEVICE, SO MY HEALTH IS NOT AT RISK, BUT THERE MAY BE OTHERS WHO ARE MISINFORMED REGARDING WHETHER THEIR DEVICE HAS BEEN RECALLED. THERE ARE MULTIPLE PROBLEMS HERE: PHILIPS LISTS THE PRODUCTS BY VARIOUS NAMES, BUT AS A CONSUMER, IT IS SOMETIMES DIFFICULT TO ASSOCIATE A NAME WITH A SERIES. FOR EXAMPLE, ON PHILIPS' WEB SITE, THEY LIST THE SYSTEM ONE 50 AND 60 SERIES, BUT IN THE RECALL LETTER, THEY LIST THE SYSTEMONE Q-SERIES. ARE THE SYSTEM ONE 50 AND 60 SERIES THE SAME AS THE SYSTEMONE Q-SERIES? MAYBE, I DON'T KNOW. HOWEVER, THESE WORDS DO NOT APPEAR ON MOST OF THE UNITS. FOR EXAMPLE, I OWN A REMSTAR AUTO A-FLEX CPAP. THE WORDS SYSTEM, ONE, Q-SERIES, AND 50 SERIES DO NOT APPEAR ANYWHERE ON THE UNIT, NOR IN THE MANUAL, BUT THAT'S WHAT IT IS. A CONSUMER COULD EASILY CONCLUDE FROM THE WEB SITE OR THE RECALL LETTER THAT THEIR UNIT IS NOT RECALLED, BECAUSE THE INTERNAL CORPORATE DESIGNATIONS USED IN THE PATIENT-FACING MATERIALS ARE NOT HOW THE UNITS ARE ACTUALLY LABELED. FOR PATIENT-FACING INFORMATION, THE MODELS SHOULD BE LISTED BY THE NAMES THE PATIENT SEES ON THE UNIT, NOT BY THE MANUFACTURER'S INTERNAL PRODUCT-LINE JARGON. THE WEB SITE AND THE RECALL LETTER CONTAIN CONFLICTING INFORMATION. ON THE WEB SITE, THE REMSTAR SE IS LISTED AS NOT RECALLED (ALONG WITH THE M-SERIES REMSTAR AND SLEEP EASY CPAPS) ON THE WEB SITE. IN THE RECALL LETTER, THE REMSTAR SE AUTO IS LISTED AS RECALLED, BUT IS NOT MENTIONED ON THE WEB SITE. IS THIS A DIFFERENT MODEL BASED ON THE "AUTO" DESIGNATION? WHY IS THE REMSTAR SE AUTO SE NOT MENTIONED ON THE WEB SITE? HOW MANY OTHER RECALLED UNITS ARE OMITTED OR MIS-DESIGNATED ON THE PATIENT-FACING WEB SITE? 3. IN ADDITION, THE WEB SITE REGISTRATION IS CONFUSING. IF THE USER ENTERS THE WRONG SERIAL NUMBER (SUCH AS PUTTING A LETTER O WHERE IT SHOULD BE THE NUMBER 0), IT RESPONDS WITH A MESSAGE THAT THE SERIAL NUMBER IS NOT RECOGNIZED AS AN AFFECTED DEVICE, AND ADVISES CHECKING THE NUMBER AGAIN. WHY DOES IT SAY "NOT RECOGNIZED AS AN AFFECTED DEVICE," WHEN IT SHOULD SAY, "THIS IS A WRONG NUMBER, YOU DON'T HAVE AN ANSWER AT ALL." THEY HAVE ALREADY CHANGED THE MESSAGE (IT WAS WORSE BEFORE) BUT IT STILL SEEMS DESTINED TO CONFUSE USERS. PHILIPS COULD EASILY PROGRAM THEIR SYSTEM TO RECOGNIZE WHETHER THE NUMBER INPUTTED IS A LEGITIMATE NUMBER OR NOT, BUT INSTEAD THEY HAVE CHOSEN TO DELIVER A MESSAGE THAT MAY CAUSE AN UNWARY CONSUMER TO BELIEVE A RECALLED DEVICE IS NOT RECALLED. THIS IS A CLASS I RECALL, AND I BELIEVE PHILIPS OWES IT TO THEIR CUSTOMERS TO PROVIDE ACCURATE AND CONSISTENT INFORMATION IN A USABLE FORM. THEY HAVE FAILED TO DO SO, IN SEVERAL WAYS THAT WILL CAUSE PATIENTS TO BELIEVE THEIR RECALLED UNITS ARE NOT RECALLED. THEIR WEB SITE AND REGISTRATION SYSTEM SHOULD BE FIXED, AND THEY SHOULD SEND A NEW RECALL LETTER WITH INTELLIGIBLE INFORMATION REGARDING WHICH MODELS ARE ACTUALLY RECALLED. SINCERELY, (B)(6), MD, PHD. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | I |  | 2021-09
2518422-2021-04141 | 2021-09-17 | 2021-07-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX500T12 | CAX500T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE MALFORMATION IN LUNGS. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE EVENT IS CURRENTLY UNKNOWN. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICE. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED MALFORMATION IN LUNGS. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE EVENT IS CURRENTLY UNKNOWN. THIS NOTIFICATION WAS REVIEWED BY THE PMS CLINICAL EXPERT DUE TO THE PATIENT REPORTING HAS A MALFORMATION ON HER LUNGS AND IS GOING TO HAVE SURGERY TO REMOVE. PATIENT ALSO REPORTED SHE HAD DEVICE FIXED TWICE AND STILL DOES NOT WORK CORRECTLY. ATTEMPTS HAVE BEEN MADE TO OBTAIN ADDITIONAL INFORMATION AND HAVE BEEN UNSUCCESSFUL. THE EVENT IS ASSESSED AS NOT RELATED TO THE RECALLED DEVICE (I.E. A FOAM DEGRADATION CONCERN) IN THIS CASE. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION D4 WAS INCORRECTLY CAPTURED IN INITIAL REPORT AND WAS UPDATED IN THIS REPORT AND SECTION H6 WAS UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5104067 | 2021-09-17 |  | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | DEVELOPED PROSTATE CA. HAD PROSTATECTOMY ON GOING CARE SUBMITTED MACHINE TO MY MEDICAL PROVIDER THEY SUPPLIED ME WITH A LOANER BUT ADVISED ME NOT TO USE IT. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5104072 | 2021-09-17 | 2020-07-10 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS DREAMSTATION AUTO-CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS/ RESPIRONICS, INC. |  |  |  | PATIENT WAS USING A PHILIPS RESPIRONIC'S DREAMSTATION AUTO-CPAP FOR ABOUT 9 MONTHS WHEN HE BEGAN TO WAKE UP IN THE MIDDLE OF THE NIGHT WITH SEVERE HEADACHE AND SINUS BURNING. THIS CONTINUED WITH THE DEVICE TO THE POINT WHERE HE WAS NO LONGER ABLE TO TOLERATE USING IT. HEADACHES AND SINUS PAIN CONTINUE DESPITE DISCONTINUATION OF THE DEVICE FOR ABOUT A YEAR. FDA SAFETY REPORT ID # (B)(4). |  | Y | N | LAY USER/PATIENT |  | 2021-09
2518422-2021-04056 | 2021-09-17 | 2021-09-07 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO, FR | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | FRX500S14 | FRX500S14 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP CHOLANGIOCARCYNOMA. THE PATIENT RECEIVED CHEMOTHERAPY AND IMMUNOTHERAPY TREATMENT IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04007 | 2021-09-17 | 2021-09-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, AN ISSUE RELATED TO THE OXYGEN BLENDING MODULE BOARD WAS OBSERVED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04009 | 2021-09-17 | 2021-08-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04104 | 2021-09-17 | 2021-08-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILS TO CHARGE ITS INTERNAL BATTERY. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF A VENTILATOR FAILING TO CHARGE ITS INTERNAL BATTERY. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT COULD NOT BE DUPLICATED. THE DEVICE WAS FOUND TO OPERATE AND ALARM AS DESIGNED. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04131 | 2021-09-17 | 2021-08-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION THAT A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04010 | 2021-09-17 | 2021-08-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE'S LCD SCREEN WAS OBSERVED TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03688 | 2021-09-17 | 2021-08-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER BECAME AWARE OF AN ALLEGATION FROM A USER THAT A THERMAL EVENT RESULTING IN SECOND DEGREE BURNS OCCURRED DURING USAGE WITH THE HEATED TUBING ATTACHED TO THE DREAMSTATION AUTO BIPAP. USER ALLEGES BURNS ON THE CHEST, ARMS, BACK OF NECK, THROAT, WRISTS AND HANDS RESULTING IN HOSPITALIZATION FOR 4 DAYS. USER HAD RECEIVED THIS TUBING IN (B)(6) OF 2021. USER ALLEGES MACHINE WAS SMOKING WHERE THE HEATED TUBING PLUGS INTO THE DEVICE. THE USER WAS DISCHARGED HOME WITH A PRESCRIPTION FOR SILVADENE CREAM 400 MG AND PAIN PILLS FOR THE BURNS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. UPON COMPLETION OF THE INVESTIGATION, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER RECEIVED THE DEVICES WITH HEATED TUBING FOR EVALUATION. THE INVESTIGATION OF THE RECEIVED DEVICES COULD NOT CONFIRM THE COMPLAINT OF THE HEATED TUBING (HT15) OVERHEATING AND CAUSING BURNS TO THE USER. THE EXTERNAL AND INTERNAL INSPECTIONS REVEALED NO THERMAL DAMAGE, MELTING, OR SMOKE DAMAGE INSIDE OR ON THE EXTERIOR OF THE DEVICES.THE DEVICES PASSED ALL TESTING AND FUNCTIONED AS DESIGNED. THERE WAS NO OVERHEATING OF THE DEVICES OR OF THE HEATED TUBING. ALL TEMPERATURES DURING WERE WITHIN SPECIFICATIONS. BLACK CONTAMINATES WERE FOUND INSIDE AND OUTSIDE OF THE MACHINE AND ON THE DEVICE FILTER, CONSISTANT WITH DUST/DIRT. PATIENT DATA INDICATES THE DATE OF LAST USAGE WAS 8/14/2021. THE REPORT ALSO INDICATES THE HEATED TUBE TEMPERATURE WAS SET TO OFF , THE HUMIDIFIER HEAT PLATE WAS SET TO OFF AND HUMIDIFIER MODE WAS SET TO ADAPTIVE. THERE WAS NO ODOR OF SMOKE ON OR IN THE DEVICES OR HEATED TUBING. BASED ON THE INFORMATION AVAILABLE, THE MANUFACTURER CONCLUDES NO FURTHER ACTION IS NECESSARY. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04142 | 2021-09-17 | 2021-07-15 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE CONGESTION. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTION FOR ANTIHISTAMINE DRUGS. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE CONGESTION. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTION FOR ANTIHISTAMINE DRUGS. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE AND FOUND WHITE AND BLACK UNKNOWN CONTAMINATION (DUST LIKE), INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM, AND SOME VERY SMALL POTENTIAL HAIRS AT THE AIR INPUT OF THE BLOWER BOX. THE MANUFACTURER ALSO FUND UNKNOWN WHITE (DUST LIKE) CONTAMINANTS ON TOP OF BLOWER MOTOR. THE MANUFACTURER VISUALLY INSPECTED THE HUMIDIFIER BOX AND FOUND UNKNOWN CREAM COLOR, BLACK CONTAMINANTS, AND SOME SMALL BLACK HAIRS, INSIDE OF HUMIDIFIER BOX. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND ZERO ERROR LOGGED. THE MANUFACTURER CONCLUDES NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. THE MANUFACTURER ALSO CONFIRMED THE PRESENCE OF MULTIPLE CONTAMINANTS THAT ARE NOT CONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM. SECTION D9, G3, H6 HAS BEEN UPDATED / CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
12487430 | 2021-09-17 | 2021-09-01 | I, n, j, u, r, y | User Facility report | V60 BIPAP | VENTILATOR, CONTINUOUS, MINIMAL VENTILATORY SUPPORT,FACILITY USE |  | V60 |  |  | CODE BLUE CALLED. PATIENT (PT) HAD BEEN ON CONTINUOUS CPAP (CONTINUOUS POSITIVE AIR PRESSURE), VIA V60+ MACHINE. PT. BEING MANUALLY BAGGED UPON ARRIVAL, NO COMPRESSIONS IN PROGRESS, SETTING UP FOR INTUBATION. PT. INTUBATED BY RT (RESPIRATORY THERAPIST) @BEDSIDE, PLACED ON VENTILATOR. PER RN (REGISTERED NURSE) @BEDSIDE, CPAP MACHINE STOPPED VENTILATING, WAS ALARMING. MACHINE PULLED, WORK ORDER PLACED FOR CTS. CTS AWARE AND MACHINE TO BE TAKEN FOR REVIEW AND DIAGNOSTICS. |  | Y | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04129 | 2021-09-17 | 2021-09-14 | D, e, a, t, h | Manufacturer report | DREAMSTATION ST30 H/HT, CN-NB | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS INC | CNX1030T17 | CNX1030T17 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT DEATH. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5104094 | 2021-09-20 | 2021-09-15 | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP / BIPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I RECENTLY WAS DIAGNOSED WITH ASTHMA AND THINK IT'S POSSIBLE THAT IT HAS TO DO WITH MY CPAP MACHINE THAT IS BEING RECALLED (PHILIPS DREAMSTATION). I NEVER HAD ANY LUNG ISSUES BEFORE THIS PAST YEAR. I AM STILL USING THE MACHINE BECAUSE THEY DO NOT HAVE A REPLACEMENT FOR ME YET. I TRIED NOT TO SLEEP WITH IT AND IT CAUSES HEADACHES AND EXTREME FATIGUE WHEN I DON'T USE IT. I WAS PUT ON AN INHALER THIS WEEK. AGAIN I'M NOT SURE IF THIS IS WHAT CAUSED THE ISSUE BUT IT MIGHT BE. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | I |  | 2021-09
MW5104102 | 2021-09-20 | 2019-12-10 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500T11 |  |  | HAD SINUS PROBLEMS, SHORTNESS OF BREATH WHICH CONTINUES TO BE A PROBLEM. TOLD THE COMPANY (B)(6) ABOUT THE PROBLEMS AND NOTHING WAS DONE. C-PAP DEVICE IS WHAT CAUSED BOTH PROBLEMS. NO MEDICAL REASON FOR SHORTNESS OF BREATH AND NEVER HAD SINUS PROBLEMS BEFORE USING THE C-PAP. MANY TESTS, WE HAVE DONE EKG, CHEST X-RAY, STRESS TESTS, HEAD CT. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03846 | 2021-09-20 | 2021-08-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | 761P | 761P |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04204 | 2021-09-20 | 2021-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SYSTEM ONE 60 SERIES BIPAP AUTO BIFLEX | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DS760HS | DS760HS |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04153 | 2021-09-20 | 2021-08-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2200X11B | DS2200X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR HAD LOW TIDAL VOLUMES. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
MW5104087 | 2021-09-20 | 2021-09-16 | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | AUTO CPAP N |  |  | I USE A PHILIPS DREAMSTATION CPAP THAT HAS BEEN REGISTERED AND CONFIRMED TO BE PART OF RECALL. I HAD NOTICED FINE BLACK PARTICLES IN WATER RESERVOIR BEFORE RECALL WHICH I THOUGHT WERE MOLD. I STEPPED UP MY CLEANING PROCEDURES AND MADE SURE I USED THE SOCLEAN OZONE CLEANER EVERY DAY. WHICH ONLY MADE IT WORSE. WHEN I LEARNED ABOUT RECALL (THROUGH NEWSPAPER) IN AUGUST, SINCE I COULD NOT SLEEP WITHOUT THE CPAP, I ADDED A BIO FILTER AT DRS RECOMMENDATION AND STOPPED USING THE SOCLEAN. I HAVE BEEN WAKING UP WITH HEADACHES FREQUENTLY AND WHEN I CHECKED THE BIO FILTER LAST NIGHT TO CHANGE IT, IT WAS VERY GRAY COMPARED TO A NEW FILTER. I HAVE STOPPED USING THE CPAP. I HAVE CALLED PHILLIPS TWICE IN THE LAST SIX WEEKS TO CHECK ON PROGRESS OF RECALL AND I'VE NEVER HEARD ANYTHING FROM THEM. I INFORMED THEM THAT MY FOAM HAD BEGUN DECOMPOSING AND ASKED FOR THEM TO EXPEDITE A REWORK FOR ME. I AM CONCERNED ABOUT WHAT I MAY HAVE BREATHED INTO MY LUNGS IF THE BIO FILTER WAS NOT ADEQUATELY FILTERING THE PARTICLES. I WILL CONTACT MY DOCTOR RIGHT AWAY ABOUT THIS. WHEN I HAD A SLEEP STUDY 3 YEARS AGO, THE STUDY HAD TO BE DISCONTINUED BECAUSE MY OXYGEN LEVEL HAD FALLEN TO 73 FOR A PERIOD OF TIME. I HAD OVER 55 AHIS DURING REM PHASE AND WAS GIVEN A STAT ORDER FOR A CPAP MACHINE THAT DAY. I HAVE USED THE CPAP EVERY NIGHT ALL NIGHT SINCE THEN AND IT HAS BEEN A LIFE SAVER. WHEN I FOUND OUT ABOUT THE RECALL, I STOPPED USING IT FOR 2 NIGHTS AND JUST COULD NOT SLEEP. I HAD TO GO BACK TO USING IT BUT ADDED THE BIO FILTER, HOPING THAT WOULD STOP ANY PARTICLES. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04067 | 2021-09-20 | 2021-08-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S MACHINE FLOW SENSOR WAS REPLACED DUE TO DUST AND DIRT CONTAMINATION. | DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'SDOWNLOADED ERROR LOG. THE DEVICE'S MACHINE FLOW SENSOR WAS REPLACED DUE TO DUST AND DIRT CONTAMINATION. THE DEVICE'S SYSTEM BOARD WAS ALSO REPLACED. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
MW5104083 | 2021-09-20 | 2021-08-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DREAMSTATION |  |  | I'VE EXPERIENCED SEVERE HEADACHES AND SINUS PROBLEMS THE PAST COUPLE OF MONTHS. I BELIEVE THIS IS DUE TO THE USE OF A PHILIPS DREAM STATION WHICH HAS BEEN RECALLED. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5104082 | 2021-09-20 | 2021-02-07 | I, n, j, u, r, y | Voluntary report | DREAMSTAATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500H11C |  |  | AS I WAS GOING TO FILL MY WATER RESERVOIR ON MY PHILIPS DREAM STATION CPAP MACHINE. I FOUND NUMEROUS PARTICLES FLOATING IN THE WATER. I RECENTLY FOUND OUT THAT THERE IS A RECALL ON THE DREAM STATION MACHINE DUE TO ISSUES WITH THE INSULATION BREAKING DOWN AND ENTERING THE WATER CHAMBER AND HOSE. I HAD USED THE SO CLEAN DEVICE ON MY CPAP MACHINE AND HAD A SERIOUS BREATHING ISSUES SHORTLY AFTER FIRST USING THE DEVICE AND STOPPED USING THE SO CLEAN ON CPAP MACHINE. I HAVE HAD CHRONIC SINUS ISSUES MOST OF 2019 AND 2020. WHEN I STOPPED USING THE DREAM STATION IN (B)(6) 2021 MY SINUS ISSUES RESOLVED ITSELF. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5104081 | 2021-09-20 | 2021-06-20 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | THE LIGHT KEEPS TELLING ME THAT MY FILTERS NEED CHANGED EVEN WHEN I JUST CHANGED THEM. MY POWER BUTTON FLASHES. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-09
MW5104086 | 2021-09-20 |  | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I WAS NOT NOTIFIED THAT MY PHILIPS RESPIRONICS CPAP "DREAMSTATION" WAS RECALLED UNTIL: (B)(6) 2021, WHEN I RECEIVED THE RECALL NOTICE IN THE MAIL FROM PHILIPS, WHY DID IT TAKE SO LONG? WEREN'T PATIENTS NOTIFIED BACK IN (B)(6)? - I STILL HAVE NOT HEARD FROM EITHER MY MEDICAL EQUIPMENT PROVIDER, NOR MY SLEEP DOCTOR; WHY? YES, MY MACHINE HAS THE PROBLEM OF THE FOAM BREAKING INTO PARTICLES, BUT, I DID NOT REALIZE THAT IT WAS THE MACHINE MALFUNCTIONING. AT FIRST, I THOUGHT IT MIGHT BE A BUG, FLOATING IN MY HUMIDIFIER TANK, BECAUSE IT WAS BLACK. THEN I GUESSED IT WAS JUST DUST FROM THE FILTER. I DID NOT KNOW, BECAUSE I HAD NOT BEEN ADVISED. NEITHER BY PHILIPS, NOR BY MY MEDICAL EQUIPMENT PROVIDER, NOR BY MY SLEEP DOCTOR. HENCE, PHILIPS SAYING IT HAS "A LOW COMPLAINT RATE". HOW WERE WE, THE PATIENTS, SUPPOSED TO KNOW TO WATCH FOR THIS? SO, I KEPT USING MY CPAP MACHINE. I NOTICED IT: MONTHS AGO. EVEN LONG BEFORE THAT, I NOTICED THAT MY CPAP MACHINE WAS GETTING LOUDER, AND LOUDER. I USED MY MACHINE A TOTAL OF 19 MONTHS (RE. OFF-GASSING, ETC.) - NO, I NEVER USED ANY OZONE CLEANER. I JUST WASHED EVERYTHING BY HAND, AS INSTRUCTED BY MY MEDICAL EQUIPMENT PROVIDERS - YES, I HAVE REGISTERED MY DREAMSTATION CPAP MACHINE WITH PHILIPS, WHO CONFIRMED: IT IS AN AFFECTED DEVICE THAT IS RECALLED. I HAVEN'T HAD ANY TESTS YET, AS MY SLEEP DOCTOR STILL HASN'T CONTACTED ME TO NOTIFY ME OF THE RECALL OF MY PHILIPS DREAMSTATION CPAP MACHINE. SHOULDN'T DOCTORS HAVE NOTIFIED THEIR PATIENTS: MONTHS AGO? NEITHER HAVE I BEEN CONTACTED BY MY MEDICAL EQUIPMENT PROVIDER EITHER. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-09
MW5104088 | 2021-09-20 | 2021-09-17 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPERONICS DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500H11C |  |  | I HAVE BEEN USING TWO OF THE RECALLED PHILIPS CPAP MACHINES SINCE 2013. I HAVE HAD NUMEROUS HEALTH ISSUES THAT SEEM TO BE RELATED TO THESE CPAP MACHINES. I HAVE A CONSTANT SORE THROAT. MY EYES ARE RED CONTINUOUSLY. THEY ALSO ITCH. MY NOSE BURNS. I HAVE A CYST ON MY THYROID. I HAVE UNEXPLAINED MUSCLE LOSS. MY ABSOLUTE LYMPHOCYTE COUNTS ARE LOW CONTINUOUSLY. I HAVE BEEN TESTED FOR ALLERGIES AND HAVE NONE. I HAVE A POTENTIAL TUMOR IN MY EAR CANAL. I HAVE BALANCE ISSUES. I HAVE A RUNNY NOSE WHEN USING THE DEVICE AND MOST OF THE DAY AFTERWARD. I HAVE NUMEROUS NASAL POLYPS DOCUMENTED BY AN MRI. I HAVE MEDICAL RECORDS TO DOCUMENT MOST OF THESE CLAIMS. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | I |  | 2021-09
2518422-2021-04091 | 2021-09-20 | 2021-07-19 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A SINUS INFECTION. THE PATIENT WAS PRESCRIBED ANTIBIOTICS AND STEROIDS IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT ALLEGED TO DEVELOP TO DEVELOP A SINUS INFECTION. THE PATIENT WAS PRESCRIBED ANTIBIOTICS AND STEROIDS IN RESPONSE TO THE REPORTED EVENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04205 | 2021-09-20 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11W | DSX500H11W |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION B5 SHOULD BE REPORTED AS THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDING TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERROR FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND UNIT SCRAPPED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03843 | 2021-09-20 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03841 | 2021-09-20 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING CHAMBER, RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER LATER, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03840 | 2021-09-20 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING /CHAMBER DIFFICULTY BREATHING/SHORT OF BREATH,DIZZY RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. A CORRECTION TO B5 WAS MADE AND SHOULD BE THE MANUFACTURER RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING /CHAMBER DIFFICULTY BREATHING/SHORT OF BREATH,DIZZY RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03837 | 2021-09-20 | 2021-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED TO HAVE BREATHING PROBLEMS/SHORTNESS OF BREATH, COUGH, FATIGUE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. AFTER THE SECOND ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, CUSTOMER STATED THE DEVICE WON'T BE AVAILABLE FOR RETURN. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04134 | 2021-09-20 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION EXPERT W/HUM/HT/CELL, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX501T12C | CAX501T12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP HEARING LOSS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 01, 2024, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP HEARING LOSS.THE REPORTED EVENT OF LOSING THEIR HEARING AND VOICE ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTIONS B1, B2, H1 AND H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR ONLY THE PRODUCT PROBLEM (BOTH ADVERSE EVENTS AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04121 | 2021-09-20 | 2021-09-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING BLACK DEPOSITS IN NOSE FILTERS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK DEPOSITS IN NOSE FILTERS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION A2 (PATIENT AGE) CORRECTED AND H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04103 | 2021-09-20 | 2021-07-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED SKIN CANCER. THE PATIENT DID RECEIVE MEDICAL INTERVENTION DIAGNOSING THE SKIN CANCER. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY SUBMITTED MDR 2518422-2021-04103-1 WITH INCORRECT SECTIONS B1, B2, H1, H6. CORRECTIONS TO PREVIOUS MDR ARE MADE IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY PRODUCT PROBLEM WAS CHECKED IN PREVIOUS MDR). SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY IT WAS BLANK). SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6- HEALTH EFFECT - IMPACT CODE WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04100 | 2021-09-20 | 2021-07-12 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED PNUEMONIA, DIFFICULTY BREATHING AND A BAD COUGH. THE PATIENT DID RECEIVE MEDICAL INTERVENTION WHICH REQUIRED HOSPITALIZATION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED PNEUMONIA, DIFFICULTY BREATHING AND A BAD COUGH. THE PATIENT DID RECEIVE MEDICAL INTERVENTION WHICH REQUIRED HOSPITALIZATION. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF UNKNOWN CONTAMINANT ON THE UI PANEL, UI KNOB, BOTH SIDE PANELS, ISO PORT PANEL, INSIDE THE ISO PORT, UNDER THE ACCESSORY MODULE FLIP DOOR, AND UNDER THE SD COVER DOOR. AN INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF A CONGEALED CONTAMINATE ON THE ISO PORT, ISO PORT PANEL, ISO PORT O-RING, INSIDE THE BLOWER BOX, AND THE BLOWER OUTPUT SEAL, A YELLOWISH UNKNOWN CONTAMINANT WAS ON THE TOP ENCLOSURE, UI PANEL, BLOWER CONNECTOR, BLOWER CAP, BLOWER BOX, ISO PORT PANEL, AND BOTTOM ENCLOSURE, UNKNOWN CONTAMINANT IN THE AIRPATH. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND NO ERRORS LOGGED. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF MULTIPLE CONTAMINANTS WITHIN THE DEVICE. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. SECTION D9, G3 AND H6 HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04102 | 2021-09-20 | 2021-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED ASTHMATIC SYMPTOMS, SHORTNESS OF BREATH. THE PATIENT DID RECEIVE MEDICAL INTERVENTION REQUIRING THE USE OF A PRESCRIBED INHALER. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 07, 2024. SECTION H11 SHOULD BE REPORTED AS THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AND CAUSED ASTHMATIC SYMPTOMS, SHORTNESS OF BREATH, STRANGE ODOR IN DEVICE, SOME THINGS IN LUNGS RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT DID RECEIVE MEDICAL INTERVENTION REQUIRING THE USE OF A PRESCRIBED INHALER. THERE IS NO REPORT OF PATIENT HARM OR INJURY. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM (ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR). SECTION B2 WAS CHANGED TO BLANK (PREVIOUSLY IT WAS OTHER SERIOUS). SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04101 | 2021-09-20 | 2021-07-22 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED LIGHTHEADEDNESS, DIZZINESS, HEADACHES , NAUSEA AND DIAGNOSED WITH A LUNG TUMOR. THE PATIENT DID RECEIVE MEDICAL INTERVENTIONS IN ORDER TO DIAGNOSE LUNG TUMOR. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING TO LIGHTHEADEDNESS, DIZZINES, HEADACHES , NAUSEA AND DIAGNOSED WITH A LUNG TUMOR RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT DID RECEIVE MEDICAL INTERVENTIONS INORDER TO DIAGNOSE LUNG TUMOR. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HEALTH EFFECT-CLINICAL CODE WAS MISSED TO CAPTURE IN PREVIOUS MDR,NOW IT IS CORRECTED AND UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04097 | 2021-09-20 | 2021-08-04 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED LUNG PROBLEMS CONSISTING OF SHORTNESS OF BREATH AND MULTIPLE SINUS INFECTIONS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION THAT CONSISTED OF EMERGENCY ROOM VISITS, LUNG X-RAYS AND OXYGEN. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED LUNG PROBLEMS CONSISTING OF SHORTNESS OF BREATH AND MULTIPLE SINUS INFECTIONS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION THAT CONSISTED OF EMERGENCY ROOM VISITS, LUNG X-RAYS AND OXYGEN. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04158 | 2021-09-20 | 2021-08-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A (CPAP, BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A (CPAP, BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN THE INITIAL REPORT SECTION B5 WAS NOT UPDATED, THE CORRECT B5 SHOULD BE - THE PATIENT HAS ALLEGED FILTERS GETTING BLACK, HEADACHES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF BEIGE DUST OR DIRT CONTAMINATE IN THE FILTER INTAKE AREA AND INTAKE CANAL, MINERAL DEPOSIT SPOTS IN THE OUTLET PORT POSSIBLY FROM HOSE RAINOUT, BROWN SMOKE CONTAMINATION COATING IN CRACKS AND CREVASSES, AND ON ALL INTERNAL ENCLOSURE PLASTIC SURFACES, AND MINOR AMOUNT OF KERATIN ON THE BLOWER BOX OUTLET PORT.. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND 5 INSTANCES OF E-22, 1 INSTANCE OF E-53 AND 2 INSTANCES OF E-130. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH BEIGE DUST OR DIRT CONTAMINATE IN THE FILTER INTAKE AREA AND INTAKE CANAL, MINERAL DEPOSIT SPOTS IN THE OUTLET PORT POSSIBLY FROM HOSE RAINOUT, BROWN SMOKE CONTAMINATION COATING IN CRACKS AND CREVASSES, AND ON ALL INTERNAL ENCLOSURE PLASTIC SURFACES, AND MINOR AMOUNT OF KERATIN ON THE BLOWER BOX OUTLET PORT. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION D9,G3,H6 HAS BEEN UPDATED/CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04140 | 2021-09-20 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTION G1 WAS CORRECTED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04163 | 2021-09-20 | 2021-01-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102994) ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A SINUS INFECTION. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON JUL 11, 2023, AND SECTION H11 SHOULD BE REPORTED AS THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102994) ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A SINUS INFECTION. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM (ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR) SECTION B2 WAS CHANGED TO BLANK (PREVIOUSLY IT WAS OTHER SERIOUS) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION SECTION H6 HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04096 | 2021-09-20 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A SMALL SORE AND A LARGE SORE IN THE NOSE. THE PATIENT DID RECEIVE MEDICAL INTERVENTION THAT REQUIRED PRESCRIBED MEDICATION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A SMALL SORE AND A LARGE SORE IN THE NOSE. THE PATIENT DID RECEIVE MEDICAL INTERVENTION THAT REQUIRED PRESCRIBED MEDICATION. THESE EVENTS WERE REVIEWED BY THE PMS CLINICAL EXPERT AND THE EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED IN THE BASE UNIT AND DUST/DIRT CONTAMINATION ON THE FLIP LID SEAL, AROUND THE HEATER PLATE, AND THROUGHTOUT THE AREAS OUTSIDE OF THE BLOWER BOX OF THE MACHINE IS CONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM WAS NOT OBSERVED THROUGHOUT DEVICE ENCLOSURE, AND AIRPATH, SUGGESTING A SOURCE EXTERNAL TO THE DEVICE. BLOWER SEAL SHOWED FINE DARK CONTAMINATION THAT APPEARS CONSISTENT WITH KERTAIN CONTAMINATION OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THE MANUFACTURER FOUND ONE ERRORS FOUND (ER 2243857). THE MANUFACTURER CONCLUDES THAT THEY CONFIRM THE CUSTOMER'S ALLEGATION AND THERE IS DUST / DIRT CONTAMINATION AND THE PRESENCE OF CONTAMINATION IN THE AIRPATH, SOURCE OF CONTAMINATIONS WERE EXTERNAL TO THE DEVICE. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04161 | 2021-09-20 | 2021-06-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103030) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SHORTNESS OF BREATH AND WHEEZING. THE PATIENT WAS PRESCRIBED STEROIDS IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON OCTOBER 09, 2023 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED AN INFORMATION FORM VOLUNTARY MEDWATCH (MW5103030), ALLEGING A PATIENT REPORTED SHORTNESS OF BREATH AND WHEEZING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT WAS PRESCRIBED STERIOIDS IN RESPONSE TO THE WHEEZING AND SHORTNESS OF BREATH. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGED EYE IRRITATION ASTHMA (NEW OR WORSENING) AND CANCER. THE SECTION E1 WAS UPDATED. SECTION H6 HEALTH EFFECTS - CLINICAL CODES WERE UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04090 | 2021-09-20 | 2021-07-22 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SHORTNESS OF BREATH. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | IN THE PREVIOUS FOLLOW-UP REPORT SECTION H10 WAS INCORRECT, IT SHOULD HAVE BEEN REPORTED AS : THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A PATIENT DEVELOPED SHORTNESS OF BREATH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04088 | 2021-09-20 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/HT/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400T11C | DSX400T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP PARTICLES IN THEIR LUNGS. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO HAVING A STRANGE PARTICLES IN LUNGS, DIFFICULTY / SHORT OF BREATH, WEAKNESS. THE MEDICAL INTERVENTION REQUIRED BY THE PATIENT WAS UNKNOWN. THE REPORTED EVENTS OF STRANGE PARTICLES IN LUNGS DIFFICULTY / SHORT OF BREATH, WEAKNESS AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THESE EVENTS ARE ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURER CONCLUDES NO FURTHER ACTION IS NECESSARY. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION AND AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTIONS B1, B2 HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- CLINICAL CODE CORRECTED AND HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04112 | 2021-09-20 | 2021-07-24 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04162 | 2021-09-20 | 2020-05-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103014) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP EYE ISSUES. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 11, 2024, AND SECTION H11 SHOULD BE REPORTED AS THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED A VOLUNTARY MEDWATCH (MW5103014) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP EYE ISSUES, SINUS, HEADACHE, HYPERSENSITIVITY AND NASAL SECRETION. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM (ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR). SECTION B2 WAS MADE BLANK (PREVIOUSLY IT WAS OTHER SERIOUS). SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04085 | 2021-09-20 | 2019-02-15 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102045) ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP THYROID NODULES. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES (B)(4). THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04147 | 2021-09-20 | 2021-08-23 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP PULMONARY FIBROSIS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP PULMONARY FIBROSIS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE REPORTED EVENT IS ASSESSED AS SERIOUS INJURY BUT NOT RELATED TO THE DEVICE. SINCE THERE IS NO CUSTOMER INFORMATION, THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04137 | 2021-09-20 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11 | DSX700H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04155 | 2021-09-20 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04089 | 2021-09-20 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP NODULES IN THEIR LUNGS. THE PATIENT WAS PRESCRIBED STEROIDS IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON 10 OCT, 2024 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGED TO DEVELOP NODULES IN THEIR LUNGS, COUGH AND WHEEZING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE REPORTED EVENT OF DIAGNOSED WITH PATIENT TO DEVELOPING PNEUMONIA TWICE WHILE USING THE DEVICE AND RECENTLY GOT OUT OF HOSPITAL. THE PATIENT WAS PRESCRIBED STEROIDS IN RESPONSE TO THE REPORTED EVENT. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECTS: CLINICAL CODES AND HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04138 | 2021-09-20 | 2021-08-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING LIGHTHEADED, BREATHING ISSUE, ALLERGIES AND INFECTION INSIDE OF NOSE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING LIGHTHEADED, BREATHING ISSUE, ALLERGIES AND INFECTION INSIDE OF NOSE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION G1-MANUFACTURER CONTACT OFFICE DETAILS WERE CORRECTLY UPDATED. SECTION H6 HAS BEEN CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04139 | 2021-09-20 | 2021-08-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING/CHAMBER DUE TO CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04159 | 2021-09-20 | 2020-08-20 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/WIFI,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500T11W | DSX500T11W |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103235) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A SINUS INFECTION. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | IN THE PREVIOUS REPORT, SECTION H10 WAS MENTIONED AS INCOMPLETE; THE CORRECT SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING HEADACHES AND DIZZINESS, DAILY EYE IRRITATION, DAILY NASAL CONGESTION WITH SEVERE THROAT PHLEGM, MULTIPLE SINUS INFECTIONS REQUIRING 4 REGIMENS OF ANTIBIOTICS, REOCCURRING COUGHS, AND ELEVATED LIVER FUNCTION TESTS. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. THE DEVICE WAS EVALUATED, AND EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN WAS NOT OBSERVED IN THE BASE UNIT. THE MANUFACTURER OBSERVED DUST/DIRT CONTAMINATION THROUGHOUT THE DEVICE ENCLOSURE AND AIRPATH, SUGGESTING A SOURCE EXTERNAL TO THE DEVICE. EVIDENCE OF WATER INGRESS ON THE BLOWER AND BLOWER BOX WAS FOUND. THE INTERNAL ASPECTS OF THE DEVICE WERE INSPECTED. THE DEVICE POWERED ON, AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THAT THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN OBSERVED IN THE BASE UNIT. THE MANUFACTURER CONFIRMED THE PRESENCE OF CONTAMINATION IN THE AIRPATH. SECTION H6 MEDICAL DEVICE PROBLEM CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, AND INVESTIGATION CONCLUSIONS HAVE BEEN CORRECTED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04164 | 2021-09-20 | 2021-01-04 | I, n, j, u, r, y | Manufacturer report | PHILLIPS RESPIRONICS DREAMSTATION BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102982) ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO REQUIRE HOSPITALIZATION. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | PREVIOUS REPORT HAS INCORRECT H3 SECTION. H3 SECTION HAS BEEN CORRECTED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04120 | 2021-09-20 | 2021-08-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX700T12 | CAX700T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04222 | 2021-09-21 | 2014-01-01 | D, e, a, t, h | Manufacturer report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102931) ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON OCT 09, 2023, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION FROM THE VOLUNTARY MEDWATCH (MW5102931) ALLEGING VISUALIZATION OF BLACK PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED COUGH, BREATHING ISSUE, NASAL/THROAT IRRITATION OR SORENESS. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGATION OF DEATH, EYE IRRITATION; KIDNEY DISEASE/TOXICITY. THE SECTION B1, B2, H1 AND H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY THE PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.). SECTION B2 WAS CORRECTED TO DEATH. (PREVIOUSLY, IT WAS BLANK.). SECTION H1 WAS CHANGED FROM MALFUNCTION TO DEATH. SECTION H6 HEALTH EFFECTS: CLINICAL CODES AND HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04221 | 2021-09-21 | 2019-07-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102973) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A LUNG ISSUE. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING UNEXPLAINED RASHES AND LUMPS UNDER THE SKIN AND WAS HOSPITALIZED FOR SOMETHING IN THE LUNGS, MUSCLE LOSS, COUGH, HEART FLUTTERING, DIZZINESS, EYE TWITCHING ON AND OFF, AND ALSO THAT THE VISION SEEMED TO BE GETTING WORSE. THE PATIENT ALSO REPORTED THAT THEY RECEIVED A NEW MACHINE, AND WHEN PUTTING NEW WATER IN THE CHAMBER, THEY FOUND BLACK LIKE PARTICLES IN THE CHAMBER. THE PATIENT WENT TO THE DOCTOR, WHO DID HEAR WHEEZING IN THEIR BREATHING AND PUT THEM ON ANTIBIOTICS. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HEALTH EFFECT - CLINICAL CODE, HEALTH EFFECT - IMPACT CODE HAS BEEN CORRECTED, AND TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED. SECTION D1, D4, G4, H4 HAS BEEN CORRECTED / UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04251 | 2021-09-21 | 2021-08-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX400H12 | CAX400H12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING/CHAMBER, PARTICLES IN AIRWAY FROM DEVICE AND CANNOT BREATHE AT NIGHT BECAUSE CONGESTED FROM THE MACHINE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE - THE PATIENT ALLEGED PARTICLES IN TUBING/CHAMBER, PARTICLES IN AIRWAY FROM DEVICE AND CANNOT BREATHE AT NIGHT BECAUSE CONGESTED FROM THE MACHINE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. CORRECTED INFORMATION WAS PROVIDED IN D4 (PREVIOUSLY UDI NUMBER CAPTURED INCORRECTLY). SECTION H6 CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04149 | 2021-09-21 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES AND BLACK SUBSTANCE. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS VISIBLE FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04169 | 2021-09-21 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION H6 HEALTH EFFECT - CLINICAL CODE & SECTION H10 CAPTURED INCORRETLY IN PREVIOUS REPORTS, IT SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED DEVICE OR POWER CORD OR SUPPLY TOO HOT TO TOUCH, STATES THAT DEVICE SHOOTS OUT PURPLE STUFF AND SPARKED WHEN PLUGGED IN. ALSO, ALLEGED TO HAVE DIFFICULTY IN BREATHING OR SHORTNESS OF BREATH. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04160 | 2021-09-21 | 2020-01-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102937) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO NEED A CHOLECYSTECTOMY THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO NEED A CHOLECYSTECTOMY THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT. SECTION B5 SHOULD BE CORRECTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED. THE PATIENT HAS ALLEGED BREATHING PROBLEMS AT NIGHT,SLEEP APNEA,ASTHMA,SEVERE HEADACHE,NOSE BLEEDS,NAUSEA,WAKING COUGHING,DIZZINESS,STOMACH ISSUES,ABDOMINAL ISSUES,HAD EMERGENCY GALL BLADDER SURGERY,THROAT IRRITATION RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04150 | 2021-09-21 | 2021-09-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | 761P | 761P |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
12505638 | 2021-09-21 | 2021-08-30 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | TRILOGY EV300, USA | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | DS2200X11B | DS2200X11B |  | ON (B)(6) 2021, A PHILIPS FIELD SERVICE REPRESENTATIVE CAME ON-SITE AND REPAIRED A EV300 VENTILATOR THAT HAD FAILED WITH THE RED SCREEN AND MESSAGE ¿VENTILATOR INOPERATIVE¿ DISPLAYED, JUST LIKE THE OTHER 3 SIMILAR INCIDENTS THAT WE HAD IN (B)(6) 2021. THIS TIME HOWEVER THE ISSUE WAS THE MAIN CIRCUIT BOARD (PCA BOARD), NOT A SOFTWARE ISSUE LIKE THE PREVIOUS ONES. THIS MAKES 4 TOTAL EV300¿S IN THE LAST 6 MONTHS THAT HAVE HAD ¿VENTILATOR INOPERABLE¿ ALARMS THAT HAD TO BE RESOLVED BY PHILIPS. MANUFACTURER RESPONSE FOR EV300, TRILOGY (PER SITE REPORTER). SOFTWARE UPDATE, FIELD SERVICE ENGINEER ONSITE REPAIR. |  | N | Y | I |  | 2021-09
2518422-2021-04225 | 2021-09-21 | 2021-08-03 | I, n, j, u, r, y | Manufacturer report | BIPAP AUTO BIFLEX, W/SD CARD, US | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DS750S | DS750S |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE ASTHMA AND OTHER RESPIRATORY PROBLEMS. THE PATIENT WAS PRESCRIBED INHALER IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGES ASTHMA AND OTHER RESPIRATORY PROBLEMS, HEADACHE, EYES HAVE BEEN BURNING. ITCHY. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 CLINICAL CODE WAS MISSED TO CAPTURE IN PREVIOUS REPORT AND UPDATED IN THIS REPORT. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04148 | 2021-09-21 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04072 | 2021-09-21 | 2021-09-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04182 | 2021-09-21 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11 | DSX400H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04207 | 2021-09-21 | 2021-08-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2100X15B | IN2100X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, THE DEVICE WAS ALARMING FOR AN OXYGEN REGULATION ISSUE. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR WAS ALARMING FOR AN OXYGEN REGULATION ISSUE AND THE OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. THE DEVICE'S OXYGEN BLENDING MODULE SUBASSEMBLY WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04236 | 2021-09-21 | 2021-08-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WOULD NOT POWER ON. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S POWER SUPPLY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04209 | 2021-09-21 | 2021-08-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS ALARMING FOR A LOW TIDAL VOLUME ALARM ISSUE. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE ALSO FAILED TEST STEPS DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04208 | 2021-09-21 | 2021-08-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS EVALUATED BY THE CUSTOMER. THE CUSTOMER REPLACED THE SYSTEM BOARD TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04070 | 2021-09-21 | 2021-09-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLANK. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE WAS EVALUATED AT THE MANUFACTURER'S SERVICE CENTER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04069 | 2021-09-21 | 2021-08-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04068 | 2021-09-21 | 2021-08-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLANK. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE WAS EVALUATED AT THE MANUFACTURER'S SERVICE CENTER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04152 | 2021-09-21 | 2021-06-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION A3 (PATIENT'S GENDER) WAS UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04258 | 2021-09-21 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | ADDITIONAL INFORMATION WAS RECEIVED ON MAR 29, 2023, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04261 | 2021-09-21 | 2021-09-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE VENTILATOR FAILED A TEST STEP DURING TESTING. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04071 | 2021-09-21 | 2021-09-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04260 | 2021-09-21 | 2021-09-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE EXTERNAL FLOW SENSOR IS NOT FUNCTIONING. THE DEVICE'S EXTERNAL FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
MW5104133 | 2021-09-21 | 2019-09-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | 00606959022478 |  |  | ASTHMA LIKE CONDITION; PATIENT BEGAN HAVING CHEST PRESSURE, TROUBLE BREATHING, SHORTNESS OF BREATH, DIFFICULTY WITH APNEA WHEN LYING FLAT, AND OCCASIONAL FLUID IN LUNGS. WORSE WITH PHYSICAL ACTIVITY AND RECUMBENCY, REMAINS AT REST. PCP, PULMONOLOGIST AND REFERRAL TO NATIONAL JEWISH HEALTH IN DENVER CO. X-RAYS, LABS, SLEEP STUDY, METHACHOLINE CHALLENGE, CT, WERE INCONCLUSIVE AS TO CAUSE. (SOME TESTS UNABLE TO PERFORM DUE TO PATIENT SENSITIVITY TO CONTRAST MEDIA). PATIENT NOTED BLACK MOLD IN CPAP MACHINE TO MULTIPLE DR'S. ALL MEDICAL RECORDS CAN BE RETRIEVED IF NEEDED. THESE ARE THE TEST THAT PATIENT REMEMBERS MOST NOTABLY. FULL MEDICAL RECORDS CAN BE RETRIEVED IF NEEDED. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04269 | 2021-09-21 | 2021-08-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED A TEST STEP DURING TESTING. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04145 | 2021-09-21 | 2021-09-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, GER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DEX500S13 | DEX500S13 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
MW5104116 | 2021-09-21 | 2020-12-17 | D, e, a, t, h | Voluntary report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | PHILIPS / RESPIRONICS, INC. |  |  |  | MY MOTHER WAS USING A TRILOGY 100 VENTILATOR AND I RECEIVED A RECALL NOTICE IN THE MAIL TODAY WHICH IS DATED 10-17-2021, AND IT LISTED A BUNCH OF SYMPTOMS FROM THE BREATHING MACHINE FAILING ON ITS FILTER SYSTEM. MY MOTHER HAD ALL THESE SYMPTOMS. BEFORE SHE WOUND UP SUCCUMBING TO PNEUMONIA WHICH I BELIEVE WAS BROUGHT ON BY KIDNEY PROBLEMS AND THE LIKE NOT SURE IF I CAN PUT A PHONE NUMBER IN HERE BUT I'M GOING TO. MY NAME IS (B)(6), MY MOTHER'S NAME WAS (B)(6), MY PHONE NUMBER IS (B)(6). THEY COULD NOT FIND ANY SYMPTOMS OF FLU OR NOTHING WITH MY MOTHER, BUT THEY HAD RECENTLY CHANGED HER MACHINE TO A NEW MACHINE AND SHE DEVELOPED ALL THESE SYMPTOMS OF ITCHING AND EVERYTHING THAT WAS LISTED ON THIS PAPER, I HAVE WHICH INCLUDES SKIN IRRITATION, HEADACHES EVERYTHING THEY LISTED. MY MOTHER WAS SUFFERING THROUGH BEFORE SHE PASSED AWAY FOR SOMEWHAT NO REASON SHE DID HAVE COPD BUT WAS DOING WELL SEEMS THAT SHE JUST ALL OF A SUDDEN. NOW THE BLUE COULDN'T FIGURE OUT WHY SHE FELL ILL, AND DIED A WEEK LATER. I DON'T HAVE ANY OF THE DATA BECAUSE WE HAD RETURNED THE MACHINE. I DO HAVE A PRINTOUT BUT I'M NOT SURE WHAT IT ALL MEANT OR WILL IT MATTER BECAUSE THE PARTICLES ARE THE PROBLEM. SHE MAY HAVE BEEN BREATHING PARTICLES FROM THE MACHINE THAT CAUSED THESE PROBLEMS. ANYTHING TO HELP IN THIS MATTER WOULD BE GREATLY APPRECIATED. I WOULD LIKE TO FIGURE OUT THE REASON WHY MY MOTHER MAY HAVE PASSED AWAY. THANK YOU. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5104117 | 2021-09-21 | 2020-10-27 | D, e, a, t, h | Voluntary report | SYSTEMONE ASV4 | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | PHILIPS / RESPIRONICS, INC. | 560P |  |  | (B)(6) RANDOMLY STARTED TO VOMIT AND FEEL ILL ONE AFTERNOON AT WORK. TWENTY FOUR HOURS LATER HE WAS VERY CONFUSED, COULD NOT ANSWER SIMPLE QUESTIONS. HE WAS TAKEN TO ER, QUICKLY ADMITTED, TESTS SHOWED SEPSIS. A FEW HOURS LATER HE NEEDED TO BE INTUBATED AS HE WAS HAVING SEVERE BREATHING PROBLEMS. DURING INTUBATION HE WENT INTO CARDIAC ARREST, HE SURVIVED. TESTS SHOWED HIS INFECTION WAS IN HIS AORTA VALVE, AND SEVERELY DAMAGED. TESTS SHOWED HE HAD NUMEROUS STROKES, HEART FAILURE, LIVER FAILURE, KIDNEY FAILURE, 10 DAYS AFTER FIRST SYMPTOM HE WAS DECLARED BRAIN DEAD AND THE DECISION WAS MADE TO REMOVE HIM FROM THE VENTILATOR, HE DIED A FEW MINUTES LATER. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5104118 | 2021-09-21 | 2021-07-29 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO BI-PAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DREAMSTNAUTOBIPAP APHUMCELL DOM | DSX700H11C |  | AFTER USE OF A RECALLED BI-PAP DEVICE, I WAS ADMITTED TO THE HOSPITAL AND HAD SURGERY TO REMOVE A CANCEROUS TUMOR FROM MY LUNG. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5104126 | 2021-09-21 | 2018-09-01 | I, n, j, u, r, y | Voluntary report | REMSTAR PRO C-FLEX SYSTEM ONE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | 469P |  |  | THE DATE IS APPROXIMATE. I USE A CPAP MACHINE THAT WAS RECALLED. I HAVE HAD INFLAMED NOSTRILS AND NASAL PASSAGES, SOME HEADACHES, AND SOME RESPIRATORY IRRITATIONS, SORT OF LIKE ALLERGY SYMPTOMS. I DID SEE AN ENT LAST SUMMER AND WAS PRESCRIBED OINTMENT FOR MY NOSE AND TOLD TO USE FLONAZE REGULARLY. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5104127 | 2021-09-21 | 2011-02-11 | I, n, j, u, r, y | Voluntary report | SYSTEM ONE Q-SERIES | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | PHILIPS / RESPIRONICS, INC. |  |  |  | TOTALLY HEALTHY BUT SNORED WIFE INSISTED I HAD SLEEP DISORDER. STARTED USING SYSTEM ONE PHILIP RESPIRONICS. WITHIN 8 MONTHS MY HEALTH STARTED DECLINING. I ASKED IF MY SHORTNESS OF BREATH WAS CAUSED BY THE SLEEP ONE DEVICE I WAS USING AND EVERY DOCTOR I ASKED DID NOT THINK SO. I USED THIS PRODUCT 6 + YEARS. I WAS FORCED TO TAKE EARLY RETIREMENT BECAUSE I COULD NO LONGER DO MY JOB. PLEASE REACH TO ME AND I WILL RELEASE ALL MY MEDICAL RECORDS TO YOU. THEY HAVE BEEN IN MY HEART MY LUNGS AND FIXED WHAT THEY THOUGHT COULD HELP. BUT BACK TO THE CPAP MACHINE THING JUST KEPT GETTING WORSE. ONE (1) YEAR AGO I WAS GIVEN A NEW MACHINE, DIFFERENT BRAND, I HAVE NOT GOT BETTER BUT HAVE NOT GOT WORSE. EIGHT (8) YEARS OF TESTING (B)(6). FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5104129 | 2021-09-21 | 2021-09-15 | I, n, j, u, r, y | Voluntary report | RESPIROINCS PHILIPS TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | PHILIPS / RESPIRONICS, INC. |  |  |  | WE HAVE 3 OUT OF OUR 8 PEDIATRIC PATIENTS ON 24/7 CHRONIC VENTILATION AND ALL 3 OF THESE PATIENTS HAVE SUFFERED UNKNOWN HARM FROM RECEIVING UP TO 10 CM OF DELIVERED VENTILATOR PRESSURE ABOVE SET LIMITS. FROM INFORMATION I HAVE TO DATE, THE HIGH PRESSURE ALARMS WERE NOT TRIGGERED. THERE IS NO OFFICIAL RECALL FOR THIS DEVICE: RESPIRONICS PHILIPS TRILOGY EVO. WE ARE EXTREMELY LIMITED IN WHAT DEVICES WE CAN PROVIDE OUR PATIENTS DUE TO JUNE RECALL OF TRILOGY 100 AND 200. WITH AID OF OUR DME COMPANIES, WE HAVE BEEN ABLE TO ALERT OUR FAMILIES, AND RESET THE VENTILATORS WITH ALLEVIATES THE ISSUE AT LEAST SHORT TERM. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04237 | 2021-09-21 | 2021-07-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN THE PREVIOUSLY SUBMITTED REPORT SECTION B5 WAS INCOMPLETE, SECTION B5 IS- THE PATIENT ALSO ALLEGED OF HAVING SINUS CONGESTION, HEADACHES, SCRATCHY THROAT, DIFFICULTY BREATHING AND DIZZINESS. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF LIQUID INGRESS. AN INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF BLACK CONTAMINANT ON FRONT PANEL, ADDITIONAL CONTAMINANT PRESENT ON THE BLOWER, INSIDE THE BLOWER, ON THE BLOWER CAP, AND INSIDE THE BLOWER BOX. THE MANUFACTURER FOUND EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. THE DEVICE WAS HOOKED UP TO POWER SUPPLY, AIRFLOW WAS VERIFIED AND THE DEVICE OPERATED PROPERLY. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND 2 INSTANCES OF E-130. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF LIQUID INGRESS, BLACK CONTAMINANT ON FRONT PANEL, ADDITIONAL CONTAMINANT ON THE BLOWER, INSIDE THE BLOWER, ON THE BLOWER CAP, AND INSIDE THE BLOWER BOX. THE MANUFACTURER CONFIRMED THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. | N | Y | LAY USER/PATIENT |  | 2021-09
MW5104135 | 2021-09-21 | 2020-04-26 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | PRIOR TO (B)(6) 2020, I HAVE NOT EXPERIENCED PROLONGED HEADACHES OF ANY TYPE. BEGINNING THIS DATE I HAVE BEEN EXPERIENCING MIGRAINES AND MULTIPLE TEST HAVE NOT BEEN ABLE TO FIND THE CAUSE. IT IS BELIEVED THAT THE USE OF A PHILIPS DREAMSTATION MAY BE THE REASON FOR THE HEADACHES. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | I |  | 2021-09
MW5104136 | 2021-09-21 | 2021-09-03 | I, n, j, u, r, y | Voluntary report | DREAMSTATION BI-LEVEL PAP DEVICE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | WAS HOSPITALIZED ON (B)(6) 2021 AND DIAGNOSED WITH BACTERIAL PNEUMONIA IN MY UPPER RIGHT LUNG (CHEST X-RAY). DUE TO A DRY COUGH AND DIFFICULT TO CONTROL T1D AND BLOOD PRESSURE, WAS KEPT ON IV ANTIBIOTICS AND PLACED IN AN INSOLATION WARD BECAUSE ONE OF THE DOCTORS THOUGHT I MIGHT HAVE TB. A CT SCAN SHOWED THE PNEUMONIA, BUT NO TB OR LUNG CANCER. I WAS RELEASED ON (B)(6) 2021 WITH 5 MORE DAYS OF ANTIBIOTICS (PILLS). I AM GOING TO HAVE ANOTHER CHEST X-RAY ON (B)(6) 2021 TO DOCUMENT THE PRESENCE/ABSENCE OF PNEUMONIA. WHEN I RETURNED HOME ON (B)(6) 2021, I RECEIVED A LETTER FROM PHILIPS RESPIRONICS STATING THAT MY DREAMSTATION BI-LEVEL PAP DEVICE WAS BEING RECALLED DUE TO DEGRADATION OF PE-PUR FOAM AND VOC EMISSIONS. THE RECALL PAPERWORK STATES THAT THE USE OF THE DEVICE IS NOT RECOMMENDED BECAUSE "THESE ISSUES CAN RESULT IN SERIOUS INJURY WHICH CAN BE LIFE-THREATENING, CAUSE PERMANENT IMPAIRMENT, AND//OR REQUIRE MEDICAL INTERVENTION TO PRECLUDE PERMANENT IMPAIRMENT". APPROXIMATELY 6 WEEKS EARLIER, I HAD COMPLAINED TO LINCARE AND PHILIPS RESPIRONICS THAT TINY BLACK PARTICULATES (FEEL RUBBERIZED) WERE ON MY DISPOSABLE FILTERS, AND I WAS UNABLE TO CLEAN THEM. ALSO, I HAD NOT RECEIVED ANY REPLACEMENT PARTS (HOSES, FILTERS, PRE-FILTERS, ETC.) ALTHOUGH I HAD MET THE MEDICARE 2 MONTH USE OF THE BI-PAP; IN COMPLIANCE. LINCARE APOLOGIZED STATING THAT THE SERIAL NUMBER HAD BEEN MISFILED WITH PHILIPS WHICH IS WHY I HAD NOT RECEIVED ANY REPLACEMENT PARTS. LINCARE STATED THAT THEY WERE "WORKING ON IT" AND THAT MY REPLACEMENT FILTERS, ETC. WOULD BE ARRIVING SHORTLY. I NEVER RECEIVED ANY PARTS BEFORE BEING ADMITTED TO THE HOSPITAL FOR WHAT TURNED OUT TO BE PNEUMONIA; NOT A SMOKER, NEVER HAD PNEUMONIA BEFORE THIS HOSPITALIZATION. NEITHER LINCARE NOR PHILIPS SAID ANYTHING ABOUT A RECALL TO ME. MY QUESTIONS TO MEDWATCH ARE: IS IT POSSIBLE TO BE DIAGNOSED WITH BACTERIAL PNEUMONIA AFTER HAVING INHALED RESPIRABLE-SIZED PE-PUR PARTICULATES FOR SEVERAL MONTHS? IF SO, IS IT POSSIBLE THAT I WILL BE DIAGNOSED WITH PNEUMONIA AGAIN? RESPIRONICS STATE THAT "WE REGRET THAT IT MAY TAKE SOME TIE TO REPLACE AFFECTED DEVICES. THE PLANNED REPAIR FOR THE AFFECTED DEVICES INVOLVES CERTAIN DESIGN CHANGES, WHICH IN SOME MARKETS MAY INCLUDE REVIEW AND/OR AUTHORIZATION BY THE RELEVANT REGULATORY AGENCIES. WE UNDERSTAND THE IMPACT OF THIS ISSUE AND WE SINCERELY REGRET THIS DISRUPTION". DO I HAVE THE RIGHT TO REFUSE THE REPAIR OR REPLACEMENT OF THE BI-PAP, AND DEMAND THAT THE MONEY PROVIDED BY MEDICARE BE REFUNDED IN FULL? ARE THERE OTHER TESTS THAT I SHOULD BE AWARE OF TO GET TO THE BOTTOM OF THIS? ANY AND ALL HELP WOULD BE GREATLY APPRECIATED. (B)(6). FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-09
MW5104137 | 2021-09-21 | 2021-09-16 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | SYSTEM ONE CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I HAVE ONE OF THE PHILLIPS-RESPIRONICS RECALLED CPAP MACHINES (SYSTEM ONE, SERIAL NO. (B)(4)). PRIOR TO WEDNESDAY (B)(6) IT HAD NOT SHOWN ANY OF THE SIGNS OF THE PROBLEMS LISTED IN THE RECALL NOTICES. IT IS NOW PRODUCING SMALL BLACK FLECKS OF MATERIAL THAT I FOUND IN THE NOSE PIECE THE AFTERNOON OF (B)(6). FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-09
MW5104138 | 2021-09-21 | 2021-02-01 | I, n, j, u, r, y | Voluntary report | CPAP MACHINE DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I'VE DEVELOPED A WORSENING COUGH FOR THE PAST SEVERAL MONTHS. I HAVE THE PHILIPS CPAP MACHINE THAT IS UNDER RECALL. I DEVELOPED A COUGH IN THE LAST FEW MONTHS. SOMETIMES I SMELL A ODOR INTERMITTENTLY WHEN BREATHING IN THE TUBE, HOWEVER, NOT SURE IF IT'S ANYTHING. I CALLED THE DME TO REPORT IT, BUT THEY SAY IT IS NO WAY TO TEST ANY MALFUNCTION. I TRIED USING A MOUTHGUARD AS AN ALTERNATIVE TO THE PHILIPS CPAP MACHINE, BUT MY CONTINUOUS OXYGEN MONITORING DEMONSTRATED LOW OXYGEN APNEA EPISODES. SUBSEQUENTLY, I CANNOT STOP USING MY CPAP MACHINE. I CANNOT REACH PHILIPS, BUT I DID REACH OUT TO MY DME COMPANY , OPTIMAL CARE /BIOSCRIPT AND THEY HAVE NOT RESPONDED YET. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-09
MW5104139 | 2021-09-21 | 2021-09-18 | I, n, j, u, r, y | Voluntary report | TRILOGY EVO | CONTINUOUS, VENTILATOR, HOME USE | PHILIPS / RESPIRONICS, |  |  |  | LIKELY DUE TO HOME TRILOGY EVO VENTILATOR PROVIDING EXCESS PEEP AND PIP ABOVE ENTERED PRESSURES (5-10 CM ABOVE SET PARAMETERS), PATIENT HAD BEEN HYPOXIC AT HOME, NEEDED TO GO TO ED AND TO BE ADMITTED. CHANGES ON CHEST XRAY (LARGER CYSTIC SPACE) THAT MAY BE DUE TO VENTILATOR MALFUNCTION. SHE NOW NEEDS TO BE SWITCHED TO A DIFFERENT TYPE OF VENT FOR HOME- WILL REQUIRE PARENT TEACHING, DAYS IN THE HOSPITAL. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5104140 | 2021-09-21 | 2021-05-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS ELECTRONICS / RESPIRONICS, INC. |  |  |  | NOTICED THAT SINUS INFECTION WAS DEVELOPING AND GETTING WORSE FROM ABOUT THIS TIME UNTIL I GOT A RECALL NOTICE ABOUT MY CPAP MACHINE. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | I |  | 2021-09
MW5104142 | 2021-09-21 | 2019-10-09 | I, n, j, u, r, y | Voluntary report | RESMED AIRCURVE 10 VAUTO CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESMED CORP. | DN 703 | R370-4123/1 |  | I WAS EXPERIENCING FACIAL AND EYE SWELLING; BACTERIAL EYE INFECTIONS, FATIGUE; SEVERE HEADACHE, SEVERE SORE THROAT; BODY PAIN TO THE POINT IT WAS PAINFUL TO MOVE; DIFFICULTY DRESSING MYSELF; GETTING IN/OUT OF BED; JUST MOVING IN GENERAL; THE FACIAL /EYE SWELLING WAS VERY SEVERE AND PROGRESSIVELY GETTING WORSE OVER A 2-YEAR PERIOD. I HAVE BEEN USING CPAP FOR 15 YEARS. DURING THAT TIME I HAVE EXPERIENCED A VARIETY OF SYMPTOMS THAT WERE DETERMINED "UNEXPLAINED". (B)(6) 2018 I WAS DIAGNOSED WITH A SINUS INFECTION; (B)(6) 2018 I STARTED GETTING EYE INFECTIONS. OVER THE NEXT SEVERAL MONTHS I WOULD EXPERIENCE SEVERE HEADACHE, SEVERE SORE THROAT, THEN MY EYE WOULD SWELL TO THE POINT I LOOKED LIKE I WAS IN A-BOXING MATCH AND LOST. I WOULD BE PUT ON ANTIBIOTIC AND THE SWELLING WOULD DECREASE BUT "SWELLING/PUFFINESS OF MY EYE-UPPER AND LOWER LIDS" AND MY FACE ALWAYS WAS SWOLLEN. DOCTORS WOULD TELL ME THAT THE BAGS UNDER MY EYES WERE "MY AGE" AND IT WAS AN ALLERGY. DURING THIS 2+ YEAR TIME PERIOD I WENT TO OVER 30 DOCTORS TRYING TO FIGURE OUT WHAT WAS CAUSING THIS PROBLEM; 12 EYE DOCTORS (3 OF THEM OCULAR PLASTIC SURGEONS), NEUROLOGIST, 4 ALLERGISTS; 4 ENT'S; RHEUMATOLOGIST; ENDOCRINOLOGISTS; CARDIOLOGISTS, PULMONOLOGIST, DERMATOLOGIST; INFECTIOUS DISEASE DOCTOR, 2 DENTISTS. THE SWELLING WAS GETTING WORSE THE FALL OF 2020. I HAD EXHAUSTED ALL MY EFFORTS. ALTHOUGH I DID FIND SCIENTIFIC ARTICLES REGARDING REACTIONS TO SILICONE WHICH IS WHAT THE MASK ARE MADE FROM AND ALSO SILICONE IS USED "INSIDE" THE RESMED CPAP MACHINES. IN (B)(6) 2020 I DECIDED TO STOP USING THE CPAP MACHINE BECAUSE THAT WAS THE ONE THING I DID NOT TRY AND NO DOCTOR WOULD ADMIT THAT COULD BE CAUSING THE PROBLEM. WITHIN WEEKS I COULD SEE THE SWELLING GOING FROM MY FACE AND EYES AND WITHIN A MONTH MY BODY PAIN WAS GONE. IT HAS TAKEN 6+ MONTHS FOR MY BODY TO RETURN TO NORMAL. I HAVE LOST WEIGHT BECAUSE I DON'T HAVE AIR PUMPED INTO MY BODY. ONE DOCTOR TOLD ME I WAS LOOKING LIKE A "CHIPMUNK" BECAUSE I HAD GOTTEN TO PUFFY. I WILL NEVE RETURN-TO USING CPAP AGAIN. IT HAS BEEN DESTRUCTIVE FOR ME. ALL THE SYMPTOMS IT WAS SUPPOSE TO CURE WERE EXACERBATED FOR ME. |  | Y | N | LAY USER/PATIENT |  | 2021-09
2518422-2021-04218 | 2021-09-21 | 2021-08-03 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103197) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP AN ALLERGY ATTACK. THE PATIENT WAS PRESCRIBED STEROIDS IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY RECEIVED A VOLUNTARY MEDWATCH (MW5103197) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP AN ALLERGY ATTACK. THE PATIENT WAS PRESCRIBED STEROIDS IN RESPONSE TO THE REPORTED EVENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL AND EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. THE HUMIDIFIER WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR EVALUATION. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04154 | 2021-09-21 | 2021-08-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2200X11B | DS2200X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE WITH A VENTILATOR'S TOUCH SCREEN. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S DISPLAY INTERFACE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04308 | 2021-09-21 | 2021-08-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S DISPLAY BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04271 | 2021-09-21 | 2021-08-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED A TEST STEP DURING TESTING. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04262 | 2021-09-21 | 2021-09-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE VENTILATOR FAILED A TEST STEP DURING TESTING. THE DEVICE'S SYSTEM BOARD, OXYGEN BLENDING MODULE BOARD AND OXYGEN BLENDING MODULE HARNESS WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04270 | 2021-09-21 | 2021-08-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED A TEST STEP DURING TESTING. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04272 | 2021-09-21 | 2021-08-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | GB2110X15B | GB2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED A TEST STEP DURING TESTING, A "SERVICE REQUIRED ALARM CONDITION OCCURRED AND AN INTERNAL BATTERY WAS FAULTY. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR WAS ALARMING FOR A SERVICE REQUIRED ALARM CONDITION, THE INTERNAL BATTERY WAS FAULTY AND THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE'S FIO2 SENSOR WAS FOUND TO BE LOOSE CAUSING THE SERVICE REQUIRED ALARM CONDTION AND WOULD HAVE CAUSED THE FAILED TEST STEP. THE DEVICE'S FIO2 SENSOR WAS REPLACED TO ADDRESS THE ISSUES. THERE WAS NO MALFUNCTION OF THE INTERNAL BATTERY NOTED. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04259 | 2021-09-21 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11 | DSX700T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING LITTLE WHITE PARTICLES IN FILTER RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. A CORRECTION TO B5 WAS MADE AND SHOULD BE THE MANUFACTURER RECEIVED INFORMATION ALLEGING LITTLE WHITE PARTICLES IN FILTER RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION G1 WAS CORRECTED WHICH WAS INCORRECT IN INITIAL REPORT. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04086 | 2021-09-21 | 2021-07-31 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP DEVICE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC |  |  |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102972) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A TRANSIENT ISCHEMIC ATTACK. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102972) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT PREVIOUSLY DEVELOP A TRANSIENT ISCHEMIC ATTACK. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04220 | 2021-09-21 | 2021-07-30 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103055) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A STROKE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED A VOLUNTARY MEDWATCH (MW5103055) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A STROKE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. NEW INFORMATION WAS PROVIDED TO SECTION D4. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04171 | 2021-09-21 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. SECTION B5 SHOULD BE PREVIOUSLY REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES IN TUBING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04166 | 2021-09-21 | 2021-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DIFFICULTY BREATHING/SHORT OF BREATH, CANNOT SLEEP WITHOUT DREAMSTATION, PARTICLES IN TANK RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. CORRECTED INFORMATION WAS PROVIDED IN H.6. SECTION (IN PREVIOUS MDR DIFFICULTY BREATHING WAS NOT UPDATED). | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04172 | 2021-09-21 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN PREVIOUS REPORT, SECTION B5 WAS INCORRECT, CORRECT B5 SHOULD BE - THE PATIENT ALLEGED VISUALIZATION OF BLACK PARTICLES. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL INSPECTION OF THE DEVICE COMPLETED BY THE MANUFACTURER AND FOUND INDENTS ON EITHER SIDE OF THE DEVICE. AN INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND DARK SPOTS OF CONTAMINATION ON ALL PARTS OF THE ENCLOSURE. THERE WERE SIGNS OF WATER INGRESS IN THE BLOWER BOX. A HUMIDIFIER WAS ALSO RETURNED TO THE MANUFACTURER ALONG WITH DEVICE. AN INTERNAL VISUAL INSPECTION OF THE HUMIDIFIER WAS COMPLETED BY THE MANUFACTURER AND FOUND MULTIPLE SPOTS OF CONTAMINATION INSIDE, CONTAMINATION ON THE HEATER PLATE AND ON THE DRY BOX SEAL. THERE WERE SIGNS OF WATER INGRESS ON THE HEATER PLATE ACCESS PORT DUE TO THE SCREWS HAVING CORROSION ON THEM. THERE WERE SIGNS OF THE PLASTIC DELAMINATING ON THE BACK OF THE FRONT ACCESS PANEL OF THE HUMIDIFIER. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND 32 ERRORS. THE MANUFACTURER CONCLUDES MULTIPLE CONTAMINATION FOUND WERE CONSISTENT WITH INDENTS ON EITHER SIDE OF THE DEVICE, DARK SPOTS OF CONTAMINATION ON ALL PARTS OF THE ENCLOSURE, WATER INGRESS IN THE BLOWER BOX, MULTIPLE CONTAMINATION IN THE HUMIDIFIER. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. IN THIS REPORT, SECTION D9, D10, G3, H3, H6 HAS BEEN UPDATED OR CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04181 | 2021-09-21 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN THE DEVICE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04174 | 2021-09-21 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURE INSPECTED EXTERNALLY OBSERVED THAT UNKNOWN WHITE, BROWN AND BLACK CONTAMINATION (DUST-LIKE), INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM, AT THE AIR INLET OF THE BLOWER BOX. LIGHT DUST-LIKE CONTAMINATION ON TOP OF THE BLOWER MOTOR AND TOP OF THE BLOWER BOX. LIGHT GRAY UNKNOWN FILM OF CONTAMINATION IN THE BOTTOM ENCLOSURE. POTENTIAL MINERAL DEPOSIT OR LIQUID SPOTS ON THE BOTTOM OF THE BLOWER MOTOR AND INSIDE OF THE BLOWER MOTOR, INDICATING POTENTIAL WATER INGRESS. BLACK CONTAMINATION, CONSISTENT WITH KERATIN, AT THE AIR OUTLET OF THE BLOWER BOX. BROWN LIQUID CONTAMINATION DEPOSITS AND POTENTIAL MINERAL DEPOSITS IN THE BOTTOM OF THE BLOWER BOX. TINY UNKNOWN BLACK (INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM), BROWN AND WHITE SPECS OF CONTAMINATION ON THE BOTTOM THE BLOWER BOX. THE MANUFACTURER USED A FITT (FOAM INTEGRITY TEST TOOL) AND WAS NOT ABLE TO CONFIRM THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION AND THERE IS NO VISIBLE DAMAGE OR FUNCTIONALITY FAILURES OF THE DEVICE. THE MANUFACTURER OBSERVED CONTAMINATION IN THE DEVICE AND ARE CONSISTENT WITH BEING FROM AN EXTERNAL SOURCE. SECTION H6 TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04191 | 2021-09-21 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP AUTOSV | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | DSX900T11C | DSX900T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING HEALTH DAMAGED, BEING VERY SICK, COUGH AND SICK EYES RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING HEALTH DAMAGED, BEING VERY SICK, COUGH AND SICK EYES RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HAS BEEN CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04188 | 2021-09-21 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04192 | 2021-09-21 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK PARTICLE OR BLACK SPECKS IN THE CHAMBER AND THROAT IRRITATION RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. CORRECTED INFORMATION PROVIDED IN SECTION H6 (HEALTH EFFECT CLINICAL CODE - THROAT IRRITATION/SORE THROAT WAS MISSED TO CAPTURE IN INITIAL REPORT, NOW IT WAS ADDED). SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04194 | 2021-09-21 | 2021-09-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04240 | 2021-09-21 | 2020-07-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | BLX500H15 | BLX500H15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED NEURO ENDOCRINE CANCER TUMORS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF SURGERY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED NEURO ENDOCRINE CANCER TUMORS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF SURGERY. SECTION G3 WAS CORRECTED TO REFLECT THE BAD WHEN NEW INFORMATION WAS RECEIVED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04177 | 2021-09-21 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11C | DSX400H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. DUST/DIRT CONTAMINATION INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM WAS OBSERVED THROUGHOUT DEVICE ENCLOSURE AND AIRPATH, SUGGESTING A SOURCE EXTERNAL TO THE DEVICE. MINERAL DEPOSITS AND EVIDENCE OF WATER INGRESS ON THE BLOWER. PIL CAN CONFIRM THAT THERE IS EVIDENCE SUPPORTING WATER LEAKING INTO THE MACHINE, WHICH SUGGESTS MISUSE (I.E. OVERFILLING WATER TANK WHILE WATER TANK IS IN HUMIDIFIER, MOVING DEVICE WITH WATER IN THE HUMIDIFIER). PIL IS ABLE TO CONFIRM THE PRESENCE OF CONTAMINATION IN THE AIRPATH. PIL CAN ALSO CONFIRM THE PRESENCE OF MINERAL DEPOSITS SUGGESTING THE USE OF NON-DISTILLED WATER. PIL FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04183 | 2021-09-21 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN PREVIOUS REPORT SECTION B5 WAS MENTIONED INCORRECTLY, CORRECT B5 SHOULD BE - THE PATIENT ALLEGED DIZZINESS, NAUSEA, SINUS OR NASAL ISSUES HEADACHES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL AND INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND WHITE UNKNOWN DUST LIKE CONTAMINATION AND HAIRS AT THE AIR INPUT OF THE BLOWER BOX, BLACK CONTAMINATION CONSISTENT WITH KERATIN AT THE AIR OUTLET OF THE BLOWER BOX. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND 1 ERRORS. THE MANUFACTURER CONCLUDES CONTAMINATION FOUND WERE CONSISTENT WITH WHITE UNKNOWN DUST LIKE CONTAMINATION AND HAIRS AND BLACK CONTAMINATION CONSISTENT WITH KERATIN ON THE BLOWER BOX. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. IN THIS REPORT, SECTION D9, G3, H3, H6 HAS BEEN UPDATED OR CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04179 | 2021-09-21 | 2021-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11W | DSX500T11W |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM.THERE WAS NO REPORT OF PATIENT HARM OR INJURY. AFTER THE SECOND ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, THE CUSTOMER STATED THAT THE DEVICE WILL NOT BE RETURNED. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. THE MANUFACTURER PREVIOUSLY NOT CAPTURED, VISUALIZATION OF PARTICLES IN SECTION B5 AND IT HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04125 | 2021-09-21 | 2021-08-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04165 | 2021-09-21 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX400H11 | DSX400H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE ONE ERROR FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. UNIT WAS SCRAPPED. SECTION D8, D9 AND H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04176 | 2021-09-21 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04180 | 2021-09-21 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY.THE PATIENT ALSO ALLEGED PARTICLES IN TUBING OR CHAMBER, YELLOW OR GOLD PARTICLE FLOTING IN THEIR TANK.THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNALLY AND INTERNALLY AND FOUND FOREIGN WHITE POWDERY SUBSTANCE WAS FOUND THROUGHOUT THE BLOWER MOTOR BOX, THE FINS OF THE BLOWER MOTOR AND OUTLET PORT OF THE BLOWER MOTOR ARE WHITE, SLIGHT BLACK CONTARMINATION AT THE BLOWER SEAL AND IT'S CONSISTANT WITH THE KERATIN CONTARMINATION. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND ONE ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION, BUT FOREIGN WHITE POWDERY SUBSTANCE WAS FOUND THROUGHOUT THE BLOWER MOTOR BOX, THE FINS OF THE BLOWER MOTOR AND OUTLET PORT OF THE BLOWER MOTOR ARE WHITE, SLIGHT BLACK CONTARMINATION AT THE BLOWER SEAL AND IT'S CONSISTANT WITH THE KERATIN CONTARMINATION WERE OBSERVED . | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04124 | 2021-09-21 | 2021-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | CAX500T12 | CAX500T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED VISUALIZATION OF PARTICLES. THE PATIENT HAS ALSO ALLEGED TO EXPERIENCE POOR BREATHING, EYE IRRITATION, EYE AND NASAL INFECTIONS, VERY DRY COUGH, HEADACHES, DIZZINESS, SEVERE FATIGUE, HYPERSOMNIA. THERE WAS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. AN INTERNAL VISUAL INSPECTION OF DEVICE WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND SCUFF MARK ON OUTSIDE OF BACK PANEL, RUST COLORED RING ON EXTERIOR PIPE, UNKNOWN RESIDUE ON INSIDE OF TOP PIECE, INSIDE OF BACK PANEL, AND INSIDE OF BOTTOM COVER AND UNIDENTIFIED BLACK MARK ON BLOWER. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERROR LOGGED. THE MANUFACTURER CONCLUDES THAT THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE MANUFACTURER ALSO CONFIRMED THAT THERE IS NO VISIBLE DAMAGE OR FUNCTIONALITY FAILURES OF THE DEVICE, WHICH SUGGEST THAT THE SOURCE OF CONTAMINATION WAS EXTERNAL TO THE DEVICE. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04126 | 2021-09-21 | 2021-08-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO AUTO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSG500S11 | DSG500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04127 | 2021-09-21 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF (B)(4). THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04173 | 2021-09-21 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION HUMID | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN PREVIOUS REPORT, SECTION B5 WAS INCORRECT. CORRECT B5 SHOULD BE - THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED PARTICLES IN THE DEVICE, NAUSEA, LIGHTHEADED. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND WATER INGRESS INSIDE THE P4 DC AND A DISCOLORATION AROUND THE P4. AN INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND DISCOLORATION AND UNKNOWN CONTAMINANT ON TOP ENCLOSURE, UI PANEL, ISO PORT PANEL, SIDE PANEL, BOTTOM EN- CLOSURE, DUST AND UNKNOWN CONTAMINANT OBSERVED BETWEEN BOTTOM ENCLOSURE AND SIDE PANEL, AND ON THE BOTTOM ENCLOSURE, LIQUID INGRESS AND MINERAL DEPOSIT WAS OBSERVED ON THE BLOWER AND THE BLOWER BOX. THE MANUFACTURER ALSO RECEIVED HUMIDIFIER AND AN INTERNAL INSPECTION COMPLETEDE BY THE MANUFACTURER AND FOUND AN UNKNOWN CONTAMINANT WAS OBSERVED ON THE HUMIDIFIER BOTTOM ASSEMBLY, HEATER PLATE, HEATER PLATE O-RING, DRY BOX SEAL, DRY BOX ASSEMBLY AND FLIP LID ENCLOSURE. THE MANUFACTURER FOUND EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND 7 ERRORS. THE MANUFACTURER CONCLUDES CONTAMINATION FOUND WERE CONSISTENT WITH UNKNOWN CONTAMINANT ON TOP ENCLOSURE, UI PANEL, ISO PORT PANEL, SIDE PANEL, BOTTOM EN- CLOSURE, DUST AND UNKNOWN CONTAMINANT OBSERVED BETWEEN BOTTOM ENCLOSURE AND SIDE PANEL, AND ON THE BOTTOM ENCLOSURE, HUMIDIFIER BOTTOM ASSEMBLY, HEATER PLATE, HEATER PLATE O-RING, DRY BOX SEAL, DRY BOX ASSEMBLY AND FLIP LID ENCLOSURE. THE MANUFACTURER ALSO CONCLUDE WATER INGRESS CONSISTENT WITH INSIDE THE P4 DC, BLOWER AND THE BLOWER BOX. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04168 | 2021-09-21 | 2021-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING/CHAMBER, DEVICE/POWER CORD OR SUPPLY MELTED, DIFFICULTY BREATHING/SHORT OF BREATH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS UPDATED REPORT WILL BE FILED. CORRECTED INFORMATION PROVIDED IN B5 AND H6, THE INFORMATION (THE MANUFACTURER RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING/CHAMBER, DEVICE/POWER CORD OR SUPPLY MELTED, DIFFICULTY BREATHING/SHORT OF BREATH RELATED TO A CPAP DEVICE) WAS NOT INCLUDED IN INITIAL REPORT. SECTION D1, D4, H4 AND H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04223 | 2021-09-21 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103131) ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DEVICE EMITTING GASES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT ALLEGED DIFFICULTY BREATHING, RESPIRATORY PROBLEM, SLEEP PROBLEM,NERVOUS PROBLEM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2 HAS UPDATED TO REFLECT AS PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE HAS BEEN UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04178 | 2021-09-21 | 2021-07-12 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE END USER INDICATED THAT OZONE OR ANOTHER UNAPPROVED CLEANING AND DISINFECTION METHOD WAS USED. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 06, 2023, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION PATIENT STATES, HAS USED SO CLEAN AND HAS SEEN PARTICLES IN WATER CHAMBER WANTS TO REPORT ALSO HAS SEEN DOCTOR FOR COUGHING AND SHORTNESS OF BREATH DOES HAVE ASTHMA. RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT DID NOT REPORT TO RECEIVE MEDICAL INTERVENTION. SECTION B1 WAS CORRECTED FOR BOTH ADVERSE EVENT AND PRODUCT PROBLEM. (PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION. TO SERIOUS INJURY. SECTION H6 WAS CORRECTED AND UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04249 | 2021-09-21 | 2021-08-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04302 | 2021-09-21 | 2021-08-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP PRO W/HUMID, LA | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | LAX600H15 | LAX600H15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED THAT A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM ALLEGEDLY BECAME DEGRADED AND THERE WAS NO PATIENT HARM OR INJURY. PATIENT ALSO ALLEGED HAVING HEADACHE, IRRITATION, INFLAMMATION, HYPERSENSITIVITY, NAUSEA AND VOMITING, THROAT IRRITATION AND SCRATCHING. THE DEVICE WAS RETURNED TO THE MANUFACTURER QUALITY PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. DURING THE INITIAL VISUAL INSPECTION OF THE DEVICE FOUND MINOR WEAR AND TEAR ON THE DEVICE. NO EVIDENCE OF DEGRADATION,CONTAMINATION, OR DAMAGE WAS OBSERVED. THE DEVICE WAS DISASSEMBLED FOR INTERNAL VISUAL INSPECTION. THE MANUFACTURER FOUND NO EVIDENCE OF ANY DEGRADATION OR DAMAGE OF THE SOUND ABATEMENT FOAM, AND ALMOST NO DIRT OR DUST IN THE PATIENT AIRWAY OR BLOWER BOX. THE MANUFACTURER OBSERVED LIQUID INGRESS TO THE BLOWER BOX. CARBONATE DEPOSIT WAS FOUND ON THE SURFACE OF THE BLOWER AND ON THE INTERIOR OF THE BLOWER BOX. THE WATER TANK ALSO HAS CARBONATE DEPOSIT ON THE INTERIOR BEVEL OF THE TANK PLATE, THOUGH IT APPEARS THAT MOST DEPOSIT HAS BEEN WASHED/CLEANED OFF. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND POWERED ON, PROVIDED AIRFLOW. THE MANUFACTURER REVIEWED THE DEVICE'S DOWNLOADED EVENT LOG AND FOUND 1 ERROR LOGGED, INSTANCE OF E-191 ERR_ALS_BIST, A CONTINUE ERROR. THE MANUFACTURER CONCLUDES THERE IS NO EVIDENCE OF FOAM DEGRADATION WITHIN THE DEVICE. THE CONTAMINATION FOUND WITHIN THEDEVICE IS CONSISTENT WITH BEING FROM AN EXTERNAL SOURCE | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04224 | 2021-09-21 | 2021-03-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102980) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A URINARY TRACT INFECTION. THE PATIENT WAS PRESCRIBED ANTIBIOTICS IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | NEW INFORMATION WAS PROVIDED TO SECTION(S) D1, D4, D10, AND G4. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04238 | 2021-09-21 | 2021-07-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11 | DSX700H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DEHYDRATION, PASSING KIDNEY STONES, SWOLLEN LIVER, COUGH , SINUS IRRITATION, DRY THROAT, NAUSEA, SWOLLEN LEGS RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. CORRECTED INFORMATION WAS PROVIDED IN SECTIONS B5, G1, H6. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04151 | 2021-09-21 | 2021-09-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04239 | 2021-09-21 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS INITIALLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE - THE PATIENT ALLEGED PARTICLES IN TUBING, NASAL AND THROAT IRRITATION AND SKIN RASHES. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. PIL IS UNABLE TO ASSESS THE SYMPTOMS DESCRIBED IN THE COMPLAINT. PIL CAN CONFIRM THE PRESENCE OF DIRT/DUST CONTAMINANTS IN THE AIR INLET. PIL CAN ALSO CONFIRM THE PRESENCE OF A GREY FILM CONTAMINANT ON THE FRONT PANEL, REAR PANEL, AND BOTTOM ENCLOSURE. PIL CAN ALSO CONFIRM THE PRESENCE OF A DARK CONTAMINANT AT THE BLOWER SEAL OF THE BLOWER BOX THAT APPEARS CONSISTENT WITH POTENTIAL KERATIN OBSERVED IN ER (B)(4)(V1). PIL FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION AND BREAKDOWN. THE MANUFACTURER CONCLUDES THE DEVICE HAS CONTAMINATION CONSISTENT WITH WATER INGRESS. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04098 | 2021-09-21 | 2021-08-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200H11C | DSX200H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DEVICE IS VERY NOISY AND WATER GETS VERY CLOUDY WHEN IN USE HAS TO CHANGE EVERY MORNING, IN WATER CHAMBER HAS A LOT OF PARTICLES INSIDE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04046 | 2021-09-21 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN THE INITIAL REPORT, SECTION B5 WAS NOT UPDATED, THE CORRECT B5 SHOULD BE - THE PATIENT HAS ALLEGED PARTICLES IN THE DEVICE, CONGESTION,BAD HEADACHES,BURNING SMELL. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE ALONG WITH A HUMIDIFIER WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF WHITE AND BLACK UNKNOWN CONTAMINATION (DUST LIKE), AND SOME VERY SMALL POTENTIAL HAIRS AT THE AIR INPUT OF THE BLOWER BOX, BLACK CONTAMINATION WITH KERATIN, AT THE AIR OUTLET OF THE BLOWER BOX, WHITE DUST-LIKE UNKNOWN CONTAMINATION, ON THE BOTTOM OF THE BLOWER BOX AND ON TOP AND ON THE BOTTOM OF THE BLOWER MOTOR, UNKNOWN BROWN AND BLACK CONTAMINANTS AND DARK-COLORED HAIR STRANDS INSIDE OF HUMIDIFIER BOX, UNKNOWN BROWN AND BLACK CONTAMINANTS AND A DARK COLORED HAIR ON AIR INLET TUBE SEAL OF HUMIDIFIER, UNKNOWN BLACK AND BROWN SUBSTANCE ON THE HUMIDIFIER LID SEAL NEAR THE OUTLET PORT. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERROR LOGGED. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH BLACK CONTAMINATION WITH KERATIN, AT THE AIR OUTLET OF THE BLOWER BOX AND WHITE AND BLACK UNKNOWN CONTAMINATION (DUST LIKE), AND SOME VERY SMALL POTENTIAL HAIRS AT THE AIR INPUT OF THE BLOWER BOX, WHITE DUST-LIKE UNKNOWN CONTAMINATION, ON THE BOTTOM OF THE BLOWER BOX AND ON TOP AND ON THE BOTTOM OF THE BLOWER MOTOR, UNKNOWN BROWN AND BLACK CONTAMINANTS AND DARK-COLORED HAIR STRANDS INSIDE OF HUMIDIFIER BOX, UNKNOWN BROWN AND BLACK CONTAMINANTS AND A DARK COLORED HAIR ON AIR INLET TUBE SEAL OF HUMIDIFIER, UNKNOWN BLACK AND BROWN SUBSTANCE ON THE HUMIDIFIER LID SEAL NEAR THE OUTLET PORT WERE INCONSISTENT WITH THE SOUND ABATEMENT FOAM. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION D9,G3,H6 HAS BEEN UPDATED/CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04099 | 2021-09-21 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11 | DSX400H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN NOSE AND WATER, HEADACHES, NASAL AND THROAT IRRITATION RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04133 | 2021-09-21 | 2021-07-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ORAL CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY SUBMITTED MDR 2518422-2021-04133-2 WITH INCORRECT INFORMATION IN SECTIONS B1, B2, H1, H6. CORRECTIONS TO PREVIOUS MDR ARE MADE IN THIS REPORT AS BELOW. SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY PRODUCT PROBLEM WAS CHECKED IN PREVIOUS MDR) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY IT WAS BLANK) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6- HEALTH IMPACT CODE WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04246 | 2021-09-21 | 2021-09-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF BLACK PARTICLES; THERMAL ISSUES; DIFFICULTY BREATHING/SHORT OF BREATH; IRRITATION AROUND MOUTH AND NOSE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. THE MANUFACTURER PREVIOUSLY NOT CAPTURED SECTION E1- EMAIL ADDRESS AND IT HAS BEEN UPDATED IN THIS REPORT. SECTION H6 HEALTH EFFECT- CLINICAL CODE, MEDICAL DEVICE PROBLEM CODE HAS BEEN CORRECTED AND TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS WERE UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04248 | 2021-09-21 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP ST30 | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | DSX1030H11 | DSX1030H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING CHAMBER RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04244 | 2021-09-21 | 2021-09-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04245 | 2021-09-21 | 2021-09-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04241 | 2021-09-21 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP ST | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS INC. | DSX1030T11C | DSX1030T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED THAT A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM ALLEGEDLY BECAME DEGRADED AND THERE WAS NO PATIENT HARM OR INJURY. THE PATIENT ALSO ALLEGED HARD TO BREATH, HAS TO USE INHALER TO BREATHE. THE DEVICE WAS RETURNED TO THE MANUFACTURER QUALITY PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. DURING THE INITIAL VISUAL INSPECTION OF THE DEVICE FOUND MINOR CORROSION ON THE HUMIDIFIER BASE SCREWS AND WHAT APPEARS TO BE CARBONATE MINERAL DEPOSITS IN THE WATER TANK. THE DEVICE WAS DISASSEMBLED FOR INTERNAL VISUAL INSPECTION. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION, WITH THE FOAM APPEARING INTACT. THE MANUFACTURER APPLIED POWER TO THE DEVICE. UNIT POWERED ON AND PROVIDED AIRFLOW. THE MANUFACTURER REVIEWED THE DEVICE'S DOWNLOADED EVENT LOG AND FOUND 1 ERROR, E-36 ERR_MOTOR_BLOCKED _OUTLET LOGGED. THE MANUFACTURER WAS UNABLE TO CONFIRM THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM. THE SOUND ABATEMENT FOAM HAD NO EVIDENCE OF DEGRADATION. THE MANUFACTURER FOUND THE PRESENCE OF MINOR CORROSION AND CARBONATE MINERAL DEPOSITS IN THE WATER TANK, SUGGESTING NON-DISTILLED WATER WAS BEING USED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04242 | 2021-09-21 | 2021-08-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | INITIAL FINAL REPORTS (CUSTOMER ALLEGATIONS): THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED TO SEEING THE YELLOW SPOT IN RESERVIOR. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04243 | 2021-09-21 | 2021-09-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04273 | 2021-09-21 | 2021-07-07 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION HUMID | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSXHCP | DSXHCP |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102990) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ACUTE ASTHMA. THE PATIENT WAS PRESCRIBED ANTIBIOTICS AND BREATHING TREATMENTS IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED A VOLUNTARY MEDWATCH (MW5102990) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ACUTE ASTHMA. THE PATIENT WAS PRESCRIBED ANTIBIOTICS AND BREATHING TREATMENTS IN RESPONSE TO THE REPORTED EVENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04195 | 2021-09-21 | 2021-09-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04226 | 2021-09-21 | 2021-09-10 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, GER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DEX500S13 | DEX500S13 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP SINUSITIS, ASTHMA AND PRESSURE IN THE CHEST. THE PATIENT HAD SURGERY PERFORMED ON NASAL CAVITY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04274 | 2021-09-21 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103082) ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A SINUS INFECTION. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON 11 NOV, 2024 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103082) ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A SINUS INFECTION, HEADACHES, THROAT PAIN AND OUS IN THROAT. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECTS: CLINICAL CODES AND HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04304 | 2021-09-21 | 2021-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF A VERY LIGHT AMOUNT OF LIGHT BEIGE DUST OR DIRT CONTAMINATE IN THE AIR INTAKE CANAL, ON THE INSIDE OF THE BLOWER BOX, ON THE INSIDE SURFACE OF THE BOTTOM ENCLOSURE, AND IN FRONT AND REAR PANEL CHANNELS, TWO DARK SPEC CONTAMINATES, AND ONE BROWN SPEC CONTAMINATE LOCATED IN FRONT PANEL CHANNEL, A VERY MINIMAL AMOUNT OF POTENTIAL KERATIN AROUND THE BLOWER BOX INNER SURFACE OF THE OUTLET PORT. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. THE HUMIDIFIER WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR EVALUATION. THE MANUFACTURER DID NOT ADDRESSED HUMIDIFIER DUE INSECT INFESTATION. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND ONE INSTANCE OF E-106. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF MULTIPLE CONTAMINANTS. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04301 | 2021-09-21 | 2021-08-03 | I, n, j, u, r, y | Manufacturer report | PHILIPS DREAMSTATION C-PAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103028) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO REQUIRE OXYGEN AND OTHER TREATMENTS. THE PATIENT WAS TRANSPORTED TO THE EMERGENCY ROOM IN RESPONSE TO THE REPORTED EVENT. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE - THE PATIENT ALLEGED OF SHALLOW BREATHING,SHORTNESS OF BREATH AND THE INABILITY TO INHALE SUFFICIENTLY.THE PATIENT WAS REQUIRED AN AMBULANCE TRANSPORT TO HOSPITAL IN THE MIDDLE OF THE NIGHT.AND ALSO CAUSED A PATIENT TO REQUIRE OXYGEN AND OTHER TREATMENTS. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL INVESTIGATION FOUND UNKNOWN WHITE DEPOSITES IN THE FILTER INLET, A MISSING LEFT FLAP AND AN ODOR THAT WAS EMITTED DURING THERAPY. ALSO DID FIND THE PRESENCE OF DEGRADED SOUND AMBATEMENT FOAM. REVIEWED DREAMSTATION ERROR LOG AND USAGE HISTORY.THE DEGRADED FOAM WAS LOCALIZED AND NOT FOUND THROUGHOUT THE AIRPATH. UNKNOWN DEBRIS CONSISTENT WITH DUST AND SMALL HAIRS WERE FOUND INSIDE THE BOTTOM OF DEVICE, UNKNOWN WHITE DEBRIS CONSISTENT WITH DUST FOUND ON TOP OF BLOWER, AND UNKNOWN RESIDUE CONSISTENT WITH WATER INGRESS WAS FOUND ON THE INSIDE OF BLOWER BOX AND ON THE BOTTOM OF BLOWER. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERORRS. THE MANUFACTURER INTERNAL INVESTIGATION AND DEVICE DISASSEMBLY REVEALED THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM. SECTION D4 HAS BEEN UPDATED IN THIS REPORT. SECTION H6 HEALTH EFFECT- CLINICAL CODE, MEDICAL DEVICE PROBLEM CODE,TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04303 | 2021-09-21 | 2021-08-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, JP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | RJPX500S16 | RJPX500S16 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION D9 WAS MISSED IN THE PREVIOUS REPORT, WHICH IS UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
MW5104144 | 2021-09-21 | 2021-01-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | REPORTER STATES THAT HE HAS BEEN HAVING HEADACHES, FATIGUE, COUGHING, PAIN, PAIN IN LEGS, AND NAUSEA. THE FIRST DREAMSTATION BROKE AND HE GOT A NEW ONE TO REPLACE IT. NOW, IT'S BEEN RECALLED. REPORTER STATES THAT HE HAD BEEN FEELING TIRED FOR SOME TIME AND THAT THE HEADACHES STARTED TWO WEEKS AGO. |  | Y | Y | I |  | 2021-09
2518422-2021-03915 | 2021-09-22 | 2021-08-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03917 | 2021-09-22 | 2021-08-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04282 | 2021-09-22 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 & H9 UPDATED IN THIS REPORT. ( IN INITIAL MDR H9 WAS MISSED TO UPDATE). | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04014 | 2021-09-22 | 2021-09-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04390 | 2021-09-22 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | SECTION H6 HEALTH EFFECT - CLINICAL CODE & SECTION H10 CAPTURED INCORRETLY IN PREVIOUS REPORTS, IT SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED HAVING ISSUES WITH BREATHING AND PARTICLES IN THE LINE. HE ALSO STATES THE AIRFLOW HAS GOTTEN WORSE AS WELL. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04284 | 2021-09-22 | 2021-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPORINOC INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | A USER CONTACTED THE MANUFACTURER REGARDING THE SOUND ABATEMENT FOAM CORRECTION/REMOVAL. THE PATIENT ALLEGED THE DEVICE SHUTS DOWN AND IS NOISY. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. DURING THE INVESTIGATION, EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN WAS OBSERVED IN THE BASE UNIT. BLACK SUBSTANCE, CONSISTENT WITH SOUND ABATEMENT FOAM DEGRADATION, WAS OBSERVED ON THE OUTSIDE BOTTOM OF THE BLOWER BOX AND WAS STUCK TO IT. MANUFACTURER WAS ABLE TO CONFIRM THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM.SECONDARY FINDIGS ARE OBSERVED THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS. THE MANUFACTURER CONCLUDES THE PRESENCE OF MULTIPLE CONTAMINANTS THAT ARE NOT CONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM FROM THE SECONDARY FINDINGS MANUFACTURE WAS ABLE TO CONFIRM THE COMPLAINT BUT WAS UNABLE TO ADDRESS THE SYMPTOMS SPECIFIED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04330 | 2021-09-22 | 2021-08-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2200X11B | DS2200X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS EVALUATED BY A FIELD SERVICE ENGINEER AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICES SYSTEM BOARD AND BLOWER NEED REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04283 | 2021-09-22 | 2021-07-13 | I, n, j, u, r, y | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A BACTERIAL INFECTION IN HER LUNGS. THE PATIENT WAS ADMITTED TO THE HOSPITAL. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A BACTERIAL INFECTION IN HER LUNGS. THE PATIENT WAS ADMITTED TO THE HOSPITAL. AFTER RECEIVING FURTHER INFORMATION FROM THE PATIENT IT IS DETERMINED THAT THE INITIAL REPORT SHOULD HAVE BEEN FILED AS ADVERSE EVENT ONLY. SECTION B1 AND B2 HAS BEEN UPDATED/CORRECTED IN THIS REPORT. | Y | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04279 | 2021-09-22 | 2021-01-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500S11 | DSX500S11 |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102722) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A COUGH. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON OCT 25, 2024, AND SECTION H11 SHOULD BE REPORTED AS THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF BLACK SPECS FLOATING IN THE HUMIDIFIER TANK AND IN THE WATER RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT ALLEGED TO DEVELOP COUGH, DIZZINESS, HEADACHE, DIFFICULTY IN BREATHING AND WAS PRESCRIBED MEDICATION IN RESPONSE TO THE COUGH. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM (ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR). SECTION B2 WAS CHANGED TO BLANK (PREVIOUSLY IT WAS OTHER SERIOUS). SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04333 | 2021-09-22 | 2021-08-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TO POWER ON. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED THE SUBMISSION DATE ENTERED IN SECTION B4 AS (B)(6) 2021. THE CORRECT SUBMISSION DATE IS (B)(6) 2021. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04331 | 2021-09-22 | 2021-08-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S INTERNAL BATTERY WAS FAULTY. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03783 | 2021-09-22 | 2021-08-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE LCD SCREEN WAS OBSERVED TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04335 | 2021-09-22 | 2021-08-30 | I, n, j, u, r, y | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2200X11B | DS2200X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR INOPERATIVE ALARM CONDITION OCCURRED AND THE CUSTOMER REPLACED THE SYSTEM BOARD AND BLOWER MOTOR PRIOR TO RETURNING FOR EVALUATION. THE REPORT HAS BEEN AMENDED TO REFLECT AN ALLEGATION OF PATIENT HARM REQUIRING INTERVENTION. A REASSESSMENT OF THE MEDICAL INTERVENTION ALLEGATION WAS REVIEWED BY THE CLINICAL EXPERT. IT WAS DETERMINED THAT BASED ON THE AVAILABLE INFORMATION, THERE IS INSUFFICIENT EVIDENCE TO PRONOUNSE THE EVENT OF NEEDING TO SWITCH TO ANOTHER DEVICE AS SERIOUS INJURY AS DEFINED IN 21 CFR 803.3(W). | Y | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04334 | 2021-09-22 | 2021-08-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLANK. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR'S LCD SCREEN WAS BLANK. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04307 | 2021-09-22 | 2021-08-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | IN2100X15B | IN2100X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD, MACHINE FLOW SENSOR NEED REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03782 | 2021-09-22 | 2021-08-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. . |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04013 | 2021-09-22 | 2021-09-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | IN2200X15B | IN2200X15B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE ASSEMBLY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03784 | 2021-09-22 | 2021-09-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BLOWER MOTOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04012 | 2021-09-22 | 2021-09-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE'S LCD SCREEN WAS OBSERVED TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04011 | 2021-09-22 | 2021-09-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | DS2200X11B | DS2200X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE ASSEMBLY WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED THAT THE DEVICE'S OXYGEN BLENDING MODULE ASSEMBLY WAS REPLACED TO ADDRESS A "SERVICE REQUIRED" CODE IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THIS DID NOT CORRECT THE ISSUE AND THE SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04332 | 2021-09-22 | 2021-08-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED A TEST STEP DURING TESTING. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04316 | 2021-09-22 | 2021-07-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP HEADACHES. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04107 | 2021-09-22 | 2021-07-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID/HT, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | AUX400T15 | AUX400T15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04276 | 2021-09-22 | 2018-10-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5102870) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG CANCER. THE PATIENT HAD SURGERY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY SUBMITTED MDR 2518422-2021-04276-1 WITH INCORRECT SECTIONS B1, B2, H1, H6. CORRECTIONS TO PREVIOUS MDR WERE MADE IN THIS REPORT AS FOLLOWS. SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY PRODUCT PROBLEM WAS CHECKED IN PREVIOUS MDR). SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY IT WAS BLANK). SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 - HEALTH EFFECT - IMPACT CODE WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04197 | 2021-09-22 | 2021-09-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11 | DSX700T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04201 | 2021-09-22 | 2021-09-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04200 | 2021-09-22 | 2021-09-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04106 | 2021-09-22 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX200T11 | DSX200T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO EXPERIENCE A CANCER. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENT OF CANCER AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04318 | 2021-09-22 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS | DSX200H11 | DSX200H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DIFFICULTY BREATHING, COUGHING BLACK MUCUS EVERY MORNING BACK IN MARCH WHICH IS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
MW5104149 | 2021-09-22 | 2018-04-15 | I, n, j, u, r, y | Voluntary report | PHILIPS CPAP DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC | DSX500T11W | DSX500H11 |  | DIZZINESS, LESS COGNITIVE, PHILIPS DREAMSTATION CAUSED ME SEVERAL SEVERE MEDICAL ISSUES. I AM A (B)(6) WOMAN, WITH MINOR HEALTH ISSUES. I WAS USING PHILIPS CPAP DREAMSTATION SINCE (B)(6) 2018, (B)(6) 2018 I BEGAN HAVING A RUNNY NOSE EVERY DAY, GUMS WERE SORE AND BLEEDING, COULD NOT TAKE A DEEP BREATH, I WAS VERY RESTRICTED IN BREATHING. MY EARS HAVE BEEN IRRITATED AND BOTHERING ME CONSTANTLY. I GET FREQUENT SINUS INFECTIONS. I WAS CONCERNED AND MENTIONED IT TO MY PRIMARY DURING MY ANNUAL PHYSICAL, MY LACK OF BEING ABLE TO TAKE A BREATH, LIKE HAVING SEVERE PNEUMONIA. BEGINNING IN (B)(6) 2020 I STARTED TO HAVE A HORSE VOICE EVERY DAY, ALL DAY AND IT ONLY GOT WORSE, A LUMP ON MY NECK THAT WOULD NOT GO AWAY AND CHANGED IN SIZE, I WAS VERY DIZZY EVERY DAY AND WHEN I GOT UP I FELL OVER OR INTO FURNITURE, BED SPINS, ETC, HEADACHES, SKIN RASHES, ACID REFLUX WAS EVERY DAY, I WAS BURPING ALL THE TIME AND I NEVER BURPED BEFORE, I MEAN NEVER, LESS COGNITIVE (WAS SPACY COULD NOT REMEMBER SIMPLE TASKS, USUALLY I WAKE UP AND PLAN MY DAY AND TASKS, I COULD DO NONE OF IT, I WAS BLANK, FORGOT PASSWORDS, TASKS, WHAT I HAD PREVIOUSLY ACCOMPLISHED). I WAS DEPRESSED OVER HOW STUPID I HAD BECOME. I STOPPED USING DREAMSTATION (B)(6) 2021 IT IS NOW (B)(6) 2021, MY NOSE IS NOT RUNNING EVERY DAY, I AM NOT DIZZY, WHEN I GET UP I DO NOT FALL OVER, OR HAVE TROUBLE NAVIGATING WALKING TO THE BATHROOM. HEADACHES ARE BETTER, MY BRAIN IS CLEAR, I CAN REMEMBER TASKS, PASSWORDS. I FEEL SO MUCH BETTER AND I CAN FUNCTION. FOR THE LAST YEAR AND 1/2 I WAS SO UPSET ABOUT MY REDUCED CAPACITIES AND ABILITIES TO ACCOMPLISH SIMPLE LIFE TASKS OR TO FUNCTION AS A NORMAL PERSON. I AM HAPPY TO BE FEELING SO MUCH BETTER. I BELIEVE THE CPAP DREAMSTATION CAUSED THE ISSUES I WAS HAVING, ESPECIALLY DUE TO THE FACT THAT A MONTH AFTER STOPPING THE USE OF THE DREAMSTATION THERE IS A NOTICEABLE DIFFERENCE IN ALL THE AREAS I MENTIONED ABOVE. DREAMSTATION SERIAL # (B)(4) AND # (B)(4), I HAD 2 OF THEM. (B)(6). I HAVE 2 DREAMSTATIONS I WAS USING: DREAMSTATION AUTO CPAP HUM DOM (B)(4) DSX500H11 2017-12-18 REV00, DREAMSTNAUTOCPAPHUMHTWIFI DOM (B)(4) 2017-01-05 REV00. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | I |  | 2021-09
2518422-2021-03916 | 2021-09-22 | 2021-08-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04189 | 2021-09-22 | 2020-09-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE HEADACHES, PNEUMONIA AND NODULES IN THROAT AND LUNGS. THE PATIENT WAS HOSPITALIZED AND DID RECEIVE MEDICAL INTERVENTION. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING TO EXPERIENCE HEADACHES, PNEUMONIA AND NODULES IN THROAT AND LUNGS,DIFFICULTY BREATHING,RESPIRATORY ISSUES RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM.THE PATIENT WAS HOSPITALIZED AND DID RECEIVE MEDICAL INTERVENTION. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HEALTH EFFECT-CLINICAL CODE WAS MISSED TO CAPTURE, NOW IT IS CORRECTED AND UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03914 | 2021-09-22 | 2021-08-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04266 | 2021-09-22 | 2021-09-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | RIN2110X15B | RIN2110X15B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE ALARMED FOR A BACKLIGHT FAILURE ISSUE. THE DEVICE'S DISPLAY BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04267 | 2021-09-22 | 2021-09-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1032804 | 1032804 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE VENTILATOR'S LCD SCREEN WAS FOUND TO BE BLANK. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04268 | 2021-09-22 | 2021-09-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2110X11B | DS2110X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04265 | 2021-09-22 | 2021-09-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04264 | 2021-09-22 | 2021-09-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04263 | 2021-09-22 | 2021-09-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S DETACHABLE BATTERY CABLE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04338 | 2021-09-22 | 2021-08-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO BIFLEX, SYS ONE 60 SRS, INT | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | IN761S | IN761S |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03686 | 2021-09-22 | 2021-07-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BLOWER NEEDS REPLACED TO ADDRESS THE ISSUE. THE DEVICE WAS RETURNED UNREPAIRED PER THE CUSTOMER REQUEST. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04199 | 2021-09-22 | 2021-09-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04202 | 2021-09-22 | 2021-09-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04196 | 2021-09-22 | 2021-09-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04093 | 2021-09-22 | 2021-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ASTHMA. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO EXPERIENCE A ASTHMA. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENT OF ASTHMA AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION AND AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTIONS B1, B2 HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04092 | 2021-09-22 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP BRONCHITIS. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED ACUTE BRONCHITIS AND ASSOCIATED COUGH AND WHEEZING. THE PATIENT ALSO ALLEGED THAT THERE WAS A "WHITE MOLD" ON THE PLATE AROUND THE WATER CONTAINER. THE REPORTED EVENT OF BRONCHITIS AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. THE PATIENT HAS SWAPPED OUT THE DEICE AT THE DISTRIBUTOR AND AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION, FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04198 | 2021-09-22 | 2021-09-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04250 | 2021-09-22 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP AUTOSV | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | DSX900H11 | DSX900H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING BLACK RESIDUE IN THE BOTTOM OF THE WATER TANK ON THE ALUMINIUM AND ALSO STATES RESIDUE IS NOT PARTICLES IT'S MORE SOLID AND COVERS AN AREA, RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04311 | 2021-09-22 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING BLACK PARTICLES IN THE WATER CHAMBER AND COUGH AND THROAT IRRITATION RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE - THE PATIENT ALLEGED BLACK PARTICLES IN THE WATER CHAMBER AND COUGH AND THROAT IRRITATION RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04277 | 2021-09-22 | 2021-08-26 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG AND HEART ISSUES. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED ON THIS DEVICE IN MDR 2518422-2021-03833. THIS REPORT WAS SUBMITTED AS A DUPLICATE REPORT OF THE PREVIOUSLY SUBMITTED REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04310 | 2021-09-22 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP ST30 | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | DSX1030H11C | DSX1030H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 WAS CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO HEADACHE; SORE THROAT; DRY MOUTH; BRAIN FOG-TROUBLE FOCUSING/DOING EVERY DAY TASKS; EXTREME FATIGUE;DIFFICULTY BREATHING,DEVICE HAS BURNT SMELL AND PARTICLES IN THE CHAMBER. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENTS OF HEADACHE; SORE THROAT; DRY MOUTH; BRAIN FOG-TROUBLE FOCUSING/DOING EVERY DAY TASKS; EXTREME FATIGUE; DIFFICULTY BREATHING AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENTS ARE ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. PIL OBSERVED THE FOLLOWING FROM INTERNAL AND EXTERNAL INSPECTION:OUTSIDE IS IN GOOD CONDITION,DUST IN FILTER INLET,WHITE DEPOSITS IN AIR OUTLET AND AIR INLET,WHITE DUST ON TOP OF OUTSIDE OF PLASTIC BLOWER BOX, ON TOP OF BLOWER, AND INSIDE OF BOTTOM OF BLOWER BOX, WATER MARKS AND DUST ON BOTTOM OF BLOWER. PIL USED A FITT (FOAM INTEGRITY TEST TOOL) AND THE ASSOCIATED PROCEDURE ER 2245460 V01 AND WAS NOT ABLE TO CONFIRM THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM. THERE IS NO VISIBLE DAMAGE OR FUNCTIONALITY FAILURES OF THE DEVICE, WHICH SUGGEST THAT THE SOURCE OF CONTAMINATION WAS EXTERNAL TO THE DEVICE. PIL WAS NOT ABLE TO CONFIRM THE COMPLAINT OR ADDRESS THE SYMPTOMS SPECIFIED. PIL WAS NOT ABLE TO GET CARE ORCHESTRATOR INFORMATION FROM DEVICE. REVIEW OF THE DEVICE LOG SHOWED: ERRORS LOGGED: 0 ERRORS WERE LOGGED.DEVICE WAS LIKELY RESET PRIOR TO PIL RECEIPT AS INDICATED BY THE DEVICE HAVING 6970.8 MACHINE HOURS, 0 BLOWER HOURS AND 0 THERAPY HOURS. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04341 | 2021-09-22 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP BRONCHITIS AND PNEUMONIA. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO EXPERIENCE NASAL / THROAT IRRITATION OR SORENESS, BRONCHITIS, AND PNEUMONIA. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENTS OF NASAL / THROAT IRRITATION OR SORENESS, BRONCHITIS, PNEUMONIA AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THESE EVENTS ARE ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION AND AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTIONS B1, B2 HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- CLINICL CODE CORRECTED. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04305 | 2021-09-22 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP AUTO SV | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | DSX900T11C | DSX900T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT BLACK/WHITE POWDERISH PARTICLES IN HER WATER RESERVOIR OVER A YEAR AGO RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, FOLLOWUP REPORT WILL BE FILED. SECTION G1 WAS CORRECTED AND H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04309 | 2021-09-22 | 2021-09-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04350 | 2021-09-22 | 2021-08-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | CAX500T12 | CAX500T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED PARTICLES IN TUBING/CHAMBER, NASAL/THROAT IRRITATION OR SORENESS. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04349 | 2021-09-22 | 2021-07-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | AUX400T15 | AUX400T15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04351 | 2021-09-22 | 2021-09-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | AUX400T15 | AUX400T15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04187 | 2021-09-22 | 2021-08-03 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE HEADACHES, TROUBLE BREATHING AND FLUID IN ONE LUNG. THE PATIENT WAS HOSPITALIZED AND DID RECEIVE MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 08, 2024, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED THE PATIENT TO HEADACHES, TROUBLE BREATHING AND FLUID IN ONE LUNG, THE PATIENT DID NOT REPORT TO RECEIVE MEDICAL INTERVENTION. SECTION B1 WAS CORRECTED FOR BOTH ADVERSE EVENT AND PRODUCT PROBLEM. ( PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION. TO SERIOUS INJURY. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04190 | 2021-09-22 | 2021-07-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11W | DSX500H11W |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO BE DIAGNOSED WITH CANCER. THE PATIENT DID NOT REPORT RECEIVING MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO BE DIAGNOSED WITH CANCER. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURER CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REPORTED BY THE PATIENT. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04186 | 2021-09-22 | 2021-07-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO BE DIAGNOSED WITH LUNG CANCER. THE PATIENT DID NOT REPORT RECEIVING MEDICAL INTERVENTION.. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DIAGNOSED WITH LUNG CANCER. THE PATIENT DID NOT REPORT RECEIVING MEDICAL INTERVENTION. THE REPORTED EVENT AND ITS SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO OTHER MEDICAL INTERVENTION/TREATMENT OR ADDITIONAL INFORMATION WAS REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-02338 | 2021-09-22 |  | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC.. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER BECAME AWARE THAT A USER ALLEGES TOOTH LOSS DUE TO HIS HUMIDIFIER NOT HEATING WITH HIS CPAP. THE HUMIDIFIER DOES NOT SHOW CONNECTED TO HIS DEVICE. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER BECAME AWARE THAT A USER ALLEGES TOOTH LOSS DUE TO HIS HUMIDIFIER NOT HEATING WITH HIS CPAP. THE HUMIDIFIER DOES NOT SHOW CONNECTED TO HIS DEVICE. EVEN THOUGH THE HUMIDIFIER BASE DID NOT HEAT, THE WATER IN THE HUMIDIFIER WOULD ACT AS A PASSOVER HUMIDIFIER. THE MANUFACTURER HAS NOT RECEIVED THE DEVICES FOR INVESTIGATION. DESPITE MULTIPLE REQUESTS, NO DEVICES HAVE BEEN RETURNED. IF FURTHER INFORMATION BECOMES AVAILABLE, AN ADDITIONAL REPORT WILL BE FILED. IF FURTHER INFORMATION BECOMES AVAILABLE, AN ADDITIONAL REPORT WILL BE FILED. THE MANUFACTURER CONCLUDES NO FURTHER ACTION IS NEEDED AT THIS TIME. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-03687 | 2021-09-23 | 2021-08-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "SERVICE REQUIRED' ERROR CODE RELATED TO HIGH TEMPERATURE. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OUTLET FLOW PATH THERMISTOR AND SENSOR BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04323 | 2021-09-23 | 2021-08-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2110X15B | IN2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLANK. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR'S LCD SCREEN WAS BLANK. THE DEVICE WAS NOT IN PATIENT USE. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04324 | 2021-09-23 | 2021-08-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TO CHARGE ITS INTERNAL BATTERY. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR FAILED TO CHARGE ITS INTERNAL BATTERY. THERE WAS NO HARM OR INJURY REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04328 | 2021-09-23 | 2021-08-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION OCCURRED RELATED TO THE OXYGEN BLENDING MODULE. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION OCCURRED RELATED TO THE OXYGEN BLENDING MODULE. THE DEVICE WAS NOT IN PATIENT USE. THERE WAS AN ADDITIONAL SERVICE REQUIRED ALARM CONDITION ALLEGED RELATED TO A SYSTEM FAILURE. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND SENSOR BOARD NEED REPLACED TO ADDRESS THE ISSUES. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04327 | 2021-09-23 | 2021-08-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION OCCURRED RELATED TO THE OXYGEN BLENDING MODULE. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION OCCURRED RELATED TO THE OXYGEN BLENDING MODULE. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04329 | 2021-09-23 | 2021-08-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2200X11B | DS2200X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A MALFUNCTION OF A VENTILATOR'S TOUCHSCREEN. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE WAS RESTARTED BY THE END USER TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04326 | 2021-09-23 | 2021-08-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION OCCURRED RELATED TO THE OXYGEN BLENDING MODULE. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION OCCURRED RELATED TO THE OXYGEN BLENDING MODULE. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION A MALFUNCTION OF THE ACITVE EXHALATION CONTROL MODULE WAS OBSERVED. THE ACTIVE EXHALATION CONTROL MODULE NEEDS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04336 | 2021-09-23 | 2021-08-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A HIGH TEMPERATURE ALARM CONDITION WAS OBSERVED IN THE DEVICE'S DOWNLOADED EVENT LOG. THE DEVICE'S OUTLET FLOWPATH THERMISTOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04352 | 2021-09-23 | 2021-08-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040000 | 1040000 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS EXPERIENCING A BOARD FAILURE. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS EXPERIENCING A BOARD FAILURE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SENSOR BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04280 | 2021-09-23 | 2021-08-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04325 | 2021-09-23 | 2021-08-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION OCCURRED RELATED TO THE OXYGEN BLENDING MODULE. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION OCCURRED RELATED TO THE OXYGEN BLENDING MODULE. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN BLENDING MODULE BOARD NEEDS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-03973 | 2021-09-23 | 2021-09-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO PATIENT HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04213 | 2021-09-23 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS RESPIRONICS, INC. | DSX500S11 | DSX500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES ON THE FILTER OF THE DEVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO EVIDENCE OF CONTAMINATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF WATER INGRESS AND AN UNKNOWN BLACK PARTICLE CONTAMINATION IN THE DRY BOX SEAL AND IN THE BLOWER MOTOR. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERROS LOGGED. THE DEVICE WAS APPLIED POWER AND THE DEVICE OPERATED PROPERLY. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH WATER INGRESS AND AN UNKNOWN BLACK PARTICLE CONTAMINATION INCONSISTENT WITH THE SOUND ABATEMENT FOAM. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-09
MW5104180 | 2021-09-23 | 2020-02-03 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAM STATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I HAVE BEEN COUGHING, CLEARING MY THROAT EXCESSIVELY HAVING A HEADACHE FOR ABOUT 2 YEARS NOW. JUST FOUND OUT MY PHILIPS CPAP MACHINE HAS BEEN RECALLED FOR A DEFECT THAT CAUSES THESE SYMPTOMS AND MORE. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04416 | 2021-09-23 | 2021-08-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED." SECTION D4 AND G4 HASS BEEN UPDATED IN THIS REPORT WHICH WAS NOT CAPTURED PREVIOUSLY. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04392 | 2021-09-23 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11 | DSX700H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES. RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. IN INITIAL MDR SECTION G3 IS IN CORRECTLY MENTIONED SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04235 | 2021-09-23 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGED COUGHING UP GREEN PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. A1 WAS MISSED TO CAPTURE, WHICH WAS UPDATED AND H6 HEALTH EFFECT- CLINICAL CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04393 | 2021-09-23 | 2021-07-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX400H11C | DSX400H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON OCT 12, 2023, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN THE WATER AND IN THE TUBING OF DEVICE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED PROBLEMS BREATHING AND ENDS UP GAGGING, RAN NUMEROUS TESTS. ADDITIONAL INFORMATION WAS RECEIVED ALLEGED OF NOSE IRRITATION HYPERSENSITIVITY REDUCED CARDIOPULMONARY RESERVE OTHER. THE SECTION E1 WAS UPDATED AND SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY THE PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECTS: CLINICAL CODES AND HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04402 | 2021-09-23 | 2021-09-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | GBX500H15 | GBX500H15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE END USER INDICATED THAT OZONE OR ANOTHER UNAPPROVED CLEANING AND DISINFECTION METHOD WAS USED. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON MAY 09, 2025 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE END USER INDICATED THAT OZONE OR ANOTHER UNAPPROVED CLEANING AND DISINFECTION METHOD WAS USED. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE THIRD PARTY SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED. THERE WERE NO ERRORS FOUND. THE THIRD PARTY SERVICE CENTER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. DEVICE WAS CORRECTED. SECTIONS D8, D9, H2, H3 AND H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04381 | 2021-09-23 | 2021-09-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | GBX500S15 | GBX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE END USER INDICATED THAT OZONE OR ANOTHER UNAPPROVED CLEANING AND DISINFECTION METHOD WAS USED. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO SERIOUS OR PERMANENT INJURY. DESPITE MULTIPLE ATTEMPTS, THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. UNIQUE IDENTIFIER(UDI) NUMBER IN SECTION D4, CODES IN SECTION H6 WERE UPDATED OR CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04383 | 2021-09-23 | 2021-09-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | GBX500S15 | GBX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE END USER INDICATED THAT OZONE OR ANOTHER UNAPPROVED CLEANING AND DISINFECTION METHOD WAS USED. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON SEP 01, 2022 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE END USER INDICATED THAT OZONE OR ANOTHER UNAPPROVED CLEANING AND DISINFECTION METHOD WAS USED. THE DEVICE WAS RETURNED TO THE THIRD PARTY SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED. THERE WERE NO ERRORS FOUND. THE THIRD PARTY SERVICE CENTER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. DEVICE WAS CORRECTED. SECTIONS D8, D9, H2, H3 AND H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04382 | 2021-09-23 | 2021-09-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | GBX500S15 | GBX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE END USER INDICATED THAT OZONE OR ANOTHER UNAPPROVED CLEANING AND DISINFECTION METHOD WAS USED. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON SEP 01, 2022 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE END USER INDICATED THAT OZONE OR ANOTHER UNAPPROVED CLEANING AND DISINFECTION METHOD WAS USED. THE DEVICE WAS RETURNED TO THE THIRD PARTY SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED. THERE WERE NO ERRORS FOUND. THE THIRD PARTY SERVICE CENTER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. DEVICE WAS CORRECTED. SECTIONS D8, D9, H2, H3 AND H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04380 | 2021-09-23 | 2021-09-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | GBX500S15 | GBX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE END USER INDICATED THAT OZONE OR ANOTHER UNAPPROVED CLEANING AND DISINFECTION METHOD WAS USED. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON SEP 01, 2022 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE END USER INDICATED THAT OZONE OR ANOTHER UNAPPROVED CLEANING AND DISINFECTION METHOD WAS USED. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED. THERE WERE NO ERRORS FOUND. THE THIRD PARTY SERVICE CENTER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. DEVICE WAS CORRECTED. SECTIONS D8, D9, H2, H3 AND H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04379 | 2021-09-23 | 2021-09-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | GBX500H15 | GBX500H15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE END USER INDICATED THAT OZONE OR ANOTHER UNAPPROVED CLEANING AND DISINFECTION METHOD WAS USED. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON SEP 01, 2022, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE END USER INDICATED THAT OZONE OR ANOTHER UNAPPROVED CLEANING AND DISINFECTION METHOD WAS USED. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED. THERE WERE NO ERRORS FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. DEVICE WAS CORRECTED. SECTIONS D8, D9, H2, H3 AND H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04398 | 2021-09-23 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. CORRECTION: THE MANUFACTURER PREVIOUSLY NOT CAPTURED, VISUALIZATION OF PARTICLES IN THE TUBING/CHAMBER AND NOTICING FLAKES IN THE MACHINE IN SECTION B5 AND IT HAS BEEN UPDATED IN THIS REPORT.SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04384 | 2021-09-23 | 2021-09-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | GBX500S15 | GBX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE END USER INDICATED THAT OZONE OR ANOTHER UNAPPROVED CLEANING AND DISINFECTION METHOD WAS USED. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER INADVERTENTLY SUBMITTED MDR 2518422-2021-04384-2. ALL RELEVANT INFORMATION HAD ALREADY BEEN CAPTURED IN MDR 2518422-2021-04384-1. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04215 | 2021-09-23 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE ASTHMA, CHRONIC BRONCHITIS AND A SINUS/UPPER RESPIRATORY INFECTION. THE PATIENT DID RECEIVE MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A PATIENT EXPERIENCED ASTHMA, CHRONIC BRONCHITIS AND A SINUS/UPPER RESPIRATORY INFECTION RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT DID RECEIVE MEDICAL INTERVENTION FOR THE REPORTED EVENTS. THE REPORTED EVENT OF AN UPPER RESPIRATORY INFECTION (SINUSITIS IS A FORM OF URI AND COUGH IS A SYMPTOM OF URI) ASSOCIATED WITH THEIR ASTHMA AND CHRONIC BRONCHITIS DIAGNOSES AND ITS REPORTED SEVERITY WAS REVIEWED BY THE PMS CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04214 | 2021-09-23 | 2021-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04217 | 2021-09-23 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING WEAKNESS, NAUSEA, NO APETITE, FEELS LIKE SOMETHING IN AIRWAYS CANNOT GET OUT OF AIRWAY, HEADACHES, DIZZINESS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED AND CORRECTED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04211 | 2021-09-23 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING COUGH, LUNG ISSUE/ STRANGE ODOR RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM.THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING COUGH, LUNG ISSUE/ STRANGE ODOR RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04212 | 2021-09-23 | 2021-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE HEADACHES, ASTHMA AND SOMETHING IN THE LUNGS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF AN X-RAY.. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4).. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT STATED DEVICE IS NOISY, A LIGHT STRANGE ODOR, ASTHMA, HEADACHES AND CHEST X-RAY AND IT SHOWED SOMETHING IN HER LUNGS HAS TO REPEAT IN 3 MONTHS. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENTS OF DEVICE IS NOISY, A LIGHT STRANGE ODOR, ASTHMA, HEADACHES AND CHEST X-RAY AND IT SHOWED SOMETHING IN HER LUNGS HAS TO REPEAT IN 3 MONTHS AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THESE EVENTS ARE ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURER CONCLUDES NO FURTHER ACTION IS NECESSARY. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION AND AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04387 | 2021-09-23 | 2021-08-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP A30 | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS,INC. | 1074683 | 1074683 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP,AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM.THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE THE PATIENT ALLEGED BLACK PARTICLES IN THE WATER IN HIS HUMIDITY CHAMBER,BURNING IN HIS MOUTH,SLIGHT INJURY IN THE MOUTH. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 (HEALTH EFFECT CLINICAL CODE) HAS BEEN CORRECTED AND (TYPE OF INVESTIGATION,INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS) HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
MW5104185 | 2021-09-23 | 2020-01-06 | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | DEVELOPED COUGH THAT HAS GRADUALLY BECOME WORSE TO THIS DATE ((B)(6) 2021). I HAVE ALSO NOTICED UPPER AIRWAY IRRITATION WHILE USING THE PHILIPS DREAMSTATION CPAP DEVISE. I HAVE BEEN CONTACTED BY PHILIPS ((B)(6) 2021) ADVISING ME OF THEIR RECALL AND HAVE COMPLETED THE INFORMATION FORMS THEY REQUIRE. |  | Y | N | LAY USER/PATIENT |  | 2021-09
MW5104179 | 2021-09-23 | 2021-09-21 | I, n, j, u, r, y | Voluntary report | DREAMSTNAUTOCPAPHUMHTCELL DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DreamStation Auto CPAP w/Humid, w/H Tube, w/Cell, US | DSX500T11C |  | I HAVE BEEN USING MY PHILIPS DREAMSTATION CPAP FOR MANY YEARS NOW. I HAVE NOTICED FOR YEARS HEAVY BLACK DEBRIS ON THE FILTER AND I HAVE TO CHANGE THE FILTER OFTEN TO THE POINT THAT I EVEN RAN OUT OF FILTERS AND HAD TO ORDER SOME ON (B)(6). WHAT WAS SENT WITH MY ORDER SHOULD HAVE BEEN ENOUGH. I NEVER THOUGHT TO LINK THIS ISSUE WITH MY CPAP, BUT I HAVE CONSTANT HEADACHES TO THE POINT I HAVE TO TAKE BC ALMOST DAILY, I HAVE A DAILY AND PERSISTENT COUGH, AND I SEEM TO HAVE A SINUS INFECTION OFTEN. I HAVE ALSO NOTICED THE PAST FEW MONTHS THAT MY CPAP MASK SEEMS TO BE BURNING MY FACE AS THE SKIN IS PEELING WHERE IT IS PLACED. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5104178 | 2021-09-23 | 2021-08-31 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | DREAMSTATION BY PHILIPS RESPRONICS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | J2245118791D9 | DSX500H11C |  | HAVE SIGNIFICANT SLEEP APNEA; I RECEIVED A RECALL LETTER FROM PHILIPS RESPRONICS - THEY CONFIRMED MY DEVICE IS RECALLED AND ADVISED NOT TO USE THE DEVICE. THEY DID NOT PROVIDE ME WITH A DATE FOR THE NEW MACHINE. AND THAT THE NEW MACHINE WOULD HAVE TO PASS CLINICAL TRIALS. I VISITED BY DOCTOR ON 9/2 AND HE MADE A PRESCRIPTION FOR ME AND SUBMITTED TO (B)(6) BY FAX. MY DOCTOR ADVISED A NEW MACHINE FROM RESMED - AN ALTERNATIVE PROVIDER. MY DOCTOR WROTE A PRESCRIPTION FOR MACHINE DAMAGED BEYOND REPAIR. I CALLED (B)(6) A WEEK LETTER AND THEY INDICATED THEY DID NOT HAVE THE FAX FROM MY DR. MY SON WENT TO THE DOCTOR AND RECEIVED A PHYSICAL COPY OF THE PRESCRIPTION - AND WE FAXED TO (B)(6). I CALLED (B)(6) TO CONFIRM RECEIPT. (B)(6) CONFIRMED. (B)(6) REP SAID IT WOULD TAKE A FEW DAYS TO PROCESS. I CALLED (B)(6) LAST WEEK TO CONFIRM PROGRESS - AND THEY TOLD ME MY PRESCRIPTION IS IN PROGRESS. NO UPDATE. I CALLED (B)(6) TODAY TO INQUIRE ON STATUS - AND SHE TOLD ME SHE COULD JUST TELL ME ITS IN PROGRESS. I ASKED WHAT IS THE STATUS OF THE PRESCRIPTION - WAS IT FAXED OR NOT? SHE SAID SHE COULD NOT TELL ME. TODAY, I CALLED (B)(6) AND THEY TOLD ME THE PRESCRIPTION WAS NOT FAXED TO THEM BY (B)(6) I WANT TO REGISTER A COMPLAINT ON (B)(6) AS WELL AS MY DR FOR LACK OF FOLLOW-UP FOR A SERIOUS HEALTH CONDITION CALLED SLEEP APNEA. I HAVE BEEN TESTED FOR SA AND HAVE BEEN TOLD MY SA RATE IF VERY HIGH AND I CANNOT SLEEP WITHOUT A MACHINE. THERE ARE COMPLIANCE RULES AROUND THIS TYPE OF CLAIM REPORTING ERROR AND LACK OF URGENCY I WANT A NEW MACHINE FROM RESMED BASED ON MY PRESCRIPTION SUBMITTED TO (B)(6). THANK YOU, (B)(6). FDA SAFETY REPORT ID# (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-09
MW5104171 | 2021-09-23 | 2016-01-01 | I, n, j, u, r, y | Voluntary report | REMSTAR AUTO A-FLEX CPAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | REMSTAR AUTO A-FLEX |  |  | I STARTED USING THE PHILIPS CPAP MACHINE, AFTER A SLEEP STUDY TEST. I HAD SLEEP APNEA AND USED THIS CPAP FOR TWO + YEARS. I ENDED UP WITH A TERMINAL LUNG DISEASE, WHICH REQUIRED A BI-LATERAL LUNG TRANSPLANT AT (B)(6) ON (B)(6) 2018. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04233 | 2021-09-23 | 2021-07-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE SINUS INFECTION, MEMORY ISSUES AND BAD BACTERIA. THE PATIENT DID RECEIVE SURGICAL INTERVENTION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING TO HAVE SINUS INFECTION, MEMORY ISSUES AND BAD BACTERIA RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE REPORTED EVENT OF DIAGNOSED WITH PATIENT TO HAVE SINUS INFECTION, MEMORY ISSUES AND BAD BACTERIA WAS REVIEWED BY THE PMS CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE.THE PATIENT DID REPORT RECEIVING MEDICAL INTERVENTION AND WAS PRESCRIBED AN INHALER. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5104169 | 2021-09-23 | 2021-09-21 | I, n, j, u, r, y | Voluntary report | PHILLIPS DREAM STATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | J237715800327 |  |  | AS (B)(6) CARE GIVER HE USES HIS CPAP MACHINE AS INSTRUCTED HASN'T USED A OZONE GENERATORS REGULARLY REPLACE FILTERS BUT HAS NOTICED BLACK DOTS IN PIPE JOINED ON HIS FACE MASK AND HEADACHE, AND HEARING DIFFICULTIES. ITCHY SKIN AND WORSENING SYMPTOMS OF BREATHLESSNESS ALSO HIS BEEN IN A LOT OF DISCOMFORT FROM LOWER BACK PAIN, SWOLLEN STOMACH, SORE THROAT SINUSES, CONGESTIVE PAIN, AND PROBLEMS IT DIURETIC RESULTS FOR HIS "EDEMA" AND LOW POTASSIUM. FATIGUE FROM WAKING IN TERRORS, FEELING CHOKED EVEN THOUGH HE IS WEARING HIS MASK, AND ERRATIC READINGS OF HIS BLOOD SUGAR AND SAYING HIS EYES AND SIGHT HEARING ARE BAD. WE ASSOCIATED THESE SYMPTOMS WITH HIS PREEXISTING HEART FAILURE SYMPTOMS. AS HIS SORE THROAT SINUSES TURNED TO A PRODUCTIVE COUGH, HE HAS REGULARLY COMPLAINING HIS TONGUE AND THROAT FELT BURNED. (B)(6) IS (B)(6) AND HAS DIABETES STAGE 3 HEART FAILURE AND SOME THESE SYMPTOMS AS KNOW ARE SOME THING WE CONTINUALLY CONTROL DAILY AND WATCH FOR BUT THESE WERE DIFFERENT. I WAS WATCHING FOR THE TALE SIGNS OF FLUID OVERLOAD, AS (B)(6) IS ON FLUID RESTRICTED AFTER WE DISCOVER THAT HIS CPAP WAS ONE AFFECTED AND DISCONTINUED USE OF SAID MASK HE IS IMPROVING, WE WANTED TO INFORM YOU SO OTHERS CAN STOP SUFFERING AS HE DID BY BEING UNAWARE AND CAN YOU PLEASE HELP ME TO CHECK THAT HE WILL NOT HAVE ANY OTHER DAMAGE OR SYMPTOMS AND WHAT TO WATCH FOR OR DO, SO HE IS OK. THANK YOU. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | I |  | 2021-09
2518422-2021-04389 | 2021-09-23 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | DREAMSTATION AUTO CPAP | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING/CHAMBER, DIFFICULTY BREATHING/SHORT OF BREATH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED PARTICLES IN TUBING/CHAMBER OF THE DEVICE AND DIFFICULTY BREATHING/SHORT OF BREATH. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION G3 WAS PREVIOUSLY INCORRECT AND SHOULD HAVE BEEN '08/09/2021'. SECTION H6 (HEALTH EFFECT - CLINICAL CODE) UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04388 | 2021-09-23 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE TWO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED DUE TO AGE. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04203 | 2021-09-23 | 2021-09-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04114 | 2021-09-23 | 2021-07-22 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY SUBMITTED MDR 2518422-2021-04114-1 WITH INCORRECT SECTIONS B1, B2, H1, H6. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04113 | 2021-09-23 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A THYROID MASS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | IN THE PREVIOUS REPORT 2518422-2021-04113-1, SECTION H1 WAS MISSED TO CAPTURE. THE FOLLOWING SECTIONS HAVE BEEN UPDATED IN THIS REPORT. SECTION B1 HAS BEEN UPDATED TO REFLECT AS PRODUCT PROBLEM. SECTION H1 HAS ALSO BEEN UPDATED TO REFLECT THE MALFUNCTION. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04337 | 2021-09-23 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04115 | 2021-09-23 | 2021-07-23 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION BIPAP AUTOSV, DS | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | DSX900S11 | DSX900S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG CANCER. THE PATIENT HAD SURGERY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON SEP 10, 2023, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING LUNG CANCER RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED ABOUT ALLEGATION OF EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, DIZZINESS AND/OR HEADACHE, CANCER, OTHER. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. SECTION D8, D9, H2, H3 AND H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04391 | 2021-09-23 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUNDABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED COMPLAINT OF PARTICLES IN WATER RESERVOIR AND A STRANGE ODOR FROM THE DEVICE. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND GOOD AIRFLOW BUT HAS ODOR, WHITE DEPOSITS IN FILTER INLET. PIL USED A FITT (FOAM INTEGRITY TEST TOOL) AND THE ASSOCIATED PROCEDURE ER 2245460 V01 AND WAS ABLE TO CONFIRM THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM. THERE ARE VISIBLE DAMAGE AND FUNCTIONALITY FAILURES OF THE DEVICE, WHICH SUGGEST THAT THE SOURCE OF CONTAMINATION WAS INTERNAL TO THE DEVICE. THIS INVESTIGATION WAS PERFORMED AS OUTLINED IN ER 2244873 (V01). PIL INITIATED THERAPY WITH THE DEVICE AND VERIFIED AIRFLOW, USING A PIL SUPPLIED POWER SUPPLY AND POWER CORD. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS. THE MANUFACTURER CONCLUDES THE DEVICE HAS CONTAMINATION INTERNAL TO THE DEVICE. THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04143 | 2021-09-23 | 2021-09-16 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL, AU | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | AUX500T15C | AUX500T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE DIZZINESS, HEADACHES, NAUSEA, VOMITING AND CHEST TIGHTNESS. THE PATIENT WAS REPORTEDLY HOSPITALIZED IN RESPONSE TO THE EVENT. AND PRESCRIBED ANTIBIOTICS. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04313 | 2021-09-23 | 2021-07-12 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP DIFFICULTY BREATHING. THE PATIENT WAS PRESCRIBED STEROIDS IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04256 | 2021-09-23 | 2021-07-07 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP RESPIRATORY INFLAMMATION. THE PATIENT WAS PRESCRIBED ANTIBIOTICS IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED RESPIRATORY INFLAMMATIONT. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENT OF RESPIRATORY INFLAMMATION AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. THE PATIENT HAS SWAPPED OUT THE DEICE AT THE DISTRIBUTOR AND AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04312 | 2021-09-23 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | CAX500T12 | CAX500T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04397 | 2021-09-23 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING CHAMBER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED.. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING CHAMBER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04116 | 2021-09-23 | 2021-07-12 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER CELLS. THE PATIENT HAD SURGERY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER CELLS. THE PATIENT HAD SURGERY IN RESPONSE TO THE REPORTED EVENT. THE REPORTED EVENT WAS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. ADDITIONAL INFORMATION WAS RECEIVED AND ADDED TO THE REPORT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION ON 02/23/2023 FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED ON 08/01/2023. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE DEVICE WAS EVALUATED, AND EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN WAS NOT OBSERVED IN THE BASE UNIT. THE MANUFACTURER OBSERVED DUST/DIRT CONTAMINATION THROUGHOUT DEVICE ENCLOSURE, AND AIRPATH, SUGGESTING A SOURCE EXTERNAL TO THE DEVICE. EVIDENCE OF WATER INGRESS ON THE BLOWER AND BLOWER BOX. INSECT WAS OBSERVED ON THE EXTERIOR OF THE BLOWER BOX. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THAT THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN OBSERVED IN THE BASE UNIT. THE MANUFACTURER CONFIRMED THE PRESENCE OF CONTAMINATION IN THE AIRPATH. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04229 | 2021-09-23 | 2021-08-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04295 | 2021-09-23 | 2021-07-30 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED CANCER. THE PATIENT DID RECEIVE MEDICAL INTERVENTION REGARDING DIAGNOSIS BUT NO OTHER INFORMATION IS KNOWN. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM AND CAUSED CANCER. THE PATIENT DID RECEIVE MEDICAL INTERVENTION REGARDING DIAGNOSIS BUT NO OTHER INFORMATION IS KNOWN. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04378 | 2021-09-23 | 2021-07-22 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE NOSE BLEEDS AND DRY MOUTH. THE PATIENT WAS ADMITTED TO THE HOSPITAL FOR DIFFICULTY BREATHING. THE DETAILS OF THE MEDICAL INTERVENTION RECEIVED BY THE PATIENT ARE CURRENTLY UNKNOWN. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT WAS ALLEGING EXPERIENCE NOSE BLEEDS AND DRY MOUTH. THE PATIENT WAS ADMITTED TO THE HOSPITAL FOR DIFFICULTY BREATHING. NO OTHER CLINICAL INFORMATION WAS REPORTED. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. ONE ERROR WAS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04347 | 2021-09-23 | 2021-09-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | JPX500S16 | JPX500S16 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK DUST ADHERED TO THE DEVICE. THERE WAS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THE MANUFACTURER CONTACT INFORMATION HAS ALSO BEEN UPDATED IN THIS FOLLOW UP REPORT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. DURING THE EVALUATION OF THE DEVICE, THE MANUFACTURER VISUALLY INSPECTED THE DEVICE AND FOUND NO EVIDENCE OF FOAM DEGRADATION. INTERNAL EVALUATION OF THE DEVICE FOUND WHITE CONTAMINATION AND INDETERMINATE BLACK PARTICULATES, INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM, AND SOME VERY SMALL POTENTIAL HAIRS AT THE AIR INPUT OF THE BLOWER BOX. BLACK CONTAMINATION, CONSISTENT WITH KERATIN, WAS OBSERVED AT THE AIR OUTLET OF THE BLOWER BOX. INDETERMINATE BLACK PARTICULATES WERE OBSERVED ALL THROUGHOUT THE ENCLOSURE AND IN THE BLOWER BOX AND DOES NOT APPEAR CONSISTENT WITH FOAM DEGRADATION. PRODUCT INVESTIGATION LAB CAN CONFIRM THE PRESENCE OF MULTIPLE CONTAMINANTS THAT ARE NOT CONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM. THERE IS NO VISIBLE DAMAGE OR FUNCTIONALITY FAILURES OF THE DEVICE, WHICH SUGGEST THAT THE SOURCE OF CONTAMINATION WAS EXTERNAL TO THE DEVICE. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND 2 ERRORS LOGGED. THE MANUFACTURER WAS UNABLE TO CONFIRM THE COMPLAINT OF BLACK DUST BUT HAS CONFIRMED MULTIPLE BLACK PARTICULATES THROUGHOUT THE DEVICE THAT DOES NOT APPEAR CONSISTENT WITH PREVIOUSLY OBSERVED FOAM DEGRADATION. PLEASE SEE SECTIONS G1, H3, AND H6 FOR THIS UPDATED CODING AND INFORMATION. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04252 | 2021-09-23 | 2021-06-15 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP HIGH BLOOD PRESSURE. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE HIGH BLOOD PRESSURE. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL PART OF THE DEVICE AND FOUND NO UNUSUAL WEAR OR TEAR. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND FOUND EVIDENCE OF CORROSION NEAR SW4. THIS CORROSION IS LIKELY DUE TO LIQUID INGRESS FROM THE THERAPY BUTTON. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND 32 ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN IN THE BASE UNIT BUT A SMALL AMOUNT OF LIQUID INGRESS WAS OBSERVED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04395 | 2021-09-23 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO THE B5 SHOULD BE : THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING/CHAMBER, RUNNY NOSE, FEELING SICK RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED / CORRECTED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04253 | 2021-09-23 | 2021-07-25 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG CANCER. THERE WAS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THAT TIME. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED BY THE MANUFACTURER. THE MANUFACTURER FOUND NO SIGNS OF CONTAMINATION ON THE DEVICE. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND SIGNS OF CONTAMINATION INSIDE THE BLOWER BOX AND ON THE BLOWER MOTOR. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERRORS LOGGED. THE DEVICE WAS HOOKED UP TO POWER SUPPLY, AIRFLOW WAS VERIFIED AND THE DEVICE OPERATED PROPERLY. THE MANUFACTURER CONCLUDES NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. THE MANUFACTURER CONFIRMED EVIDENCE OF CONTAMINATION IN THE BLOWER BOX AND ON THE BLOWER MOTOR. SECTION D9, G3, H6 HAS BEEN UPDATED / CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04342 | 2021-09-23 | 2021-06-15 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THE PATIENT WAS PRESCRIBED CHEMOTHERAPY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON OCT 09, 2023, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING CANCER. THE PATIENT WAS PRESCRIBED CHEMOTHERAPY IN RESPONSE TO THE REPORTED EVENT. RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGED EYE IRRITATION, LUNG DISEASE. SECTIONS E1 AND SECTION H6 HEALTH EFFECTS: CLINICAL CODES HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04254 | 2021-09-23 | 2021-07-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11 | DSX400H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP EYE ISSUES. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 12, 2024, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES THE MANUFACTURER RECEIVED INFORMATION ALLEGING AND CAUSED THE PATIENT TO HAVE A NASAL/THROAT IRRITATION, EYE DRYNESS, ITCHY AND DOCTOR HAS HER ON TWO DIFFERENT MEDICATIONS RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTIONS B1, B2, H1 AND H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR ONLY THE PRODUCT PROBLEM (BOTH ADVERSE EVENTS AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.). SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.). SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04348 | 2021-09-23 | 2021-09-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | RJPX500S16 | RJPX500S16 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER COMPLETED AN INTERNAL VISUAL INSPECTION. THE MANUFACTURER FOUND EVIDENCE OF AN UNKNOWN BLACK PARTICULATES, VERY SMALL POTENTIAL HAIRS AND BLACK CONTAMINATION CONSISTENT WITH KERATIN IN THE AIR INPUT OF THE BLOWER BOX, BLOWER BOX AND AIR OUTLET OF THE BLOWER BOX. THE MANUFACTURER ALSO FOUND EVIDENCE OF GREY (DUST-LIKE) CONTAMINATION AND BIO-CONTAMINATION IN THE BLOWER MOTOR, BLOWER BOX AND NOSE MASK. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION /BREAKDOWN. THE DEVICE'S EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND THE ERROR LOG SHOWED, 1 ERROR LOGGED. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH BLACK PARTICULATE , KERATIN, DUST AND SMALL POTENTIAL HAIRS , CONTAMINATION INCONSISTENT WITH THE SOUND ABATEMENT FOAM. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. SECTION G3, H3 AND H6 HAS BEEN UPDATED / CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04228 | 2021-09-23 | 2021-08-23 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500S11 | DSX500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE PULMONARY DISTRESS. THE PATIENT WAS REPORTEDLY HOSPITALIZED IN RESPONSE TO THE EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED BLACK MOLDS IN THE DEVICE AND WAS HOSPITALIZE SEVERAL TIMES BECAUSE OF PULMONARY DISTRESS, DEVELOPED ASTHMA, REACTIVE AIRWAY DISEASE, SORE THROAT AND COUGH. THE PATIENT WAS REPORTEDLY HOSPITALIZED IN RESPONSE TO THE EVENT AND UNDERGOING THE MEDICAL INTERVENTION. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5104165 | 2021-09-23 | 2021-06-14 | I, n, j, u, r, y | Voluntary report | PHILIPS RSPIRONCS DREAM STATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I HAVE A PHILIPS RESPIRONICS DREAM STATION CPAP, A MODEL THAT WAS INCLUDED IN THE RECALL. I WAS HOSPITALIZED WITH COVID-19 PNEUMONIA IN (B)(6) OF 2020. I USED MY CPAP WHILE IN A REHABILITATION FACILITY IN (B)(6), AND CONTINUED TO USE IT, AS NORMAL AFTERWARDS. WHEN I CAME HOME FROM THE REHAB FACILITY, I WAS HAVING SIGNIFICANT PROBLEMS WITH SHORTNESS OF BREATH. THESE IMPROVED NOTICEABLY OVER THE SUBSEQUENT MONTHS, THEN I BEGIN TO SEE THIS SYMPTOM (SHORTNESS OF BREATH) WORSENING. I ALSO BEGIN TO EXPERIENCE A GREAT DEAL OF THROAT AND AIRWAYS IRRITATION. VERY MUCH LIKE ONE WOULD EXPECT IF YOU WERE COMING DOWN WITH A BAD COLD, ONLY I NEVER GOT THE COLD, JUST A CONTINUAL SIGNIFICANT IRRITATION OFT HE THROAT AND AIRWAYS. WHEN I HEARD OF THE RECALL, I STOPPED USING MY CPAP IMMEDIATELY. THE THROAT AND AIRWAY IRRITATION STOPPED WITHIN TWO DAYS, AND HAS NOT RETURNED. I AM NOW USING AN AIRSENSE 10, WITH NO IRRITATION. MY SHORTNESS OF BREATH HAS IMPROVED AGAIN, SLIGHTLY, BUT NOT TO THE POINT IT WAS AT PRIOR TO THE IRRITATION OF THROAT AND AIR PASSAGES, AND THE WORSENING OF SHORTNESS OF BREATH THAT OCCURRED IN EARLY (B)(6). EQUIPMENT PROVIDER (B)(6) REPORTED DEVICE FOR RECALL TO PHILIPS RESPIRONICS. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04396 | 2021-09-23 | 2021-08-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED (B)(4). | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING USING AN OZONE BASED DISINFECTION DEVICE AND WHITE SPECKS IN THE WATER TANK RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04314 | 2021-09-24 | 2020-12-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG NODULES. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE SECTIONS B1, B2, H1 AND H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM. (IN THE PREVIOUS MDR BOTH ADVERSE EVENTS AND PRODUCT PROBLEM WAS CHECKED.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CORRECTED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04345 | 2021-09-24 | 2021-08-03 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION BIPAP AUTOSV H/HT/C, DS | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS INC | DSX900T11C | DSX900T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP RESPIRATORY FAILURE. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04472 | 2021-09-24 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11 | DSX400H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING COUGH,CHEST FEELS HEAVY AND DIFFICULTY BREATHING/SHORT OF BREATH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04289 | 2021-09-24 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04292 | 2021-09-24 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES, NASAL/THROAT IRRITATION OR SORENESS, BREATHING PROBLEMS, MILKY YUCKY SUBSTANCE INSIDE THE WATER CHAMBER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION D1 (BRAND NAME), D4 (MODEL #, CATALOG #, SERIAL #, UNIQUE IDENTIFIER (UDI) #), G4 (PMA/510(K)), H4 (DEVICE MANUFACTURE DATE), H6 (HEALTH EFFECT - CLINICAL CODE) CORRECTED IN THIS REPORT. SECTION H6 (TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS) UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04294 | 2021-09-24 | 2021-07-07 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING THE PATIENT REPORTED COUGH, COLD COMING ON, DIFFICULTY BREATHING AND SKIN CANCER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT DID REPORT TO RECEIVE MEDICAL INTERVENTION AND HAD SOME SKICANCER REMOVED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 HAS BEEN UPDATED IN THIS REPORT. SECTION G4 HAS BEEN CORRECTED IN THIS REPORT, TO REFLECT THE CORRECT INFORMATION. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04291 | 2021-09-24 | 2021-08-12 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON OCT 24, 2024, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES AND LUNG DISEASES, ASTHMA RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THERE WAS NO REPORT OF PERMANENT OR SERIOUS PATIENT HARM OR INJURY. THE PATIENT DID NOT REPORT TO RECEIVE MEDICAL INTERVENTION. SECTION B1 WAS CORRECTED FOR BOTH ADVERSE EVENT AND PRODUCT PROBLEM. (PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION. TO SERIOUS INJURY. SECTION H6 WAS CORRECTED AND UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04290 | 2021-09-24 | 2021-07-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DEVICE GIVING OFF BLACK SUBSTANCE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS UPDATED REPORT WILL BE FILED. THE MANUFACTURER PREVIOUSLY NOT CAPTURED IN SECTION H4 AND IT HAS BEEN UPDATED IN THIS REPORT SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04317 | 2021-09-24 | 2021-07-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX700H11 | DSX700H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. IN THE INITIAL REPORT, H10 SECTION WAS MARK INCORRECTLY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF LIQUID INGRESS, MINERAL DEPOSITS ON THE BLOWER, KERATIN CONTAMINATION AROUND THE BLOWER BOX, DUST CONTAMINATION IN THE BLOWER AND BLOWER BOX SEAL. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERROR LOGGED. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH LIQUID INGRESS, MINERAL DEPOSITS ON THE BLOWER, KERATIN CONTAMINATION AROUND THE BLOWER BOX, DUST CONTAMINATION IN THE BLOWER AND BLOWER BOX SEAL. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION D9,G3,H6 HAS BEEN UPDATED/CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04297 | 2021-09-24 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04367 | 2021-09-24 | 2021-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04344 | 2021-09-24 | 2021-09-03 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP,AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AND CAUSED A PATIENT TO DEVELOP LUNG CANCER RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04359 | 2021-09-24 | 2019-09-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AUX400T15 | AUX400T15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION B5 WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT SHOULD BE REPORTED AS FOLLOWS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED SINUSITIS, RECURRENT LOWER RESPIRATORY TRACK INFECTIONS AND HEADACHES. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. SECTION H6 HEALTH EFFECT - CLINICAL CODE WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT HAS BEEN CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04369 | 2021-09-24 | 2021-03-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTIONS H6 AND H11 WERE INCORRECTLY CAPTURED IN THE PREVIOUS REPORT. IT SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED OF HEADACHE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04371 | 2021-09-24 | 2021-04-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04473 | 2021-09-24 | 2021-09-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11 | DSX700H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING OILY BLACK SUBSTANCE GOING THROUGH THE TUBING ALL THE WAY ON HIS NOSE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING OILY BLACK SUBSTANCE GOING THROUGH THE TUBING ALL THE WAY ON HIS NOSE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04445 | 2021-09-24 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF WHITE PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF WHITE PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04446 | 2021-09-24 | 2021-08-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING, PARTICLES IN AIRWAY, CONSTANTLY COUGHING AND BREATHING PROBLEM RELATED TO A CPAP SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION PARTICLES IN AIRWAY, CONSTANTLY COUGHING AND BREATHING PROBLEM RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04450 | 2021-09-24 | 2021-08-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TESTING. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BLOWER MOTOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04357 | 2021-09-24 | 2021-08-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WOULD NOT POWER ON. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | IT WAS INITIALLY REPORTED, THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WOULD NOT POWER ON. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S POWER SUPPLY WAS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04449 | 2021-09-24 | 2021-08-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR FAILED TESTING. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S PROPORTIONAL VALVE WAS REPLACED TO ADDRESS THE ISSUE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04451 | 2021-09-24 | 2021-09-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040007 | 1040007 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04322 | 2021-09-24 | 2021-08-25 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | GB2110X15B | GB2110X15B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED A THIRD PARTY SERVICE CENTER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04368 | 2021-09-24 | 2021-08-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04415 | 2021-09-24 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04298 | 2021-09-24 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF BLACK PARTICLES IN FILTER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE EVENT IS CURRENTLY UNKNOWN. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04285 | 2021-09-24 | 2021-07-19 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | IS THE ORIGINAL REPORT ASSOCIATED WITH THIS DEVICE. A DUPLICATE REPORT HAS BEEN SUBMITTED FOR THIS DEVICE UNDER REPORT MDR 2518422-2022-04399. THIS FOLLOW-UP REPORT IS BEING FILED FOR AWARENESS OF THIS DUPLICATE REPORT. . | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04432 | 2021-09-24 | 2021-07-22 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP AN UPPER RESPIRATORY TRACT INFECTION. THE PATIENT WAS PRESCRIBED ANTIBIOTICS IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP AN UPPER RESPIRATORY TRACT INFECTION. THE PATIENT WAS PRESCRIBED ANTIBIOTICS IN RESPONSE TO THE REPORTED EVENT. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL.THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION.THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04255 | 2021-09-24 | 2021-07-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG CANCER. THE PATIENT WAS PRESCRIBED CHEMOTHERAPY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY SUBMITTED MDR 2518422-2021-04255-1 WITH INCORRECT SECTIONS B1, B2, H1, H6. CORRECTIONS TO PREVIOUS MDR ARE MADE IN THIS REPORT AS FOLLOWS. SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY PRODUCT PROBLEM WAS CHECKED IN PREVIOUS MDR). SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY IT WAS BLANK). SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6- HEALTH IMPACT CODE WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04420 | 2021-09-24 | 2021-08-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY.THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING COUGHING OF FLUID WITH COLOUR EVERY MORNING AND FEEL WEAK AND DULL IN THE MORNING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILE CORRECTED B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE USER ALLEGING COUGHING OF FLUID WITH COLOUR EVERY MORNING AND FEEL WEAK AND DULL IN THE MORNING. THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE PATIENT DID NOT REPORT TO RECEIVE MEDICAL INTERVENTION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04417 | 2021-09-24 | 2021-09-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11 | DSX700T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04412 | 2021-09-24 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING STRANGE ODOR AND DISCOLOURED FILTER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING STRANGE ODOR AND DISCOLOURED FILTER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTION G1- MANUFACTURER CONTACT OFFICE INFO WAS PREVIOUSLY INCORRECT AND HAS NOW BEEN CORRECTLY UPDATED. SECTION G3 WAS PREVIOUSLY INCORRECT. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04413 | 2021-09-24 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP ST30 | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | DSX1030T11C | DSX1030T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING BLACK PARTICLES IN TUBING AND INLINE FILTER RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK PARTICLES IN TUBING AND INLINE FILTER RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTION G1-MANUFACTURER CONTACT OFFICE INFO WAS PREVIOUSLY INCORRECT AND HAS NOW BEEN CORRECTLY UPDATED. SECTION G3 WAS PREVIOUSLY INCORRECT. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04320 | 2021-09-24 | 2021-09-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, JP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | JPX500S16 | JPX500S16 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED FOR EVALUATION AND THE CUSTOMER CANCELLED REPAIR. THE DEVICE WAS RETURNED UNREPAIRED. THE REPORTED PHENOMENON WAS NOT DUPLICATED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04343 | 2021-09-24 | 2021-08-03 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX700H11 | DSX700H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SINUS INFECTIONS. THE PATIENT WAS PRESCRIBED ANTIBIOTICS IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SINUS INFECTIONS. A CORRECTION TO THE B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM AND CAUSED A PATIENT TO DEVELOP SINUS INFECTIONS, WHEEZING AND COUGHING UP MUCUS, CHEST CONGESTION WITH COUGH. THE PATIENT WAS PRESCRIBED ANTIBIOTICS IN RESPONSE TO THE REPORTED EVENT. THE DEVICE WAS RETURNED TO THE PRODUCT INVESTIGATION LAB FOR EVALUATION. THE INTERNAL AND EXTERNAL ASPECT OF THE DEVICE WAS INSPECTED AND THE MANUFACTURER OBSERVED DUST/DIRT CONTAMINATION INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM THROUGHOUT DEVICE ENCLOSURE AND AIRPATH, SUGGESTING A SOURCE EXTERNAL TO THE DEVICE. THE MANUFACTURER ALSO OBSERVED EVIDENCE OF WATER INGRESS ON THE BLOWER AND BLOWER BOX OF THE DEVICE. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY THE MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION BUT EVIDENCE OF DUST / DIRT CONTAMINATION AND WATER INGRESS IN THE DEVICE WAS OBSERVED BY THE MANUFACTURER, WHICH ARE CONSISTENT WITH BEING FROM AN EXTERNAL SOURCE. | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5104188 | 2021-09-24 |  | D, e, a, t, h | Voluntary report | RESPIRONICS REMSTAR AUTO M SERIES A FLEX | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIROICS, INC. | 1038863 |  |  | ON (B)(6) 2020 MY MOTHER (B)(6) WAS UNABLE TO AWAKE FROM SLEEP AND SHOWED SIGNS OF CONFUSION AND LETHARGY, SHE WAS UNABLE TO SPEAK OR ANSWER QUESTIONS APPROPRIATELY. WE HAD CONCERNS THERE WAS AN ISSUE WITH HER CPAP AT THE TIME BECAUSE OF HER EXHIBITING CONFUSION. SHE WAS TRANSPORTED BY AMBULANCE, INTUBATED AND THEN ADMITTED TO THE HOSPITAL. SHE WAS HOSPITALIZED A FEW MORE TIMES FOR THIS SAME TYPE OF ISSUE OVER THE COURSE OF A FEW MONTHS. IN (B)(6) 2020 SHE WAS DIAGNOSED WITH PANCREATIC AND LIVER CANCER. SHE PASSED AWAY ON (B)(6) 2021. EARLIER THIS MONTH I WAS MADE AWARE OF A RECALL OF PHILLIPS CPAP MACHINES DUE TO THE PE-PER FOAM USED TO MINIMIZE NOISE. THIS IS OF GREAT CONCERN TO ME BECAUSE MY MOM HAD USED A PHILLIPS CPAP MACHINE FOR MANY YEARS (REPIRONICS REMSTAR AUTO M SERIES). SHE HAD SEVERE SLEEP APNEA AND COULD NOT SLEEP UNLESS SHE USED HER C-PAP MACHINE SO SHE USED IT EVERY TIME FOR SEVERAL HOURS, AND DONE SO FOR AT LEAST 5 YEARS. I DO NOT BELIEVE THE MACHINE SHE USED IS PART OF THE RECALL BECAUSE IT IS NO LONGER IN CIRCULATION SO I AM UNABLE TO FIND OUT IF THIS TYPE OF FOAM WAS USED IN HER MACHINE. I HAVE HER OLD MACHINE IN MY POSSESSION AND AM WILLING TO HAVE IT ANALYZED TO SEE IF THIS FOAM IS IN IT. I AM CONCERNED BECAUSE HER PHYSICIANS WERE NOT ABLE TO FIND AN ETIOLOGY FOR HER CANCER BUT DID DETERMINE IT WAS NOT GENETIC. IN ADDITION SHE HAD EXHIBITED CONFUSION ON AND OFF FOR A COUPLE OF YEARS WHILE USING THE PHILLIPS MACHINE, SHE HAD SEVERAL TEST DONE PRIOR TO (B)(6) 2020 TRYING TO FIND OUT WHAT WAS CAUSING THE INTERMITTENT CONFUSION BUT WAS UNABLE TO FIND THE CAUSE OF HER ISSUE. WITH THE RECENT DEVELOPMENT OF PHILLIPS USING CANCER CAUSING FOAM IN THEIR MACHINES, AND MY MOM BEING VERY COMPLIANT WITH THE USE OF HER CPAP AND THEN BEING DX WITH CANCER OF NO GENETIC ORIGIN I BELIEVE THIS SHOULD BE RESEARCHED BUT I AM NOT SURE HOW TO GO ABOUT GETTING THIS DONE. THERE ARE MULTIPLE TESTS SHE HAD RUN AND SEVERAL HOSPITALIZATIONS OVER THE COURSE OF A FEW YEARS. AS I EXPLAINED WE HAD TRIED FOR A FEW YEARS TO DETERMINE THE CAUSE OF HER CONFUSION, WEAKNESS AND LETHARGY. SHE HAD APPOINTMENTS WITH MULTIPLE VARIOUS SPECIALISTS FROM 2015 - 2020 BECAUSE WE WERE UNABLE TO FIND OUR WHAT WAS CAUSING HER CONFUSION. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5104193 | 2021-09-24 | 2020-06-06 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS DREAMSTATION BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I WOULD WAKE IN THE MIDDLE OF THE NIGHT WITH AN UNCONTROLLABLE COUGH. I HAD TO USE A RESCUE INHALER TO CALM DOWN MY COUGH. I AM UNSURE OF WHAT TESTS YOU ARE REFERRING TO. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04411 | 2021-09-24 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT STATES BLACK GOOEY SUBSTANCE AT THE BOTTOM OF WATER CHAMBER. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT STATES BLACK GOOEY SUBSTANCE AT THE BOTTOM OF WATER CHAMBER. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION G3 NEEDS TO CORRECT IN INITIAL REPORT WHICH WAS CAPTURED WRONGLY AND SHOULD BE 8/10/2021. | N | Y | LAY USER/PATIENT |  | 2021-09
MW5104207 | 2021-09-24 | 2021-09-23 | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | REMSTAR AUTO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | EXPOSED TO "HAZ." MATERIAL; EXPOSURE TO PARTICULATE FOAM / OUTGASSING FROM FOAM IN PHILIPS REMSTAR CPAP (S/N (B)(4).) CONFIRMED DEVICE IS ON RECALL BY PHILIPS THROUGH THEIR WEBSITE (AND SUBMITTED INFORMATION TO THEM.) CONFIRMED PRESENCE OF BLACK PARTICULATE MATTER IN TUBE RUNNING FROM CPAP TO FACEMASK. PREVIOUSLY USED THIS DEVICE FOR APPROX 10 YEARS (DEVICE WAS REPLACED EARLIER THIS YEAR BEFORE RECALL WAS ANNOUNCED.) HAVE HAD CHRONIC SINUS INFLAMMATION FOR YEARS ALONG WITH BORDERLINE METABOLIC PANEL RESULTS INDICATING POSSIBLE LIVER / KIDNEY ISSUES DURING THE SAME PERIOD. FDA SAFETY REPORT ID # (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-09
MW5104201 | 2021-09-24 | 2014-06-01 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS SYSTEMONE 60 CPAP REMSTAR PRO DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | REMSTAR PRO DOM | 460P |  | PHILIPS RESPIRONIC CPAP MEDICAL DEVICE RECALL (SYSTEM ONE 60 SERIES) POLYESTER-BASED POLYURETHANE (PE-PUR) NOTICE 09/21/2021. WAS DIAGNOSED WITH BILATERAL RENAL CANCER IN (B)(6) 2014 AT (B)(6). RESULTED IN RADICAL NEPHRECTOMY OF LEFT KIDNEY AND PARTIAL RADICAL NEPHRECTOMY OF 20% OF RIGHT KIDNEY. REOCCURRENCE OF RENAL CANCER IN 2019 IN RIGHT KIDNEY RESULTING IN CT ABLATION (FREEZING) OF TUMOR. STARTED USE OF MY PHILIPS RESPIRONIC SYSTEM ONE 60 SERIES CPAP YEARS BEFORE CANCER WAS DIAGNOSED. FULL DETAILED MEDICAL FILES AVAILABLE. ADDITIONAL CT SCANS, MRIS, AND PATHOLOGY PERFORMED IN PREPARATION OF, POST-SURGICAL PROCEDURES. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04418 | 2021-09-24 | 2021-09-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO AUTO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSG500S11 | DSG500S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04419 | 2021-09-24 | 2021-09-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11 | DSX700T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04414 | 2021-09-24 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A (CPAP, BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING SEEING WHITE FLECKS IN WATER CHAMBER AFTER USAGE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION G1 HAS BEEN CORRECTED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04286 | 2021-09-24 | 2021-08-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO BIFLEX W/SD CARD, HUMID, US | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DS750HS | DS750HS |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOISY, BURNING/SMOKE/ELECTRICAL ODOR, BLACK PARTICLES IN UNIT, STOPS BREATHING, KIDNEY ISSUES RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. AFTER THE FINAL ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION, CUSTOMER DECLINED TO RESPOND TO THE GFE RELATED QUESTIONS AND TERMINATED THE CALL. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04315 | 2021-09-24 | 2021-09-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP S/T C SERIES INTL 30, FLOW GEN | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS INC | 1061T | 1061T |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
MW5104215 | 2021-09-24 | 2021-09-14 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS SYSTEM ONE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS INC. | SYSTEM ONE | REF 560P |  | APPOINTMENT WITH (B)(6) SLEEP DEPARTMENT CONCERNING MY CPAP MACHINE, THE BLACK FLECKS IN MY NASAL PIECE, THE RECALL ON THE PRODUCT, SIDE EFFECTS I'VE NOTICED AND POSSIBILITIES OF REPLACING MY MACHINE VIA MY INSURANCE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04232 | 2021-09-26 | 2021-09-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL, AU | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | AUX500T15C | AUX500T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT HAD A CT SCAN DONE THAT REVEALED AN OPAQUE AREA IN THE LUNG BEING INVESTIGATED FOR FURTHER DIAGNOSES. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON 12 NOV, 2024 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND THE PATIENT HAD A CT SCAN DONE THAT REVEALED AN OPAQUE AREA IN THE LUNG BEING INVESTIGATED FOR FURTHER DIAGNOSES. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04579 | 2021-09-27 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUNDABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PARTICLES IN THE DEVICE AND SHORTNESS OF BREATH RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. SECTION H7 AND H9 CORRECTED IN THIS REPORT. THE MANUFACTURER PREVIOUSLY NOT CAPTURED, PARTICLES IN THE DEVICE AND SHORTNESS OF BREATH IN SECTION B5 AND IT HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04580 | 2021-09-27 | 2021-08-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK PARTICLES IN FILTER RELATED TO TROLOGY VENTILATOR'S SOUND ABATEMENT FOAM. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK PARTICLES IN FILTER RELATED TO TROLOGY VENTILATOR'S SOUND ABATEMENT FOAM. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04497 | 2021-09-27 | 2021-09-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR HAD POWER SOURCE ISSUES. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE TO POWER ON WAS OBSERVED. THE DEVICE'S POWER SUPPLY NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04170 | 2021-09-27 | 2021-09-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04281 | 2021-09-27 | 2021-09-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04572 | 2021-09-27 | 2021-09-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04306 | 2021-09-27 | 2021-08-10 | I, n, j, u, r, y | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT'S BLOOD OXYGEN SATURATION LEVEL DECREASED WHILE ON A VENTILATOR. THE PATIENT WAS MANUALLY VENTILATED AND PLACED ON ANOTHER VENTILATOR WITHOUT INCIDENT. THE MANUFACTURER IS CURRENTLY INVESTIGATING THIS INCIDENT AND A FOLLOW-UP REPORT WILL BE FILED WHEN THE INVESTIGATION IS COMPLETE. | IT WAS INITIALLY REPORTED, THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT'S BLOOD OXYGEN SATURATION LEVEL DECREASED WHILE ON A VENTILATOR. THE PATIENT WAS MANUALLY VENTILATED AND PLACED ON ANOTHER VENTILATOR WITHOUT INCIDENT. THE DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE PASSED ALL TESTING AND WAS FOUND TO OPERATE AND ALARM AS DESIGNED. | Y | N | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04499 | 2021-09-27 | 2021-09-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 | 1054260 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A "SERVICE REQUIRED" ALARM CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE BY THE CUSTOMER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE CUSTOMER REPLACED THE DEVICE'S INTERNAL BATTERY TO ADDRESS THE ISSUE. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE DEVICE FAILED TEST STEPS DURING TESTING. THE DEVICE'S SENSOR BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUES. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04428 | 2021-09-27 | 2021-09-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2200X11B | DS2200X11B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S MACHINE FLOW SENSOR NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04500 | 2021-09-27 | 2021-09-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04429 | 2021-09-27 | 2021-09-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | FR2110X14B | FR2110X14B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04426 | 2021-09-27 | 2021-09-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2200X11B | DS2200X11B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, THE DEVICE'S LCD SCREEN WAS PHYSICALLY DAMAGED AND UNABLE TO BE READ. THE DEVICE'S DISPLAY BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04424 | 2021-09-27 | 2021-09-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | FR2110X14B | FR2110X14B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SOFTWARE WAS RELOADED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04470 | 2021-09-27 | 2021-09-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04427 | 2021-09-27 | 2021-09-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | DS2200X11B | DS2200X11B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04425 | 2021-09-27 | 2021-09-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | IN2200X15B | IN2200X15B |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, A "VENTILATOR INOPERATIVE" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S SYSTEM BOARD NEEDS TO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04576 | 2021-09-27 | 2021-09-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THERE WAS NO PATIENT HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE FAILED A TEST STEP DURING TESTING. THE DEVICE'S SENSOR BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04507 | 2021-09-27 | 2021-08-30 | I, n, j, u, r, y | Manufacturer report | BIPAP AUTOSV ADV | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | DS960HS | DS960HS |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON MAY 18, 2024, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT WAS ALLEGING TO DEVELOP CANCER. NO MEDICAL INTERVENTION WAS REPORTED. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGED THYROID CANCER, HASIMOTO . SECTION H6 HEALTH EFFECTS: CLINICAL CODES WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04476 | 2021-09-27 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | SYSTEM ONE BIPAP AUTOSV ADV | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC | DS960HS | DS960HS |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF LITTLE BLACK AND WHITE PARTICLE AND PARTICLE IN TUBING CHAMBER RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS DIZZY, LIGHT HEADED, HARD TIME BREATHING THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 CORRECTED AND UPDATED IN THIS REPORT | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04574 | 2021-09-27 | 2021-09-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04479 | 2021-09-27 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP PRO BIFLEX | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DS650HS | DS650HS |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON APR 07, 2025, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES AND DEVICE EMITS ELICTRICAL ODOR RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTION H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04498 | 2021-09-27 | 2021-09-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | RJP2100X16B | RJP2100X16B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WAS NOT CHARGING ITS BATTERY TO FULL CAPACITY. THERE WAS NO HARM OR INJURY REPORTED. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S BATTERY CHARGE LIMIT WAS RESET BY RECALIBRATING THE DEVICE AND THE SYSTEM BOARD WILL ALSO BE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04573 | 2021-09-27 | 2021-09-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE OF THE DEVICE'S LCD SCREEN WAS OBSERVED. THE DEVICES LCD DISPLAY NEEDS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04501 | 2021-09-27 | 2021-08-12 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04457 | 2021-09-27 | 2021-08-05 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION EXPERT W/HUM/HT/CELL, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | CAX501T12C | CAX501T12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED ASTHMA WITH SECRETIONS, BLOOD COMING FROM THE NOSE AND ALLERGIES. THE PATIENT DID NOT REPORT RECEIVING MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED ASTHMA WITH SECRETIONS, BLOOD COMING FROM THE NOSE AND ALLERGIES. THE PATIENT DID NOT REPORT RECEIVING MEDICAL INTERVENTION. THE REPORTED EVENT AND ITS SEVERITY WERE REVIEWED BY THE MANUFACTURER'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. HENCE, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO THE ALLEGED SERIOUS INJURY. UPON REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION, THE CUSTOMER RESPONDED THAT THEY COULD NOT RETURN THE DEVICE AS THE CUSTOMER IS NOW IN CANADA. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. IN SECTION D4 UNIQUE IDENTIFIER (UDI) NUMBER HAS BEEN CORRECTED / UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04436 | 2021-09-27 | 2021-07-12 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE LUNG CANCER. THE PATIENT DID NOT REPORT RECEIVING ANY MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON OCT 09, 2023, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED LUNG CANCER. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGED CANCER AND LUNG CANCER. THE MANUFACTURER PREVIOUSLY SUBMITTED MDR 2518422-2021-04436-1 WITH INCORRECT SECTIONS B1, B2, H1, H6. CORRECTIONS TO PREVIOUS MDR ARE MADE IN THIS REPORT AS FOLLOWS. SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY THE PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECTS: CLINICAL CODES AND HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04434 | 2021-09-27 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE A COUGH. THE PATIENT DID RECEIVE ANTIBIOTICS AND STEROIDS AS MEDICAL INTERVENTION FOR COUGH.. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 07, 2024. SECTION H11 SHOULD BE REPORTED AS THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED ALLEGATION OF AN ISSUE RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE A COUGH. THE PATIENT DID RECEIVE ANTIBIOTICS AND STEROIDS AS MEDICAL INTERVENTION FOR COUGH. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM (ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR). SECTION B2 WAS CHANGED TO BLANK (PREVIOUSLY IT WAS OTHER SERIOUS). SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION . SECTION H6 HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04481 | 2021-09-27 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04504 | 2021-09-27 | 2021-09-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | RJPX500S16 | RJPX500S16 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED RUNNY NOSE, NASAL/THROAT IRRITATION OR SORENESS. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL AND INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND THAT THERE WAS DEBRIS ON TOP OF THE INSIDE OF THE HUMIDIFIER, CONSISTENT WITH MOLD, DEBRIS ON BOTTOM OF THE INSIDE OF THE DEVICE, CONSISTENT WITH MOLD, DEBRIS ON INSIDE OF THE BACK PANEL CONSISTENT WITH MOLD, UNKNOWN BLACK AND WHITE DEBRIS ON THE HUMIDIFIER VALVE TO DEVICE, CONSISTEN WITH KERATIN, DEBRIS UNDER RIGHT FLAP OF DEVICE BY SD INSERT AND FILTER, UNKNOWN BROWN DEBRIS AND HAIR ON THE INSIDE OF BOTTOM OF DEVICE, BLACK RESIDUE LOCATED ON BLOWER OUTLET ON PLASTIC, CONSISTENT WITH KERATIN, WHITE DUST INSIDE BLOWER EXIT. WERE NO SIGNS OF CONTAMINATION ON THE OUTSIDE OF THE DEVICE. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER AND FOUND THAT THERE WAS EVIDENCE OF TAP WATER BEING USED IN THE DEVICE. THERE WERE MINERAL DEPOSITS ON THE INSIDE OF THE BLOWER BOX AS WELL AS ON THE BLOWER MOTOR. THERE ARE ALSO SIGNS THAT THERE WAS WATER INGRESS IN THE BLOWER BOX. THERE IS NO VISIBLE DAMAGE OR FUNCTIONALITY FAILURES OF THE DEVICE, WHICH SUGGEST THAT THE SOURCE OF CONTAMINATION WAS EXTERNAL TO THE DEVICE. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION H6 WERE UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04503 | 2021-09-27 | 2021-08-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURYTHIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. SECTION B5 SHOULD BE PREVIOUSLY REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED DEVELOPED THROAT AND CHEST PAIN AND IRRITATION THAT SEEM TO BE GETTING WORSE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 (HEALTH EFFECT - CLINICAL CODE) HAS BEEN CORRECTED AND (TYPE OF INVESTIGATION,INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS) HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04509 | 2021-09-27 | 2021-08-30 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A SINUS INFECTION. THE PATIENT WAS PRESCRIBED ANTIBIOTICS IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGED TO TO DEVELOP A SINUS INFECTION RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM.THERE WAS NO REPORT OF PATIENT HARM OR INJURY.THE PATIENT WAS PRESCRIBED ANTIBIOTICS IN RESPONSE TO THE REPORTED EVENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED EXTERNALLY/INTERNALLY AND UNKNOWN DUST/DIRT CONTAMINATION INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM WAS OBSERVED ON THE OUTER SURFACE,CRACKS AND CREVICES SUGGESTING A SOURCE EXTERNAL TO THE DEVICE.MARKS /FINGERPRINTS FOUND ON THE BOTTOM SURFACE OF THE PCA. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THEY CONFIRMED CONTAMINATION ON THE BLOWER BOX AND THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04474 | 2021-09-27 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. SECTION B5 SHOULD BE PREVIOUSLY REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING/CHAMBER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04475 | 2021-09-27 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04520 | 2021-09-27 | 2021-12-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION FROM VOLUNTARY MEDWATCH ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO EXPERIENCE A TROUBLE BREATHING, SPITTING OUT BLACK/GRAY STUFF THROUGHOUT THE NIGHT, THROWING UP BLACK TINTED PHLEGM, EVERY MORNING AFTER USING CPAP, FEELING LIKE CHOKING, NOT FEELING LIKE SLEPT AND HAVING TO INCREASE EMERGENCY ASTHMA INHALER USAGE, NOSEBLEEDS, BLACK DEBRIS STARTED TO GATHER AROUND MOUTH, GUMS AND SALIVA AND MAKING FEEL SO SICK. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM (ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR). SECTION B2 WAS MADE BLANK (PREVIOUSLY IT WAS OTHER SERIOUS). SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 CODE HEALTH EFFECT- IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04486 | 2021-09-27 | 2021-03-31 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04521 | 2021-09-27 | 2021-09-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION GO AUTO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | GBG502S15 | GBG502S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04401 | 2021-09-27 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUMID/HT, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX700T12 | CAX700T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED DIFFICULTY BREATHING, FILTER ON THE SIDE OF THE MACHINE TURNS COMPLETELY BLACK. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. EXTERNAL EXAMINATION OBSERVED FINE, WHITE POWDER IN AND AROUND THE AIR INLET AREA. PIL LOCATED WHAT APPEARS TO BE A DRIED LIQUID AROUND THE MODEM AREA. PIL RETRIEVED THE ERROR LOG WHICH CONTAINED 0 ERRORS. PIL USED THE FITT (FOAM INTEGRITY TEST TOOL) TOOK PER FITT DOCUMENT (ER 2244873 V1) TO TEST THE SOUND ABATEMENT FOAM FOR DEGRADATION AND BREAKDOWN. PIL FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. PIL PERFORMED AN INTERNAL INVESTIGATION AND LOCATED A GREY CONTAMINANT ON THE TOP ENCLOSURE, FRONT PANEL, AND BOTTOM ENCLOSURE. A FINE, WHITE POWDER WAS FOUND ON THE PCA, BLOWER CAP SEAL AND IN THE BLOWER BOX CHIMNEY. A FINE, WHITE POWDER WAS FOUND ON THE BLOWER BOX, BLOWER, BLOWER BLADES, ISO PORT, AND REAR PANEL ISO PORT SEAL. WATER INGRESS WAS LOCATED AT THE BOTTOM OF THE BLOWER AND IN THE BOTTOM OF THE BLOWER BOX. A DARK CONTAMINANT THAT APPEARS CONSISTENT WITH POTENTIAL KERATIN OBSERVED IN ER 2243857 (V1) WAS FOUND AT THE BLOWER SEAL OF THE BLOWER BOX AND IN THE BOTTOM OF THE BLOWER BOX. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04400 | 2021-09-27 | 2021-08-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION EXPERT W/HUM/HT/CELL, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX501T12C | CAX501T12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. SECTION B5 SHOULD BE PREVIOUSLY REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES IN AIRWAY FROM DEVICE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT HAS APNEA AND BREATHING ISSUES. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED. SECTION D4 UNIQUE IDENTIFIER (UDI) # CORRECTED IN THIS REPORT. SECTION G1 CONTACT OFFICE (AND MANUFACTURING SITE FOR DEVICES) OR COMPOUNDING OUTSOURCING FACILITY CORRECTED IN THIS REPORT. SECTION H6 HEALTH EFFECT - CLINICAL CODE CORRECTED. SECTION H6 TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04505 | 2021-09-27 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES,BURNING ODOR AND OVERHEATING OF DEVICE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM.THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. THE MANUFACTURER PREVIOUSLY NOT CAPTURED, VISUALIZATION OF PARTICLES,BURNING ODOR AND OVERHEATING OF DEVICE IN SECTION B5 AND IT HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04399 | 2021-09-27 | 2021-09-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUMID/HT, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | AUX400T15 | AUX400T15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04480 | 2021-09-27 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP AUTOSV | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | DSX900H11C | DSX900H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING/CHAMBER AND SINUS INFECTION RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04526 | 2021-09-27 | 2021-07-07 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION BIPAP AUTOSV H/HT/C, DS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX900T11C | DSX900T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SINUS AND RESPIRATORY INFECTIONS. THE PATIENT WAS PRESCRIBED ANTIBIOTICS AND STEROIDS IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SINUS AND RESPIRATORY INFECTIONS. THE PATIENT DID NOT REPORT TO RECEIVE MEDICAL INTERVENTION. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04435 | 2021-09-27 | 2021-07-07 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX400H11 | DSX400H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE BRONCHITIS, SHORTNESS OF BREATH, HEADACHES AND AN INFECTION. THE PATIENT DID RECEIVE ANTIBIOTICS AND STEROID AS MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE BRONCHITIS, SHORTNESS OF BREATH, HEADACHES AND AN INFECTION. THE PATIENT DID RECEIVE ANTIBIOTICS AND STEROID AS MEDICAL INTERVENTION. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04523 | 2021-09-27 | 2021-07-07 | D, e, a, t, h | Manufacturer report | DREAMSTATION ST30 H/HT, CN-NB | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS INC | CNX1030T17 | CNX1030T17 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO BE HOSPITALIZED. THE PATIENT EXPIRED. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. SECTION H1 HAS BEEN CORRECTED TO REFLECT THE INFORMATION REPORTED IN B2 AND "DEATH" WAS ACCURATLEY CHECKED IN SECTION H1. THIS ACTION WAS REPORTED TO FDA PER 21 CFR PART 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5104242 | 2021-09-27 | 2021-08-16 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | AS RESULT OF USING PHILIPS CPAP DREAMSTATION, NOTED BLACK RESIDUE TO MACHINE INSIDE. LARYNGITIS, BRONCHITIS, SINUS INFECTION, REQUIRING ANTIBIOTICS, INHALER AND STEROIDS. TO DATE (B)(6) 2021 UNRESOLVED. STILL UNDER MD CARE. |  | Y | Y | I |  | 2021-09
2518422-2021-04440 | 2021-09-27 | 2021-01-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE CYSTS ON THYROID, KIDNEY/LIVER DAMAGE, HEMANGIOMAS TUMOR ON SPINE, NERVE DAMAGE, LEG NUMBNESS AND LESIONS ON LEFT KIDNEY. THE PATIENT DID NOT REPORT RECEIVING ANY MEDICAL INTERVENTION THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO EXPERIENCE CYSTS ON THYROID, KIDNEY/LIVER DAMAGE, HEMANGIOMAS TUMOR ON SPINE, NERVE DAMAGE, LEG NUMBNESS, LESIONS ON LEFT KIDNEY, COUGHING UP PHLEGM, ACID REFLUX AND PRESCRIBED PROTONIX, THYROID GETTING LARGER AND INHIBITS SWALLOWING 2 MASSES BOTH BENIGN, WORSENED ABILITY TO SLEEP, SIT AND LIE DOWN, SEEING UROLOGIST, NEUROLOGIST, AND NEUROSURGEON, AND PHYSICAL THERAPIST. THE REPORTED EVENT OF ACID REFLUX AND PRESCRIBED PROTONIX, THYROID GETTING LARGER AND INHIBITS SWALLOWING 2 MASSES BOTH BENIGN, WORSENED ABILITY TO SLEEP, SIT AND LIE DOWN, SEEING UROLOGIST, NEUROLOGIST, AND NEUROSURGEON, AND PHYSICAL THERAPIST AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04466 | 2021-09-27 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP ST30 H, DS | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | DSX1030H11 | DSX1030H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP ANAPHYLACTIC SHOCK. THE PATIENT REQUIRED EPIPEN IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED. THE PATIENT ALLEGED TO EXPERIENCE ANAPHYLACTIC SHOCK, NASAL / THROAT IRRITAION AND SORENESS, DIFFICULTY BREATHING, LIGHTHEADEDNESS, DIZZINESS, HEADACHE. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENT OF ANAPHYLACTIC SHOCK, NASAL / THROAT IRRITAION AND SORENESS, DIFFICULTY BREATHING, LIGHTHEADEDNESS, DIZZINESS, HEADACHE AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THE PATIENT DID RECEIVE MEDICAL INTERVENTION. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED ANAPHYLACTIC SHOCK, NASAL / THROAT IRRITAION AND SORENESS, DIFFICULTY BREATHING, LIGHTHEADEDNESS, DIZZINESS, HEADACHE. THE PATIENT REQUIRED EPIPEN IN RESPONSE TO THE REPORTED EVENT. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04467 | 2021-09-27 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP TUMORS IN THEIR BREAST. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 8, 2024, AND SECTION H11 SHOULD BE REPORTED AS THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER WAS INITIALLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT ALLEGED TUMORS IN THEIR BREAST. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THERE IS NO PATIENT HARM OR INJURY. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM (ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR) SECTION B2 WAS CHANGED TO BLANK (PREVIOUSLY IT WAS OTHER SERIOUS) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION SECTION H6 HEALTH EFFECT - CLINICAL CODE, HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04468 | 2021-09-27 | 2021-08-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A RASH. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 07 2024 AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO DEVELOP A RASH, DRY COUGH, DRY MOUTH AND EXTREMELY SHORT WINDED FROM THE DEVICE. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENTS OF TO DEVELOP A RASH, DRY COUGH, DRY MOUTH AND EXTREMELY SHORT WINDED AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THESE EVENTS ARE ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURER CONCLUDES NO FURTHER ACTION IS NECESSARY. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04447 | 2021-09-27 | 2021-08-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL AND INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN. THE DEVICE'S EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND THE ERROR LOG SHOWED, 8 INSTANCES OF: E-53, 6 INSTANCES OF: E-191 AND 1 INSTANCE OF: E-130. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. SECTION D9, G3, H3 AND H6 HAS BEEN UPDATED / CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04485 | 2021-09-27 | 2021-08-03 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400T11C | DSX400T11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103208) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP BRONCHITIS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED AND ADDED TO THE REPORT. THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED BRONCHITIS, DIFFICULTY BREATHING, HEADACHE, NAUSEA, DRY COUGH, DIZZY,VOMITING . THE PATIENT DID NOT REPORT RECEIVING ANY MEDICAL INTERVENTION. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY ON (B)(6) 2023 FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED ON (B)(6) 2023. THE MANUFACTURE INSPECTED INTERNALLY AND EXTERNALLY OBSERVED THAT CRACK ON THE UPPER ENCLOSURE, WHITE POWDER IN AND AROUND THE AIR INLET CANAL, DARK DUST/DIRT AROUND THE ISO PORT. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND ONE ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS SOUND ABATEMENT FOAM DEGRADATION THROUGH OUT THE BLOWER BOX. AND THERE IS A BLACK TAR-LIKE SUBSTANCE, WHITE POWDER, SLIGHT YELLOW DISCOLORATION IN BLOWER BOX. SECTION D1, D4, D8, D9, G4, H2, H3, H4 AND H6 WERE UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04469 | 2021-09-27 | 2021-08-06 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11 | DSX700T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A COUGH. THE PATIENT WAS PRESCRIBED ANTIBIOTICS AND STEROIDS IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND DEVELOP A COUGH. THE PATIENT DID NOT REPORT TO RECEIVE MEDICAL INTERVENTION. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED DUE TO AGE. SECTION H6 WERE UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04452 | 2021-09-27 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION B5 SHOULD BE REPORTED AS THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES, HEADACHE, SNEEZING RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04453 | 2021-09-27 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11 | DSX400H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE 2 ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. UNIT SCRAPED DUE TO AGE. SECTION H6 WAS UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04454 | 2021-09-27 | 2021-08-03 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP,AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING M IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT IS ALSO ALLEGING LUNG DAMAGE, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, CHRONIC BRONCHITIS, IRRITATION TO EYES, COUGH,CHEST PRESSURE, SINUS INFECTIONS, AND DIFFICULTY BREATHING. WHILE USING THE DEVICE. PATIENT ALSO ALLEGED VISUALIZATION OF DIRTY/BLACK FILTERS IN THE DEVICE. THIS NEW INFORMATION HAS MADE THIS A SERIOUS INJURY ALLEGATION.RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM.THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04448 | 2021-09-27 | 2021-09-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED ON THIS DEVICE IN MDR 2518422-2021-03980 . THIS REPORT WAS SUBMITTED AS A DUPLICATE REPORT OF THE PREVIOUSLY SUBMITTED REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04494 | 2021-09-27 | 2021-08-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. IN PREVIOUS REPORT SECTION B5 WAS INCORRECT, CORRECT B5 SHOULD BE - THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103195) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL AND INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO CONTAMINATION. AN INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND SLIGHT DISCOLORATION TO THE PANELS IN THE CASE. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND THREE INSTANCES OF ERRORS. THE DEVICE WAS APPLIED POWER AND THE DEVICE OPERATED PROPERLY. THE MANUFACTURER CONCLUDES SLIGHT DISCOLORATION TO THE PANELS IN THE CASE. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. IN THIS REPORT, SECTION D9, G3, H3, H6 HAS BEEN UPDATED OR CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04495 | 2021-09-27 | 2021-09-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP AVAPS30 | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | DSX1130T11C | DSX1130T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04496 | 2021-09-27 | 2021-09-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP AVAPS30 | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | DSX1130T11C | DSX1130T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04506 | 2021-09-27 | 2021-07-07 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED LUNG CANCER. THE PATIENT DID NOT REPORT TO RECEIVE MEDICAL INTERVENTION. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION D4 AND H4 UPDATED AS ADDITIONAL INFORMATION RECEIVED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04439 | 2021-09-27 | 2021-07-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE NASAL/THROAT IRRITATION AND SORENESS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF RECONSTRUCTIVE THROAT SURGERY AND TRACHEOSTOMY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). | ADDITIONAL INFORMATION WAS RECEIVED ON OCTOBER 09, 2023, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE NASAL/THROAT IRRITATION AND SORENESS. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF RECONSTRUCTIVE THROAT SURGERY AND TRACHEOSTOMY. THE PATIENT ALSO ALLEGED VISUALIZATION OF PARTICLES AND DIFFICULTY IN BREATHING. EXPERIENCE NASAL/THROAT IRRITATION AND SORENESS. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGED KIDNEY DISEASE/TOXICITY AND LUNG DISEASE. | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5104229 | 2021-09-27 |  | M, a, l, f, u, n, c, t, i, o, n | Voluntary report | PHILIPS RESPIRONICS DREAM STATION | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSXHCP |  |  | CPAP MACHINE PHILLIPS RESPIRING SO WAS RECALLED MONTHS AGO AND GETTING NO SATISFACTION FROM PHILLIPS AS TO WHEN THE REPAIR OR REPLACEMENT WILL BE COMPLETED. THEY ARE SAYING THAT THEY ARE WAITING FOR FDA APPROVAL ON A REPAIR/REPLACEMENT. THIS IS ABSOLUTELY RIDICULOUS THAT IT WOULD TAKE THIS LONG TO GET ANY SORT OF RESPONSE, SO I'M GOING TO YOU TO FIND OUT WHAT THE HOLDUP IS, AND WHEN MY HUSBAND CPAP MACHINE CAN BE REPLACED OR REPAIRED, IT HAS BEEN MONTHS. MY HUSBAND OBVIOUSLY NEEDS THIS MACHINE AND CANNOT GO WITHOUT IT AND DOES HE HAVE LUNG ISSUES SO TO HAVE TO WAIT THIS LONG IS UNACCEPTABLE. FDA SAFETY REPORT ID# (B)(4). |  | N | Y | I |  | 2021-09
2518422-2021-04502 | 2021-09-27 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS MISSED TO CAPTURE IN SECTION B5, IT SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED OF BLOCKAGE ON HIS NOSE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 (CLINICAL CODES) HAS BEEN CORRECTED IN THIS REPORT. SECTION H6 (TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS) HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04356 | 2021-09-28 | 2021-09-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE WIRE HARNESS WAS OBSERVED TO BE DISCONNECTED. THE WIRE HARNESS WAS REATTACHED AND THE DEVICE WAS RECALIBRATED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04355 | 2021-09-28 | 2021-09-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EVO | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | DS2110X11B | DS2110X11B |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR INOPERATIVE CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S MACHINE FLOW SENSOR AND SYSTEM BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04543 | 2021-09-28 | 2021-09-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY EV300 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | DS2200X11B | DS2200X11B |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ISSUE RELATED TO THE MACHINE FLOW SENSOR WAS OBSERVED. THE DEVICE'S MACHINE FLOW SENSOR WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04354 | 2021-09-28 | 2021-09-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054096 | 1054096 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY AND POWER MANAGEMENT BOARD WERE REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04544 | 2021-09-28 | 2021-09-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040007 | 1040007 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A SERVICE REQUIRED CONDITION OCCURRED. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S OXYGEN BLENDING MODULE WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04598 | 2021-09-28 | 2021-09-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THERE WAS NO HARM OR INJURY REPORTED. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD AND MANIFOLD NEED REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04599 | 2021-09-28 | 2021-09-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1054096 | 1054096 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WOULD NOT CYCLE PROPERLY. THE DEVICE WAS NOT IN PATIENT USE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER'S COMPLAINT WAS CONFIRMED. THE DEVICE'S POWER MANAGEMENT BOARD WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04458 | 2021-09-28 | 2021-09-17 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION CPAP PRO, GER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DEX400S13 | DEX400S13 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSING A HEART ATTACK. THE PATIENT DID RECEIVE MEDICAL INTERVENTION REQUIRING HOSPITALIZATION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | THIS CASE WAS PREVIOUSLY REPORTED IN ERROR IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THIS WAS INCORRECT AND FOLLOWING FURTHER REVIEW OF THE COMPLAINT INFORMATION AND INVESTIGATION THIS CASE IS SHOULD NOT BE CONSIDERED AS PART OF THE VOLUNTARY RECALL. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE DREAMSTATION DEVICE WAS SET TO 12 MBAR DURING CUSTOMER USE, HOWEVER THE OUTPUT OF THE DEVICE WAS STATED TO ONLY DELIVER AT 8-9MBAR. THE PATIENT WAS ALLEGED TO SUFFER A HEART ATTACK, OCCURRENCE DATE AND TIME WERE NOT PROVIDED, AND REQUIRED MEDICAL INTERVENTION. LIMITED INFORMATION HAS BEEN PROVIDED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. IN THE INITIAL REPORT, SECTIONS H7 AND H9 WAS CAPTURED INCORRECTLY AND IT SHOULD BE BLANK. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04550 | 2021-09-28 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING FOAM PARTICLES IN DEVICE AND LOUD RINGING NOISE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION B5 HAS BEEN CORRECTED AND SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING FOAM PARTICLES IN DEVICE AND LOUD RINGING NOISE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04546 | 2021-09-28 | 2021-09-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04547 | 2021-09-28 | 2021-07-07 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON MAY 22, 2024, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED ABOUT THE ALLEGED NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, CHRONIC COUGHING, STRESS ON LUNG TRANSPLANT OR HEART. THE SECTIONS B1, B2, H1, AND H6 HAVE BEEN UPDATED OR CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY THE PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECTS: CLINICAL CODES AND HEALTH EFFECTS - IMPACT CODE WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04549 | 2021-09-28 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING DIFFICULTY BREATHING, BLACK PARTICLES IN THE TUBING, AND A LIGHT MALFUNCTIONING RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. ADDITIONAL INFORMATION WAS RECEIVED AND ADDED TO THE REPORT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER ON (MENTION D9 DATE (04/21/2023) FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED ON (MENTION G3 DATE (06/06/2023)). THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. THE DEVICE WAS EVALUATED. THERE WAS MENTION OF DUST/DIRT CONTAMINATION INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM WAS OBSERVED THROUGHOUT DEVICE ENCLOSURE, AND AIRPATH, SUGGESTING A SOURCE EXTERNAL TO THE DEVICE, EVIDENCE OF WATER INGRESS ON THE BLOWER AND BLOWER BOX UPON EXTERNAL AND INTERNAL ASPECT OF THE DEVICE.THERE WAS NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT PIL WAS UNABLE TO ADDRESS THE ALLEGATION OF UNABLE TO ADDRESS THE SYMPTOMS DESCRIBED. PIL CAN CONFIRM THE PRESENCE OF CONTAMINATION IN THE AIRPATH. PIL DID NOT OBSERVE ANY LIGHT MALFUNCTIONING. SECTION H6 WERE UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04545 | 2021-09-28 | 2021-09-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11 | DSX700T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04542 | 2021-09-28 | 2021-09-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 202 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC. | 1040000 | 1040000 |  | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, THE DEVICE FAILED TO POWER ON WITH AC POWER. THE POWER SUPPLY WAS REPLACED DUE TO DAMAGE TO COMPONENT U1. DURING THE EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04353 | 2021-09-28 | 2021-09-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 200 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC. | 1054260 | 1054260 |  | DURING A CHECK-OUT PROCEDURE AT THE MANUFACTURER'S SERVICE CENTER, A VENTILATOR'S LCD SCREEN WAS OBSERVED TO BE BLANK. THE DEVICE WAS NOT IN PATIENT USE. THE DEVICE'S LCD SCREEN WAS REPLACED TO ADDRESS THE ISSUE. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04548 | 2021-09-28 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED AFFECTING HIS SINUSES AND PARTICLES ON FILTERS. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO CONTAMINATION. SECONDARY FINDINGS INCLUDED CONTAMINATION ON TOP ENCLOSURE, BOTTOM ENCLOSURE, FRONT PANEL, AND REAR PANEL. AN INTERNAL INVESTIGATION SHOWED THAT THERE WERE SIGNS OF AN UNKNOWN DARK CONTAMINATION ON THE FILTER PORT. THERE WERE NO SIGNS OF SOUND ABATEMENT FOAM DEGRADATION IN THE DEVICE. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE - THE PATIENT ALLEGED AFFECTING HIS SINUSES AND PARTICLES ON FILTERS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04575 | 2021-09-28 | 2021-09-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY 100 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054655 | 1054655 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S POWER MANAGEMENT BOARD WAS REPLACED TO ADDRESS THE ISSUE. | THE MANUFACTURER PREVIOUSLY REPORTED A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE AND DURING THE EVALUATION OF THE DEVICE, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG AND THE DEVICE'S POWER MANAGEMENT BOARD WAS REPLACED TO ADDRESS THE ISSUE. SERVICE. THE DEVICE'S SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04461 | 2021-09-28 | 2021-07-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG NODULES. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM AND CAUSED A PATIENT TO DEVELOP LUNG NODULES. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04460 | 2021-09-28 | 2021-07-07 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CANCER. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE /RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM THAT BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE CANCER. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL AND INTERNAL PARTS OF THE DEVICE. EXTERNAL INVESTIGATION OBSERVED INDETERMINATE CONTAMINANTS ON THE BOTTOM ENCLOSURE, ALONG AND BEHIND THE FLIP-COVERS, AND ALONG THE BACK EDGE OF THE UPPER ENCLOSURE. INTERNAL INVESTIGATION OBSERVED DUST ON AND IN THE BLOWER. PIECES OF DEGRADED SOUND ABATEMENT FOAM WERE OBSERVED IN THE BLOWER BOX AND SIGNIFICANT SAGGING OF THE FOAM WAS OBSERVED. DARK CONTAMINANT WAS OBSERVED AROUND THE BLOWER SEAL WHICH APPEARS CONSISTENT WITH THE KERATIN CONTAMINATION OBSERVED IN (B)(4). THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND ONE ERROR CODE. THE MANUFACTURER CONCLUDES THEY COULD CONFIRM THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM IN THE AIRPATH AND ALSO OBSERVED UNIDENTIFIED EXTERNAL CONTAMINATION AND CONTAMINATION ON AND IN THE BLOWER. SECTION H6- TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04551 | 2021-09-28 | 2021-09-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | INX700S15 | INX700S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04462 | 2021-09-28 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A MYOCARDIAL INFARCTION. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | AFTER REVIEW, IT WAS DETERMINED THAT IN PREVIOUS REPORT CORRECTION WAS MISSED TO UPDATE. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. SECTION(S) B1, B2 HAS UPDATED TO REFLECT AS PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04463 | 2021-09-28 | 2021-07-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION BIPAP AUTOSV H/C, DS | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS, INC. | DSX900H11C | DSX900H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP LUNG CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES ALSO PATIENT ALLEGED TO HAVE DIFFICULTY BREATHING AND DEVELOP LUNG CANCER RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM.THE REPORTED EVENT TO DEVELOP LUNG CANCER WAS REVIEWED BY THE PMS CLINICAL EXPERT. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS INSPECTED INTERNAL AND EXTERNALLY AND SLIGHT WHITE AND BLACK CONTAMINATION DUST LIKE ,INCONSISTENT WITH SOUND ABATEMENT FOAM AT THE AIR INPUT OF THE BLOWER BOX .LIGHT WHITE DUST LIKE CONTAMINATION ,TINY UNKNOWN BLACK , FEW TINY PIECES OF POTENTIAL HAIR AND WHITE SPECS OF CONTAMINATION ON THE BOTTOM OF THE BLOWER BOX ,WHICH SUGGEST THE SOURCE OF CONTAMINATION WAS EXTERNAL TO THE DEVICE. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WAS ONE ERROR (#193 ERR_USB_HOST_NO_DEVICE) LOGGED. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE IS DUST LIKE CONTAMINATION AND THE SOURCE OF CONTAMINATIONS WERE EXTERNAL TO THE DEVICE. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04408 | 2021-09-28 | 2021-07-07 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP AN UNSPECIFIED ILLNESS. THE PATIENT WAS HOSPITALIZED IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO THE DEVICE'S SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP AN UNSPECIFIED ILLNESS. THE PATIENT WAS HOSPITALIZED FOR THREE DAYS WITH A TERRIBLE COUGH. THE REPORTED EVENT AND ITS SEVERITY WERE REVIEWED BY THE MANUFACTURER'S CLINICAL EXPERT. BASED ON INFORMATION AVAILABLE AT THIS TIME, THERE IS NOT ENOUGH EVIDENCE TO SUPPORT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE ALLEGED SERIOUS INJURY. THERE IS INSUFFICIENT INFORMATION RELATED TO DEVICE USE AND MAINTENANCE, TIMELINE OF EVENTS, RELEVANT PAST MEDICAL HISTORY, AND MEDICAL INTERVENTION INFORMATION. UPON REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION, THE CUSTOMER RESPONDED THAT THE AFFECTED DEVICES WERE NOT AVAILABLE FOR RETURN TO PHILLIPS. SECTION H6 HEALTH EFFECT - CLINICAL CODE HAS BEEN CORRECTED AND TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04405 | 2021-09-28 | 2020-01-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103161) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP BREATHING DIFFICULTIES AND HEADACHES. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO CPAP DEVICE'S SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE DEVELOP BREATHING DIFFICULTIES AND HEADACHES. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THERE IS NO CUSTOMER INFORMATION AND WE CANNOT REACH OUT TO THE CUSTOMER, HENCE, ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION CAN NOT BE MADE AT THIS TIME. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04406 | 2021-09-28 | 2017-11-01 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION C-PAP MACHINE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103162) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP REACTIVE AIRWAY DISEASE. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. SECTION B5 SHOULD BE PREVIOUSLY REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT WAS ALLEGING TO DEVELOP REACTIVE AIRWAY DISEASE, VIRAL RESPIRATORY INFECTION AND HYPERSENSITIVITY. ALSO EXPERIENCED SYMPTOMS WHICH INCLUDED SHORTNESS OF BREATH, NON-PRODUCTIVE COUGH, PRONOUNCED FATIGUE, FLU-LIKE SYMPTOMS, DIZZY AFTER MILD EXERTION, SEVERELY SWOLLEN SINUSES WITHOUT MUCH CONGESTION, FORM OF RESPIRATORY DISTRESS, MID-FACIAL HEADACHES, AND REDUCED CARDIO-VASCULAR FUNCTIONING. IT WAS ALSO ALLEGED THAT HE PREVIOUSLY RECEIVED IMMUNIZATION THERAPY AND HAD BEEN DIAGNOSED WITH CHRONIC NON-ALLERGIC RHINITIS-SINUSITIS. PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. SINCE THERE IS NO CUSTOMER INFORMATION, THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS A FOLLOW-UP REPORT WILL BE FILED. SECTION G1 UPDATED. SECTION H6 HEALTH EFFECT - CLINICAL CODES WHICH WAS MISSED IN PREVIOUS REPORT WAS CAPTURED. SECTION H6 TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04594 | 2021-09-28 | 2021-07-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/HT/CELL, AUS | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | AUX400T15C | AUX400T15C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION B5 WAS CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT SHOULD BE REPORTED AS FOLLOWS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION PATIENT NOTICED HE FEELS DIZZY WHEN HE WAKES UP AFTER USING THE MACHINE. HE ALSO FOUND BLACK SUBSTANCE ON HIS HOSE AND MASK IN THE MORNING. FOUND INSIDE THE DEVICE. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. NO DEVICE HAS BEEN RETURNED. NO FURTHER EVALUATION IS POSSIBLE AT THIS TIME. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. SECTION H6 CODES CAPTURED INCORRECTLY IN PREVIOUS REPORT, IT HAS BEEN CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04482 | 2021-09-28 | 2021-07-07 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP HIGH BLOOD PRESSURE. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | THE MANUFACTURER PREVIOUSLY SUBMITTED MDR 2518422-2021-04482-1 WITH INCORRECT SECTIONS B1, B2, H1, H6. CORRECTIONS TO PREVIOUS MDR ARE MADE IN THIS REPORT AS FOLLOWS. SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY PRODUCT PROBLEM WAS CHECKED IN PREVIOUS MDR) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY IT WAS BLANK) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6- HEALTH IMPACT CODE WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04483 | 2021-09-28 | 2021-08-25 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION DEVICE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP MELANOMA. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM AND BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE MELANOMA OF THE FACE. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE EVENT IS CURRENTLY UNKNOWN. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD CONFIRMED THE DUST CONTAMINATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. SECTION H6 WAS UPDATED AND D4,G4,H4 WAS CORRECTED IN THIS UPDATE. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04409 | 2021-09-28 | 2021-07-22 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A BACTERIAL INFECTION. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION WAS CAPTURED INCORRECTLY IN THE PREVIOUS REPORT. IT SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A BACTERIAL INFECTION, EYE IRRITATION, AND HEADACHES. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. DURING EXTERIOR INVESTIGATION: THERE IS UNKNOWN DUST/DIRT CONTAMINATION PRESENT ON ALL SURFACES OF THE BASE UNIT. THERE IS AN UNKNOWN DUST CONTAMINATE PRESENT AT THE AIR INLET WHERE THE FILTER WOULD BE AND AT THE ISO PORT ENTRANCE. DURING INTERIOR INVESTIGATION: THERE IS UNKNOWN DEBRIS IN THE TRACKS OF THE UI PANEL. THERE IS AN UNKNOWN WHITE DUST CONTAMINATION FOUND ON THE BOTTOM ENCLOSURE, BLOWER BOX HOUSING, BLOWER MOTOR, BLOWER IMPELLER, AND REAR PANEL O-RING. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED IN THE BASE UNIT. THE UNKNOWN DUST/DIRT CONTAMINATION AND FIBERS FOUND ON THE BLOWER CASING/IMPELLER AND BLOWER BOX WAS INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM CONTAMINATION. THE PRESENCE OF CONTAMINATION THROUGHOUT THE AIRPATH SUGGESTS A SOURCE EXTERNAL TO THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES UNABLE TO DIRECTLY ADDRESS THE SYMPTOMS OUTLINED IN THE COMPLAINT. CAN CONFIRM THAT THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN OBSERVED IN THE BASE UNIT. CONFIRMED THE PRESENCE OF CONTAMINATION IN THE AIRPATH. IN SECTION H6 HEALTH EFFECTS: CLINICAL CODE WAS MISSED TO CAPTURE IN THE PREVIOUS REPORT. IT HAS BEEN UPDATED OR CORRECTED IN THIS REPORT. IN THIS REPORT, SECTION HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04319 | 2021-09-28 | 2021-08-31 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | ADDITIONAL INFORMATION WAS RECEIVED ON SEP 18, 2022 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. THE MANUFACTURER PREVIOUSLY RECEIVED MISSED ALLEGATIONS HEADACHE, FATIGUE, DIZZINESS AND COUGH. THE DEVICE WAS RETURNED TO THE THIRD PARTY SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED. THERE WERE NO ERROR FOUND. THE THIRD PARTY SERVICE CENTER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04577 | 2021-09-28 | 2021-09-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY O2 | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040004 | 1040004 |  | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTATIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S OXYGEN BLENDING MODULE BOARD, INTERFACE BOARD, AND POWER MANAGEMENT BOARD WERE REPLACED TO ADDRESS THE ISSUES. |  | N | Y | HEALTH PROFESSIONAL |  | 2021-09
2518422-2021-04552 | 2021-09-28 | 2021-09-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04422 | 2021-09-28 | 2021-09-21 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, JP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | JPX500S16 | JPX500S16 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO HAVE AN INFECTION. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTION FOR ANTIBIOTICS. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. PATIENT WAS ALLEGING TO DEVELOP HAVE AN INFECTION SUFFERING FROM SORE THROAT AND HAD YELLOW SPUTUM WHEN USING THIS UNIT. PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF PRESCRIPTION FOR ANTIBIOTICS. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED IN THIS DEVICE. CONTAMINATION UNDER THE DRY BOX WAS OBSERVED. THE EXTERNAL INVESTIGATION SHOWED THAT THERE WERE NO SIGNS OF CONTAMINATION ON THE OUTSIDE OF THE DEVICE. THE DEVICE WAS HOOKED UP TO A KNOWN GOOD POWER SUPPLY AND CORD. AIRFLOW WAS VERIFIED TO BE OPERATING CORRECTLY. THE CARE ORCHESTRATOR DATA WAS UNAVAILABLE. THE ERROR LOG SHOWED THAT THERE WERE NO ERRORS ON THE DEVICE. THE INTERNAL INVESTIGATION OF THE HUMIDIFIER SHOWED THAT THERE WERE SIGNS OF CONTAMINATION UNDER THE DRY BOX. THE INTERNAL INVESTIGATION SHOWED THAT THERE WERE NO SIGNS OF CONTAMINATION INSIDE OF THE DEVICE. INVESTIGATION FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. SECTION G1 UPDATED. SECTION D8, D9, H2, H3 AND H6 HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04433 | 2021-09-28 | 2021-07-07 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A RECURRENT SINUS INFECTION. THE PATIENT WAS PRESCRIBED ANTIBIOTICS IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A RECURRENT SINUS INFECTION. THE PATIENT WAS PRESCRIBED ANTIBIOTICS IN RESPONSE TO THE REPORTED EVENT. IN THE INITIAL REPORT SECTION B5 WAS NOT UPDATED, THE CORRECT B5 SHOULD BE - THE PATIENT HAS ALLEGED BLACK OILY STUFF, COUGH, RESPIRATORY ISSUES, SINUS INFECTION, UNDERWENT LUNG X- RAY AND TAKING ANTIBIOTICS FOR SINUS INFECTION. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF DISCOLORATION DUE TO CORROSION FROM WATER INGRESS, DUST AND FAINT BLUE-GREEN CONTAMINATION ON THE BLOWER, WATER INGRESS TO BLOWER BOX AND ON THE BLOWER. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND NO ERROR LOGGED. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH DISCOLORATION DUE TO CORROSION FROM WATER INGRESS, DUST AND FAINT BLUE-GREEN CONTAMINATION ON THE BLOWER, WATER INGRESS TO BLOWER BOX AND ON THE BLOWER. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION D9,G3,H6 HAS BEEN UPDATED/CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04442 | 2021-09-28 | 2021-08-06 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700H11C | DSX700H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO BE DIAGNOSED WITH LUNG CANCER. THE PATIENT DID NOT REPORT RECEIVING ANY MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGING SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO BE DIAGNOSED WITH LUNG CANCER. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND UNIT PASSED FINAL TEST. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04443 | 2021-09-28 | 2021-08-03 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE AN KIDNEY FAILURE, UPPER RESPIRATORY DISTRESS, SINUS INFECTION AND HEADACHES. THE PATIENT DID GO TO THE HOSPITAL TO RECEIVE MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04441 | 2021-09-28 | 2021-08-04 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | DREAMSTATION AUTO CPAP | RESPORINOC INC | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO BE DIAGNOSED WITH ASTHMA. THE PATIENT DID NOT REPORT RECEIVING MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON JUL 27, 2023, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO BE DIAGNOSED WITH ASTHMA. THE PATIENT DID NOT REPORT RECEIVING MEDICAL INTERVENTION. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04437 | 2021-09-28 | 2020-08-31 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE INFECTION OF THE COLON. THE PATIENT DID RECEIVE MEDICAL INTERVENTION IN THE FORM OF HOSPITALIZATION AND WAS PRESCRIBED ANTIBIOTICS. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5104267 | 2021-09-28 | 2021-09-01 | I, n, j, u, r, y | Voluntary report | B-PAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | CONFIRMED I WAS USING A PHILLIPS B-PAP THAT WAS RECALLED. (CONFIRMED MY ENTERING SERIAL # ON THE RECALL WEBSITE.) I'M EXPERIENCING HEADACHES, SEVERE CONGESTION, WATERY EYES, ITCHY SKIN AND EYES, SINUS INFECTION ETC. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | I |  | 2021-09
MW5104263 | 2021-09-28 | 2021-09-24 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRTRONICS CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX500H11C |  |  | HEADACHES AND STUFFY NOSE. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04530 | 2021-09-28 | 2021-07-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. |  |  |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN THE DEVICE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING PARTICLES IN THE DEVICE S RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION G3 WAS INCORRECTLY CAPTURED IN THE INITIAL REPORT. SECTION G4 [(PMA/510(K)] WAS INCORRECTLY CAPTURED IN THE INITIAL REPORT WAS CORRECTED AND SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
MW5104262 | 2021-09-28 | 2021-01-01 | I, n, j, u, r, y | Voluntary report | PHILIPS RESPIRONICS DREAMSTATION AUTOCPAPHUMHT DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | REF DSX500T11 |  |  | I PUT (B)(6) OF 2021 AS A REFERENCE POINT. I STARTED HAVING DIZZINESS/HEADACHES/FLUCTUATING BLOOD PRESSURE, ISSUES WITH COUGHING AND HAVE HAD A SLIGHTLY ELEVATED CO2 IN MY BLOOD FOR OVER 3 YEARS. I WAS TALKING WITH MY DOCTOR ABOUT THESE SYMPTOMS AND TRYING TO FIGURE OUT WHAT WAS GOING ON WHEN I RECEIVED THE RESPIRONICS CPAP RECALL. I USED A DREAM STATION FOR 5 YEARS AND ALSO USED THE SOCLEAN TO CLEAN THE MACHINE WITH THE MFG. ATTACHMENT. RECEIVING THE RECALL NOTICE WAS EXTREMELY UPSETTING FOR ME AND IN THE PROCESS OF SWITCHING MACHINES TO ONE THAT WAS NOT DEFECTIVE, I SUFFERED AN AFIB EPISODE AND WAS HOSPITALIZED. ALL RESPIRONICS AND THE FDA ARE TALKING ABOUT IS HOW TO "SWITCH OUT" A DEFECTIVE MACHINE FOR ONE THAT IS NOT DEFECTIVE HOWEVER; NO ONE IS TELLING ME HOW TO DETERMINE IF THE MACHINE I USED FOR 5 YEARS IS ACTUALLY DEFECTIVE AND MAY HAVE ACTUALLY HARMED MY HEALTH. WHAT IS THE FDA DOING ABOUT HELPING INDIVIDUALS FIND OUT IF THE RECALLED CPAP MACHINE THEY WERE USING WAS IN FACT DEFECTIVE AND MAY HAVE HARMED THEIR HEALTH? WHAT TESTS SHOULD BE DONE ON THE MACHINE AND THE INDIVIDUAL TO DETERMINE THAT NO HARM HAS COME TO THEM? DID THE OFF GASSES FROM THE RECALLED CPAP DAMAGE MY LUNGS, KIDNEYS, ETC., AND/OR, IN THE FUTURE, WILL MY HEALTH SUFFER? THE FEAR OF THE UNKNOWN IS IN ITSELF AFFECTING MY HEALTH. PLEASE PROVIDE ME WITH INFORMATION TO DETERMINE IF THE RECALLED CPAP MACHINE I USED FOR 5 YEARS WAS DEFECTIVE AND IF SO, HOW IT MAY HAVE AFFECTED MY HEALTH GIVEN THE SYMPTOMS I HAVE BEEN EXPERIENCING FOR SOMETIME NOW. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5104254 | 2021-09-28 | 2021-07-09 | I, n, j, u, r, y | Voluntary report | CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | IN (B)(6) OF 2020, EPISODE OF DOUBLE VISION. MRI SHOWED A LESION IN THE BRAIN. IN (B)(6) 2021 MRI DETECTED A GROWTH IN THE LEFT NASAL PASSAGE. ON (B)(6) 2021 GROWTH WAS REMOVED AND BIOPSIED. MARGINAL ZONE LYMPHOMAS WAS THE DIAGNOSIS. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | I |  | 2021-09
MW5104253 | 2021-09-28 | 2021-09-20 | I, n, j, u, r, y | Voluntary report | PHILIPS DREAMSTATION CPAP DEVICE | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. |  |  |  | I HAVE BEEN USING THE PHILIPS DREAMSTATION CPAP DEVICE SINCE (B)(6) 2019. THERE WAS RECENTLY A RECALL ON THIS DEVICE. I DEVELOPED A COUGH AND SHORTNESS OF BREATH LAST WEEK AND WENT TO URGENT CARE ON (B)(6) 2021 AND THE EMERGENCY DEPARTMENT ON (B)(6) 2021. FDA SAFETY REPORT ID# (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04477 | 2021-09-28 | 2021-07-26 | I, n, j, u, r, y | Manufacturer report | BIPAP A40 | VENTILATOR, CONTINUOUS, MINIMAL VENTILATORY SUPPORT, FACILITY USE |  | 1111169 | 1111169 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE UPPER NASAL ISSUES MARKED BY INTENSIVE ANTIBIOTIC THERAPY AND AN ADMISSION INTO THE INTENSIVE CARE UNIT (ICU). THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | CORRECTED DATA: E1 INITIAL REPORTER NAME AND ADDRESS: PREVIOUSLY REPORTED AS (B)(6) ; UPDATED TO (B)(6). (B)(6) G3 DATE RECEIVED BY MANUFACTURER: PREVIOUSLY REPORTED AS 21-SEP-2021 ; UPDATED TO 14-SEP-2021. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04584 | 2021-09-28 | 2021-09-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP S/T | VENTILATOR, CONTINUOUS, NON-LIFE-SUPPORTING | RESPIRONICS INC. | 1061457 | 1061457 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04363 | 2021-09-28 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | BIPAP AUTO BIFLEX | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | CA761TS | CA761TS |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES, HEADACHE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 (HEALTH EFFECT - CLINICAL CODE) CORRECTED IN THIS REPORT. SECTION H6 (TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS) UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04438 | 2021-09-28 | 2021-07-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX600H11C | DSX600H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PARTICLES IN THE CHAMBER OF THE DEVICE AND FEELS PIECES OF SOMETHING HITTING HER IN THE FACE WHEN DEVICES ON AND WAKEUP WITH HEADACHES AFTER USING DEVICE, THER HAS BEEN A BURNING SMELL THAT DOESN'T LAST LONG TWICE BEFORE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, FOLLOWUP REPORT WILL BE FILED. SECTION G1 CORRECTED, G3 WAS INCORRECTLY CAPTURED AND H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
MW5104260 | 2021-09-28 | 2021-08-12 | I, n, j, u, r, y | Voluntary report | TRILOGY VENTILATOR | VENTILATOR, CONTINUOUS, FACILITY USE | PHILIPS / RESPIRONICS, INC. |  |  |  | I WENT TO (B)(6) RADIOLOGY FOR A CT SCAN FOR ANNUAL CANCER SCREENING. A BACTERIAL PNEUMONIA WAS DERIVED. I WAS SICK FOR ABOUT A MONTH. I RECEIVED ANTIBIOTIC MEDICATION FROM MY PULMONOLOGIST WHICH EVENTUALLY CLEARED THE PROBLEM. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04410 | 2021-09-28 | 2017-01-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP EPISTAXIS, HEADACHES, NAUSEA AND TOOTH LOSS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 11, 2024 AND SECTION H11 SHOULD BE REPORTED AS THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP NOSEBLEEDS, HEADACHES, NAUSEA AND TOOTH LOSS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED TO PRODUCT PROBLEM (ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR). SECTION B2 WAS CHANGED TO BLANK (PREVIOUSLY IT WAS OTHER SERIOUS). SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION. SECTION H6 HEALTH EFFECT - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04478 | 2021-09-28 | 2021-09-03 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE NOSE BLEEDS, EXACERBATION OF PERIODONTITIS, INFLAMMATION OF THE ROOT TIP, JAW PAIN, HEADACHE, NAUSEA AND LACK OF APPETITE. THE PATIENT UNDERWENT TOOTH EXTRACTION AS A RESULT OF PERIODONTITIS. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04375 | 2021-09-28 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04465 | 2021-09-28 | 2021-07-07 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A SINUS INFECTION. THE PATIENT WAS PRESCRIBED ANTIBIOTICS IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 11, 2024 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A SINUS INFECTION. THE PATIENT WAS PRESCRIBED ANTIBIOTICS IN RESPONSE TO THE REPORTED EVENT. THE SECTIONS B1, B2, H1, H6 HAVE BEEN CORRECTED IN THIS REPORT AS FOLLOWS: SECTION B1 WAS CORRECTED FOR ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY THE PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY, IT WAS BLANK.) SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6 HEALTH EFFECTS - IMPACT CODE WAS CORRECTED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04553 | 2021-09-28 | 2021-09-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04554 | 2021-09-28 | 2021-09-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11 | DSX700T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04587 | 2021-09-28 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY IN THE INITIAL REPORT SECTION B5 WAS NOT UPDATED, THE CORRECT B5 SHOULD BE - THE PATIENT HAS ALLEGED PARTICLES IN THE DEVICE. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL VISUAL INPSECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO EVIDENCE OF CONTAMINATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF WATER INGRESS (MINERAL DEPOSITS) IN BLOWER BOX, CONTAMINATION IN BLOWER BOX, CONTAMINATION ON THE INLET PORT OF THE DEVICE. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. A HUMIDIFIER WAS ALSO RETURNED TO THE MANUFACTURER ALONG WITH DEVICE. AN INTERNAL VISUAL INSPECTION OF THE HUMIDIFIER WAS COMPLETED BY THE MANUFACTURER AND FOUND CONTAMINATION ON THE BOTTOM OF THE INSIDE OF THE HUMIDIFIER AND ON THE FLIP LID, CONTAMINATION ON THE RESERVOIR SEAL. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND E-191 ON THE ERROR LOG. THE MANUFACTURER CONCLUDES THE CONTAMINATES FOUND WERE CONSISTENT WITH WATER INGRESS (MINERAL DEPOSITS) IN BLOWER BOX, CONTAMINATION IN BLOWER BOX, CONTAMINATION ON THE INLET PORT OF THE DEVICE, CONTAMINATION ON THE BOTTOM OF THE INSIDE OF THE HUMIDIFIER AND ON THE FLIP LID, CONTAMINATION ON THE RESERVOIR SEAL OF THE HUMIDIFIER. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION D9,D10,G3,H6 HAS BEEN UPDATED/CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04586 | 2021-09-28 | 2021-08-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES, HEADACHE, DIZZINESS, SHORTNESS OF BREATH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDING TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERROR FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT SCRAPPED DUE TO AGE. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04376 | 2021-09-28 | 2020-09-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY SUBMITTED MDR 2518422-2021-04376 WITH INCORRECT SECTIONS H6. CORRECTIONS TO PREVIOUS MDR ARE MADE IN THIS REPORT AS FOLLOWS. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN HEADACHE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTION H6- HEALTH EFFECT - CLINICAL CODE WAS UPDATED. ADDITIONAL INFORMATION WAS RECEIVED ON OCT 19, 2022, AND SECTION H11 SHOULD BE REPORTED AS: THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED. THERE WERE NO ERRORS FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. DEVICE WAS CORRECTED. SECTIONS D8, D9, H2, H3 AND H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04585 | 2021-09-28 | 2021-08-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500T11W | DSX500T11W |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04364 | 2021-09-28 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX500T12 | CAX500T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. ADDITIONAL INFORMATION WAS RECEIVED AND ADDED TO THE REPORT. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER ON 04/04/2023 FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED ON 05/11/2023. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND DEVICE WAS SCRAPPED. SECTION D8, D9, H2, H3 AND H6 WERE UPDATED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04366 | 2021-09-28 | 2021-04-21 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04373 | 2021-09-28 | 2021-09-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04370 | 2021-09-28 | 2021-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04372 | 2021-09-28 | 2021-08-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04374 | 2021-09-28 | 2021-09-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED IRRITATION IN UPPER THROAT, DEVICE EMITS DUST PARTICLES WHILE USING THE DEVICE. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE THIRD PARTY SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THEY COULDN'T CONFIRM THE CUSTOMER COMPLAINT AND NO VISIBLE FOAM PARTICLES. THE DEVICE WAS SCRAPPED. SECTION G4 UDI NUMBER HAS BEEN CORRECTED IN THIS REPORT. SECTION H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04583 | 2021-09-28 | 2021-09-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04529 | 2021-09-28 | 2021-07-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION B5 WAS CORRECTED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04607 | 2021-09-29 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400T11C | DSX400T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER PREVIOUSLY REPORTED AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT ALSO ALLEGED VISUALIZATION OF PARTICLES.THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL PART OF THE DEVICE AND FOUND WHITE DUST/PARTICLES LOCATED ON THE OUTER SURFACES, CRACKS, CREVICES, SD CARD AREA, MODEM AREA, AROUND THE OUTLET PORT, AROUND THE AIR INLET AND IN THE AIR INLET CANAL, OUTER SURFACE OF THE BLOWER AND WHERE THE BLOWER RESTS ON THE BOWLER BOX. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND DIDNOT FIND ANY SIGNS OF DUST OR CONTAMINATION. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION AND DUST CONTAMINATION WERE OBSERVED AND ARE CONSISTENT WITH BEING FROM AN EXTERNAL SOURCE. IN INITIAL REPORT, SECTION G1 WRONGLY CAPTURED, AND SECTION G3 DATE WAS CAPTURED AS (B)(6)2021, SHOULD BE (B)(6) 2021, CORRECTED AND UPDATED THESE FEILDS IN THIS REPORT. SECTION H6 WERE UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04640 | 2021-09-29 | 2021-08-30 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, ND | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | NDX500S15 | NDX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP KIDNEY STONES AND PULMONARY FIBROSIS. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE /RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM THAT BECAME DEGRADED AND CAUSED THE PATIENT TO DEVELOP KIDNEY STONES AND PULMONARY FIBROSIS. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. SINCE THERE IS NO CUSTOMER INFORMATION, THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, FOLOOW-UP REPORT WILL BE FILED. SECTION D4 CORRECTED AND SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04650 | 2021-09-29 | 2021-09-22 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP, GB | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | GBX500S15 | GBX500S15 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP AN UNSPECIFIED ILLNESS. THE PATIENT HAS REQUIRED SURGERY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP AN UNSPECIFIED ILLNESS. AFTER FIRST ATTEMPT TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION, THE CUSTOMER HAS RESPONDED ON (B)(6) 2021 BUT THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04564 | 2021-09-29 | 2021-08-12 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO BE DIAGNOSED WITH BONE MARROW CANCER. THE PATIENT DID NOT REPORT RECEIVING ANY MEDICAL INTERVENTIONS. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO BE DIAGNOSED WITH BONE MARROW CANCER. THE PATIENT DID NOT REPORT RECEIVING ANY MEDICAL INTERVENTIONS. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURER CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REPORTED BY THE PATIENT. SINCE THERE IS NO CUSTOMER INFORMATION, THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04565 | 2021-09-29 | 2021-08-12 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11W | DSX700T11W |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO BE DIAGNOSED WITH BREAST CANCER. THE PATIENT DID NOT REPORT RECEIVING ANY MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04559 | 2021-09-29 | 2021-06-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS,INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE SHORTNESS OF BREATH, HEADACHES AND HIVES. THE PATIENT WAS PRESCRIBED MEDICATIONS AS A FORM OF MEDICAL INTERVENTION. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED CAUSED THE PATIENT TO EXPERIENCE SLEEP APNEA,SHORTNESS OF BREATH,HEADACHES,IRRITABLE,SKIN ISSUES,LACK OF FOCUS AND ENERGY AND HIVES. THE PATIENT WAS PRESCRIBED MEDICATIONS AS A FORM OF MEDICAL INTERVENTION.THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE.THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENT OF CAUSED THE PATIENT TO EXPERIENCE SLEEP APNEA,SHORTNESS OF BREATH, HEADACHES,IRRITABLE,SKIN ISSUES,LACK OF FOCUS AND ENERGY AND HIVES. THE PATIENT WAS PRESCRIBED MEDICATIONS AS A FORM OF MEDICAL INTERVENTION.THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE.BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION.IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04563 | 2021-09-29 | 2021-08-12 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO BE DIAGNOSED WITH CANCER. THE PATIENT DID NOT REPORT RECEIVING ANY MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY SUBMITTED MDR 2518422-2021-04563-1 WITH INCORRECT SECTIONS B1, B2, H1, H6. CORRECTIONS TO PREVIOUS MDR ARE MADE IN THIS REPORT AS FOLLOWS. SECTION B1 WAS CORRECTED TO ADVERSE EVENT AND PRODUCT PROBLEM. (ONLY PRODUCT PROBLEM WAS CHECKED IN PREVIOUS MDR). SECTION B2 WAS CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. (PREVIOUSLY IT WAS BLANK). SECTION H1 WAS CHANGED FROM MALFUNCTION TO SERIOUS INJURY. SECTION H6- HEALTH IMPACT CODE WAS UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04560 | 2021-09-29 | 2021-08-07 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX700H11 | DSX700H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE INFLAMED KNEES AND KIDNEY CANCER. THE PATIENT DID NOT REPORT RECEIVING ANY MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER (B)(4). | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. SECTION H6 INVESTIGATION FINDINGS WAS MISSED TO CAPTURE IN FOLLOW UP REPORT (2518422-2021-04560-1), NOW ITS CORRECTED AND UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04562 | 2021-09-29 | 2021-08-12 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO BIPAP/HUM/HT/CELL | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX700T11C | DSX700T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A BIPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED THE PATIENT TO EXPERIENCE PARTICLES IN THE LUNGS. THE PATIENT WENT TO THE HOSPITAL TO RECEIVE MEDICAL INTERVENTION. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A PATIENT TO EXPERIENCE PARTICLES IN THE LUNGS RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT WENT TO THE HOSPITAL TO RECEIVE MEDICAL INTERVENTION. THE REPORTED EVENT AND ITS SEVERITY WERE REVIEWED BY THE MANUFACTURER'S CLINICAL EXPERT. BASED ON INFORMATION AVAILABLE AT THIS TIME, THERE IS NOT ENOUGH EVIDENCE TO SUPPORT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE ALLEGED SERIOUS INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04662 | 2021-09-29 | 2020-09-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPORINOC INC | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MW5103155) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A SINUS INFECTION. THE PATIENT WAS PRESCRIBED MEDICATION IN RESPONSE TO THE REPORTED EVENT. THE MANUFACTURER'S INVESTIGATION IN ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON NOV 13 , 2024 AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED A VOLUNTARY MEDWATCH (MW5103155) ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A SINUS INFECTION, SEVERE NASAL IRRITATION, AND HEADACHE.THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. SECTION B1 WAS CORRECTED FOR PRODUCT PROBLEM. (BOTH ADVERSE EVENT AND PRODUCT PROBLEM WAS CHECKED IN THE PREVIOUS MDR.) SECTION B2 WAS CORRECTED TO BLANK. (PREVIOUSLY, IT WAS OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS.) SECTION H1 WAS CHANGED FROM SERIOUS INJURY TO MALFUNCTION SECTION H6 WAS CORRECTED AND UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04689 | 2021-09-29 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED PARTICLES IN AIRWAY, BREATHING PROBLEMS, DRY MOUTH,. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR FURTHER INVESTIGATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04688 | 2021-09-29 | 2021-06-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. IN PREVIOUS REPORT SECTION B5 WAS INCORRECT, CORRECT B5 SHOULD BE THE PATIENT ALLEGED VISUALIZATION OF BLACK PARTICLES, BURNING IN HIS SINUS CAVITY AND THROAT, COUGH. THERE WAS NO REPORT SERIOUS OR PERMANENT PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL AND INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND NO CONTAMINATION. AN INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND SIGNS OF SOUND ABATEMENT FOAM DEGRADATION WITHIN THE DEVICE, DEGRADED FOAM STUCK TO THE END OF THE FITT TOOL. THE MANUFACTURER RECEIVED HUMIDIFIER AND FOUND CONTAMINATION ON THE BOTTOM OF THE HUMIDIFIER AND UNDER THE DRY BOX. THE MANUFACTURER FOUND EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED BY THE MANUFACTURER AND FOUND THREE INSTANCES OF ERRORS. THE DEVICE WAS APPLIED POWER AND THE DEVICE OPERATED PROPERLY. THE MANUFACTURER CONCLUDES CONTAMINATION FOUND WERE CONSISTENT WITH CONTAMINATION ON THE BOTTOM OF THE HUMIDIFIER AND UNDER THE DRY BOX. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. IN THIS REPORT, SECTION D9, D10, G3, H3, H6 HAS BEEN UPDATED OR CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04687 | 2021-09-29 | 2021-09-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700T11 | DSX700T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04613 | 2021-09-29 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. . | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04658 | 2021-09-29 | 2021-09-28 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04659 | 2021-09-29 | 2021-09-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION EXPERT | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX501H12 | CAX501H12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04661 | 2021-09-29 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX600H11C | DSX600H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT STATES OVER TIME HAS HAD ADDITIONAL PROBLEMS WITH BREATHING, SINUS PROBLEMS, SPOKE WITH DR AND HE HAS ORDER A CHEST XRAY DOES HAVE PARTICLES IN THE WATER AND TUBE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED AND FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION OR BREAKDOWN. LIMITED INVESTIGATION WAS PERFORMED DUE TO THE DISCOVERY OF INSECT CONTAMINATION. DUST/DIRT CONTAMINATION BUILDUP AROUND THE AIR INLET. 2,655.1 BLOWER HOURS, AND 2,605.3 THERAPY HOURS WERE LOGGED. 2,655.1 MACHINE HOURS WERE LOGGED. 26 INSTANCES OF STOP ERROR E-200 (ERR_RTOS_ABORT_ERROR) WERE LOGGED PRIOR TO PIL RECEIPT AND 5 INSTANCES OF REBOOT ERROR E-190 (ERR_LCD_ID_ERROR) AS WELL AS 1 INSTANCE OF REBOOT ERROR E-20 (ERR_MOTOR_SPINUP_FLUX_LOW) WERE LOGGED DURING TESTING. RECOMMENDED ACTION IF THE UNIT WAS NOT INFESTED WITH BUGS WOULD BE TO REPLACE THE PCA. CARE ORCHESTRATOR DATA WAS UNAVAILABLE. CONTAMINATION APPEARING CONSISTENT WITH INSECT INFESTATION WAS OBSERVED ON THE INTERIOR OF THE DEVICE ENCLOSURE AND BLOWER BOX. EVIDENCE OF WATER INGRESS OBSERVED IN THE BLOWER BOX, SUGGESTING THE USE OF NON-DISTILLED WATER IN THE HUMIDIFIER WATER CHAMBER. DARK CONTAMINATION APPEARING CONSISTENT WITH KERATIN CONTAMINATION OBSERVED IN ER (B)(4) V1. CONFIRM THE PRESENCE OF CONTAMINATION THROUGHOUT THE AIRPATH SUGGESTING AN EXTERNAL SOURCE. THE LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE 02 ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04660 | 2021-09-29 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION BIPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC. | DSX600H11C | DSX600H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04612 | 2021-09-29 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DNX500H11C | DNX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING FILTERS GETTING BLACK RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04611 | 2021-09-29 | 2021-08-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04488 | 2021-09-29 | 2021-02-17 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP CHONDROSARCOMA OF THE CRICOID. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED A VOLUNTARY MEDWATCH (MW5103409) ALLEGING A PATIENT DEVELOPED CHONDROSARCOMA OF THE CRICOID. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04487 | 2021-09-29 | 2021-08-06 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/HT/CELL,DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP A LUNG INFECTION. THE PATIENT REQUIRED SURGERY IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. IN INITIAL REPORTS SECTION B5 MENTIONED INCOMPLETE, CORRECT B5 SHOULD BE: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING LUNG INFECTION, COUGH, SHORT OF BREATH RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04513 | 2021-09-29 | 2021-09-13 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX500T12 | CAX500T12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGED TO HAVE SMALL BLACK DEBRIS WITHIN THEIR AFFECTED DEVICE, SEVERE HEADACHES AND THEY ARE MORE TIRED DURING THE DAY RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04511 | 2021-09-29 | 2021-08-10 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/HT/CELL, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX400T12C | CAX400T12C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED SINUS CONGESTION, SORE THROAT AND COUGH. THE PATIENT DID RECEIVE MEDICAL INTERVENTION AND STATES HE IS ON MEDICATION FOR THESE SYMPTOMS. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO EXPERIENCE A SINUS CONGESTION, COUGH, AND SORE THROAT . THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE REPORTED EVENTS OF SINUS CONGESTION, COUGH, AND SORE THROAT AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THIS EVENT IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. THE PATIENT HAS SWAPPED OUT THE DEICE AT THE DISTRIBUTOR AND AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION(S) B1, B2, HAS CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE, TYPE OF INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04512 | 2021-09-29 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID, CAN | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | CAX500H12 | CAX500H12 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED WHEEZING AND COUGHING. THE PATIENT DID RECEIVE MEDICAL INTERVENTION THAT RESULTED IN A PRESCRIPTION FOR AN INHALER. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. ADDITIONAL INFORMATION WAS RECEIVED AND SECTION B5 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CPAP DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED WHEEZING AND COUGHING. THE MEDICAL INTERVENTION THAT THE PATIENT RECEIVED IN RESPONSE TO THE EVENT IS CURRENTLY UNKNOWN. THE REPORTED EVENTS OF WHEEZING AND COUGHING AND ITS REPORTED SEVERITY WAS REVIEWED BY THE MANUFACTURE'S CLINICAL EXPERT. THESE EVENTS ARE NOT ASSESSABLE DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE IN THIS CASE. BASED ON THE AVAILABLE INFORMATION, THE MANUFACTURE CONCLUDES NO FURTHER ACTION IS NECESSARY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LAB FOR FURTHER EVALUATION. THE EXTERNAL AND INTERNAL ASPECTS OF THE DEVICE WERE INSPECTED, AND THE MANUFACTURER OBSERVED AN UNKNOWN DUST CONTAMINATION INSIDE THE BLOWER BOX AT THE AIR INLET OF THE DEVICE. AN UNKNOWN CONTAMINATION WAS OBSERVED ON THE ISO PORT OF THE REAR PANEL. DUST CONTAMINATION WAS FOUND ON THE TOP ENCLOSURE, FRONT PANEL, REAR PANEL, BOTTOM ENCLOSURE, BLOWER, BLOWER SEAL, AND BLOWER BOX. EVIDENCE OF LIQUID INGRESS WITH A CONTAMINANT CONSISTENT WITH MINERAL DEPOSITS WAS OBSERVED ON THE BLOWER AND BLOWER BOX. A KERATIN-LIKE SUBSTANCE WAS OBSERVED ON THE BLOWER BOX OUTLET. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE NO ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THERE WAS NO PRESENCE OF DEGRADED SOUND ABATEMENT FOAM IN THE DEVICE BUT EVIDENCE OF DUST/ DIRT AND LIQUID INGRESS WAS OBSERVED IN THE DEVICE. SECTIONS B1, B2 HAVE CHANGED RELATED TO THE COMPLAINT CHANGING FROM THE REPORTED ADVERSE EVENT TO A PRODUCT PROBLEM. SECTION H1 HAS CHANGED TO REFLECT A MALFUNCTION. SECTION H6 HEALTH EFFECT- IMPACT CODE (2199) HAS BEEN UPDATED. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04555 | 2021-09-29 | 2021-08-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11C | DSX400H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGING ISSUES. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WAS NO MENTION OF VISUAL FINDINGS TO THE EXTERNAL PART OF THE DEVICE. THE INTERNAL ASPECT OF THE DEVICE WAS INSPECTED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE FOUR ERRORS FOUND. THE MANUFACTURER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT TO BE SCRAPPED DUE TO MULTIPLE E53 ERROR CODES. SECTION H6 WERE UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04604 | 2021-09-29 | 2021-08-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO BIPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX700S11 | DSX700S11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO SOUND ABATEMENT FOAM. UPON FURTHER REVIEW, THIS DEVICE WAS A REPAIRED DEVICE AND DID NOT CONTAIN SOUND ABATEMENT FOAM THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY AND IS NOT IN SCOPE OF RES 88058. THEREFORE, THERE IS NO ALLEGATION OF A REPORTABLE EVENT ASSOCIATED WITH THE DEVICE AT THIS TIME. SECTION G1 CORRECTED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04609 | 2021-09-29 | 2021-07-08 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE PATIENT REPORTED USING AN OZONE BASED DISINFECTION DEVICE. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES (B)(4). |  | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04605 | 2021-09-29 | 2021-08-04 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PARTICLES IN TUBING/CHAMBER, BURNING/SMOKE/ELECTRICAL ODOR, DEVICE UNIT RUNNING HOT RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04508 | 2021-09-29 | 2021-06-15 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUMID/HT, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11 | DSX500T11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SHORTNESS OF BREATH AND A COUGH. THE PATIENT WAS PRESCRIBED ANTIBIOTICS AND AN INHALER IN RESPONSE TO THE REPORTED EVENT. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING A DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SHORTNESS OF BREATH AND A COUGH. THE PATIENT WAS PRESCRIBED ANTIBIOTICS AND AN INHALER IN RESPONSE TO THE REPORTED EVENT.THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL AND INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF CONTAMINATION ON FINE DUST CONTAMINATION OBSERVED ON AND IN THE BLOWER, WEAR AND TEAR INCLUDING SCUFFS, DARK MARKS, AND INDETERMINATE CONTAMINATION ON EXTERIOR OF DEVICE AND DARK CONTAMINANT INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM, OTHER UNIDENTIFIED CONTAMINATES WERE OBSERVED ON THE EXTERIOR OF THE MASK. THE MANUFACTURER WAS ALSO FOUND EVIDENCE OF WATER INGRESS TO THE BLOWER BOX ON THE BLOWER. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION.THE HUMIDIFIER WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR EVALUATION. AN INTERNAL VISUAL INSPECTION WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF CONTAMINATION CONSISTENT WITH KERATIN CONTAMINATION AND/OR BIO-CONTAMINATION WAS OBSERVED ON THE HUMIDIFIER FLIP LID SEAL.THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND 4 INSTANCES OF E-53 AND 1 INSTANCES OF E-130.THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF CONTAMINATION CONSISTENT WITH KERATIN CONTAMINATION AND/OR BIO-CONTAMINATION WAS OBSERVED ON THE HUMIDIFIER FLIP LID SEAL, CONTAMINATION ON FINE DUST CONTAMINATION OBSERVED ON AND IN THE BLOWER, WEAR AND TEAR INCLUDING SCUFFS, DARK MARKS, AND INDETERMINATE CONTAMINATION ON EXTERIOR OF DEVICE AND DARK CONTAMINANT INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM, OTHER UNIDENTIFIED CONTAMINATES WERE OBSERVED ON THE EXTERIOR OF THE MASK, EVIDENCE OF WATER INGRESS TO THE BLOWER BOX ON THE BLOWER. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN.IN THE PREVIOUSLY SUBMITTED REPORT SECTION G3 WAS INCORRECT, THE CORRECT G3 SECTION WAS 06/15/2021.SECTION B3, D9, D10, G3 AND H6 HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04637 | 2021-09-29 | 2021-08-10 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500T11C | DSX500T11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING MODERATE PULMONARY HYPERTENSION, NASAL/THROAT IRRITATION, SORENESS ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION B1, B2 AND H1 ARE CORRECTED. SECTION H6 UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
MW5104290 | 2021-09-29 | 2019-01-01 | I, n, j, u, r, y | Voluntary report | DREAMSTATION CPAP WITH HUMIDIFIER | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | PHILIPS / RESPIRONICS, INC. | DSX400H11C |  |  | FROM ABOUT (B)(6) 2019 TO PRESENT, I STARTED TO NOTICE A BLACK MOLD-LIKE FILM ON MY WATER RESERVOIR AND HOSE. THE BLACK SUBSTANCE IS DIFFICULT TO REMOVE. I HAVE EXPERIENCED DIFFICULTY BREATHING ON AND OFF UPON WAKING UP AND DIAGNOSED WITH 3RD STAGE KIDNEY DISEASE. ALSO, I WAKE UP WITH HEADACHES ALMOST EVERY DAY WITH EYE IRRITATION. I OWN TWO PHILIPS CPAP MACHINES. PRIOR TO 2019 MY KIDNEY FUNCTION WAS ABOVE NORMAL. GRADUALLY KIDNEY FUNCTION HAS DECLINED ALONG WITH OTHER MEDICAL ISSUES DISCUSSED ABOVE. FDA SAFETY REPORT ID # (B)(4). |  | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04642 | 2021-09-29 | 2021-09-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP, JP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | RJPX500S16 | RJPX500S16 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL AND INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER. THE MANUFACTURER FOUND EVIDENCE OF HOUSEHOLD DUST/DIRT CONTAMINATION. THE MANUFACTURER FOUND NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED. THE MANUFACTURER FOUND NO ERRORS LOGGED. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF HOUSEHOLD DUST/DIRT CONTAMINATION. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. SECTION D9, G3 AND H6 HAS BEEN UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04489 | 2021-09-29 | 2021-08-04 | I, n, j, u, r, y | Manufacturer report | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX500H11C | DSX500H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) DEVICE'S SOUND ABATEMENT FOAM BECAME DEGRADED AND CAUSED A PATIENT TO DEVELOP SKIN CANCER. THERE IS NO REPORT OF THE MEDICAL INTERVENTION THAT THE PATIENT HAS RECEIVED AT THIS TIME. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGATION ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED TO DEVELOP SKIN CANCER. PATIENT ALSO ALLEGED DIFFICULTY BREATHING/SHORT OF BREATH, BREATHING PROBLEMS/ CAN'T BREATHE, SKIN CANCER, CHOKING, KIDNEY MALFUNCTIONING DUE TO MACHINE, STABBING PAIN IN EYES, VOMITING, HEADACHES, ASTHMA, NASAL/THROAT IRRITATION OR SORENESS, WAKES UP TIRED, NAUSEA, UPPER AIRWAY IRRITATION, AND NOT ABLE TO EAT SOLID FOOD DUE TO NOT ABLE TO BREATHE. DOCTORS PUT HER ON A LIQUID DIET. PATIENT HAD TO VISIT THE ER ROOM FOR BREATHING ISSUES. PATIENT IS HAVING ISSUES WITH KIDNEYS AS WELL AS STATES SHE HAS SKIN IRRITATION. WAS HOSPITALIZED WITH HYPERTENSION FROM THE MACHINE. HER QUALITY OF LIFE HAS RAPIDLY DECREASED. THE ABOVE REPORTED EVENT IS ASSESSED AS POSSIBLY RELATED TO THE DEVICE IN THIS CASE. HENCE, THE DEVICE MAY HAVE CAUSE OR CONTRIBUTE TO THE ALLEGED SERIOUS INJURY. THE PATIENT DID NOT REPORT TO RECEIVE MEDICAL INTERVENTION REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, A FOLLOW-UP REPORT WILL BE FILED. SECTION H6 (HEALTH EFFECT - CLINICAL CODE) HAS BEEN UPDATED/CORRECTED AND (TYPE OF INVESTIGATION,INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS) HAS BEEN UPDATED IN THIS REPORT. | Y | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04610 | 2021-09-29 | 2021-07-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION AUTO CPAP | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS INC | DSX500H11 | DSX500H11 |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09
2518422-2021-04606 | 2021-09-29 | 2021-08-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | DREAMSTATION CPAP PRO | VENTILATOR, NON-CONTINUOUS (RESPIRATOR) | RESPIRONICS, INC. | DSX400H11C | DSX400H11C |  | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THIS ISSUE WAS REPORTED TO THE FDA PER 21 CFR 806. THE DEVICE WILL BE CORRECTED PER RES 88058. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY REPORTED AN ALLEGATION OF AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGED BLACK DEBRIS IN DEVICE AND UPPER RESPIRATORY ISSUES THAT LED TO HOSPITALISATION. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER VISUALLY INSPECTED THE EXTERNAL PART OF THE DEVICE AND FOUND DUST OR DIRT FOUND ON OUTER SURFACES, CRACKS, AND CREVICES AND LIGHT AMOUNT OF DUST OR DIRT FOUND ON THE OUTER SURFACE OF THE BLOWER. THE MANUFACTURER VISUALLY INSPECTED THE DEVICE INTERNALLY AND FOUND BLACK CONTAMINANT FOUND ON THE INSIDE OF THE BLOWER . THE HUMIDIFIER'S DRY BOX SEAL SHOWED SLIGHT YELLOW DISCOLORATION. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION BUT WAS ABLE TO CONFIRM THE PRESENCE OF CONTAMINANT ON THE BLOWER. A CORRECTION TO B5 WAS MADE AND SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK DEBRIS IN DEVICE AND UPPER RESPIRATORY ISSUES THAT LED TO HOSPITALISATION RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. SECTION G1-MANUFACTURING CONTACT OFFICE INFO WAS PREVIOUSLY INCORRECT AND HAS BEEN CORRECTLY UPDATED. SECTION H6 WAS CORRECTED AND UPDATED IN THIS REPORT. | N | Y | LAY USER/PATIENT |  | 2021-09