## All Reports
report_number | date_received | date_of_event | event_type | report_source_code | brand_name | generic_name | manufacturer_d_name | model_number | catalog_number | device_class | event_description | manufacturer_narrative | adverse_event_flag | product_problem_flag | device_operator | manufacturer_g1_name
1219977-2016-00029 | 2016-02-22 |  | Injury | Manufacturer report | C-QUR MESH | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION |  |  |  | A THOROUGH INVESTIGATION WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT CODE, LOT NUMBER, OR SAMPLE WAS PROVIDED. THIS REPORT SHALL NOT BE CONSIDERED AS AN ADMISSION BY ATRIUM MEDICAL THAT THE PRODUCT DESCRIBED IN THE LAWSUIT AND DESCRIBED HEREIN ARE OR WERE DEFECTIVE, OR THAT IT HAD ANY CASUAL RELATIONSHIP TO ANY INJURIES ALLEGEDLY SUSTAINED BY THE PLAINTIFF. |  | Y | N | PHYSICIAN
1219977-2016-00049 | 2016-03-28 |  | Injury | Manufacturer report | C-QUR MESH | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION |  |  |  | A COMPLETE INVESTIGATION WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT CODE, LOT NUMBER OR SAMPLE WAS PROVIDED. PER THE STUDY, ADHESION CHARACTERISTICS OF MESH PLACED INTRAPERITONEALLY AND ADHESIOLYSIS-RELATED COMPLICATIONS DURING LAPAROSCOPIC REPAIR ARE ASSOCIATED WITH UNIQUE PROPERTIES OF THE MESH AND/OR BARRIER. |  | Y | N | PHYSICIAN
1219977-2016-00060 | 2016-04-07 | 20160321 | Injury | Manufacturer report | C-QUR TACSHIELD MESH | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION | 31633 | 31633 |  | A REVIEW OF LOT HISTORY AND STERILIZATION RECORDS WAS PERFORMED. THE FINISHED GOOD AND SUBASSEMBLY LOTS PASSED ALL IN-PROCESS INSPECTIONS AND TESTING INCLUDING BUT NOT LIMITED TO FOURIER TRANSFORM INFRARED SPECTROSCOPY AND SEAL STRENGTH ON PRE AND POST STERILE POUCHES. ALL RAW MATERIALS MET ALL INCOMING TESTING AND INSPECTIONS. REVIEW SHOWS THAT THIS LOT MET ALL STERILIZATION AND MANUFACTURING ACCEPTANCE CRITERIA FOR LOT RELEASE. CLINICAL EVALUATION: CLINICAL EVALUATION: A RASH IS A NOTICEABLE CHANGE IN THE TEXTURE OR COLOR OF THE SKIN. THE SKIN MAY BECOME SCALY, BUMPY, ITCHY, OR OTHERWISE IRRITATED. THERE ARE NUMEROUS CAUSES FOR RASHES, INCLUDING: EXPOSURE TO AN IRRITANT OR ANOTHER FOREIGN SUBSTANCE. POSSIBLE IRRITANTS COULD BE LAUNDRY DETERGENT, PERSONAL HYGIENE PRODUCTS, CHEMICALS OR POISONOUS PLANTS. THE RESULTING RASH MAY BE ITCHY AND RED OR INFLAMED. A RASH MAY ALSO INDICATE INFECTION. MICROORGANISMS CAN INFECT A SURGICAL WOUND THROUGH VARIOUS FORMS OF CONTACT, SUCH AS FROM THE TOUCH OF A CONTAMINATED CAREGIVER OR SURGICAL INSTRUMENT, THROUGH MICROORGANISMS IN THE AIR OR IN YOUR BODY. ANY INFECTION WILL CAUSE REDNESS, DELAYED HEALING, FEVER, PAIN, TENDERNESS, WARMTH, OR SWELLING. OTHER RISKS FOR INFECTION INCLUDE A COMPROMISED PATIENT, ELDERLY AND/OR OVERWEIGHT PATIENTS, WEAKENED IMMUNE SYSTEMS AND DIABETES. MOST SURGICAL SITE INFECTIONS CAN BE TREATED SUCCESSFULLY WITH ANTIBIOTIC MEDICATIONS. IT STATES IN ADVERSE REACTIONS, "COMPLICATIONS THAT MAY OCCUR WITH THE USE OF ANY SURGICAL MESH INCLUDE, BUT ARE NOT LIMITED TO INFLAMMATION AND INFECTION. |  | Y | N | PHYSICIAN
3011175548-2016-00001 | 2016-04-13 |  | Injury | Manufacturer report | C-QUR MOSAIC | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORP | 31106 | 31106 |  | A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION INTO THIS EVENT. |  | Y | N | PHYSICIAN
1219977-2016-00092 | 2016-05-06 | 20160419 | Injury | Manufacturer report | C-QUR V-PATCH | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION | 31202 | 31202 |  | A REVIEW OF LOT HISTORY AND STERILIZATION RECORDS WAS PERFORMED. THE FINISHED GOOD AND SUBASSEMBLY LOTS PASSED ALL IN-PROCESS INSPECTIONS AND TESTING INCLUDING BUT NOT LIMITED TO FOURIER TRANSFORM INFRARED SPECTROSCOPY AND SEAL STRENGTH ON PRE AND POST STERILE POUCHES. ALL RAW MATERIALS MET ALL INCOMING TESTING AND INSPECTIONS. REVIEW SHOWS THAT THIS LOT MET ALL STERILIZATION AND MANUFACTURING ACCEPTANCE CRITERIA FOR LOT RELEASE. THE INSPECTION OF THE MESH REVEALED SEVERAL AREAS OF COATING DELAMINATION ALTHOUGH THE STITCHING AROUND THE PERIMETER WAS INTACT. THE RING IN THE CENTER OF THE DEVICE WAS INTACT, THE CORRECT SIZE AND LAID FLAT. CONSEQUENTLY, THE C-QUR V-PATCH MESH SAMPLE WAS STILL CONFORMING AND WOULD HAVE BEEN ACCEPTABLE FOR USE. WHEN THE PRODUCT WAS RETURNED THE RING WAS LYING FLAT. DUE TO LIMITED INFORMATION PROVIDED, THE CAUSE OF THE COMPLAINT CANNOT BE DETERMINED. CLINICAL EVALUATION: WHEN A PROCEDURE IS PLANNED AND PREPARATIONS ARE MADE, THE REQUIRED PRODUCT IS INTRODUCED TO THE PROCEDURE TABLE PRIOR TO THE START OF THE PROCEDURE. PRODUCT, EXPIRATION DATE, SIZE AND STERILITY/PACKAGE INTEGRITY ARE INSPECTED AND CONFIRMED WITH TWO CLINICIANS AT THAT TIME. IF THERE WAS A DEFECT IN THE PRODUCT IT WOULD BE NOTICED AT THAT TIME AND A REPLACEMENT WOULD BE LOCATED. THIS WOULD REPRESENT A DELAY FOR THE PATIENT WHO IS CONFINED TO AN OPERATING TABLE AND UNDER SEDATION OR ANESTHESIA, INCREASING THE RISK OF IMMOBILITY AND PROLONGED SEDATION TIME. THE COMPLAINT STATES THAT THE RING OF THE MESH WAS TWISTED AND WOULD NOT LAY FLAT. A PRODUCT IS USED OR DISCARDED AT THE DISCRETION OF THE PHYSICIAN. A DEFECTIVE PRODUCT IS NOT EXPECTED TO BE USED IN A PROCEDURE. THE INSTRUCTIONS FOR USE WARN, PLEASE INSPECT PACKAGING TO ENSURE IT IS INTACT AND NOT DAMAGED PRIOR TO USE. |  | Y | N | PHYSICIAN
1219977-2016-00120 | 2016-06-06 |  | Injury | Manufacturer report | C-QUR MESH | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION |  |  |  | A COMPLETE INVESTIGATION WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT CODE, LOT NUMBER OR SAMPLE WAS PROVIDED. CLINICAL EVALUATION: THERE IS A LIMITED AMOUNT OF INFORMATION AVAILABLE ABOUT THE DETAILS OF THIS ISSUE INCLUDING LOCATION, ORGANISM, HISTORY AND TREATMENT PLAN. GENERALLY SPEAKING ANY SURGERY THAT CAUSES A BREAK IN THE SKIN CAN LEAD TO AN INFECTION. MICROORGANISMS CAN INFECT A SURGICAL SITE THROUGH VARIOUS FORMS OF CONTACT, SUCH AS FROM THE TOUCH OF A CONTAMINATED CAREGIVER OR SURGICAL INSTRUMENT, THROUGH MICROORGANISMS IN THE AIR, OR THROUGH MICROORGANISMS THAT ARE ALREADY ON OR IN YOUR BODY. CONTRIBUTING FACTORS ARE OBESITY AND HYGIENE PRACTICES. INFECTIONS MAY CAUSE REDNESS, DELAYED HEALING, FEVER, PAIN, TENDERNESS, WARMTH, OR SWELLING. PATIENTS THAT SMOKE, DRINK EXCESSIVELY, HAVE OBESITY AND DIABETES HAVE INCREASED RISK FOR INFECTION AND HEALING PROBLEMS. MOST INFECTIONS CAN BE TREATED SUCCESSFULLY WITH ANTIBIOTIC MEDICATIONS IF THE ORGANISM IS CORRECTLY IDENTIFIED. OCCASIONALLY ADDITIONAL SURGERY OR PROCEDURES MAY BE REQUIRED TO TREAT THE SURGICAL INFECTION. |  | Y | N | PHYSICIAN
1219977-2016-00122 | 2016-06-08 | 20150619 | Injury | Manufacturer report | C-QUR TACSHIELD | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION | 31633 | 31633 |  | A THOROUGH REVIEW OF THE DEVICE HISTORY AND STERILIZATION RECORDS INDICATES THAT THIS LOT OF HERNIA MESH PASSED ALL QUALITY INSPECTIONS AND PERFORMANCE INSPECTIONS. BASED ON THE DETAILS OF THE COMPLAINT AND ACCEPTABLE LOT QUALIFICATION RESULTS ATRIUM MEDICAL CAN FIND NO FAULT WITH THE PRODUCT IN QUESTION. |  | Y | N | PHYSICIAN
1219977-2016-00123 | 2016-06-08 | 20130515 | Injury | Manufacturer report | C-QUR V-PATCH MESH | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION | 31202 | 31202 |  | A THOROUGH REVIEW OF THE DEVICE HISTORY AND STERILIZATION RECORDS INDICATES THAT THIS LOT OF HERNIA MESH PASSED ALL QUALITY INSPECTIONS AND PERFORMANCE INSPECTIONS. BASED ON THE DETAILS OF THE COMPLAINT AND ACCEPTABLE LOT QUALIFICATION RESULTS ATRIUM MEDICAL CAN FIND NO FAULT WITH THE PRODUCT IN QUESTION. |  | Y | N | PHYSICIAN
1219977-2016-00124 | 2016-06-08 | 20150619 | Injury | Manufacturer report | C-QUR TACSHIELD | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION | 31628 | 31628 |  | A THOROUGH REVIEW OF THE DEVICE HISTORY AND STERILIZATION RECORDS INDICATES THAT THIS LOT OF HERNIA MESH PASSED ALL QUALITY INSPECTIONS AND PERFORMANCE INSPECTIONS. BASED ON THE DETAILS OF THE COMPLAINT AND ACCEPTABLE LOT QUALIFICATION RESULTS ATRIUM MEDICAL CAN FIND NO FAULT WITH THE PRODUCT IN QUESTION. |  | Y | N | PHYSICIAN
1219977-2016-00126 | 2016-06-10 | 20120522 | Injury | Manufacturer report | C-QUR MESH | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION | 31537 | 31537 |  | ALLEGEDLY, PLAINTIFF ALSO EXPERIENCED INCARCERATION, ADHESIONS, HEMATOMA, INFECTION, INFLAMMATION, GRANULATION, AND FOREIGN BODY REACTION. |  | Y | N | PHYSICIAN
1219977-2016-00133 | 2016-06-15 |  | Injury | Manufacturer report | C-QUR MOSAIC MESH | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION | 31105 | 31105 |  | A FOLLOW UP REPORT WILL BE SUBMITTED ON COMPLETION OF THE INVESTIGATION INTO THIS EVENT. |  | Y | N | PHYSICIAN
1219977-2016-00131 | 2016-06-17 |  | Injury | Manufacturer report | ATRIUM C-QUR V-PATCH | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION | 31201 | 31201 |  | ASSOCIATED FILE: 1219977-2016-00132. ON COMPLETION OF THE INVESTIGATION A FOLLOW UP REPORT WILL BE SUBMITTED. |  | Y | N | PHYSICIAN
1219977-2016-00132 | 2016-06-17 |  | Injury | Manufacturer report | ATRIUM C-QUR MESH | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION | 31544 | 31544 |  | ASSOCIATED FILE: 1219977-2016-00131. ON COMPLETION OF THE INVESTIGATION A FOLLOW UP REPORT WILL BE SUBMITTED. |  | Y | N | PHYSICIAN
3011175548-2016-00002 | 2016-07-07 | 20160518 | Injury | Manufacturer report | C-QUR V-PATCH MESH | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORP | 31201 | 31201 |  | A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION INTO THIS EVENT. RELATED MDRS: 3011175548-2016-00003, 3011175548-2016-00004. |  | Y | N | PHYSICIAN
3011175548-2016-00003 | 2016-07-07 |  | Injury | Manufacturer report | C-QUR V-PATCH MESH | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORP | 31200 | 31200 |  | A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION INTO THIS EVENT. RELATED MDRS: 3011175548-2016-00002, 3011175548-2016-00004. |  | Y | N | PHYSICIAN
3011175548-2016-00004 | 2016-07-07 |  | Injury | Manufacturer report | C-QUR V-PATCH MESH | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORP | 31200 | 31200 |  | A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION INTO THIS EVENT. RELATED MDRS: 3011175548-2016-00002, 3011175548-2016-00003. |  | Y | N | PHYSICIAN
1219977-2016-00167 | 2016-07-25 |  | Injury | Manufacturer report | C-QUR V-PATCH | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION |  |  |  | ADDITIONAL INFORMATION RECEIVED STATED THAT THE DOCTOR DID NOT HAVE PATIENTS WITH ALLERGIC REACTION TO THE FISH OIL. THE DOCTOR HAD INSTEAD ONE PATIENT THAT HAD SEROMA. WE ARE UNABLE TO FULLY INVESTIGATE THIS EVENT AS THE USER FACILTIY HAS NOT PROVIDED PRODUCT CODE OR LOT NUMBER USED IN THIS EVENT. CLINICAL EVALUATION: A SEROMA IS A COLLECTION OF FLUID UNDER THE SKIN. USUALLY IT OCCURS AT THE SITE OF A SURGICAL INCISION AND MAY FORM ABOUT ONE TO TWO WEEKS AFTER SURGERY. IT MAY LOOK LIKE A SWOLLEN LUMP AND FEEL TENDER OR SORE AND THERE MAY BE DRAINAGE. A SMALL SEROMA IS NOT DANGEROUS. HOWEVER, A LARGE SEROMA THAT IS CAUSING PAIN MAY NEED TO BE DRAINED. OCCASIONALLY A PROCEDURE MAY NEED TO BE PERFORMED TO REMOVE THE SEROMA AND/OR EXPLANT THE GRAFT. C-QUR MESH HAS AN O3FA COATING THAT IS DERIVED FROM HIGHLY PURIFIED PHARMACEUTICAL GRADE FISH OIL CONSISTING OF A UNIQUE BLEND OF TRIGLYCERIDES AND OMEGA 3 FATTY ACIDS. HIGHLY REFINED OILS MAY BE DERIVED FROM A PRIORITY FOOD ALLERGEN SOURCE; HOWEVER, BECAUSE OF THE DEGREE TO WHICH THEY HAVE BEEN REFINED, THEY CONTAIN NO PROTEIN. BECAUSE OMEGA THREE FATTY ACIDS ARE HIGHLY REFINED THEY ARE NOT ALLERGENIC. THE INSTRUCTIONS FOR USE STATE THAT COMPLICATIONS THAT MAY OCCUR WITH THE USE OF ANY SURGICAL MESH INCLUDE, BUT ARE NOT LIMITED TO, INFLAMMATION, INFECTION, SEROMA, HEMATOMA, FISTULA FORMATION OR MECHANICAL DISRUPTION OF THE TISSUE AND/OR MESH MATERIAL, POSSIBLE ADHESIONS WHEN PLACED IN DIRECT CONTACT WITH THE VISCERA (INTESTINES) AND ORGANS. SEROMA POST IMPLANT OF ANY MESH PRODUCT MAY BE CAUSED BY IMPROPER FIXATION METHODS. CAREFUL ATTENTION TO PROPER FIXATION TECHNIQUES AND SIZING OF THE PRODUCT SHOULD BE TAKEN TO HELP PREVENT EXCESSIVE TENSION OR DISRUPTION BETWEEN THE GRAFT MATERIAL AND TISSUE. |  | Y | N | PHYSICIAN
1219977-2016-00146 | 2016-07-26 |  | Injury | Manufacturer report | C-QUR MESH | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION | 31533 | 31533 |  | ALLEGEDLY PLAINTIFF ALSO EXPERIENCED BOWEL RESECTION, ADHESIONS, PURULENT DRAINAGE, INFLAMMATION, FOREIGN BODY REACTION, AND PAIN. |  | Y | N | PHYSICIAN
1219977-2016-00147 | 2016-07-26 | 20160309 | Injury | Manufacturer report | C-QUR MESH | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION | 31540 | 31540 |  | ALLEGEDLY PLAINTIFF ALSO EXPERIENCED IRECURRENT HERNIA, INCORPORATION OF MESH, MASS AND ADHESIONS. |  | Y | N | PHYSICIAN
1219977-2016-00148 | 2016-07-26 | 20130921 | Injury | Manufacturer report | C-QUR V-PATCH | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION | 31201 | 31201 |  | A REVIEW OF ALL MANUFACTURING LOT HISTORY AND STERILIZATION RECORDS WAS CONDUCTED. ALL IN-PROCESS SPECIFICATIONS AND RELEASE CRITERIA WERE MET, INCLUDING SEAL STRENGTH TESTING ON BOTH THE PRE-AND-POST-STERILE PACKAGING. BALL BURST AND SUTURE RETENTION TESTING WAS ALSO CONDUCTED ON THE MESH AT INCOMING AND FOURIER TRANSFORM INFRARED SPECTROSCOPY WAS PERFORMED ON THE CURED COATED MESH PANELS. |  | Y | N | PHYSICIAN
1219977-2016-00149 | 2016-07-26 | 20150213 | Injury | Manufacturer report | C-QUR MESH | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION | 31533 | 31533 |  | ALLEGEDLY PLAINTIFF ALSO EXPERIENCED INFLAMMATION AND BLOOD LOSS. |  | Y | N | PHYSICIAN
1219977-2016-00144 | 2016-07-26 | 20140730 | Injury | Manufacturer report | C-QUR MESH | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION | 31533 | 31533 |  | A REVIEW OF ALL MANUFACTURING LOT HISTORY AND STERILIZATION RECORDS WAS CONDUCTED. ALL IN-PROCESS SPECIFICATIONS AND RELEASE CRITERIA WERE MET, INCLUDING SEAL STRENGTH TESTING ON BOTH THE PRE-AND-POST-STERILE PACKAGING. BALL BURST AND SUTURE RETENTION TESTING WAS ALSO CONDUCTED ON THE MESH AT INCOMING AND FOURIER TRANSFORM INFRARED SPECTROSCOPY WAS PERFORMED ON THE CURED COATED MESH PANELS. |  | Y | N | PHYSICIAN
1219977-2016-00151 | 2016-07-26 | 20141111 | Injury | Manufacturer report | C-QUR V-PATCH | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION | 31201 | 31201 |  | PLAINTIFF ALSO ALLEGEDLY EXPERIENCED ABDOMINOPLASTY. |  | Y | N | PHYSICIAN
1219977-2016-00150 | 2016-07-26 | 20100917 | Injury | Manufacturer report | C-QUR MESH | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION | 31542 | 31542 |  | PLAINTIFF ALSO ALLEGEDLY EXPERIENCED PAIN, ABDOMINAL WALL ABSCESS, INFLAMMATION, RECURRENT HERNIA AND ADHESIONS. |  | Y | N | PHYSICIAN
1219977-2016-00153 | 2016-07-26 | 20150313 | Injury | Manufacturer report | C-QUR TACSHIELD | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION | 31626 | 31626 |  | ALLEGEDLY, PLAINTIFF ALSO EXPERIENCED ABDOMINAL PAIN, ADHESIONS, UMBILICAL RECURRENT HERNIA, FAT NECROSIS AND SMALL BOWEL OBSTRUCTION. |  | Y | N | PHYSICIAN
1219977-2016-00152 | 2016-07-26 | 20150119 | Injury | Manufacturer report | C-QUR MESH | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION | 31528 | 31528 |  | PLAINTIFF ALSO ALLEGEDLY EXPERIENCED ABDOMINOPLASTY. |  | Y | N | PHYSICIAN
1219977-2016-00154 | 2016-07-26 |  | Injury | Manufacturer report | C-QUR CENTRIFX | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION | 31396 | 31396 |  | PLAINTIFF ALSO ALLEGEDLY EXPERIENCED RECURRENT RIGHT INGUINAL HERNIA. |  | Y | N | PHYSICIAN
1219977-2016-00156 | 2016-07-26 | 20151024 | Injury | Manufacturer report | C-QUR V-PATCH | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION | 31201 | 31201 |  | ALLEGEDLY, PLAINTIFF ALSO EXPERIENCED INCARCERATED UMBILICAL HERNIA, CHRONIC GRANULATION, ABDOMINAL WALL ABSCESS AND PURULENT FLUID. |  | Y | N | PHYSICIAN
1219977-2016-00155 | 2016-07-26 |  | Injury | Manufacturer report | C-QUR CENTRIFX | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION | 31396 | 31396 |  | ALLEGEDLY, PLAINTIFF EXPERIENCED RECURRENT HERNIA POTENTIALLY CAUSED BY SEVERE SPASMS AND COUGHING. |  | Y | N | PHYSICIAN
1219977-2016-00157 | 2016-07-26 | 20150917 | Injury | Manufacturer report | C-QUR MESH | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION | 31536 | 31536 |  | A REVIEW OF ALL MANUFACTURING LOT HISTORY AND STERILIZATION RECORDS WAS CONDUCTED. ALL IN-PROCESS SPECIFICATIONS AND RELEASE CRITERIA WERE MET, INCLUDING SEAL STRENGTH TESTING ON BOTH THE PRE-AND-POST-STERILE PACKAGING. BALL BURST AND SUTURE RETENTION TESTING WAS ALSO CONDUCTED ON THE MESH AT INCOMING AND FOURIER TRANSFORM INFRARED SPECTROSCOPY WAS PERFORMED ON THE CURED COATED MESH PANELS. |  | Y | N | PHYSICIAN
1219977-2016-00159 | 2016-07-26 | 20110806 | Injury | Manufacturer report | C-QUR MESH | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION | 31528 | 31528 |  | ALLEGEDLY PATIENT ALSO EXPERIENCED CHRONIC INFECTION, SEPSIS, ADHESION TO BOWEL, FISTULA AND CHRONIC PAIN. |  | Y | N | PHYSICIAN
1219977-2016-00158 | 2016-07-26 |  | Injury | Manufacturer report | C-QUR MESH | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION |  |  |  | WE ARE UNABLE TO FULLY INVESTIGATE THIS REPORT AS NO PRODUCT NUMBER, LOT NUMBER OR SAMPLE WAS PROVIDED. |  | Y | N | PHYSICIAN
1219977-2016-00160 | 2016-07-26 | 20150923 | Injury | Manufacturer report | C-QUR MESH | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION | 31542 | 31542 |  | ALLEGEDLY PATIENT ALSO EXPERIENCED CHRONIC INFECTION, SEPSIS, ADHESION TO BOWEL, FISTULA AND CHRONIC PAIN. |  | Y | N | PHYSICIAN
1219977-2016-00161 | 2016-07-26 |  | Injury | Manufacturer report | C-QUR MESH | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORP. | 31533 | 31533 |  | PLAINTIFF ALLEGEDLY ALSO EXPERIENCED SEVERE AND CHRONIC PAIN, SMALL BOWEL OBSTRUCTION, ADHESIONS AND MENTAL PAIN. |  | Y | N | PHYSICIAN
1219977-2016-00162 | 2016-07-26 |  | Injury | Manufacturer report | C-QUR MESH | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION | 31528 | 31528 |  | A REVIEW OF ALL MANUFACTURING LOT HISTORY AND STERILIZATION RECORDS WAS CONDUCTED. ALL IN-PROCESS SPECIFICATIONS AND RELEASE CRITERIA WERE MET, INCLUDING SEAL STRENGTH TESTING ON BOTH THE PRE-AND-POST-STERILE PACKAGING. BALL BURST AND SUTURE RETENTION TESTING WAS ALSO CONDUCTED ON THE MESH AT INCOMING AND FOURIER TRANSFORM INFRARED SPECTROSCOPY WAS PERFORMED ON THE CURED COATED MESH PANELS. |  | Y | N | PHYSICIAN
1219977-2016-00171 | 2016-08-08 | 20080228 | Injury | Manufacturer report | C-QUR EDGE | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION | 31233 | 31233 |  | A REVIEW OF ALL MANUFACTURING LOT HISTORY AND STERILIZATION RECORDS WAS CONDUCTED. ALL IN-PROCESS SPECIFICATIONS AND RELEASE CRITERIA WERE MET, INCLUDING SEAL STRENGTH TESTING ON BOTH THE PRE-AND-POST-STERILE PACKAGING. BALL BURST AND SUTURE RETENTION TESTING WAS ALSO CONDUCTED ON THE MESH AT INCOMING AND FOURIER TRANSFORM INFRARED SPECTROSCOPY WAS PERFORMED ON THE CURED COATED MESH PANELS. |  | Y | N | PHYSICIAN
1219977-2016-00172 | 2016-08-08 | 20150809 | Injury | Manufacturer report | C-QUR EDGE | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION | 31228 | 31228 |  | A THOROUGH REVIEW OF THE DEVICE HISTORY RECORDS INDICATES THAT THIS LOT OF HERNIA MESH PASSED ALL QUALITY INSPECTIONS AND PERFORMANCE INSPECTIONS. BASED ON THE DETAILS OF THE COMPLAINT AND ACCEPTABLE LOT QUALIFICATION RESULTS, ATRIUM MEDICAL CAN FIND NO FAULT WITH THE PRODUCT IN QUESTION. |  | Y | N | PHYSICIAN
1219977-2016-00184 | 2016-08-15 |  | Injury | Manufacturer report | C-QUR V-PATCH | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORPORATION |  |  |  | WE ARE UNABLE TO FULLY INVESTIGATE THIS REPORT AS NO PRODUCT NUMBER, LOT NUMBER OR SAMPLE WAS PROVIDED. THERE IS NO INDICATION THIS EVENT IS DUE TO A DEVICE FAILURE. CLINICAL EVALUATION: UMBILICAL HERNIA OCCURS WHEN A SECTION OF THE INTESTINE PUSHES THROUGH A WEAK SPOT IN THE ABDOMINAL MUSCLES, NEAR THE BELLY BUTTON. CONDITIONS THAT LEAD TO UMBILICAL HERNIA ARE INCREASED INTRA-ABDOMINAL PRESSURE AND WEAKENED FASCIA AT THE LEVEL OF THE UMBILICUS. ALSO, OBESITY, ASCITES, MULTIPLE PREGNANCIES, AND LARGE ABDOMINAL TUMORS ALL CONTRIBUTE TO THE DEVELOPMENT OF UMBILICAL HERNIAS. ANY SURGERY THAT CAUSES A BREAK IN THE SKIN CAN LEAD TO A POSTOPERATIVE INFECTION. DOCTORS CALL THESE INFECTIONS SURGICAL SITE INFECTIONS (SSIS) BECAUSE THEY OCCUR ON THE PART OF THE BODY WHERE THE SURGERY TOOK PLACE. MICROORGANISMS CAN INFECT A SURGICAL WOUND THROUGH VARIOUS FORMS OF CONTACT, SUCH AS FROM THE TOUCH OF A CONTAMINATED CAREGIVER OR SURGICAL INSTRUMENT, THROUGH MICROORGANISMS IN THE AIR, OR THROUGH MICROORGANISMS THAT ARE ALREADY ON OR IN YOUR BODY AND THEN SPREAD INTO THE WOUND. ANY SSI MAY CAUSE REDNESS, DELAYED HEALING, FEVER, PAIN, TENDERNESS, WARMTH, OR SWELLING. |  | Y | N | PHYSICIAN
3011175548-2016-00005 | 2016-08-18 | 20160808 | Malfunction | Manufacturer report | C-QUR V-PATCH | MESH, SURGICAL, POLYMERIC | ATRIUM MEDICAL CORP | 31200 | 31200 |  | UPON THE COMPLETION OF THE INVESTIGATION INTO THIS EVENT A FOLLOW UP REPORT WILL BE SUBMITTED. |  | N | Y | PHYSICIAN

## Monthly Summary
year_month | count
2016-02 | 1
2016-03 | 1
2016-04 | 2
2016-05 | 1
2016-06 | 8
2016-07 | 22
2016-08 | 4

## Event Types
Event Type | Count
Injury | 38
Malfunction | 1

## Brand Names
Brand Name | Count
C-QUR MESH | 16
C-QUR V-PATCH | 7
C-QUR V-PATCH MESH | 4
C-QUR TACSHIELD | 3
C-QUR CENTRIFX | 2
C-QUR EDGE | 2
C-QUR TACSHIELD MESH | 1
C-QUR MOSAIC | 1
C-QUR MOSAIC MESH | 1
ATRIUM C-QUR V-PATCH | 1
ATRIUM C-QUR MESH | 1

## Case Metadata
Case Name | MDL Number | Filing Date | Court | Status | Settlement | Data Window Start | Data Window End | Total Records | Search Queries Used
Atrium C-QUR Mesh | 2753 | 2016-08-23 | D. New Hampshire | Active | N/A | 2016-02-22 | 2016-08-23 | 39 | 3