## All Reports
report_number | date_received | date_of_event | event_type | report_source_code | brand_name | generic_name | manufacturer_d_name | model_number | catalog_number | device_class | event_description | manufacturer_narrative | adverse_event_flag | product_problem_flag | device_operator | manufacturer_g1_name
4597991 | 2015-02-20 | 20150212 | No answer provided | User Facility report | RECOVERY CONE | FILTER, INTRAVASCULAR, CARDIOVASCULAR | BARD PERIPHERAL VASCULAR, INC. | UNK | UNKNOWN |  |  |  | N | * | INVALID DATA
2020394-2015-00331 | 2015-04-14 | 20150212 | Injury | Manufacturer report | RECOVERY FILTER SYSTEM | DTK | BARD PERIPHERAL VASCULAR, INC. |  | RF048F |  | IT WAS REPORTED THAT APPROXIMATELY 11 YEARS POST VENA CAVA FILTER DEPLOYMENT; TWO DETACHED FILTER LIMBS WERE DISCOVERED. THE FILTER AND ONE DETACHED LIMB WERE SUCCESSFULLY CAPTURED AND RETRIEVED, THE SECOND DETACHED FILTER ARM REMAINS EMBEDDED IN THE RIGHT LUNG. THERE ARE NO PLANS TO REMOVE THE DETACHED LIMB IN THE LUNG. THERE WAS NO REPORTED PATIENT INJURY. |  | Y | Y | HEALTH PROFESSIONAL
2020394-2015-00616 | 2015-05-18 | 20050830 | Injury | Manufacturer report | RECOVERY FILTER SYSTEM | VENA CAVA FILTER | BARD PERIPHERAL VASCULAR, INC. |  | RF048F |  | IT WAS REPORTED THAT APPROXIMATELY NINE YEARS POST VENA CAVA FILTER DEPLOYMENT, THE FILTER WAS RETRIEVED SUCCESSFULLY; HOWEVER, A DETACHED FILTER LIMB REMAINS EMBEDDED. NO ATTEMPT WAS MADE TO CAPTURE AND RETRIEVE THE DETACHED FILTER LIMB. THE PATIENT STATUS AT THIS TIME IS UNKNOWN. |  | Y | Y | HEALTH PROFESSIONAL
2020394-2015-00972 | 2015-06-26 | 20130626 | Injury | Manufacturer report | RECOVERY FILTER SYSTEM | VENA CAVA FILTER | BARD PERIPHERAL VASCULAR, INC. |  | RF048F |  | THE LOT NUMBER WAS NOT PROVIDED; THEREFORE, THE DEVICE HISTORY RECORDS COULD NOT BE REVIEWED. THE INVESTIGATION IS CURRENTLY UNDERWAY. THE INFO PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFO AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFO, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. |  | Y | Y | HEALTH PROFESSIONAL
2020394-2015-01022 | 2015-07-09 | 20131213 | Injury | Manufacturer report | RECOVERY FILTER SYSTEM | VENA CAVA FILTER | BARD PERIPHERAL VASCULAR, INC. |  | RF048F |  | IT WAS REPORTED THAT APPROXIMATELY NINE YEARS POST IMPLANTATION OF A VENA CAVA FILTER; THE FILTER LIMB DETACHED AND MIGRATED WITHIN THE PATIENT'S BODY. REPORTEDLY THE FILTER WAS SURGICALLY REMOVED; HOWEVER, A DETACHED LIMB REMAINS IMPLANTED. PATIENT STATUS AT THIS TIME IS UNKNOWN. |  | Y | Y | HEALTH PROFESSIONAL
2020394-2015-01170 | 2015-07-17 | 20150617 | Injury | Manufacturer report | RECOVERY CONE REMOVAL SYSTEM |  | BARD PERIPHERAL VASCULAR, INC. |  | FBRC |  | NO DEVICE, HOSPITAL/MEDICAL RECORDS OR MEDICAL IMAGES HAVE BEEN MADE AVAILABLE TO THE MFR. AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS BEING PERFORMED. THE INVESTIGATION OF THE REPORTED EVENT IS UNDERWAY. |  | Y | Y | HEALTH PROFESSIONAL
2020394-2015-01247 | 2015-07-24 | 20140416 | Injury | Manufacturer report | RECOVERY FILTER SYSTEM | VENA CAVA FILTER | BARD PERIPHERAL VASCULAR, INC. |  | RF048F |  | IT WAS REPORTED THAT APPROXIMATELY 10 YEARS POST VENA CAVA FILTER DEPLOYMENT; ALLEGEDLY SEVERAL FILTER FRAGMENTS WERE DISCOVERED. ONE DETACHED FILTER FRAGMENT REPORTEDLY WAS LOCATED IN THE AORTA. THE VENA CAVA FILTER WAS REMOVED SUCCESSFULLY. ALLEGEDLY, SOME FILTER FRAGMENTS WERE NOT REMOVED. THE PATIENT STATUS AT THIS TIME IS UNKNOWN. |  | Y | Y | HEALTH PROFESSIONAL
2020394-2015-01258 | 2015-07-29 | 20150630 | Malfunction | Manufacturer report | RECOVERY CONE REMOVAL SYSTEM | RECOVERY CONE REMOVAL SYSTEM | BARD PERIPHERAL VASCULAR, INC. |  | FBRC |  | THE LOT NUMBER HAS BEEN PROVIDED AND THE DEVICE HISTORY RECORDS ARE BEING REVIEWED. THE INVESTIGATION IS CURRENTLY UNDERWAY. THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. |  | N | Y | HEALTH PROFESSIONAL
MW5045040 | 2015-07-29 | 20101214 | Injury | Voluntary report | GOPUMP RAPID RECOVERY SYSTEM | ELASTOMERIC PUMP | SYMBIOS |  |  |  | THE SYMBIOS GOPUMP RAPID RECOVERY SYSTEM WAS USED FOR MY AUXILIARY NODE DISSECTION SURGERY. THE SURGEON PLACE THE DOUBLE CHAMBER PUMP WITH FLOW RATE OF 6ML/HR DURING THE SURGERY. I AM A (B)(6) FEMALE, LEAN BODY MASS. I WAS SENT HOME THE SAME DAY, AND BY EVENING WAS FEELING EXTREMELY SHAKY AND WAS EXPERIENCING HEART PALPITATIONS AND RINGING IN MY EARS. DURING THE EARLY MORNING HOURS, MY HUSBAND AND I CALLED THE HELP NUMBER ON THE PUMP BROCHURE WE WERE GIVEN TO SEE WHAT TO DO. THE HOTLINE PERSON ADVISED ME TO CLAMP OFF THE PUMP HOSE TO STOP THE FLOW OF MARCAINE ANESTHESIA UNTIL I COULD SEE DOCTOR IN THE MORNING. FIRST THING NEXT MORNING, THE NURSE AT THE SURGEON OFFICE PULLED OUT THE TUBE DELIVERING THE ANESTHESIA. THE EXTREME SHAKINESS AND HEART PALPS GOT BETTER AS THE MARCAINE SLOWLY WORE OFF, HOWEVER, MY EARS HAVE NEVER RECOVERED. I BELIEVE IT WAS TOO MUCH DELIVERED INTO MY SYSTEM BY THE PUMP AND POOR PLACEMENT OF THE TUBE AS WELL. UPDATE TO PRESENT, I HAVE HAD ISSUES IN RANGE OF MOTION TO THAT ARM AND SHOULDER AND WORRY THAT THE MEDICATION DELIVERED BY THE PUMP DAMAGED OTHER BONE AND VESSELS IN THAT AREA. THIS IS THE FIRST I SAW OF THE PUMP RECALL, AND AM NOT SURPRISED. I KEPT MY BROCHURE ON IT BECAUSE I FELT IT CAUSED ME THESE PROBLEMS AND I AM THANKFUL I HAD THE PUMP CLAMPED OFF AS SOON AS I DID DUE TO THE HEART PALPS OR IN THINK IT COULD HAVE BEEN FATAL. |  | Y | N | LAY USER/PATIENT
3005075853-2015-00871 | 2015-02-05 | 20150102 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | (B)(4). ADDITIONAL INFORMATION: THE DEVICE WAS RECEIVED THE UPPER JAW DAMAGED AT THE PROXIMAL SIDE. THE DEVICE WAS CONNECTED TO THE GENERATOR AND IT WAS RECOGNIZED. BECAUSE THE JAW WAS DAMAGED NOT ALL FUNCTIONAL TESTING COULD BE PERFORMED WITH THE GENERATOR. THE CONDITION OF THE JAWS PREVENTED THE FUNCTIONALITY OF THE JAW. THE DEVICE WAS UNABLE TO CYCLE OPEN AND CLOSE. A PROBABLE CAUSE OF THE DAMAGE TO THE JAW COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-00873 | 2015-02-05 | 20150109 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | (B)(4). INFORMATION ASKED FOR BUT UNKNOWN OR NOT PROVIDED DURING INITIAL CONTACT. INFORMATION IS UNAVAILABLE; DEVICE NOT RETURNED. NO DEVICE RECEIVED FOR ANALYSIS AT TIME OF SUBMISSION OF 3500A. WHEN ADDITIONAL INFORMATION IS RECEIVED AND/OR THE DEVICE ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL MEDWATCH WILL BE SENT. |  | N | Y | HEALTH PROFESSIONAL
3004604967-2015-00020 | 2015-02-03 | 20141231 | Injury | Manufacturer report | STELLAR 100 - EUR G2 | VENTILATOR, CONTINUOUS, MIMIMAL VENT | RESMED RHT | 24154 | 24154 |  | (B)(4). |  | Y | N | LAY USER/PATIENT
3007573469-2015-00020 | 2015-02-03 | 20141231 | Injury | Distributor report | STELLAR 100 - EUR G2 | VENTILATOR, CONTINUOUS, MINIMAL VENT | RESMED RHT | 24154 | 24154 |  | IT WAS REPORTED TO RESMED FINLAND THAT THE PATIENT'S PRESSURE SETTINGS CHANGED DURING TREATMENT. IT WAS REPORTED THAT THE PATIENT'S OXYGEN SATURATION WAS REPORTED THAT THE PATIENT'S OXYGEN SATURATION WAS DECREASING PRIOR TO THE SETTINGS CHANGE AND IT CONTINUED TO DECREASE UNTIL IT WAS NOTICED BY HOSPITAL STAFF. PER THE REPORTER, THE PATIENT'S CARDIOLOGIST STATED THAT "IN MY OPINION, THIS PATIENT SAFETY INCIDENT DID NOT HAVE A SIGNIFICANT EFFECT ON THE PROGRESS OF THE DISEASE." THERE WAS NO PERMANENT INJURY TO THE PATIENT. MFR. #3004604967-2015-00020. |  | Y | N | LAY USER/PATIENT
3005075853-2015-00881 | 2015-02-05 | 20141203 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE INSTRUMENT DIDN¿T WORK PROPERLY. CUSTOMER CANNOT PROVIDE ANYMORE DETAILS ABOUT THE CIRCUMSTANCES IN WHICH THE ISSUE OCCURED. NO PATIENT CONSEQUENCE WAS REPORTED . |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-00913 | 2015-02-06 | 20141119 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | IT WAS REPORTED THAT DURING A LAPAROSCOPIC ASSISTED VAGINAL HYSTERECTOMY (LAVH) PROCEDURE, ONE DEVICE WAS NOT WORKING. THERE WERE NO PATIENT CONSEQUENCES. |  | N | Y | HEALTH PROFESSIONAL
2020394-2015-00038 | 2015-02-04 | 20150111 | Injury | Manufacturer report | G2 FILTER SYSTEM |  | BARD PERIPHERAL VASCULAR, INC. |  | UNKNOWN G2 |  | IT WAS REPORTED THAT UNSCHEDULED FILTER RETRIEVAL PROCEDURE WAS PERFORMED FOR COMPLAINT OF CHEST PAIN AND THREE DETACHED FILTER LIMBS. THE FILTER AND ONE DETACHED LIMB WERE RETRIEVED SUCCESSFULLY; HOWEVER, ONE FILTER LIMB REMAINS IN THE PULMONARY ARTERY AND THE SECOND LIMB REMAINS EMBEDDED IN THE IVC WALL. THE PATIENT WAS REPORTED TO BE ASYMPTOMATIC FOLLOWING THE PROCEDURE. |  | Y | Y | HEALTH PROFESSIONAL
3005075853-2015-01000 | 2015-02-11 | 20141122 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | (B)(4). SHOULD THE INFORMATION BE PROVIDED LATER, A SUPPLEMENTAL MEDWATCH WILL BE SENT. HE DEVICE WAS RECEIVED THE UPPER JAW DAMAGED AT THE PROXIMAL SIDE. THE DEVICE WAS CONNECTED TO THE GENERATOR AND IT WAS RECOGNIZED. BECAUSE THE JAW WAS DAMAGED NOT ALL FUNCTIONAL TESTING COULD BE PERFORMED WITH THE GENERATOR. THE CONDITION OF THE JAWS PREVENTED THE FUNCTIONALITY OF THE JAW. THE DEVICE WAS UNABLE TO CYCLE OPEN AND CLOSE. A PROBABLE CAUSE OF THE DAMAGE TO THE JAW COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01001 | 2015-02-11 | 20150119 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C25 |  | IT WAS REPORTED THAT DURING AN OPEN GASTRECTOMY, THE JAW BECAME NOT TO CLOSE AFTER SEVERAL ACTUATIONS. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00018 | 2015-02-06 | 20150101 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M5725INT0600010 |  | SURGICAL INTERVENTION OCCURRED TO ADDRESS ARTHROFIBROSIS. POLY INSERTS WERE EXCHANGED DURING THE PROCEDURE. |  | Y | N | HEALTH PROFESSIONAL
3005075853-2015-01034 | 2015-02-12 | 20150208 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | (B)(4). INFORMATION ASKED FOR BUT UNKNOWN OR NOT PROVIDED DURING INITIAL CONTACT. INFORMATION IS UNAVAILABLE; DEVICE NOT RETURNED. NO DEVICE RECEIVED FOR ANALYSIS AT TIME OF SUBMISSION OF 3500A. WHEN ADDITIONAL INFORMATION IS RECEIVED AND/OR THE DEVICE ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL MEDWATCH WILL BE SENT. |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00015 | 2015-02-06 | 20141201 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M5725INT0600010 |  | LOOSENING OF THE TOTAL KNEE REPLACEMENT WAS REPORTED. REVISION SURGERY OCCURRED. |  | Y | N | HEALTH PROFESSIONAL
2954740-2015-00048 | 2015-02-12 | 20130204 | Injury | Manufacturer report | ORBIT GALAXY G2 MICROCOIL DELIVERY SYSTEM | ARTIFICIAL EMBOLIZATION DEVICE | CODMAN AND SHURTLEFF, INC |  | 641CX2535 |  | CONCOMITANT MEDICATIONS: ASPIRIN 325 MG TABLET -TAKE 325 MG DAILY, ATORVASTATIN 20 MG DAILY, METOPROLOL XL 50 MG DAILY, OXYCODONE-ACETAMINOPHEN (PERCOCET) 5-325 MG EVERY 4 (FOUR) HOURS AS NEEDED, SENNOSIDES-DOCUSATE SODIUM 8.6-50 MG TABLET TWICE DAILY. DEVICES USED DURING INITIAL COIL EMBOLIZATION: ORBIT G2 COMPLEX FILL COILS (641CF0510/ C11298, 641CF0306/ C11716, 641CF2505/ C11605, 641CX2535/ C11536). COMPLAINT CONCLUSION: THE DEVICES WERE IMPLANTED AND NOT AVAILABLE FOR ANALYSIS. A REVIEW OF THE MANUFACTURING DOCUMENTATION ASSOCIATED WITH THIS LOT PRESENTED NO ISSUES DURING THE MANUFACTURING OR INSPECTION PROCESS THAT CAN BE RELATED TO THE REPORTED COMPLAINT. COIL COMPACTION IS A KNOWN LONG TERM EVENT ASSOCIATED WITH COIL EMBOLIZATION. FACTORS WHICH MAY HAVE A CORRELATION WITH RECANALIZATION POST COIL EMBOLIZATION INCLUDE NECK SIZE, PACKING DENSITY, AND INFLOW ANGLE. PROCEDURAL FACTORS AND VESSEL/ANEURYSM CHARACTERISTICS MAY HAVE CONTRIBUTED TO THE REPORTED ANEURYSM RE-CANALIZATION. WITH REVIEW OF THE INFORMATION, THERE IS NO INDICATION OF ANY DEVICE MANUFACTURING ISSUES RELATED TO THE EVENT. THEREFORE, NO CORRECTIVE ACTIONS WILL BE TAKEN. THIS IS 1 OF 4 MDRS SUBMITTED FOR THIS COMPLAINT WITH ASSOCIATED MANUFACTURE REPORT NUMBERS OF: 2954740-2015-00048, 2954740-2015-00049, 2954740-2015-00050, AND 2954740-2015-00051. THIS IS AN INITIAL/FINAL MDR. |  | N | Y | HEALTH PROFESSIONAL
2954740-2015-00050 | 2015-02-12 | 20130204 | Injury | Manufacturer report | ORBIT GALAXY G2 MICROCOIL DELIVERY SYSTEM | ARTIFICIAL EMBOLIZATION DEVICE | CODMAN AND SHURTLEFF, INC |  | 641CF0306 |  | CONCOMITANT MEDICATIONS: ASPIRIN 325 MG TABLET -TAKE 325 MG DAILY, ATORVASTATIN 20 MG DAILY, METOPROLOL XL 50 MG DAILY, OXYCODONE-ACETAMINOPHEN (PERCOCET) 5-325 MG EVERY 4 (FOUR) HOURS AS NEEDED, SENNOSIDES-DOCUSATE SODIUM 8.6-50 MG TABLET TWICE DAILY. SUBARACHNOID HEMORRHAGE AND ANEURYSM OF INTERNAL CAROTID ARTERY ((B)(6) 2012), HYPERLIPIDEMIA, HYPERTENSIVE DISORDER, OCCLUSIVE CORONARY ARTERY DISEASE, CEREBROVASCULAR DISEASE, MYOCARDIAL INFARCTION, CAROTID ARTERY STENOSIS ((B)(6) 2012), UNSTABLE ANGINA PECTORIS. DEVICES USED DURING INITIAL COIL EMBOLIZATION: ORBIT G2 COMPLEX FILL COILS (641CF0510/ C11298, 641CF0306/ C11716, 641CF2505/ C11605, 641CX2535/ C11536). COMPLAINT CONCLUSION: THE DEVICES WERE IMPLANTED AND NOT AVAILABLE FOR ANALYSIS. A REVIEW OF THE MANUFACTURING DOCUMENTATION ASSOCIATED WITH THIS LOT PRESENTED NO ISSUES DURING THE MANUFACTURING OR INSPECTION PROCESS THAT CAN BE RELATED TO THE REPORTED COMPLAINT. COIL COMPACTION IS A KNOWN LONG TERM EVENT ASSOCIATED WITH COIL EMBOLIZATION. FACTORS WHICH MAY HAVE A CORRELATION WITH RECANALIZATION POST COIL EMBOLIZATION INCLUDE NECK SIZE, PACKING DENSITY, AND INFLOW ANGLE. PROCEDURAL FACTORS AND VESSEL/ANEURYSM CHARACTERISTICS MAY HAVE CONTRIBUTED TO THE REPORTED ANEURYSM RE-CANALIZATION. WITH REVIEW OF THE INFORMATION, THERE IS NO INDICATION OF ANY DEVICE MANUFACTURING ISSUES RELATED TO THE EVENT. THEREFORE, NO CORRECTIVE ACTIONS WILL BE TAKEN. THIS IS 1 OF 4 MDRS SUBMITTED FOR THIS COMPLAINT WITH ASSOCIATED MANUFACTURE REPORT NUMBERS OF: 2954740-2015-00048, 2954740-2015-00049, 2954740-2015-00050, AND 2954740-2015-00051. THIS IS AN INITIAL/FINAL MDR. |  | N | Y | HEALTH PROFESSIONAL
2954740-2015-00051 | 2015-02-12 | 20130204 | Injury | Manufacturer report | ORBIT GALAXY G2 MICROCOIL DELIVERY SYSTEM | ARTIFICIAL EMBOLIZATION DEVICE | CODMAN AND SHURTLEFF, INC |  | 641CF0510 |  | A PATIENT EXPERIENCED RECURRENT LEFT SUPRACLINOID INTERNAL CAROTID ARTERY ANEURYSM DUE TO COIL COMPACTION APPROXIMATELY 5 MONTHS AFTER PLACEMENT OF FOUR ORBIT G2 COMPLEX FILL COILS (641CF0510/ C11298, 641CF0306/ C11716, 641CF2505/ C11605, 641CX2535/ C11536). INITIALLY THE PATIENT PRESENTED WITH A 6 X 5MM SUPRACLINOID LEFT INTERNAL CAROTID ARTERY ANEURYSM WITH A 3MM NECK AND SUBARACHNOID HEMORRHAGE. FOLLOWING COIL EMBOLIZATION, THERE WAS 95% OCCLUSION WITH MINIMAL RESIDUAL AT THE BASE. THERE WERE NO PRODUCT MALFUNCTIONS OR ADVERSE EVENTS AT THAT TIME. APPROXIMATELY 5 MONTHS AFTER COIL EMBOLIZATION, CEREBRAL ANGIOGRAM SHOWED AN INCREASE IN THE NECK REMNANT DUE TO COIL COMPACTION AND 10 MONTHS LATER, ANGIOGRAM REVEALED INCREASE IN NECK REMNANT. APPROXIMATELY TWO AND A HALF YEARS AFTER THE EMBOLIZATION, THE COIL COMPACTION WAS TREATED WITH STENT ASSISTED COIL EMBOLIZATION INVOLVING PLACEMENT OF AN ENTERPRISE STENT (ENF451412/10420180), WITH NO RESIDUAL AND NO PRODUCT MALFUNCTION. THE PROCEDURE WAS COMPLICATED BY RIGHT PHARYNGEAL WALL HEMATOMA, EDEMA (RIGHT TONSILLAR PILLAR) AND DYSPHAGIA SECONDARY TO INTUBATION. THE PATIENT WAS DISCHARGED 3 DAYS LATER WITHOUT TREATMENT. IT WAS LATER CONFIRMED THAT THE PATIENT WAS EATING AND DRINKING WITHOUT DIFFICULTY. THE PHYSICIAN CONFIRMED THAT ALL OF THE SYMPTOMS WERE RELATED TO INTUBATION. |  | N | Y | HEALTH PROFESSIONAL
2954740-2015-00049 | 2015-02-12 | 20130204 | Injury | Manufacturer report | ORBIT GALAXY G2 MICROCOIL DELIVERY SYSTEM | ARTIFICIAL EMBOLIZATION DEVICE | CODMAN AND SHURTLEFF, INC |  | 641CF2505 |  | A PATIENT EXPERIENCED RECURRENT LEFT SUPRACLINOID INTERNAL CAROTID ARTERY ANEURYSM DUE TO COIL COMPACTION APPROXIMATELY 5 MONTHS AFTER PLACEMENT OF FOUR ORBIT G2 COMPLEX FILL COILS (641CF0510/ C11298, 641CF0306/ C11716, 641CF2505/ C11605, 641CX2535/ C11536). INITIALLY THE PATIENT PRESENTED WITH A 6 X 5MM SUPRACLINOID LEFT INTERNAL CAROTID ARTERY ANEURYSM WITH A 3MM NECK AND SUBARACHNOID HEMORRHAGE. FOLLOWING COIL EMBOLIZATION, THERE WAS 95% OCCLUSION WITH MINIMAL RESIDUAL AT THE BASE. THERE WERE NO PRODUCT MALFUNCTIONS OR ADVERSE EVENTS AT THAT TIME. APPROXIMATELY 5 MONTHS AFTER COIL EMBOLIZATION, CEREBRAL ANGIOGRAM SHOWED AN INCREASE IN THE NECK REMNANT DUE TO COIL COMPACTION AND 10 MONTHS LATER, ANGIOGRAM REVEALED INCREASE IN NECK REMNANT. APPROXIMATELY TWO AND A HALF YEARS AFTER THE EMBOLIZATION, THE COIL COMPACTION WAS TREATED WITH STENT ASSISTED COIL EMBOLIZATION INVOLVING PLACEMENT OF AN ENTERPRISE STENT (ENF451412/10420180), WITH NO RESIDUAL AND NO PRODUCT MALFUNCTION. THE PROCEDURE WAS COMPLICATED BY RIGHT PHARYNGEAL WALL HEMATOMA, EDEMA (RIGHT TONSILLAR PILLAR) AND DYSPHAGIA SECONDARY TO INTUBATION. THE PATIENT WAS DISCHARGED 3 DAYS LATER WITHOUT TREATMENT. IT WAS LATER CONFIRMED THAT THE PATIENT WAS EATING AND DRINKING WITHOUT DIFFICULTY. THE PHYSICIAN CONFIRMED THAT ALL OF THE SYMPTOMS WERE RELATED TO INTUBATION. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01117 | 2015-02-17 | 20150212 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | (B)(4). ADDITIONAL INFORMATION: THE DEVICE WAS RECEIVED THE JAWS STUCK CLOSED AS THE UPPER JAW WAS DAMAGED AT THE PROXIMAL SIDE, THE HINGES WERE PROTRUDING. THE DEVICE WAS CONNECTED TO THE GENERATOR AND IT WAS RECOGNIZED. BECAUSE OF THE CONDITION OF THE DEVICE NOT ALL FUNCTIONAL TESTING COULD BE PERFORMED WITH THE GENERATOR. AS IT WAS OBSERVED THE INSTRUMENT WAS RECEIVED WITH TISSUE STUCK IN THE DAMAGED JAWS SO CAUTION SHOULD BE TAKEN BY ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01118 | 2015-02-17 | 20150210 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35A |  | (B)(4); DEVICE NOT RETURNED. NO DEVICE RECEIVED FOR ANALYSIS AT TIME OF SUBMISSION OF 3500A. WHEN ADDITIONAL INFORMATION IS RECEIVED AND/OR THE DEVICE ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL MEDWATCH WILL BE SENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01122 | 2015-02-17 |  | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S25 |  | IT WAS REPORTED THAT DURING A LAPAROSCOPIC SIGMA PROCEDURE, THE JAW BROKEN AFTER A FEW TIMES; PART FELL INTO SITUS BUT COULD BE REMOVED. NO FURTHER INFORMATION IS AVAILABLE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01124 | 2015-02-17 |  | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S25 |  | (B)(4). SHOULD THE INFORMATION BE PROVIDED LATER, A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE DEVICE WAS RECEIVED WITH THE TOP OF THE I-BLADE FRACTURED AND SLIGHTLY LIFTED. IN ADDITION THE UPPER JAW WAS DETACHED AND RETURNED. THE DEVICE WAS CONNECTED TO THE GENERATOR AND IT WAS RECOGNIZED. BECAUSE OF THE CONDITION OF THE DEVICE NOT ALL FUNCTIONAL TESTING COULD BE PERFORMED WITH THE GENERATOR. THE CONDITION OF THE BLADE PREVENTED THE FUNCTIONALITY OF THE JAW. THE JAW COULD NOT BE HELD IN PLACE DUE TO THE BREAKAGE. A PROBABLE CAUSE OF THE DAMAGE TO THE I BLADE COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01179 | 2015-02-19 | 20150210 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C45A |  | (B)(4). DEVICE NOT RETURNED. NO DEVICE RECEIVED FOR ANALYSIS AT TIME OF SUBMISSION OF 3500A WHEN ADDITIONAL INFORMATION IS RECEIVED AND/OR THE DEVICE ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION REQUESTED, BUT NOT RECEIVED: DID THE SURGEON ATTEMPT TO MANUALLY OPEN THE JAWS WITH HIS FINGERS? IF NO: ¿ WAS THE DEVICE ON A VESSEL/ARTERY WHEN IT WOULD NOT OPEN? ¿ IF YES, HOW WAS THE DEVICE REMOVED? (I.E. CUT OFF, FORCED OPENED) ¿ IF CUT OFF, DID THIS CAUSE A CHANGE IN THE PLAN OF CARE FOR THE PATIENT? ¿ DID THE REMOVAL CAUSE ANY DAMAGE TO THE VESSEL/ARTERY? ¿ IF YES, WHAT IS THE DAMAGE AND HOW WAS THE DAMAGE REPAIRED? ¿ DID THE SURGEON CONTINUE THE CASE WITH THIS SAME DEVICE? |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01187 | 2015-02-19 | 20150115 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C45 |  | (B)(4): INFORMATION ASKED FOR BUT UNKNOWN OR NOT PROVIDED DURING INITIAL CONTACT. SHOULD THE INFORMATION BE PROVIDED LATER, A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE DEVICE WAS RECEIVED IN GOOD VISUAL CONDITION. UPON VISUAL INSPECTION UNDER MAGNIFICATION IT WAS NOTICED THAT THE UPPER JAW WAS SLIGHTLY DAMAGED AT THE RETENTION PIN INTERPHASE. RESULTING IN THE JAW BEEN LOOSE AND DIFFICULTY IN OPENING AND CLOSING OF THE JAWS. HOWEVER THE DAMAGE WAS NOT SIGNIFICANT ENOUGH TO PREVENT THE DEVICE FROM CYCLING. THE DEVICES WAS TESTED ON THE GENERATOR AND PASSED ALL FUNCTIONAL TESTING. THE ENERGY OUTPUT DELIVERED FROM THE DEVICE WAS VERIFIED. ALL THREE TONES WERE HEARD DURING FUNCTIONAL TESTING (TONE 1 IS HEARD WHEN THE ENERGY ACTIVATION BUTTON IS PRESSED; TONE 2 IS HEARD WHEN TISSUE IMPEDANCE THRESHOLD IS REACHED; AND TONE 3 IS HEARD WHEN THE CYCLE IS COMPLETE). A PROBABLE CAUSE OF THE DAMAGE TO THE JAW COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01191 | 2015-02-19 | 20140128 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35 |  | IT WAS REPORTED THAT THEY WERE HAVING AN ERROR AND WOULD NOT RESET DURING A BOWEL SURGERY. CUSTOMER HAD TRIED FIVE DIFFERENT DEVICES. PROCEDURE WAS COMPLETED USING LIGASURE. NO PATIENT CONSEQUENCES. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01193 | 2015-02-19 | 20140128 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35 |  | (B)(4) - STUCK IN ACK 1 OF 3 STATUS. RE-SUBMITTING (B)(4) 2015 AS A SUPPLEMENTAL. INFORMATION ASKED FOR BUT UNKNOWN OR NOT PROVIDED DURING INITIAL CONTACT. SHOULD THE INFORMATION BE PROVIDED LATER, A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE DEVICE WAS RETURNED WITH THE UPPER JAW DETACHED AND NOT RETURNED. IN ADDITION, THE CABLE WAS CUT OFF. DUE TO THE RETURNED CONDITION OF THE DEVICE, NO FUNCTIONAL TESTING COULD BE PERFORMED WITH THE GENERATOR. BECAUSE THE JAW WAS DETACHED NOT ALL MECHANICAL TESTING COULD BE PERFORMED. A PROBABLE CAUSE OF THE DAMAGE TO THE JAW COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. IT SHOULD BE NOTED THAT AS PART OF OUR QUALITY PROCESS, EACH DEVICE IS VISUALLY INSPECTED AND FUNCTIONALLY TESTED DURING MANUFACTURING TO ENSURE THE DEVICE MEETS THE REQUIRED SPECIFICATIONS PRIOR TO SHIPMENT. |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00020 | 2015-02-06 | 20150101 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M5725060XE20 |  | SURGICAL INTERVENTION OCCURRED TO REMOVE SCAR TISSUE AND RESTORE RANGE OF MOTION. POLY INSERTS WERE EXCHANGED DURING THE PROCEDURE. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00016 | 2015-02-06 | 20150109 | Injury | Manufacturer report | IDUO G2 | BICOMPARTMENTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57230600240 |  | PATIENT WITH BICOMPARTMENTAL KNEE IMPLANTED EXPERIENCED LOSS OF EXTENSION AND FLEXION. ARTHROSCOPIC LYSIS OF ADHESION, LATERAL MENISCECTOMY, AND MANIPULATION UNDER ANESTHESIA OCCURRED. THE ISSUE WAS RESOLVED. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00019 | 2015-02-06 | 20150101 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57250600010 |  | REVISION SURGERY IS PLANNED TO EXCHANGE THE POLY INSERTS. REASON FOR REVISION IS UNKNOWN AT THIS TIME. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. |  | Y | N | HEALTH PROFESSIONAL
2020394-2015-00041 | 2015-02-09 | 20100112 | Malfunction | Manufacturer report | G2 FILTER SYSTEM - FEMORAL | VENA CAVA FILTER | BARD PERIPHERAL VASCULAR, INC. |  | RF310F |  | MANUFACTURING REVIEW: A MANUFACTURING REVIEW COULD NOT BE CONDUCTED AS THE LOT NUMBER WAS NOT PROVIDED. VISUAL/MICROSCOPIC INSPECTION: AS THE DEVICE WAS NOT RETURNED, AN INSPECTION COULD NOT BE PERFORMED. FUNCTIONAL/PERFORMANCE EVALUATION: AS THE DEVICE WAS NOT RETURNED, AN EVALUATION COULD NOT BE PERFORMED. MEDICAL RECORDS REVIEW: THE VENA CAVA FILTER WAS SUCCESSFULLY DEPLOYED IN THE INFRARENAL IVC PROPHYLACTICALLY PRIOR TO ABDOMINAL SURGERY. APPROXIMATELY FIVE YEARS POST FILTER DEPLOYMENT, CT SCANS DEMONSTRATED AN INCIDENTAL FINDING OF SEVERAL FILTER LIMBS EXTENDING BEYOND THE LUMEN OF THE IVC WALL: TWO WITHIN THE WALL OF THE ABDOMINAL AORTA; ONE WITHIN THE ANTERIOR PORTION OF THE L4 VERTEBRAL BODY. THERE WAS NO EVIDENCE OF AORTIC PSEUDOANEURYSM AND NO CHANGES IN THE POSITION OF THE FILTER OVER THE LAST TWO YEARS. THE PHYSICIAN FELT THAT NO INTERVENTION WAS NECESSARY AT THAT TIME AND WAS DOUBTFUL THAT THE PATIENT¿S SYMPTOMS OF INCREASING BACK AND CHEST PAIN WERE RELATED TO THE FILTER. APPROXIMATELY EIGHT YEARS POST FILTER DEPLOYMENT, A CT SCAN DEMONSTRATED PERFORATION OF FILTER LIMBS THROUGH THE IVC WALL UNCHANGED IN APPEARANCE FROM PREVIOUS EXAMS. NO FLUID COLLECTIONS OR ANY OTHER ABNORMALITY REGIONAL TO THE IVC WAS IDENTIFIED. THE PATIENT HAD NO ABDOMINAL PAIN OR SYMPTOMS RELATED TO THE FILTER AND THE PHYSICIAN DID NOT RECOMMEND REMOVAL OF THE FILTER AT THAT TIME. APPROXIMATELY NINE YEARS THREE MONTHS POST FILTER DEPLOYMENT, A CT SCAN DEMONSTRATED PERFORATION OF FILTER LIMBS THROUGH THE IVC WALL UNCHANGED IN APPEARANCE THAT DID NOT APPEAR TO CREATE ANY RISK OF INJURY. NO FOCAL ABNORMALITY WAS IDENTIFIED REGIONAL TO THE FILTER. THE PHYSICIAN FELT THAT FILTER LIMB PERFORATION OF THE IVC WALL WAS A COMMON APPEARANCE, OF DOUBTFUL CLINICAL SIGNIFICANCE AND NO INTERVENTION WAS NECESSARY AT THAT TIME. IMAGE/PHOTO REVIEW: NO MEDICAL IMAGES HAVE BEEN MADE AVAILABLE TO THE MANUFACTURER. CONCLUSION: THE DEVICE WAS NOT RETURNED. IMAGES WERE NOT PROVIDED. MEDICAL RECORDS WERE PROVIDED FOR REVIEW. A VENA CAVA FILTER WAS SUCCESSFULLY DEPLOYED. FIVE YEARS AND FOUR MONTHS POST FILTER DEPLOYMENT, A CT SCAN IDENTIFIED SEVERAL FILTER LIMBS EXTENDING BEYOND THE LUMEN OF THE IVC WALL. SUBSEQUENT CT SCANS CONTINUED TO IDENTIFIED THE PERFORATION OF FILTER LIMBS THROUGH THE IVC WALL, WHICH SHOWED NO CHANGE FROM INITIAL IDENTIFICATION OF PERFORATION. BASED ON THE MEDICAL RECORDS, FILTER LIMB PERFORATION OF THE IVC WALL CAN BE CONFIRMED. THE INVESTIGATION HOWEVER IS INCONCLUSIVE FOR FILTER MIGRATION BASED ON THE PROVIDED MEDICAL RECORDS. BASED UPON THE AVAILABLE INFORMATION THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. LABELING REVIEW: THE CURRENT IFU (INSTRUCTIONS FOR USE) STATES: WARNINGS/POTENTIAL COMPLICATIONS: MOVEMENT, MIGRATION OR TILT OF THE FILTER ARE KNOWN COMPLICATIONS OF VENA CAVA FILTERS. MIGRATION OF FILTERS TO THE HEART OR LUNGS HAS BEEN REPORTED. THERE HAVE ALSO BEEN REPORTS OF CAUDAL MIGRATION OF THE FILTER. MIGRATION MAY BE CAUSED BY PLACEMENT IN IVCS WITH DIAMETERS EXCEEDING THE APPROPRIATE LABELED DIMENSIONS SPECIFIED IN THIS IFU. MIGRATION MAY ALSO BE CAUSED BY IMPROPER DEPLOYMENT, DEPLOYMENT INTO CLOTS AND/OR DISLODGEMENT DUE TO LARGE CLOT BURDENS. PERFORATION OR OTHER ACUTE OR CHRONIC DAMAGE OF THE IVC WALL. THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. |  | Y | Y | HEALTH PROFESSIONAL
3004153240-2015-00023 | 2015-02-12 | 20150101 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M5725060XE20 |  | PATIENT HAD A PATELLA TENDON RUPTURE DURING PHYSICAL THERAPY. SURGERY OCCURRED TO REPAIR THE TENDON AND EXCHANGE THE POLY INSERTS. AT THE TIME OF SURGERY, IT WAS DETERMINED THAT THE PCL HAD AVULSED AND THE KNEE WAS NOT STABLE. REVISION TO A POSTERIOR STABILIZED IMPLANT IS PLANNED. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. |  | Y | N | HEALTH PROFESSIONAL
2020394-2015-00055 | 2015-02-13 | 20100509 | Injury | Manufacturer report | G2 FILTER SYSTEM |  | BARD PERIPHERAL VASCULAR, INC. |  | UNK G2 |  | A LOT HISTORY REVIEW COULD NOT BE PERFORMED AS THE LOT NUMBER WAS NOT PROVIDED. THE DEVICE WAS NOT RETURNED. IMAGES WERE NOT PROVIDED. THE COMPLAINT INVESTIGATION IS INCONCLUSIVE. UNABLE TO DETERMINE THE ROOT CAUSE BASED UPON THE AVAILABLE INFORMATION. |  | Y | Y | HEALTH PROFESSIONAL
3005075853-2015-01207 | 2015-02-19 | 20150128 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C14 |  | (B)(4): INFORMATION ASKED FOR BUT UNKNOWN OR NOT PROVIDED DURING INITIAL CONTACT. SHOULD THE INFORMATION BE PROVIDED LATER, A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE DEVICE WAS RECEIVED THE UPPER JAW DAMAGED AT THE PROXIMAL SIDE. IN ADDITION, THE PTC WAS DAMAGED. THE DAMAGE ON THE JAWS CAUSED A SHORT THAT DAMAGED THE PTC AND RESULTED IN AN ALERT SCREEN "REPOSITION JAWS AND REACTIVE", THIS IS ADVISING THAT THE INSTRUMENT IS BEING ACTIVATED ON LOW IMPEDANCE (THIN) TISSUE OR METAL (SUCH AS STAPLES, CLIPS, RETRACTORS, OR CLAMPS), THE "REPLACE INSTRUMENT" ALERT SCREEN WOULD BE DISPLAYED AFTER RECEIVING THE "REPOSITION AND REACTIVATE" ALERT SCREEN TWICE. IT IS POSSIBLE THAT THE REPORTED ¿INSTRUMENT ERROR¿ ALERT SCREEN MIGHT HAVE BEEN DISPLAYED BY THE GENERATOR DUE TO THE DESCRIBED CONDITION. A PROBABLE CAUSE OF THE DAMAGE TO THE JAWS COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01211 | 2015-02-19 | 20150106 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE TIP OF THE SCALPEL BECAME DEFORMED AND THE JAW COULD NOT BE CLOSED. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01235 | 2015-02-20 | 20141218 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35A |  | IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE WOULD NOT SEAL OR CUT. NO OTHER INFORMATION REGARDING THE INCIDENT WAS PROVIDED. |  | N | Y | HEALTH PROFESSIONAL
2954740-2015-00058 | 2015-02-20 | 20141209 | Malfunction | Manufacturer report | ORBIT GALAXY G2 MICROCOIL DELIVERY SYSTEM | ARTIFICIAL EMBOLIZATION DEVICE | CODMAN AND SHURTLEFF, INC |  | 641CF0515 |  | (B)(4). THE UNIDENTIFIED MICROCATHETER AND RHV WERE NOT RETURNED. CONCERNING CLEANLINESS ONLY, THE MICROCOIL SYSTEM WAS RETURNED IN ALMOST PRISTINE CONDITION. THE SYSTEM WAS EITHER NOT USED OR WAS CLEANED/RINSED BEFORE BEING RETURNED WHICH MAY HAVE PRODUCED FURTHER DAMAGE. IT IS ALSO UNKNOWN IF THE DEVICE WAS FURTHER MANIPULATED AND/OR INSPECTED POST-PROCEDURALLY. ANY TRACE EVIDENCE THAT MAY HAVE BEEN COMPLAINT RELATED MAY HAVE BEEN ALTERED OR REMOVED PRIOR TO BEING RETURNED. LOCATED 38.0 CM OFF THE DISTAL TIP OF THE GREEN INTRODUCER, THE CORE WIRE PROTRUDES OUTSIDE THE SHEATH FOR THE REMAINDER OF ITS LENGTH. NO MECHANICAL SHEATH DAMAGE TO THE SHEATH RESULTING IN AN OPENED SKIVE WAS FOUND. OUTSIDE VIEW OF THE SOCKET RING JUNCTION AND THE STRETCHED PROXIMAL SECTION OF THE COIL. THE REMAINDER OF THE COIL IS UNDAMAGED. LOCATED ON THE TOP PROXIMAL END OF THE RESHEATHING TOOL IN THE OPEN CUTOUT SECTION, THE V NOTCH HAS BEEN FRACTURED AND THE DAMAGED EDGES HAVE BEEN RAISED ABOVE THE SURFACE PLANE. THE LOCKING MECHANISM HAS COMPRESSION AND STRETCHING DAMAGE. NO MANUFACTURING DEFECTS WERE FOUND. A REVIEW OF THE MANUFACTURING DOCUMENTATION ASSOCIATED WITH THIS LOT PRESENTED NO ISSUES DURING THE MANUFACTURING OR INSPECTION PROCESS THAT CAN BE RELATED TO THE REPORTED COMPLAINT. THE PROXIMAL SECTION OF THE COIL HAS BEEN STRETCHED. LOCATED 38.0 CENTIMETERS OFF THE DISTAL TIP OF THE GREEN INTRODUCER, THE CORE WIRE PROTRUDES OUTSIDE THE SHEATH FOR THE REMAINDER OF ITS LENGTH. NO MECHANICAL SHEATH DAMAGE TO THE SHEATH RESULTING IN AN OPENED SKIVE WAS FOUND. THE EVIDENCE SUGGESTS THAT THE COIL WAS STRETCHED AND DAMAGED DURING THE PROTRUSION OUTSIDE THE SHEATH. THE REMAINDER OF THE COIL IS UNDAMAGED. LOCATED ON THE TOP PROXIMAL END OF THE RESHEATHING TOOL IN THE OPEN CUTOUT SECTION, THE V NOTCH HAS BEEN FRACTURED AND THE DAMAGED EDGES HAVE BEEN RAISED ABOVE THE SURFACE PLANE. THE LOCKING MECHANISM HAS COMPRESSION AND STRETCHING DAMAGE. NO MANUFACTURING DEFECTS WERE FOUND. THE MOST LIKELY INDICATION OF THE EVENT IS RELATED TO PROCEDURAL FACTORS. A REVIEW OF THE MANUFACTURING DOCUMENTATION ASSOCIATED WITH THIS LOT PRESENTED NO ISSUES DURING THE MANUFACTURING OR INSPECTION PROCESS THAT CAN BE RELATED TO THE REPORTED COMPLAINT. |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00025 | 2015-02-14 | 20140521 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57250600020 |  | PATIENT WAS REVISED TO A REVISION TOTAL KNEE SYSTEM. REASON FOR REVISION IS UNKNOWN. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. |  | Y | N | HEALTH PROFESSIONAL
2020394-2015-00084 | 2015-02-18 | 20140919 | Injury | Manufacturer report | G2 FILTER SYSTEM |  | BARD PERIPHERAL VASCULAR, INC. |  | UNKNOWN G2 |  | A COMPLETE MANUFACTURING REVIEW COULD NOT BE COMPLETED BECAUSE THE LOT NUMBER WAS NOT PROVIDED. THE DEVICE WAS NOT RETURNED FOR EVALUATION OR IMAGES WERE NOT PROVIDED FOR REVIEW. THE COMPLAINT INVESTIGATION IS INCONCLUSIVE. UNABLE TO DETERMINE THE ROOT CAUSE BASED UPON THE AVAILABLE INFORMATION. |  | Y | Y | HEALTH PROFESSIONAL
3004153240-2015-00027 | 2015-02-20 | 20150101 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M5725060XE10 |  | PT DEVELOPED SCAR TISSUE FOLLOWING TOTAL KNEE REPLACEMENT WHICH LIMITED RANGE OF MOTION. REVISION SURGERY OCCURRED TO RESECT 2MM ADD'L BONE FROM THE TIBIA AND IMPLANT A NEW TIBIAL TRAY AND POLY INSERTS. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED BY SPEC. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00026 | 2015-02-20 | 20150101 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M5725INT0600010 |  | LABELS WERE INCORRECTLY APPLIED TO THE POLY INSERT POUCHES. THE POUCH LABELED AS 8MM MEDIAL POLY INSERT CONTAINED THE LATERAL "A" POLY INSERT. THE POUCH LABELED AS LATERAL "A" POLY INSERT CONTAINED THE 8MM MEDIAL INSERT. THE CORRECT POLY INSERTS WERE PROVIDED. THE SURGERY WAS COMPLETED SUCCESSFULLY. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPEC. LABEL RETAINS INDICATE THAT THE CORRECT POLY INSERT LABELS WERE PRESENT. IT APPEARS THAT THE LABELS WERE APPLIED TO THE WRONG POUCHES WITHIN THE PRODUCT KIT. |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00022 | 2015-02-12 | 20150101 | Malfunction | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M75250600020 |  | STERILE BARRIER OF THE POUCH CONTAINING THE FEMORAL COMPONENT WAS COMPROMISED. THE IMPLANT WAS STERILIZED AT THE HOSP AND THE CASE WAS COMPLETED SUCCESSFULLY. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01340 | 2015-02-25 | 20150203 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35 |  | IT WAS REPORTED THAT DURING A TOTAL LAPAROSCOPIC HYSTERECTOMY (TLH) PROCEDURE, THE TOP JAW AND WHAT APPEARS TO BE PART OF THE KNIFE BLADE BROKE OFF OF THE DEVICE. THERE WERE NO PATIENT CONSEQUENCES. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01442 | 2015-03-02 | 20150205 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S14 |  | IT WAS REPORTED THAT DURING A JUGULAR PLASTIC SURGERY PROCEDURE, THE I-BLADE WAS BROKEN OFF AND THE UPPER JAW WAS CAME OFF OUTSIDE THE PATIENT. NO PIECES WERE LEFT INSIDE THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01453 | 2015-03-03 | 20150225 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C45 |  | (B)(4). INFORMATION ASKED FOR BUT UNKNOWN OR NOT PROVIDED DURING INITIAL CONTACT. INFORMATION IS UNAVAILABLE; DEVICE NOT RETURNED. NO DEVICE RECEIVED FOR ANALYSIS AT TIME OF SUBMISSION OF 3500A. WHEN ADDITIONAL INFORMATION IS RECEIVED AND/OR THE DEVICE ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL MEDWATCH WILL BE SENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01463 | 2015-03-03 | 20150123 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35 |  | IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, DURING THE USE OF OUR DEVICE, THE UPPER PART OF THE I-BLADE WENT FORTY-FIVE DEGREES UPWARD. AS A RESULT, IT COULD NOT BE USED PROPERLY. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01470 | 2015-03-03 | 20141203 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S14 |  | IT WAS REPORTED THAT PRIOR TO AN UNKNOWN PROCEDURE, ¿REACTIVATES AGAIN." IT IS UNKNOWN WHAT WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01507 | 2015-03-04 | 20150226 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C45 |  | IT WAS REPORTED THAT DURING A ROBOTIC LAPAROSCOPIC RADICAL PROSTATECTOMY PROCEDURE, ON INITIAL FIRING OF THE DEVICE, ERROR MESSAGE 'REPOSITION JAWS' AND 'REACTIVATE'. THE JAWS WOULD NOT RELEASE THE TISSUE. HANDLES PRIED OPEN TO RELEASE TISSUE. THEY UNPLUGGED THE HAND PIECE AND RE-ATTACHED. IMMEDIATE ERROR CODE 'REPLACE HAND PIECE'. CASE COMPLETED WITH ANOTHER DEVICE OF THE SAME PRODUCT CODE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. . |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01509 | 2015-03-04 | 20150212 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S25 |  | IT WAS REPORTED THAT DURING A WHIPPLE PROCEDURE, THE DEVICE WOULD NOT WORK WHEN ACTIVATED. THERE WERE NO PATIENT CONSEQUENCES. CASE COMPLETED WITH ANOTHER DEVICE OF THE SAME PRODUCT CODE. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01528 | 2015-03-05 | 20150106 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | PER USER FACILITY MEDWATCH REPORT FORM #3302450000-2015-8002, DURING AN UNKNOWN PROCEDURE; THE DOCTOR WAS INSERTING THE DEVICE INTO THE PATIENT AND THE JAWS OF THE DEVICE KEPT STICKING TOGETHER. IT IS NOT KNOWN HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01530 | 2015-03-05 | 20150205 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION PROCEDURE, THE SURGEON GRASPED TISSUE IN THE JAWS OF THE DEVICE. ERROR MESSAGES WERE REPORTED WHEN THE SURGEON ATTEMPTED TO ACTIVATE THE ENERGY. THE ERROR MESSAGE WAS NOT NOTED BY THE STAFF. WHEN ATTEMPTING TO GRASP THE TISSUE AGAIN THE SURGEON NOTICED THAT THE ANVIL WAS LOOSE AND APPEARS TO HAVE COME UNHINGED FROM THE JAW. THERE WERE NO PATIENT CONSEQUENCES. CASE COMPLETED WITH ANOTHER DEVICE OF THE SAME PRODUCT CODE. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01538 | 2015-03-05 | 20150110 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S45 |  | IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE JAW BROKE OFF THE DEVICE. THE JAW WAS RETRIEVED FROM THE PATIENT. A SECOND DEVICE WAS PULLED TO COMPLETE THE PROCEDURE WITH NO PATIENT CONSEQUENCE. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01595 | 2015-03-10 | 20150302 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | (B)(4). ADDITIONAL INFORMATION: THE NSLG2C35 PACKAGE WAS VISUALLY INSPECTED. THE TYVEK LID HAD A LARGE RIP THAT WAS CAUSED BY AN IMPACT FROM THE OUTSIDE. THE EVENT DESCRIPTION STATES THAT THE BOX WAS REPORTED AS BEING RIPPED BUT CONDITION OF THE CARTON AND SHIPPER BOX COULD NOT BE VERIFIED BECAUSE NEITHER WAS RETURNED. IT IS SUSPECTED THAT THE DAMAGE WAS CAUSED BY FORKLIFT TINES THRUST THROUGH THE CARTON AND INTO THE PACKAGE THAT WAS NEAREST THE OUTER EDGE OF THE CARTON. THE SIZE AND SHAPE OF THE DAMAGE WAS CONSISTENT WITH FORKLIFT TINES. ALL PACKAGES ARE 100% INSPECTED IMMEDIATELY AFTER SEALING BEFORE PLACING THEM INTO THE SALES UNIT CARTON. RANDOM QUALITY INSPECTIONS ARE PERFORMED. THE DAMAGE THAT OCCURRED IS INCONSISTENT WITH INTERNAL PACKAGING PROCESS AND HANDLING. THE DAMAGE IS NOT CONSISTENT WITH DROPPED PRODUCT, WHICH WOULD HAVE CRUSHING TYPE OF INJURY AND WOULD NOT LIKELY HAVE RIPPED TYVEK. FINISHED PRODUCT IS NEVER HANDLED USING A FORKLIFT AT THE PACKAGING FACILITY. COMPLAINT EVENT IS CONFIRMED AND IS ATTRIBUTABLE TO HANDLING EXTERNAL TO THE PACKAGING FACILITY. |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00032 | 2015-03-04 | 20140201 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57250600020 |  | POSSIBLE INFECTION WAS REPORTED FOR PATIENT WITH A TOTAL KNEE IMPLANT. REVISION SURGERY OCCURRED TO EXCHANGE THE POLY INSERTS. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00031 | 2015-03-04 | 20150201 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M5725060XE20 |  | TIBIAL IMPLANT LOOSENING WAS REPORTED. SURGERY IS PLANNED TO REVISE THE TIBIAL IMPLANTS. |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00028 | 2015-03-04 | 20141015 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57250600020 |  | PATIENT REPORTED POST-OPERATIVE PAIN AND STIFFNESS. MANIPULATION UNDER ANESTHESIA OCCURRED. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00030 | 2015-03-04 | 20150201 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57250600020 |  | POST-OPERATIVE LAXITY WAS REPORTED. REVISION SURGERY OCCURRED TO EXCHANGE THE POLY INSERTS WITH THICKER SIZES. CAUSE OF LAXITY IS UNKNOWN. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00033 | 2015-03-04 | 20150201 | Injury | Manufacturer report | IUNI G2 | UNICONDYLAR KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57220600220 |  | PT DEVELOPED AN INFECTION. REVISION SURGERY OCCURRED TO EXCHANGE THE POLY INSERT. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00029 | 2015-03-04 | 20150201 | Malfunction | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57250600030 |  | THE DISPOSABLE FEMORAL TRIAL BROKE DURING SURGERY. SURGERY WAS COMPLETED SUCCESSFULLY. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01610 | 2015-03-10 | 20150303 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | IT WAS REPORTED THAT DURING A LAPAROSCOPIC SIGMOID COLECTOMY PROCEDURE, THE JAWS WILL NOT OPEN; REMOVED FROM TISSUE BY MANUALLY OPENING. THERE WERE NO PATIENT CONSEQUENCES. CASE COMPLETED WITH ANOTHER DEVICE OF THE SAME PRODUCT CODE. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01614 | 2015-03-10 | 20150302 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S25 |  | (B)(4). BATCH #L90C8M. THE DEVICE WAS RECEIVED IN GOOD VISUAL CONDITION. UPON VISUAL INSPECTION UNDER MAGNIFICATION IT WAS NOTICED THAT THE UPPER JAW WAS SLIGHTLY DAMAGED. RESULTING IN THE JAW BEEN LOOSE AND DIFFICULTY IN OPENING AND CLOSING OF THE JAWS. HOWEVER THE DAMAGE WAS NOT SIGNIFICANT ENOUGH TO PREVENT THE DEVICE FROM CYCLING. THE DEVICES WAS TESTED ON THE GENERATOR AND PASSED ALL FUNCTIONAL TESTING. THE ENERGY OUTPUT DELIVERED FROM THE DEVICE WAS VERIFIED. ALL THREE TONES WERE HEARD DURING FUNCTIONAL TESTING (TONE 1 IS HEARD WHEN THE ENERGY ACTIVATION BUTTON IS PRESSED; TONE 2 IS HEARD WHEN TISSUE IMPEDANCE THRESHOLD IS REACHED; AND TONE 3 IS HEARD WHEN THE CYCLE IS COMPLETE). A PROBABLE CAUSE OF THE DAMAGE TO THE JAW COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. THE BATCH HISTORY RECORDS WERE REVIEWED WITH NO ANOMALIES NOTED DURING THE MANUFACTURING PROCESS. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01615 | 2015-03-10 | 20150204 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35 |  | IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE BLADE OF THE DEVICE JAMMED BECAUSE THE JAW WAS OUT OF ALIGNMENT AFTER A FULL BITE OF TISSUE. THERE WAS TISSUE ON THE JAW, AND THE JAW COULD NOT BE OPENED BECAUSE OF THE JAMMED BLADE. THE TISSUE SURROUNDING THE JAWS WAS REMOVED WITH AN EES SCISSOR. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01616 | 2015-03-10 | 20150303 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | (B)(4). BATCH #L92533. THE DEVICE WAS RETURNED WITH THE I-BLADE LOWER TIP ON THE WRONG SIDE OF THE JAW; THE LOWER JAW CHANNEL WAS DAMAGED. THE DEVICE WAS CONNECTED TO THE GENERATOR AND IT WAS RECOGNIZED. BECAUSE OF THE CONDITION OF THE DEVICE NOT ALL FUNCTIONAL TESTING COULD BE PERFORMED WITH THE GENERATOR. THE JAW WAS UNABLE TO CYCLE OPEN AND CLOSE. A PROBABLE CAUSE OF THE DAMAGE TO THE JAW COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. THE BATCH HISTORY RECORDS WERE REVIEWED WITH NO ANOMALIES NOTED DURING THE MANUFACTURING PROCESS. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01620 | 2015-03-10 | 20150204 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35 |  | (B)(4). ADDITIONAL INFORMATION: THE DEVICE WAS RECEIVED WITH THE UPPER JAW DETACHED NOT RETURNED AND THE CABLE CUT OFF. DUE TO THE RETURNED CONDITION OF THE DEVICE, NO FUNCTIONAL TESTING COULD BE PERFORMED WITH THE GENERATOR. THE CONDITION OF THE DEVICE PREVENTED THE FUNCTIONALITY OF THE JAW. THE JAW WAS UNABLE TO CYCLE OPEN AND CLOSE. A PROBABLE CAUSE OF THE DAMAGE TO THE JAW COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01621 | 2015-03-10 | 20150204 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35 |  | (B)(4). DEVICE NOT RETURNED. NO DEVICE RECEIVED FOR ANALYSIS AT TIME OF SUBMISSION OF 3500A. WHEN ADDITIONAL INFORMATION IS RECEIVED AND/OR THE DEVICE ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL MEDWATCH WILL BE SENT. |  | N | Y | HEALTH PROFESSIONAL
2020394-2015-00121 | 2015-03-05 | 20080623 | Malfunction | Manufacturer report | G2 FILTER SYSTEM - FEMORAL | VENA CAVA FILTER | BARD PERIPHERAL VASCULAR, INC. |  | RF310F |  | RECEIVED A (B)(4) ABOUT AN G2 FILTER, THE CATALOG NUMBER AND LOT NUMBER WERE NOT PROVIDED. A MANUFACTURING REVIEW COULD NOT BE PERFORMED AS THE LOT NUMBER WAS NOT PROVIDED. THE PATIENT'S WEIGHT WAS NOT PROVIDED BY THE REPORT SOURCE. THE DEVICE WAS NOT RETURNED. IMAGES WERE NOT PROVIDED. THE COMPLAINT INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED EVENT. PER THE REPORTED EVENT DETAILS, AN UNSUCCESSFUL ATTEMPT TO RETRIEVE THE FILTER WAS PERFORMED IN 2008. IT IS POSSIBLE THE LIMBS DETACHED DURING THE RETRIEVAL ATTEMPT. HOWEVER, THE DEFINITIVE ROOT CAUSE IS UNKNOWN. (B)(4). |  | Y | Y | HEALTH PROFESSIONAL
2020394-2015-00122 | 2015-03-05 | 20140626 | Injury | Manufacturer report | G2 FILTER SYSTEM - FEMORAL | VENA CAVA FILTER | BARD PERIPHERAL VASCULAR, INC. |  | RF310F |  | IT WAS REPORTED THAT APPROXIMATELY SEVEN YEARS POST FILTER DEPLOYMENT SOME DETACHED FILTER LIMBS WERE DISCOVERED. THE FILTER AND DETACHED LIMBS WERE REMOVED; ALTHOUGH, ONE DETACHED FILTER LIMB REMAINS IN THE PATIENT. NO OTHER INFORMATION HAS BEEN PROVIDED; THE PATIENT STATUS AT THIS TIME IS UNKNOWN. |  | Y | Y | HEALTH PROFESSIONAL
3004153240-2015-00035 | 2015-03-07 | 20150205 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600020 |  | SURGEON IDENTIFIED A LUCENT LINE UNDERNEATH THE TIBIAL IMPLANT. REVISION OF THE TIBIAL IMPLANTS IS PLANNED. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. |  | Y | N | HEALTH PROFESSIONAL
2020394-2015-00125 | 2015-03-05 | 20110405 | Injury | Manufacturer report | G2 FILTER SYSTEM - FEMORAL | VENA CAVA FILTER | BARD PERIPHERAL VASCULAR, INC. |  | RF310F |  | A MANUFACTURING REVIEW WAS NOT PERFORMED AS THE LOT NUMBER WAS NOT PROVIDED. THE PLAINTIFF'S ATTORNEY DID NOT PROVIDE PATIENT AGE OR WEIGHT. THE PRODUCT CATALOG NUMBER OR LOT NUMBER WAS NOT PROVIDED. A REQUEST LETTER WAS SENT TO THE ATTORNEY TO OBTAIN INFORMATION ABOUT THE EVENT DETAILS. NO RESPONSE FROM THE ATTORNEY WAS RECEIVED. THE DEVICE WAS NOT RETURNED. IMAGES WERE NOT PROVIDED. THE COMPLAINT INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED EVENT. BASED UPON THE AVAILABLE INFORMATION THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. |  | Y | Y | HEALTH PROFESSIONAL
3004153240-2015-00036 | 2015-03-07 | 20141201 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600020 |  | MANIPULATION UNDER ANESTHESIA OCCURRED FOR PATIENT WITH A TOTAL KNEE IMPLANT. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. |  | Y | N | HEALTH PROFESSIONAL
3005075853-2015-01692 | 2015-03-13 | 20150224 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C45A |  | IT WAS REPORTED THAT DURING A ROBOTIC PROSTATECTOMY PROCEDURE, THE JAWS WOULDN'T OPEN ON THE INSTRUMENT. THE CASE WAS COMPLETED USING CLIPS. NO PATIENT CONSEQUENCE WAS REPORTED. |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00034 | 2015-03-07 | 20150201 | Malfunction | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M5725INT0600010 |  | THE FEMORAL TRIAL AND FEMORAL IMPLANT DID NOT MATCH THE CUT BONE. THEY APPEARED TO BE LARGER THAN THE CUT BONE. CONVERSION TO AN ALTERNATE IMPLANT SYSTEM OCCURRED INTRAOPERATIVELY. THIS LED TO INCREASED SURGERY TIME AND PAIN FOR THE PATIENT. |  | N | Y | HEALTH PROFESSIONAL
MW5041385 | 2015-03-05 | 20150119 | Injury | Voluntary report | G2 | FILTER SYSTEM | BARD |  |  |  | FRACTURED, MIGRATED FILTER NOTED ON IMAGING. ONE LEG IN LEFT PULMONARY ARTERY BRANCH. ONE IN SOFT TISSUES OF RETROPERITONEUM. IN RETROSPECT WAS PRESENT LONG BEFORE, AT LEAST 2013. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01746 | 2015-03-16 | 20150202 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35A |  | IT WAS REPORTED THAT DURING A HYSTERECTOMY PROCEDURE, WHILE CUTTING THE VAGINAL VAULT THE MOVING JAW OF THE DEVICE WAS PUSHED AGAINST THE UTERINE MOBILIZER SO THE MOVING JAW RETRACTED AGAINST THE DEVICES SHAFT. THE SURGEON FINISHED THE HYSTERECTOMY WITH A MONOPOLAR DEVICE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01749 | 2015-03-16 | 20150309 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C45A |  | (B)(4). ADDITIONAL INFORMATION: THE DEVICE WAS RECEIVED IN GOOD PHYSICAL CONDITION. THE DEVICE WAS TESTED ON THE GENERATOR AND PASSED ALL FUNCTIONAL TESTING. THE ENERGY OUTPUT DELIVERED FROM THE DEVICE WAS VERIFIED. ALL THREE TONES WERE HEARD DURING FUNCTIONAL TESTING (TONE 1 IS HEARD WHEN THE ENERGY ACTIVATION BUTTON IS PRESSED; TONE 2 IS HEARD WHEN TISSUE IMPEDANCE THRESHOLD IS REACHED AND TONE 3 IS HEARD WHEN THE CYCLE IS COMPLETE). NO ISSUES WERE NOTED WITH OPENING AND CLOSING OF THE JAWS. NO ALERT SCREENS WERE DISPLAYED DURING TESTING. THERE WERE NO ANOMALIES NOTED WITH THE FUNCTIONALITY OF THE DEVICE. THIS REPORT IS NOT INTENDED TO DENY THAT YOU EXPERIENCED A PROBLEM WITH THE DEVICE. THERE MAY HAVE BEEN OTHER CIRCUMSTANCES OR ISSUES THAT OCCURRED DURING THE USE OF THE DEVICE THAT WE WERE UNABLE TO DUPLICATE DURING OUR LABORATORY ANALYSIS. |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00040 | 2015-03-12 |  | Injury | Manufacturer report | IDUO G2 | BICOMPARTMENTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | UNK |  | IT WAS REPORTED THAT REVISION SURGERY OCCURRED FOR PATIENT WITH A BICOMPARTMENTAL KNEE IMPLANT. REASON FOR REVISION IS UNKNOWN. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00037 | 2015-03-12 |  | Injury | Manufacturer report | IUNI G2 | UNICONDYLAR KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | UNK |  | IT WAS REPORTED THAT REVISION SURGERY OCCURRED FOR PATIENT WITH A UNICONDYLAR KNEE IMPLANT DUE TO PERSISTENT PAIN. |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00039 | 2015-03-12 |  | Injury | Manufacturer report | IDUO G2 | BICOMPARTMENTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | UNK |  | IT WAS REPORTED THAT REVISION SURGERY OCCURRED FOR PATIENT WITH A BICOMPARTMENTAL KNEE IMPLANT. REASON FOR REVISION IS UNKNOWN. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00038 | 2015-03-12 |  | Injury | Manufacturer report | IUNI G2 | UNICONDYLAR KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | UNK |  | IT WAS REPORTED THAT REVISION SURGERY OCCURRED FOR PATIENT WITH A UNICONDYLAR KNEE IMPLANT DUE TO PERSISTENT PAIN. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00042 | 2015-03-13 | 20150201 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57250600020 |  | SURGERY OCCURRED TO ADDRESS SCAR TISSUE IN THE KNEE. POLY INSERTS WERE AVAILABLE FOR POSSIBLE EXCHANGE DURING THE PROCEDURE. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00041 | 2015-03-13 | 20150101 | Injury | Manufacturer report | IUNI G2 | UNICONDYLAR KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57220600210 |  | FEMORAL IMPLANT LOOSENING WAS REPORTED. REVISION SURGERY OCCURRED. |  | Y | N | HEALTH PROFESSIONAL
3005075853-2015-01770 | 2015-03-17 | 20150225 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | (B)(4). INFORMATION ASKED FOR BUT UNKNOWN OR NOT PROVIDED DURING INITIAL CONTACT. INFORMATION IS UNAVAILABLE; DEVICE NOT RETURNED. NO DEVICE RECEIVED FOR ANALYSIS AT TIME OF SUBMISSION OF 3500A. WHEN ADDITIONAL INFORMATION IS RECEIVED AND/OR THE DEVICE ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL MEDWATCH WILL BE SENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01773 | 2015-03-17 |  | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S25 |  | IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE PRODUCT WAS DEFECTIVE; THE TOP JAW POPPED OFF UPON INITIAL USE. THE JAW FELL INTO THE PATIENT AND WAS RETRIEVED. IT WAS NOT KNOWN HOW THE CASE WAS COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED. |  | N | Y | HEALTH PROFESSIONAL
2020394-2015-00139 | 2015-03-13 | 20150123 | Injury | Manufacturer report | G2 EXPRESS FILTER |  | BARD PERIPHERAL VASCULAR, INC. |  | RF400F |  | IT WAS REPORTED THAT DURING A VENA CAVA FILTER RETRIEVAL PROCEDURE, THE APEX FILTER WAS REPORTEDLY EMBEDDED IN THE IVC WALL AND COULD NOT BE SUCCESSFULLY REMOVED. A SECONDARY RETRIEVAL FILTER ATTEMPT WAS MADE APPROXIMATELY ONE MONTH POST INITIAL FILTER RETRIEVAL ATTEMPT. A SNARE DEVICE WAS USED TO SUCCESSFULLY CAPTURE RETRIEVE THE FILTER WITHOUT INCIDENT. THERE WAS NO REPORTED PATIENT INJURY. |  | Y | Y | HEALTH PROFESSIONAL
2020394-2015-00138 | 2015-03-13 | 20150212 | Injury | Manufacturer report | G2 FILTER SYSTEM - FEMORAL | VENA CAVA FILTER | BARD PERIPHERAL VASCULAR, INC. |  | RF310F |  | MANUFACTURING REVIEW: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A MANUFACTURING REVIEW COULD NOT BE PERFORMED. VISUAL/MICROSCOPIC INSPECTION: AS THE DEVICE WAS NOT RETURNED, AN INSPECTION COULD NOT BE PERFORMED. FUNCTIONAL/PERFORMANCE EVALUATION: AS THE DEVICE WAS NOT RETURNED, AN EVALUATION COULD NOT BE PERFORMED. MEDICAL RECORDS REVIEW: THE PATIENT WITH HISTORY OF DEEP VENOUS THROMBOSIS AND PULMONARY EMBOLISM WAS SCHEDULED FOR FILTER PLACEMENT. THE RIGHT COMMON FEMORAL VEIN WAS ACCESSED AND A VENOGRAM DEMONSTRATED NO EVIDENCE OF THROMBUS AND IDENTIFIED THE RENAL VEINS. THE FILTER WAS SUCCESSFULLY DEPLOYED IN AN INFRARENAL LOCATION AND COMPLETION VENOGRAM DEMONSTRATED A CENTRALLY WELL PLACED FILTER. THE PATIENT TOLERATED THE PROCEDURE WELL. APPROXIMATELY SIX YEARS NINE MONTHS POST FILTER DEPLOYMENT, THE PATIENT WAS DIAGNOSED WITH NEW-ONSET LEFT VENTRICULAR DYSFUNCTION AND WAS SCHEDULED FOR LEFT HEART CATHETERIZATION. ANGIOGRAPHY DEMONSTRATED A DETACHED FILTER LIMB WITHIN THE RIGHT VENTRICLE. THE PATIENT WAS RECOMMENDED TO CONSULT INTERVENTIONAL RADIOLOGY FOR EVALUATION OF THE DETACHED FILTER LIMB. THE PATIENT WAS SCHEDULED FOR FILTER RETRIEVAL APPROXIMATELY SIX YEARS ELEVEN MONTHS POST FILTER DEPLOYMENT. THE RIGHT INTERNAL JUGULAR VEIN WAS ACCESSED AND A CONE REMOVAL SYSTEM WAS ADVANCED TO SUCCESSFULLY CAPTURE AND REMOVE THE FILTER. A SNARE DEVICE WAS ADVANCED INTO THE RIGHT VENTRICLE AND POSITIONED OVER THE FILTER FRAGMENT. THE DETACHED FILTER LIMB WAS UNABLE TO BE CAPTURED AND REMOVED. DUE TO THE PATIENT¿S SIZE, IT WAS NOT POSSIBLE TO KNOW THE ANTERIOR-POSTERIOR LOCATION OF THE SNARE. THE DECISION WAS MADE TO CONCLUDE THE PROCEDURE RATHER THAN MAKE FURTHER ATTEMPTS AT RETRIEVAL. THERE WERE NO KNOWN COMPLICATIONS. IMAGE/PHOTO REVIEW: AS MEDICAL IMAGES WERE NOT PROVIDED, A REVIEW COULD NOT BE PERFORMED. CONCLUSION: THE DEVICE WAS NOT RETURNED. IMAGES WERE NOT PROVIDED. MEDICAL RECORDS WERE PROVIDED AND REVIEWED. APPROXIMATELY SIX YEARS AND NINE MONTHS POST FILTER DEPLOYMENT, ANGIOGRAPHY DEMONSTRATED A DETACHED FILTER LIMB WITHIN THE RIGHT VENTRICLE. TWO MONTHS LATER THE FILTER WAS SUCCESSFULLY REMOVED, HOWEVER THE DETACHED FILTER COULD NOT BE RETRIEVED. BASED ON THE MEDICAL RECORDS, THE INVESTIGATION CAN BE CONFIRMED FOR ONE DETACHED FILTER LIMB. PER THE REPORTED EVENT DETAILS, THE PATIENT WAS IDENTIFIED AS MORBIDLY OBESE. PER THE IFU, "THE SAFETY AND EFFECTIVENESS OF THIS DEVICE HAS NOT BEEN ESTABLISHED FOR MORBIDLY OBESE PATIENTS. THERE HAVE BEEN REPORTS OF COMPLICATIONS, INCLUDING DEATH, ASSOCIATED WITH THE USE OF VENA CAVA FILTERS IN MORBIDLY OBESE PATIENTS. THE RISK/BENEFIT RATIO OF ANY OF THESE COMPLICATIONS SHOULD BE WEIGHED AGAINST THE INHERENT RISK/BENEFIT RATIO FOR A PATIENT WHO IS AT RISK OF PULMONARY EMBOLISM WITHOUT INTERVENTION." THEREFORE, IT IS UNKNOWN IF PATIENT ISSUES CONTRIBUTED TO THE REPORTED EVENT. THE DEFINITIVE ROOT CAUSE IS UNKNOWN. LABELING REVIEW: THE CURRENT IFU (INSTRUCTIONS FOR USE) STATES: WARNINGS/POTENTIAL COMPLICATIONS: FILTER FRACTURES ARE A KNOWN COMPLICATION OF VENA CAVA FILTERS. THERE HAVE BEEN SOME REPORTS OF SERIOUS PULMONARY AND CARDIAC COMPLICATIONS WITH VENA CAVA FILTERS REQUIRING THE RETRIEVAL OF THE FRAGMENT UTILIZING ENDOVASCULAR AND/OR SURGICAL TECHNIQUES. ALL OF THE ABOVE COMPLICATIONS HAVE BEEN ASSOCIATED WITH SERIOUS ADVERSE EVENTS SUCH AS MEDICAL INTERVENTION AND/OR DEATH. THERE HAVE BEEN REPORTS OF COMPLICATIONS, INCLUDING DEATH, ASSOCIATED WITH THE USE OF VENA CAVA FILTERS IN MORBIDLY OBESE PATIENTS. THE RISK/BENEFIT RATIO OF ANY OF THESE COMPLICATIONS SHOULD BE WEIGHED AGAINST THE INHERENT RISK/BENEFIT RATIO FOR A PATIENT WHO IS AT RISK OF PULMONARY EMBOLISM WITHOUT INTERVENTION. PRECAUTIONS: THE SAFETY AND EFFECTIVENESS OF THIS DEVICE HAS NOT BEEN ESTABLISHED FOR MORBIDLY OBESE PATIENTS. (B)(4). THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. |  | Y | Y | HEALTH PROFESSIONAL
3005075853-2015-01783 | 2015-03-18 | 20150127 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C25 |  | IT WAS REPORTED THAT DURING A HYSTERECTOMY, THE I-BLADE WAS DAMAGED IN ABOUT 3 HOURS. NO PIECES FELL INTO THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01788 | 2015-03-18 | 20150304 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | IT WAS REPORTED THAT DURING A LAPAROSCOPIC COLON PROCEDURE, THE SURGEON WAS USING THE DEVICE ON MESENTERY AND TOP PART OF THE JAW POPPED OFF. IT IS UNKNOWN HOW THE CASE WAS COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01837 | 2015-03-20 | 20141208 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C45 |  | IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE GENERATOR KEPT DISPLAYING AN ERROR MESSAGE 'REPOSITION JAW.' THE CAUSE IS NOT IDENTIFIED. IT IS UNKNOWN WHAT WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01891 | 2015-03-24 | 20140926 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE UPPER JAW COMPARTMENT WAS BROKEN WHILE ATTEMPTING TO GRASP TISSUE. THE BROKEN JAW COMPARTMENT WAS RETRIEVED. IT IS UNKNOWN WHAT WAS USED TO COMPLETE THE PROCEDURE. |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00047 | 2015-03-18 | 20141208 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600020 |  | PT WAS REVISED TO A REVISION TOTAL KNEE SYSTEM. REASON FOR REVISION IS UNK. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00045 | 2015-03-18 | 20140801 | Injury | Manufacturer report | IDUO G2 | BIOCOMPARTMENTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M5723INT0600220 |  | PT WAS HOSPITALIZED DUE TO PATELLA PAIN. SURGERY OCCURRED TO RESURFACE THE PATELLA. THE PATELLA WAS NOT RESURFACED DURING THE PRIMARY KNEE REPLACEMENT SURGERY. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00048 | 2015-03-18 | 20150201 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600010 |  | REVISION SURGERY OCCURRED FOR PT WITH A TOTAL KNEE IMPLANT. REVISION WAS REQUIRED DUE TO INJURY SUSTAINED IN AN AUTOMOBILE ACCIDENT. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00046 | 2015-03-18 |  | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600010 |  | REVISION SURGERY OCCURRED FOR PT WITH A TOTAL KNEE IMPLANT. POLYETHYLENE SYNOVITIS, OSTEOLYSIS, AND TIBIAL IMPLANT LOOSENING WERE OBSERVED AT THE TIME OF REVISION SURGERY. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00044 | 2015-03-18 | 20140101 | Injury | Manufacturer report | IDUO G2 | BICOMPARTMENTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M5723INT0600220 |  | PT WAS HOSPITALIZED DUE TO PATELLA PAIN. SURGERY OCCURRED TO RESURFACE THE PATELLA. THE PATELLA WAS NOT RESURFACED DURING THE PRIMARY KNEE REPLACEMENT SURGERY. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00043 | 2015-03-18 | 20141201 | Injury | Manufacturer report | IDUO G2 | BICOMPARTMENTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M5723INT0600210 |  | PT WAS HOSPITALIZED DUE TO PATELLA PAIN. SURGERY OCCURRED TO RESURFACE THE PATELLA. THE PATELLA WAS NOT RESURFACED DURING THE PRIMARY KNEE REPLACEMENT SURGERY. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. |  | Y | N | HEALTH PROFESSIONAL
3005075853-2015-01917 | 2015-03-25 | 20150315 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | (B)(4). DEVICE NOT RETURNED. NO DEVICE RECEIVED FOR ANALYSIS AT TIME OF SUBMISSION OF 3500A. WHEN ADDITIONAL INFORMATION IS RECEIVED AND/OR THE DEVICE ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL MEDWATCH WILL BE SENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01928 | 2015-03-25 | 20150218 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C45 |  | (B)(4). SHOULD THE INFORMATION BE PROVIDED LATER, A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE INSTRUMENT WAS RECEIVED WITH THE LOWER JAW SLIGHTLY BENT. IT IS POSSIBLE THAT DUE TO THE EXTRA FRICTION BETWEEN THE I-BLADE LOWER TIP AND THE JAW, THE USER COULD HAVE HEARD ABNORMAL NOISES. THIS WAS NOT CONFIRMED DURING THE ANALYSIS. THE DEVICE WAS CONNECTED TO THE GENERATOR AND IT WAS RECOGNIZED. BECAUSE THE JAW WAS DAMAGED NOT ALL FUNCTIONAL TESTING COULD BE PERFORMED WITH THE GENERATOR. THE CONDITION OF THE JAWS PREVENTED THE FUNCTIONALITY OF THE JAW. THE DEVICE WAS UNABLE TO CYCLE OPEN AND CLOSE. A PROBABLE CAUSE OF THE DAMAGE TO THE JAW COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01921 | 2015-03-25 | 20150313 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35A |  | (B)(4). ADDITIONAL INFORMATION : THE DEVICE RETURNED HAD A DIFFERENT MANUFACTURING BATCH NUMBER THAN REPORTED. THE DEVICE WAS RETURNED WITH THE I-BLADE IN GOOD CONDITION AND THE PTC WAS GOOD CONDITION AND NOT DETACHED AS REPORTED. THE DEVICE¿S JAWS WERE MISALIGNED AND THE CABLE CUT OFF. UPON INSPECTION UNDER MAGNIFICATION OF THE JAW IT WAS NOTED THAT ONE OF THE RETENTION PINS THAT HOLDS THE JAWS IN POSITION WAS SHEARED OFF, CAUSING THE MISALIGNMENT. IT IS POSSIBLE THAT DUE TO THE PIN DAMAGE THE JAW MISALIGNED WOULD HAVE RESULTED IN A SHORT WITH THE ELECTRODE AND A ¿REPOSITION JAWS AND REACTIVATE¿ ALERT SCREEN WHEN THE JAWS WERE FULLY OR PARTIALLY CLOSED. THE "REPLACE INSTRUMENT" ALERT SCREEN WOULD BE DISPLAYED AFTER RECEIVING THE "REPOSITION AND REACTIVATE" ALERT SCREEN TWICE. THIS WAS NOT CONFIRMED DURING THE ANALYSIS. BECAUSE THE CABLE WAS CUT OFF NO FUNCTIONAL TESTING COULD BE PERFORMED WITH THE GENERATOR. A PROBABLE CAUSE OF THE DAMAGE TO THE RETENTION PINS COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01934 | 2015-03-25 | 20150317 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | IT WAS REPORTED THAT DURING A LAPAROSCOPIC LIVER RESECTION THE FIRST INSTRUMENT'S TOP JAW BROKE OFF COMPLETELY AS THE I-BLADE WAS ADVANCED. BOTH THE INSTRUMENT AND THE COMPETITIVE, 5 MM PLASTIC TROCAR, HAD TO BE REMOVED BECAUSE THE INSTRUMENT COULDN'T BE REMOVED THROUGH THE TROCAR. THE BROKEN PIECE OF JAW WAS REMOVED, AS WELL. THE PATIENT HAD TO BE REINSUFLATED, WHICH RESULTED IN A TWENTY MINUTE DELAY IN THE PROCEDURE. THE PROCEDURE WAS COMPLETED USING A COMPETITIVE DEVICE WITH NO CONSEQUENCE TO THE PATIENT. |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00050 | 2015-03-24 | 20150201 | Injury | Manufacturer report | IUNI G2 | UNICONDYLAR KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M5722INT0600210 |  | PATIENT DEVELOPED A RASH FOLLOWING KNEE REPLACEMENT SURGERY AND REPORTED TESTING POSITIVE FOR COBALT CHROME ALLERGY. SURGERY IS PLANNED TO REMOVE THE IMPLANT. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00052 | 2015-03-24 | 20150201 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600020 |  | REVISION SURGERY IS PLANNED TO EXCHANGE THE POLY INSERTS DUE TO SOFT TISSUE LAXITY. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00053 | 2015-03-24 | 20150201 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M5725INT060XE10 |  | REVISION SURGERY IS PLANNED TO EXCHANGE THE POLY INSERTS. REASON FOR REVISION IS UNKNOWN. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00054 | 2015-03-24 | 20150301 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600010 |  | REVISION SURGERY IS PLANNED FOR PATIENT WITH A TOTAL KNEE IMPLANT. REASON FOR REVISION IS UNKNOWN. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. |  | Y | N | HEALTH PROFESSIONAL
3005075853-2015-01988 | 2015-03-30 | 20150317 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35A |  | (B)(4). ADDITIONAL INFORMATION: THE DEVICE WAS RETURNED WITH THE UPPER JAW DETACHED AND NOT RETURNED. IN ADDITION, THE CABLE WAS CUT OFF. UPON VISUAL INSPECTION TO THE DEVICE NO ANOMALIES WERE FOUND WITH THE ELECTRODE AND CERAMIC AS REPORTED IN THE EVENT DESCRIPTION. DUE TO THE JAW WAS DETACHED AND THE CABLE WAS CUT OFF NOT ALL FUNCTIONAL TESTING COULD BE PERFORMED WITH THE GENERATOR. A PROBABLE CAUSE OF THE DAMAGE TO THE JAW COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00051 | 2015-03-24 | 20150201 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M5725INT0600020 |  | PATIENT DEVELOPED MID-FLEXION INSTABILITY. REVISION SURGERY IS PLANNED TO EXCHANGE THE POLY INSERTS. |  | Y | N | HEALTH PROFESSIONAL
3005075853-2015-01989 | 2015-03-30 | 20150226 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C45A |  | IT WAS REPORTED THAT DURING A ROBOTIC PROSTATECTOMY PROCEDURE, THE JAWS WOULD NOT OPEN. THERE WAS TISSUE IN THE JAWS. NO INFORMATION WAS KNOWN ON HOW THE JAWS WERE OPENED. CASE COMPLETED WITH ANOTHER DEVICE OF THE SAME PRODUCT CODE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01990 | 2015-03-30 | 20150304 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | IT WAS REPORTED THAT DURING A LAPAROSCOPIC HYSTERECTOMY WITH THE INSTRUMENT AND THE JAWS OF THE INSTRUMENT WOULDN'T OPEN AFTER RESECTING THE TISSUE. THE DOCTOR FORCEFULLY OPENED THE JAWS AND CONTINUED WITH THE PROCEDURE. WHILE ATTEMPTING TO ADVANCE THE KNIFE BLADE ON ANOTHER RESECTION, THE KNIFE BLADE WOULDN'T ADVANCE. THE DOCTOR MANUALLY FORCED THE JAWS OPEN AND REMOVED THE INSTRUMENT FROM THE PATIENT. THE DOCTOR COMPLETED THE PROCEDURE USING A COMPETITORS BIPOLAR DEVICE WITH NO CONSEQUENCE TO THE PATIENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-01992 | 2015-03-30 |  | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35A |  | IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE JAW WAS STICKING; UNABLE TO OPEN. IT IS UNKNOWN HOW THE CASE WAS COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-02000 | 2015-03-30 | 20150302 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35 |  | (B)(4). ADDITIONAL INFORMATION: THE DEVICE WAS RETURNED WITH THE UPPER JAW DETACHED AND IT WAS RETURNED WITH THE DEVICE. AN IMAGE OF THE INSTRUMENT WAS ALSO SENT, IT MATCHES THE PHYSICAL ANALYSIS. THE DEVICE WAS CONNECTED TO THE GENERATOR AND IT WAS RECOGNIZED. BECAUSE THE JAW WAS DETACHED NOT ALL FUNCTIONAL TESTING COULD BE PERFORMED WITH THE GENERATOR. A PROBABLE CAUSE OF THE DAMAGE TO THE JAW COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-02003 | 2015-03-30 | 20150210 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S45 |  | IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE STOPPED WORKING. ANOTHER DEVICE WAS PULLED TO COMPLETE THE PROCEDURE WITH NO PATIENT CONSEQUENCE. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-02013 | 2015-03-30 | 20150210 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S45 |  | IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE JAW BROKE ON THE DEVICE. ANOTHER DEVICE WAS PULLED TO COMPLETE THE PROCEDURE WITH NO PATIENT CONSEQUENCE. |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00055 | 2015-03-24 | 20150301 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57250600020 |  | REVISION SURGERY TO EXCHANGE THE POLY INSERTS IS PLANNED DUE TO LAXITY IN THE KNEE. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00049 | 2015-03-24 | 20150201 | Malfunction | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600040 |  | A DISPOSABLE CUTTING GUIDE INSTRUMENT BROKE FOLLOWING IMPACTION ONTO THE BONE DURING SURGERY. THE SURGEON ALTERED THE TECHNIQUE SLIGHTLY TO COMPLETE THE SURGERY. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-02050 | 2015-03-31 | 20150323 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | (B)(4). INFORMATION IS UNAVAILABLE; DEVICE NOT RETURNED. NO DEVICE RECEIVED FOR ANALYSIS AT TIME OF SUBMISSION OF 3500A. WHEN ADDITIONAL INFORMATION IS RECEIVED AND/OR THE DEVICE ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL MEDWATCH WILL BE SENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-02101 | 2015-04-03 | 20150319 | Injury | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | IT WAS REPORTED THAT AFTER A THORACOSCOPIC ESOPHAGECTOMY WITH OPEN GASTRIC TUBE RECONSTRUCTION ON (B)(6) 2015, POST-OPERATIVE BLEEDING OCCURRED FROM THE SHORT GASTRIC ARTERY THREE DAYS AFTER THE INITIAL OPERATION. REOPERATION, OPENING THE ABDOMEN, WAS PERFORMED TO STOP THE MAJOR LEAK BY CLIPS AND ELECTROSURGICAL SCALPEL. DURING THE THORACOSCOPIC ESOPHAGECTOMY, THE SALES REP ATTENDED AND CHECKED THAT THE DOCTOR USED THE DEVICE PROPERLY. THE SALES REP CAME OUT OF THE OR BEFORE THE BEGINNING, THE RECONSTRUCTION. THE PATIENT HAD HAD A RIGHT HEMICOLECTOMY A FEW YEARS AGO. DURING THE RECONSTRUCTION, THE DOCTOR FOUND THAT THE RIGHT GASTRO-OMENTAL ARTERY HAD BEEN CUT AT THE TIME. THEREFORE, A VASCULAR BYPASS SURGERY WITH INTERNAL THORACIC ARTERY WAS REQUIRED WITH THE RECONSTRUCTION, TO KEEP THE BLOOD FLOWING. DURING THE BYPASS SURGERY, THE DEVICE WAS USED TO SEAL WITHOUT CLIPS ON THE GASTRIC ARTERY THAT BLEEDING OCCURRED FROM. BLOOD TRANSFUSION WAS REQUIRED, BUT THE VOLUME OF BLEEDING AND TRANSFUSION WERE UNKNOWN. AS OF NOW, THE PATIENT IS IN THE HOSPITAL AND THE PATIENT¿S CONDITION IS STABLE. NO DEVICE WILL BE RETURNING. |  | Y | N | HEALTH PROFESSIONAL
3005075853-2015-02105 | 2015-04-03 | 20150305 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35 |  | (B)(4). SHOULD THE INFORMATION BE PROVIDED LATER, A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE DEVICE WAS RECEIVED THE UPPER JAW DAMAGED AT THE PROXIMAL SIDE. ONE OF THE HINGES WAS PROTRUDING. THE DEVICE WAS CONNECTED TO THE GENERATOR AND IT WAS RECOGNIZED. BECAUSE THE JAW WAS DAMAGED NOT ALL FUNCTIONAL TESTING COULD BE PERFORMED WITH THE GENERATOR. THE CONDITION OF THE JAWS PREVENTED THE FUNCTIONALITY OF THE JAW. THE DEVICE WAS UNABLE TO CYCLE OPEN AND CLOSE. A PROBABLE CAUSE OF THE DAMAGE TO THE JAW COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-02104 | 2015-04-03 | 20150326 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | IT WAS REPORTED THAT THE CUSTOMER STATED THAT THE JAW FELL OFF INTO THE PATIENT DURING A TOTAL LAP HYSTERECTOMY. DEVICE WAS RETRIEVED THROUGH THE TROCAR. NO PATIENT CONSEQUENCES. PROCEDURE COMPLETED USING A SECOND DEVICE. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-02187 | 2015-04-07 | 20150331 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35 |  | IT WAS REPORTED THAT DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY PROCEDURE THE I-BLADE DID NOT REACH TO THE END OF THE JAW. THE BLADE WAS STUCK DUE TO THE TISSUE AND TOP JAW WAS BROKEN WHEN THE SURGEON TRIED TO ADVANCE THE I-BLADE. THE BROKEN PART WAS REMOVED FROM THE PATIENT¿S BODY. IT IS NOT KNOWN HOW THE PROCEDURE WAS COMPLETED. |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00059 | 2015-04-02 | 20150101 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600020 |  | PT WAS REVISED TO A REVISION TOTAL KNEE SYSTEM. REASON FOR REVISION IS UNK. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00058 | 2015-04-03 | 20150301 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57250600020 |  | PT WITH A TOTAL KNEE IMPLANT HAD A TIGHT KNEE POST OPERATIVELY. REVISION SURGERY OCCURRED TO CLEAN THE POSTERIOR KNEE AND EXCHANGE THE POLY INSERTS. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00057 | 2015-04-01 | 20150301 | Malfunction | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57250600020 |  | STERILE BARRIER OF THE INNER POUCH CONTAINING THE TIBIAL IMPLANT WAS COMPROMISED. THE OUTER POUCH WAS INTACT. THE SURGERY WAS COMPLETED SUCCESSFULLY. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. ALL PROCESSING STEPS WERE COMPLETED SUCCESSFULLY. |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00056 | 2015-04-01 | 20150301 | Malfunction | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M5725INT0600020 |  | THE SURGEON WAS UNABLE TO INSERT THE LATERAL POLY INSERTS INTO THE TIBIAL TRAY. THE INSERTS APPEARED TO BE TOO WIDE TO FIT THE TIBIAL TRAY IMPLANT. THE KNEE WAS CONVERTED TO AN OFF-THE-SHELF KNEE IMPLANT SYSTEM. SURGERY TIME WAS EXTENDED BY APPROXIMATELY 30 MINUTES. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-02252 | 2015-04-09 | 20150219 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35 |  | IT WAS REPORTED THAT DURING A LAPAROSCOPIC HYSTERECTOMY PROCEDURE, THE SURGEON PULLS THE TISSUE AND CLOSED THE JAWS TO SEND ENERGY TO THE STRUCTURE. WHEN THE DEVICE REACHED THE SECOND TONE, THE SURGEON FEELS THAT THE DEVICE STOPPED WORKING PROPERLY. LOOKING THROUGH THE LENS, HE REALIZED THAT THE UPPER PART OF THE JAW IS LOOSE. IT WAS IMPOSSIBLE TO RETRIEVE THE DEVICE FROM THE TROCAR. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-02282 | 2015-04-10 | 20150402 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C25 |  | (B)(4). INFORMATION IS UNAVAILABLE; DEVICE NOT RETURNED. NO DEVICE RECEIVED FOR ANALYSIS AT TIME OF SUBMISSION OF 3500A. WHEN ADDITIONAL INFORMATION IS RECEIVED AND/OR THE DEVICE ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL MEDWATCH WILL BE SENT. |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00060 | 2015-04-08 | 20150301 | Malfunction | Manufacturer report | IUNI G2 | UNICONDYLAR KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57220600220 |  | IT WAS REPORTED THAT THE DISPOSABLE FEMORAL CUTTING GUIDE DIDN'T CONFORM WELL WITH THE FEMORAL BONE. SURGEON ELECTED TO COMPLETE THE SURGERY USING AN ALTERNATE IMPLANT SYSTEM. A SMALL DELAY IN SURGERY WAS EXPERIENCED. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. |  | N | Y | HEALTH PROFESSIONAL
4706399 | 2015-04-09 | 20150401 | Malfunction | User Facility report | ENSEAL G2 | ELECTROSURGICAL, CUTTING AND COAGULATION | ETHICON ENDO-SURGERY, INC | * | NSLG2C35A |  | DEVICE WAS FIRING INTERMITTENTLY AND TRIGGER WAS STICKING, ALSO MACHINE KEPT GIVING ERROR CODE TO REPLACE DEVICE.======================MANUFACTURER RESPONSE FOR ENSEAL ARTICULATING DEVICE, (BRAND NOT PROVIDED) (PER SITE REPORTER).======================REP WAS NOTIFIED. |  | N | Y | PHYSICIAN
3004153240-2015-00065 | 2015-04-16 | 20150301 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600010 |  | INSTABILITY WAS REPORTED FOR PATIENT WITH A TOTAL KNEE IMPLANT. REVISION SURGERY IS PLANNED TO EXCHANGE THE POLY INSERTS. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00063 | 2015-04-16 | 20150301 | Injury | Manufacturer report | IUNI G2 | UNICONDYLAR KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57220600230 |  | PATIENT PRESENTED WITH MINOR PAIN AND RADIOGRAPHIC EVIDENCE OF FEMORAL COMPONENT LOOSENING. REVISION SURGERY IS PLANNED. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00062 | 2015-04-16 | 20150301 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600010 |  | PATIENT HAS LIMITED RANGE OF MOTION AND IS UNABLE TO FULLY EXTEND THE LEG. REVISION SURGERY OCCURRED TO PERFORM A LATERAL RELEASE AND EXCHANGE THE POLY INSERTS. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00064 | 2015-04-16 | 20150301 | Injury | Manufacturer report | IUNI G2 | UNICONDYLAR KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57220600230 |  | PATIENT PRESENTED WITH MINOR PAIN AND RADIOGRAPHIC EVIDENCE OF FEMORAL COMPONENT LOOSENING. REVISION SURGERY IS PLANNED. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00061 | 2015-04-16 | 20150301 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600020 |  | PATIENT REPORTED PAIN AND POST-OPERATIVE STAPH INFECTION FOLLOWING TOTAL KNEE REPLACEMENT SURGERY. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. |  | Y | N | HEALTH PROFESSIONAL
1818910-2015-19054 | 2015-04-21 | 20130422 | Injury | Manufacturer report | G2 PRESS FIT STEM SZ8L | HIP FEMORAL STEM/SLEEVE | DEPUY IRELAND 9616671 |  | 900000128 |  | THE PATIENT WAS SUBJECTED TO AN UNKNOWN INTERVENTION IN WHICH WAS IMPLANTED A DEVICE CALLED 'DEPUY G2 LATERAL NR 8- TESTINA MM. 36+1.5'. AT BLOOD TESTS , THE LEVELS OF CO-CR WERE VERY HIGH.. THE PATIENT WAS CONSTRICTED TO A HOSPITALIZATION TO REMOVE THE DEVICE. FOLLOWING THE RE-INTERVENTION, THE PATIENT 'WAS CONSTRICTED AT A LONG PERIOD OF REHABILITATION. UPDATE REC¿D: 21/04/2015 - PATIENT'S TRANSLATED CLINICAL DOCUMENT WAS RECEIVED. MEDICAL RECORDS WERE REVIEWED FOR MDR REPORTABILITY. RECORDS INDICATE UPON REVISION THE PATIENT WAS FOUND TO HAVE OSTEOLYSIS ON THE FEMORAL METAPHYSIS, A PSEUDOTUMOR, AND METALLOSIS. METALLOSIS, AND OSTEOLYSIS ARE BEING ADDED TO THE PATIENT CODES. THERE IS NO NEW INFORMATION THAT WOULD CHANGE THE EXISTING MDR DECISION. DOR: (B)(6) 2013. THE COMPLAINT WAS UPDATED ON: 05/02/2015. UPDATE REC¿D 23 AND 25 MARCH 2015 - THE PATIENT'S MEDICAL RECORDS WERE RECEIVED. MEDICAL RECORDS WERE REVIEWED FOR MDR REPORTABILITY. ACCORDING TO THE MEDICAL RECORDS UPON REVISION BROWN HAEMATIC FLUID WAS ALSO ENCOUNTERED. RECORDS ALSO INDICATE THE PATIENT HAD ELEVATED COBALT AND CHROMIUM LEVELS. AT THIS TIME THE FEMORAL STEM IS BEING REPORTED. THE COMPLAINT WAS UPDATED ON: 20/04/2015. |  | Y | Y | HEALTH PROFESSIONAL
1818910-2015-19102 | 2015-04-21 | 20150325 | Injury | Manufacturer report | G2 PRESS FIT STEM SZ8L | HIP FEMORAL STEM/SLEEVE | DEPUY IRELAND 9616671 |  | 900000128 |  | THIS COMPLAINT IS STILL UNDER INVESTIGATION. DEPUY WILL NOTIFY THE FDA OF THE RESULTS OF THIS INVESTIGATION ONCE IT HAS BEEN COMPLETED. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL MEDWATCH, A FOLLOW-UP MEDWATCH WILL BE FILED AS APPROPRIATE. (B)(4). |  | Y | N | HEALTH PROFESSIONAL
2954740-2015-00094 | 2015-04-21 | 20150326 | Malfunction | Manufacturer report | ORBIT GALAXY G2 MICROCOIL DELIVERY SYSTEM | ARTIFICIAL EMBOLIZATION DEVICE | CODMAN AND SHURTLEFF, INC |  | 641CX0408 |  | DEVICES USED WERE PRESIDIO (PC410051730/C27882), ORBIT GALAXY G2 (641CX0408/C29862) AND ECHELON (LOT/CATALOG UNK) MICROCATHETER. IT IS ANTICIPATED THAT THE DEVICE WILL BE RETURNED FOR ANALYSIS; HOWEVER, THE DEVICE HAS NOT YET BEEN RETURNED. INFORMATION REGARDING PATIENT AGE, GENDER, WEIGHT, AND CONCOMITANT MEDICATIONS WERE NOT PROVIDED. ADDITIONAL INFORMATION WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. THIS IS 1 OF 2 MDRS SUBMITTED FOR THIS COMPLAINT, WITH ASSOCIATED REFERENCE NUMBERS OF 2954740-2015-00094 AND 2954740-2015-00093. |  | N | Y | HEALTH PROFESSIONAL
2954740-2015-00095 | 2015-04-21 | 20150304 | Malfunction | Manufacturer report | ORBIT GALAXY G2 MICROCOIL DELIVERY SYSTEM | ARTIFICIAL EMBOLIZATION DEVICE | CODMAN AND SHURTLEFF, INC | N/A | 641HX0206 |  | THE CONTACT FROM THE HOSPITAL REPORTED THAT DURING EMBOLIZATION OF AN INTERNAL CAROTID ARTERY ANEURYSM, THE ORBIT GALAXY G2 COIL (641HX0206/C28215) WAS ATTEMPTED TO BE ADVANCED THROUGH THE ECHELON MICROCATHETER (DETAILS UNKNOWN) BUT FACED WITH STIFF RESISTANCE AT THE HUB. THE VESSEL WAS NOT CALCIFIED OR TORTUOUS. THE DOCTOR WAS UNABLE TO PUSH THE COIL THROUGH THE MICRO CATHETER. THEN THE DOCTOR WITHDREW THE G2 DELIVERY SYSTEM AND COIL AND USED ANOTHER COVIDIEN COIL (DETAILS UNKNOWN) OF THE SAME SIZE AND IT DEPLOYED WELL INSIDE THE ANEURYSM. AT THE TIME OF INITIAL CONTACT THE PRODUCT WAS AVAILABLE FOR RETURN. THERE WAS NO REPORTED PATIENT IMPACT ASSOCIATED WITH THIS COMPLAINT. THERE WAS NO SIGNIFICANT CLINICAL DELAY TO THE PROCEDURE DUE TO THE ISSUE. THE DEVICE DID NOT KINK OR BEND AT ANY TIME PRIOR TO THE RESISTANCE/FRICTION. CONCOMITANT DEVICES DID NOT KINK OR BEND AT ANY TIME. AN ADEQUATE CONTINUOUS FLUSH WAS MAINTAINED THROUGH THE CATHETER AT ALL TIMES. THE REPLACEMENT COIL WAS ADVANCED THROUGH THE SAME MICROCATHETER. IT WAS DISCOVERED DURING PRODUCT ANALYSIS THAT THERE WAS STRETCHING, COMPRESSION AND KINKING DAMAGE. |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00066 | 2015-04-16 | 20150301 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600030 |  | INFECTION WAS REPORTED FOR PATIENT WITH A TOTAL KNEE IMPLANT. REVISION SURGERY OCCURRED TO EXCHANGE THE POLY INSERTS. |  | Y | N | HEALTH PROFESSIONAL
3005075853-2015-02533 | 2015-04-22 | 20150414 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35A |  | IT WAS REPORTED THAT DURING A LAPAROSCOPIC LIVER RESECTION PROCEDURE, THE DEVICE JAMMED IN THE TISSUE AND WOULD NOT OPEN. TISSUE HAD TO BE CUT FROM THE JAWS OF THE DEVICE TO REMOVE THE DEVICE FROM THE PATIENT. THEY FIGURED THE DEFECT OF THE PRODUCT AND CARRY ON AND FINISHED THE PROCEDURE. IT IS UNKNOWN WHAT WAS USED TO COMPLETE THE PROCEDURE. THE PATIENT IS FINE AND IN STABLE CONDITION. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. |  | N | Y | HEALTH PROFESSIONAL
2954740-2015-00096 | 2015-04-23 | 20150404 | Malfunction | Manufacturer report | ORBIT GALAXY G2 MICROCOIL DELIVERY SYSTEM | ARTIFICIAL EMBOLIZATION DEVICE | CODMAN AND SHURTLEFF, INC |  | 641HX0204 |  | THE DOCTOR TRIED TO DEPLOY ORBIT GALAXY G2 THROUGH THE ECHELON MICROCATHETER, BUT THE MICROCATHETER AND OPERATOR FACED PROBLEM IN COILING THE ACOM ANEURYSM AS THE COIL WAS NOT CONFORMING WELL AND THE CATHETER WAS KICKING OUT. THE ISSUE WAS NOT DIFFICULTY POSITIONING THE COIL CAUSE BY THE ECHELON MC NOT MAINTAINING TARGET SITE, NOR THE COIL BEING TOO LARGE FOR THE ANEURYSM. THE COIL HAD POOR CONFORMABILITY, AND ALL GUIDELINES FOLLOWED DURING THE EVENT. THE COIL REMAINED ATTACHED TO THE DELIVERY SYSTEM. THERE WAS NO ADVERSE EVENT AND THE COMPONENT WILL BE RETURNED FOR ANALYSIS. NAME OF SURGERY-ACOM ANEURYSM. DATE OF SURGERY- (B)(6) 2015. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-02563 | 2015-04-23 | 20150408 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35 |  | (B)(4). INFORMATION WAS NOT PROVIDED BY THE INITIAL CONTACT. INFORMATION IS UNAVAILABLE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. AT THE TIME OF THIS SUBMISSION, THE DEVICE HAS NOT BEEN RETURNED FOR ANALYSIS. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-02570 | 2015-04-23 | 20150415 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C45 |  | IT WAS REPORTED THAT DURING A LAVH PROCEDURE, THE ARTICULATING DEVICE BECAME STUCK ON TISSUE. CUSTOMER TRIED A NON-ARTICULATING DEVICE AND THE JAWS BECAME STUCK CLOSED ALSO BUT NOT ON TISSUE. CUSTOMER WAS ABLE TO CUT AROUND THE ARTICULATING DEVICE USING A HARMONIC ACE DEVICE TO FREE IT FROM THE TISSUE. PROCEDURE COMPLETED USING THE HARMONIC ACE DEVICE. NO PATIENT CONSEQUENCES |  | N | Y | HEALTH PROFESSIONAL
MW5042209 | 2015-04-13 | 20150413 | Injury | Voluntary report | G2 | IVC FILTER | BARD | G2 |  |  | G2 FILTER FRACTURED AND EMBOLIZED. FOUND INCIDENTALLY WHEN PT PRESENTED WITH UNRELATED MI. |  | N | Y | HEALTH PROFESSIONAL
2020394-2015-00522 | 2015-04-22 | 20120511 | Injury | Manufacturer report | G2 FILTER SYSTEM - FEMORAL | VENA CAVA FILTER | BARD PERIPHERAL VASCULAR, INC. |  | RF310F |  | IT WAS REPORTED THAT APPROXIMATELY 4 YRS POST FILTER DEPLOYMENT, A CT SCAN DEMONSTRATED A TILTED FILTER THAT PERFORATED THE IVC AND 2 DETACHED FILTER LIMBS. SURGERY WAS PERFORMED TO SUCCESSFULLY REMOVE THE VENA CAVA FILTER FROM THE IVC. THE PATIENT STATUS AT THIS TIME IS UNK. |  | Y | Y | HEALTH PROFESSIONAL
3004153240-2015-00070 | 2015-04-22 | 20150301 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57250600010 |  | PATIENT IS HAVING TROUBLE GAINING FULL EXTENSION. REVISION SURGERY IS PLANNED TO EXCHANGE THE POLY INSERTS. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00069 | 2015-04-22 | 20150330 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600010 |  | PATIENT REQUIRED MANIPULATION UNDER ANESTHESIA FOLLOWING TOTAL KNEE REPLACEMENT TO ADDRESS A FLEXION CONTRACTURE AND INCREASE RANGE OF MOTION. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00071 | 2015-04-22 | 20150401 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600010 |  | PATIENT DEVELOPED AN INFECTION. REVISION SURGERY OCCURRED TO EXCHANGE THE POLY INSERTS. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00067 | 2015-04-22 | 20150301 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600020 |  | PATELLA WASN'T TRACKING WELL FOR PATIENT WITH A TOTAL KNEE IMPLANT. REVISION SURGERY OCCURRED TO EXCHANGE THE PATELLA IMPLANT AND POLY INSERTS. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00068 | 2015-04-22 | 20150301 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57250600020 |  | PATIENT FELL AND EXPERIENCED A PATELLA TENDON TEAR. SURGERY OCCURRED TO PERFORM A LATERAL PATELLA TENDON RELEASE AND REMOVE SCAR TISSUE THAT WAS INHIBITING PROPER PATELLA TRACKING. ALL IMPLANTED COMPONENTS WERE LEFT IN PLACE. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. |  | Y | N | HEALTH PROFESSIONAL
3005075853-2015-02700 | 2015-04-29 | 20150313 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | IT WAS REPORTED THAT DURING A TOTAL LAPAROSCOPIC HYSTERECTOMY WITH THE ENSEAL INSTRUMENT AND THE DOCTOR REPORTED THE JAWS OF THE INSTRUMENT WERE DIFFICULT TO CLOSE AND THE I-BLADE WAS DIFFICULT TO ADVANCE TO CUT THE TISSUE. THE PROCEDURE WAS COMPLETED BY CONTINUALLY CYCLING THE POWER ON THE GENERATOR WITH NO CONSEQUENCE TO THE PATIENT. |  | N | Y | HEALTH PROFESSIONAL
2954740-2015-00100 | 2015-04-29 | 20150406 | Malfunction | Manufacturer report | ORBIT GALAXY G2 MICROCOIL DELIVERY SYSTEM | ARTIFICIAL EMBOLIZATION DEVICE | CODMAN AND SHURTLEFF, INC |  | 641CX2505 |  | THE CONTACT AT THE HOSPITAL REPORTED THAT THE DOCTOR TRIED TO DEPLOY THE PRESIDIO COIL (DETAILS UNKNOWN) AND ORBIT GALAXY G2 (641CX2505/C15306) THROUGH THE ECHELON MICROCATHETER (DETAILS UNKNOWN) BUT THE MICROCATHETER FACED STIFF RESISTANCE AND THE COIL DID NOT BREAK PROPERLY INSIDE THE ANEURYSM AND AFTER TWO LOOPS, THE CATHETER KICKED BACK. THE DOCTOR USED A SIMILAR SIZED TARGET COIL FROM STRYKER AND IT DEPLOYED VERY WELL. THERE WAS NO REPORTED PATIENT IMPACT OR SIGNIFICANT PROCEDURAL DELAY. AT THE TIME OF INITIAL CONTACT THE ORBIT GALAXY G2 WAS AVAILABLE FOR RETURN. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-02745 | 2015-04-30 | 20150415 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | IT WAS REPORTED THAT DURING A TOTAL LAPAROSCOPIC HYSTERECTOMY PROCEDURE, PART OF THE JAW BROKE AND FELL INTO THE PATIENT DURING ACTIVATION. THE SURGEON WAS ABLE TO RETRIEVE THE BROKEN PIECE FROM THE PATIENT. ANOTHER DEVICE LIKE WAS OPEN. CASE COMPLETED WITH ANOTHER DEVICE OF THE SAME PRODUCT CODE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-02776 | 2015-05-04 | 20150225 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35 |  | RECEIVED USER FACILITY MEDWATCH REPORT FORM #(B)(4), ADVISING THAT DURING A LAPAROSCOPIC LOWER ANTERIOR RESECTION WITH COLONOSCOPY PROCEDURE; THE TIP OF THE DEVICE BROKE OFF INTO THE PATIENT¿S PNEUMOPERITONEUM. THE DOCTOR WAS ABLE TO RETRIEVE THE TIP FROM THE PATIENT¿S PNEUMOPERITONEUM. THE DEVICE WAS REPLACED AND THE PROCEDURE CONTINUED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. |  | N | Y | HEALTH PROFESSIONAL
2020394-2015-00533 | 2015-04-28 | 20150407 | Injury | Manufacturer report | G2 FILTER SYSTEM |  | BARD PERIPHERAL VASCULAR, INC. |  | UNKNOWN G2 |  | RECEIVED A MEDWATCH REPORT (MW5042209) ON 07/07/2015 FROM THE FDA. |  | Y | Y | HEALTH PROFESSIONAL
3004153240-2015-00073 | 2015-04-30 | 20150302 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600020 |  | PATIENT REQUIRED MANIPULATION UNDER ANESTHESIA FOLLOWING TOTAL KNEE REPLACEMENT TO INCREASE RANGE OF MOTION. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00072 | 2015-04-30 | 20150313 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600020 |  | PATIENT REQUIRED MANIPULATION UNDER ANESTHESIA FOLLOWING TOTAL KNEE REPLACEMENT TO INCREASE RANGE OF MOTION. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. |  | Y | N | HEALTH PROFESSIONAL
3005075853-2015-02905 | 2015-05-07 | 20150325 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C25 |  | IT WAS REPORTED THAT DURING AN OPEN GASTRECTOMY, THE JAW DID NOT OPEN FULLY. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00074 | 2015-05-05 | 20140401 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57250600010 |  | PT IS EXPERIENCING PAIN. BONE SCAN DEMONSTRATED INCREASED ACTIVITY UNDERNEATH THE LATERAL TIBIAL PLATEAU. CT SCAN DEMONSTRATED SOME RESORPTION OF A POSTERIOR OSTEOARTICULAR ALLOGRAFT UNDERNEATH THE LATERAL TIBIAL PLATEAU. REVISION SURGERY IS PLANNED. |  | Y | N | HEALTH PROFESSIONAL
3005075853-2015-02933 | 2015-05-11 | 20150129 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35 |  | RECEIVED USER FACILITY MEDWATCH REPORT FORM #(B)(4), ADVISING THAT DURING A LAPAROSCOPIC COLECTOMY WITH LOW PELVIC ANASTOMOSIS AND MOBILIZATION OF THE SPLENIC FLEXURE AND SBR PROCEDURE; THE DEVICE MALFUNCTIONED AND STOPPED OPERATING DURING THE COURSE OF THE PROCEDURE. THE DEVICE TIP FELL OFF AFTER BEING REMOVED FROM THE PATIENT AND NO FRAGMENTS REMAINED IN THE PATIENT. IT IS NOT KNOWN HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-02938 | 2015-05-11 | 20150421 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35 |  | IT WAS REPORTED THAT DURING A LAPAROSCOPIC MYOMECTOMY, THE JAW BECAME NOT TO OPEN AFTER ONE SEALING ACTUATION. ANOTHER COMPETITIVE DEVICE WAS USED TO COMPLETE THE CASE. THERE IS NO INFORMATION HOW THE JAWS WERE OPENED BUT THE DOCTOR SAID THAT THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. THERE WAS A POSSIBILITY THAT THE DEVICE HAD CLAMPED VERY THICK TISSUE. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-02941 | 2015-05-11 | 20150313 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | IT WAS REPORTED THAT DURING A LAPAROSCOPIC HEPATECTOMY PROCEDURE, WHILE SEPARATING THE LIVER TISSUE TO REMOVE THE TUMOR, THE JAW OF DEVICE BROKE WHEN TAKING A SECOND BITE. NO PIECES OF THE DEVICE BROKE OFF. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-02951 | 2015-05-11 | 20150326 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35A |  | (B)(4). SHOULD THE INFORMATION BE PROVIDED LATER, A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE DEVICE C WAS RECEIVED WITH THE TOP OF THE I-BLADE UPPER TIP BROKEN OFF NOT RETURNED. THE DEVICE WAS CONNECTED TO THE GENERATOR AND IT WAS RECOGNIZED. BECAUSE THE I BLADE WAS DAMAGED NOT ALL FUNCTIONAL TESTING COULD BE PERFORMED WITH THE GENERATOR. THE CONDITION OF THE BLADE PREVENTED THE FUNCTIONALITY OF THE JAW. THE JAW WAS UNABLE TO CYCLE OPEN AND CLOSE. A PROBABLE CAUSE OF THE DAMAGE TO THE I BLADE COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-02954 | 2015-05-11 | 20150326 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35A |  | IT WAS REPORTED THAT DURING A ROBOTIC PROSTATECTOMY PROCEDURE, THE STAFF REPORTED THAT THE BLADE OF THE DEVICE BROKE OFF FROM THE UPPER JAW COMPONENT DURING USE. THE DEVICE IS SEEN TO BE MISSING PART OF THE GOLD BLADE ON TOP JAW. ANOTHER DEVICE WAS OPENED TO COMPLETE THE CASE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00077 | 2015-05-08 | 20150401 | Injury | Manufacturer report | IUNI G2 | UNICONDYLAR KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M5722INT0600210 |  | PATIENT DEVELOPED AN INFECTION. REVISION SURGERY OCCURRED TO EXCHANGE THE POLY INSERT. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00081 | 2015-05-08 | 20150401 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600030 |  | THE PATIENT HAS LIMITED RANGE OF MOTION POST-OPERATIVELY. THE SURGEON MANIPULATED THE KNEE AND DEBRIDED ARTHROSCOPICALLY. SURGERY IS PLANNED TO OPEN THE KNEE AND EXCHANGE THE POLY INSERTS. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00080 | 2015-05-08 | 20150909 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | MS7250600010 |  | THE PT IS IN A FEW DEGREES OF VARUS POST-OPERATIVELY. REVISION SURGERY IS PLANNED TO EXCHANGE THE POLY INSERTS. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00082 | 2015-05-08 | 20150401 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600020 |  | PT WITH A TOTAL KNEE REPLACEMENT WAS REVISED TO A REVISION TOTAL KNEE SYSTEM. REASON FOR REVISION IS UNK. |  | Y | N | HEALTH PROFESSIONAL
3005075853-2015-02991 | 2015-05-12 | 20150326 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C45 |  | IT WAS REPORTED THAT PRIOR TO AN UNKNOWN PROCEDURE, THE JAW COMPARTMENT WAS FOUND TO BE BROKEN WHEN THE PACKAGE WAS OPENED. IT IS UNKNOWN WHAT WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-02993 | 2015-05-12 | 20150403 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S25 |  | IT WAS REPORTED THAT THE JAW WAS BROKEN DURING A TOTAL VAGINAL HYSTERECTOMY. CUSTOMER USED A SECOND DEVICE TO COMPLETE THE PROCEDURE. NO PATIENT CONSEQUENCES. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-02992 | 2015-05-12 | 20150428 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35A |  | IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE JAWS OF THE DEVICE WOULD NOT OPEN. THE DEVICE WAS NOT ON PATIENT TISSUE WHEN THIS OCCURRED. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. NO PATIENT CONSEQUENCES WERE REPORTED. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-02999 | 2015-05-12 | 20150403 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35 |  | IT WAS REPORTED THAT THE JAWS OF THE DEVICE WOULD NOT CLOSE DURING A LAP SIGMOID RESECTION. THE PROCEDURE WAS COMPLETED WITH A SECOND SAME DEVICE. NO PATIENT CONSEQUENCE. |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00075 | 2015-05-05 | 20150401 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600010 |  | AT THE END OF THE TRIALING PHASE DURING A TOTAL KNEE REPLACEMENT PROCEDURE, IT WAS IDENTIFIED THAT THE KNEE WAS TOO LOOSE LATERALLY. CONVERSION TO AN ALTERNATE IMPLANT SYSTEM OCCURRED INTRAOPERATIVELY. THERE WAS A ONE HOUR DELAY IN SURGERY BECAUSE THE INSTRUMENTS FOR THE ALTERNATE IMPLANT SYSTEM WERE NOT AVAILABLE. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00079 | 2015-05-05 | 20150401 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600010 |  | PT HAS FORMED ADHESIONS. SYNOVECTOMY PROCEDURE IS PLANNED. POLY INSERTS MAY BE EXCHANGED DURING THE PROCEDURE. |  | Y | N | HEALTH PROFESSIONAL
3005075853-2015-03019 | 2015-05-13 | 20150429 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35A |  | (B)(4) . NO DEVICE RECEIVED FOR ANALYSIS AT TIME OF SUBMISSION OF 3500A. WHEN ADDITIONAL INFORMATION IS RECEIVED AND/OR THE DEVICE ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL MEDWATCH WILL BE SENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03020 | 2015-05-13 | 20150429 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35A |  | IT WAS REPORTED THAT DURING A LAPAROSCOPIC TOTAL HYSTERECTOMY PROCEDURE, THE JAWS WOULD NOT OPEN AFTER TEN MINUTES OF USE. IT WAS UNKNOWN IF THE DEVICE WAS ON TISSUE WHEN IT WOULD NOT OPEN. ANOTHER DEVICE OF THE SAME PRODUCT CODE WAS USED TO COMPLETE THE PROCEDURE. NO PATIENT CONSEQUENCES WERE REPORTED. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03024 | 2015-05-13 | 20150420 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35A |  | (B)(4). ADDITIONAL INFORMATION: THE DEVICE WAS RECEIVED IN GOOD PHYSICAL CONDITION. THE DEVICE WAS TESTED ON THE GENERATOR AND PASSED ALL FUNCTIONAL TESTING. THE ENERGY OUTPUT DELIVERED FROM THE DEVICE WAS VERIFIED. ALL THREE TONES WERE HEARD DURING FUNCTIONAL TESTING (TONE 1 IS HEARD WHEN THE ENERGY ACTIVATION BUTTON IS PRESSED; TONE 2 IS HEARD WHEN TISSUE IMPEDANCE THRESHOLD IS REACHED AND TONE 3 IS HEARD WHEN THE CYCLE IS COMPLETE). NO ISSUES WERE NOTED WITH OPENING AND CLOSING OF THE JAWS. THERE WERE NO ANOMALIES NOTED WITH THE FUNCTIONALITY OF THE DEVICE. |  | N | Y | HEALTH PROFESSIONAL
2954740-2015-00115 | 2015-05-13 | 20150422 | Injury | Manufacturer report | ORBIT GALAXY G2 MICROCOIL DELIVERY SYSTEM | ARTIFICIAL EMBOLIZATION DEVICE | CODMAN AND SHURTLEFF, INC | N/A | 641CX0308 |  | THE COIL WAS RETURNED HOWEVER VERY LIMITED INFORMATION WAS RECEIVED. THE DEVICE POSITIONING UNIT (DPU) WAS NOT RETURNED. THE PROXIMAL SECTION OF THE SL-10 MICROCATHETER WAS SEVERED AND NOT RETURNED. THE COIL WAS RETURNED PARTIALLY PROTRUDING OUTSIDE THE MICROCATHETER. THE PROXIMAL SECTION OF THE COIL WAS STRETCHED. THE DISTAL SECTION OF THE COIL WAS ALSO DAMAGED. BASED ON THE EVIDENCE RECEIVED AND FROM THE END USER¿S NARRATIVE, THE MOST LIKELY CONTRIBUTING FACTOR TO THE COIL¿S RESISTANCE MAY HAVE OCCURRED DURING THE POSITIONING OF THE COIL INSIDE THE ANEURYSM DUE TO DISTAL INTERFERENCE FROM EITHER THE COILS ALREADY DWELLING INSIDE THE ANEURYSM (IF PREVIOUSLY PLACED), ON ITSELF, OR FROM THE DISTAL TIP OF THE MICROCATHETER IF PLACED AT AN ACUTE ANGLE. THE COIL STRETCHING MOST LIKELY OCCURRED WHEN THE COIL BECAME TEMPORARILY ANCHORED ON COILS ALREADY DWELLING INSIDE THE ANEURYSM, ON ITSELF, OR FROM THE DISTAL TIP OF THE MICROCATHETER. WHEN THE ANCHORED COIL WAS RETRACTED EITHER FOR REPOSITIONING INSIDE THE ANEURYSM OR DURING REMOVAL FROM THE PATIENT, THE COIL THEN STRETCHED WHICH MAY HAVE CAUSED THE UNINTENDED COIL DETACHMENT, HOWEVER WITHOUT THE RETURN OF THE DPU USED IN THE PROCEDURE, THE EXACT CONTRIBUTING FACTORS TO THE COIL¿S RESISTANCE, ENTANGLEMENT, ANCHORING, STRETCHING, AND UNINTENDED DETACHMENT CANNOT BE DEFINITIVELY DETERMINED NOR CAN THE STIFFNESS IN THE DETACHMENT ZONE WHICH MAY HAVE INFLUENCED THE MICROCATHETER¿S POSITION AS REPORTED BE DETERMINED AS TO THE EXACT ROOT CAUSE. THE END USER¿S NARRATIVE IN THE COMPLAINT EVENT DOES AGREE WITH BOTH THE EVIDENCE AND THE LABORATORY¿S FINDINGS. IT WAS NOT STATED IF A ONE TO ONE MOVEMENT WAS OBSERVED OR IF THE MICROCATHETER WAS AN ACUTE ANGLE TO THE ANEURYSM¿S NECK. IF THIS WAS THE CASE, THEN FOR OPTIMUM PRODUCT PERFORMANCE AND TO PREVENT POTENTIAL COMPLICATIONS, THE INSTRUCTIONS FOR USE (IFU) RECOMMENDS, ¿CAUTION: IF REPOSITIONING OF THE MICROCOIL IS NECESSARY, CAREFULLY OBSERVE THE MOTION OF THE MICROCOIL IN RESPECT TO THE DPU WIRE WHILE RETRACTING THE MICROCOIL UNDER FLUOROSCOPY. IF THE MICROCOIL MOVEMENT IS NOT ONE-TO-ONE WITH THE DPU WIRE, OR IF REPOSITIONING IS DIFFICULT, THE MICROCOIL MAY HAVE BECOME STRETCHED AND COULD POSSIBLY BREAK. GENTLY REMOVE AND DISCARD THE MICROCOIL SYSTEM. CAUTION: IF THE MICROCOIL IS POSITIONED AT A RELATIVE SHARP ANGLE TO THE MICROCATHETER, A MICROCOIL MAY STRETCH OR BREAK AS IT IS BEING WITHDRAWN. BY REPOSITIONING THE DISTAL TIP OF THE CATHETER AT OR SLIGHTLY INSIDE THE OSTIUM OF THE ANEURYSM, THE MICROCOIL MAY BE MORE EASILY FUNNELED BACK INTO THE MICROCATHETER.¿ IN ADDITION, WITHOUT THE RETURN OF THE DPU USED IN THE PROCEDURE, IT CANNOT BE DETERMINED IF THIS COMPONENT HAD ANY ADDITIONAL CONTRIBUTIONS TO THE COMPLAINT EVENT. DHR: A REVIEW OF THE MANUFACTURING DOCUMENTATION ASSOCIATED WITH THIS LOT PRESENTED NO ISSUES DURING THE MANUFACTURING OR INSPECTION PROCESS THAT CAN BE RELATED TO THE REPORTED COMPLAINT. THE COMPLAINT THAT THE COIL WAS STRETCHED WAS CONFIRMED DURING PRODUCT INVESTIGATION DURING VISUAL ANALYSIS. THE CAUSE OF THE STRETCH COULD BE RELATED TO ANOTHER FAILURE THE USER REPORTED, WHICH IS ENTANGLEMENT OF THE COIL WITH PREVIOUSLY PLACED COILS. WITHOUT THE RETURN OF THE DPU THE EXACT ROOT CAUSE OF THE ALLEGED ISSUES COULD NOT BE DETERMINED HOWEVER PROCEDURAL FACTORS OUTLINED IN THE IFU LIKELY CONTRIBUTED TO THE EVENT. ADDITIONALLY, REVIEW OF THIS LOT PRESENTED NO ISSUES DURING THE MANUFACTURING OR INSPECTION PROCESS THAT CAN BE RELATED TO THE REPORTED COMPLAINT. THEREFORE, NO CORRECTIVE ACTIONS WILL BE TAKEN AT THIS TIME. |  | Y | Y | HEALTH PROFESSIONAL
3005075853-2015-03056 | 2015-05-14 | 20150429 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | IT WAS REPORTED THAT DURING A TOTAL LAPAROSCOPIC HYSTERECTOMY (TLH) PROCEDURE, THE DEVICE BROKE PRIOR TO BEING FIRED BY THE SURGEON. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03059 | 2015-05-14 | 20150506 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | (B)(4). NO DEVICE RECEIVED FOR ANALYSIS AT TIME OF SUBMISSION OF 3500A. WHEN ADDITIONAL INFORMATION IS RECEIVED AND/OR THE DEVICE ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION REQUESTED: WERE THE JAWS DIFFICULT TO OPEN AND CLOSE? DID THE JAWS NOT OPEN (JAWS WERE STUCK CLOSED)? IF YES, WERE THE JAWS STUCK ON TISSUE WHEN THEY WOULD NOT OPEN? HOW WAS THE DEVICE REMOVED FROM THE TISSUE? WAS THERE ANY CHANGE TO THE PATIENT¿S POST-OP CARE AS A RESULT OF THIS ISSUE? |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03063 | 2015-05-14 | 20150506 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S45 |  | (B)(4). NO DEVICE RECEIVED FOR ANALYSIS AT TIME OF SUBMISSION OF 3500A. WHEN ADDITIONAL INFORMATION IS RECEIVED AND/OR THE DEVICE ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION REQUESTED DID THE BROKEN PIECE FALL INTO THE PATIENT? IF YES, WAS THE PIECE RETRIEVED? |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03094 | 2015-05-15 | 20150331 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35A |  | IT WAS REPORTED THAT DURING A LAPAROSCOPIC ASSISTED VAGINAL HYSTERECTOMY PROCEDURE, THE DEVICE WAS ACTIVATED A FEW TIMES ON UNKNOWN STAPLE LINE FROM A PREVIOUS SURGERY. FOLLOWING ERROR MESSAGES AND REACTIVATING, THE SURGEON CONTINUED TO USE THROUGHOUT THE CASE WITHOUT PROBLEM. LATE IN THE CASE THE SURGEON WAS ATTEMPTING TO ACTIVATE ON THICK VASCULAR TISSUE, GRASPED WITHIN THE JAWS AND EVEN WITH COMPRESSION AND ACTIVATION WAS UNABLE TO ADVANCE THE BLADE. THE SURGEON THEN ATTEMPTED SAME WITH LESS TISSUE IN THE JAW AS INSTRUCTED AND WAS STILL UNABLE TO ADVANCE THE BLADE.. A COMPETITOR PRODUCT WAS OPENED AND USED TO COMPLETE THE CASE. THERE WAS NO ADVERSE CONSEQUENCE TO THE PATIENT. |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00088 | 2015-05-15 | 20150401 | Injury | Manufacturer report | IUNI G2 | UNICONDYLAR KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57220600210 |  | REVISION SURGERY IS PLANNED FOR PATIENT WITH A UNICONDYLAR KNEE IMPLANT DUE TO SUSPECTED TIBIAL COMPONENT LOOSENING. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00086 | 2015-05-15 | 20150401 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M5725INT0600010 |  | THE PATIENT EXPERIENCED A FALL AND NOW HAS LATERAL INSTABILITY. CAUSE IS UNKNOWN. REVISION SURGERY IS PLANNED TO EXCHANGE THE POLY INSERTS. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00087 | 2015-05-15 | 20150401 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M5725INT0600020 |  | THE PATIENT HAS JOINT INSTABILITY POST-OPERATIVELY. SURGERY IS PLANNED TO DO A MEDIAL RELEASE AND EXCHANGE THE POLY INSERTS. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00085 | 2015-05-15 | 20150401 | Injury | Manufacturer report | IDUO G2 | BICOMPARTMENTAL KNEE REPLACEMENT | CONFORMIS, INC. |  | M57230600210 |  | THE TIBIAL COMPONENT HAS SHIFTED. REVISION SURGERY IS PLANNED TO EXCHANGE THE TIBIAL TRAY AND POLY INSERTS. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00089 | 2015-05-15 | 20150401 | Malfunction | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57250600010 |  | THE DISPOSABLE FEMORAL IMPACTOR TIP IJIG BROKE DURING USE. A SMALL PIECE BROKE OFF THE SURFACE BUT REMAINED ATTACHED TO THE MAIN BODY OF THE IMPACTOR TIP. THE SURGERY WAS COMPLETED SUCCESSFULLY. |  | N | Y | HEALTH PROFESSIONAL
1818910-2015-21115 | 2015-05-20 | 20150514 | Injury | Manufacturer report | G2 PRESS FIT STEM SZ2L | HIP FEMORAL STEM/SLEEVE | DEPUY ORTHOPAEDICS INC US |  | 900000122 |  | NO DEVICE ASSOCIATED WITH THIS REPORT WAS RECEIVED FOR EXAMINATION. BASED ON PREVIOUS INVESTIGATIONS THIS COMPLICATION OF JOINT REPLACEMENT IS UNLIKELY TO HAVE BEEN THE RESULT OF A DEVICE FAILING TO MEET REQUIRED SPECIFICATIONS. THE INFORMATION RECEIVED WILL BE RETAINED FOR POTENTIAL SERIES INVESTIGATIONS IF TRIGGERED BY TREND ANALYSIS, POST MARKET SURVEILLANCE, OR OTHER EVENTS WITHIN THE QUALITY SYSTEM. CORRECTIVE ACTION WAS NOT INDICATED. DEPUY CONSIDERS THE INVESTIGATION CLOSED. SHOULD ADDITIONAL INFORMATION BE RECEIVED, THE INFORMATION WILL BE REVIEWED AND THE INVESTIGATION WILL BE RE-OPENED AS NECESSARY. |  | Y | N | HEALTH PROFESSIONAL
2020394-2015-00617 | 2015-05-18 | 20110307 | Malfunction | Manufacturer report | G2 FILTER SYSTEM - JUGULAR | VENA CAVA FILTER | BARD PERIPHERAL VASCULAR, INC. |  | RF320J |  | IT WAS REPORTED THAT APPROXIMATELY TWO YEARS POST VENA CAVA FILTER DEPLOYMENT, FILTER MIGRATION IN THE IVC WAS REPORTED; HOWEVER, THE FINAL LOCATION OF THE FILTER IN THE IVC WAS NOT PROVIDED. NO ATTEMPT WAS MADE TO RETRIEVE THE FILTER ND REMAINS IMPLANTED. THE PATIENT STATUS IS UNKNOWN AT THIS TIME. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03175 | 2015-05-21 | 20150414 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35A |  | IT WAS REPORTED THAT THE DEVICE CAUSED MULTIPLE REACTIVATE ERRORS FOLLOWED BY A REPLACE INSTRUMENT ERROR EVERY TIME THE CUSTOMER ATTEMPTED TO ACTIVATE THE DEVICE ON ANY TYPE OF TISSUE DURING A LAP RIGHT CHOLECYSTECTOMY. THE PROCEDURE WAS COMPLETED WITH A SECOND SAME DEVICE. NO PATIENT CONSEQUENCE. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03187 | 2015-05-22 | 20150317 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | IT WAS REPORTED THAT DURING A LAPAROSCOPIC LIVER RESECTION THE TOP JAW BROKE OFF COMPLETELY AS THE I-BLADE WAS ADVANCED. BOTH THE INSTRUMENT AND THE COMPETITIVE, 5 MM PLASTIC TROCAR, HAD TO BE REMOVED BECAUSE THE INSTRUMENT COULDN'T BE REMOVED THROUGH THE TROCAR. THE BROKEN PIECE OF JAW WAS REMOVED, AS WELL. THE PATIENT HAD TO BE RE-INSUFFLATED, WHICH RESULTED IN A TWENTY MINUTE DELAY IN THE PROCEDURE. THE PROCEDURE WAS COMPLETED USING A COMPETITIVE DEVICE WITH NO CONSEQUENCE TO THE PATIENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03198 | 2015-05-26 | 20150205 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | (B)(4). SHOULD THE INFORMATION BE PROVIDED LATER, A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE DEVICE WAS RECEIVED WITH THE TOP OF THE I-BLADE UPPER TIP BROKEN OFF NOT RETURNED. UPON VISUAL INSPECTION TO THE DEVICE NO ANOMALIES WERE OBSERVED WITH THE HANDLE COMPONENT AS REPORTED. THE DEVICE WAS CONNECTED TO THE GENERATOR AND IT WAS RECOGNIZED. BECAUSE THE I BLADE WAS DAMAGED NOT ALL FUNCTIONAL TESTING COULD BE PERFORMED WITH THE GENERATOR. THE CONDITION OF THE BLADE PREVENTED THE FUNCTIONALITY OF THE JAW. THE JAW WAS UNABLE TO CYCLE OPEN AND CLOSE. A PROBABLE CAUSE OF THE DAMAGE TO THE I BLADE COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03201 | 2015-05-26 | 20150217 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35A |  | (B)(4). SHOULD THE INFORMATION BE PROVIDED LATER, A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE DEVICE WAS RECEIVED THE UPPER JAW BROKEN WITH ONE OF THE SIDE DISLODGED FROM THE DEVICE AND THE OTHER STILL ATTACHED TO THE DEVICE. IN ADDITION THE I-BLADE WAS DAMAGED AND THE CABLE WAS CUT OFF. DUE TO THE RETURNED CONDITION OF THE DEVICE, NO FUNCTIONAL TESTING COULD BE PERFORMED WITH THE GENERATOR. HOWEVER, THE DEVICE WAS MECHANICALLY TESTED AND NO ANOMALIES WERE NOTED. IT SHOULD BE NOTED THAT AS PART OF OUR QUALITY PROCESS, EACH DEVICE IS VISUALLY INSPECTED AND FUNCTIONALLY TESTED DURING MANUFACTURING TO ENSURE THE DEVICE MEETS THE REQUIRED SPECIFICATIONS PRIOR TO SHIPMENT. A PROBABLE CAUSE OF THE DAMAGE TO THE JAW AND THE I-BLADE COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00090 | 2015-05-21 | 20150420 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57250600020 |  | KNEE STIFFNESS WAS REPORTED AT TWO MONTHS POST-OP. MANIPULATION UNDER ANESTHESIA OCCURRED TO INCREASE RANGE OF MOTION. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00092 | 2015-05-21 | 20150401 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57250600010 |  | PATIENT WITH A TOTAL KNEE IMPLANT HAS SOFT TISSUE LAXITY. REVISION SURGERY IS PLANNED TO EXCHANGE THE POLY INSERTS. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. |  | Y | N | HEALTH PROFESSIONAL
3005075853-2015-03208 | 2015-05-26 | 20150415 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C25 |  | IT WAS REPORTED THAT DURING AN OPEN CYSTECTOMY PROCEDURE, I-BLADE HAD BROKEN DURING CUTTING AND ACTIVATING. THE TIP OF I-BLADE DROPPED TO THE PELVIS. THEY FOUND THAT PIECE DURING THE PROCEDURE. ALSO THE JAWS DAMAGED AT THE SAME TIME. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00091 | 2015-05-21 | 20150401 | Malfunction | Manufacturer report | ITTOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M5725INT0600010 |  | ONE OF THE THREE LATERAL POLY INSERT THICKNESSES PROVIDED IN THE IMPLANT KIT WAS MISSING FROM THE KIT. THIS WAS IDENTIFIED UPON OPENING THE KIT. SURGERY WAS COMPLETED SUCCESSFULLY USING THE AVAILABLE INSERT THICKNESSES. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03213 | 2015-05-26 | 20150505 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35 |  | (B)(4). ADDITIONAL INFORMATION: THE DEVICE WAS RETURNED WITH THE UPPER JAW DETACHED RETURNED AND THE TOP OF THE I-BLADE UPPER TIP BROKEN OFF BUT RETURNED. THE DEVICE WAS CONNECTED TO THE GENERATOR AND IT WAS RECOGNIZED. BECAUSE THE I BLADE AND THE JAW WERE DAMAGED NOT ALL FUNCTIONAL TESTING COULD BE PERFORMED WITH THE GENERATOR. THE CONDITION OF THE BLADE PREVENTED THE FUNCTIONALITY OF THE JAW. THE JAW WAS UNABLE TO CYCLE OPEN AND CLOSE. A PROBABLE CAUSE OF THE DAMAGE TO THE JAW COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03215 | 2015-05-26 | 20150506 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C25 |  | (B)(4). INFORMATION IS UNAVAILABLE; DEVICE NOT RETURNED. NO DEVICE RECEIVED FOR ANALYSIS AT TIME OF SUBMISSION OF 3500A. WHEN ADDITIONAL INFORMATION IS RECEIVED AND/OR THE DEVICE ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL MEDWATCH WILL BE SENT. |  | N | Y | HEALTH PROFESSIONAL
2954740-2015-00117 | 2015-05-28 | 20150422 | Malfunction | Manufacturer report | ORBIT GALAXY G2 MICROCOIL DELIVERY SYSTEM | ARTIFICIAL EMBOLIZATION DEVICE | CODMAN AND SHURTLEFF, INC |  | 641HX0204 |  | DURING COIL EMBOLIZATION OF AN ANTERIOR COMMUNICATING ARTERY ANEURYSM, A GALAXY G2 (641HX0204/ C19389) FAILED TO DETACH. THE COIL WAS ALMOST THE LAST COIL TO BE PLACED IN THE ANEURYSM. UPON ATTEMPTING TO DETACH THE COIL, NO SIGNAL WAS SEEN ON THE DETACHMENT BOX USED. THE DETACHMENT CABLE WAS EXCHANGED AND AGAIN THERE WAS NO SIGNAL DURING ATTEMPTED DETACHMENT. THEREFORE, THE PHYSICIAN RECAPTURED THE COIL AND USED ANOTHER COIL TO CONTINUE THE PROCEDURE. PRIOR TO LOADING THE NEW COIL INTO THE MICROCATHETER, HE CONNECTED THE PROXIMAL DETACHMENT PART OF THE COIL DELIVERY WIRE TO THE CONNECTING CABLE TO CONFIRM THAT THE DEVICE WAS WORKING, AND THERE WERE WAS A SIGNAL ON THE DETACHMENT BOX. THE SAME CABLE AND DETACHMENT CONTROL BOX USED WITH THE COMPLAINT PRODUCT WAS USED TO COMPLETE THE PROCEDURE. THE EVENT RESULTED IN A 5 MINUTE DELAY IN THE PROCEDURE, BUT NO PATIENT ADVERSE EVENT. THE PROCEDURE WAS CONTINUED WITH A SIMILAR PRODUCT. IT WAS REPORTED THAT A PRE-DEPLOYMENT ELECTRICAL CHECK HAD NOT BEEN PERFORMED. THE ACTUAL COIL DID NOT APPEAR DAMAGED. THE DEVICE HAD BEEN STORED AS PER THE LABELING INSTRUCTIONS. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03273 | 2015-06-01 | 20150415 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | IT WAS REPORTED THAT DURING A LAPAROSCOPIC HAND ASSIST COLECTOMY PROCEDURE, WHILE USING THE DEVICE ON THE COLON THAT WAS EXTRACORPOREALIZED, THE JAWS WOULD NOT CLOSE. THE DEVICE WORKED FINE AND THEN SUDDENLY STOPPED. ENERGY WOULD STILL ACTIVATE BUT NO MOVEMENT OF THE JAWS. THERE WERE NO PATIENT CONSEQUENCES. CASE COMPLETED WITH ANOTHER DEVICE OF THE SAME PRODUCT CODE. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03280 | 2015-06-01 | 20150514 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | UNK | NSLG2S35A |  | (B)(4). THE DEVICE WAS RETURNED WITH THE UPPER JAW DETACHED AND NOT RETURNED. IN ADDITION, THE DEVICE WAS RECEIVED WITH SKIVING OF THE HEAT SHRINK (SHAFT COVER) MATERIAL. ALL OF THE HEAT SHRINK MATERIAL APPEARED TO HAVE BEEN RETURNED. THE DEVICE WAS CONNECTED TO THE GENERATOR AND IT WAS RECOGNIZED. BECAUSE THE JAW WAS DETACHED NOT ALL FUNCTIONAL TESTING COULD BE PERFORMED WITH THE GENERATOR. A PROBABLE CAUSE OF THE DAMAGE TO THE JAW COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. THE IFU WARNS "PRIOR TO INSERTING OR REMOVING THE DEVICE FROM THE TROCAR, ENSURE THAT THE DEVICE IS IN THE STRAIGHT, NON-ARTICULATED POSITION BY CONFIRMING THAT THE INDICATOR ARROW ON THE ARTICULATION WHEEL IS ALIGNED WITH THE TOP OF THE INSTRUMENT." THE IFU ALSO WARNS THE USER TO DISCONTINUE USE IF BLACK HEAT SHRINK COVERING IS DAMAGED. THE BATCH HISTORY RECORDS WERE REVIEWED AND CERTIFIED BY EXTERNAL MANUFACTURING THAT THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THE EQUIPMENT. THE CERTIFICATE RECORDS ARE ACCESSIBLE THROUGH EXTERNAL MANUFACTURING. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03281 | 2015-06-01 | 20150312 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | IT WAS REPORTED THAT AT AN UNKNOWN TIME DURING AN UNKNOWN PROCEDURE, THE OPENING OF THE TERMINAL BRANCHES OF THE INSTRUMENT IS DEFECTIVE. THE TIPS DO NOT OPEN. IT IS UNKNOWN WHAT WAS USED TO COMPLETE THE PROCEDURE. THE PROCEDURE WAS PROLONGED BY UNKNOWN MINUTES. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03283 | 2015-06-01 | 20150515 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | IT WAS REPORTED THAT DURING A TOTAL LAPAROSCOPIC HYSTERECTOMY PROCEDURE, THE TOP PART OF THE JAW DISLODGED FROM THE DEVICE AND FELL INTO THE PATIENT. PIECE WAS RETRIEVED FROM THE PATIENT. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT. |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00102 | 2015-05-29 | 20150501 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLCEMENT SYSTEM | CONFORMIS |  | M5725INT060XE20 |  | IT WAS REPORTED THAT THE PT HAS PCL INSTABILITY/RUPTURE AND REQUIRES ADDITIONAL STABILITY. REVISION SURGERY TO AN OFF-THE-SHELF KNEE IMPLANT SYSTEM IS PLANNED. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00098 | 2015-05-29 | 20150401 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M5725INT0600020 |  | KNEE INSTABILITY SWAS REPORTED FOLLOWING A TRAUMATIC INCIDENT. REVISION SURGERY IS PLANNED TO EXCHANGE THE POLY INSERTS. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00099 | 2015-05-29 | 20150401 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57250600010 |  | TIBIAL LOOSENING WAS REPORTED FOR PATIENT WITH A TOTAL KNEE IMPLANT. REVISION SURGERY OCCURRED TO EXCHANGE THE TIBIAL IMPLANTS. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00101 | 2015-05-29 | 20150501 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M5725INTO600020 |  | ARTHROFIBROSIS AND LIMITED RANGE OF MOTION WERE REPORTED FOR PATIENT WITH A TOTAL KNEE IMPLANT. SURGERY IS PLANNED TO ADDRESS THE ARTHROFIBROSIS AND EXCHANGE THE POLY INSERTS. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00097 | 2015-05-29 | 20150401 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57250600020 |  | PAIN AND INSTABILITY WAS REPORTED FOR PATIENT WITH A TOTAL KNEE IMPLANT. X-RAY EXAMINATION REVEALED THAT THE TIBIAL IMPLANT IS IN VALGUS. REVISION SURGERY IS PLANNED TO EXCHANGE THE TIBIAL IMPLANTS. |  | Y | N | HEALTH PROFESSIONAL
4817672 | 2015-05-27 | 20150513 | No answer provided | User Facility report | BARD G2 | FILTER, INTRAVASCULAR, CARDIOVASCULAR | BARD PERIPHERAL VASCULAR, INC. | * | * |  | PATIENT PRESENTED FOR IVC FILTER REMOVAL. THE DEVICE WAS PLACED AT AN OUTSIDE FACILITY. DUE TO THE PLACEMENT OCCURRING AT ANOTHER FACILITY, THE FILTER'S MODEL, LOT, AND SERIAL NUMBERS ARE NOT AVAILABLE. PRIOR TO REMOVAL ATTEMPT, SPOT FILM IMAGING OF THE FILTER SHOWED IT TO BE TILTED TO THE PATIENT'S RIGHT AND THAT THERE WAS A FRACTURED ARM. FLUOROSCOPIC EXAMINATION OF THE CHEST DEMONSTRATED A BROKEN ARM IN THE INTRACARDIAC POSITION. AN ATTEMPT MADE TO RETRIEVE FRACTURED ARM WAS UNSUCCESSFUL. THE PATIENT WAS NOTIFIED OF THE RETAINED PIECE. PATIENT TO FOLLOW-UP WITH CARDIOTHORACIC SURGERY REGARDING RETAINED PIECE. |  | N | * | INVALID DATA
3005075853-2015-03395 | 2015-06-08 | 20150501 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | (B)(4). WHEN ADDITIONAL INFORMATION IS RECEIVED AND/OR THE DEVICE ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL MEDWATCH WILL BE SENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03402 | 2015-06-08 | 20150516 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | (B)(4). NO DEVICE RECEIVED FOR ANALYSIS AT TIME OF SUBMISSION OF 3500A. WHEN ADDITIONAL INFORMATION IS RECEIVED AND/OR THE DEVICE ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS THE BROKEN JAW RETRIEVED FROM THE PATIENT? YES WAS THERE ANY CHANGED TO THE PATIENTS ROUTINE POST-OP CARE? NO IS THE BROKEN JAW BEING RETURNED WITH THE DEVICE? UNSURE, OR, WAS THE BROKEN JAW DISCARDED? UNSURE, WILL HAVE DEVICE IN MY POSSESSION IN A FEW DAYS. |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00105 | 2015-06-04 | 20150501 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M5725INT0600010 |  | REVISION SURGERY IS PLANNED TO EXCHANGE THE POLY INSERTS. REASON FOR REVISION IS UNK. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00108 | 2015-06-04 | 20150501 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M5725INT0600020 |  | REVISION SURGERY IS PLANNED TO EXCHANGE THE POLY INSERTS. REASON FOR REVISION IS UNK. |  | Y | N | HEALTH PROFESSIONAL
2020394-2015-00869 | 2015-06-04 | 20150504 | Injury | Manufacturer report | G2 FILTER SYSTEM |  | BARD PERIPHERAL VASCULAR, INC. |  | UNK G2 |  | IT WAS REPORTED THAT APPROXIMATELY SEVEN YEARS POST VENA CAVA FILTER DEPLOYMENT THE PT PRESENTED WITH A PERICARDIAL EFFUSION. DURING THE EVALUATION A DETACHED FILTER LIMB WAS DISCOVERED IN THE RIGHT VENTRICLE OF THE HEART. SURGERY WAS PERFORMED TO SUCCESSFULLY REMOVE THE DETACHED LIMB AND DRAIN THE PERICARDIUM WITHOUT INCIDENT. THE PT IS REPORTED TO BE IN STABLE CONDITION. |  | Y | Y | HEALTH PROFESSIONAL
3004153240-2015-00107 | 2015-06-04 | 20150501 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M5725INT0600010 |  | PT DEVELOPED AN INFECTION. REVISION SURGERY IS PLANNED TO EXCHANGE TO POLY INSERTS. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00104 | 2015-06-04 | 20150501 | Malfunction | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600010 |  | THE PRODUCT KIT WAS PROVIDED WITHOUT THE DISPOSABLE 5 DEGREE TIBIAL CUTTING GUIDE IJIG. TWO 10 DEGREE TIBIAL CUTTING GUIDE IJIGS WERE PROVIDED. THE SURGEON PERFORMED THE TIBIAL SLOPE CUT WITHOUT USE OF A CUTTING GUIDE. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03427 | 2015-06-09 | 20150511 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35 |  | (B)(4). ADDITIONAL INFORMATION: THE DEVICE WAS RECEIVED WITH THE CABLE CUT OFF. DUE TO THE RETURNED CONDITION OF THE DEVICE, NO FUNCTIONAL TESTING COULD BE PERFORMED WITH THE GENERATOR. HOWEVER, THE DEVICE WAS MECHANICALLY TESTED AND NO ANOMALIES WERE NOTED WITH THE OPENING AND CLOSING OF THE DEVICE. IT SHOULD BE NOTED THAT AS PART OF OUR QUALITY PROCESS, EACH DEVICE IS VISUALLY INSPECTED AND FUNCTIONALLY TESTED DURING MANUFACTURING TO ENSURE THE DEVICE MEETS THE REQUIRED SPECIFICATIONS PRIOR TO SHIPMENT. THE BATCH HISTORY RECORDS WERE REVIEWED WITH NO ANOMALIES NOTED DURING THE MANUFACTURING PROCESS. |  | N | Y | HEALTH PROFESSIONAL
2954740-2015-00129 | 2015-06-11 | 20150519 | Malfunction | Manufacturer report | ORBIT GALAXY G2 MICROCOIL DELIVERY SYSTEM | ARTIFICIAL EMBOLIZATION DEVICE | CODMAN AND SHURTLEFF, INC | N/A | 641CF0620 |  | THE CONTACT AT THE HOSPITAL REPORTED THAT THE DOCTOR TRIED TO FILL AN ANEURYSM AT THE ANTERIOR COMMUNICATING ARTERY WITH AN ORBIT GALAXY G2 (641CF0620/C26273) HOWEVER HAD TROUBLE POSITIONING IT AND FOUND THAT THE COIL MAY HAVE BEEN DETACHED WHEN TRYING TO REPOSITION. ADDITIONAL INTERVENTION WAS ATTEMPTED TO RETRIEVE THE COIL. THE ANEURYSM WAS SACCULAR, 11 MM WIDE AND 16 MM LONG WITH A NECK WIDTH OF 4 MM. AN ENVOY DA (67125695D/LOT NOT PROVIDED) AND A HEADWAY 17 MICROCATHETER (MC172150S/LOT NOT PROVIDED) WERE USED IN THE PROCEDURE. HE TRIED TO BEST POSITION IT, BUT IT WAS PROTRUDING TO THE VESSEL. WHEN HE TRIED TO PULL BACK THE WIRE WITH COIL FOR THE THIRD TIME HE NOTICED THAT THERE WAS NO MOVEMENT OF THE COIL CORRESPONDING TO PULLING BACK ON THE PROXIMAL PART OF THE DPU WIRE. DESPITE NOT FEELING ANY FRICTION, TENSION OR ANY KIND OF OBSTACLE, HE REALIZED THE COIL MUST HAVE BEEN DETACHED FROM THE WIRE. HE REPORTEDLY DID NOT EVEN TOUCH THE DETACHMENT CONTROL BOX SO HE WAS SURE HE DID NOT DETACH IT HIMSELF. THE DOCTOR HAD TO STOP THE EMBOLIZATION AND TRIED TO RETRIEVE THE COIL WITH A LASSO DEVICE BUT WAS NOT SUCCESSFUL. PART OF THE COIL REMAINED IN THE ANEURYSM AND PART OF IT PROTRUDED INTO THE VESSEL. THE IMAGE SEEN DURING THE PROCEDURE SEEMS TO SHOW A SHADOW, WHICH LOOKED LIKE A KIND OF FABRIC OR GENTLE PLASTIC FIBER CONNECTED TO THE COIL. THE PRODUCT WAS STORED ACCORDING TO LABELING INSTRUCTIONS. THE PRODUCT WAS AVAILABLE FOR RETURN. THERE WAS NO SIGNIFICANT DELAY TO THE PROCEDURE AS A RESULT OF THE ISSUE. THERE WAS NO FLOW RESTRICTION AS A RESULT OF THE ISSUE. THE BLEEDING FROM THE TARGET ANEURYSM STOPPED AFTER THE PROCEDURE. THERE WAS NO ADDITIONAL ADVERSE OUTCOME TO THE PATIENT. TWO OTHER ANEURYSMS WERE LEFT UNTREATED ALTHOUGH UNRELATED TO THE PRODUCT ISSUE. NO NECK REMODELING WAS UTILIZED. AN ADEQUATE CONTINUOUS FLUSH WAS MAINTAINED THROUGH THE MICROCATHETER AT ALL TIMES. THERE WAS NO RESISTANCE BETWEEN THE GUIDEWIRE AND MICROCATHETER WHEN ACCESSING THE TARGET SITE. THERE WAS NO RESISTANCE AT ANY TIME DURING ADVANCEMENT OF THE COIL THROUGH THE MICROCATHETER. PRIOR TO THIS COIL, 7 COILS (DETAILS UNKNOWN) WERE PLACED IN THE ANEURYSM. THERE WERE NO KINKS IN THE MICROCATHETER THAT MAY HAVE CONTRIBUTED TO THE EVENT. THERE WAS NO RESISTANCE WHEN THE COIL WAS ADVANCED/REPOSITIONED/WITHDRAWN. THE MICROCATHETER WAS NOT REPOSITIONED OVER THE COIL WHILE THE COIL WAS DEPLOYED OR PARTIALLY DEPLOYED OUT OF THE DISTAL END OF THE MICROCATHETER. A ONE TO ONE RELATIONSHIP BETWEEN THE COIL AND DELIVERY TUBE WAS VERIFIED WITH FLUORO PRIOR TO REPOSITIONING. |  | Y | Y | HEALTH PROFESSIONAL
2020394-2015-00881 | 2015-06-10 | 20141121 | Injury | Manufacturer report | G2 FILTER SYSTEM - FEMORAL | VENA CAVA FILTER | BARD PERIPHERAL VASCULAR, INC. |  | RF310F |  | A LOT HISTORY REVIEW COULD NOT BE PERFORMED AS THE LOT NUMBER WAS NOT PROVIDED. UNABLE TO PERFORM A SAMPLE EVALUATION AS THE DEVICE WAS NOT RETURNED AND IMAGES WERE NOT PROVIDED. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. |  | Y | Y | HEALTH PROFESSIONAL
3004153240-2015-00109 | 2015-06-10 | 20150422 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57250600230 |  | REVISION SURGERY OCCURRED FOR PATIENT WITH A TOTAL KNEE IMPLANT. REASON FOR REVISION IS UNKNOWN. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE IS MANUFACTURED TO SPECIFICATION. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00110 | 2015-06-10 | 20140101 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57250600020 |  | REVISION SURGERY OCCURRED FOR PATIENT WITH A TOTAL KNEE IMPLANT. REASON FOR REVISION IS UNKNOWN. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00113 | 2015-06-10 |  | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57250600010 |  | INFECTION WAS REPORTED FOR PATIENT WITH A TOTAL KNEE IMPLANT. REVISION SURGERY IS PLANNED TO EXCHANGE THE POLY INSERTS. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00111 | 2015-06-10 | 20150506 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57250600020 |  | REVISION SURGERY OCCURRED FOR PATIENT WITH A TOTAL KNEE IMPLANT. REASON FOR REVISION IS UNKNOWN. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00112 | 2015-06-10 |  | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57250600010 |  | PATIENT REPORTED ANTERIOR KNEE PAIN FOLLOWING TOTAL KNEE REPLACEMENT SURGERY. REVISION SURGERY IS PLANNED. |  | Y | N | HEALTH PROFESSIONAL
3005075853-2015-03568 | 2015-06-15 | 20150527 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C45 |  | (B)(4). BATCH # M90H9Z. THE DEVICE WAS RECEIVED WITH THE TOP JAW OF THE DEVICE ATTACHED (NOT BROKEN OFF AS REPORTED) BUT THE TOP OF THE I-BLADE UPPER TIP WAS BROKEN OFF AND NOT RETURNED. UPON VISUAL INSPECTION TO THE DEVICE NO ANOMALIES WERE OBSERVED WITH THE JAW OF THE DEVICE AS REPORTED IN THE EVENT DESCRIPTION THE DEVICE WAS CONNECTED TO THE GENERATOR AND IT WAS RECOGNIZED. BECAUSE THE I BLADE WAS DAMAGED NOT ALL FUNCTIONAL TESTING COULD BE PERFORMED WITH THE GENERATOR. THE CONDITION OF THE BLADE PREVENTED THE FUNCTIONALITY OF THE JAW. THE JAW WAS UNABLE TO CYCLE OPEN AND CLOSE. A PROBABLE CAUSE OF THE DAMAGE TO THE I BLADE COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. THE BATCH HISTORY RECORDS WERE REVIEWED WITH NO ANOMALIES NOTED DURING THE MANUFACTURING PROCESS. |  | N | Y | HEALTH PROFESSIONAL
1818910-2015-23658 | 2015-06-16 | 20150613 | Injury | Manufacturer report | G2 PRESS FIT STEM SZ6L | HIP FEMORAL STEM/SLEEVE | DEPUY (IRELAND) 9616671 |  | 900000126 |  | ADDITIONAL NARRATIVE: THIS COMPLAINT IS STILL UNDER INVESTIGATION. DEPUY WILL NOTIFY THE FDA OF THE RESULTS OF THIS INVESTIGATION ONCE IT HAS BEEN COMPLETED. (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL MEDWATCH, A FOLLOW-UP MEDWATCH WILL BE FILED AS APPROPRIATE. |  | Y | N | HEALTH PROFESSIONAL
3005075853-2015-03624 | 2015-06-16 | 20150430 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35A |  | IT WAS REPORTED THAT THE JAWS WOULD NOT CLOSE PROPERLY DURING A TOTAL LAP HYSTERECTOMY. CUSTOMER USED A SECOND DEVICE TO COMPLETE THE PROCEDURE. NO PATIENT CONSEQUENCES. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03623 | 2015-06-16 | 20150508 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C25 |  | IT WAS REPORTED THAT DURING A LAP LOBECTOMY, THE JAWS BECAME NOT TO OPEN AND CLOSE DURING USE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03625 | 2015-06-16 | 20150513 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35A |  | IT WAS REPORTED THAT DURING AN LAVH PROCEDURE, THE DEVICE GAVE A REPLACE INSTRUMENT ERROR AS SOON AS IT WERE PLUGGED IN. THIS OCCURRED ON TWO DIFFERENT GENERATORS. NO ADDITIONAL INFORMATION WAS AVAILABLE. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03642 | 2015-06-17 |  | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35A |  | IT WAS REPORTED THE INSTRUMENT WAS BEING USED IN AN UNKNOWN PROCEDURE AND STOPPED ACTIVATING DURING THE PROCEDURE. IT WAS UNKNOWN HOW THE PROCEDURE WAS COMPLETED BUT THERE WAS NO CONSEQUENCE TO THE PATIENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03643 | 2015-06-17 | 20150506 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35A |  | IT WAS REPORTED THAT DURING A TOTAL LAPAROSCOPIC HYSTERECTOMY (TLH) PROCEDURE, THE SURGEON COMPLAINED THAT THE JAWS WOULD NOT CLOSE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT. |  | N | Y | HEALTH PROFESSIONAL
2954740-2015-00134 | 2015-06-17 | 20150519 | Malfunction | Manufacturer report | ORBIT GALAXY G2 MICROCOIL DELIVERY SYSTEM | ARTIFICIAL EMBOLIZATION DEVICE | CODMAN AND SHURTLEFF, INC |  | 641CF0412 |  | THE PRODUCT WAS RECEIVED FOR ANALYSIS. |  | N | Y | HEALTH PROFESSIONAL
2954740-2015-00135 | 2015-06-17 | 20150519 | Malfunction | Manufacturer report | ORBIT GALAXY G2 MICROCOIL DELIVERY SYSTEM | ARTIFICIAL EMBOLIZATION DEVICE | CODMAN AND SHURTLEFF, INC |  | 641CF0510 |  | THE PRODUCT IS AVAILABLE FOR EVALUATION AND TESTING, HOWEVER THE ANALYSIS HAS NOT BEEN COMPLETED. ADDITIONAL INFORMATION WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. (B)(4). |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03714 | 2015-06-18 | 20150603 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35 |  | IT WAS REPORTED THAT DURING A TOTAL LAPAROSCOPIC HYSTERECTOMY (TLH) PROCEDURE, THE TOP JAW FELL OFF WHILE MOBILIZING THE UTERUS. THE DEVICE WAS DIFFICULT TO CLOSE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03764 | 2015-06-19 | 20150423 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE DID NOT FUNCTION CORRECTLY. THE CASE WAS COMPLETED BY USING ANOTHER INSTRUMENT. CUSTOMER CANNOT PROVIDE ANYMORE DETAILS ABOUT THE CIRCUMSTANCES IN WHICH THE ISSUE OCCURRED. NO PATIENT CONSEQUENCE WAS REPORTED . |  | N | Y | HEALTH PROFESSIONAL
2954740-2015-00138 | 2015-06-19 | 20150528 | Malfunction | Manufacturer report | ORBIT GALAXY G2 MICROCOIL DELIVERY SYSTEM | ARTIFICIAL EMBOLIZATION DEVICE | CODMAN AND SHURTLEFF, INC |  | 641CX3505 |  | IT WAS REPORTED BY THE REP THAT DURING AN ANEURYSM COILING PROCEDURE, WE PUT THE G2 COMPLEX XTRASOFT COIL 3.5X5 (641CX3505/C14765) INSIDE THE MICROCATHETER (DETAILS UNKNOWN) BUT IT STARTED TO HAVE SOME RESISTANCE CLOSE TO THE DISTAL PORTION OF THE MICROCATHETER. WE DECIDE TO REMOVE IT AND IT CAME STRETCHED. IT WOULD BE THE SECOND COIL SO THERE WERE NO ISSUES WITH THE MICROCATHETER AND WE PUT A MICROWIRE AGAIN TO SEE IF THERE WERE NO KINKS OR LOOP WITH IT AND IT NAVIGATED NORMALLY AS THE FOLLOWING COILS THAT WE USED UNTIL THE END. THE PROCEDURE WAS COMPLETED WITH A SAME/LIKE COIL (DETAILS UNKNOWN). THERE WAS NO ADVERSE CONSEQUENCE TO THE PATIENT. ONE DEVICE WILL BE RETURNED. THE PROCEDURE WAS DELAYED FOR 5 MINUTES AS A RESULT OF THE DIFFICULTIES ENCOUNTERED WITH THE DEVICE. THE PRODUCT WAS STORED ACCORDING TO LABELING INSTRUCTIONS. AN ADEQUATE CONTINUOUS FLUSH WAS MAINTAINED THROUGH THE MICROCATHETER. |  | N | Y | HEALTH PROFESSIONAL
2020394-2015-00894 | 2015-06-15 | 20060209 | Injury | Manufacturer report | G2 VENA CAVA FILTER- FEMORAL | VENA CAVA FILTER | BARD PERIPHERAL VASCULAR, INC. |  | RF310F |  | MEDICAL RECORD REVIEW: A VENA CAVA FILTER WAS DEPLOYED AT THE REQUEST OF THE PATIENT APPROXIMATELY THREE MONTHS PRIOR TO GASTRIC BYPASS SURGERY. LOCAL ANESTHESIA WAS ADMINISTERED AND THE RIGHT COMMON FEMORAL VEIN WAS ACCESSED. THE INFERIOR VENA CAVA WAS WITHOUT THROMBUS AND VARIANT ANATOMY. THE VENA CAVA FILTER WAS DEPLOYED SUCCESSFULLY IN AN INFRARENAL POSITION. THE PATIENT TOLERATED THE PROCEDURE WELL. OUTCOMES ATTRIBUTED TO ADVERSE EVENT: LIFE-THREATENING; REPORT SOURCE: OTHER-ATTORNEY; RESULTS: NO RESULTS AVAILABLE SINCE NO EVALUATION PERFORMED. THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. |  | Y | Y | HEALTH PROFESSIONAL
3005075853-2015-03755 | 2015-06-19 | 20150604 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35 |  | (B)(4). ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS THE DEVICE ON A VESSEL/ARTERY WHEN IT WOULD NOT OPEN? YES. IF YES, HOW WAS THE DEVICE REMOVED? (I.E. CUT OFF, FORCED OPENED) HE HAD TO PUT CLIPS ON BOTH SIDES OF THE DEVICE JAW AND CUT THE VESSEL WITH THE JAW CLOSED ON IT. IF CUT OFF, DID THIS CAUSE A CHANGE IN THE PLAN OF CARE FOR THE PATIENT? IT DID NOT CHANGE THE PLAN OF CARE. DID THE REMOVAL CAUSE ANY DAMAGE TO THE VESSEL/ARTERY? NO. IF YES, WHAT IS THE DAMAGE AND HOW WAS THE DAMAGE REPAIRED? NA. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03761 | 2015-06-19 | 20150604 | Injury | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35A |  | (B)(4). INFORMATION IS UNAVAILABLE; DEVICE NOT RETURNED. NO DEVICE RECEIVED FOR ANALYSIS AT TIME OF SUBMISSION OF 3500A. WHEN ADDITIONAL INFORMATION IS RECEIVED AND/OR THE DEVICE ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION REQUESTED: WHAT IS THE RATIONALE FOR THE CONVERT TO OPEN? TO RETRIEVE THE BLACK SHAFT COVER? OR WAS THERE ANOTHER REASON FOR THE CONVERT TO OPEN, SUCH AS: PATIENT¿S ANATOMY? PHYSICIAN PREFERENCE? WAS THE BLACK COATING ON THE SHAFT OR WAS IT THE PTC IN THE UPPER JAW OF DEVICE? WHERE REUSABLE TROCARS BEING USED? WAS THE DEVICE ARTICULATED WHEN WITHDRAWN FROM THE TROCAR? THE BLACK COATING IS NOT INSULATION AS THE COATING WOULD BE ON A BI-POLAR TYPE DEVICE. THE COATING IS A COVER FOR THE SHAFT ONLY. |  | Y | N | HEALTH PROFESSIONAL
3005075853-2015-03793 | 2015-06-22 | 20150414 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | IT WAS REPORTED THAT THE DEVICE STOPPED WORKING DURING A COLON PROCEDURE. THE PROCEDURE WAS COMPLETED WITH A SECOND SAME DEVICE. NO PATIENT CONSEQUENCE. |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00115 | 2015-06-18 | 20150501 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600010 |  | PT IS HAVING RECURRENT EFFUSIONS. THE SURGEON BELIEVES IT IS DUE TO POLY INSERT WEAR. REVISION SURGERY IS PLANNED TO EXCHANGE THE POLY INSERTS. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00114 | 2015-06-18 | 20150501 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600020 |  | PT HAS JOINT LAXITY THAT WAS NOT APPARENT AT INITIAL IMPLANTATION. REVISION SURGERY OCCURRED TO EXCHANGE THE POLY INSERTS. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00117 | 2015-06-18 | 20150501 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M5725INT0600020 |  | MID-FLEXION INSTABILITY WAS REPORTED FOR PT WITH A TOTAL KNEE IMPLANT. REVISION SURGERY OCCURRED TO EXCHANGE THE POLY INSERTS. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00118 | 2015-06-18 | 20150501 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600010 |  | TIBIAL LOOSENING OCCURRED FOLLOWING A PT FALL. REVISION SURGERY IS PLANNED TO REPLACE THE TIBIAL IMPLANTS. |  | Y | N | HEALTH PROFESSIONAL
2954740-2015-00147 | 2015-06-25 | 20150603 | Malfunction | Manufacturer report | ORBIT GALAXY G2 MICROCOIL DELIVERY SYSTEM | ARTIFICIAL EMBOLIZATION DEVICE | CODMAN AND SHURTLEFF, INC |  | 641CX0306 |  | THE G2 COMPLEX (641CX0306, LOT C13527) WILL NOT BE RETURNED, THEREFORE, THE ROOT CAUSE OF THE RESHEATHING DIFFICULTY AND THE COIL DAMAGE CANNOT BE DETERMINED. A REVIEW OF THE MANUFACTURING DOCUMENTATION ASSOCIATED WITH THIS LOT PRESENTED NO ISSUES DURING THE MANUFACTURING OR INSPECTION PROCESS THAT CAN BE RELATED TO THE REPORTED COMPLAINT. THE PRODUCT WAS NOT RECEIVED FOR ANALYSIS, AND WITHOUT THE PRODUCT, THE REPORTED EVENTS COULD NOT BE CONFIRMED. REVIEW OF THE LOTS REVEALED NO ANOMALIES DURING THE MANUFACTURING AND INSPECTION PROCESSES THAT CAN BE ASSOCIATED WITH THE REPORTED COMPLAINTS. THEREFORE, NO CORRECTIVE ACTIONS WILL BE TAKEN AT THIS TIME. (B)(4). |  | N | Y | HEALTH PROFESSIONAL
2020394-2015-00958 | 2015-06-22 | 20120602 | Injury | Manufacturer report | G2 FILTER SYSTEM - FEMORAL | VENA CAVA FILTER | BARD PERIPHERAL VASCULAR, INC. |  | RF310F |  | IT WAS REPORTED THAT APPROXIMATELY THREE YEARS POST VENA CAVA FILTER DEPLOYMENT REPORTEDLY THE FILTER WAS TILTED WITH TOW FILTER LIMBS PERFORATING THE IVC WALL. A DETACHED LIMB WAS REPORTED IN THE HEPATIC VEIN. AN ATTEMPT TO CAPTURE AND RETRIEVE THE FILTER AND DETACHED LIMB WAS UNSUCCESSFUL. THE PATIENT STATUS AT THIS TIME IS UNKNOWN. |  | Y | Y | HEALTH PROFESSIONAL
3005075853-2015-03903 | 2015-06-29 | 20150609 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | IT WAS REPORTED THAT DURING AN UNKNOWN LAP PROCEDURE, THE JAWS BECAME NOT TO OPEN IN CLAMPING THE TARGET TISSUE. THE TISSUE AROUND THE JAW WAS LIGATED AND THEN, THE DEVICE WAS REMOVED FROM THE PATIENT BY CUTTING THE TISSUE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO DEVICE WILL BE RETURNING. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03906 | 2015-06-29 | 20150610 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | IT WAS REPORTED THAT DURING A TOTAL LAPAROSCOPIC HYSTERECTOMY PROCEDURE; THE RIGHT SIDE OF THE UTERUS WAS MOBILIZED WITH NO ISSUES. THE LEFT SIDE WAS MOBILIZED BY THE ASSISTING DOCTOR WHO NOTED THE TISSUE SEEMED THICKER ON HIS SIDE. HE WAS MOBILIZING CLOSE TO THE CERVIX AND HEARD A CRUNCH. WHILE REMOVING THE INSTRUMENTS, PART OF THE JAW REMAINED ON THE TISSUE AND ONE OF THE WIRES STILL ATTACHED TO THE INSTRUMENT WAS EXPOSED. THE DOCTOR USED A GRASPER TO GET THE BROKEN JAW AND PULLED THE PIECE AND THE TROCAR OUT. THEY CONTINUED THE PROCEDURE WITH A DIFFERENT DEVICE. THE SCRUB TECH THEN NOTED THAT A PIECE OF THE ORIGINAL BROKEN JAW WAS STILL MISSING. THEY SEARCHED THE ABDOMEN, BUT COULD NOT FIND IT. FLUORO WAS USED TO SEARCH FOR THE MISSING PIECE. THEY THOUGHT THEY FOUND IT NEAR THE LIVER, BUT COULD NOT RETRIEVE IT WITH A LAP INSTRUMENT. A HAND PORT WAS THEN PLACED IN THE RIGHT UPPER QUADRANT AND THE SURGEON PLACED HIS HAND THROUGH AND DETERMINED THE ITEM THEY SAW WAS AN ENCAPSULATED CLIP FROM THE PATIENT¿S PREVIOUS LAP GASTRIC BYPASS. THEY SEARCHED AROUND THE ABDOMEN ONE MORE TIME AND SAW THE PIECE OF THE JAW WAS WEDGED BETWEEN THE TROCAR AND THE FASCIA. THE PIECE WAS RETRIEVED. THE PROCEDURE WAS COMPLETED USING A DIFFERENT DEVICE. THE PROCEDURE WAS PROLONGED BY 90-100 MINUTES. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03896 | 2015-06-29 | 20150514 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | (B)(4). INFORMATION WAS NOT PROVIDED BY THE CONTACT. SHOULD THE INFORMATION BE PROVIDED LATER, A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE DEVICE WAS RECEIVED THE UPPER JAW DAMAGE AT THE PROXIMAL SIDE. THE DEVICE WAS CONNECTED TO THE GENERATOR AND IT WAS RECOGNIZED. BECAUSE THE JAW WAS DAMAGED NOT ALL FUNCTIONAL TESTING COULD BE PERFORMED WITH THE GENERATOR. THE CONDITION OF THE JAWS PREVENTED THE FUNCTIONALITY OF THE JAW. THE DEVICE WAS UNABLE TO CYCLE OPEN AND CLOSE. A PROBABLE CAUSE OF THE DAMAGE TO THE JAW COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. THE BATCH HISTORY RECORDS WERE REVIEWED AND CERTIFIED BY EXTERNAL MANUFACTURING THAT THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THE EQUIPMENT. THE CERTIFICATE RECORDS ARE ACCESSIBLE THROUGH EXTERNAL MANUFACTURING. |  | N | Y | HEALTH PROFESSIONAL
2954740-2015-00155 | 2015-06-30 | 20150616 | Malfunction | Manufacturer report | ORBIT GALAXY G2 MICROCOIL DELIVERY SYSTEM | ARTIFICIAL EMBOLIZATION DEVICE | CODMAN AND SHURTLEFF, INC |  | 641CF0410 |  | DURING TREATMENT OF NORMAL ANATOMY WITHOUT EXCESSIVE TORTUOSITY, THE FIRST COIL, STRYCKER 360 SOFT 6X20, IS DEPLOYED WITHOUT DIFFICULTY. HOWEVER, THE SECOND ORBIT GALAXY G2 MICROCOIL DELIVERY SYSTEM COIL WAS PLACED IN THE ANEURYSMAL CAVITY WITHOUT RESISTANCE, BUT AS THE PHYSICIAN WAS NOT SATISFIED WITH THE PLACEMENT, AND WANTED TO REMOVE IT FOR REPLACING, BUT THE COIL WAS LOCKED AT ONCE. IT WAS IMPOSSIBLE TO GET IT BACK. THE COIL STRETCHED WITHOUT APPARENT REASON. FORTUNATELY, AS THE COIL WAS IN THE ANEURYSMAL CAVITY, THE DOCTOR WAS ABLE TO DEPLOYED IT WITHOUT CONSEQUENCE. THEN, THE PROCEDURE WAS COMPLETED SUCCESSFULLY WITH OTHER COILS STRYCKER WITHOUT ANY DIFFICULTY. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03939 | 2015-07-01 | 20150518 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C25 |  | (B)(4). UPON VISUAL INSPECTION UNDER MAGNIFICATION IT WAS NOTICED THAT THE UPPER JAW WAS SLIGHTLY DAMAGED AT THE RETENTION PIN INTERPHASE. RESULTING IN THE JAW BEEN LOOSE AND DIFFICULTY IN OPENING AND CLOSING OF THE JAWS. HOWEVER THE DAMAGE WAS NOT SIGNIFICANT ENOUGH TO PREVENT THE DEVICE FROM CYCLING. THE DEVICES WAS TESTED ON THE GENERATOR AND PASSED ALL FUNCTIONAL TESTING. THE ENERGY OUTPUT DELIVERED FROM THE DEVICE WAS VERIFIED. ALL THREE TONES WERE HEARD DURING FUNCTIONAL TESTING (TONE 1 IS HEARD WHEN THE ENERGY ACTIVATION BUTTON IS PRESSED; TONE 2 IS HEARD WHEN TISSUE IMPEDANCE THRESHOLD IS REACHED; AND TONE 3 IS HEARD WHEN THE CYCLE IS COMPLETE). A PROBABLE CAUSE OF THE DAMAGE TO THE JAW COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. IT IS POSSIBLE THAT DUE TO THE DAMAGE AT THE JAW IT COULD BE DISPLAYED A "REPOSITION JAWS AND REACTIVATED" ALERT SCREEN. HOWEVER, THIS WAS NOT SEEN DURING THE ANALYSIS. THE BATCH RECORD WAS REVIEWED AND NO ANOMALIES WERE NOTED DURING THE MANUFACTURING PROCESS. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03944 | 2015-07-01 | 20150508 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | IT WAS REPORTED THAT THE TOP JAW OF THE DEVICE BROKE DURING A TOTAL LAP HYSTERECTOMY. NO PIECES FELL IN TO THE PATIENT. THE PROCEDURE WAS COMPLETED WITH A HARMONIC DEVICE. NO PATIENT CONSEQUENCE. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03943 | 2015-07-01 | 20150521 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C14 |  | IT WAS REPORTED THAT DURING AN OPEN LOBECTOMY, THE JAWS BECAME NOT TO CLOSE IN THE BEGINNING OF THE OPERATION. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03945 | 2015-07-01 | 20150508 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | (B)(4). THE DEVICE WAS RECEIVED THE JAWS CLOSED AND THE UPPER JAW DAMAGED AT THE PROXIMAL SIDE. THE DEVICE WAS CONNECTED TO THE GENERATOR AND IT WAS RECOGNIZED. BECAUSE THE JAW WAS DAMAGED NOT ALL FUNCTIONAL TESTING COULD BE PERFORMED WITH THE GENERATOR. THE CONDITION OF THE JAWS PREVENTED THE FUNCTIONALITY OF THE JAW. THE DEVICE WAS UNABLE TO CYCLE OPEN AND CLOSE. A PROBABLE CAUSE OF THE DAMAGE TO THE JAW COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. THE BATCH HISTORY RECORDS WERE REVIEWED AND CERTIFIED BY EXTERNAL MANUFACTURING THAT THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THE EQUIPMENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03954 | 2015-07-01 |  | Injury | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35A |  | (B)(4). INFORMATION NOT AVAILABLE, DEVICE NOT RETURNED FOR ANALYSIS. SHOULD THE INFORMATION BE PROVIDED LATER, A SUPPLEMENTAL MEDWATCH WILL BE SENT ADDITIONAL INFORMATION REQUESTED: HOW LONG AFTER THE ORIGINAL PROCEDURE DID THE BLEEDING OCCUR (PROVIDED NUMBER OF HOURS, DAYS, ETC.) WHERE WAS THE BLEEDING FROM? (STATE WHERE) HOW MUCH BLOOD WAS LOST? (CC¿S) WAS THERE A SECOND PROCEDURE TO CONTROL THE BLEEDING OR DID THE PATIENT JUST RECEIVE A TRANSFUSION? IF YES, THERE WAS A SECOND PROCEDURE: WAS THE BLEEDING FROM AN AREA WHERE THE ENSEAL DEVICE WAS USED? HOW WAS THE BLEEDING CONTROLLED IN THE 2ND PROCEDURE? WHAT WAS THE SIZE OF THE VESSEL OR ARTERY? IF NO, THERE WAS NOT A SECOND PROCEDURE: HOW IS THE SURGEON ATTRIBUTING THE BLEEDING TO THE USE OF THE ENSEAL DEVICE? DURING THE LAP. HYSTERECTOMY, WHAT OTHER DEVICES WERE USED IN THE CASE? HOW LONG HAS THE SURGEON BEEN USING THE NSLG2C35A? WERE THERE ANY HEMO ISSUES IN FIRST PROCEDURE? IF YES, HOW WAS THIS ADDRESSED? WHAT ARE THE SURGEON¿S TYPICAL USAGE AS TO THE I-BLADE ADVANCEMENT AND HEARING TONE 2 DOES THE SURGEON BELIEVE THAT THE ENSEAL HAD ANYTHING TO DO WITH THE POST-OP BLEEDING? WAS THE PATIENT TAKING ANY ANTICOAGULANTS, SUCH AS ASPIRIN, ANTICOAGULANTS, OR ANTIPLATELET AGENTS)? IF YES, SPECIFY PRESCRIBED MEDICATION. HAS THE PATIENT UNDERGONE ANY RADIATION/CHEMOTHERAPY THERAPY? IF YES, PLEASE SPECIFY RADIATION, CHEMO, OR BOTH. DOES THE PATIENT HAS A KNOWN COAGULATION DISORDER? HAS THE PATIENT TAKEN ANY STEROIDS? WHAT WAS THE TOTAL BLOOD LOSS? WHAT WAS THE PATIENT¿S PRE-OP HEMOGLOBIN AND HEMATOCRIT? WHAT WAS THE PATIENT¿S POST OPERATIVE HEMOGLOBIN AND HEMATOCRIT? ADDITIONAL INFORMATION RECEIVED: THE PATIENT RECEIVED A TRANSFUSION AND SECOND PROCEDURE WAS NOT NECESSARY. I HAVE REPORTED ALL INFORMATION AND HAVE NO FURTHER INFORMATION. WE DID NOT RECEIVE A BATCH OR LOT NUMBER FOR THE PRODUCT INVOLVED IN THIS COMPLAINT. THEREFORE, WE WERE UNABLE TO CHECK MANUFACTURING RECORDS FOR ANY RELATED NON-CONFORMANCES. |  | Y | N | HEALTH PROFESSIONAL
3005075853-2015-03991 | 2015-07-02 |  | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | (B)(4). THE FOLLOWING INFORMATION WAS REQUESTED, BUT UNAVAILABLE: DID THE TOP JAW DETACHED COMPLETELY FROM THE DEVICE? IF YES DID ANY PIECES FALL INTO THE PATIENT? IS YES, WERE THEY RETRIEVED? IF RETRIEVED, WILL THE BROKEN PIECE BE RETURNING WITH THE DEVICE? WHAT TYPE OF PROCEDURE WAS THE DEVICE USED IN? WHAT WAS THE PROCEDURE DATE? WHAT IS THE LOT NUMBER FOR THE DEVICE? WHAT IS THE NAME OF THE PERSON WHO USED THE DEVICE? THE DEVICE WAS RETURNED WITH THE UPPER JAW DETACHED AND NOT RETURNED. IN ADDITION, THE TOP OF THE I-BLADE WAS FRACTURED AND SLIGHTLY LIFTED. THE DEVICE WAS CONNECTED TO THE GENERATOR AND IT WAS RECOGNIZED. BECAUSE THE JAW WAS DETACHED NOT ALL FUNCTIONAL TESTING COULD BE PERFORMED WITH THE GENERATOR. A PROBABLE CAUSE OF THE DAMAGE TO THE JAW AND THE I BLADE COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. THE BATCH HISTORY RECORDS WERE REVIEWED AND CERTIFIED BY EXTERNAL MANUFACTURING THAT THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THE EQUIPMENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03984 | 2015-07-02 | 20150129 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | (B)(4): CERAMIC. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: PLEASE CLARIFY THE PARTS THAT DETACHED FROM THE SHAFT OF THE DEVICE? NO INFORMATION THE DEVICE WAS RECEIVED WITH THE ELECTRODE, CERAMIC AND ACTIVE ROD SEPARATED FROM THE DEVICE AS ONE PIECE AND IT WAS RETURNED WITH THE DEVICE. IN ADDITION, THE CERAMIC WAS PARTIALLY MISSING. THE CERAMIC WAS DAMAGED BY THE SEPARATION OF THE CERAMIC FROM THE LOWER JAW. UPON VISUAL INSPECTION TO THE DEVICE, THE JAW WAS FULLY ATTACHED TO THE DEVICE AND NO ANOMALIES WERE FOUND WITH THE JAW. THE DEVICE WAS TESTED ON THE GENERATOR, THIS RESULTED IN THE ¿CLOSE JAWS ON TISSUE AND REACTIVATE¿ ALERT SCREEN, THIS IS ADVISING THE GENERATOR CANNOT DELIVER ENERGY. THEN THE "REPLACE INSTRUMENT" SCREEN WOULD BE DISPLAYED AFTER RECEIVING THE "CLOSE JAWS ON TISSUE AND REACTIVATE" MESSAGE TWICE. THERE IS INSUFFICIENT EVIDENCE RELATED TO WHAT CAUSED THE DAMAGE ON THE DEVICE. THE BATCH HISTORY RECORDS WERE REVIEWED WITH NO ANOMALIES NOTED DURING THE MANUFACTURING PROCESS. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-03985 | 2015-07-02 | 20150612 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | IT WAS REPORTED THAT DURING A VAGINAL REPAIR ANTERIOR AND POSTERIOR PROCEDURE; WHEN BEING USED DURING THE FINAL PART OF THE PROCEDURE, WHEN FIRED IT WILL NOT OPEN AND COMPLETELY ARTICULATED. THE PROCEDURE WAS COMPLETED USING SAME THIS DEVICE. NO ADVERSE PATIENT CONSEQUENCES WERE REPORTED. |  | N | Y | HEALTH PROFESSIONAL
4890957 | 2015-06-24 | 20150624 | Malfunction | User Facility report | ENSEAL G2 | ELECTROSURGICAL, CUTTING AND COAGULATION | ETHICON ENDO-SURGERY, INC | * | NSLG2C35A |  | COATING CAME OFF TIP OF DEVICE, REMOVED FROM PATIENT.======================MANUFACTURER RESPONSE FOR ENSEAL, (BRAND NOT PROVIDED) (PER SITE REPORTER).======================THIS IS LIKELY NOT A DEFECTIVE PRODUCT, IT WAS PROBABLY NOT STRAIGHTENED OUT PRIOR TO REMOVING FROM THE TROCAR. WE WILL SEND A KIT FOR RETURN. |  | N | Y | PHYSICIAN
3005075853-2015-04038 | 2015-07-06 | 20150511 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35 |  | IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE BROKE OFF THE I-BLADE AND IT FELL INSIDE THE PATIENT. THE FOREIGN BODY WAS RETRIEVED BEFORE COMPLETING THE CASE. NO PATIENT CONSEQUENCES. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04039 | 2015-07-06 | 20150505 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | IT WAS REPORTED THAT DURING AN OPEN THORACOTOMY PROCEDURE, THE DEVICE STOPPED WORKING UNEXPECTEDLY. THEY DID NOT REMEMBER ANY ERROR CODE. THE SURGEON CONTINUED THE OPERATION TO THE END WITH DIATHERMIA. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04065 | 2015-07-06 | 20150611 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | IT WAS REPORTED THAT DURING A LAP NEPHRECTOMY PROCEDURE, THE TIP OF DEVICE BROKE OFF, FELL INTO PATIENT BUT WAS RETRIEVED THROUGH THE TROCAR. A 2ND DEVICE WAS USED COMPLETE THE PROCEDURE. NO REPORTED PATIENT CONSEQUENCE. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04084 | 2015-07-07 | 20150603 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35A |  | IT WAS REPORTED THAT DURING A SINGLE INCISION LAPAROSCOPIC HYSTERECTOMY WITH SALPINGO OPHERECTOMY PROCEDURE, WHEN DECIDING TISSUE NEAR THE CERVIX, THE DEVICE BEGAN GENERATING THE ERROR CODE "REPOSITION AND REACTIVATE". THIS HAPPENED SIX TO SEVEN TIMES AND THEN THE CODE CHANGED TO "REPLACE DEVICE". CASE COMPLETED WITH A COMPETING DEVICE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04085 | 2015-07-07 | 20150515 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35A |  | IT WAS REPORTED THAT DURING A LAPAROSCOPIC GYN PROCEDURE, A REPLACE INSTRUMENT ERROR WAS PROVIDED ON THE GENERATOR. NO OTHER DETAILS REGARDING APPLICATION OR PROCEDURE WERE PROVIDED. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04093 | 2015-07-07 | 20150525 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | IT WAS REPORTED THAT DURING A LAP TOTAL PROSTATECTOMY, THE DEVICE BECAME NOT TO CUT THE TARGET TISSUE WHEN TRYING TO MOVE I-BLADE IN ABOUT 2 HOURS THOUGH THE DEVICE WAS ACTIVATED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00126 | 2015-07-02 | 20150601 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57250600010 |  | TIBIAL IMPLANT LOOSENING WAS REPORTED. REVISION SURGERY IS PLANNED TO EXCHANGE THE TIBIAL IMPLANTS. |  | Y | N | HEALTH PROFESSIONAL
2954740-2015-00158 | 2015-07-07 | 20150618 | Malfunction | Manufacturer report | ORBIT GALAXY G2 MICROCOIL DELIVERY SYSTEM | ARTIFICIAL EMBOLIZATION DEVICE | CODMAN AND SHURTLEFF, INC |  | 641CX2505 |  | IT WAS REPORTED THAT TWO G2 COILS HERNIATED FROM THE ANEURYSM AND HAD RESHEATHING FAILURE. DURING COILING OF AN UNRUPTURED 5MM CAROTID ANEURYSM NEAR THE PCOM ARTERY USING G2 COILS, WHILE ATTEMPTING TO PLACE 4X10 GALAXY XTRASOFT AS THE FIRST COIL, IT WOULD NOT STAY IN THE ANEURYSM CORRECTLY. THE COIL WAS PULLED OUT AND TRIED TO RE-SHEATH IT FOR USE LATER IN THE CASE, BUT THE DELIVERY SHEATH DID NOT RECAPTURE THE COIL (IT WAS THEN DISCARDED). THEN, IT WAS DECIDED TO TRY A 4X10 GALAXY FILL, BUT AGAIN THE COIL WOULD NOT SIT CORRECTLY IN THE ANEURYSM. ONCE AGAIN TRIED TO RECAPTURE THE COIL FOR LATER USE, AND AGAIN THE COIL COULD NOT BE RE-SHEATHED AND NEED TO BE DISCARDED. AFTER PLACING A STENT ACROSS THE ANEURYSM, THE PHYSICIAN WAS ABLE TO INSERT 5 OTHER G2 COILS SUCCESSFULLY EMBOLIZE THE ANEURYSM. NO HARM OR INCREASED PROCEDURE TIME WAS OBSERVED AS A RESULT OF THE COILS NOT BEING ABLE TO BE RE-SHEATHED. ADDITIONAL INFORMATION INDICATED THAT ISSUE WAS THAT THE NECK WAS TOO LARGE TO HOLD COILS WITHOUT A STENT, AND THE COILS HERNIATED OUT OF THE ANEURYSM DURING POSITIONING. THE CAROTID WAS STRAIGHT AT THE POINT WHERE THE ANEURYSM WAS LOCATED (VERY CLOSE TO PCOM). ALL GUIDELINES WERE FOLLOWED DURING RECAPTURING OF THE COILS. AFTER THE EVENT, NO DAMAGES WERE NOTICED ON THE DEVICES (KINK, BEND, STRETCHED, FRACTURE, SEPARATE, ETC.), AND THE COILS REMAINED ATTACHED TO THE DELIVERY SYSTEMS. (B)(4). THE UNIDENTIFIED MICROCATHETER WAS NOT RETURNED. CONCERNING CLEANLINESS ONLY, THE MICROCOIL SYSTEM WAS RETURNED IN ALMOST PRISTINE CONDITION. THE SYSTEM WAS EITHER NOT USED OR WAS CLEANED/RINSED BEFORE BEING RETURNED WHICH MAY HAVE PRODUCED FURTHER DAMAGE. IT IS ALSO UNKNOWN IF THE DEVICE WAS FURTHER MANIPULATED AND/OR INSPECTED POST-PROCEDURALLY. ANY TRACE EVIDENCE THAT MAY HAVE BEEN COMPLAINT RELATED MAY HAVE BEEN ALTERED OR REMOVED PRIOR TO BEING RETURNED. AS VIEWED INTO THE RETURNED PACKAGING IT WAS FOUND THAT THE COIL WAS ENTANGLED. LOCATED AT THE DISTAL TIP OF THE SKIVE THE CORE WIRE PROTRUDES OUTSIDE THE SHEATH. PROXIMAL, BUT ADJACENT TO THE PROTRUSION SITE THE CORE WIRE IS KINKED. DUE TO POST-PROCEDURAL HANDING AND BEING UNREPORTED, THE CIRCUMSTANCES OF HOW AND WHEN THE CORE WIRE WAS KINKED CANNOT BE DETERMINED. THE ENTANGLED COIL WAS RETURNED WITH SEVERE INTERMITTENT DAMAGE ALONG ITS ENTIRE LENGTH. DUE TO POST-PROCEDURAL HANDING, THE CIRCUMSTANCES OF HOW AND WHEN THE COIL WAS DAMAGED CANNOT BE DETERMINED. MECHANICAL SHEATH DAMAGE CAUSED BY THE RESHEATHING TOOL WAS FOUND AT THE DISTAL TIP OF THE SKIVE. LOCATED ON THE TOP PROXIMAL END OF THE RESHEATHING TOOL IN THE OPEN CUTOUT SECTION, THE V NOTCH HAS BEEN SEVERELY FRACTURED WITH THE DAMAGED EDGE ELEVATED ABOVE THE SURFACE PLANE. THE LOCKING MECHANISM HAS COMPRESSION AND STRETCHING DAMAGE. NO MANUFACTURING DEFECTS WERE FOUND. CONCERNING THE POSITIONING DIFFICULTY AND THE COILS HERNIATING OUTSIDE THE ANEURYSM, IT WAS REPORTED THAT THE ANEURYSM¿S NECK WAS TOO LARGE AND THAT A STENT WAS REQUIRED TO COMPLETE THE COILING PROCESS. A REVIEW OF THE MANUFACTURING DOCUMENTATION ASSOCIATED WITH THIS LOT PRESENTED NO ISSUES DURING THE MANUFACTURING OR INSPECTION PROCESS THAT CAN BE RELATED TO THE REPORTED COMPLAINT. THE REPORTED EVENT OF RESHEATHING WAS CONFIRMED WITH ANALYSIS, BUT THE COIL HERNIATION COULD NOT BE CONFIRMED. HOWEVER, FOR THE COIL HERNIATION COULD BE RELATED TO THE ANEURYSM NECK SIZE AS INDICATED A STENT WAS USED TO COMPLETE THE PROCEDURE. WITH THE AVAILABLE INFORMATION, IT APPEARS THAT BOTH FAILURES WERE PROCEDURAL RELATED. ADDITIONALLY, REVIEW OF THE MANUFACTURING DOCUMENTATION ASSOCIATED WITH THIS LOT PRESENTED NO ISSUES DURING THE MANUFACTURING OR INSPECTION PROCESS THAT CAN BE RELATED TO THE REPORTED COMPLAINT. THEREFORE, NO CORRECTIVE ACTIONS WILL BE TAKEN AT THIS TIME. |  | N | Y | HEALTH PROFESSIONAL
2954740-2015-00159 | 2015-07-07 | 20150618 | Malfunction | Manufacturer report | ORBIT GALAXY G2 MICROCOIL DELIVERY SYSTEM | ARTIFICIAL EMBOLIZATION DEVICE | CODMAN AND SHURTLEFF, INC |  | 641CF0410 |  | THE PRODUCT WILL NOT BE RETURNED FOR ANALYSIS, AND ADDITIONAL INFORMATION WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. (B)(4). |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00125 | 2015-07-02 | 20150601 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M5725060XE20 |  | PATIENT DEVELOPED AN INFECTION. REVISION SURGERY IS PLANNED TO EXCHANGE THE POLY INSERTS. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00127 | 2015-07-02 | 20150601 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M5725060XE20 |  | PATIENT DEVELOPED AN INFECTION. REVISION SURGERY IS PLANNED TO EXCHANGE THE POLY INSERTS. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. |  | Y | N | HEALTH PROFESSIONAL
3005075853-2015-04113 | 2015-07-08 | 20150616 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | (B)(4). WHEN ADDITIONAL INFORMATION IS RECEIVED AND/OR THE DEVICE ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL MEDWATCH WILL BE SENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04146 | 2015-07-09 | 20150519 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | (B)(4). BATCH # IYZC79306. THE DEVICE WAS RECEIVED THE UPPER JAW DAMAGED AT THE PROXIMAL SIDE. IN ADDITION, THE DEVICE WAS RECEIVED WITH SKIVING OF THE HEAT SHRINK (SHAFT COVER) MATERIAL. ALL OF THE HEAT SHRINK MATERIAL APPEARED TO HAVE BEEN RETURNED. THE DEVICE WAS CONNECTED TO THE GENERATOR AND IT WAS RECOGNIZED. BECAUSE THE JAW WAS DETACHED NOT ALL FUNCTIONAL TESTING COULD BE PERFORMED WITH THE GENERATOR. A PROBABLE CAUSE OF THE DAMAGE TO THE JAW COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. THE IFU WARNS "PRIOR TO INSERTING OR REMOVING THE DEVICE FROM THE TROCAR, ENSURE THAT THE DEVICE IS IN THE STRAIGHT, NON-ARTICULATED POSITION BY CONFIRMING THAT THE INDICATOR ARROW ON THE ARTICULATION WHEEL IS ALIGNED WITH THE TOP OF THE INSTRUMENT." THE IFU ALSO WARNS THE USER TO DISCONTINUE USE IF BLACK HEAT SHRINK COVERING IS DAMAGED. THE BATCH HISTORY RECORDS WERE REVIEWED WITH NO ANOMALIES NOTED DURING THE MANUFACTURING PROCESS. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04178 | 2015-07-10 | 20150619 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | IT WAS REPORTED THAT DURING A TOTAL LAPAROSCOPIC HYSTERECTOMY, THE JAW BECAME NOT TO OPEN AND CLOSE WHEN THE DEVICE CUT THE VAGINAL WALL. THE VAGINAL WALL WAS THICK. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00122 | 2015-07-07 | 20141101 | Injury | Manufacturer report | IUNI G2 | UNICONDYLAR KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57220600210 |  | PATIENT SUSTAINED A PROXIMAL TIBIAL FRACTURE AT ONE DAY POST-OP. SURGERY OCCURRED TO PLATE THE TIBIA. IT WAS REPORTED THAT THE IMPLANTS ARE STABLE BUT THE TIBIA IS COLLAPSED. REVISION SURGERY IS PLANNED. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00123 | 2015-07-07 | 20150601 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600020 |  | PATIENT WITH A TOTAL KNEE IMPLANT TRIPPED AND TWISTED HIS KNEE. MEDIAL TIBIAL PLATEAU FRACTURE AND TIBIAL COMPONENT LOOSENING WERE IDENTIFIED. REVISION SURGERY TO A POSTERIOR STABILIZED IMPLANT SYSTEM OCCURRED. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00121 | 2015-07-07 | 20150501 | Injury | Manufacturer report | I DUO G2 | BICOMPARTMENTAL KNEE REPLACEMENT | CONFORMIS |  | M57230600210 |  | REVISION SURGERY IS OCCURRING FOR PATIENT WITH AN IDUO IMPLANT DUE TO CONTINUAL PAIN. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. |  | Y | N | HEALTH PROFESSIONAL
3005075853-2015-04182 | 2015-07-10 | 20150619 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | (B)(4). CONCLUSION: THE DEVICE WAS RECEIVED IN GOOD PHYSICAL CONDITION. UPON MECHANICAL TESTING IT WAS NOTED THAT WHEN THE DEVICE JAWS DID NOT OPEN UNAIDED ABOVE THE MINIMUM REQUIRED 20 DEGREES. THIS WAS CONFIRMED AND MEASURED WITH A CHART. IT SHOULD BE NOTED THAT THE DEVICE WAS TESTED WITH TEST MEDIA AND THERE WERE NO ADDITIONAL FINDINGS. THE JAWS DID OPEN AND CLOSE. THE DEVICE WAS TESTED ON THE GENERATOR AND PASSED ALL FUNCTIONAL TESTING. THE ENERGY OUTPUT DELIVERED FROM THE DEVICE WAS VERIFIED. ALL THREE TONES WERE HEARD DURING FUNCTIONAL TESTING (TONE 1 IS HEARD WHEN THE ENERGY ACTIVATION BUTTON IS PRESSED; TONE 2 IS HEARD WHEN TISSUE IMPEDANCE THRESHOLD IS REACHED AND TONE 3 IS HEARD WHEN THE CYCLE IS COMPLETE). SINCE NO DAMAGE WAS OBSERVED ON THE JAWS AND THE I BLADE NO CONCLUSION COULD BE DRAW AS WHAT MAY HAVE CAUSED THIS ISSUE. THE BATCH HISTORY RECORDS WERE REVIEWED WITH NO ANOMALIES NOTED DURING THE MANUFACTURING PROCESS. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04185 | 2015-07-10 | 20150316 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35 |  | IT WAS REPORTED THAT DURING A LAPAROSCOPIC TOTAL COLECTOMY PROCEDURE, THE JAWS BROKE INSIDE THE PATIENT. IT WAS STATED THAT ALL PIECES OF THE DEVICE WERE RETRIEVED WITHOUT ADDITIONAL MEDICAL INTERVENTION. NO PATIENT CONSEQUENCE WAS REPORTED. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04167 | 2015-07-13 | 20150617 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | IT WAS REPORTED THAT DURING A LAVH, THE TISSUE (ADIPO OR SCAR TISSUE) WAS CAUGHT IN THE DISTAL END OF THE JAWS WHILE THE DEVICE WAS BEING USED FOR THE VAGINAL WALL. WHEN THE DOCTOR REMOVED THE DEVICE FROM THE STUCK TISSUE, THE ELECTRODE PEELED AWAY. BESIDES, WHEN THE DEVICE WAS REMOVED VIA A TROCAR, THE ELECTRODE GOT LODGED IN THE EDGE OF THE DISTAL END OF THE TROCAR. THE ELECTRODE AND THE ACTIVE ROD CAME OFF AND WERE LEFT INSIDE THE TROCAR. THEN, THE TROCAR WAS REMOVED FROM THE PATIENT TO AVOID FALLING THE DETACHED PARTS INTO THE PATIENT. NO PIECES FELL INTO THE PATIENT. THE GENERATOR WAS UNKNOWN. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04209 | 2015-07-13 | 20150529 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35A |  | IT WAS REPORTED THAT DURING A TOTAL LAPAROSCOPIC HYSTERECTOMY PROCEDURE; THE GENERATOR DISPLAYED ¿REPLACE THE INSTRUMENT¿ ERROR AFTER TRYING WITH TWO DIFFERENT GENERATORS. THE PROCEDURE WAS COMPLETED USING SAME LIKE DEVICE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04211 | 2015-07-13 | 20150520 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35A |  | (B)(4). THE DEVICE WAS RECEIVED THE UPPER JAW DAMAGED AT THE RETENTION PIN AREA. THE DEVICE WAS CONNECTED TO THE GENERATOR AND IT WAS RECOGNIZED. BECAUSE THE JAW WAS DAMAGED NOT ALL FUNCTIONAL TESTING COULD BE PERFORMED WITH THE GENERATOR. THE CONDITION OF THE JAWS PREVENTED THE FUNCTIONALITY OF THE JAW. THE DEVICE WAS UNABLE TO CYCLE OPEN AND CLOSE. A PROBABLE CAUSE OF THE DAMAGE TO THE JAW COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. |  | N | Y | HEALTH PROFESSIONAL
2020394-2015-01016 | 2015-07-08 | 20110216 | Injury | Manufacturer report | G2 X FILTER SYSTEM - FEMORAL | VENA CAVA FILTER | BARD PERIPHERAL VASCULAR, INC. |  | RF400F |  | IT WAS REPORTED THAT APPROXIMATELY FIFTEEN MONTHS POST VENA CAVA FILTER DEPLOYMENT, THE PATIENT PRESENTED WITH COMPLAINT OF CHEST PAIN. DIAGNOSTIC IMAGING DEMONSTRATED A DETACHED LIMB. REPORTEDLY MULTIPLE ATTEMPTS AND CAREFUL MANIPULATION WERE REQUIRED TO SUCCESSFULLY CAPTURE AND RETRIEVE THE VENA CAVA FILTER; HOWEVER, A DETACH FILTER LIMB REMAINS EMBEDDED. NO ATTEMPT WAS MADE TO CAPTURE AND RETRIEVE THE DETACHED FILTER LIMB. THE PATIENT STATUS AT THIS ITEM IS UNKNOWN. |  | Y | Y | HEALTH PROFESSIONAL
2020394-2015-01017 | 2015-07-08 | 20150513 | Injury | Manufacturer report | G2 FILTER SYSTEM |  | BARD PERIPHERAL VASCULAR, INC. |  | UNKNOWN G2 |  | IT WAS REPORTED THAT APPROXIMATELY SEVEN YEARS POST VENA CAVA FILTER DEPLOYMENT, AT THE TIME OF SCHEDULED RETRIEVAL, SPOT FILM IMAGING DEMONSTRATED THE FILTER TO BE TILTED TO THE PATIENT'S RIGHT WITH A DETACHED ARM IN THE INTRACARDIAC POSITION. THE FILTER WAS SUCCESSFULLY RETRIEVED; HOWEVER, HE DETACHED ARM COULD NOT BE RETRIEVED. PATIENT IS TO FOLLOW UP WITH A CARDIOTHORACIC PHYSICIAN REGARDING THE RETAINED DETACHED LIMB. |  | Y | Y | HEALTH PROFESSIONAL
3004153240-2015-00128 | 2015-07-10 | 20150601 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600020 |  | POST-OPERATIVE KNEE PAIN WAS REPORTED. REVISION SURGERY IS PLANNED. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00129 | 2015-07-10 | 20150601 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M5725INT0600020 |  | KNEE LAXITY WAS REPORTED FOR PATIENT WITH A TOTAL KNEE REPLACEMENT. REVISION SURGERY IS PLANNED TO EXCHANGE THE POLY INSERTS. |  | Y | N | HEALTH PROFESSIONAL
3005075853-2015-04294 | 2015-07-14 | 20150625 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. AT THE TIME OF THIS SUBMISSION, THE DEVICE HAS NOT BEEN RETURNED FOR ANALYSIS. BATCH # UNK. |  | N | Y | HEALTH PROFESSIONAL
2954740-2015-00164 | 2015-07-14 | 20150617 | Malfunction | Manufacturer report | ORBIT GALAXY G2 MICROCOIL DELIVERY SYSTEM | ARTIFICIAL EMBOLIZATION DEVICE | CODMAN AND SHURTLEFF, INC |  | 641CX0410 |  | CONCLUSION UPDATED WITH ANALYSIS: IT WAS REPORTED BY THE HOSPITAL CONTACT THAT DURING AN ANEURYSM EMBOLIZATION WHEN THE PHYSICIAN TRIED TO PASS THE MICROCOIL (641CX0410/C11578) THROUGH THE PROWLER SELECT LP ES MICROCATHETER (606S155JX/15965185) IT DIDN´T PROGRESS. THEY HAD TO OPEN ANOTHER MICROCATHETER (DETAILS UNKNOWN) AND ANOTHER MICROCOIL (DETAILS UNKNOWN). THE DOCTOR MENTIONED THAT THIS EVENT CAUSED A DELAY IN THE PROCEDURE. IT IS UNKNOWN IF A CONTINUOUS FLUSH SOLUTION WAS MAINTAINED THROUGH THE MICROCATHETER. IT IS UNKNOWN IF THERE WAS ANY DIFFICULTY ENCOUNTERED WHEN INSERTING THE GUIDEWIRE INTO THE MICROCATHETER. THERE WAS NO RESISTANCE/KINKING WHEN INSERTING THE MICROCATHETER. THERE WERE NO DAMAGES NOTED ON THE MICROCATHETER. IT IS UNKNOWN IF IT WAS VISUALLY CONFIRMED THAT THE INTRODUCER WAS SEATED CORRECTLY AND THAT THE COIL WAS CO-ASXILLY ENTERING THE MICROCATHETER. IT IS UNKNOWN IF OTHER COILS HAD PREVIOUSLY GONE THROUGH THE MICROCATHETER WITHOUT DIFFICULTY. THE PRODUCTS WILL BE RETURNED FOR ANALYSIS. AS VIEWED THROUGH THE INTRODUCER SHEATH IT WAS FOUND THAT THE COIL WAS DETACHED AND NOT RETURNED. THIS WAS UNREPORTED DAMAGE. AS ADVANCED OUTSIDE THE DISTAL TIP OF THE GREEN INTRODUCER IT WAS FOUND THAT THE COIL WAS MECHANICALLY DETACHED AS NO SIGNS OF HEAT AND MELTING WERE FOUND. LOCATED 32.5 CENTIMETERS OFF THE PROXIMAL END THE CORE WIRE PROTRUDES OUTSIDE THE SHEATH. NO MECHANICAL SHEATH DAMAGE WAS FOUND AT THE PROTRUSION SITE THAT WOULD HAVE OPENED UP THE SKIVE. LOCATED ON THE TOP PROXIMAL END OF THE RESHEATHING TOOL IN THE OPEN CUTOUT SECTION, THE V NOTCH HAS BEEN FRACTURED WITH ONE SIDE OF THE DAMAGED EDGE RAISED ABOVE THE SURFACE PLANE. THE LOCKING MECHANISM HAS COMPRESSION AND STRETCHING DAMAGE. NO MANUFACTURING DEFECTS WERE FOUND. WHILE THE EXACT ROOT CAUSE CANNOT BE DETERMINED, MULTIPLE CONTRIBUTING FACTORS WERE FOUND THAT MAY HAVE CAUSED THE COILS RESISTANCE AND POSSIBLE UNINTENDED MECHANICAL DETACHMENT. THE ROOT CAUSE CANNOT BE DETERMINED CONCERNING THE UNREPORTED UNINTENDED MECHANICAL DETACHMENT OF THE NON-RETURNED COIL. THE PRIMARY CONTRIBUTING FACTOR TO THE COILS RESISTANCE MAY HAVE OCCURRED WHEN THE MICROCOIL SYSTEM WAS FIRST UNLOCKED FOR USE AND THE SHEATH WAS RETRACTED STRAIGHT BACK INSTEAD OF UP AT A FORTY-FIVE DEGREE ANGLE AND THEN BACK. WHEN THE SHEATH WAS PULLED STRAIGHT BACK, THE LOCKING MECHANISM CAUGHT THE INSIDE OF THE V NOTCH OF THE RESHEATHING TOOL AND BECAME EMBEDDED. IN ADDITION, THE LOCKING MECHANISM MAY NOT HAVE BEEN FULLY DISENGAGED OFF THE CORE WIRE. THE SHEATH ALSO CAUGHT THE V NOTCHES EXTENDED EDGES. THIS PRODUCED A BINDING ACTION BETWEEN THE DEVICE POSITIONING UNIT (DPU), THE SHEATH, AND THE COIL. THIS BINDING ACTION PRODUCED SIGNIFICANT RESISTANCE WHICH CAUSED THE CORE WIRE TO PROTRUDE OUTSIDE THE SHEATH AND MAY HAVE BEEN A FACTOR CONCERNING THE COILS UNINTENDED MECHANICAL DETACHMENT. IN THIS CONDITION SEVERE RESISTANCE WILL BE ENCOUNTERED WHEN ATTEMPTING TO ADVANCE THE COIL INSIDE THE MICROCATHETER. FOR OPTIMUM PRODUCT PERFORMANCE AND TO PREVENT POTENTIAL COMPLICATIONS, THE INSTRUCTIONS FOR USE (IFU) RECOMMENDS, ¿HOLD THE INTRODUCER SHEATH (LOOSELY-LOOPED) IN THE LEFT HAND. KEEPING THE INTRODUCER TIP NEAR THE RE-SHEATHING TOOL, GRASP THE DISTAL END OF THE RE-SHEATHING TOOL BETWEEN YOUR LEFT THUMB AND FOREFINGER. GRASP THE CLEAR TAB NEAR THE END OF THE INTRODUCER SHEATH BODY WITH THE THUMB AND FOREFINGER OF YOUR OTHER HAND. GENTLY PULL THE CLEAR TAB OF THE INTRODUCER SHEATH OUT AND AWAY FROM THE RE SHEATHING TOOL AT A 45-DEGREE ANGLE TO UNLOCK THE MICROCOIL. CONTINUE TO PULL THE TAB UNTIL AN ADDITIONAL 0.5 TO 1.0 INCHES (1.3 TO 2.5 CM) OF THE TRANSLUCENT MATERIAL IS EXPOSED. GENTLY FOLD THE TRANSLUCENT TAB TOWARDS THE DISTAL END, AND FIRMLY GRASP THE DISTAL END OF THE RE SHEATHING TOOL AND THE TRANSLUCENT TAB BETWEEN YOUR THUMB AND FOREFINGER, AS SHOWN IN FIGURE 3¿¿ CAUTION: IF UNUSUAL FRICTION IS NOTICED DURING ADVANCEMENT OR RETRACTION OF THE MICROCOIL SYSTEM, VERIFY THE LOCKING MECHANISM, OR CLEAR TAB IS UNLOCKED AND PULLED OUT FROM THE RESHEATHING TOOL APPROXIMATELY 1IN. (2-3CM).¿ IN ADDITION, WITHOUT THE RETURN OF THE COIL, THE PROWLER SELECT LP MICROCATHETER, AND THE ROTATING HEMOSTATIC VALVE (RHV) USED IN THE PROCEDURE, IT CANNOT BE DETERMINED IF THESE COMPONENTS CONTRIBUTED TO THE COMPLAINT EVENT. THE SECONDARY CONTRIBUTING FACTOR TO THE COILS RESISTANCE INSIDE THE MICROCATHETER MAY HAVE BEEN DUE TO DISTAL INTERFERENCE INSIDE THE UNIDENTIFIED MICROCATHETER. PER COMPLAINT, ¿¿IT IS UNKNOWN IF A CONTINUOUS FLUSH SOLUTION WAS MAINTAINED THROUGH THE MICROCATHETER¿¿ A SUB-COMPONENT TO THIS FACTOR MAY HAVE BEEN THE LACK OF A CONSISTENT FLUSH BASED ON THE BLOOD MIXTURE ADHERING TO THE COIL WHICH HAD A COPIOUS AMOUNT OF PROTEIN. THE EXACT SOURCE OF THIS INTERFERENCE; WHETHER OF A FIXED OR DETACHED NATURE CANNOT BE DETERMINED. IF A CONSISTENT FLUSH WAS NOT UTILIZED DURING THE PROCEDURE, THEN FOR OPTIMUM PRODUCT PERFORMANCE AND TO PREVENT POTENTIAL COMPLICATIONS, THE INSTRUCTIONS FOR USE (IFU) RECOMMENDS, ¿TO ACHIEVE OPTIMAL PERFORMANCE OF THE CODMAN MICROCOIL SYSTEM, IT IS IMPORTANT THAT A CONTINUOUS INFUSION OF AN APPROPRIATE FLUSH SOLUTION BE MAINTAINED. FIGURE 2 ILLUSTRATES THE CONNECTIONS NECESSARY FOR THE CODMAN MICROCOIL DELIVERY SYSTEM INCLUDING A TYPICAL CONTINUOUS SALINE FLUSH SET UP WITH PRESSURE BAG FOR THE CATHETER SYSTEMS.¿ THE THIRD CONTRIBUTING FACTOR TO THE COILS RESISTANCE INSIDE THE MICROCATHETER MAY HAVE BEEN OCCURRED DURING THE INSERTION OF THE COIL INTO THE MICROCATHETER. PER COMPLAINT, ¿¿IT IS UNKNOWN IF IT WAS VISUALLY CONFIRMED THAT THE INTRODUCER WAS SEATED CORRECTLY AND THAT THE COIL WAS CO-ASXILLY ENTERING THE MICROCATHETER¿¿ IF A MISALIGNMENT OCCURRED THEN FOR OPTIMUM PRODUCT PERFORMANCE AND TO PREVENT POTENTIAL COMPLICATIONS, THE INSTRUCTIONS FOR USE (IFU) RECOMMENDS, ¿¿ADDITIONALLY, IF THE INTRODUCER TIP AND MICROCATHETER HUB ARE MISALIGNED, DAMAGE MAY OCCUR TO THE MICROCOIL AS IT PASSES THROUGH THIS TRANSITION¿¿ BASED ON THE INFORMATION AND THE ANALYSIS, THE EVENT COULD NOT BE CONFIRMED, HOWEVER PROCEDURAL FACTORS MAY HAVE CONTRIBUTED TO THE EVENT. ADDITIONALLY, A REVIEW OF THE MANUFACTURING DOCUMENTATION ASSOCIATED WITH THIS LOT PRESENTED NO ISSUES DURING THE MANUFACTURING OR INSPECTION PROCESS THAT CAN BE RELATED TO THE REPORTED COMPLAINT. THEREFORE, NO CORRECTIVE ACTIONS WILL BE TAKEN AT THIS TIME. THIS IS 1 OF 2 REPORTS ASSOCIATED WITH REPORT # 1058196-2015-00147. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04331 | 2015-07-16 | 20150625 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | IT WAS REPORTED THAT DURING A COLON RESECTION WITH THE ENSEAL INSTRUMENT AND THE TOP JAW BROKE OFF DURING THE RESECTION. IT WASN'T INDICATED IF THE BROKEN PART OF THE JAW FELL INTO THE PATIENT. A SECOND ENSEAL INSTRUMENT WAS USED TO COMPLETE THE PROCEDURE WITH NO CONSEQUENCE TO THE PATIENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04332 | 2015-07-16 | 20150630 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C14 |  | (B)(4). INFORMATION ANTICIPATED, BUT UNAVAILABLE AT THIS TIME. BATCH # NI. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04334 | 2015-07-16 | 20150625 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C14 |  | IT WAS REPORTED THE TOP JAW BROKE OFF OF THE DEVICE. ADDITIONAL PRODUCT INFORMATION WAS NOT AVAILABLE AND THERE WERE NO PATIENT CONSEQUENCES REPORTED. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04333 | 2015-07-16 | 20150625 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C14 |  | (B)(4). ADDITIONAL INFORMATION: BATCH # M90R56. THE INSTRUMENT WAS RECEIVED WITH THE ENERGY BUTTON DETACHED AND RETURNED WITH THE INSTRUMENT. IN ADDITION, THE ELECTRODE WAS SLIGHTLY DETACHED FROM THE CERAMIC. UPON VISUAL INSPECTION TO THE DEVICE NO ANOMALIES WERE FOUND WITH THE TOP JAW AS REPORTED IN THE EVENT DESCRIPTION. THE BUTTON WAS REATTACHED TO THE INSTRUMENT AND TESTED WITH THE GEN11. TESTING FOUND A SHORT ON THE DEVICE WHEN THE JAWS ARE FULLY OR PARTIALLY CLOSED, RESULTING IN AN ALERT SCREEN "REPOSITION JAWS AND REACTIVE". THIS IS ADVISING THAT THE INSTRUMENT IS BEING ACTIVATED ON LOW IMPEDANCE (THIN) TISSUE OR METAL (SUCH AS STAPLES, CLIPS, RETRACTORS, OR CLAMPS). THE "REPLACE INSTRUMENT" ALERT SCREEN WOULD BE DISPLAYED AFTER RECEIVING THE "REPOSITION AND REACTIVATE" ALERT SCREEN TWICE. ALTHOUGH NO CONCLUSION COULD BE REACHED AS TO HOW THE BUTTON DETACHED FROM THE INSTRUMENT IT SHOULD BE NOTED THAT OUR MANUFACTURING PROCESS VERIFIES THE PRESENCE AND FUNCTIONALITY OF THE INSTRUMENT PRIOR TO SHIPPING. THERE IS INSUFFICIENT EVIDENCE RELATED TO WHAT CAUSED THE DAMAGE ON THE DEVICE. THE BATCH HISTORY RECORDS WERE REVIEWED WITH NO ANOMALIES NOTED DURING THE MANUFACTURING PROCESS. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04359 | 2015-07-17 | 20150625 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35A |  | (B)(4). ADDITIONAL INFORMATION: BATCH # M91G42. THE DEVICE WAS RECEIVED IN GOOD PHYSICAL CONDITION. THE DEVICE WAS TESTED ON THE GENERATOR AND PASSED ALL FUNCTIONAL TESTING. THE ENERGY OUTPUT DELIVERED FROM THE DEVICE WAS VERIFIED. ALL THREE TONES WERE HEARD DURING FUNCTIONAL TESTING (TONE 1 IS HEARD WHEN THE ENERGY ACTIVATION BUTTON IS PRESSED; TONE 2 IS HEARD WHEN TISSUE IMPEDANCE THRESHOLD IS REACHED AND TONE 3 IS HEARD WHEN THE CYCLE IS COMPLETE). NO ISSUES WERE NOTED WITH OPENING AND CLOSING OF THE JAWS. THERE WERE NO ANOMALIES NOTED WITH THE FUNCTIONALITY OF THE DEVICE. THE BATCH HISTORY RECORDS WERE REVIEWED WITH NO ANOMALIES NOTED DURING THE MANUFACTURING PROCESS. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04360 | 2015-07-17 | 20150604 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | (B)(4). BATCH # L92V4L. THE DEVICE WAS RECEIVED WITH THE TOP JAW ATTACHED AND NOT MISSING AS REPORTED. THE TIP OF THE I-BLADE WAS BROKEN OFF AND NOT RETURNED WITH THE DEVICE. THE DEVICE WAS CONNECTED TO THE GENERATOR AND IT WAS RECOGNIZED. BECAUSE THE I BLADE WAS DAMAGED NOT ALL FUNCTIONAL TESTING COULD BE PERFORMED WITH THE GENERATOR. THE CONDITION OF THE BLADE PREVENTED THE FUNCTIONALITY OF THE JAW. THE JAW WAS UNABLE TO CYCLE OPEN AND CLOSE. A PROBABLE CAUSE OF THE DAMAGE TO THE I BLADE COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. THE BATCH HISTORY RECORDS WERE REVIEWED WITH NO ANOMALIES NOTED DURING THE MANUFACTURING PROCESS. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04370 | 2015-07-17 | 20150626 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35 |  | (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. AT THE TIME OF THIS SUBMISSION, THE DEVICE HAS NOT BEEN RETURNED FOR ANALYSIS. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IS THE PACKAGING AVAILABLE FOR RETURN? IF YES, PLEASE, RETURN THE DEVICE WITH THE PACKAGING. IS THERE A PHOTO AVAILABLE OF THE PACKAGING/DEVICE? WHAT IS THE PACKAGING LOT NUMBER? |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04378 | 2015-07-17 | 20150625 | Injury | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35A |  | (B)(4). INFORMATION NOT AVAILABLE, DEVICE NOT RETURNED FOR ANALYSIS SHOULD THE INFORMATION BE PROVIDED LATER, A SUPPLEMENTAL MEDWATCH WILL BE SENT. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00134 | 2015-07-16 | 20150601 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600020 |  | PATIENT DEVELOPED AN INFECTION. REVISION SURGERY WAS SCHEDULED TO EXCHANGE THE POLY INSERTS. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00131 | 2015-07-16 | 20150601 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600010 |  | PATIENT WITH A TOTAL KNEE IMPLANT DEVELOPED AN INFECTION. THE IMPLANT WAS REMOVED AND A SPACER WAS PUT IN. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00130 | 2015-07-16 | 20150827 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M5725INT0600010 |  | IT WAS REPORTED THAT THE PATIENT HAD PAIN AND LOOSENING OF THE LATERAL POLY INSERT. A REVISION SURGERY WAS PERFORMED TO EXCHANGE THE POLY INSERTS. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THE DEVICE WAS MANUFACTURED TO SPECIFICATION. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00132 | 2015-07-16 | 20150601 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600010 |  | THE PATIENT FELL AS STRETCHED OUT THEIR LIGAMENTS. SUBSEQUENT REHABILITATION CAUSED ARTHROFIBROSIS. REVISION SURGERY OCCURRED TO CLEAN OUT THE KNEE AND EXCHANGE THE POLY INSERTS. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00133 | 2015-07-16 | 20150601 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600020 |  | REVISION SURGERY IS PLANNED FOR PATIENT WITH A TOTAL KNEE IMPLANT DUE TO POSSIBLE INFECTION. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. ALL STERILIZATION REQUIREMENTS WERE MET. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00136 | 2015-07-16 | 20150601 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600020 |  | PATIENT DEVELOPED AN INFECTION. REVISION SURGERY WAS SCHEDULED TO EXCHANGE THE POLY INSERTS. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. ALL STERILIZATION REQUIREMENTS WERE MET. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00135 | 2015-07-16 | 20150601 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600010 |  | PATIENT DEVELOPED AN INFECTION. REVISION SURGERY IS PLANNED TO WASH OUT THE KNEE AND EXCHANGE THE POLY INSERTS. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. ALL STERILIZATION REQUIREMENTS WERE MET. |  | Y | N | HEALTH PROFESSIONAL
3005075853-2015-04433 | 2015-07-20 | 20150629 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING STRAIGHT | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S45A |  | (B)(4). INFORMATION WAS NOT PROVIDED BY THE CONTACT. INFORMATION ANTICIPATED, BUT UNAVAILABLE AT THIS TIME. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT WAS THE PROCEDURE? UNKNOWN. WHEN WAS THE PROCEDURE? UNKNOWN. WHO WAS THE SURGEON? IS THE DEVICE AVAILABLE TO BE RETURNED FOR EVALUATION? YES. WERE THERE ANY PATIENT CONSEQUENCES AS A RESULT OF THIS ISSUE? WHAT WAS THE ISSUE WITH THIS DEVICE? DID THE TOP JAW BREAK OFF OF THE DEVICE? DID THE TOP JAW FALL INTO THE PATIENT? WAS THE TOP JAW RETRIEVED FROM THE PATIENT? DID THIS CAUSE A CHANGE IN THE PLAN OF CARE FOR THE PATIENT? ADDITIONAL INFORMATION: NO PATIENT CONSEQUENCES REPORTED. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04436 | 2015-07-21 | 20150601 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | (B)(4). INFORMATION WAS NOT PROVIDED BY THE CONTACT. (B)(4). THE DEVICE WAS RECEIVED WITH THE TOP OF THE I-BLADE UPPER TIP BROKEN OFF NOT RETURNED. THE DEVICE WAS CONNECTED TO THE GENERATOR AND IT WAS RECOGNIZED. BECAUSE THE I BLADE WAS DAMAGED NOT ALL FUNCTIONAL TESTING COULD BE PERFORMED WITH THE GENERATOR. THE CONDITION OF THE BLADE PREVENTED THE FUNCTIONALITY OF THE JAW. THE JAW WAS UNABLE TO CYCLE OPEN AND CLOSE. THE INSTRUMENT WAS DISASSEMBLED AND NO EVIDENCE WAS FOUND AS TO WHAT COULD HAVE CONTRIBUTED TO THE ACTIVATION ISSUES. A PROBABLE CAUSE OF THE DAMAGE TO THE I BLADE COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. THE BATCH HISTORY RECORDS WERE REVIEWED WITH NO ANOMALIES NOTED DURING THE MANUFACTURING PROCESS. |  | N | Y | HEALTH PROFESSIONAL
1818910-2015-26087 | 2015-07-21 | 20130916 | Injury | Manufacturer report | G2 PRESS FIT STEM SZ5 | HIP FEMORAL STEM/SLEEVE | DEPUY IRELAND 9616671 |  | 900000105 |  | ASR REVISION, ASR XL- LEFT, REASON(S) FOR REVISION: ALVAL / SOFT TISSUE REACTION, PAIN. 16 JULY 2015 - (B)(6) LEGAL LETTER, ADDED PATIENT NAME, DOB, GENDER, NEW REASON FOR REVISION: HIGH METAL IONS. ADDED EXP DATES TO ALL PRODUCTS, (B)(4) 21/07/2015. |  | Y | N | HEALTH PROFESSIONAL
3005075853-2015-04623 | 2015-07-27 | 20150706 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | (B)(4). INFORMATION ASKED FOR BUT UNKNOWN OR NOT PROVIDED DURING INITIAL CONTACT. INFORMATION ANTICIPATED, BUT UNAVAILABLE AT THIS TIME WHEN ADDITIONAL INFORMATION IS RECEIVED AND/OR THE DEVICE ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL MEDWATCH WILL BE SENT. BATCH # NI ADDITIONAL INFORMATION RECEIVED: PATIENT WAS MORBIDLY OBESE AND THERE WERE A LOT OF REMOVALS AND REINSERTIONS OF THE INSTRUMENT INTO TROCAR PORTS, NO ERRORS FROM THE GENERATOR DURING THE PROCEDURE. THE RISK MANAGEMENT DEPARTMENT WILL NOT RELEASE THE DEVICE, AT THIS TIME. |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00140 | 2015-07-21 | 20150601 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M5725INT0600020 |  | IT WAS REPORTED THAT THE PATIENT'S LEG IS SLIGHTLY VARUS AND THE LATERAL JOINT SPACE IS NARROW. REVISION SURGERY IS PLANNED TO RESURFACE THE PATELLA AND EXCHANGE THE POLY INSERTS. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00138 | 2015-07-21 | 20150601 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS | M57250600010 |  |  | THE PATIENT IS EXPERIENCING PAIN. REVISION SURGERY IS PLANNED TO REVISE THE PATELLA AND EXCHANGE THE POLY INSERTS. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00137 | 2015-07-22 | 20150601 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600020 |  | LAXITY IN THE KNEE WAS REPORTED FOR PATIENT WITH A TOTAL KNEE IMPLANT. REVISION SURGERY IS PLANNED TO EXCHANGE THE POLY INSERTS. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00139 | 2015-07-22 | 20150601 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600020 |  | REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00141 | 2015-07-21 | 20150601 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57250600010 |  | PATIENT PRESENTED WITH AN INFECTION. CAUSE IS UNKNOWN. REVISION SURGERY OCCURRED TO EXCHANGE THE POLY INSERTS. |  | Y | N | HEALTH PROFESSIONAL
2020394-2015-01205 | 2015-07-21 | 20150119 | Injury | Manufacturer report | G2 FILTER SYSTEM |  | BARD PERIPHERAL VASCULAR, INC. |  | UNKNOWN G2 |  | IT WAS REPORTED BY THE HEALTH CARE PROVIDER THAT DIAGNOSTIC IMAGING NOTED THE FILTER HAD MIGRATED AND TWO DETACHED LEGS, ONE FILTER LEG LOCATED IN THE LEFT PULMONARY BRANCH AND ONE FILTER LEG LOCATED IN THE SOFT TISSUES OF THE RETROPERITONEUM. REPORTEDLY, UPON RETROSPECTIVE IMAGE REVIEW, IT WAS DISCOVERED THAT THIS HAS BEEN PRESENT SINCE 2013. CURRENT FILTER AND PATIENT STATUS ARE UNKNOWN. |  | Y | Y | HEALTH PROFESSIONAL
2954740-2015-00182 | 2015-07-28 | 20150616 | Malfunction | Manufacturer report | ORBIT GALAXY G2 MICROCOIL DELIVERY SYSTEM | ARTIFICIAL EMBOLIZATION DEVICE | CODMAN AND SHURTLEFF, INC |  | 641CX2505 |  | IT WAS INITIALLY REPORTED THAT THE FAILURE OCCURRED DURING THE PRE-DEPLOYMENT ELECTRICAL CHECK. ADDITIONAL INFORMATION FOUND DURING PRODUCT ANALYSIS INDICATED THAT THE DEVICE WAS USED IN THE PATIENT. THE COMPLAINT MET MDR REPORTING CRITERIA ON (B)(6) 2015 WHEN THE ANALYSIS WAS COMPLETED. CONCOMITANT PRODUCTS: ENPOWER CONNECTING CABLE (LOT P10196), UNSPECIFIED DETACHMENT CONTROL BOX. COMPLAINT CONCLUSION: THE DEVICE WAS RETURNED FOR ANALYSIS. AS VIEWED INTO THE RETURNED BOX IT WAS FOUND THAT THE COIL WAS RETURNED FULLY UNSHEATHED, ENTANGLED, AND STRETCHED. THE PROXIMAL SECTION OF THE COIL WAS RETURNED STRETCHED. LOCATED OFF THE DISTAL END OF THE STRETCHED COIL, THE REMAINDER OF THE COIL IS UNDAMAGED. THE CIRCUMSTANCES OF HOW AND WHEN THIS UNREPORTED DAMAGE OCCURRED CANNOT BE DETERMINED. BLOOD AND CONTRAST WERE FOUND ON THE INTRODUCER SHEATH AND INSIDE THE OUTER SHEATH AND THE COIL WAS FULLY UNSHEATHED. THIS CANNOT BE EXPLAINED AS THE INSTRUCTIONS FOR USE (IFU) RECOMMEND THAT THE PRE-DEPLOYMENT ELECTRICAL CHECK OCCUR WITH THE SYSTEMS STILL RESIDING INSIDE THE HOOP DISPENSER. PER IFU, ¿¿VERIFY THE FUNCTIONALITY OF THE CODMAN MICROCOIL DELIVERY SYSTEM BEFORE PROCEEDING WITH MICROCOIL PLACEMENT. THIS NEEDS TO BE DONE WITH THE MICROCOIL STILL IN THE HOOP¿¿ UPON RECEIPT, WHEN THE CONTROL CABLE WAS CONNECTED TO A TEST MICROCOIL SYSTEM AND THE DCB, THE ENPOWER¿S SYSTEMS GO GREEN LIGHT ILLUMINATED. THE ENPOWER REMOTE CABLE WAS TESTED TO (B)(4) AND PASSED ELECTRICAL SPECIFICATIONS BY QCI. THE RETURNED ECB CONTAINED IN THIS COMPLAINT PASSED FUNCTIONALITY TESTING BY DETACHING TEN MICROCOILS ON THE FIRST DETACHMENT CYCLE. NO MANUFACTURING DEFECTS WERE FOUND ON THE CONTROL CABLE. THE DETACHMENT FIBER IS STILL INTACT AND DID NOT RECEIVE HEAT AND MELT. NO MANUFACTURING DEFECTS WERE FOUND ON THE MICROCOIL SYSTEM. THE DEVICE POSITIONING UNIT (DPU) FAILED ELECTRICAL TESTING WITH RESISTANCE FAILING WITH READINGS FLOATING UPWARD >28.85 MEGA AND THE ENPOWER SYSTEMS GO GREEN LIGHT FAILED TO ILLUMINATE. COMPRESSION ON THE RESISTIVE HEATING COIL SECTION CAUSED THE DPU TO PASS ELECTRICAL TESTING WITH RESISTANCE AT 54.5 OHMS AND THE ENPOWER SYSTEMS GO GREEN LIGHT ILLUMINATED. A REVIEW OF THE MANUFACTURING DOCUMENTATION ASSOCIATED WITH THIS LOT PRESENTED NO ISSUES DURING THE MANUFACTURING OR INSPECTION PROCESS THAT CAN BE RELATED TO THE REPORTED COMPLAINT. THE COIL DETACHMENT FAILURE WAS CONFIRMED DURING PRODUCT ANALYSIS SINCE THE DPU FAILED ELECTRICAL TESTING. THE MOST LIKELY CONTRIBUTING FACTOR TO THE MICROCOIL SYSTEM FAILING THE PRE-DEPLOYMENT ELECTRICAL CHECK MAY HAVE BEEN DUE TO A FRACTURE OF THE SOLDER JOINT CONNECTION INSIDE THE RESISTIVE HEATING COIL. THE CIRCUMSTANCES OF HOW AND WHERE THIS DAMAGE OCCURRED CANNOT BE DETERMINED AS ALL MICROCOIL SYSTEMS ARE ELECTRICALLY TESTED PRIOR TO FINAL PACKAGING. ALTHOUGH IT WAS REPORTED THAT THE EVENT OCCURRED DURING THE PRE-DEPLOYMENT ELECTRICAL CHECK, THE CONDITION OF THE RETURNED DEVICE (BLOOD AND CONTRAST ON THE DEVICE) PROVIDES EVIDENCE THAT THE FAILURE TO DETACH OCCURRED DURING USE IN THE PATIENT. SINCE THE DEVICE WAS RETURN UNSHEATHED AND ENTANGLED, THE COIL DAMAGE WAS MOST LIKELY RELATED TO POST-PROCEDURE HANDLING/SHIPPING. THERE WAS NO EVIDENCE OF A MANUFACTURING ISSUE RELATED TO THE EVENT; THEREFORE, NO CORRECTIVE ACTIONS WILL BE TAKEN AT THIS TIME. THIS IS AN INITIAL/FINAL MDR. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04666 | 2015-07-28 | 20150707 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | (B)(4). WE RECEIVED AN IMAGE TO REVIEW FOR THIS ANALYSIS. THE INSTRUMENT WAS NOT RETURNED. THIS ANALYSIS IS ONLY ASSOCIATED WITH THE PICTURES PROVIDED BY THE ACCOUNT. A REVIEW OF THE IMAGE SHOWS WHAT APPEARS TO BE A ENSEAL G2 ART STRA SEALER 35CM (NSLG2S35A). ONE OF THE IMAGES REVIEWED WAS OF THE DISTAL PORTION OF THE INSTRUMENT, THE JAW. THE IMAGE SHOWED THE UPPER JAW OF THE INSTRUMENT PRESENT BUT DETACHED FROM THE JAW ASSEMBLY. IT ALSO SHOWED THAT THE UPPER PORTION OF THE I-BLADE WAS LIFTED OR BROKEN. THE DAMAGE TO THE I-BLADE IS WHAT ALLOWED THE UPPER JAW TO DETACH. NO CONCLUSION COULD BE MADE AS TO THE REPORTED ISSUE BASED ON THE IMAGES PROVIDED. THE INFORMATION YOU PROVIDED IS COMPILED MONITORED AND REVIEWED BY UPPER MANAGEMENT ON A ROUTINE BASIS FOR ANY ASSOCIATED TRENDS. WE VALUE YOUR ASSISTANCE IN PROVIDING US AN OPPORTUNITY TO EVALUATE THE REPORTED EVENT AS ALL INFORMATION REPORTED TO OUR COMPANY IS CRITICAL TO OUR CONTINUOUS IMPROVEMENT EFFORTS. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04676 | 2015-07-28 | 20150527 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C25 |  | (B)(4). INFORMATION ASKED FOR BUT UNKNOWN OR NOT PROVIDED DURING INITIAL CONTACT. SHOULD THE INFORMATION BE PROVIDED LATER, A SUPPLEMENTAL MEDWATCH WILL BE SENT. (B)(4). THE DEVICE WAS RECEIVED WITH THE TOP OF THE I-BLADE UPPER TIP BROKEN OFF NOT RETURNED. THE DEVICE WAS CONNECTED TO THE GENERATOR AND IT WAS RECOGNIZED. THE DEVICE ACTIVATED ONCE THE ENERGY BUTTON WAS PRESSED. BECAUSE THE I BLADE WAS DAMAGED NOT ALL FUNCTIONAL TESTING COULD BE PERFORMED WITH THE GENERATOR. THE CONDITION OF THE BLADE PREVENTED THE FUNCTIONALITY OF THE JAW. THE JAW WAS UNABLE TO CYCLE OPEN AND CLOSE. THE INSTRUMENT WAS DISASSEMBLED AND NO EVIDENCE WAS FOUND AS TO WHAT COULD HAVE CONTRIBUTED TO THE ACTIVATION ISSUES. A PROBABLE CAUSE OF THE DAMAGE TO THE I BLADE COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04679 | 2015-07-28 | 20150702 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35 |  | (B)(4). INFORMATION ASKED FOR BUT UNKNOWN OR NOT PROVIDED DURING INITIAL CONTACT. INFORMATION IS UNAVAILABLE; DEVICE NOT RETURNED. NO DEVICE RECEIVED FOR ANALYSIS AT TIME OF SUBMISSION OF 3500A. WHEN ADDITIONAL INFORMATION IS RECEIVED AND/OR THE DEVICE ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL MEDWATCH WILL BE SENT. BATCH # UNK. ADDITIONAL INFORMATION: THE ENTIRE UPPER JAW CAME OFF THE DEVICE. SURGEON DOESN'T KNOW IF HE WAS TOO ROUGH WITH THE DEVICE OR NOT. HE BELIEVES HE WASN'T AS HE USED IT THE SAME WAY HE ALWAYS USES IT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04682 | 2015-07-28 | 20150630 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35A |  | (B)(4). THE DEVICE WAS RECEIVED IN GOOD VISUAL CONDITION. UPON VISUAL INSPECTION UNDER MAGNIFICATION IT WAS NOTICED THAT THE UPPER JAW WAS SLIGHTLY DAMAGED AT THE RETENTION PIN INTERPHASE. RESULTING IN THE JAW BEEN LOOSE. HOWEVER THE DAMAGE WAS NOT SIGNIFICANT ENOUGH TO PREVENT THE DEVICE FROM CYCLING. THE DEVICES WAS TESTED ON THE GENERATOR AND PASSED ALL FUNCTIONAL TESTING. THE ENERGY OUTPUT DELIVERED FROM THE DEVICE WAS VERIFIED. ALL THREE TONES WERE HEARD DURING FUNCTIONAL TESTING (TONE 1 IS HEARD WHEN THE ENERGY ACTIVATION BUTTON IS PRESSED; TONE 2 IS HEARD WHEN TISSUE IMPEDANCE THRESHOLD IS REACHED; AND TONE 3 IS HEARD WHEN THE CYCLE IS COMPLETE). A PROBABLE CAUSE OF THE DAMAGE TO THE JAW COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. THE BATCH HISTORY RECORDS WERE REVIEWED WITH NO ANOMALIES NOTED DURING THE MANUFACTURING PROCESS. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04681 | 2015-07-28 | 20150630 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35A |  | (B)(4). THE DEVICE WAS RECEIVED IN GOOD PHYSICAL CONDITION. THE DEVICE WAS TESTED ON THE GENERATOR AND PASSED ALL FUNCTIONAL TESTING. THE ENERGY OUTPUT DELIVERED FROM THE DEVICE WAS VERIFIED. ALL THREE TONES WERE HEARD DURING FUNCTIONAL TESTING (TONE 1 IS HEARD WHEN THE ENERGY ACTIVATION BUTTON IS PRESSED; TONE 2 IS HEARD WHEN TISSUE IMPEDANCE THRESHOLD IS REACHED AND TONE 3 IS HEARD WHEN THE CYCLE IS COMPLETE). NO ISSUES WERE NOTED WITH OPENING AND CLOSING OF THE JAWS. THERE WERE NO ANOMALIES NOTED WITH THE FUNCTIONALITY OF THE DEVICE. THE BATCH HISTORY RECORDS WERE REVIEWED WITH NO ANOMALIES NOTED DURING THE MANUFACTURING PROCESS. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04680 | 2015-07-29 | 20150707 | No answer provided | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | (B)(4). ADDITIONAL INFORMATION: BATCH #M90T07. THE DEVICE WAS RECEIVED WITH THE TOP OF THE I-BLADE UPPER TIP BROKEN OFF NOT RETURNED. THE DEVICE WAS CONNECTED TO THE GENERATOR AND IT WAS RECOGNIZED. BECAUSE THE I BLADE WAS DAMAGED NOT ALL FUNCTIONAL TESTING COULD BE PERFORMED WITH THE GENERATOR. THE CONDITION OF THE BLADE PREVENTED THE FUNCTIONALITY OF THE JAW. THE JAW WAS UNABLE TO CYCLE OPEN AND CLOSE. A PROBABLE CAUSE OF THE DAMAGE TO THE I BLADE COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. THE BATCH HISTORY RECORDS WERE REVIEWED WITH NO ANOMALIES NOTED DURING THE MANUFACTURING PROCESS. |  | N | N | HEALTH PROFESSIONAL
3005075853-2015-04708 | 2015-07-29 | 20150513 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C25 |  | IT WAS REPORTED THAT DURING A LAP EMPHYSEMATOUS BULLA, ¿REPLACE INSTRUMENT¿ WAS DISPLAYED AFTER A FEW ACTUATIONS. ALTHOUGH THE DEVICE WAS RE-ASSEMBLED AND THE GEN11 WAS REBOOTED, THE EVENT CONTINUED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04720 | 2015-07-29 | 20150609 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C45 |  | IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY PROCEDURE; THE SURGEON MENTIONED THAT THE JAWS OF THE DEVICE DID NOT CLOSE PROPERLY. THE DOCTOR CONTINUED TO USE THE DEVICE TO COMPLETE THE PROCEDURE. NO ADVERSE PATIENT CONSEQUENCES WERE REPORTED. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04730 | 2015-07-29 | 20150618 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | (B)(4). INFORMATION ASKED FOR BUT UNKNOWN OR NOT PROVIDED DURING INITIAL CONTACT. SHOULD THE INFORMATION BE PROVIDED LATER, A SUPPLEMENTAL MEDWATCH WILL BE SENT. (B)(4). THE DEVICE WAS RECEIVED THE UPPER JAW DAMAGED AT THE PROXIMAL SIDE. IN ADDITION, THE DEVICE WAS RECEIVED WITH SKIVING OF THE HEAT SHRINK (SHAFT COVER) MATERIAL. ALL OF THE HEAT SHRINK MATERIAL APPEARED TO HAVE BEEN RETURNED. THE DEVICE WAS CONNECTED TO THE GENERATOR AND IT WAS RECOGNIZED. BECAUSE THE JAW WAS DAMAGED NOT ALL FUNCTIONAL TESTING COULD BE PERFORMED WITH THE GENERATOR. THE CONDITION OF THE JAWS PREVENTED THE FUNCTIONALITY OF THE JAW. THE DEVICE WAS UNABLE TO CYCLE OPEN AND CLOSE. A PROBABLE CAUSE OF THE DAMAGE TO THE JAW COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. THE IFU WARNS "PRIOR TO INSERTING OR REMOVING THE DEVICE FROM THE TROCAR, ENSURE THAT THE DEVICE IS IN THE STRAIGHT, NON-ARTICULATED POSITION BY CONFIRMING THAT THE INDICATOR ARROW ON THE ARTICULATION WHEEL IS ALIGNED WITH THE TOP OF THE INSTRUMENT." THE IFU ALSO WARNS THE USER TO DISCONTINUE USE IF BLACK HEAT SHRINK COVERING IS DAMAGED. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04732 | 2015-07-29 | 20150603 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35A |  | (B)(4). SHOULD THE INFORMATION BE PROVIDED LATER, A SUPPLEMENTAL MEDWATCH WILL BE SENT. BATCH #IYZD20245 THE DEVICE WAS RECEIVED IN GOOD VISUAL CONDITION. UPON VISUAL INSPECTION UNDER MAGNIFICATION IT WAS NOTICED THAT THE UPPER JAW WAS SLIGHTLY DAMAGED AT THE RETENTION PIN INTERPHASE. RESULTING IN THE JAW BEEN LOOSE AND DIFFICULTY IN OPENING AND CLOSING OF THE JAWS. HOWEVER THE DAMAGE WAS NOT SIGNIFICANT ENOUGH TO PREVENT THE DEVICE FROM CYCLING. THE DEVICES WAS TESTED ON THE GENERATOR AND PASSED ALL FUNCTIONAL TESTING. THE ENERGY OUTPUT DELIVERED FROM THE DEVICE WAS VERIFIED. ALL THREE TONES WERE HEARD DURING FUNCTIONAL TESTING (TONE 1 IS HEARD WHEN THE ENERGY ACTIVATION BUTTON IS PRESSED; TONE 2 IS HEARD WHEN TISSUE IMPEDANCE THRESHOLD IS REACHED; AND TONE 3 IS HEARD WHEN THE CYCLE IS COMPLETE). A PROBABLE CAUSE OF THE DAMAGE TO THE JAW COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04736 | 2015-07-29 | 20150512 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35 |  | (B)(4). INFORMATION ASKED FOR BUT UNKNOWN OR NOT PROVIDED DURING INITIAL CONTACT. SHOULD THE INFORMATION BE PROVIDED LATER, A SUPPLEMENTAL MEDWATCH WILL BE SENT. BATCH #M90F4W. THE DEVICE WAS RETURNED WITH THE UPPER JAW DETACHED AND RETURNED WITH THE DEVICE. THE ELECTRODE AND CERAMIC WERE INTACT AND NOT DETACHED AS REPORTED. THE DEVICE WAS CONNECTED TO THE GENERATOR AND IT WAS RECOGNIZED. BECAUSE THE JAW WAS DETACHED NOT ALL FUNCTIONAL TESTING COULD BE PERFORMED WITH THE GENERATOR. A PROBABLE CAUSE OF THE DAMAGE TO THE JAW COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. |  | N | Y | HEALTH PROFESSIONAL
2020394-2015-01252 | 2015-07-24 | 20130823 | Malfunction | Manufacturer report | G2 EXPRESS FILTER SYSTEM - FEMORAL | VENA CAVA FILTER | BARD PERIPHERAL VASCULAR, INC. |  | RF400F |  | THE BRAND NAME AND CATALOG NUMBER HAVE BEEN CHANGED TO UNKNOWN. THE COMPLAINANT ALLEGES BOTH A G2 AND G2 EXPRESS; THEREFORE, THE REPORT WILL REFLECT UNKNOWN VENA CAVA FILTER. MANUFACTURING REVIEW: A MANUFACTURING REVIEW COULD NOT BE CONDUCTED AS THE LOT NUMBER WAS NOT PROVIDED. VISUAL/MICROSCOPIC INSPECTION: AS THE DEVICE WAS NOT RETURNED, AN INSPECTION COULD NOT BE PERFORMED. FUNCTIONAL/PERFORMANCE EVALUATION: AS THE DEVICE WAS NOT RETURNED, AN EVALUATION COULD NOT BE PERFORMED. MEDICAL RECORDS REVIEW: AS MEDICAL RECORDS WERE NOT PROVIDED, A REVIEW COULD NOT BE PERFORMED. IMAGE/PHOTO REVIEW: NO MEDICAL IMAGES HAVE BEEN MADE AVAILABLE TO THE MANUFACTURER. CONCLUSION: THE INVESTIGATION IS INCONCLUSIVE FOR THE ALLEGED LIMB DETACHMENT AND PERFORATION OF THE IVC WALL. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. LABELING REVIEW: THE CURRENT IFU (INSTRUCTIONS FOR USE) STATES: WARNINGS/POSSIBLE COMPLICATIONS INCLUDE, BUT ARE NOT LIMITED TO, THE FOLLOWING: - FILTER FRACTURES ARE A KNOWN COMPLICATION OF VENA CAVA FILTERS. THERE HAVE BEEN SOME REPORTS OF SERIOUS PULMONARY AND CARDIAC COMPLICATIONS WITH VENA CAVA FILTERS REQUIRING THE RETRIEVAL OF THE FRAGMENT UTILIZING ENDOVASCULAR AND/OR SURGICAL TECHNIQUES. - PERFORATION OR OTHER ACUTE OR CHRONIC DAMAGE OF THE IVC WALL. EXPLANTED (APPROX (B)(6) 2013). |  | Y | Y | HEALTH PROFESSIONAL
2020394-2015-01250 | 2015-07-24 |  | Malfunction | Manufacturer report | G2 FILTER SYSTEM |  | BARD PERIPHERAL VASCULAR, INC. |  | UNKNOWN G2 |  | THE LOT NUMBER WAS NOT PROVIDED; THEREFORE, THE DEVICE HISTORY RECORDS COULD NOT BE REVIEWED. THE INVESTIGATION IS CURRENTLY UNDERWAY. THE INFORMATION PROVIDED BY BARD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BARD. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04742 | 2015-07-29 | 20150714 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | (B)(4). BATCH #IYZC50456. ADDITIONAL INFORMATION RECEIVED: DID THE MOVEABLE TOP JAW BREAK OFF? YES. DID THE BLACK PTC BREAK OFF OF THE JAW (ATTACHED TO MOVEABLE TOP JAW)? NO. DID THE CERAMIC (ON THE LOWER NON-MOVEABLE JAW) BREAK OFF? NO. DID THE ELECTRODE (ON THE LOWER NON-MOVEABLE JAW) BREAK OFF? NO. IS THE BROKEN PIECE BEING RETURNED WITH THE DEVICE? UNKNOWN. OR WAS THE BROKEN PIECE DISCARDED? UNKNOWN THE DEVICE WAS RECEIVED WITH THE TOP JAW FULLY ATTACHED AND NOT BROKEN OFF AS REPORTED. THE DEVICE'S HEAT SHRINK (SHAFT COVER) MATERIAL WAS SKIVED AND ALL OF THE HEAT SHRINK MATERIAL APPEARED TO HAVE BEEN RETURNED. THE DEVICE WAS MECHANICALLY TESTED AND NO ANOMALIES WERE NOTED. THE DEVICE WAS TESTED ON THE GENERATOR AND PASSED ALL FUNCTIONAL TESTING. THE ENERGY OUTPUT DELIVERED FROM THE DEVICE WAS VERIFIED. ALL THREE TONES WERE HEARD DURING FUNCTIONAL TESTING (TONE 1 IS HEARD WHEN THE ENERGY ACTIVATION BUTTON IS PRESSED; TONE 2 IS HEARD WHEN TISSUE IMPEDANCE THRESHOLD IS REACHED AND TONE 3 IS HEARD WHEN THE CYCLE IS COMPLETE). THE IFU WARNS "PRIOR TO INSERTING OR REMOVING THE DEVICE FROM THE TROCAR, ENSURE THAT THE DEVICE IS IN THE STRAIGHT, NON-ARTICULATED POSITION BY CONFIRMING THAT THE INDICATOR ARROW ON THE ARTICULATION WHEEL IS ALIGNED WITH THE TOP OF THE INSTRUMENT." THE IFU ALSO WARNS THE USER TO DISCONTINUE USE IF BLACK HEAT SHRINK COVERING IS DAMAGED. THE BATCH HISTORY RECORDS WERE REVIEWED WITH NO ANOMALIES NOTED DURING THE MANUFACTURING PROCESS. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04760 | 2015-07-30 | 20150702 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE PTC DETACHED FROM THE JAW. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT. ONE DEVICE WAS DISCARDED. |  | N | Y | HEALTH PROFESSIONAL
1818910-2015-27004 | 2015-07-31 | 20150706 | Injury | Manufacturer report | G2 PRESS FIT STEM SZ8L | HIP FEMORAL STEM/SLEEVE | DEPUY IRELAND 9616671 |  | 900000128 |  | THIS COMPLAINT IS STILL UNDER INVESTIGATION. DEPUY WILL NOTIFY THE FDA OF THE RESULTS OF THIS INVESTIGATION ONCE IT HAS BEEN COMPLETED. |  | Y | N | HEALTH PROFESSIONAL
3005075853-2015-04845 | 2015-07-31 | 20150716 | Malfunction | Manufacturer report | ENSEAL G2 ARTICULATING | ELECTROSURGICAL, CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S35A |  | IT WAS REPORTED THAT DURING A LEFT COLECTOMY PROCEDURE, THE RESIDENT WAS IN THE FINAL STAGES OF USING THE ENSEAL TO MOBILIZE THE COLON. AT THIS TIME THE TOP JAW OF THE ENSEAL BROKE OFF IN THE PATIENT¿S ABDOMINAL CAVITY. THE PHYSICIAN RETRIEVED THE BROKEN PIECE AND PLACED WITH DEFECTIVE INSTRUMENT. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04848 | 2015-07-31 | 20150717 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | (B)(4). INFORMATION WAS NOT PROVIDED BY THE INITIAL CONTACT. INFORMATION IS UNAVAILABLE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. AT THE TIME OF THIS SUBMISSION, THE DEVICE HAS NOT BEEN RETURNED FOR ANALYSIS. BATCH # UNK. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04726 | 2015-07-31 | 20150603 | Malfunction | Manufacturer report | ENSEAL G2 CURVED JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2C35 |  | (B)(4). BATCH # M90E1Y. THE DEVICE WAS RECEIVED IN GOOD VISUAL CONDITION. UPON VISUAL INSPECTION UNDER MAGNIFICATION IT WAS NOTICED THAT THE UPPER JAW WAS SLIGHTLY DAMAGED AT THE RETENTION PIN INTERPHASE. RESULTING IN THE JAW BEEN LOOSE AND DIFFICULTY IN OPENING AND CLOSING OF THE JAWS. HOWEVER THE DAMAGE WAS NOT SIGNIFICANT ENOUGH TO PREVENT THE DEVICE FROM CYCLING. THE DEVICES WAS TESTED ON THE GENERATOR AND PASSED ALL FUNCTIONAL TESTING. THE ENERGY OUTPUT DELIVERED FROM THE DEVICE WAS VERIFIED. ALL THREE TONES WERE HEARD DURING FUNCTIONAL TESTING (TONE 1 IS HEARD WHEN THE ENERGY ACTIVATION BUTTON IS PRESSED; TONE 2 IS HEARD WHEN TISSUE IMPEDANCE THRESHOLD IS REACHED; AND TONE 3 IS HEARD WHEN THE CYCLE IS COMPLETE). A PROBABLE CAUSE OF THE DAMAGE TO THE JAW COULD BE NOT ALLOWING THICK AND FIBROUS TISSUES TO DENATURE PRIOR TO I-BLADE ADVANCEMENT. THE BATCH HISTORY RECORDS WERE REVIEWED WITH NO ANOMALIES NOTED DURING THE MANUFACTURING PROCESS. |  | N | Y | HEALTH PROFESSIONAL
3005075853-2015-04856 | 2015-07-31 | 20150716 | Malfunction | Manufacturer report | ENSEAL G2 STRAIGHT JAW | ELECTROSURGICAL CUTTING & COAGULATION & ACCESSORIES | ETHICON ENDO-SURGERY, LLC. | NA | NSLG2S45 |  | (B)(4). INFORMATION IS UNAVAILABLE; DEVICE NOT RETURNED. NO DEVICE RECEIVED FOR ANALYSIS AT TIME OF SUBMISSION OF 3500A. WHEN ADDITIONAL INFORMATION IS RECEIVED AND/OR THE DEVICE ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL MEDWATCH WILL BE SENT. BATCH #UNK. |  | N | Y | HEALTH PROFESSIONAL
3004153240-2015-00143 | 2015-07-31 | 20150701 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M57250600010 |  | LAXITY IN THE KNEE WAS REPORTED. REVISION SURGERY IS SCHEDULED TO EXCHANGE THE POLY INSERTS. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00148 | 2015-07-31 | 20150701 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M5725INT0600010 |  | REVISION SURGERY IS SCHEDULED FOR PT WITH A TOTAL KNEE IMPLANT. REASON FOR REVISION IS UNK AT THIS TIME. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00146 | 2015-07-31 | 20150701 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M5725INT0600020 |  | THE PT HAS RESIDUAL PAIN BEHIND THE PATELLA. SURGERY IS SCHEDULED TO IMPLANT A PATELLA IMPLANT AND EXCHANGE THE POLY INSERTS. THE NATIVE PATELLA WAS NOT RESURFACED DURING THE PRIMARY PROCEDURE. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00144 | 2015-07-31 | 20150701 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M5725INT0600020 |  | THE PT HAS AN INFECTION. REVISION SURGERY IS SCHEDULED TO EXCHANGE THE POLY INSERTS. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00145 | 2015-07-31 | 20150701 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M5725INT0600020 |  | ARTHROFIBROSIS WAS REPORTED FOR PT WITH A TOTAL KNEE IMPLANT. SURGERY IS SCHEDULED TO RESOLVE THE FIBROSIS AND EXCHANGE THE POLY INSERTS. REVIEW OF THE DEVICE HISTORY RECORD INDICATES THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATIONS. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00142 | 2015-07-30 | 20150701 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS |  | M5725INT0600010 |  | ARTHROFIBROSIS WAS REPORTED FOR PATIENT WITH A TOTAL KNEE IMPLANT. SURGERY IS SCHEDULED TO RESOLVE THE FIBROSIS AND EXCHANGE THE POLY INSERTS. |  | Y | N | HEALTH PROFESSIONAL
3004153240-2015-00149 | 2015-08-03 | 20150701 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57250600020 |  | PATIENT HAS LIMITED RANGE OF MOTION POST-OPERATIVELY. THE PATIENT HAD PNEUMONIA POST SURGERY AND DID NOT DO THERAPY FOR AWHILE AS A RESULT. MANIPULATION OCCURRED BUT DID NOT RESTORE RANGE OF MOTION. REVISION SURGERY IS PLANNED TO DEBRIDE SCAR TISSUE, WASH OUT THE JOINT, AND EXCHANGE THE POLY INSERTS. |  | Y | N | PHYSICIAN
3004153240-2015-00150 | 2015-08-03 | 20150701 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M57250600020 |  | PATIENT DEVELOPED AN INFECTION. REVISION SURGERY OCCURRED TO EXCHANGE THE POLY INSERTS. |  | Y | N | PHYSICIAN
3004153240-2015-00151 | 2015-08-03 | 20150701 | Injury | Manufacturer report | ITOTAL G2 | TOTAL KNEE REPLACEMENT SYSTEM | CONFORMIS, INC. |  | M5725060XE20 |  | THE TIBIAL TRAY IMPLANT IS NOT SEATED PROPERLY POSTERIORLY. CEMENT WAS NOT EXTRUDED AND THE TIBIAL IMPLANT WAS NOT ADEQUATELY SEATED DURING THE PRIMARY PROCEDURE. REVISION SURGERY IS PLANNED TO RE-POSITION THE TIBIAL TRAY AND EXCHANGE THE POLY INSERTS. |  | Y | N | PHYSICIAN
2020394-2015-01268 | 2015-07-30 | 20100225 | Injury | Manufacturer report | G2 FILTER SYSTEM - FEMORAL |  | BARD PERIPHERAL VASCULAR, INC. |  | RF310F |  | AFTER FURTHER CLINICAL REVIEW, THIS EVENT HAS BEEN DETERMINED TO MEET THE DEFINITION OF A SERIOUS INJURY MDR PER 21 CFR 803. AS A RESULT OF THE RE-REVIEW OF THE IMAGES, THE PRODUCT CATALOG NUMBER HAS BEEN IDENTIFIED AS RF310F. ONE CD OF PROCEDURES WERE RE-REVIEWED AND THE IMAGES DEMONSTRATED THAT ON (B)(6) 2009 AN ANGIOPLASTY WAS PERFORMED ON THE LEFT ILIAC VEIN, THE ILIAC VEIN WAS APPROXIMATELY 85% OCCLUDED. INFLATION OF A PTA BALLOON WAS DEMONSTRATED IN THE LEFT COMMON ILIAC VEIN. IMAGES DEMONSTRATED A PATENT ILIAC VEIN WITH GOOD BLOOD FLOW. TWO IMAGES WERE PROVIDED OF THE IVC, ONE IMAGE WAS A VENA CAVA GRAM; HOWEVER, THE RENAL TAKEOFFS COULD NOT BE IDENTIFIED IN THIS IMAGE. THE SECOND IMAGE DEMONSTRATED A VENA CAVA FILTER UPRIGHT POSITION, THE EXACT LOCATION COULD NOT BE DETERMINED BASED ON A DSA VENA CAVA IMAGE PROVIDED. A SNARE HOOK COULD NOT BE IDENTIFIED ON THE APEX OF THIS FILTER WHICH INDICATED THE FILTER IS A G2 FILTER. THE DEPLOYMENT SYSTEM IS DEMONSTRATED VIA FEMORAL VEIN ACCESS; THEREFORE, THE FILTER WAS A G2 FEMORAL FILTER. THERE ARE NO IMAGES PROVIDED THAT INDICATED A SNARE DEVICE WAS USED TO CAPTURE THE FILTER. DUE TO THE POOR IMAGE QUALITY THE NUMBER OF LIMBS OF THE FILTER COULD NOT BE COUNTED. THE APEX OF THE FILTER WAS AGAINST THE WALL OF THE IVC. BASED ON THE RE-REVIEW OF THE IMAGES, A TILTED FILTER IN THE IVC CAN BE CONFIRMED; HOWEVER, SUCCESSFUL FILTER RETRIEVAL CANNOT BE CONFIRMED. MULTIPLE FOLLOW UP ATTEMPTS WERE MADE WITH THE FACILITY TO OBTAIN ANY INFORMATION PERTAINING TO THE PATIENT, PRODUCT, AND/OR PROCEDURAL DETAILS (E.G. DATE OF THE EVENT, RELEVANT TEST DATA, RELEVANT HISTORY, LOT #, CATALOG #, IMPLANT AND/OR EXPLANTED DATES, AND CONCOMITANT PRODUCT(S) OR THERAPY) THAT WERE NOT PREVIOUSLY OBTAINED DURING THE INITIAL INVESTIGATION. |  | Y | Y | HEALTH PROFESSIONAL
4972565 | 2015-07-16 | 20150710 | Malfunction | User Facility report | ENSEAL TISSUE SEALER G2 | ELECTROSURGICAL CUTTING COAG INSTRUMENT | ETHICON ENDO SURGERY | * | * |  | WHEN USING DEVICE, IT WAS CLICKING AND NOT CLOSING PROPERLY. ANOTHER DEVICE OBTAINED. PROCEDURE CONTINUED. NO HARM TO PATIENT.====================== MANUFACTURER RESPONSE FOR LAPAROSCOPIC TISSUE SEALER G2, LAPAROSCOPIC TISSUE SEALER G2 (PER SITE REPORTER).======================NOTHING YET. |  | N | Y | PHYSICIAN

## Monthly Summary
year_month | count
2015-02 | 41
2015-03 | 71
2015-04 | 40
2015-05 | 51
2015-06 | 47
2015-07 | 95
2015-08 | 3

## Event Types
Event Type | Count
Malfunction | 194
Injury | 151
No answer provided | 3

## Brand Names
Brand Name | Count
ITOTAL G2 | 98
ENSEAL G2 ARTICULATING | 68
ENSEAL G2 CURVED JAW | 57
ENSEAL G2 STRAIGHT JAW | 38
ORBIT GALAXY G2 MICROCOIL DELIVERY SYSTEM | 21
IUNI G2 | 11
G2 FILTER SYSTEM - FEMORAL | 9
IDUO G2 | 7
G2 FILTER SYSTEM | 7
RECOVERY FILTER SYSTEM | 5
G2 PRESS FIT STEM SZ8L | 3
STELLAR 100 - EUR G2 | 2
RECOVERY CONE REMOVAL SYSTEM | 2
G2 | 2
ENSEAL G2 | 2
ENSEAL G2 ARTICULATING STRAIGHT | 1
G2 FILTER SYSTEM | 1
G2 X FILTER SYSTEM - FEMORAL | 1
RECOVERY CONE | 1
G2 PRESS FIT STEM SZ5 | 1

## Case Metadata
Case Name | MDL Number | Filing Date | Court | Status | Settlement | Data Window Start | Data Window End | Total Records | Search Queries Used
IVC Filters (C.R. Bard) | 2641 | 2015-08-03 | D. Arizona | Active | N/A | 2015-02-01 | 2015-08-03 | 348 | 3