## All Reports
report_number | date_received | date_of_event | event_type | report_source_code | brand_name | generic_name | manufacturer_d_name | model_number | catalog_number | device_class | event_description | manufacturer_narrative | adverse_event_flag | product_problem_flag | device_operator | manufacturer_g1_name
MW5035189 | 2014-03-20 | 20140219 | Malfunction | Voluntary report | CELECT | FILTER | COOK MEDICAL | IGTCFS-65-UNI-CELECT | G52915 |  | THE INNER DILATOR OF THE SHEATH SPLIT APART AFTER COMPLETING A POWER INJECTION AT 450 PSI. THERE WAS NO PT INJURY. |  | N | Y | HEALTH PROFESSIONAL
3002808486-2014-00008 | 2014-03-27 | 20140210 | Malfunction | Manufacturer report | COOK CELECT FEMORAL VENA CAVA FILTER SET | DTK FILTER, INTRAVASCULAR, CARDIOVASCULAR | WILLIAM COOK EUROPE |  | IGTCFS-65-FEM-CELECT-PERM |  | DESCRIPTION ACCORDING TO CLAIMANT: "A ROUTINE X-RAY WAS PERFORMED ON THE PATIENT. THE X-RAY RESULTS ILLUSTRATED A LEG OF THE FILTER BROKE-OFF". THE PATIENT DID NOT REQUIRE ANY ADDITIONAL PROCEDURES DUE TO THIS OCCURRENCE. ACCORDING TO THE INITIAL REPORTER, THE PATIENT DID NOT EXPERIENCE ANY ADVERSE EFFECTS DUE TO THIS OCCURRENCE. |  | N | Y | HEALTH PROFESSIONAL
3002808486-2014-00010 | 2014-04-10 | 20140313 | Malfunction | Manufacturer report | COOK CELECT NAVALIGN FEMORAL VENA CAVA FILTER SET | DTK FILTER, INTRAVASCULAR, CARDIOVASCULAR | WILLIAM COOK EUROPE |  | IGTCFS-65-1-FEM-CELECT |  | DESCRIPTION OF EVENT ACCORDING TO COMPLAINANT: DEVICE INITIALLY PLACED SUCCESSFULLY ON (B)(6) 2011. URETER PENETRATION NOTED IN 2012, CAUSING THE PATIENT TO RETURN PERIODICALLY TO FLUID DRAINAGE. ON (B)(6) 2014, A PRIMARY LEG FRACTURE WAS NOTED UPON A PERIODIC FLUID DRAINAGE. FILTER WILL NOT BE RETRIEVED. FRACTURED PIECE OF PRIMARY LEG IS ASYMPTOMATIC AND WILL NOT BE RETRIEVED. PATIENT OUTCOME: THE PATIENT NOT REQUIRE ANY ADDITIONAL PROCEDURES DUE TO THIS OCCURRENCE. ACCORDING TO THE INITIAL REPORTER, THE PATIENT DID NOT EXPERIENCE ANY ADVERSE EFFECTS DUE TO THIS OCCURRENCE. |  | N | Y | HEALTH PROFESSIONAL
3002808486-2014-00017 | 2014-05-22 |  | Injury | Manufacturer report | NAVALIGN COOK CELECT JUGULAR VENA CAVA FILTER SET | DTK FILTER, INTRAVASCULAR, CARDIOVASCULAR | WILLIAM COOK EUROPE |  | IGTCFS-65-1-JUG-CELECT |  | (B)(4). UNKNOWN AS INFORMATION WAS NOT PROVIDED BY REPORTER. UNKNOWN AS LOT# IS UNKNOWN. AS NO DEVICE, IMAGING STUDIES OR HOSPITAL OR MEDICAL RECORDS HAVE BEEN AVAILABLE, THE INVESTIGATION IS LIMITED TO THE ALLEGATIONS IN THE EVENT DESCRIPTION. CONSEQUENTLY, BASED ON VERY LIMITED INFORMATION IT IS DIFFICULT TO COMMENT ON THE ALLEGED FILTER FRACTURE AND MIGRATION. INSTRUCTIONS FOR USE LIST DAMAGE TO THE VENA CAVA, AND VENA CAVA PERFORATION AS POTENTIAL ADVERSE EVENTS. SUCH EVENTS HAVE ALSO BEEN REPORTED IN THE PUBLISHED MEDICAL AND SCIENTIFIC LITERATURE. CATALOG #KNOWN, HOWEVER LOT# WAS NOT PROVIDED, WHY DEVICE HISTORY RECORD CANNOT BE INVESTIGATED. HOWEVER, NOTHING INDICATES THE DEVICE WAS NOT MANUFACTURED ACCORDING TO SPECIFICATION. THE EXACT ROOT CAUSE FOR THE ALLEGED FILTER FRACTURE CANNOT BE DETERMINED BASED ON THE LIMITED INFORMATION PROVIDED. MONITORING OF SIMILAR REPORTS WILL CONTINUE. |  | Y | Y | HEALTH PROFESSIONAL
3002808486-2014-00015 | 2014-05-14 | 20140415 | Malfunction | Manufacturer report | COOK CELECT NAVALIGN JUGULAR & FEMORAL VENA CAVA FILTER SET | DTK FILTER, INTRAVASCULAR, CARDIOVASCULAR | WILLIAM COOK EUROPE |  | IGTCFS-65-1-UNI-CELECT |  | DESCRIPTION OF EVENT ACCORDING TO COMPLAINANT: ATTEMPTED A FILTER RETRIEVAL. 1 PRIMARY AND 1 SECONDARY STRUT HAD FRACTURED AND MIGRATED. THE LOCATION IS NOT KNOWN AT THIS TIME. THE FILTER WAS NOT RETRIEVED AS A RESULT OF FRACTURE. UPDATE FROM SALES REPRESENTATIVE RECEIVED ON 05/01/2014: THE PATIENT WENT TO THE ER ON (B)(6) 2014 FOR AN EVENT UNRELATED TO THE FILTER. DURING PATIENT'S VISIT, A CT SCAN WAS PERFORMED AND THE PHYSICIAN SAW THE FRACTURED FILTER, THUS REFERRING THE PATIENT TO AN INTERVENTIONAL RADIOLOGIST AT THE HOSPITAL. THE RADIOLOGIST ATTEMPTED TO RETRIEVE THE FILTER AND WAS SUCCESSFUL. THE FRACTURED STRUTS REMAIN IN THE PATIENT. ONE OF THE FRACTURED STRUTS IS LOCATED IN THE KIDNEY AND THE OTHER ONE IS LOCATED IN THE BOWEL. THE PHYSICIAN STATED THE STRUTS WILL REMAIN IN THE PATIENT AND NOT BE REMOVED. PATIENT OUTCOME: THE PATIENT WAS ASYMPTOMATIC RELATED TO THE FILTER. ACCORDING TO THE INITIAL REPORTER. THE PATIENT DID NOT EXPERIENCE ANY ADVERSE EFFECTS DUE TO THIS OCCURRENCE. |  | Y | Y | HEALTH PROFESSIONAL
3002808486-2014-00018 | 2014-05-30 | 20120510 | Injury | Manufacturer report | COOK CELECT FEMORAL AND JUGULAR VENA CAVA FILTER | DTK FILTER, INTRAVASCULAR, CARDIOVASCULAR | WILLIAM COOK EUROPE |  | IGTCFS-65-UNI-CELECT- |  | DESCRIPTION ACCORDING TO CLAIMANTS ATTORNEY: IT IS ALLEGED THAT "ON (B)(6) 2012, PT WAS ADMITTED TO (B)(6) EMERGENCY ROOM COMPLAINING OF RIGHT FLANK PAIN. IT WAS ALLEGEDLY DISCOVERED THAT THE CELECT HAD "MULTIPLE LIMBS PENETRATING THE WALL OF THE INFERIOR VENA CAVA. ONE LIMB OF THE FILTER APPEARED TO HAVE PENETRATED THE RIGHT URETER AND ANOTHER LIMB APPEARED TO HAVE PENETRATED THE DUODENUM. ON (B)(6) 2012, PT UNDERWENT AN "EXPLORATORY LAPAROTOMY" AND OPEN CAVOGRAM WITH REMOVAL OF THE IVC FILTER AND REPAIR TO HIS DUODENUM". PATIENT OUTCOME REQUESTED BUT NOT PROVIDED. |  | Y | Y | HEALTH PROFESSIONAL
MW5036703 | 2014-06-16 | 20131113 | Injury | Voluntary report | CELECT IVC FILTER | IVC FILTER | COOK, INC. | * | * |  | CELECT IVC FILTER WAS IMPLANTED ON (B)(6), 2013. THREE MONTHS LATER, DR. TRIED TO REMOVE FILTER. UNSUCCESSFUL. TWO WEEKS LATER, ANOTHER DR. TRIED TO REMOVE FILTER. UNSUCCESSFUL. (B)(6), I WENT TO ANOTHER HOSPITAL AND THE FILTER WAS REMOVED. THE FIRST 2 ATTEMPTS MANGLED THE FILTER AND BROKE OFF 2 SPOKES, OR LEGS OF THE DEVICE. THEY WERE LOCATED IN THE HEART AND LUNG, UNABLE TO REMOVE. THEY ARE STILL THERE. THAT IS 4 SURGERIES AND I AM LEFT WITH 2 PIECES OF METAL IN 2 VITAL ORGANS. MANUFACTURER REPORTED AS BARD. |  | N | N | HEALTH PROFESSIONAL
MW5036784 | 2014-06-18 | 20140606 | Injury | Voluntary report | COOK FILTER CELECT UNI | IVC FILTER | COOK MEDICAL |  | G52915 |  | AFTER APPROPRIATE POSITIONING, A CELECT FILTER WAS ADVANCED. THIS FILTER WAS UNABLE TO DEPLOY. SEVERAL ATTEMPTS AT DEPLOYMENT DID NOT RELEASE ALL OF THE TIMES. THE FILTER WAS REMOVED AND A TULIP FILTER WAS REINSERTED WITHOUT INCIDENT. |  | N | Y | HEALTH PROFESSIONAL
3002808486-2014-00019 | 2014-05-30 | 20110329 | Malfunction | Manufacturer report | COOK CELECT FEMORAL AND JUGULAR VENA CAVA FILTER | DTK FILTER, INTRAVASCULAR, CARDIOVASCULAR | WILLIAM COOK EUROPE |  | IGTCFS-65-UNI-CELECT-PERM |  | (B)(4). COMMON DEVICE NAME UNKNOWN BUT REFERRED TO AS CELECT FILTER. LOT NUMBER UNKNOWN AS INFORMATION WAS NOT PROVIDED BY REPORTER. CATALOG NUMBER UNKNOWN AS INFORMATION WAS NOT PROVIDED BY REPORTER BUT REFERRED TO AS CELECT FILTER. EXPIRATION DATE UNKNOWN AS LOT NUMBER IS UNKNOWN. MFG DATE UNKNOWN AS LOT NUMBER IS UNKNOWN. RPN AND LOT NUMBER WERE NOT PROVIDED; THEREFORE, THE INVESTIGATION OF DEVICE HISTORY RECORD HAS NOT BEEN POSSIBLE. HOWEVER, NOTHING INDICATES THE DEVICE WAS NOT MANUFACTURED ACCORDING TO SPECIFICATION. AS NO DEVICE, IMAGING STUDIES OR HOSPITAL OR MEDICAL RECORDS HAVE BEEN AVAILABLE, THE INVESTIGATION IS LIMITED TO THE ALLEGATIONS IN THE LITIGATION. CONSEQUENTLY, BASED ON VERY LIMITED INFORMATION IT IS DIFFICULT TO COMMENT ON THE ALLEGED FILTER TILT AND DIFFICULT RETRIEVAL ATTEMPTS. THE EXACT ROOT CAUSE FOR THE ALLEGED FILTER TILT AND DIFFICULT RETRIEVAL ATTEMPTS CANNOT BE DETERMINED BASED ON THE LIMITED INFORMATION PROVIDED. THE INSTRUCTIONS FOR USE LIST DAMAGE TO THE VENA CAVA, AND VENA CAVA PERFORATION AS POTENTIAL ADVERSE EVENTS. SUCH EVENTS HAVE ALSO BEEN REPORTED IN THE PUBLISHED MEDICAL AND SCIENTIFIC LITERATURE. MONITORING OF SIMILAR REPORTS WILL CONTINUE. |  | N | Y | HEALTH PROFESSIONAL
3002808486-2014-00023 | 2014-07-11 | 20140530 | Injury | Manufacturer report | COOK CELECT PLATINUM NAVALIGN FEMORAL VENA CAVA FILTER SET |  | WILLIAM COOK EUROPE |  | IGTCFS-65-1-FEM-CELEC |  | DESCRIPTION ACCORDING TO COMPLAINANT: ON (B)(6) 2014, THE FILTER WAS DEPLOYED. IT WAS PLACED RIGHT AT RENALS. THE SEAT WAS INFRARENAL. THE HOOK WAS SUPRARENAL. THERE WAS MINIMAL TILT AT DEPLOYMENT. ON (B)(6) 2014, THE PATIENT HAD AN ANKLE SURGERY. ON (B)(6) 2014, IT WAS DETERMINED AT THE ER THAT THE PATIENT HAD A PULMONARY EMBOLISM. ON (B)(6) 2014, A CHEST CT SCAN REVEALED THAT THE FILTER HAD TILTED TO THE RIGHT AND THAT 1 SECONDARY LEG WAS ABOVE THE RENAL VEIN. PATIENT OUTCOME: THE PATIENT WILL REQUIRE AN ADDITIONAL FILTER REMOVAL PROCEDURE DUE TO THIS OCCURRENCE. ACCORDING TO THE INITIAL REPORTER, THE PATIENT DID EXPERIENCE THE ADVERSE EFFECT OF A PULMONARY EMBOLISM DUE TO THIS OCCURRENCE. |  | Y | N | HEALTH PROFESSIONAL
3002808486-2014-00032 | 2014-07-18 | 20140629 | Death | Manufacturer report | COOK CELECT NAVALIGN JUGULAR & FEMORAL VENA CAVA FILTER SET | DTK FILTER, INTRAVASCULAR, CARDIOVASCULAR | WILLIAM COOK EUROPE |  | IGTCFS-65-1-UNI-CELECT |  | DESCRIPTION ACCORDING TO COMPLAINANT: THE PATIENT PRESENTED WITH A MALPOSITIONED FILTER IN THE LEFT VENTRICLE, 10 DAYS POST IMPLANT. ON EXAM, THE STAFF REPORTED THE PATIENTS IVC MEASURED TWICE WHAT IT DID DURING THE INITIAL PLACEMENT PROCEDURE. RETRIEVAL WAS UNSUCCESSFUL PERCUTANEOUSLY. THE PATIENT WAS TAKEN TO THE OPERATING ROOM AND AN OPEN PROCEDURE WAS PERFORMED TO REMOVE THE DEVICE. ONCE REMOVED, THE DEVICE WAS FOUND TO BE ENCAPSULATED BY A HUGE CLOT BURDEN, IN AND AROUND THE FILTER. PATIENT OUTCOME: (B)(6) 2014 ACCORDING TO PHYSICIAN THE PATIENT WAS NOT DOING WELL AND WAS NOT REACTING TO NORMAL STIMULI. ADDITIONAL INFORMATION REQUESTED BUT NOT PROVIDED. |  | Y | N | HEALTH PROFESSIONAL
MW5037367 | 2014-07-21 | 20140710 | Injury | Voluntary report | FEM-CELECT-PEM | FEM-CELECT-PEM | COOK |  | IGTCFS-65-FEM-CELECT-FEM |  | A (B)(6) MALE WITH DVT POOR CANDIDATE FOR LONG-TERM ANTICOAGULATION THUS AN IVC FILTER PLACEMENT WAS INDICATED. THE IVC FILTER PERFORATED AND SHEATH WITHIN THE RIGHT COMMON FEMORAL VEIN. THE FILTER HAD TO BE DEPLOYED WITHIN THE RIGHT COMMON ILIAC VEIN. HE DID HAVE A SECOND FILTER PLACED IN THE NORMAL INFRARENAL POSITION, PROTECTING HIM FROM DVT. THE FOLLOWING DAY, THE PT WAS BROUGHT BACK TO SURGERY TO ATTEMPT TO REMOVE THE FILTER. THE HOOK APPEARED TO BE EMBEDDED IN THE WALL OF THE VEIN SO IT COULD NOT BE REMOVED. |  | N | Y | HEALTH PROFESSIONAL
3002808486-2014-00030 | 2014-07-16 | 20110627 | Injury | Manufacturer report | COOK CELECT FEMORAL VENA CAVA FILTER SET | DTK: FILTER, INTRAVASCULAR, CARDIOVASCULAR | WILLIAM COOK EUROPE |  | IGTCFS-65-FEM-CELECT |  | ALLEGATIONS ACCORDING TO COMPLAINANT: IT IS ALLEGED THAT ON (B)(6) 2011, THE PT WAS INVOLVED IN AN AUTOMOBILE ACCIDENT AND SUSTAINED A FRACTURE OF HIS BACK WITH PARAPLEGIA. ON (B)(6) 2011, PT UNDERWENT THE PROPHYLACTIC PLACEMENT OF A COOK CELECT IVC FILTER. ON (B)(6) 2011, REMOVAL OF THE CELECT FILTER WAS ALLEGEDLY ATTEMPTED AT (B)(6) MEDICAL CTR. THE DEVICE WAS UNABLE TO BE RETRIEVED ALLEGEDLY DUE TO "THE HOOK OF THE COOK CELECT FILTER WAS LIKELY COVERED BY INTIMAL HYPERPLASIA CAUSING IT TO BE EMBEDDED IN THE WALL OF THE INFERIOR VENA CAVA." IT IS ALLEGED THAT PLAINTIFF'S PHYSICIANS ARE "UNWILLING TO ATTEMPT MORE AGGRESSIVE MEASURE(S) TO REMOVE THE EMBEDDED FILTER" DUE TO THE RISK INVOLVED. |  | Y | Y | HEALTH PROFESSIONAL
3002808486-2014-00021 | 2014-06-30 | 20140529 | Malfunction | Manufacturer report | COOK CELECT FEMORAL AND JUGULAR VENA CAVA FILTER | DTK FILTER, INTRAVASCULAR, CARDIOVASCULAR | WILLIAM COOK EUROPE |  | IGTCFS-65-UNI-CELECT-PERM |  | DESCRIPTION OF EVENT ACCORDING TO COMPLAINANT: FILTER PLACED AT (B)(6) HOSPITAL ON (B)(6) 2011. CT WAS DONE IN 2012 WHICH SHOWED 2 LEGS BROKEN. THE PT HAS NOT HAD ANY TRAUMA OR SURGERY SINCE PLACEMENT. THE PT RECENTLY HAD ABDOMINAL PAIN WHEN A PHYSICIAN AT (B)(6) DID CT SCAN AND SHOWED LEG DETACHED. THE FILTER WAS RETRIEVED, BUT IT IS BELIEVED ONE LEG IS STILL IN HEAD OF PANCREAS. THE OTHER LEG WAS NOT LOCATED. IT WAS QUESTIONABLE WHEN READ LAST CT WHETHER THE OTHER LEG WAS LOCATED AT HEAD OF VERTEBRAL LEVEL. HOWEVER, IT DID NOT LOOK THAT WAY. RATHER, IT WAS CALCIUM (REALLY THICK) AND WAS NOT CONVINCED IT WAS OTHER LEG. CELECT FILTER WAS RETRIEVED WITH SHEATH AND FORCEPS. DETACHED LEG LOCATED WAS NOT RETRIEVED AND REMAINED IN PT. |  | N | Y | HEALTH PROFESSIONAL
3002808486-2014-00012 | 2014-04-28 |  | Malfunction | Manufacturer report | COOK CELECT PLATINUM NAVALIGN FEMORAL VENA CAVA FILTER SET | DTK: FILTER, INTRAVASCULAR, CARDIOVASCULAR | WILLIAM COOK EUROPE |  | IGTCFS-65-1-FEM-FT-CELECT-PT |  | DESCRIPTION OF EVENT ACCORDING TO COMPLAINANT: PATIENT CAME IN FOR AN ISSUE UNRELATED TO FILTER. IT WAS NOTED ON CT THAT 1 CONFIRMED (POTENTIALLY MULTIPLE) LEG, PENETRATED IVC WALL. THE CONFIRMED PENETRATION IS ASYMPTOMATIC AND THE FILTER WILL NOT BE RETRIEVED. THE PATIENT DID NOT REQUIRE ANY ADDITIONAL PROCEDURES DUE TO THIS OCCURRENCE. ACCORDING TO THE INITIAL REPORTER, THE PATIENT DID NOT EXPERIENCE ANY ADVERSE EFFECTS DUE TO THIS OCCURRENCE. NO PRODUCT RECEIVED FOR INVESTIGATION AS FILTER IS STILL IMPLANTED. FEW IMAGE HAVE BEEN AVAILABLE FOR THE INVESTIGATION. IMAGING CONFIRMS THAT ALL FOUR CELECT PLATINUM LEGS HAVE PERFORATED THE IVC. ALSO, IT WAS NOTED THAT ONE OF THE SECONDARY FILTER LEG IS EITHER ENTANGLED OR ONE HAS FRACTURED AND EMBOLIZED OUTSIDE THE FIELD OF VIEW. DUODENAL PERFORATION IS UNLIKELY BUT CANNOT BE EXCLUDED ON THIS IMAGING PARTICULARLY GIVEN THE HISTORY OF GI BLEED. THE CAVA IS NARROWED AT THE LEVEL OF THE HOOK AND RENAL VEINS FROM LOW INTRAVASCULAR VOLUME AND STENOSIS. FILTER PERFORATION OF THE VENA CAVA WALL IS A WELL KNOWN RISK. SEVERAL CASE REPORTS PUBLISHED IN ARTICLES, DESCRIBE FILTER PERFORATION OF THE VENA CAVA WALL. DESPITE PERFORATION THE MAJORITY ENDS UP IN SUCCESSFUL FILTER RETRIEVAL. |  | N | Y | HEALTH PROFESSIONAL
3002808486-2014-00039 | 2014-07-25 |  | Injury | Manufacturer report | COOK CELECT NAVALIGN FEMORAL VENA CAVA FILTER SET | DTK FILTER, INTRAVASCULAR, CARDIOVASCULAR | WILLIAM COOK EUROPE |  | IGTCFS-65-1-FEM-CELECT |  | DESCRIPTION ACCORDING TO COMPLAINANT: VENA CAVA FILTER DID NOT OPEN WHILE THE FILTER WAS BEING PLACED INSIDE THE PATIENT. AS FACILITY WAS CONCERNED THAT THE UNDEPLOYED FILTER MIGHT MIGRATE ANOTHER FILTER WAS DEPLOYED ADJACENT TO THE NON-FUNCTIONING FILTER. THE COMPLAINT DEVICE/FILTER STILL REMAINS INSIDE THE PATIENT. POST PROCEDURE THE PATIENT WAS FINE. |  | Y | Y | HEALTH PROFESSIONAL
3002808486-2014-00040 | 2014-08-11 | 20140710 | Other | Manufacturer report | COOK CELECT FEMORAL VENA CAVA FILTER SET | FILTER, INTRAVASCULAR, CARDIOVASCULAR | WILLIAM COOK EUROPE |  | IGTCFS-65-FEM-CELECT-PERM |  | (B)(4): UNK AS INFO WAS NOT PROVIDED. THE INVESTIGATION IS IN PROGRESS. |  | Y | N | HEALTH PROFESSIONAL
3002808486-2014-00041 | 2014-08-13 | 20140711 | Malfunction | Manufacturer report | COOK CELECT NAVALIGN FEMORAL VENA CAVA FILTER SET | FILTER, INTRAVASCULAR, CARDIOVASCULAR | WILLIAM COOK EUROPE |  | IGTCFS-65-1-FEM-CELECT |  | (B)(4). THE INVESTIGATION IS IN PROGRESS. |  | N | Y | HEALTH PROFESSIONAL
3002808486-2014-00044 | 2014-08-22 |  | Injury | Manufacturer report | COOK CELECT JUGULAR VENA CAVA FILTER SET | DTK FILTER, INTRAVASCULAR, CARDIOVASCULAR | WILLIAM COOK EUROPE |  | IGTCFS-65-JUG-CELECT |  | DESCRIPTION OF EVENT ACCORDING TO COMPLAINANT'S ATTORNEY. ON OR ABOUT (B)(6) 2007, PT WAS IMPLANTED WITH AN IVC FILTER KNOWN AS THE COOK CELECT VENA CAVA FILTER. IT IS ALLEGED THAT SOMETIME AFTER PLACEMENT, THE DEVICE FAILED, INCLUDING TILTING AND PERFORATING INTO THE VESSEL WALL. IN 2012, PT HAD IMAGING STUDIES CONDUCTED WHICH REVEALED THAT SEVERAL STRUTS OF THE FILTER HAD PERFORATED AND/OR PENETRATED HER IVC WALL WITH AT LEAST ONE STRUT PIERCING HER AORTA. FILTER REMAINS INSIDE THE PT. THE PT'S DOCTOR BELIEVE THE FILTER CAN ONLY BE REMOVED VIA AN INVASIVE AND RISKY SURGICAL PROCEDURE. |  | Y | Y | HEALTH PROFESSIONAL
4116864 | 2014-08-14 | 20140724 | No answer provided | User Facility report | COOK CELECT IVC FILTER | FILTER, INTRAVASCULAR, CARDIOVASCULAR | COOK MEDICAL | * | * |  | PATIENT PRESENTED FOR IVC FILTER REMOVAL. THE DEVICE WAS PLACED AT AN OUTSIDE HOSPITAL AND REPORTEDLY, ONE FAILED REMOVAL ATTEMPT WAS MADE AT THAT FACILITY. BECAUSE IT WAS PLACED AT ANOTHER FACILITY, WE DO NOT HAVE ACCESS TO THE FILTER'S MODEL NUMBER, SERIAL NUMBER OR LOT NUMBER.PRIOR TO REMOVAL ATTEMPT AT OUR FACILITY, SPOT FILM IMAGING OF THE FILTER SHOWED THAT THE FRACTURED/MISSING FILTER ARM WAS IN THE RIGHT LUNG BASE. THE FILTER WAS SUCCESSFULLY REMOVED INTACT EXCEPT FOR THE PREVIOUSLY NOTED FRACTURED AND EMBOLIZED FILTER ARM.MD DISCUSSED THE RISKS AND BENEFITS OF RETRIEVAL OF THE REMAINING FRAGMENT WITH THE PATIENT. PATIENT DECLINED TO HAVE REMOVAL ATTEMPT PERFORMED. |  | N | * | INVALID DATA
1820334-2014-00072 | 2014-03-07 | 20120725 | Injury | Manufacturer report | GUNTHER TULIP FEMORAL VENA CAVA FILTER SET | DTK FILTER, INTRAVASCULAR, CARDIOVASCULAR | COOK, INC. | NA | IGTCFS-65-FEM |  | THE FOLLOWING INFORMATION WAS REPORTED THROUGH SUIT FILED IN THE (B)(6) DISTRICT OF (B)(4), (B)(4) DIVISION; CASE NO. (B)(4); IT IS ALLEGED THAT ON OR ABOUT (B)(6) 2006, THE PLAINTIFF WAS INVOLVED IN A MOTOR VEHICLE ACCIDENT ON A SNOWMOBILE. IT WAS DETERMINED PLAINTIFF WOULD BE IMPLANTED WITH A TEMPORARY/RETRIEVABLE IVC FILTER KNOWN AS THE GUNTHER TULIP VENA CAVA FILTER. THE FILTER WAS INSERTED INTO THE PLAINTIFF. THERE WERE NO COMPLICATIONS AT THE TIME OF IMPLANTATION. IT IS ALLEGED THAT ON OR ABOUT (B)(6) 2011, PLAINTIFF SUFFERED FROM CHRONIC BACK PAIN AND ABDOMINAL PAIN WHICH WAS BELIEVED TO BE RELATED TO THE FILTER. ON OR ABOUT (B)(6) 2012, PLAINTIFF WAS GIVEN AN INTERIOR VENACAVOGRAM WHICH SHOWED THAT THE INFERIOR VENA CAVA THROMBOSED BELOW THE FILTER AND HAS NUMEROUS SMALL RECANALIZED VENOUS COLLATERALS. AS THERE WAS NO EVIDENCE OF THROMBUS BELOW THE FILTER AND OCCLUSION OF THE INFERIOR VENA CAVA AND BILATERAL ILIACS PRECLUDED THE POSSIBILITY OF MOBILE THROMBUS, THE DECISION WAS MADE TO REMOVE THE FILTER. ON OR ABOUT (B)(6) 2012, AN ATTEMPT WAS MADE TO REMOVE THE FILTER MECHANICALLY. THE FILTER WAS FIRMLY ADHERED TO THE INFERIOR VENA CAVA AND FIBROTIC SEGMENTS AND THEREFORE, COULD NOT BE REMOVED. ON OR ABOUT (B)(6) 2012, ANOTHER ATTEMPT WAS MADE TO REMOVE THE FILTER BUT WAS UNSUCCESSFUL. THE PLAINTIFF WAS THEN REFERRED TO ANOTHER INSTITUTION WHICH RECOMMENDED AN OPEN PROCEDURE. ON OR ABOUT (B)(6) 2012, THE PLAINTIFF UNDERWENT AN OPEN INFERIOR VENA CAVA FILTER REMOVAL WHICH WAS SUCCESSFUL. |  | Y | N | HEALTH PROFESSIONAL
3002808486-2014-00005 | 2014-03-10 | 20140208 | Injury | Manufacturer report | GUNTHER TULIP NAVALIGN FEMORAL VENA CAVA FILTER SET | DTK, FILTER, INTRAVASCULAR, CARDIOVASCULAR | WILLIAM COOK EUROPE |  | IGTCFS-65-A-FEM-TULIP |  | DURING A FILTER PLACEMENT, THE FILTER ONLY DEPLOYED HALF WAY. REPORTER STATED INJURY DUE TO DEPLOYMENT FAILURE WAS A LACERATION. THE USER HAD TO PERFORM A CUT DOWN TO COMPLETE THE PROCEDURE. CUT DOWN WAS REQUIRED TO COMPLETE THE PROCEDURE. PATIENT IS IN STABLE CONDITION. ACCORDING TO THE INITIAL REPORTER, THE PATIENT DID NOT EXPERIENCE ANY ADVERSE EFFECTS DUE TO THIS OCCURRENCE. |  | Y | N | HEALTH PROFESSIONAL
1820334-2014-00106 | 2014-03-18 | 20140220 | Injury | Manufacturer report | GUNTHER TULIP VENA CAVA FILTER RETRIEVAL SET | DTK FILTER, INTRAVASCULAR, CARDIOVASCULAR | COOK, INC. | NA | GTRS-200-RB |  | AFTER ACCESSING JUGULAR VEIN, THE PHYSICIAN ATTEMPTED TO RETRIEVE THE IVC FILTER. THE PHYSICIAN FULLY EXPANDED THE SNARE LOOP ON THE GUNTHER TULIP VENA CAVA FILTER RETRIEVAL SET; HOWEVER, DID NOT LIKE POSITIONING AND ATTEMPTED TO RETRACT THE SNARE LOOP. AS THE PHYSICIAN PULLED "PIN-VICE" BACK, THE SNARE LOOP BROKE FREE FROM THE CATHETER. THE SNARE LOOP REMAINED IN THE PATIENT UNTIL THE PHYSICIAN ACCESSED VIA FEMORAL VEIN AND RETRIEVED THE COOK SNARE LOOP WITH AN ENSNARE. AFTER THE BROKE SNARE LOOP WAS REMOVED FROM THE PATIENT, A NEW RETRIEVAL SET WAS USED TO PERFORM A TRADITIONAL IVC FILTER REMOVAL VIA JUGLAR VEIN. ACCORDING TO THE INITIAL REPORTER, THE PATIENT DID NOT EXPERIENCE ANY ADVERSE EFFECTS DUE TO THIS OCCURRENCE. |  | Y | N | HEALTH PROFESSIONAL
1820334-2014-00171 | 2014-04-21 | 20140324 | Malfunction | Manufacturer report | GUNTHER TULIP VENA CAVA FILTER RETRIEVAL SET | DTK FILTER, INTRAVASCULAR, CARDIOVASCULAR | COOK INC | NA | GTRS-200-RB |  | AN (B)(6) FEMALE PT UNDERWENT IVC FILTER RETRIEVAL ON (B)(6) 2014. WHILE ATTEMPTING TO RETRIEVE A TULIP IVC FILTER, THE SINGLE SNARE LOOP ON THE RETRIEVAL DEVICE DISCONNECTED FROM THE DEVICE INSIDE THE PTS IVC. THE PHYSICIAN USED ANOTHER RETRIEVAL DEVICE TO SUCCESSFULLY RETRIEVE THE COOK LOOP SNARE W/O HARM TO THE PT. A SECTION OF THE DEVICE DID NOT REMAIN INSIDE THE PT'S BODY. THE PT DID NOT REQUIRE ANY ADD'L PROCEDURES DUE TO THIS OCCURRENCE. ACCORDING TO THE INITIAL RPTR, THE PT DID NOT EXPERIENCE ANY ADVERSE EFFECTS DUE TO THIS OCCURRENCE. |  | N | Y | HEALTH PROFESSIONAL
1820334-2014-00223 | 2014-05-27 | 20140427 | Injury | Manufacturer report | GUNTHER TULIP FEMORAL VENA CAVA FILTER SET | DTK FILTER, INTRAVASCULAR, CARDIOVASCULAR | COOK, INC. | NA | IGTCFS-65-FEM |  | A PT UNDERWENT FILTER PLACEMENT ON (B)(6) 2104. THE LEGS OF GUNTHER TULIP FEMORAL VENA CAVA FILTER FAILED TO EXPAND AFTER BEING DEPLOYED IN THE PT, LEADING TO THE FILTER MIGRATING TO THE RIGHT ATRIUM OF THE HEART. A RETRIEVAL KIT WAS USED TO SUCCESSFULLY SNARE THE FILTER FOR REMOVAL. ANOTHER MANUFACTURER'S DEVICE WAS USED TO COMPLETE THE PROCEDURE AND NO HARM WAS DONE TO THE PT. A SECTION OF THE DEVICE DID NOT REMAIN INSIDE THE PT'S BODY. ACCORDING TO THE INITIAL REPORTER, THE PT DID NOT EXPERIENCE ANY ADVERSE EFFECTS DUE TO THIS OCCURRENCE. |  | Y | N | HEALTH PROFESSIONAL
1820334-2014-00093 | 2014-03-20 | 20140217 | Malfunction | Manufacturer report | GUNTHER TULIP VENA CAVA FILTER RETRIEVAL SET | DTK FILTER, INTRAVASCULAR, CARDIOVASCULAR | COOK, INC. | NA | GTRS-200-RB |  | THE WHITE TORQUING HANDLE BROKE OFF, DETACHING THE SNARE, MAKING THE SNARE LOOSE IN THE SHEATH. THE PHYSICIAN USED A HEMOSTAT TO REMOVE THE SNARE FROM THE SHEATH. ANOTHER / DIFFERENT SNARE WAS ALREADY OPENED. THAT SNARE WAS USED TO COMPLETE THE PROCEDURE. A SECTION OF THE DEVICE DID NOT REMAIN INSIDE THE PATIENT'S BODY. THE PATIENT DID NOT REQUIRE ANY ADDITIONAL PROCEDURES DUE TO THIS OCCURRENCE. ACCORDING TO THE INITIAL REPORTER, THE PATIENT DID NOT EXPERIENCE ANY ADVERSE EFFECTS DUE TO THIS OCCURRENCE. |  | N | Y | HEALTH PROFESSIONAL
1820334-2014-00236 | 2014-05-29 | 20140428 | Injury | Manufacturer report | GUNTHER TULIP VENA CAVA FILTER RETRIEVAL SET | DTK FILTER, INTRAVASCULAR, CARDIOVASCULAR | COOK, INC. | N/A | GTRS-200-RB |  | ON (B)(6) 2014, DURING AN IVC RETRIEVAL PROCEDURE, THE SNARE BROKE OFF WHILE ATTEMPTING TO RETRIEVE FILTER. A NEW SNARE WAS OPENED TO RETRIEVE THE SNARE AND THEN FILTER. |  | Y | N | HEALTH PROFESSIONAL
1820334-2014-00246 | 2014-06-13 | 20120713 | Injury | Manufacturer report | GUNTHER TULIP FEMORAL VENA CAVA FILTER SET | DTK FILTER, INTRAVASCULAR, CARDIOVASCULAR | COOK, INC. | N/A | IGTCFS-65-FEM |  | PER LAWSUIT FILED, IT IS ALLEGED THAT ON OR ABOUT (B)(6) 2008, PLAINTIFF PRESENTED AND WAS DIAGNOSED WITH RIGHT ILIOFEMORAL VEIN THROMBOSIS. THE ALLEGATIONS OF THEIR LAWSUIT ARE: IT WAS DETERMINED THAT A RETRIEVABLE IVC FILTER BE IMPLANTED. THE GUNTHER TULIP VENA CAVA FILTER WAS IMPLANTED. THERE WERE NO COMPLICATIONS AT THE TIME OF IMPLANTATION. ON OR ABOUT (B)(6) 2012, A CT SCAN WAS TAKEN OF THE PLAINTIFF'S ABDOMEN AND PELVIC, WHICH IS WAS FIRST DISCOVERED THAT A METALLIC STRUT OF THE FILTER PROTRUDING INTO THE THIRD PORTION OF THE DUODENUM. THE DECISION WAS MADE TO REMOVE THE IVC FILTER. ON OR ABOUT (B)(6) 2012, PLAINTIFF PRESENTED FOR SURGERY CONSISTING OF REPAIR TO THEIR INFERIOR VENA CAVA TO DUODENAL FISTULA WITH REMOVAL OF THE IVC FILTER AND INFERIOR VENA CAVA LITIGATION. DURING THE REMOVAL PROCEDURE, IT WAS NOTED THAT THERE WERE AT LEAST THREE DIFFERENT STRUTS OF THE FILTER THAT HAD PROTRUDED INTO THE DUODENUM AND CREATED A SMALL LEAK. THE FILTER WAS ALSO DISCOVERED TO BE DEEPLY EMBEDDED INTO THE WALLS OF THE INFERIOR VENA CAVA. IT WAS REPORTED THAT THE PLAINTIFF HAD REMOVAL OF THE IVC FILTER AND REPAIR OF THE IVC TO DUODENAL FISTULA. |  | Y | N | HEALTH PROFESSIONAL
1820334-2014-00248 | 2014-06-13 | 20140514 | Injury | Manufacturer report | GUNTHER TULIP VENA CAVA FILTER RETRIEVAL SET | DTK FILTER, INTRAVASCULAR, CARDIOVASCULAR | COOK, INC. | NA | GTRS-200-RB |  | THE FILTER WAS DEPLOYED IN AN UNSATISFACTORY POSITION. THE FILTER WAS REPOSITIONED THREE TIMES AND ON THE THIRD TRY IT PREMATURELY DEPLOYED IN AN UNSATISFACTORY POSITION. A REMOVAL SET WAS USED TO REMOVE THE FILTER AND PUT IN ANOTHER ONE. THE SNARE ENGAGED THE FILTER AT THE CLIP BUSHING THAT HOLDS THE LEGS TOGETHER. THE PHYSICIAN COULD NOT GET IT TO DISENGAGE. UPON LOOKING AT THE FILMS, THE ANGLE AT WHICH THE FILTER WAS DEPLOYED VS. THE ANGLE AT WHICH THE SNARE CAME THROUGH THE HEART INTO THE INFERIOR VENA CAVA WAS DIFFICULT AND IN THE OPINION OF THE DISTRICT MANAGER LED TO THE PREMATURE DEPLOYMENT AND THE SNARE GETTING CAUGHT. THE PHYSICIAN WORKED FOR A WHILE AND ENDED UP BREAKING THE SNARE AND LEAVING IT IN THE PATIENT ATTACHED TO THE FILTER. THE FILTER WAS DAMAGED, BUT STILL INTACT. THE PATIENT WAS REMOVED FROM THE TABLE WITH THE SNARE STILL ATTACHED TO THE FILTERS. THE DEVICE REMAINS IN THE PATIENT ATTACHED TO THE SNARE. |  | Y | N | HEALTH PROFESSIONAL
1820334-2014-00245 | 2014-06-12 | 20130206 | Injury | Manufacturer report | GUNTHER TULIP FEMORAL VENA CAVA FILTER SET | DTK FILTER, INTRAVASCULAR, CARDIOVASCULAR | COOK INC | N/A | IGTCFS-45-FEM |  | COMPLAINT RECEIVED VIA LAWSUIT-COMPLAINT RECEIVED BY (B)(6). THE ALLEGATIONS OF THEIR LAWSUIT ARE: ON OR ABOUT (B)(6) 2005, THE PT PRESENTED AND WAS DIAGNOSED WITH LEFT LOWER EXTREMITY DEEP VEIN THROMBOSIS. IT WAS DETERMINED THAT SHE WOULD BE IMPLANTED WITH A TEMPORARY/RETRIEVABLE IVC FILTER KNOWN AS THE GUNTHER TULIP VENA CAVA FILTER. THE FILTER WAS INSERTED INTO THE PT. THERE WERE NO COMPLICATIONS AT THE TIME OF IMPLANTATION. ON OR ABOUT (B)(6) 2011, THE PT PRESENTED FOR A CT SCAN AT WHICH POINT IT WAS DISCOVERED THAT THE LIMBS OF THE FILTER HAD PENETRATED THROUGH THE VENA CAVA INTO THE SURROUNDING RETROPERITONEAL FAT. ON AND OFF OVER THE NEXT COUPLE OF YEARS, THE PLAINTIFF EXPERIENCED SHORTNESS OF BREATH, CHEST PAINS, AND OTHER SHARP STABBING PAIN THAT WAS ASSOCIATED WITH THE PERFORATED LIMBS OF THE FILTER. THE FILTER WAS SUCCESSFULLY REMOVED IN (B)(6) 2013. NO ADD'L INFO HAS BEEN PROVIDED. |  | Y | N | HEALTH PROFESSIONAL
1820334-2014-00251 | 2014-06-18 | 20140520 | Injury | Manufacturer report | GUNTHER TULIP VENA FILTER RETRIEVAL SET | DTK FILTER, INTRAVASCULAR, CARDIOVASCULAR | COOK, INC. | NA | GTRS-200-RB |  | DURING AN IVC FILTER RETRIEVAL PROCEDURE THE SNARE SHEARED OFF IN THE PATIENT. THE PHYSICIAN STATED WAS PULLING VERY HARD AND TWISTING AND THE WELD POINT ON THE SNARE FAILED. AN EV3 ENSNARE WAS USED TO CAPTURE AND REMOVE THE PIECE OF SNARE THAT SHEARED OFF AND SUBSEQUENTLY THE FILTER WAS REMOVED AS WELL. THE PATIENT IS DOING FINE AND NO FURTHER COMPLICATIONS OR PROCEDURES ARE NECESSARY. |  | Y | N | HEALTH PROFESSIONAL
1820334-2014-00250 | 2014-06-20 | 20140520 | Injury | Manufacturer report | GUNTHER TULIP VENA CAVA FILTER RETRIEVAL SET | FILTER, INTRAVASCULAR, CARDIOVASCULAR | COOK, INC. | NA | GTRS-200-RB |  | DURING IVC DENALI FILTER RETRIEVAL, RIGHT JUGULAR ACCESS ON A (B)(6) MALE PT, AFTER ACCESS WAS GAINED AND THE COOK INTRODUCER SET WAS IN PLACE IN THE VENA CAVA, THE FILTER WAS SUCCESSFULLY SNARED. THE PHYSICIAN TIGHTENED THE SNARE AROUND THE FILTER HOOK AND PROCEEDED TO COVER THE FILTER WITH THE INNER INTRODUCER SHEATH. HE WAS ABLE TO COVER THE FILTER TO A POINT JUST ABOVE THE FEET OF THE FILTER. THE PHYSICIAN THEN TRIED TO COVER THE TRAPPED FILTER WITH THE OUTER INTRODUCER SHEATH BUT THE FILTER GOT STUCK WITH THE FEET OUTSIDE THE SHEATH. WHILE HE ATTEMPTED TO PULL THE FILTER INTO THE INTRODUCER SHEATH, THE PROXIMAL WIRE KNOB BROKE OFF. PLACING A CLAMP ON THE PROXIMAL END OF THE WIRE SNARE THE PHYSICIAN PULLED ON THE WIRE, TRYING TO PULL THE REMAINING PORTION OF THE FILTER INTO THE INTRODUCER SHEATH. WHILE PULLING WITH SOME FORCE, THE WIRE SNARE BROKE. HE PULLED THE WHOLE SYSTEM OUT TO SKIN LEVEL BUT DID NOT WANT TO YANK THE BARE FILTER OUT THROUGH THE VESSEL AND SKIN. THE PHYSICIAN WAS ABLE TO REMOVE THE GTRS SYSTEM COMPLETELY, LEAVING THE FILTER STUCK IN THE PT'S NECK, AND THEN PLACED A CORDIS BRITE TIP SHEATH OVER THE FILTER, SUTURING IT IN PLACE. A VASCULAR SURGERY CONSULT WAS CALLED. THE PT WAS TAKEN TO THE OPERATING ROOM WHERE A VASCULAR SURGEON PERFORMED A FOREIGN BODY REMOVAL PROCEDURE AND THE FILTER WAS ENTIRELY REMOVED. |  | Y | N | HEALTH PROFESSIONAL
1820334-2014-00284 | 2014-06-27 | 20140411 | Injury | Manufacturer report | GUNTHER TULIP VENA CAVA FILTER RETRIEVAL SET | DTK FILTER, INTRAVASCULAR, CARDIOVASCULAR | COOK, INC. | NA | GTRS-200-RB |  | (B)(4). EVENT EVALUATION: STILL UNDER INVESTIGATION. |  | Y | N | HEALTH PROFESSIONAL
MW5037141 | 2014-07-07 | 20100208 | Injury | Voluntary report | GUNTHER TULIP | IVC FILTER | COOK MEDICAL |  |  |  | I WAS GIVEN A COOK GUNTHER TULIP IVC FILTER IN 2010. IT WAS SUPPOSED TO BE A RETRIEVABLE TEMPORARY FILTER. HOWEVER IN (B)(6), THE FILTER WAS UNSUCCESSFUL REMOVED. I HAVE BEEN SUFFERING FROM PAIN SINCE 2010. IT HAS NOW PERFORATED THE IVC AND HIS TOUCHING MY VERTEBRA AT L4, WE ARE AWAITING THE MEANS OF GETTING IT REMOVED-HOPEFULLY THIS TIME. MY MEDICAL RECORD SHOWS THAT THE COOK GUNTHER TULIP WAS DEFECTIVE AND I AM NOW NEEDING A FOURTH SURGERY TO HAVE IT REMOVED FROM THE IVC. |  | Y | N | 0
1820334-2014-00312 | 2014-07-11 | 20120820 | Malfunction | Manufacturer report | GUNTHER TULIP FEMORAL VENA CAVA FILTER SET | DTK FILTER, INTRAVASCULAR, CARDIOVASCULAR | COOK INC | N/A | IGTCFS-65-FEM |  | (B)(4). EVENT EVALUATION: EVALUATION IS BASED ON EVENT DESCRIPTION ONLY, SINCE NO DEVICE, MED RECORDS OR IMAGING HAVE BEEN MADE AVAILABLE TO ASSIST IN THE INVESTIGATION. THE FILTER WAS IMPLANTED ON (B)(6) 2011. (B)(4) DEVICES WERE PRODUCED ON THE DEVICE LOT AND NO OTHER COMPLAINTS ON THE SAME LOT HAVE BEEN RECEIVED. THERE WERE NO RELEVANT NOTES ON THE WORK ORDER, NOR ON THE FILTER LOTS. THERE IS NO EVIDENCE TO SUGGEST THAT THIS DEVICE WAS NOT MANUFACTURED ACCORDING TO SPECIFICATIONS. THE DEVICE IS SHIPPED WITH INSTRUCTIONS FOR USE WHICH STATES, "EXCESSIVE FORCE SHOULD NOT BE USED TO RETRIEVE THE FILTER." THE EVALUATION IS BASED ON EVENT DESCRIPTION ONLY, SINCE NO DEVICE, MED RECORDS OR IMAGING HAS BEEN RETURNED. TO DATE, THERE IS NO EVIDENCE TO SUGGEST THAT THIS DEVICE WAS NOT MANUFACTURED ACCORDING TO SPECIFICATIONS. DIFFICULT FILTER REMOVAL DUE TO THE EMBEDMENT OF FILTER LEGS IN THE IVC WALL IS WELL-KNOWN RISK THAT HAS BEEN DOCUMENTED IN THE SCIENTIFIC AND MED LITERATURE. SEVERAL CASE REPORTS PUBLISHED IN ARTICLES DESCRIBE DIFFICULT FILTER REMOVAL FROM THE VENA CAVA WALL (WILLIAM T. KUO, ET AL: PHOTOTHERMAL ABLATION WITH THE EXCIMER LASER SHEATH TECHNIQUE FOR EMBEDDED INFERIOR VENA CAVA FILTER REMOVAL: INITIAL RESULTS FROM A PROSPECTIVE STUDY). WITHOUT IMAGES IT IS DIFFICULT TO COMMENT ON THE TOTAL OCCLUSION/SEVERE STENOSIS AT THE INFERIOR VENA CAVA AT THE LEVEL OF THE FEET OF THE FILTER. BASED ON THE ABOVE, THE EXACT REASON FOR THE DIFFICULT REMOVAL OF THE EMBEDDED FILTER REMAINS UNK. FURTHER RISK REDUCTION WILL NOT BE COMPLETED AT THIS TIME. THE MED BENEFIT OF PREVENTION OF RECURRENT PULMONARY EMBOLISMS OBTAINED BY PLACING AN IVC FILTER OUTWEIGHS THE RESIDUAL RISK OF THE DEVICE BEING DIFFICULT TO REMOVE. WE WILL CONTINUE TO MONITOR FOR SIMILAR COMPLAINTS AND HAVE NOTIFIED THE APPROPRIATE INTERNAL PERSONNEL. |  | Y | Y | HEALTH PROFESSIONAL
1820334-2014-00332 | 2014-08-01 | 20140629 | Injury | Manufacturer report | GUNTHER TULIP VENA CAVA FILTER RETRIEVAL SET | DTK FILTER, INTRAVASCULAR, CARDIOVASCULAR | COOK, INC. | NA | GTRS-200-RB |  | DURING A FILTER RETRIEVAL PROCEDURE ON A (B)(6) PATIENT, THE SNARES BROKE OFF OF BOTH DEVICES (1820334-2014-00331 AND 00332) DURING THE SAME PROCEDURE. THE SNARE OF ONE OR BOTH DEVICES IS STILL INSIDE THE PATIENT. IT IS NOT PRESENTLY KNOWN IF THE PATIENT WILL REQUIRE ANY ADDITIONAL PROCEDURES DUE TO THIS OCCURRENCE. ACCORDING TO THE INITIAL REPORTER, THE PATIENT DID NOT EXPERIENCE ANY ADVERSE EFFECTS DUE TO THIS OCCURRENCE. IT IS NOT KNOWN HOW OR IF THE PROCEDURE WAS COMPLETED. |  | Y | N | HEALTH PROFESSIONAL
1820334-2014-00331 | 2014-07-31 | 20140629 | Injury | Manufacturer report | GUNTHER TULIP VENA CAVA FILTER RETRIEVAL SET | DTK FILTER, INTRAVASCULAR, CARDIOVASCULAR | COOK, INC. | NA | GTRS-200-RB |  | DURING A FILTER RETRIEVAL PROCEDURE ON A (B)(6) MALE PATIENT, THE SNARES BROKE OFF OF BOTH DEVICES (1820334-2014-00331 AND 00332) DURING THE SAME PROCEDURE. THE SNARE OF ONE OR BOTH DEVICES IS STILL INSIDE THE PATIENT. IT IS NOT PRESENTLY KNOWN IF THE PATIENT WILL REQUIRE ANY ADDITIONAL PROCEDURES DUE TO THIS OCCURRENCE. ACCORDING TO THE INITIAL REPORTER, THE PATIENT DID NOT EXPERIENCE ANY ADVERSE EFFECTS DUE TO THIS OCCURRENCE. IT IS NOT KNOWN HOW OR IF THE PROCEDURE WAS COMPLETED. |  | N | N | HEALTH PROFESSIONAL
1820334-2014-00391 | 2014-08-21 | 20140708 | Injury | Manufacturer report | GUNTHER TULIP JUGULAR VENA CAVA FILTER ET | DTK FILTER, INTRAVASCULAR, CARDIOVASCULAR | COOK, INC. | NA | IGTCFS-65-JUG |  | A (B)(6)-YEAR OLD MALE PATIENT PRESENTED WITH A MEDICAL HISTORY OF CURRENT PULMONARY EMBOLISM (PE) AND HYPERTENSION TO PARTICIPATE IN A STUDY UNDERWENT INFERIOR VENA CAVA FILTER REMOVAL ON (B)(6) 2014. CORE LAB REPORTED THE FOLLOWING FROM THE FILTER PRE-RETRIEVAL CT: THE NUMBER OF FILTER LEGS WITH IVC WALL INTERACTION WAS EIGHT WITH GRADE ZERO, ONE WITH GRADE ONE, ONE WITH GRADE TWO, AND TWO WITH GRADE THREE. THE GRADE THREE INTERACTION AFFECTED THE DUODENUM. THERE WAS NO HEMORRHAGE, HEMATOMA, OR OTHER CLINICAL FINDING. THE EVENT IS THE TWO FILTER LEGS WITH GRADE 3 INTERACTION. GRADE 3 INTERACTION IS DEFINED AS FOLLOWS IN APPENDIX C OF THE PROTOCOL (P. 53): "FILTER STRUT IS TOUCHING, IMPRESSING, OR PERFORATING ANOTHER ORGAN (E.G., LIVER, BOWEL, AORTA, PSOAS MUSCLE, VERTEBRAL BODY, LYMPH NODES)." ON (B)(6) 2014 (ONE DAY PRE-PROCEDURE), THE PATIENT WAS ADMITTED TO THE HOSPITAL. ON THIS DAY, A COMMUTED TOMOGRAPHY (CT) WAS PERFORMED. THE CT REVEALED THE PE. THE PATIENT REMAINS IN THE STUDY AND IS IN FOLLOW-UP POST FILTER RETRIEVAL. ON THE DAY OF THE PROCEDURE, THE PATIENT WAS (B)(6). THE PATIENT WAS ON PROPHYLACTIC LOW MOLECULAR WEIGHT (LMW) HEPARIN. THE PRE-PLACEMENT ULTRASOUND WAS NEGATIVE FOR THROMBUS IN THE INFERIOR VENA CAVA (IVC) AND THE PELVIC VEINS. THERE WAS THROMBUS IN THE LEFT LOWER EXTREMITY VEINS. CORE LAB ANALYSIS OF THE PRE-PLACEMENT ULTRASOUND REVEALED NO THROMBUS IN THE IVC. THERE WAS OCCLUDING THROMBUS IN THE LEFT PELVIC VEINS AND THE LEFT LOWER EXTREMITY VEINS. THE PRIMARY REASON FOR THE FILTER REPLACEMENT WAS THE CURRENT DEEP VEIN THROMBOSIS (DVT) WITH THE ADDITIONAL REASON OF FAILURE OF ANTICOAGULATION. THE FILTER PLACEMENT WAS ANTICIPATED TO BE TEMPORARY AND WAS ACCOMPLISHED WITH THE USE OF A VENACAVAGRAM. THE IVC DIAMETER AT THE INTENDED FILTER LOCATION WAS 18.5 MM. THERE WAS NO IVC ANATOMIC ANOMALY OR PRESENCE OF THROMBUS IN THE IVC PRIOR TO FILTER PLACEMENT. USING THE RIGHT INTERNAL JUGULAR VEIN ACCESS SITE, A GUNTHER TULIP FILTER WAS DEPLOYED AT THE INTENDED LOCATION OF THE IVC INFRARENAL SITE. THERE WAS NO EVIDENCE OF FILTER FRACTURE, DEFORMATION, MIGRATION, OR FILTER TILT. THE FILTER WAS NEITHER DEPLOYED PREMATURELY NOR DID THE FILTER JUMP UPON DEPLOYMENT. THERE WAS NO EXTRAVASATION OF CONTRAST OR FILTER LEG(S) APPEARING OUTSIDE THE COLUMN OF CONTRAST AFTER FILTER PLACEMENT. CORE LAB ANALYSIS OF THE PLACEMENT PROCEDURE VENACAVAGRAM REVEALED THE IVC DIAMETER AT THE INTENDED FILTER LOCATION WAS 18.0 MM. THE FILTER WAS PLACED AT THE IVC INFRARENAL SITE. THERE WAS NO EVIDENCE OF FILTER FRACTURE, DEFORMATION, OR MIGRATION. THE FILTER TILT IN THE AP VIEW WAS 1.2 DEGREES. CORE LAB COMMENTED, "FILTER REPOSITIONED DURING PROCEDURE WITH EXCELLENT RESULTS." THE POST-PROCEDURE X-RAY REVEALED NO EVIDENCE OF FILTER FRACTURE, EMBOLIZATION, OR MIGRATION. FILTER TILT WAS ZERO DEGREES. CORE LAB ANALYSIS REVEALED NO EVIDENCE OF FILTER FRACTURE, DEFORMATION, EMBOLIZATION, OR MIGRATION. FILTER TILT WAS NOT ASSESSED. ON (B)(6) 2014 (THREE DAYS POST-PROCEDURE), THE PATIENT WAS DISCHARGED. NO PORTION OF THE HOSPITALIZATION WAS DIRECTLY RELATED TO THE PLACEMENT PROCEDURE. THE PATIENT WAS DISCHARGED ON PROPHYLACTIC LOW MOLECULAR WEIGHT (LMW) HEPARIN. ON (B)(6) 2014 (34 DAYS POST-PROCEDURE), THE ONE MONTH CLINICAL ASSESSMENT WAS COMPLETED. A RETRIEVAL PROCEDURE WAS PLANNED AT THIS TIME. THE PATIENT WAS TAKING WARFARIN AND HAD NO SIGNIFICANT MEDICAL PROBLEMS OR HOSPITALIZATIONS SINCE THE LAST VISIT. ON (B)(6) 2014 (48 DAYS POST-PROCEDURE), THE PRE-RETRIEVAL ULTRASOUND WAS COMPLETED AND WAS NEGATIVE TOR THROMBUS IN THE IVC AND THE TOWER EXTREMITY VEINS. THERE WAS THROMBUS IN THE LEFT PELVIC VEINS. CORE LAB ANALYSIS NOTED NO THROMBUS IN THE IVC. THERE WAS PRE-EXISTING, IMPROVED, BUT OCCLUDING THROMBUS IN THE LEFT PELVIC VEINS. THERE WAS PRE-EXISTING, IMPROVED, NON-OCCLUDING THROMBUS IN THE LEFT LOWER EXTREMITY. CORE LAB COMMENTED, "SUBTOTAL LEFT FEMORAL VEIN OCCLUSION. RECANALIZATION IS OCCURRING. OVERALL, INTERVAL IMPROVEMENT IN THROMBUS IN LEFT SIDED PELVIC AND LOWER EXTREMITY VEINS." ON (B)(6) 2014 (67 DAYS POST-PROCEDURE), THE PRE-RETRIEVAL CT AND X-RAY WERE PERFORMED. THE CT REVEALED NO EVIDENCE PE OR FILER EMBOLIZATION, AND THE HIGHEST GRADE OF FILTER LEG INTERACTION WITH THE IVC WALL WAS TWO. CORE LAB ANALYSIS REVEALED EVIDENCE OF THE PRE-EXISTING LEFT PE, BUT NO FILTER EMBOLIZATION. FILTER TILT IN THE SAGITTAL PLANE WAS 4.25 DEGREES AND IN THE CORONAL PLANE WAS 5.05 DEGREES. THE NUMBER OF FILTER LEGS WITH JVC WALL INTERACTION WAS EIGHT WITH GRADE ZERO, ONE WITH GRADE ONE, ONE WITH GRADE TWO, AND TWO WITH GRADE THREE. THE GRADE THREE INTERACTION AFFECTED THE DUODENUM. THERE WAS NO HEMORRHAGE, HEMATOMA, OR OTHER CLINICAL FINDING. THE X-RAY REVEALED NO EVIDENCE OF FILTER FRACTURE, EMBOLIZATION, OR MIGRATION. THERE IS NO INFORMATION ON FILTER TILT. CORE LAB ANALYSIS REVEALED NO EVIDENCE OF FILTER FRACTURE, DEFORMATION, EMBOLIZATION, OR MIGRATION. FILTER TILT WAS NOT ASSESSED. CORE LAB COMMENTED, "THIS IS A PA AND LAT CHEST, NOT AN ABDOMEN. ONLY THE SUPERIOR PORTION OF THE FILTER IS VISIBLE." CORE LAB ALSO NOTED, "QUESTIONS ON FILTER MIGRATION, DEFORMATION, AND FRACTURE WERE OBTAINED BY ANALYZING THE CT OF THE ABDOMEN FROM (B)(6) 2014. AN ABDOMINAL RADIOGRAPH WAS NOT SUBMITTED." ON (B)(6) 2014 (70 DAYS POST-PROCEDURE), THE ENDOVASCULAR RETRIEVAL PROCEDURE WAS ACCOMPLISHED SUCCESSFULLY USING A GUNTHER TULIP RETRIEVAL SET. NO FILTER LEGS APPEARED OUTSIDE THE COLUMN OF CONTRAST BEFORE THE RETRIEVAL, AND NO THROMBUS WAS PRESENT IN THE FILTER. CORE LAB ANALYSIS OF THE RETRIEVAL PROCEDURE REVEALED NO THROMBUS PRESENT IN THE FILTER. THERE WAS FILTER LEG(S) OUTSIDE THE COLUMN OF CONTRAST BEFORE RETRIEVAL. FILTER TILT PRIOR TO RETRIEVAL IN THE AP VIEW WAS 4.0 DEGREES. FILTER TILT IN THE LAT AND OBLIQUE VIEWS WERE NOT ASSESSED. THERE WAS NO EXTRAVASATION OF CONTRAST AFTER THE RETRIEVAL. |  | Y | N | HEALTH PROFESSIONAL
1820334-2014-00327 | 2014-07-22 |  | Malfunction | Manufacturer report | GUNTHER TULIP VENA CAVA FILTER RETRIEVAL SET | FILTER, INTRAVASCULAR, CARDIOVASCULAR | COOK, INC. | NA | GTRS-200-RB |  | EVENT EVAL: STILL UNDER INVESTIGATION. |  | N | Y | HEALTH PROFESSIONAL
1820334-2014-00328 | 2014-07-22 | 20140624 | Malfunction | Manufacturer report | GUNTHER TULIP VENA CAVA FILTER RETRIEVAL SET | FILTER, INTRAVASCULAR, CARDIOVASCULAR | COOK, INC. | NA | GTRS-200-RB |  | A PT UNDERWENT IVC FILTER RETRIEVAL ON (B)(6) 2014. THE RETRIEVAL KIT WOULD NOT COLLAPSE FILTER. THE PHYSICIAN FORCIBLY REMOVED FILTER WITH A CATHETER. A SECTION OF THE DEVICE DID NOT REMAIN INSIDE THE PT'S BODY. THE PT DID NOT REQUIRE ADD'L PROCEDURES DUE TO THIS OCCURRENCE. ACCORDING TO THE INITIAL REPORTER, THE PT DID NOT EXPERIENCE AND ADVERSE EFFECTS DUE TO THIS OCCURRENCE. |  | N | Y | HEALTH PROFESSIONAL

## Monthly Summary
year_month | count
2014-03 | 6
2014-04 | 3
2014-05 | 6
2014-06 | 9
2014-07 | 10
2014-08 | 6

## Event Types
Event Type | Count
Injury | 24
Malfunction | 13
Death | 1
Other | 1
No answer provided | 1

## Brand Names
Brand Name | Count
GUNTHER TULIP VENA CAVA FILTER RETRIEVAL SET | 11
GUNTHER TULIP FEMORAL VENA CAVA FILTER SET | 5
COOK CELECT NAVALIGN FEMORAL VENA CAVA FILTER SET | 3
COOK CELECT FEMORAL AND JUGULAR VENA CAVA FILTER | 3
COOK CELECT FEMORAL VENA CAVA FILTER SET | 3
COOK CELECT NAVALIGN JUGULAR & FEMORAL VENA CAVA FILTER SET | 2
COOK CELECT PLATINUM NAVALIGN FEMORAL VENA CAVA FILTER SET | 2
CELECT | 1
GUNTHER TULIP | 1
GUNTHER TULIP VENA FILTER RETRIEVAL SET | 1
GUNTHER TULIP NAVALIGN FEMORAL VENA CAVA FILTER SET | 1
FEM-CELECT-PEM | 1
COOK CELECT IVC FILTER | 1
COOK CELECT JUGULAR VENA CAVA FILTER SET | 1
COOK FILTER CELECT UNI | 1
CELECT IVC FILTER | 1
NAVALIGN COOK CELECT JUGULAR VENA CAVA FILTER SET | 1
GUNTHER TULIP JUGULAR VENA CAVA FILTER ET | 1

## Case Metadata
Case Name | MDL Number | Filing Date | Court | Status | Settlement | Data Window Start | Data Window End | Total Records | Search Queries Used
Cook Medical IVC Filters | 2570 | 2014-08-28 | S.D. Indiana | Active | N/A | 2014-02-26 | 2014-08-28 | 40 | 3