## All Reports
report_number | date_received | date_of_event | event_type | report_source_code | brand_name | generic_name | manufacturer_d_name | model_number | catalog_number | device_class | event_description | manufacturer_narrative | adverse_event_flag | product_problem_flag | device_operator | manufacturer_g1_name
2649622-2008-00831 | 2008-02-20 | 20071224 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THE LEAD WAS EXPERIENCING HIGH IMPEDANCE AND OVERSENSING. IT WAS EXPLANTED AND REPLACED. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | N | LAY USER/PATIENT
2649622-2008-00829 | 2008-02-07 | 20070118 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00842 | 2008-02-25 | 20071221 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-00843 | 2008-02-25 | 20080201 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT ALERT TONE SOUNDED DUE TO HIGH SVC IMPEDANCE OF GREATER THAN 200 OHMS. REVIEW OF SAVE TO DISK DATA SHOWED SVC IMPEDANCE RANGING FROM 188 OHMS TO NEARLY 6000 OHMS OVER THE PAST TWO MONTHS. IT WAS ALSO NOTED THAT THE PATIENT HAD TWIDDLERS SYNDROME. THE LEAD WAS REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED REPORTING THAT THAT THERE WAS ALSO NOISE ON THE PACE/SENSE PORTION OF THE LEAD, AND THE LEAD WAS EXPLANTED. |  | N | Y | LAY USER/PATIENT
2649622-2008-00844 | 2008-02-28 | 20080212 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES FOUND; PROXIMAL SEGMENT RETURNED AND ANALYZED. IT WAS REPORTED THAT THE PATIENT RECEIVED INAPPROPRIATE SHOCKS, THERE WAS OVERSENSING, HIGH IMPEDANCE OF GREATER THAN 3000 OHMS, AN APPARENT LEAD FRACTURE, AND A LEAD DISLODGEMENT. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ACTUAL DEVICE INVOLVED IN INCIDENT WAS EVALUATED ELECTRICAL TESTS PERFORMED, MECHANICAL TESTS PERFORMED, VISUAL EXAMINATION DEVICE PERFORMED ACCORDING TO SPECIFICATIONS NO CONCLUSION CAN BE DRAWN DISLODGED IMPEDANCE, HIGH LEAD(S), FRACTURE OF OVERSENSING SHOCK, INAPPROPRIATE. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00846 | 2008-03-04 | 20070124 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THE LEAD WAS EXPLANTED AND REPLACED DUE TO FRACTURE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION RECEIVED STATES THAT A LAWSUIT ALLEGES THE PATIENT RECEIVED INAPPROPRIATE SHOCKS AND "SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH, ECONOMIC LOSSES AND OTHER DAMAGES" DUE TO THE LEAD. |  | Y | Y | LAY USER/PATIENT
1010517 | 2008-02-07 | 20080114 | Malfunction | User Facility report | SPRINT FIDELIS | LEAD, ICD | MEDTRONIC INC. | LFJ159302V | * |  | DUE TO THE PATIENT'S PAST HISTORY OF A HEART BLOCK AND PACER DEPENDENCY THE PHYSICIAN REPLACED THE RECALLED LEAD.====================== MANUFACTURER RESPONSE FOR LEAD, ICD, SPRINT FIDELIS======================KNOWN RECALL |  | N | Y | INVALID DATA
1012109 | 2008-02-15 | 20080207 | Malfunction | User Facility report | SPRINT FIDELIS | LEAD, ICD | MEDTRONIC INC. | 6949-65 | * |  | PATIENT PRESENTED TO THE ED WITH SEVERAL SHOCKS THIS MORNING. THE MEDTRONIC REPRESENTATIVE INTERROGATED THE DEVICE AND FOUND THAT THE RV LEAD WAS FRACTURED RESULTING IN INAPPROPRIATE SHOCKS. |  | N | Y | INVALID DATA
1012988 | 2008-02-25 | 20080213 | Malfunction | User Facility report | FIDELIS LEAD | LEAD, ICD | MEDTRONIC, INC. | 6949 | * |  | THIS PATIENT, WITH A HISTORY OF HYPERTROPHIC CARDIOMYOPATHY, IS FOLLOWED BY A CARDIOLOGIST IN BOSTON. THREE YEARS AGO, HE HAD A DUAL-CHAMBER ICD IMPLANTED WITH FIDELIS LEAD. THIS FRACTURED, RESULTING IN MULTIPLE SHOCKS, WHICH WERE ONLY CONTROLLED BY REMOVAL OF THE DEVICE. THE LEAD WAS RETAINED AND ANOTHER FIDELIS LEAD IMPLANTED. THIS LEAD WAS SUBSEQUENTLY RECALLED. THE PATIENT WAS BROUGHT TO THE OPERATING ROOM WHERE STANDARD MONITORING AND GENERAL ANESTHESIA WERE SAFELY AND UNEVENTFULLY INDUCED.THE SURGEON FOUND A CAPPED, ABANDONED, AND RETAINED FIDELIS LEAD IMPLANTED ORIGINALLY THREE YEARS AGO, AND HE FOUND ANOTHER FIDELIS LEAD WITHIN THE HEADER OF THE DEVICE. THIS WAS IMPLANTED TWO YEARS LATER. BOTH OF THESE LEADS WERE FREED AND DISSECTED DOWN TO THEIR INSERTION SITES. THE OLDER OF THE TWO LEADS, WAS REMOVED IN ITS ENTIRETY WITHOUT DIFFICULTY. THE NEWER OF THE TWO LEADS, COULD NOT BE REMOVED IN THIS FASHION. ACCORDINGLY, THE LEAD WAS AMPUTATED AND A #1 LASER LOCKING STYLET WAS PASSED ON THE LENGTH OF THE LEAD. USING A LASER, THE SURGEON WAS ABLE TO REMOVE THIS LEAD IN ITS ENTIRETY WITHOUT DIFFICULTY. |  | N | Y | INVALID DATA
2649622-2008-00967 | 2008-03-10 | 20071211 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2182208-2008-00049 | 2008-03-10 | 20070306 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC, INC. | 6931 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00956 | 2008-03-10 | 20071116 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD IN SEGMENTS RETURNED FOR ANALYSIS. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00957 | 2008-03-10 | 20071119 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | IT WAS REPORTED THAT OVERSENSING WAS NOTED. THE PATIENT DID NOT RECEIVE ANY INAPPROPRIATE THERAPIES. LEAD REPLACEMENT WAS PLANNED. FOLLOW-UP INFORMATION RECEIVED INDICATES THE LEAD IS STILL IN USE. IT WAS SUCCESSFULLY REPOSITIONED ON 20-NOV-2007, BUT WAS NOT REPLACED. THE PATIENT NOW HAS ROUTINE FOLLOW-UP EVERY THREE MONTHS. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-00958 | 2008-03-10 | 20071120 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-00959 | 2008-03-10 | 20051111 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00961 | 2008-03-10 | 20071120 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE PATIENT ALLEGES THAT ONE MONTH AFTER IMPLANT, HE WENT TO SEE THE DOCTOR, AND ONE OF THE LEADS WAS "NOT DOING RIGHT." THE PATIENT'S WIFE ALLEGED THE LEAD WAS BROKEN. A "CAR ACCIDENT" WAS MENTIONED, WITH NO FURTHER INDICATION OF WHEN IT OCCURRED. THE LEAD WAS REPLACED. ADDITIONAL INFORMATION WAS SUBSEQUENTLY RECEIVED REPORTING THAT THE PATIENT PRESENTED TO THE EMERGENCY ROOM WITH CONGESTIVE HEART FAILURE. DEVICE INTERROGATION REVEALED HIGH THRESHOLDS AND VARIABLE R-WAVE SENSING. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-00963 | 2008-03-10 | 20071119 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00964 | 2008-03-10 | 20071204 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT WAS SHOCKED 24 TIMES. THE LEAD WAS REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. A COMPANY REPRESENTATIVE REPORTED THAT THE LEAD CRACKED AND LEAD MALFUNCTION WAS REPORTED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00955 | 2008-03-10 | 20071207 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00965 | 2008-03-10 | 20071024 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00949 | 2008-03-10 | 20071206 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE PHYSICIAN STATED TACHY THERAPIES WERE TURNED OFF BECAUSE "THERE'S NO INDICATION FOR TACHY THERAPIES ANYMORE." PHYSICIAN ALSO STATED THAT THREE SHOCKS WERE DELIVERED WHICH HE DEEMED INAPPROPRIATE DUE TO OVERSENSING, AND THERE WAS AN APPARENT LEAD FRACTURE. THE LEAD WAS CAPPED AND REPLACED WITH A BRADY LEAD. NO OTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00950 | 2008-03-10 | 20071121 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT RV PACING IMPEDANCE VARIED WITH AUTOMATIC TESTING, FROM GREATER THAN 2500 OHMS TO LESS THAN 200 OHMS, SINCE IMPLANT, TRIGGERING THE PATIENT ALERT. HVB & SVC IMPEDANCES WERE STABLE. THE LV IMPEDANCE ALSO VARIED, EVEN THOUGH IT WAS PROGRAMMED LV TIP TO RV COIL. WITH MANUAL TESTING, IMPEDANCES WERE ALWAYS IN RANGE. NO SENSING OR THRESHOLD ISSUES WERE NOTED. THE DEVICE AND LEAD WERE BOTH REPLACED, SINCE IT WAS UNABLE TO BE DETERMINED IF THE LEAD OR DEVICE WAS THE CAUSE OF THE EVENT. THE PATIENT'S WIFE CALLED A WEEK AFTER REPLACEMENT, UPSET AND CONCERNED, STATING HER HUSBAND "MAY DIE" AFTER HAVING SURGERY TO REPLACE A "DEFECTIVE LEAD" AND DEVICE. A FAMILY FRIEND CALLED 21 DAYS AFTER REPLACEMENT TO REPORT THE PATIENT WENT INTO A COMA AFTER "SURGERY BECAUSE OF YOUR PROBLEM DEVICE" AND HAD EXPIRED 11 DAYS AFTER REPLACEMENT. THE DEVICE AND LEAD WERE SUBSEQUENTLY RETURNED WITH REPORT OF SENSING DIFFICULTY, UNDERSENSING, AND FRACTURE. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00951 | 2008-03-10 | 20071205 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00952 | 2008-03-10 | 20071113 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00947 | 2008-03-10 | 20071219 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT RECEIVED 6 INAPPROPRIATE THERAPIES DUE TO OVERSENSING AND NOISE. IT WAS ALSO NOTED THAT THE IMPEDANCE HAD INCREASED FROM 400 OHMS TO 1088 OHMS. THE LEAD WAS SUBSEQUENTLY EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00954 | 2008-03-10 | 20071207 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | N | LAY USER/PATIENT
2649622-2008-00936 | 2008-03-10 | 20070921 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED BY THE PATIENT THAT THE LEAD WAS DOUBLE COUNTING. HE HAD SURGERY TO CORRECT THE "FAULTY" LEAD. THE PATIENT WAS NOT AWARE IF THE LEAD WAS REPLACED. ADDITIONAL INFORMATION FROM THE HCP (HEALTH CARE PROFESSIONAL) REPORTED THE PATIENT HAD BEEN EXPERIENCING SHOCKS, DUE TO OVERSENSING, DURING STRENUOUS ACTIVITY. THE LEAD WAS CAPPED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00937 | 2008-03-10 | 20071106 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED BY THE PATIENT'S DAUGHTER THAT THE PATIENT COMPLAINED OF CHEST PAINS AND WENT IN TO HAVE HER DEVICE CHECKED. THE TECHNICIAN INDICATED THAT THE DEVICE HAD DELIVERED THERAPIES THREE TIMES. THE PATIENT WAS IN THE HOSPITAL WHEN THE DEVICE DELIVERED ANOTHER THERAPY AND HER HEART STOPPED RIGHT AFTER. THE PATIENT PASSED AWAY 5 DAYS AFTER GOING IN TO HAVE HER DEVICE CHECKED. THE DAUGHTER IS QUESTIONING IF THE LEAD HAS ANYTHING TO DO WITH THE PATIENT'S DEATH. ADDITIONAL INFORMATION RECEIVED FROM THE HCP REPORTS THE PATIENT DID HAVE SOME VT, THEN SUSTAINED VT/VF REQUIRING APPROPRIATE SHOCKS. THE CAUSE OF DEATH WAS SEPTIC SHOCK RESULTING IN ASYSTOLE. THERE WAS NORMAL DEVICE/SYSTEM FUNCTION AND THE DEATH WAS NOT DEVICE OR LEAD RELATED. |  | Y | N | LAY USER/PATIENT
2649622-2008-00938 | 2008-03-10 | 20071121 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00939 | 2008-03-10 | 20070309 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | AN ATTORNEY ALLEGED THAT THE PATIENT HAS BEEN "REPEATEDLY HOSPITALIZED TO CORRECT ITS DEFECTIVE LEADS." A LAWSUIT FURTHER ALLEGED THAT THE LEAD DISLODGED AND WAS "DEFECTIVE", AND THAT THE PATIENT WAS HOSPITALIZED DURING THE REPLACEMENT SURGERIES. THE LEAD WAS REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-00940 | 2008-03-10 | 20071119 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT, WHO WAS PACER DEPENDENT, PRESENTED TO THE ER (EMERGENCY ROOM) AFTER EXPERIENCING A SYNCOPAL EPISODE AND FALLING, SUFFERING A HEAD CONTUSION. DEVICE INTERROGATION INDICATED NOISE CORRESPONDING TO THE PATIENT'S FALL. A PAUSE IN PACING OUTPUT, DUE TO OVERSENSING, WAS ABLE TO BE REPRODUCED WHEN THE PATIENT RAISED HER ARM ABOVE HER BODY. DURING THE REPLACEMENT PROCEDURE, MULTIPLE PAUSES IN PACING, DUE TO OVERSENSING, WERE OBSERVED, AGAIN CAUSING ASYSTOLE. THESE WERE INDUCED BY MANIPULATION OF THE POCKET DURING IMPLANT PREPARATION. A FRACTURE WAS NOTED ON FLUOROSCOPY. WHEN THE POCKET WAS OPENED, THE IMPEDANCE BECAME HIGH, RISING TO GREATER THAN 4000 OHMS. THE LEAD WAS REPLACED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00941 | 2008-03-10 | 20071121 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00942 | 2008-03-10 | 20071102 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00943 | 2008-03-10 | 20071126 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00944 | 2008-03-10 |  | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00945 | 2008-03-10 |  | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00946 | 2008-03-10 | 20071210 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00928 | 2008-03-10 | 20071119 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00929 | 2008-03-10 | 20071213 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00930 | 2008-03-10 | 20060222 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00932 | 2008-03-10 | 20071106 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED BY THE PATIENT THAT THE LEAD BROKE, AND HE WAS SHOCKED TWENTY TIMES. THE LEAD WAS REPLACED. INFORMATION WAS SUBSEQUENTLY RECEIVED FROM AN ATTORNEY ALLEGING "PLAINTIFF HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, AND ECONOMIC LOSSES AND CONSEQUENTIAL DAMAGES..." NO SPECIFIC DETAIL REGARDING ALLEGED INJURIES WAS RECEIVED, AND NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00933 | 2008-03-10 | 20071026 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE PATIENT REPORTED THAT HIS SPRINT FIDELIS LEAD WAS REMOVED. UPON FOLLOW-UP, IT WAS DETERMINED THAT THERE WAS RISING RV IMPEDANCE AND VENTRICULAR OVERSENSING, AND THAT THE LEAD WAS REMOVED FOR THESE REASONS AS WELL AS THE FIDELIS RECALL. FURTHER ALLEGED THAT THE PATIENT "EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING" AND A LEAD FRACTURE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00934 | 2008-03-10 | 20071120 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-00935 | 2008-03-10 | 20071114 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE PATIENT'S WIFE REPORTED THAT THE LEAD WAS REPLACED BECAUSE IT WAS "BAD." NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION RECEIVED FROM THE PHYSICIAN REPORTED THE LEAD WAS REPLACED DUE TO 2 SEPARATE IMPEDANCE SPIKES TO OVER 1000 OHMS. |  | N | Y | LAY USER/PATIENT
2649622-2008-00927 | 2008-03-10 | 20071107 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-00917 | 2008-03-10 | 20071119 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-00918 | 2008-03-10 | 20071114 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00919 | 2008-03-10 | 20071110 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00920 | 2008-03-10 | 20071126 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00922 | 2008-03-10 | 20071203 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00923 | 2008-03-10 | 20071101 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00924 | 2008-03-10 | 20071102 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00925 | 2008-03-10 | 20071105 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00926 | 2008-03-10 | 20071106 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT AT FOLLOW-UP, THE LEAD SHOWED AN INCREASING IMPEDANCE AND SIC (SENSING INTEGRITY COUNTER), INDICATIVE OF OVERSENSING. LEAD FRACTURE AND INAPPROPRIATE SHOCKS WERE ALSO NOTED. THE LEAD WAS REPLACED. THE REPLACEMENT PROCEDURE WAS REPORTED TO BE SUCCESSFUL, AND NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00916 | 2008-03-10 | 20061203 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00908 | 2008-03-10 | 20071106 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00909 | 2008-03-10 | 20070410 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00910 | 2008-03-10 | 20071023 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED |  | Y | Y | LAY USER/PATIENT
2649622-2008-00911 | 2008-03-10 | 20071025 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00912 | 2008-03-10 | 20071221 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00914 | 2008-03-10 | 20071030 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE LEAD WAS RETURNED AND ANALYZED AND SUBSEQUENTLY TESTED OUT OF SPECIFICATION. A LAWSUIT ALLEGED THAT THE PATIENT SUFFERED PHYSICAL AND OTHER INJURY AS A RESULT OF THE RECALLED LEAD, AND THAT THE PATIENT "EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY REQUIRING MEDICAL INTERVENTION." NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00904 | 2008-03-10 | 20061212 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00905 | 2008-03-10 | 20071004 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00906 | 2008-03-10 | 20071023 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-00895 | 2008-03-10 | 20060530 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00896 | 2008-03-10 | 20071101 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00900 | 2008-03-10 | 20070319 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00902 | 2008-03-10 | 20071107 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00894 | 2008-03-10 | 20071129 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: DISTAL CONDUCTOR FRACTURED; FULL LEAD IN SEGMENTS RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00890 | 2008-03-10 | 20080114 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00892 | 2008-03-10 |  | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00886 | 2008-03-10 | 20071204 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE PATIENT PRESENTED WITH HIGH PACING IMPEDANCE. THE IMPEDANCE TREND REVEALED MINOR FLUCTUATIONS BEGINNING 15 MONTHS PRIOR. THE LEAD WAS REPLACED. ADDITIONAL INFORMATION RECEIVED REPORTED THE LEAD HAD SENSING DIFFICULTY. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-00887 | 2008-03-10 | 20080116 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00888 | 2008-03-10 | 20071217 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THROUGH TECHNICAL SERVICES THAT THE SVC IMPEDANCE HAD BEEN INFINITE AND THEN DECLINED TO THE 1700 OHM RANGE. THE PATIENT LATER REPORTED SWEATING AND FEELING VERY TIRED. HE SAID THE LEADS WERE FRACTURED. BOTH LEADS WERE SUBSEQUENTLY REPLACED. ADDITIONAL INFORMATION WAS LATER RECEIVED IN A LAWSUIT ALLEGING THAT THE PATIENT "EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY REQUIRING MEDICAL INTERVENTION INCLUDING BUT NOT LIMITED TO SURGERY, REMOVAL AND/OR REPLACEMENT OF THE DEFECTIVE SPRINT FIDELIS LEAD" AND THAT THE PATIENT SUFFERED "PHYSICAL AND OTHER INJURY". NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00889 | 2008-03-10 | 20070810 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00882 | 2008-03-10 | 20071205 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | IT WAS REPORTED THAT THE RV LEAD WAS EXPLANTED DUE TO VARYING HIGH LEAD IMPEDANCE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ATTORNEY ALLEGES THE PLAINTIFF HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, AND SEVERE EMOTIONAL DISTRESS AS A RESULT OF THE LEAD. |  | Y | N | LAY USER/PATIENT
2649622-2008-00884 | 2008-03-10 | 20071214 | Injury | Manufacturer report | SPRINT FIDELIS | DEFIBRILLATOR, AUTOMATIC IMPLANTABLE CARDIOVERTER | MEDTRONIC PUERTO RICO, INC. | 6949 | 6949 |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-00880 | 2008-03-10 | 20071022 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. DETAILED ANALYSIS OF THE LEAD WAS NOT PERFORMED AT THIS TIME DUE TO PENDING LITIGATION. THEREFORE, WE ARE UNABLE TO FULLY EVALUATE THE REPORTED EVENT. EVALUATION SUMMARY: NO ANOMALIES WERE FOUND. |  | N | N | LAY USER/PATIENT
2649622-2008-00875 | 2008-03-10 | 20071204 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00877 | 2008-03-10 | 20071005 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00878 | 2008-03-10 | 20070917 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00872 | 2008-03-10 | 20070829 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THAT THE PATIENT RECEIVED MULTIPLE INAPPROPRIATE SHOCKS DUE TO NOISE. THE LEAD WAS REPLACED. THE LEAD WAS INSPECTED AFTER EXPLANT AND APPEARED TO HAVE SOME IRREGULARITY OF THE EXTERNAL PLASTIC NEAR THE SITE OF THE TIE DOWNS. ADDITIONAL INFORMATION WAS LATER RECEIVED FROM AN ATTORNEY ALLEGING THE PATIENT "EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY REQUIRING MEDICAL INTERVENTION INCLUDING, BUT NOT LIMITED TO, SURGERY, REMOVAL AND/OR REPLACEMENT" OF THE DEFECTIVE LEAD. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00871 | 2008-03-10 | 20071206 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE RV LEAD WAS EXPLANTED DUE TO INCONSISTENT HIGH SVC COIL IMPEDANCE OF GREATER THAN 200 OHMS, AND NOISE SEEN ON BOTH HIGH VOLTAGE COILS WITH MANIPULATION. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-00866 | 2008-03-10 | 20071201 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT RECEIVED INAPPROPRIATE THERAPY DUE TO OVERSENSING, AND THERE WAS AN APPARENT LEAD FRACTURE. THE LEAD WAS TO BE REPLACED, HOWEVER TO THE PHYSICIAN WAS UNABLE TO REMOVE IT SO HIGH VOLTAGE THERAPIES WERE PROGRAMMED OFF, AND THE RV LEAD IS BEING USED FOR PACING. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED FURTHER REPORTING THAT THERE WAS ALSO HIGH IMPEDANCE, NOISE, AND ELEVATED THRESHOLDS. THE PACEMAKER WAS PROGRAMMED TO DDD MODE. ADDITIONAL INFORMATION RECEIVED REPORTING FRACTURE, HIGH IMPEDANCE OF GREATER THAN 2500 OHMS, AND THAT THE LEAD WILL BE CAPPED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00862 | 2008-03-10 | 20050315 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00861 | 2008-03-10 | 20060224 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT HAD BEEN REPORTED THE LEAD WAS OVERSENSING. THE HCP WAS UNABLE TO PASS A WIRE THROUGH THE MID COIL AND THE DISTAL TIP OF THE PACE/SENSE PORTION. THE LEAD WAS PARTIALLY REMOVED, CAPPED, AND REPLACED DUE TO AN APPARENT FRACTURE. ADDITIONAL INFORMATION RECEIVED REPORTS THE PATIENT RECEIVED SHOCKS. THE LEAD ALSO SHOWED OVERSENSING, HIGH IMPEDANCE AND FOUR SECOND PAUSES. NO FURTHER PATIENT COMPLICATIONS REPORTED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00857 | 2008-03-10 |  | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00850 | 2008-03-10 | 20050818 | Injury | Manufacturer report | SPRINT FIDELIS | DEFIBRILLATOR, AUTOMATIC IMPLANTABLE CARDIOVERTER | MEDTRONIC PUERTO RICO, INC. | 6949 | 6949 |  | IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00852 | 2008-03-10 | 20071219 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00853 | 2008-03-10 | 20080110 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00854 | 2008-03-10 |  | No answer provided | Manufacturer report | SPRINT FIDELIS | ASKU | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THAT THE SHORT INTERVAL COUNT HAD INCREASED, AND VENTRICULAR PACING IMPEDANCE HAD INCREASED FROM 500 OHMS TO 1654 OHMS. IT WAS FURTHER REPORTED THAT THERE WAS OVERSENSING AND A LEAD FRACTURE. A NEW LEAD WAS PLACED FOR PACING/SENSING PURPOSES. THE HIGH VOLTAGE PORTION OF THE LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | 0
2649622-2008-00855 | 2008-03-10 | 20071213 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | IT WAS REPORTED THAT THE PATIENT EXPERIENCED INAPPROPRIATE SHOCKS DUE TO OVERSENSING OF NOISE. THESE EVENTS WERE SEEN DURING FOLLOW-UP AND WERE NOTED TO HAVE OCCURRED ONE YEAR AGO. AN IMPEDANCE INCREASE TO 1000 OHMS AND A POSSIBLE LEAD FRACTURE WERE ALSO NOTED. THE PATIENT HAS REFUSED TO HAVE THE LEAD REVISED, SO IT HAS BEEN ABANDONNED BY TURNING OFF VF DETECT. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
1013891 | 2008-02-15 | 20080207 | Malfunction | User Facility report | SPRINT FIDELIS | LEAD, ICD | MEDTRONIC INC. | 6949-65 | * |  | THE PATIENT PRESENTED TO THE ELECTROPHYSIOLOGY CLINIC FOR A ROUTINE ICD FOLLOW UP. THE PATIENT STATED HE HAD RECEIVED A SHOCK FROM HIS ICD THIS MORNING. THE INTERROGATION OF DEVICE REVEALED A FRACTURED RIGHT VENTRICULAR LEAD. |  | N | Y | INVALID DATA
2649622-2008-01999 | 2008-03-24 | 20080219 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THIS INFORMATION WAS RECEIVED FROM A COMPETITOR. ALL DATA PERTINENT TO THE EVENT IS PROVIDED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES FOUND; DISTAL SEGMENT RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02000 | 2008-03-24 |  | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
MW5005985 | 2008-03-21 | 20080311 | Injury | Voluntary report | MEDTRONIC/SPRINT FIDELIS | DEFIBRILLATOR | MEDTRONIC, INC. |  |  |  | MY FATHER HAD A MEDTRONIC DEFIBRILLATOR PLACED APPROXIMATELY 3 YEARS AGO. UP UNTIL 2008 HAD NO PROBLEMS AND NO CONCERNS OF A FAULTY DEVICE. ON THE SAME DAY AT APPROXIMATELY 9:30 PM, THE DEFIBRILLATOR DELIVERED 30 SHOCKS WITHIN APPROXIMATELY A 30 MINUTE TIME PERIOD UNTIL THE AMBULANCE ARRIVED AT THE EMERGENCY ROOM AND THEY COULD SHUT IT OFF WITH A MAGNET. A MEDTRONIC TECHNICIAN CAME TO THE HOSPITAL AND EXPLAINED THAT HE DID NOT HAVE A HEART ATTACK, THAT HE WAS ONE OF THE PATIENTS THAT HAD RECEIVED A SPRINT FIDELIS LEAD THAT HAD BEEN RECALLED AND IT HAD FRACTURED CAUSING HIS DEVICE TO THINK HIS HEART RATE WAS 300-500 BPM, SO IT CONTINUALLY SHOCKED HIM. THE MEDTRONIC DEFIBRILLATOR DEVICE HAD FRACTURED LEAD AND INAPPROPRIATELY SHOCKED THE PATIENT -MY FATHER- 30 TIMES IN A 30 MINUTE TIMEFRAME. THE PATIENT HAD ONE OF THE RECALLED MODEL NUMBERS OF THE SPRINT FIDELIS LEADS THAT WERE REPORTED TO HAVE POSSIBLE PROBLEMS IN 2007 IN A LETTER TO PATIENTS AND DOCTORS. DIAGNOSIS OR REASON FOR USE: CHF, RISK OF MI. |  | N | Y | LAY USER/PATIENT
MW5006147 | 2008-04-08 | 20080405 | Injury | Voluntary report | SPRINT FIDELIS LEAD | NONE | MEDTRONIC | 694965 |  |  | I WAS IMPLANTED WITH A MEDTRONIC ICD IN 2007, SPRINT FIDELIS LEAD LFJ MODEL # 694965. THE DEVICE FIRED WHILE I WAS AT REST. VERY DISTRESSING. |  | Y | N | 0
2649622-2008-02009 | 2008-04-09 | 20080229 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: DISTAL CONDUCTOR WAS FRACTURED; FULL LEAD WAS RETURNED FOR ANALYSIS. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02010 | 2008-04-09 | 20080317 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED. FULL LEAD IN SEGMENTS RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02011 | 2008-04-09 | 20080312 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THE LEAD WAS EXPLANTED DUE TO APPARENT FRACTURE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THE EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-02014 | 2008-04-16 | 20080325 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE LEAD WAS EXPLANTED AND REPLACED DUE TO HIGH IMPEDANCE, OVERSENSING, AN APPARENT FRACTURE AND INAPPROPRIATE THERAPY. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02015 | 2008-04-16 | 20080204 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
1033270 | 2008-04-11 | 20080303 | Malfunction | User Facility report | SPRINT FIDELIS | LEAD, ICD | MEDTRONIC, INC. CARDIAC RHYTHM MANAGEMENT | 6949-58 | 6949 |  | ELECTIVE REPLACEMENT OF IMPLANTABLE CARDIOVERTER-DEFIBRILLATOR AND ABANDONED RECALLED FIDELIS LEAD. |  | N | Y | INVALID DATA
1033271 | 2008-04-11 | 20080305 | No answer provided | User Facility report | SPRINT FIDELIS DEFIBRILLATOR LEAD | LEAD, ICD | MEDTRONIC, INC. CARDIAC RHYTHM MANAGEMENT | 6931 | NA |  | PATIENT IMPLANTABLE CARDIOVERTER-DEFIBRILLATOR LEAD IS A MEDTRONIC SPRINT FIDELIS THAT IS ON THE VOLUNTARY RECALL LIST. PATIENT AND HER FAMILY DESIRE REPLACEMENT DUE TO PATIENT'S HISTORY OF AN ARREST. |  | N |  | INVALID DATA
2649622-2008-02041 | 2008-04-28 |  | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: DISTAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
1035206 | 2008-04-10 | 20080403 | Injury | User Facility report | SPRINT FIDELIS | LEAD, ICD | MEDTRONIC, INC. | 6949 | * |  | FRACTURE OF CHRONIC INDWELLING DEFIBRILLATOR LEAD WHICH RESULTED IN A HIGH IMPEDANCE ALARM SOUNDING ON A DAILY BASIS. PATIENT WAS EXPERIENCING CHEST PAIN WITH EXERTION. |  | Y | Y | INVALID DATA
MW5006500 | 2008-04-30 | 20071205 | Injury | Voluntary report | SPRINT FIDELIS | ICD LEAD | MEDTRONIC, INC. | 6949 |  |  | MEDTRONIC ICD LEAD SPRINT FIDELIS FRACTURED/MALFUNCTIONED RESULTING IN INAPPROPRIATE ICD SHOCKS. ICD LEAD WAS REPLACED IN 2007. |  | Y | Y | HEALTH PROFESSIONAL
MW5006501 | 2008-04-30 | 20080102 | Injury | Voluntary report | SPRINT FIDELIS | ICD LEAD | MEDTRONIC INC. | 6949 |  |  | MEDTRONIC 6949 ICD MALFUNCTION RESULTING INAPPROPRIATE ICD SHOCKS. ICD LEAD WAS REPLACED IN 2008. |  | Y | Y | HEALTH PROFESSIONAL
MW5006502 | 2008-04-30 | 20080108 | Injury | Voluntary report | MEDTRONIC SPRINT FIDELIS | ICD LEAD | MEDTRONIC INC. | 6949 |  |  | INAPPROPRIATE ICD SHOCKS DUE TO MEDTRONIC FIDELIS 6949 LEAD MALFUNCTION. ICD LEAD WAS REPLACED IN 2008. |  | Y | Y | HEALTH PROFESSIONAL
MW5006503 | 2008-04-30 | 20071022 | Injury | Voluntary report | MEDTRONIC SPRINT FIDELIS | ICD LEAD | MEDTRONIC INC. | 6949 |  |  | PATIENT RECEIVED SALVOES OF INAPPROPRIATE ICD SHOCKS DUE TO A FRACTURED MEDTRONIC SPRINT FIDELIS LEAD. ICD LEAD REPLACED IN 2007. DATES OF USE: 2005 - 2007. |  | Y | Y | HEALTH PROFESSIONAL
MW5006504 | 2008-04-30 | 20071102 | Injury | Voluntary report | SPRINT FIDELIS | ICD LEAD | MEDTRONIC INC. | 6949 |  |  | MEDTRONIC SPRINT FIDELIS ICD LEAD FRACTURE RESULTED IN ICD SHOCKS. ICD LEAD WAS REPLACED IN 2007. |  | Y | Y | HEALTH PROFESSIONAL
MW5006505 | 2008-04-30 | 20071212 | Injury | Voluntary report | SPRINT FIDELIS | MEDTRONIC | MEDTRONIC INC. | 6949 |  |  | MEDTRONIC SPRINT FIDELIS FRACTURE RESULTED IN INAPPROPRIATE ICD SHOCKS. ICD LEAD WAS EXPLANTED AND A NEW LEAD WAS IMPLANTED. |  | Y | Y | HEALTH PROFESSIONAL
2649622-2008-02043 | 2008-05-05 | 20080314 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02033 | 2008-04-25 | 20080219 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE LEAD WILL BE REMOVED, DUE TO INAPPROPRIATE THERAPY CAUSED BY A LEAD FRACTURE. REPLACEMENT AT THE TIME OF REPORT WAS NOT POSSIBLE DUE TO PATIENT'S OTHER HEALTH ISSUES, SO A LIFEVEST WAS APPLIED UNTIL IT COULD BE DONE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED REPORTING THAT THE LEAD HAS BEEN REPLACED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02034 | 2008-04-25 | 20080114 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02035 | 2008-04-25 | 20080122 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02036 | 2008-04-25 | 20080207 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02038 | 2008-04-25 | 20080228 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: NO ANOMALIES FOUND; PROXIMAL SEGMENT RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02040 | 2008-04-25 | 20080311 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT OVERSENSING CAUSED INAPPROPRIATE SHOCKS. HIGH IMPEDANCE AND LEAD FRACTURE WAS ALSO NOTED. THE FAMILY ELECTED NOT TO REPLACE THE LEAD. THE HIGH VOLTAGE PORTION OF THE DEVICE WAS TURNED OFF, BUT THE LOW VOLTAGE FUNCTIONALITY WAS LEFT ON. IT WAS FURTHER REPORTED THAT THE PATIENT COMPLAINED OF "DISCOMFORT AT THE SITE". A SECOND MEDWATCH FORM WAS LATER RECEIVED REPORTING DOCUMENTED MALFUNCTIONING AND INAPPROPRIATE SHOCKING. THE LEAD WAS LATER REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02032 | 2008-04-25 | 20071001 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02021 | 2008-04-25 | 20071205 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02022 | 2008-04-25 | 20071219 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02023 | 2008-04-25 | 20080122 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; DISTAL SEGMENT RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02024 | 2008-04-25 | 20071022 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02025 | 2008-04-25 | 20080205 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | EVALUATION SUMMARY: NO ANOMALIES FOUND; PROXIMAL SEGMENT RETURNED AND ANALYZED. OTHER: IT WAS INITIALLY REPORTED THE LEAD APPEARED TO BE KINKED. A MEDWATCH 3500A WAS SUBSEQUENTLY RECEIVED AND REPORTED THE PATIENT WAS ADMITTED WITH SYNCOPE. ADDITIONAL INFORMATION WAS LATER RECEIVED REPORTING THE PATIENT RECEIVED NUMEROUS INAPPROPRIATE SHOCKS. LEAD FRACTURE WAS NOTED. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ACTUAL DEVICE INVOLVED IN INCIDENT WAS EVALUATED. ELECTRICAL TESTS PERFORMED; MECHANICAL TESTS PERFORMED. VISUAL EXAMINATION: DEVICE PERFORMED ACCORDING TO SPECIFICATIONS. NO CONCLUSION CAN BE DRAWN. KINK, LEAD(S), FRACTURE OF, SHOCK, INAPPROPRIATE. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02026 | 2008-04-25 | 20080207 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02027 | 2008-04-25 | 20080109 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02028 | 2008-04-25 | 20071129 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02029 | 2008-04-25 | 20080114 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02030 | 2008-04-25 | 20080227 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THERE WAS FLUCTUATING AND OUT OF RANGE IMPEDANCE ON THE PACE/SENSE PORTION OF THE LEAD, AND A LEAD FRACTURE WAS SUSPECTED. THE LEAD WAS TO BE EXPLANTED USING A LASER. DURING THE PROCEDURE, THE LASER/SHEATH COULD ONLY BE ADVANCED HALFWAY OVER THE SVC COIL, THEN APPEARED TO BE STUCK. WHILE TRYING TO ADVANCE THE SHEATH, IT BECAME CLEAR THAT THE PATIENT WAS IN ACUTE DISTRESS. IT WAS SUSPECTED THE SVC HAD BEEN PUNCTURED/TORN BY THE LASER. AFTER MUCH TIME AND WORK BY THE PHYSICIANS AND STAFF, THE PATIENT EXPIRED. A MEDWATCH 3500A WAS SUBSEQUENTLY RECEIVED REPORTING ABNORMAL IMPEDANCE, POSSIBLY INDICATIVE OF LEAD FRACTURE/BREAK. THE PATIENT BECAME HYPOTENSIVE AND BRADYCARDIC WHILE TRYING TO REMOVE THE LEAD, AND WENT INTO ELECTRICAL MECHANICAL DISSOCIATION. THE PATIENT DIED. IT WAS SUSPECTED THAT THE DEATH WAS RELATED TO LASER PERFORATION OF VESSEL OR MI (MYOCARDIAL INFARCTION). IT IS NOT KNOWN IF THE LEAD WAS EXPLANTED, AND AN AUTOPSY WAS NOT PERFORMED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02020 | 2008-04-25 | 20071211 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02031 | 2008-04-25 | 20080117 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS REPLACED DUE TO POOR SENSING. ADDITIONALLY, ALLEGED THE PATIENT "SUFFERED PHYSICAL AND OTHER INJURY" AND "EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY REQUIRING MEDICAL INTERVENTION INCLUDING BUT NOT LIMITED TO SURGERY, REMOVAL AND/OR REPLACEMENT OF THE DEFECTIVE [LEAD]". IT WAS ALSO ALLEGED THAT THE PATIENT SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH AND SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH. |  | N | Y | LAY USER/PATIENT
2649622-2008-02019 | 2008-04-24 | 20080111 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02069 | 2008-05-08 | 20071127 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02065 | 2008-05-08 | 20071112 | Injury | Manufacturer report | SPRINT FIDELIS | DEFIBRILLATOR, AUTOMATIC IMPLANTABLE CARDIOVERTER | MPRI | 6949 | 6949 |  | THE LEAD WAS RETURNED TO THE MANUFACTURER, ANALYZED, AND TESTED OUT OF SPECIFICATION. APPROXIMATELY THREE MONTHS PRIOR TO EXPLANT, IT WAS REPORTED THAT THE PATIENT HEARD A HIGH URGENCY ALERT AT 4 AND 10AM; HOWEVER, THERE WERE NO ALERTS IN THE ALERT LOG. IT WAS ALSO REPORTED THAT THE PATIENT RECEIVED A SHOCK DURING THE NIGHT. NO INFORMATION WAS RECEIVED INDICATING WHETHER OR NOT THE SHOCK WAS APPROPRIATE. AT THAT TIME, THE PATIENT WAS INSTRUCTED TO KEEP A LOG, AND THE LEAD REMAINED IN USE. ADDITIONAL INFORMATION SUBSEQUENTLY RECEIVED ALLEGED THAT THE PLAINTIFF "EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, OR OTHER INJURY REQUIRING MEDICAL INTERVENTION INCLUDING BUT NOT LIMITED TO SURGERY, REMOVAL AND/OR REPLACEMENT OF THE DEFECTIVE [LEAD]." |  | Y | Y | LAY USER/PATIENT
2649622-2008-02063 | 2008-05-08 | 20071129 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: DEFIB CONDUCTOR FRACTURED; FULL LEAD ANALYZED. |  | N | Y | LAY USER/PATIENT
2182208-2008-00102 | 2008-05-08 | 20071005 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC, INC. | 6948 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02057 | 2008-05-08 | 20080229 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: DISTAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02058 | 2008-05-08 | 20071113 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02051 | 2008-05-08 | 20080126 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | IT WAS REPORTED THAT THE PATIENT HAD MULTIPLE REVISIONS AFTER INITIAL IMPLANT. MULTIPLE SHOCKS, FRACTURE, AND ALERT TONES WAS NOTED. THE LEAD WAS EXPLANTED AND REPLACED, WITH THE SYSTEM UPGRADED TO DUAL CHAMBER AT THAT TIME. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02052 | 2008-05-08 | 20061229 | Injury | Manufacturer report | SPRINT FIDELIS | DEFIBRILLATOR, AUTOMATIC IMPLANTABLE CARDIOVERTER | MPRI | 6931 | 6931 |  | IT WAS REPORTED THE LEAD WAS EXPLANTED AND REPLACED DUE TO 13 INAPPROPRIATE SHOCKS, NOT SENSING OR PACING AND HIGH IMPEDANCE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION RECEIVED STATED THE LEAD WAS REPLACED DUE TO A FRACTURE. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02053 | 2008-05-08 | 20061130 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ATTORNEY ALLEGES 'CLAIM FOR DAMAGES ARISING OUT OF DEFECTS IN AND REMOVAL OF THE ABOVE-REFERENCED EQUIPMENT." THE LEAD WAS EXPLANTED AND REPLACED. ADDITIONAL INFORMATION FROM THE ATTORNEY REPORTS THAT BECAUSE OF THE LEAD DEFECT, THE PATIENT RECEIVED UNNECESSARY AND INAPPROPRIATE SHOCKS. IT WAS FURTHER REPORTED BY THE ATTORNEY THE PATIENT SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02046 | 2008-05-08 | 20080318 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-02047 | 2008-05-08 | 20080208 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE RV (RIGHT VENTRICULAR) LEAD WAS EXPLANTED AND REPLACED DUE TO AN APPARENT LEAD FRACTURE, AND THAT THE PATIENT RECEIVED INAPPROPRIATE THERAPY DUE TO T-WAVE OVERSENSING. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02048 | 2008-05-08 | 20051016 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | A LAWSUIT ALLEGES THAT DUE TO THE "DESIGN, MANUFACTURE, ASSEMBLY, MARKETING AND SALES OF THE SPRINT FIDELIS LEAD," THE PATIENT "HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, AND SEVERE EMOTIONAL DISTRESS." ADDITIONAL INFORMATION RECEIVED REPORTED THE LEAD HAD BEEN IMPLANTED FOR TWO DAYS AND WAS REPLACED, DUE TO A REPORT OF HIGH THRESHOLDS, NO CAPTURE, SENSING DIFFICULTY, AND BLOOD INGRESS. AT THAT TIME, THE LEAD WAS RETURNED AND TESTED WITHIN SPECIFICATION. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. SENSING DIFFICULTY, HIGH THRESHOLDS, AND BLOOD INGRESS NEAR RV COIL ONE-DAY POST IMPLANT. LEAD EXPLANTED DUE TO POOR R-WAVE SENSING. IT WAS NOTED THAT LEAD HAD BLOOD INGRESS NEAR RV COIL. ON (B)(6) 2008: A LAWSUIT ALLEGES THAT DUE TO THE "DESIGN, MANUFACTURE, ASSEMBLY, MARKETING AND SALES OF THE SPRINT FIDELIS LEAD," THE PATIENT "HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, AND SEVERE EMOTIONAL DISTRESS." ATTORNEY LATER ALLEGES THE PATIENT IS DECEASED AND THAT THERE WAS A LEAD FRACTURE AND/OR EXPLANT. REVIEW OF MANUFACTURER'S DATABASE REVEALED THE LEAD WAS EXPLANTED PRIOR TO THE PATIENT'S DEATH. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02049 | 2008-05-08 | 20080303 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: NO ANOMALIES FOUND; PROXIMAL SEGMENT WAS RETURNED AND ANALYZED. DISCLAIMER: SUBMISSION OF INFORMATION BY MEDTRONIC UNDER THE MEDICAL DEVICE REPORTING REGULATION DOES NOT CONSTITUTE AN ADMISSION THAT THE DEVICE(S) HAS MALFUNCTIONED OR THAT THERE IS ANY CAUSAL CONNECTION BETWEEN THE PERFORMANCE OF THE DEVICE AND ANY INJURY THAT MAY HAVE OCCURRED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-03191 | 2008-05-09 | 20080330 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03176 | 2008-05-09 | 20080326 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-03177 | 2008-05-09 | 20080311 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE, DUE TO CONFIDENTIALITY CONCERNS. EVALUATION SUMMARY: OUTER TUBING BREACH; FULL LEAD RETURNED. THE BLOOD OBSERVED UNDER THE OVERLAY TUBING ENTERED THROUGH THE BREACH. THIS WOULD NOT AFFECT LEAD PERFORMANCE. THE MELTS IN THE OVERLAY TUBING APPEAR TO BE FROM CAUTERY. |  | N | Y | LAY USER/PATIENT
2649622-2008-03178 | 2008-05-09 | 20060517 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THE LEAD WAS REPLACED DUE TO OVERSENSING. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-03179 | 2008-05-09 | 20080310 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THE PATIENT RECEIVED 2 INAPPROPRIATE SHOCKS. HIGH IMPEDANCE AND SMALL R-WAVES WERE ALSO INDICATED. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-03180 | 2008-05-09 | 20080314 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-03181 | 2008-05-09 | 20080318 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-03183 | 2008-05-09 | 20080328 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE PATIENT ARRIVED AT THE HOSPITAL WITH BRADYCARDIA, DUE TO A LOSS OF PACING. ADDITIONALLY, THE LEAD EXHIBITED OVERSENSING, HIGH THRESHOLDS, AND HIGH IMPEDANCES. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-03184 | 2008-05-09 | 20080320 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-03185 | 2008-05-09 | 20080228 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03186 | 2008-05-09 | 20080306 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-03188 | 2008-05-09 | 20080329 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03167 | 2008-05-09 | 20080323 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-03168 | 2008-05-09 | 20080306 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY, AND TESTED OUT OF SPECIFICATION CONSISTENT WITH THAT ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR WAS FRACTURED: FULL LEAD IN SEGMENTS WAS RETURNED FOR ANALYSIS. |  | Y | Y | LAY USER/PATIENT
2649622-2008-03169 | 2008-05-09 | 20080325 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THERE WAS A LEAD FRACTURE, THAT PART OF THE CONDUCTOR CAME OUT OF THE LEAD BODY WHEN CUT DURING EXPLANT PROCEDURE, AND THAT THE HELIX WOULD NOT RETRACT USING THE ROTATION TOOL BUT ONLY WHEN THERE WAS CONSTANT PULLING PRESSURE ON THE LEAD. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED. ADDITIONAL INFORMATION WAS RECEIVED ALLEGING THE PATIENT "SUFFERED PHYSICAL AND OTHER INJURY", AND THAT THE PATIENT "EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY REQUIRING MEDICAL INTERVENTION INCLUDING BUT NOT LIMITED TO SURGERY, REMOVAL AND/OR REPLACEMENT OF THE DEFECTIVE [LEAD]". IT ALSO ALLEGED THAT THE PATIENT "SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH, ECONOMIC LOSSES AND OTHER DAMAGES. [PATIENT] HAS FURTHER SUFFERED PHYSICAL INJURIES AND VARIOUS PHYSICAL MANIFESTATIONS OF EMOTIONAL DISTRESS". |  | Y | Y | LAY USER/PATIENT
2649622-2008-03170 | 2008-05-09 | 20080326 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THERE WAS OVERSENSING AND HIGH IMPEDANCE GREATER THAN 3000 OHMS. THE LEAD WAS EXPLANTED AND REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. FURTHER ALLEGED THAT THE PATIENT SUFFERED PHYSICAL INJURY, INCLUDING BUT NOT LIMITED TO REPEATED AND UNNECESSARY HIGH-VOLTAGE SHOCKS OF ELECTRICITY THROUGH THE CHEST AND SUFFERED PHYSICAL AND OTHER INJURIES, AND THE LEAD WAS DEFECTIVE AND SUFFERED PHYSICAL AND EMOTIONAL INJURIES, INCLUDING BUT NOT NECESSARILY LIMITED TO DEATH, EMERGENCY AND ADDITIONAL SURGERIES TO REMOVE OR REPLACE THE DEFECTIVE LEADS, UNNECESSARY SHOCKING, ADDITIONAL MEDICAL MONITORING, AND VARIOUS PHYSICAL MANIFESTATIONS OF EXTREME EMOTIONAL DISTRESS. |  | Y | Y | LAY USER/PATIENT
2649622-2008-03172 | 2008-05-09 | 20080213 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-03173 | 2008-05-09 |  | No answer provided | Manufacturer report | SPRINT FIDELIS | ASKU | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | 0
2649622-2008-03174 | 2008-05-09 | 20080310 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-03166 | 2008-05-09 | 20080328 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES FOUND; FULL LEAD ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-03175 | 2008-05-09 | 20080301 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: DISTAL CONDUCTOR FRACTURED; PARTIAL LEAD IN SEGMENTS RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-03156 | 2008-05-09 | 20080321 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THERE WAS HIGH IMPEDANCE GREATER THAN 2000 OHMS AND THAT THE PATIENT RECEIVED 42 INAPPROPRIATE SHOCKS DUE TO OVERSENSING. IT WAS ALSO REPORTED THAT THERE WAS A LEAD FRACTURE. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-03157 | 2008-05-09 | 20080320 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03158 | 2008-05-09 | 20080330 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03159 | 2008-05-09 | 20080321 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03162 | 2008-05-09 | 20080324 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-03163 | 2008-05-09 | 20080321 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-03164 | 2008-05-09 | 20080321 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03165 | 2008-05-09 | 20080321 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-03146 | 2008-05-09 | 20080310 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-03147 | 2008-05-09 | 20080225 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03149 | 2008-05-09 | 20080310 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-03150 | 2008-05-09 | 20070926 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED BY THE ATTORNEY THAT THE PATIENT WAS HOSPITALIZED AND HAD A LEAD REPLACED. THERE WERE NO REPORTED PATIENT COMPLICATIONS AS A RESULT OF THIS EVENT. ATTORNEY LATER ALLEGES "ON (B)(6) 2007, (PATIENT) EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY REQUIRING MEDICAL INTERVENTION INCLUDING BUT NOT LIMITED TO SURGERY, REMOVAL AND/OR REPLACEMENT OF THE DEFECTIVE SPRINT FIDELIS LEAD, MODEL 6949" AND AS A DIRECT RESULT OF THE LEAD, PATIENT "HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH, ECONOMIC LOSSES, AND OTHER DAMAGES." REVIEW OF MANUFACTURER'S DATABASE NOTED PATIENT DEATH AND ALSO THAT THE PATIENT DID NOT HAVE A SPRINT FIDELIS LEAD MODEL IMPLANTED AT THE TIME OF DEATH. THERE IS NO ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH HAS BEEN RESEARCHED AND NOT RECEIVED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-03152 | 2008-05-09 | 20080319 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT HEARD HIS PATIENT ALERT SOUND AND WENT TO THE CLINIC, AND THAT THERE WAS OVERSENSING, HIGH IMPEDANCE GREATER THAN 1900 OHMS, AND A POSSIBLE LEAD FRACTURE. IT WAS FURTHER REPORTED THAT THERE WAS AN IMPEDANCE INCREASE AND SENSING DIFFICULTY. THE LEAD WAS EXPLANTED AND REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-03154 | 2008-05-09 | 20080321 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03144 | 2008-05-09 | 20060928 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03138 | 2008-05-09 |  | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-03141 | 2008-05-09 | 20080304 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-03123 | 2008-05-09 | 20071109 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-03124 | 2008-05-09 | 20071216 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03125 | 2008-05-09 | 20080305 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03126 | 2008-05-09 | 20080110 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-03130 | 2008-05-09 | 20060306 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-03132 | 2008-05-09 | 20071128 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03122 | 2008-05-09 | 20080128 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: MIDDLE INSULATION BREACHED; DISTAL SEGMENT ANALYZED. DISTAL CONDUCTOR FRACTURED; FULL LEAD ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-03112 | 2008-05-09 | 20080221 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THERE WERE HIGH THRESHOLDS. THE PATIENT FURTHER REPORTED THAT HIS DOCTOR TOLD HIM "THE LEAD IS NOT CAPTURING PROPERLY AND NOT SENSING PROPERLY". THE LEAD WAS EXPLANTED AND REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | N | N | LAY USER/PATIENT
2649622-2008-03115 | 2008-05-09 | 20080229 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-03120 | 2008-05-09 | 20080314 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-03104 | 2008-05-09 | 20080208 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED BY THE PATIENT THAT HE HAD A LEAD REPLACED DUE TO IT "GOING HAYWIRE." ADDITIONAL INFORMATION RECEIVED FROM THE PHYSICIAN OFFICE STATED THE LEAD "MALFUNCTIONED AND THE PATIENT IS NOT GETTING SHOCKS ANY MORE." THE LEAD WAS REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED. ADDITIONAL INFORMATION WAS RECEIVED ALLEGING THAT THE PATIENT "SUFFERED PHYSICAL AND OTHER INJURY" DUE TO THE LEAD, AND THAT THE PATIENT SUSTAINED "INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY REQUIRING MEDICAL INTERVENTION INCLUDING BUT NOT LIMITED TO SURGERY, REMOVAL AND/OR REPLACEMENT OF THE DEFECTIVE [LEAD]". IT ALSO ALLEGED THAT THE PATIENT HAS "SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH, ECONOMIC LOSSES AND OTHER DAMAGES". [PATIENT] HAS FURTHER SUFFERED PHYSICAL INJURIES AND VARIOUS PHYSICAL MANIFESTATIONS OF EMOTIONAL DISTRESS". |  | Y | Y | LAY USER/PATIENT
2649622-2008-03105 | 2008-05-09 | 20070813 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03106 | 2008-05-09 | 20080219 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-03107 | 2008-05-09 | 20080229 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. |  | N | Y | LAY USER/PATIENT
2649622-2008-03108 | 2008-05-09 | 20080205 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-03091 | 2008-05-09 | 20080220 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03092 | 2008-05-09 | 20080228 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-03094 | 2008-05-09 | 20080226 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-03095 | 2008-05-09 | 20080228 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-03096 | 2008-05-09 | 20080221 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03098 | 2008-05-09 | 20080218 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: NO ANOMALIES FOUND; PROXIMAL SEGMENT RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-03099 | 2008-05-09 | 20080314 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-03090 | 2008-05-09 | 20080229 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03102 | 2008-05-09 |  | Injury | Manufacturer report | SPRINT FIDELIS | ASKU | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT BOTH THE RA AND THE RV LEADS WERE EXPLANTED AND REPLACED DUE TO HIGH IMPEDANCE AND OVERSENSING. BOTH LEADS WERE RETURNED TO THE MANUFACTURER, ANALYZED, AND SUBSEQUENTLY TESTED OUT OF SPECIFICATION. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | 0
2649622-2008-03085 | 2008-05-09 | 20080220 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03086 | 2008-05-09 | 20080229 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS REMOVED AND REPLACED DUE TO OVERSENSING, SENSING DIFFICULTY, AND INAPPROPRIATE THERAPIES. THE PATIENT FURTHER REPORTED THAT HIS LEAD BROKE, AND THAT HE WAS SHOCKED 4 TIMES WHILE DRIVING. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-03089 | 2008-05-09 | 20080219 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THAT THE PATIENT RECEIVED INAPPROPRIATE SHOCKS. A LEAD FRACTURE WAS SUSPECTED. THE LEAD WAS REPLACED. THERE WERE NO FURTHER REPORTED PATIENT COMPLICATIONS AS A RESULT OF THIS EVENT. THE PATIENT'S STATUS WAS REPORTED AS "OK". |  | Y | Y | LAY USER/PATIENT
2649622-2008-03074 | 2008-05-09 | 20080219 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03075 | 2008-05-09 | 20080214 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03076 | 2008-05-09 | 20080213 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | IT WAS REPORTED THAT THERE WAS INCREASED THRESHOLD. THE LEAD WAS EXPLANTED AND REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT, AND THE PATIENT'S STATUS WAS REPORTED AS "OK". IT WAS REPORTED THAT THE ATRIAL LEAD WAS EXPLANTED AND REPLACED DUE TO DISLODGEMENT. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | N | N | LAY USER/PATIENT
2649622-2008-03078 | 2008-05-09 | 20080225 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-03079 | 2008-05-09 | 20080118 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-03073 | 2008-05-09 | 20080211 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03063 | 2008-05-09 | 20080211 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03064 | 2008-05-09 | 20080211 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-03065 | 2008-05-09 | 20080212 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-03066 | 2008-05-09 | 20071031 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-03067 | 2008-05-09 | 20080208 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-03062 | 2008-05-09 | 20080211 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-03072 | 2008-05-09 | 20080213 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03052 | 2008-05-09 | 20080208 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-03053 | 2008-05-09 | 20080213 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: NO ANOMALIES FOUND; FULL LEAD RETURNED. |  | N | Y | LAY USER/PATIENT
2649622-2008-03054 | 2008-05-09 | 20080215 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-03055 | 2008-05-09 | 20080211 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-03056 | 2008-05-09 | 20080212 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: NO ANOMALIES FOUND; THERE IS ONLY ONE SET OF SETSCREW MARKS ON THE IS-1 CONNECTOR PIN AND IT IS TOO PROXIMAL WHICH INDICATES THE LEAD WAS NOT FULLY INSERTED INTO THE DEVICE; FULL LEAD ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-03058 | 2008-05-09 | 20071206 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03059 | 2008-05-09 | 20071108 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03051 | 2008-05-09 | 20071029 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THERE WAS NOISE ON THE LEAD. FURTHER ALLEGED THAT THE PATIENT "EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING" AND LEAD FRACTURE. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-03041 | 2008-05-09 | 20080204 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-03042 | 2008-05-09 | 20080129 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-03044 | 2008-05-09 | 20080207 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-03047 | 2008-05-09 | 20080221 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THE PATIENT ALERT SOUNDED FOR HIGH HV LEAD IMPEDANCES. THE LEAD HAD AN APPARENT FRACTURE. THE LEAD WAS EXPLANTED AND REPLACED. ADDITIONAL INFORMATION RECEIVED ALLEGES THAT THE PATIENT "SUFFERED PHYSICAL AND OTHER INJURY" AS A RESULT OF THE LEAD. IT ALSO ALLEGED THE PATIENT "EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY REQUIRING MEDICAL INTERVENTION INCLUDING BUT NOT LIMITED TO SURGERY, REMOVAL AND/OR REPLACEMENT OF THE DEFECTIVE [LEAD]. ALSO ALLEGED WAS THAT THE PATIENT" HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH, ECONOMIC LOSSES AND OTHER DAMAGES. [PATIENT] HAS FURTHER SUFFERED PHYSICAL INJURIES AND VARIOUS PHYSICAL MANIFESTATIONS OF EMOTIONAL DISTRESS". LEAD "FAILURE" WAS ALSO NOTED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-03050 | 2008-05-09 | 20071030 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: NO ANOMALIES FOUND; FULL LEAD RETURNED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-03028 | 2008-05-09 | 20080204 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THERE WAS T-WAVE OVERSENSING. THE LEAD WAS EXPLANTED AND REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. THE PATIENT FURTHER REPORTING THAT HE WAS WALKING AND WAS "HIT" BY THE DEVICE, AND WAS TOLD THE LEAD WAS "BAD". FURTHER ALLEGED THAT THE PATIENT "EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING" AND A LEAD FRACTURE. |  | Y | Y | LAY USER/PATIENT
2649622-2008-03029 | 2008-05-09 | 20080206 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03037 | 2008-05-09 | 20080206 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS EXPLANTED AND REPLACED DUE TO LOW IMPEDANCE, OVER/UNDER SENSING, AND HIGH PACING THRESHOLDS. ATTORNEY LATER ALLEGES PATIENT "SUFFERED PHYSICAL AND OTHER INJURY AS A RESULT OF THE RECALLED LEAD." FURTHER ALLEGES THAT "ON OR ABOUT (B) (6) 2008, (PATIENT) EXPERIENCED A MALFUNCTION AND PROBABLE FRACTURE OF THE MEDTRONIC SPRINT FIDELIS LEAD WIRE, REQUIRING MEDICAL INTERVENTION INCLUDING BUT NOT LIMITED TO SURGERY, REMOVAL AND/OR REPLACEMENT OF THE DEFECTIVE SPRINT FIDELIS LEAD, MODEL 6949" AND "SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH, ECONOMIC LOSSES AND OTHER DAMAGES." |  | Y | Y | LAY USER/PATIENT
2649622-2008-03038 | 2008-05-09 | 20080207 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03027 | 2008-05-09 | 20080205 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03039 | 2008-05-09 | 20080129 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-03021 | 2008-05-09 | 20080122 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-03022 | 2008-05-09 | 20080129 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-03024 | 2008-05-09 | 20080201 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-03025 | 2008-05-09 | 20080201 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03026 | 2008-05-09 | 20080201 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03016 | 2008-05-09 | 20080115 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-03018 | 2008-05-09 | 20080109 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-03019 | 2008-05-09 |  | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-03014 | 2008-05-09 | 20071022 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-03013 | 2008-05-09 |  | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03012 | 2008-05-09 | 20071024 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6930 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-03010 | 2008-05-09 | 20080213 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-03011 | 2008-05-09 | 20060803 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-03002 | 2008-05-09 | 20080306 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES FOUND; FULL LEAD IN SEGMENTS RETURNED AND ANALYZED. |  | N | N | LAY USER/PATIENT
2649622-2008-03008 | 2008-05-09 | 20080229 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | INCREASED/HIGH IMPEDANCE AND OVERSENSING WAS REPORTED. THE LEAD WAS EXPLANTED AND REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-03005 | 2008-05-09 | 20080108 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-03196 | 2008-05-12 | 20071027 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03197 | 2008-05-12 | 20071019 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | AN ATTORNEY ALLEGES THE PATIENT SUSTAINED INJURY AS A RESULT OF THE LEAD. ATTORNEY LATER ALLEGES PATIENT "SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH, ECONOMIC LOSSES AND OTHER DAMAGES" AND "SUFFERED PHYSICAL INJURIES AND VARIOUS PHYSICAL MANIFESTATIONS OF EMOTIONAL DISTRESS ASSOCIATED WITH ONE OR MORE OF THE FOLLOWING: THE IMPLANTATION, RECALL, FAILURE, REMOVAL/REPLACEMENT, AND/OR INABILITY TO HAVE THE DEFECTIVE SPRINT FIDELIS LEAD REMOVED OR REPLACED. (PATIENT) DIED AS A DIRECT AND PROXIMATE RESULT OF DEFECTS IN DEFENDANTS' SPRINT FIDELIS LEADS." THERE IS NO ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH HAS BEEN RESEARCHED AND NOT RECEIVED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-03217 | 2008-05-13 | 20070730 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | MRSA ENDOCARDITIS AND AN ECHODENSITY ON THE RV DEFIBRILLATION LEAD AND THE AORTIC VALVE WERE REPORTED. THE DEVICE AND LEAD WERE EXPLANTED. THE PATIENT EXPIRED APPROXIMATELY ONE MONTH LATER. THERE IS NO ALLEGATION BY A HEALTHCARE PROFESSIONAL THAT THE DEATH WAS DEVICE RELATED. |  | Y | N | LAY USER/PATIENT
2649622-2008-03221 | 2008-05-13 | 20080507 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE PATIENT RECEIVED 28 SHOCKS DUE TO NOISE. THE LEAD WAS REPLACED. IT WAS ALSO REPORTED THAT THE PATIENT WAS TRAUMATIZED BY THE EPISODE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION RECEIVED REPORTED THAT PATIENT WAS SHOCKED DUE TO APPARENT LEAD FRACTURE. |  | Y | Y | LAY USER/PATIENT
2649622-2008-03213 | 2008-05-13 | 20061202 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-03203 | 2008-05-13 | 20071028 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | AN ATTORNEY ALLEGES THE PATIENT'S LEAD REPEATEDLY SHOCKED HIM, RESULTING IN HIS DEATH. IT WAS LATER ALLEGED BY ATTORNEY THAT THE PATIENT SUFFERED PHYSICAL AND OTHER INJURY AS A RESULT OF THE RECALLED LEAD. IT IS FURTHER ALLEGED THE PATIENT EXPERIENCED "INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY REQUIRING MEDICAL INTERVENTION INCLUDING BUT NOT LIMITED TO SURGERY, REMOVAL AND/OR REPLACEMENT OF THE DEFECTIVE SPRINT FIDELIS LEAD." ATTORNEY ALSO ALLEGES AS A RESULT OF THE LEAD, THE PATIENT "SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH." REVIEW OF MANUFACTURER'S DATABASE VERIFIED PATIENT'S DEATH. THERE IS NO ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH HAS BEEN RESEARCHED AND NOT RECEIVED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-03209 | 2008-05-13 |  | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-03210 | 2008-05-13 | 20051226 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-03211 | 2008-05-13 | 20050531 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-03212 | 2008-05-13 | 20060104 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2182208-2008-00146 | 2008-05-14 | 20070921 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC, INC. | 6948 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03223 | 2008-05-14 | 20080313 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THERE APPEARED TO BE OVERSENSING, AN IMPEDANCE INCREASE/SUBSEQUENT HIGH IMPEDANCE, AND QUESTIONABLE CAPTURE ON THE LEAD INDICATING POSSIBLE FRACTURE. FOLLOW-UP WITH THE CLINIC INDICATES THAT THE PATIENT'S SON CALLED TO CANCEL AN APPOINTMENT BECAUSE HIS FATHER DIED THE DAY BEFORE. HE ALLEGES THAT THE LEAD MALFUNCTIONED. CAUSE OF DEATH WAS REQUESTED, BUT THE CLINIC DID NOT HAVE IT AVAILABLE. ADDITIONAL INFORMATION RECEIVED FROM AN ARTICLE INDICATES THE PATIENT COLLAPSED, HAD INAPPROPRIATE SHOCKS, AND EPISODES OF VF/VT. THE PATIENT PRESENTED TO THE EMERGENCY ROOM, BUT COULD NOT BE RESUSCITATED AND WAS PRONOUNCED DEAD. THE LEAD AN APPARENT LEAD FRACTURE. THE AUTHOR FEELS THAT THE PATIENT'S DEATH WAS DUE TO THE "INHIBITION OF PACING," FROM THE "TRANSIENT OR INTERMITTENT BRADYCARDIA/ASYSTOLE," VF EPISODES, AND THE SUDDEN INCREASE IN IMPEDANCE. |  | Y | Y | LAY USER/PATIENT
2649622-2008-03224 | 2008-05-14 |  | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03227 | 2008-05-15 | 20050101 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-03228 | 2008-05-15 | 20080111 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT PRESENTED TO THE EMERGENCY ROOM WITH SYMPTOMS OF NAUSEA AND VOMITING, WEAKNESS, AND SHORTNESS OF BREATH THAT HAD INCREASED SLIGHTLY OVER THE PAST SEVERAL MONTHS. CHEST X-RAY CONFIRMED A PLEURAL EFFUSION. THE PATIENT WAS GIVEN MEDICATION TO TREAT NAUSEA, AND WAS DISCHARGED HOME. TWO DAYS LATER THE PATIENT WAS FOUND DEAD IN BED. THE SYSTEM WAS NOT EXPLANTED. IT WAS UNKNOWN WHETHER THE PATIENT'S DEATH WAS SYSTEM RELATED. ADDITIONAL INFORMATION WAS RECEIVED REPORTING THAT THE CAUSE OF DEATH ON THE PATIENT'S DEATH CERTIFICATE IS CONGESTIVE HEART FAILURE DUE TO ATHEROSCLEROTIC HEART DISEASE. |  | Y | N | LAY USER/PATIENT
2649622-2008-02997 | 2008-05-08 | 20080328 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02985 | 2008-05-08 | 20070915 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02986 | 2008-05-08 | 20070904 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02989 | 2008-05-08 | 20080328 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6930 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR WAS FRACTURED; FULL LEAD WAS RETURNED FOR ANALYSIS. |  | N | Y | LAY USER/PATIENT
2649622-2008-02990 | 2008-05-08 | 20080312 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THAT THE PATIENT RECEIVED 14 INAPPROPRIATE SHOCKS, AND PACING IMPEDANCE WAS INCREASED/HIGH. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT, AND THE PATIENT'S STATUS WAS REPORTED AS "FINE" FOLLOWING THE REPLACEMENT PROCEDURE. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02992 | 2008-05-08 | 20080310 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES FOUND. OUTER INSULATION BREACHED (CLAVICLE-RIB CRUSH); DISTAL CONDUCTOR FRACTURED; LEAD APPEARS TO HAVE BEEN IMPLANTED WITH A BEND IN IT AT THE FRACTURE SITE; FULL LEAD ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02968 | 2008-05-08 | 20080328 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6930 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02969 | 2008-05-08 | 20080121 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6930 | ASKU |  | IT WAS REPORTED THAT THE PATIENT RECEIVED INAPPROPRIATE SHOCKS. THE LEAD WAS EXPLANTED AND REPLACED. THERE WERE NO FURTHER REPORTED PATIENT COMPLICATIONS AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02970 | 2008-05-08 | 20080212 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02971 | 2008-05-08 | 20071201 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02967 | 2008-05-08 | 20080327 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02955 | 2008-05-08 | 20070831 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02956 | 2008-05-08 | 20080227 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02957 | 2008-05-08 | 20080305 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY AND TESTED OUT OF SPECIFICATION CONSISTENT WITH THE ADVISORY. EVALUATION SUMMARY: DISTAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02958 | 2008-05-08 | 20080306 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. EVALUATION SUMMARY: NO ANOMALIES FOUND; FULL LEAD ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02959 | 2008-05-08 | 20080306 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE LEAD HAD INCREASING AND SUBSEQUENT HIGH HV IMPEDANCE. IT WAS ALSO REPORTED THAT DURING A TEST SHOCK, THE DEVICE DELIVERED LESS THAN EXPECTED SHOCK JOULES. THE PHYSICIAN DECIDED TO REPLACE THE LEAD. THERE WERE NO REPORTED PATIENT COMPLICATIONS AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-02960 | 2008-05-08 | 20080324 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02961 | 2008-05-08 | 20080317 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02962 | 2008-05-08 | 20080317 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED IN SEGMENTS AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02963 | 2008-05-08 | 20080320 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. EVALUATION SUMMARY: NO ANOMALIES FOUND; PARTIAL LEAD IN SEGMENTS RETURNED AND ANALYZED. THERE IS ONLY ONE SET OF SETSCREW MARKS ON THE IS-1 PIN AND IT IS TOO PROXIMAL, INDICATING THE LEAD WAS NOT FULLY INSERTED INTO THE DEVICE. |  | N | Y | LAY USER/PATIENT
2649622-2008-02964 | 2008-05-08 | 20080208 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02965 | 2008-05-08 | 20080212 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02966 | 2008-05-08 | 20080208 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02941 | 2008-05-08 | 20080320 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE, DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: DISTAL CONDUCTOR FRACTURED; THE ANALYST NOTED THAT THE LEAD APPEARED TO HAVE BEEN IMPLANTED WITH A BEND IN IT AT THE FRACTURE SITE; FULL LEAD RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02942 | 2008-05-08 | 20080320 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES FOUND; FULL LEAD RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02943 | 2008-05-08 | 20070514 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02946 | 2008-05-08 | 20070829 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02948 | 2008-05-08 | 20080331 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT RECEIVED INAPPROPRIATE SHOCKS DUE TO OVERSENSING. FURTHER FOLLOW-UP INDICATED THAT THERE WAS A LEAD FRACTURE. THE LEAD WAS REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02952 | 2008-05-08 | 20080307 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES FOUND; THE ANALYST NOTED THAT THE ONLY DAMAGE FOUND THAT MAY HAVE BEEN CONSIDERED INSULATION DAMAGE IS THE CAUTERY MELT OF THE OVERLAY TUBING, WHICH BY DESIGN, IS NOT CONSIDERED INSULATION, AND A BREACH WOULD NOT AFFECT ELECTRICAL PERFORMANCE. PROXIMAL SEGMENT RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02940 | 2008-05-08 | 20080319 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02928 | 2008-05-08 | 20080311 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02929 | 2008-05-08 | 20080326 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02930 | 2008-05-08 | 20080326 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THAT THE PATIENT RECEIVED 3 SHOCKS DUE TO OVERSENSING. A LEAD FRACTURE WAS SUSPECTED. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT, AND THE PATIENT'S STATUS WAS REPORTED AS "OK". |  | Y | Y | LAY USER/PATIENT
2649622-2008-02931 | 2008-05-08 | 20080213 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02932 | 2008-05-08 | 20060206 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-02933 | 2008-05-08 | 20080403 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT WAS TAKEN BY AMBULANCE TO HOSPITAL FOR LEAD REPLACEMENT. THE LEAD SHOWED OVERSENSING WITH INCREASING AND SUBSEQUENT HIGH IMPEDANCE. A LEAD FRACTURE WAS SUSPECTED. THERE WERE NO REPORTED PATIENT COMPLICATIONS AS A RESULT OF THIS EVENT. FURTHER INFORMATION RECEIVED FROM AN ATTORNEY ALLEGES PATIENT SUFFERED PHYSICAL AND OTHER INJURY AS A RESULT OF THE RECALLED LEAD. PATIENT EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY REQUIRING MEDICAL INTERVENTION INCLUDING BUT NOT LIMITED TO SURGERY, REMOVAL AND/OR REPLACEMENT OF THE DEFECTIVE SPRINT FIDELIS LEAD. ATTORNEY ALSO ALLEGES THE PATIENT "HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH". |  | Y | Y | LAY USER/PATIENT
2649622-2008-02934 | 2008-05-08 | 20080311 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6930 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02935 | 2008-05-08 | 20080313 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02936 | 2008-05-08 | 20080313 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02937 | 2008-05-08 | 20080315 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THAT THERE WAS OVERSENSING AND A SUDDEN JUMP IN IMPEDANCE TO 1696 OHMS. THE LEAD WAS EXPLANTED AND REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT, AND THE PATIENT'S STATUS WAS REPORTED AS "OK" FOLLOWING THE LEAD REPLACEMENT PROCEDURE. |  | N | Y | LAY USER/PATIENT
2649622-2008-02938 | 2008-05-08 | 20080317 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED. OTHER: IT WAS REPORTED THAT THE PATIENT HEARD THE PATIENT ALERT AND WENT TO THE HOSPITAL, WHEREUPON IT WAS DETERMINED THERE WAS HIGH IMPEDANCE GREATER THAN 2000 OHMS. THE LEAD WAS EXPLANTED AND REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT, AND THE PATIENT'S STATUS WAS REPORTED AS "NORMAL" FOLLOWING THE REPLACEMENT PROCEDURE. ACTUAL DEVICE INVOLVED IN INCIDENT WAS EVALUATED. ELECTRICAL TESTS PERFORMED; MECHANICAL TESTS PERFORMED. VISUAL EXAMINATION: COMPONENT/SUBASSEMBLY FAILURE. DEVICE WAS OUT OF SPECIFICATION IN A MANNER THAT RELATES TO EVENT, IMPEDANCE, HIGH. |  | N | Y | LAY USER/PATIENT
2649622-2008-02939 | 2008-05-08 | 20080318 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY AND TESTED OUT OF SPECIFICATION CONSISTENT WITH THE ADVISORY. EVALUATION SUMMARY: DISTAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02926 | 2008-05-08 | 20080322 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02914 | 2008-05-08 | 20070925 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02916 | 2008-05-08 | 20080326 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THERE WAS A POSSIBLE LEAD FRACTURE AND THAT THE PATIENT RECEIVED 35 INAPPROPRIATE THERAPIES DUE TO OVERSENSING. IT WAS FURTHER REPORTED THAT THERE WAS HIGH IMPEDANCE GREATER THAN 2500 OHMS. THE LEAD WAS CAPPED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. NO CONCLUSION CAN BE DRAWN IMPEDANCE, HIGH LEAD(S), FRACTURE OF OVERSENSING SHOCK, INAPPROPRIATE. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02917 | 2008-05-08 | 20080328 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02918 | 2008-05-08 | 20080324 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02919 | 2008-05-08 | 20080331 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THE PATIENT ALERT SOUNDED FOR LOW IMPEDANCE. THE LEAD WAS EXPLANTED AND REPLACED DUE TO AN APPARENT LEAD FRACTURE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-02920 | 2008-05-08 | 20080331 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02921 | 2008-05-08 | 20080310 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE, DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02922 | 2008-05-08 | 20080312 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE, DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES FOUND; FULL LEAD IN SEGMENTS ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02923 | 2008-05-08 | 20080331 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02924 | 2008-05-08 | 20080303 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | IT WAS REPORTED THAT THERE WAS HIGH IMPEDANCE, OVERSENSING, NOISE, AND A LIKELY ISSUE WITH A SETSCREW. THE DOCTOR DECIDED TO REPLACE THE LEAD IN VIEW OF THE FIDELIS ADVISORY. THE LEAD WAS EXPLANTED AND REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT, AND THE PATIENT'S STATUS WAS REPORTED AS 'WELL'. |  | N | Y | LAY USER/PATIENT
2649622-2008-02925 | 2008-05-08 | 20080320 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE, DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02913 | 2008-05-08 | 20080327 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT ALERT SOUNDED, AND THE PATIENT WENT TO THE HOSPITAL. INCREASED IMPEDANCE AND SIC (SENSING INTEGRITY COUNTER) COUNT OF 13 SHORT V-V INTERVALS WITHIN ONE DAY WAS NOTED. THE ICD WAS DISABLED PENDING LEAD REVISION. THE LEAD WAS SUBSEQUENTLY EXPLANTED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-02905 | 2008-05-08 |  | No answer provided | Manufacturer report | SPRINT FIDELIS | ASKU | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | 0
2649622-2008-02906 | 2008-05-08 | 20080328 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES FOUND; PROXIMAL SEGMENT RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02909 | 2008-05-08 | 20080229 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-02910 | 2008-05-08 | 20080307 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-02911 | 2008-05-08 | 20080313 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | IT WAS REPORTED THE LEAD WAS EXPLANTED AND REPLACED DUE TO INAPPROPRIATE THERAPIES, OVERSENSING AND HIGH AND VARYING IMPEDANCES. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02912 | 2008-05-08 | 20080318 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | N | LAY USER/PATIENT
2649622-2008-02902 | 2008-05-08 | 20080303 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE RIGHT VENTRICULAR LEAD WAS EXPLANTED DUE TO HIGH IMPEDANCE AND AN APPARENT LEAD FRACTURE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED IN A LAWSUIT ALLEGING THAT THE PATIENT 'SUFFERED PHYSICAL AND OTHER INJURY AS A RESULT OF THE RECALLED LEAD AND DEFENDANT'S CONDUCT. [PATIENT] EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY REQUIRING MEDICAL INTERVENTION INCLUDING BUT NOT LIMITED TO SURGERY, REMOVAL AND/OR REPLACEMENT' OF THE 'DEFECTIVE' LEAD. IT FURTHER ALLEGED THAT THE PATIENT HAS 'SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH, ECONOMIC LOSSES AND OTHER DAMAGES. [PATIENT] HAS SUFFERED PHYSICAL INJURIES AND VARIOUS PHYSICAL MANIFESTATIONS OF EMOTIONAL DISTRESS'. LEAD 'FAILURE' WAS ALSO REPORTED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02892 | 2008-05-08 | 20080311 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES FOUND; FULL LEAD ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02893 | 2008-05-08 | 20080326 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02895 | 2008-05-08 | 20080312 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02896 | 2008-05-08 | 20080327 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02897 | 2008-05-08 | 20080228 | Injury | Manufacturer report | SPRINT FIDELIS | DEFIBRILLATOR, AUTOMATIC IMPLANTABLE CARDIOVERTER | MEDTRONIC PUERTO RICO, INC. | 694965 | 694965 |  | THIS DEVICE WAS REPORTED AS INCLUDED IN THE FIELD ACTION. BASED ON THE INFORMATION RECEIVED AND WITHOUT THE RETURN OF THE PRODUCT, IT CANNOT BE CONFIRMED THAT THIS DEVICE PERFORMED AS DESCRIBED IN THE FIELD ACTION. IT IS INCLUDED IN THE FIELD ACTION IN THE ABUNDANCE OF CAUTION. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02899 | 2008-05-08 | 20080324 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES FOUND; FULL LEAD ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02901 | 2008-05-08 | 20080326 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | OVERSENSING AND NOISE WERE REPORTED. THE LEAD WAS EXPLANTED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. IT WAS FURTHER REPORTED THAT THE PATIENT HAD PRESENTED TO THE ED AFTER RECEIVING SEVERAL SHOCKS. THE PATIENT HAD RECEIVED A TOTAL OF 55 SHOCKS. IT APPEARED THAT THE PATIENT HAD NO PACING DUE TO OVERSENSING. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02881 | 2008-05-08 | 20080324 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS EXPLANTED AND RETURNED DUE TO OVERSENSING, INAPPROPRIATE THERAPIES AND AN APPARENT LEAD FRACTURE. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED REPORTING AN ALLEGATION FROM AN ATTORNEY INDICATED THE PATIENT IS DECEASED AND FURTHER INDICATED THAT THE LEAD HAD EXHIBITED A FRACTURE AND/OR WAS EXPLANTED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02885 | 2008-05-08 | 20080325 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. ANALYSIS OF THE LEAD IS IN PROCESS; RESULTS WILL BE FORWARDED WHEN AVAILABLE. |  | N | N | LAY USER/PATIENT
2649622-2008-02886 | 2008-05-08 | 20071212 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02887 | 2008-05-08 | 20080316 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02888 | 2008-05-08 | 20080313 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02889 | 2008-05-08 | 20080123 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ATTORNEY ALLEGES 'CLAIM FOR DAMAGES ARISING OUT OF DEFECTS IN AND REMOVAL OF THE ABOVE-REFERENCED EQUIPMENT." THE LEAD WAS EXPLANTED AND REPLACED. ADDITIONAL INFORMATION FROM THE ATTORNEY ALLEGES THAT BECAUSE OF THE DEFECTIVE LEAD, THE PATIENT "RECEIVED UNNECESSARY AND INAPPROPRIATE SHOCKS." IT WAS FURTHER ALLEGED BY THE ATTORNEY THE PATIENT "SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH."ATTORNEY ALLEGES 'CLAIM FOR DAMAGES ARISING OUT OF DEFECTS IN AND REMOVAL OF THE ABOVE-REFERENCED EQUIPMENT." THE LEAD WAS EXPLANTED AND REPLACED. ADDITIONAL INFORMATION FROM THE ATTORNEY REPORTS THAT BECAUSE OF THE LEAD DEFECT, THE PATIENT "RECEIVED UNNECESSARY AND INAPPROPRIATE SHOCKS." IT WAS FURTHER REPORTED BY THE ATTORNEY THE PATIENT "SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH." |  | Y | N | LAY USER/PATIENT
2649622-2008-02890 | 2008-05-08 | 20080320 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. EVALUATION SUMMARY: THE ACTUAL LEAD WAS NOT RECEIVED FOR EVALUATION, HOWEVER, PERFORMANCE DATA WAS COLLECTED FROM THE DEVICE AND ANALYZED. ANALYSIS OF THAT DATA REVEALED AN ELEVATED VENTRICULAR PACING IMPEDANCE OF 1232 OHMS AND AN ELEVATED SENSING INTEGRITY COUNTER. A HIGH VALUE OF THIS COUNT CAN INDICATE A POSSIBLE INTERMITTENCY IN THE SYSTEM. ADDENDUM: THE LEAD HAS SINCE BEEN RECEIVED FOR ANALYSIS. DETAILED ANALYSIS OF THE LEAD WAS NOT PERFORMED AT THIS TIME DUE TO PENDING LITIGATION. THEREFORE, WE ARE UNABLE TO FULLY EVALUATE THE REPORTED EVENT. OTHER: A CALL WAS RECEIVED THROUGH TECHNICAL SERVICES REPORTING THE PATIENT HAD RECEIVED MORE THAN 30 SHOCKS BEFORE DETECTIONS WERE TURNED OFF AT THE HOSPITAL. BECAUSE THE HOSPITAL ONLY HAD SMALL MAGNETS, THREE WERE USED TO SUSPEND THERAPIES. REVIEW OF THE PATIENT'S HISTORY INDICATED THE PATIENT HAD HEARD THE PATIENT ALERT 6 DAYS BEFORE THE 1ST SHOCK BUT DID NOT REPORT TO THE HOSPITAL AS INSTRUCTED. THE LEAD WAS EXPLANTED AND REPLACED. ADDITIONAL INFORMATION WAS SUBSEQUENTLY RECEIVED REPORTING RV (RIGHT VENTRICULAR) PACING IMPEDANCE INCREASE, SHORT V-V INTERVALS, INDICATIVE OF OVERSENSING, AND SUSPECTED LEAD FRACTURE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. IMPEDANCE, HIGH, LEAD(S), FRACTURE OF, OVERSENSING; SHOCK, INAPPROPRIATE. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02880 | 2008-05-08 | 20060829 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. THE DEVICE IS PART OF THE ADVISORY FOR THIS MODEL. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02867 | 2008-05-08 | 20080331 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED BY THE PATIENT'S WIFE THAT THE PATIENT EXPERIENCED TWO SHOCKS WHILE MAKING THE BED. SHE ALSO REPORTS PATIENT IS FEELING OK EXCEPT FOR SORENESS IN THE IMPLANT AREA. ADDITIONAL INFORMATION RECEIVED INDICATES THAT THE LEAD WAS REPLACED DUE TO AN APPARENT FRACTURE. THE PHYSICIAN REPORTED THAT THE PATIENT WAS IN A "SERIOUS ATV ACCIDENT WHICH MAY HAVE PREFACED THE FRACTURE." THERE WERE NO FURTHER REPORTED PATIENT COMPLICATIONS AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02868 | 2008-05-08 | 20080321 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02869 | 2008-05-08 | 20080311 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02870 | 2008-05-08 | 20080311 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; PARTIAL LEAD IN SEGMENTS RETURNED AND ANALYZED |  | Y | Y | LAY USER/PATIENT
2649622-2008-02874 | 2008-05-08 | 20080218 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. OTHER : IT WAS REPORTED THAT THE PATIENT RECEIVED MULTIPLE SHOCKS. HIGH IMPEDANCE AND SENSING DIFFICULTY WAS INDICATED. THERE WERE MANY V-V INTERVALS SENSED, INDICATIVE OF OVERSENSING. THE LEAD WAS EXPLANTED AND REPLACED. THERE WERE NO FURTHER REPORTED PATIENT COMPLICATIONS AS A RESULT OF THIS EVENT. ACTUAL DEVICE INVOLVED IN INCIDENT WAS EVALUATED ELECTRICAL TESTS PERFORMED MECHANICAL TESTS PERFORMED; VISUAL EXAMINATION COMPONENT/SUBASSEMBLY FAILURE. LEAD CONDUCTOR DEVICE WAS OUT OF SPECIFICATION IN A MANNER THAT RELATES TO EVENT IMPEDANCE, HIGH OVERSENSING SENSITIVITY SHOCK, INAPPROPRIATE. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02875 | 2008-05-08 | 20080319 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02877 | 2008-05-08 | 20080311 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02878 | 2008-05-08 | 20080317 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6930 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02879 | 2008-05-08 | 20080319 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. DISTAL CONDUCTOR FRACTURED; FULL LEAD IN SEGMENTS RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02866 | 2008-05-08 | 20060731 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02857 | 2008-05-08 | 20080305 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. DISTAL CONDUCTOR FRACTURED; DISTAL SEGMENT RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02858 | 2008-05-08 | 20080215 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE RV LEAD HAD MALFUNCTIONED. THE LEAD WAS EXPLANTED AND REPLACED. THERE WERE NO REPORTED PATIENT COMPLICATIONS AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-02859 | 2008-05-08 | 20080317 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: NO ANOMALIES FOUND. PROXIMAL SEGMENT RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02860 | 2008-05-08 | 20070920 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | A CLAIM, SUBMITTED BY AN ATTORNEY, ALLEGED THE PATIENT' DEATH WAS "CAUSED BY THE [MALFUNCTIONING] OF THE [IMPLANTED DEVICES]." NO ADDITIONAL AND/OR SPECIFIC PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02861 | 2008-05-08 | 20080325 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  |  |  | N | N | LAY USER/PATIENT
2649622-2008-02863 | 2008-05-08 | 20080317 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT RECEIVED MULTIPLE SHOCKS DUE TO A LEAD FRACTURE. THE LEAD WAS REPLACED. THERE WERE NO FURTHER REPORTED PATIENT COMPLICATIONS AS A RESULT OF THIS EVENT. IT WAS LATER ALLEGED BY AN ATTORNEY THAT THE LEAD HAD EXHIBITED A FRACTURE AND/OR WAS EXPLANTED. ADDITIONALLY IT WAS ALLEGED THE PATIENT IS DECEASED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02865 | 2008-05-08 | 20070616 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02845 | 2008-05-08 | 20080128 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02847 | 2008-05-08 | 20080321 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: DISTAL CONDUCTOR FRACTURED; FULL LEAD RETURNED IN SEGMENTS AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02848 | 2008-05-08 | 20080229 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-02849 | 2008-05-08 | 20080317 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02852 | 2008-05-08 | 20080317 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02853 | 2008-05-08 | 20080214 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02854 | 2008-05-08 | 20080318 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02830 | 2008-05-08 | 20080128 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: NO ANOMALIES FOUND; FULL LEAD ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02831 | 2008-05-08 | 20071106 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES FOUND; FULL LEAD RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02832 | 2008-05-08 | 20080331 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD IN SEGMENTS RETURNED AND ANALYZED. IT WAS REPORTED THAT THERE WAS OVERSENSING. IT WAS FURTHER REPORTED THAT THERE WAS AN ABRUPT CHANGE IN THE IMPEDANCE TREND. THE LEAD WAS EXPLANTED AND REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ACTUAL DEVICE INVOLVED IN INCIDENT WAS EVALUATED ELECTRICAL TESTS PERFORMED, MECHANICAL TESTS PERFORMED, VISUAL EXAMINATION LEAD CONDUCTOR COMPONENT/SUBASSEMBLY FAILURE DEVICE WAS OUT OF SPECIFICATION IN A MANNER THAT RELATES TO EVENT OVERSENSING. |  | N | Y | LAY USER/PATIENT
2649622-2008-02833 | 2008-05-08 | 20080229 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE LEAD SHOWED OVERSENSING. THE LEAD WAS EXPLANTED AND REPLACED. THERE WERE NO REPORTED PATIENT COMPLICATIONS AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-02834 | 2008-05-08 | 20080226 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT DURING DEVICE REPLACEMENT, A HIGH VOLTAGE SHOCK WAS INEFFECTIVE AND THE RV LEAD IMPEDANCE WAS HIGH, > 200 OHMS. THE LEAD WAS EXPLANTED AND REPLACED. THERE WERE NO PATIENT COMPLICATIONS REPORTED AS A RESULT OF THIS EVENT. THE PATIENT STATUS WAS NOTED AS 'OK' |  | N | Y | LAY USER/PATIENT
2649622-2008-02835 | 2008-05-08 | 20080306 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02836 | 2008-05-08 | 20080307 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS EXPLANTED AND REPLACED DUE TO INAPPROPRIATE SHOCKS, HIGH IMPEDANCE (>3000 OMS), OVERSENSING, AND SUSPECTED LEAD FRACTURE. ADDITIONALLY, DURING THE LEAD EXTRACTION THE PROXIMAL VENTRICULAR SUBCLAVIA WAS PUNCTURED AS THE STYLET COULD NOT BE PUSHED FORWARD. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02837 | 2008-05-08 | 20080307 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02838 | 2008-05-08 | 20080311 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02839 | 2008-05-08 | 20080310 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. NO ANOMALIES FOUND; PROXIMAL SEGMENT RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02840 | 2008-05-08 | 20080206 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02841 | 2008-05-08 | 20080311 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE PATIENT WAS ADMITTED TO THE INTENSIVE CARE UNIT DUE TO SEVERAL INAPPROPRIATE SHOCKS. IN ADDITION, THE LEAD EXHIBITED HIGH IMPEDANCES, HIGH THRESHOLDS, AND OVERSENSING. THE LEAD WAS EXPLANTED AND REPLACED DURING THE PATIENT'S HOSPITALIZATION. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02829 | 2008-05-08 | 20080305 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02817 | 2008-05-08 | 20080117 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: NO ANOMALIES FOUND; FULL LEAD RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02818 | 2008-05-08 | 20080311 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02819 | 2008-05-08 | 20080312 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT RECEIVED SEVERAL SHOCKS DUE TO LEAD FRACTURE. IT WAS FURTHER REPORTED THAT THERE WAS "ARTIFACT" AND HIGH IMPEDANCE GREATER THAN 3000 OHMS. THE LEAD WAS CAPPED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02820 | 2008-05-08 | 20080313 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02821 | 2008-05-08 | 20080204 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED BY THE PATIENT'S SON THAT THE PATIENT HAD SUSTAINED "SEVERAL ELECTROCUTIONS." THE LEAD WAS REPLACED. ADDITIONALLY ALLEGED THAT THE [PATIENT] "WAS IMPLANTED WITH A [LEAD WIRE SYSTEM], AND SUFFERED PHYSICAL AND OTHER INJURY AS A RESULT OF THE RECALLED LEAD." ALSO ALLEGED WAS THAT THE [PATIENT] EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY." FURTHER ALLEGED WAS THAT THE PATIENT" HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH, ECONOMIC LOSSES AND OTHER DAMAGES." THE [PATIENT] HAS FURTHER"SUFFERED PHYSICAL INJURIES AND VARIOUS PHYSICAL MANIFESTATIONS OF EMOTIONAL DISTRESS ASSOCIATED WITH ONE OR MORE OF THE FOLLOWING: THE IMPLANTATION, RECALL, FAILURE, REMOVAL/REPLACEMENT, AND/OR INABILITY TO HAVE THE DEFECTIVE [LEAD] REMOVED OR REPLACED." |  | Y | Y | LAY USER/PATIENT
2649622-2008-02822 | 2008-05-08 | 20050413 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02823 | 2008-05-08 | 20080318 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02824 | 2008-05-08 | 20080320 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT ALERT HAD BEEN GOING OFF FOR A COUPLE DAYS. INTERROGATION INDICATED THAT THE LEAD SVC IMPEDANCE WAS > 200 OHMS. THE LEAD WAS REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION RECEIVED REPORTED COIL FRACTURE. |  | N | Y | LAY USER/PATIENT
2649622-2008-02828 | 2008-05-08 | 20080317 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02816 | 2008-05-08 | 20080315 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02804 | 2008-05-08 | 20070208 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02805 | 2008-05-08 | 20080312 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE PATIENT HAD SMALL R WAVES. A NON INVASIVE PROGRAMMED STIMULATION WAS PERFORMED THAT CAUSED A LOT OF NOISE ON THE LEAD. THE LEAD WAS PARTIALLY EXPLANTED AND CAPPED AND REPLACED. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION RECEIVED REPORTED THERE WERE ISSUES WITH OVERSENSING ON THE LEAD PRIOR TO REPLACEMENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-02806 | 2008-05-08 | 20080313 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02807 | 2008-05-08 | 20080316 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02808 | 2008-05-08 | 20080312 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES FOUND; PROXIMAL SEGMENT RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02809 | 2008-05-08 | 20080310 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: DISTAL CONDUCTOR FRACTURED; DISTAL SEGMENT RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02811 | 2008-05-08 | 20080308 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02812 | 2008-05-08 | 20080310 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THERE WAS OVERSENSING, ELEVATED IMPEDANCE, AND THAT THE PATIENT RECEIVED TWO INAPPROPRIATE THERAPIES. IT WAS ALSO REPORTED THAT AT THE TIME OF EXPLANT, THE HELIX WOULD NOT RETRACT. IT WAS FURTHER REPORTED THAT THERE WAS AN APPARENT LEAD FRACTURE. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02813 | 2008-05-08 | 20080313 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS EXPLANTED AND REPLACED DUE TO HIGH IMPEDANCE, OVERSENSING, AND SUSPECTED FRACTURE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-02814 | 2008-05-08 | 20071031 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02564 | 2008-05-08 | 20080226 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT PRESENTED TO THE EMERGENCY DEPARTMENT DUE TO THE PATIENT ALERT BEING TRIGGERED. INTERROGATION SHOWED HIGH IMPEDENCE, OF GREATER THAN 2500 OHMS, AND THE SIC (SHORT INTERVAL COUNTER) VALUE INDICATED OVERSENSING. THE LEAD WAS REMOVED AND REPLACED. THE LEAD WAS SUBSEQUENTLY RETURNED REPORTING "LEAD'S INSULATION WAS BREAKING." NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02566 | 2008-05-08 | 20080122 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: DEFIB CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02567 | 2008-05-08 | 20071105 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-02568 | 2008-05-08 | 20080226 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02570 | 2008-05-08 | 20080221 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-02571 | 2008-05-08 | 20080218 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THAT THE PATIENT HEARD THE PATIENT ALERT AND RECEIVED SEVERAL SHOCKS DUE TO NOISE WITHIN ONE DAY. THE LEAD WAS REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. A (B) (6) FURTHER ALLEGED THAT THE PATIENT EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING AND LEAD FRACTURE. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02573 | 2008-05-08 | 20080221 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02562 | 2008-05-08 | 20080229 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THERE WAS NOISE ON THE LEAD AND THE RV PACING IMPEDANCE HAD "SPIKED." THE PATIENT PRESENTED TO THE EMERGENCY DEPARTMENT COMPLAINING OF MULTIPLE SHOCKS. THE LEAD WAS REPLACED. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION RECEIVED REPORTED THE LEAD HAD BEEN OVERSENSING AND A FRACTURE WAS SUSPECTED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02574 | 2008-05-08 | 20080226 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | A CALL WAS RECEIVED THROUGH TECHNICAL SERVICES REPORTING THAT THE PATIENT ALERT SOUNDED DUE TO HIGH IMPEDANCE GREATER THAN 1000 OHMS; VARYING IMPEDANCE BETWEEN GREATER THAN 2000 OHMS AND LESS THAN 700 OHMS. THE LEAD WAS REPLACED. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-02549 | 2008-05-08 | 20080211 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS EXPLANTED DUE TO INAPPROPRIATE SHOCKS. FURTHER ALLEGED THAT THE PATIENT "EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING" AND THAT THERE WAS A LEAD FRACTURE. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02551 | 2008-05-08 | 20080220 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02553 | 2008-05-08 | 20080220 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02554 | 2008-05-08 |  | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02556 | 2008-05-08 | 20080219 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-02557 | 2008-05-08 | 20080219 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02558 | 2008-05-08 | 20080219 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-02559 | 2008-05-08 | 20080222 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THERE WAS OVERSENSING LEADING TO INAPPROPRIATE SHOCKS. FURTHER ALLEGED THAT THE PATIENT "EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING" AND THAT THERE WAS A LEAD FRACTURE. THE LEAD WAS REPLACED. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02560 | 2008-05-08 | 20080218 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02537 | 2008-05-08 | 20080228 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | . |  | Y | Y | LAY USER/PATIENT
2649622-2008-02538 | 2008-05-08 | 20080226 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS EXPLANTED DUE TO HIGH IMPEDANCE, OVERSENSING, AND APPARENT FRACTURE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-02540 | 2008-05-08 |  | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02541 | 2008-05-08 | 20080220 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02545 | 2008-05-08 | 20080222 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02547 | 2008-05-08 | 20080218 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02536 | 2008-05-08 | 20080225 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS EXPLANTED AND REPLACED DUE TO 18 INAPPROPRIATE SHOCKS, OVERSENSING, AND FRACTURE. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02525 | 2008-05-08 | 20080205 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02526 | 2008-05-08 | 20080227 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02527 | 2008-05-08 | 20080221 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | IT WAS REPORTED THERE WERE INCREASED THRESHOLDS, AND THE PATIENT HAD BEEN HOSPITALIZED TWICE FOR SYNCOPE, DUE TO LOSS OF CAPTURE. THE SYSTEM WAS REPLACED AND MOVED TO THE RIGHT SIDE BECAUSE THE PATIENT DIDN'T FEEL COMFORTABLE WITH THE DEVICE ON THE LEFT SIDE. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT, AND THE PATIENT STATUS WAS NOTED TO BE OK FOLLOWING THE REPLACEMENT PROCEDURE. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02528 | 2008-05-08 | 20080218 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02529 | 2008-05-08 | 20071221 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02530 | 2008-05-08 | 20080313 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02531 | 2008-05-08 | 20080215 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-02532 | 2008-05-08 | 20080303 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02524 | 2008-05-08 | 20080225 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; DISTAL SEGMENT RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02535 | 2008-05-08 | 20080221 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES FOUND; FULL LEAD RETURNED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02513 | 2008-05-08 | 20080214 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02514 | 2008-05-08 | 20080218 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02516 | 2008-05-08 | 20080219 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02517 | 2008-05-08 | 20080229 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02518 | 2008-05-08 | 20080229 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (NON-CRT) | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  | IT WAS REPORTED THAT THE PATIENT ALERT SOUNDED DUE TO HIGH IMPEDANCE. THE LEAD WAS EXPLANTED AND REPLACED. ADDITIONAL INFORMATION RECEIVED PER THE SALES REP IS THE NOTATION OF A FRACTURE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02519 | 2008-05-08 | 20080303 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02520 | 2008-05-08 | 20080220 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE, DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; PARTIAL LEAD IN SEGMENTS RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02521 | 2008-05-08 | 20080220 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02522 | 2008-05-08 | 20080219 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE, DUE TO CONFIDENTIALITY CONCERNS. EVALUATION SUMMARY: NO ANOMALIES FOUND; FULL LEAD RETURNED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02512 | 2008-05-08 | 20080214 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THE LEAD WAS CAPPED AND REPLACED DUE TO HIGH IMPEDANCES, OVERSENSING AND INAPPROPRIATE THERAPIES. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02499 | 2008-05-08 | 20071127 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY DISTAL CONDUCTOR FRACTURED; PARTIAL LEAD RETURNED IN SEGMENTS. |  | N | Y | LAY USER/PATIENT
2649622-2008-02500 | 2008-05-08 | 20080221 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES FOUND; PARTIAL LEAD IN SEGMENTS RETURNED AND ANALYZED. EVALUATION SUMMARY: PARTIAL LEAD IN SEGMENTS. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02501 | 2008-05-08 | 20070525 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02502 | 2008-05-08 | 20080114 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02510 | 2008-05-08 | 20061114 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02511 | 2008-05-08 | 20080208 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. EVALUATION SUMMARY: NO ANOMALIES FOUND; DISTAL SEGMENT RETURNED. IT WAS REPORTED THAT THE PATIENT RECEIVED INAPPROPRIATE SHOCKS, AND THERE WAS A FAILURE TO SHOCK DURING 5 SUCCESSIVE THERAPIES. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT, AND THE PATIENT'S STATUS WAS REPORTED AS "STABLE" FOLLOWING THE REPLACEMENT PROCEDURE. ACTUAL DEVICE INVOLVED IN INCIDENT WAS EVALUATED ELECTRICAL TESTS PERFORMED, MECHANICAL TESTS PERFORMED, VISUAL EXAMINATION DEVICE PERFORMED, ACCORDING TO SPECIFICATIONS NO CONCLUSION CAN BE DRAWN SHOCK, INAPPROPRIATE DELIVER, FAILURE TO. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02488 | 2008-05-08 | 20071205 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02489 | 2008-05-08 | 20070801 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02493 | 2008-05-08 | 20080222 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: NO ANOMALIES FOUND; FULL LEAD RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02495 | 2008-05-08 | 20080304 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: DISTAL CONDUCTOR FRACTURED; DISTAL SEGMENT RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02496 | 2008-05-08 | 20071009 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE LEAD HAD HIGH IMPEDANCE AND OVERSENSING. THE PACE/SENSE PORTION OF THE LEAD WAS CAPPED AND REPLACED. ADDITIONAL INFORMATION RECEIVED ALLEGES THAT THE PATIENT "SUFFERED PHYSICAL AND OTHER INJURY" AS A RESULT OF THE LEAD. ALSO ALLEGES THAT THE PATIENT "EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY REQUIRING MEDICAL INTERVENTION INCLUDING BUT NOT LIMITED TO SURGERY, REMOVAL AND/OR REPLACEMENT OF THE DEFECTIVE [LEAD]." IT ALSO ALLEGED THE PATIENT "HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH, ECONOMIC LOSSES AND OTHER DAMAGES. [PATIENT] HAS FURTHER SUFFERED PHYSICAL INJURIES AND VARIOUS PHYSICAL MANIFESTATIONS OF EMOTIONAL DISTRESS". LEAD "FAILURE" WAS ALSO NOTED. |  | Y | Y | HEALTH PROFESSIONAL
2649622-2008-02497 | 2008-05-08 | 20070812 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS OVERSENSING, HAD HIGH IMPEDANCE AND THE PATIENT RECEIVED MULTIPLE SHOCKS. IN ADDITION, AN ATTORNEY ALLEGED THE PATIENT WAS SHOCKED 16 TIMES BEFORE HE COULD GET TO THE HOSPITAL TO DISABLE THERAPIES. THE PACE/SENSE PORTION OF THE LEAD WAS CAPPED AND REPLACED. ADDITIONAL INFORMATION RECEIVED INDICATES THAT A LAWSUIT ALLEGES THAT THE PATIENT "EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY." THERE WERE NO FURTHER PATIENT COMPLICATIONS AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02475 | 2008-05-08 | 20080227 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES FOUND; PROXIMAL SEGMENT RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02476 | 2008-05-08 | 20080218 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02477 | 2008-05-08 | 20080226 | Injury | Manufacturer report | SPRINT FIDELIS | DEFIBRILLATOR, AUTOMATIC IMPLANTABLE CARDIOVERTER | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS REPLACED DUE TO APPARENT FRACTURE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION RECEIVED REPORTED 3 INAPPROPRIATE SHOCKS. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02478 | 2008-05-08 | 20080214 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02479 | 2008-05-08 | 20080214 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE LEAD WAS "NON-FUNCTIONAL" AND THAT THERE WAS NO CAPTURE. AN ATTORNEY LETTER FURTHER ALLEGED THAT THE PATIENT SUSTAINED "INJURY". FURTHER FOLLOW-UP INDICATED THAT THERE WAS A LEAD FRACTURE. THE LEAD WAS REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02480 | 2008-05-08 | 20061127 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY, AND TESTED OUT OF SPECIFICATION CONSISTENT WITH THAT ADVISORY. EVALUATION SUMMARY: DISTAL CONDUCTOR WAS FRACTURED; PARTIAL LEAD IN SEGMENTS WAS RETURNED FOR ANALYSIS. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02481 | 2008-05-08 | 20080226 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT RECEIVED SEVEN INAPPROPRIATE SHOCKS. LEAD FRACTURE, OVERSENSING, AND PACING IMPEDANCE OF 1152 OHMS WAS ALSO NOTED. THE LEAD WAS PARTIALLY EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02482 | 2008-05-08 | 20080227 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THERE WAS OVERSENSING AND HIGH IMPEDANCE. THE PATIENT'S HUSBAND FURTHER REPORTED THAT THERE WAS A LEAD FRACTURE. THE LEAD WAS EXPLANTED AND REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-02483 | 2008-05-08 | 20080206 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02484 | 2008-05-08 | 20080207 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02485 | 2008-05-08 |  | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02464 | 2008-05-08 | 20080111 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02465 | 2008-05-08 | 20080222 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. DETAILED ANALYSIS OF THE DEVICE WAS NOT PERFORMED AT THIS TIME DUE TO PENDING LITIGATION. THEREFORE, WE ARE UNABLE TO FULLY EVALUATE THE REPORTED EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-02466 | 2008-05-08 | 20080226 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; PARTIAL LEAD IN SEGMENTS RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02467 | 2008-05-08 | 20080214 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02468 | 2008-05-08 | 20080220 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02469 | 2008-05-08 | 20080220 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02470 | 2008-05-08 | 20080211 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02471 | 2008-05-08 | 20080205 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THERE WAS OVERSENSING. IT WAS FURTHER REPORTED THAT THE PATIENT WAS PACER DEPENDENT AND THAT THE OVERSENSING WAS CAUSING PACING BELOW THE LOWER RATE. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED ALLEGING THAT THE PATIENT "EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, OR FAILURE TO RECEIVE THERAPEUTIC SHOCKS. EACH OF WHICH HAVE REQUIRED THEM TO SEEK MEDICAL INTERVENTION RESULTING IN REPLACEMENT OF THE DEFECTIVE [LEAD]". IT ALSO ALLEGED THAT THE PATIENT "SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH, ECONOMIC LOSSES AND OTHER DAMAGES". |  | Y | Y | LAY USER/PATIENT
2649622-2008-02472 | 2008-05-08 | 20080221 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02462 | 2008-05-08 | 20080213 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED. IT WAS OBSERVED DURING ANALYSIS THAT THERE WAS ONLY ONE SET OF SETSCREW MARKS ON THE IS-1 PIN, AND IT WAS TOO PROXIMAL, WHICH INDICATES THE LEAD WAS NOT FULLY INSERTED INTO THE DEVICE. DISCLAIMER: SUBMISSION OF INFORMATION BY MEDTRONIC UNDER THE MEDICAL DEVICE REPORTING REGULATION DOES NOT CONSTITUTE AN ADMISSION THAT THE DEVICE(S) HAS MALFUNCTIONED OR THAT THERE IS ANY CAUSAL CONNECTION BETWEEN THE PERFORMANCE OF THE DEVICE AND ANY INJURY THAT MAY HAVE OCCURRED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02453 | 2008-05-08 | 20071119 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02454 | 2008-05-08 | 20071109 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES WERE FOUND; FULL LEAD WAS RETURNED FOR ANALYSIS. |  | N | Y | LAY USER/PATIENT
2649622-2008-02458 | 2008-05-08 | 20080214 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE LEAD WAS REMOVED AND REPLACED DUE TO IMPEDANCE INCREASE, HIGH/VARYING IMPEDANCE, OVERSENSING, AND NOISE ON THE LEAD. A LEAD FRACTURE WAS SUSPECTED. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT, AND THE PATIENT'S STATUS WAS REPORTED TO BE WELL FOLLOWING THE REPLACEMENT PROCEDURE. |  | N | Y | LAY USER/PATIENT
2649622-2008-02459 | 2008-05-08 | 20080206 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY, AND TESTED OUT OF SPECIFICATION CONSISTENT WITH THAT ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR WAS FRACTURED; DISTAL SEGMENT RETURNED FOR ANALYSIS. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02460 | 2008-05-08 | 20080118 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02449 | 2008-05-08 | 20080207 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS EXHIBITING HIGH IMPEDANCE AND OVERSENSING. IT WAS EXPLANTED AND REPLACED. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-02461 | 2008-05-08 | 20080122 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY, AND TESTED OUT OF SPECIFICATION CONSISTENT WITH THAT ADVISORY. EVALUATION SUMMARY: DISTAL CONDUCTOR WAS FRACTURED; FULL LEAD IN SEGMENTS WAS RETURNED FOR ANALYSIS. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02437 | 2008-05-08 |  | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02438 | 2008-05-08 | 20070529 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02439 | 2008-05-08 | 20070703 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02440 | 2008-05-08 | 20070321 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02441 | 2008-05-08 | 20070913 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02442 | 2008-05-08 | 20080123 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02443 | 2008-05-08 | 20080207 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES FOUND; PROXIMAL SEGMENT RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02444 | 2008-05-08 | 20080212 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02445 | 2008-05-08 | 20080207 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02447 | 2008-05-08 | 20080214 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02436 | 2008-05-08 | 20080208 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. EVALUATION SUMMARY: NO ANOMALIES WERE FOUND; FULL LEAD WAS RETURNED FOR ANALYSIS. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02448 | 2008-05-08 | 20071102 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02424 | 2008-05-08 | 20080224 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE PATIENT EXPIRED FOLLOWING INAPPROPRIATE SHOCKS THAT STARTED WHEN SHE WAS IN THE SHOWER. THE PT WAS SENT TO THE ER IN A DECOMPENSATED STATE. CPR WAS PERFORMED FOR 45 MIN TO 1 HOUR. THE DEVICE WAS INTERROGATED AND INAPPROPRIATE SHOCKS, OVERSENSING OF NOISE, LOSS OF CAPTURE, AND LEAD IMPEDANCE GREATER THAN 3000 OHMS WERE NOTED. POSSIBLE LEAD FRACTURE WAS SUSPECTED. LEAD IMPEDANCE TAKEN BY THE DEVICE ON THE PREVIOUS DAY WAS NORMAL. PACING INHIBITION WAS ALSO NOTED, DUE TO NOISE. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02427 | 2008-05-08 | 20080212 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02429 | 2008-05-08 | 20080226 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02430 | 2008-05-08 | 20071120 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES FOUND; FULL LEAD IN SEGMENTS RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02431 | 2008-05-08 | 20080111 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THERE WAS INCREASED IMPEDANCE. THE LEAD WAS REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. FURTHER ALLEGED THAT THE PATIENT "SUFFERED PHYSICAL AND OTHER INJURY", THAT THE PATIENT "EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY REQUIRING MEDICAL INTERVENTION INCLUDING BUT NOT LIMITED TO SURGERY, REMOVAL AND/OR REPLACEMENT OF THE DEFECTIVE [LEAD]", AND THAT THE PATIENT "SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH, ECONOMIC LOSSES AND OTHER DAMAGES. [PATIENT] HAS FURTHER SUFFERED PHYSICAL INJURIES AND VARIOUS PHYSICAL MANIFESTATIONS OF EMOTIONAL DISTRESS ASSOCIATED WITH ONE OR MORE OF THE FOLLOWING: THE IMPLANTATION, RECALL, FAILURE, REMOVAL/REPLACEMENT, AND/OR INABILITY TO HAVE THE DEFECTIVE [LEAD] REMOVED OR REPLACED." |  | Y | Y | LAY USER/PATIENT
2649622-2008-02432 | 2008-05-08 | 20070908 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE PATIENT REPORTED THAT HE RECEIVED SHOCKS, AND STATED THAT HE RECEIVED A "BAD LEAD". THE PACE/SENSE PORTION OF THE LEAD WAS CAPPED, AND A NEW PACE/SENSE LEAD WAS IMPLANTED. FURTHER INFORMATION WAS RECEIVED VIA A TECHNICAL SERVICES CALL THAT THERE WAS AN IMPEDANCE INCREASE AND FLUCTUATING IMPEDANCE, AND THAT THE HIGH VOLTAGE COIL INTEGRITY WAS SUSPECT. THE ENTIRE LEAD WAS REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02433 | 2008-05-08 | 20080220 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02434 | 2008-05-08 | 20071123 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE PATIENT REPORTED THAT HE RECEIVED 4 OR 5 SHOCKS AND WENT TO THE ER. IT WAS FURTHER REPORTED THAT THERE WAS OVERSENSING, HIGH IMPEDANCE GREATER THAN 4000 OHMS, AND A LEAD FRACTURE. THE LEAD WAS CAPPED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02423 | 2008-05-08 | 20080220 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; DISTAL SEGMENT RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02435 | 2008-05-08 | 20080218 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02411 | 2008-05-08 | 20080211 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02412 | 2008-05-08 | 20080211 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02414 | 2008-05-08 | 20080213 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: DISTAL CONDUCTOR FRACTURED; PARTIAL LEAD IN SEGMENTS RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02415 | 2008-05-08 | 20080212 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS EXPLANTED AND REPLACED DUE TO ELEVATED THRESHOLD AND DECREASED SENSING. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-02416 | 2008-05-08 | 20080213 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS EXPLANTED AND REPLACED DUE TO OVERSENSING AND HIGH IMPEDANCE. WHEN THE NEW LEAD WAS ATTACHED TO THE OLD DEVICE, NOISE WAS NOTED ON THE RV TIP TO RV RING DESPITE RECONNECTING THREE TIMES. THE HCP THEN EXPLANTED AND REPLACED THE DEVICE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-02417 | 2008-05-08 | 20080208 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02419 | 2008-05-08 | 20080212 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THERE WAS OVERSENSING AND INAPPROPRIATE THERAPIES. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02420 | 2008-05-08 | 20080215 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02421 | 2008-05-08 | 20080204 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02410 | 2008-05-08 | 20071220 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02422 | 2008-05-08 | 20080208 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02398 | 2008-05-08 |  | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02399 | 2008-05-08 | 20071025 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02400 | 2008-05-08 | 20060710 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-02401 | 2008-05-08 | 20051207 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-02402 | 2008-05-08 | 20061009 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02403 | 2008-05-08 | 20060210 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-02404 | 2008-05-08 | 20071219 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02405 | 2008-05-08 | 20080219 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT DURING SYSTEM TESTING, THE PATIENT DID NOT JUMP WITH A SHOCK AT 1 OR 2 JOULES, AND THAT A 5 JOULE SHOCK "RESULTED IN LESS THAN 20 OHMS ON HV LEAD IMPEDANCE". IT WAS ALSO REPORTED THAT THE LEAD WAS ALREADY IN QUESTION DUE TO SMALL R-WAVES. FURTHER TESTING, AFTER A NEW LEAD WAS IMPLANTED, RESULTED IN A CHARGE CIRCUIT TIMEOUT, AND EVENTUALLY, CHARGE CIRCUIT INACTIVE STATUS. WHEN VF WAS FINALLY INDUCED, THE DEVICE TOOK 22 SECONDS TO CHARGE, AND EXTERNAL RESCUE WAS NEEDED. THE DEVICE AND LEAD WERE EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02406 | 2008-05-08 | 20070726 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02407 | 2008-05-08 | 20070709 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02397 | 2008-05-08 | 20080129 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02386 | 2008-05-08 | 20070209 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02387 | 2008-05-08 | 20080201 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02388 | 2008-05-08 |  | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02389 | 2008-05-08 | 20080123 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THE LEAD WAS EXPLANTED AND REPLACED DUE TO HIGH IMPENDANCE, NOISE AND OVERSENSING. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-02391 | 2008-05-08 | 20071231 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02392 | 2008-05-08 | 20080201 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. EVALUATION SUMMARY: NO ANOMALIES WERE FOUND. |  | N | Y | LAY USER/PATIENT
2649622-2008-02394 | 2008-05-08 | 20080215 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED FOR ANALYSIS. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02395 | 2008-05-08 | 20080205 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE, DUE TO CONFIDENTIALITY CONCERNS. EVALUATION SUMMARY: NO ANOMALIES WERE FOUND; FULL LEAD WAS RETURNED FOR ANALYSIS. |  | N | Y | LAY USER/PATIENT
2649622-2008-02385 | 2008-05-08 | 20080214 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THAT THE LEAD WAS REPLACED DUE TO INCREASING AND FLUCTUATING IMPEDANCE. IT WAS NOTED AT EXPLANT THAT THE LEAD WAS DAMAGED. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. IT WAS REPORTED THAT THE LEAD WAS REPLACED DUE TO IMPEDANCE RAISE AND FLUCTUATION. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-02396 | 2008-05-08 | 20080111 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-02374 | 2008-05-08 | 20080201 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02376 | 2008-05-08 | 20080211 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02377 | 2008-05-08 | 20071211 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02378 | 2008-05-08 | 20080212 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THERE WAS HIGH IMPEDANCE OF GREATER THAN 3000 OHMS. A LAWSUIT FURTHER ALLEGED THAT THE PATIENT "SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH, AND ECONOMIC LOSSES AND OTHER DAMAGES". THE LEAD WAS EXPLANTED AND REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02379 | 2008-05-08 | 20080211 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: DEFIB CONDUCTOR FRACTURED; THERE IS ONLY ONE SET OF SETSCREW MARKS ON THE IS-1 CONNECTOR PIN AND IT IS TOO PROXIMAL WHICH INDICATES THE LEAD WAS NOT FULLY INSERTED INTO THE DEVICE; PARTIAL LEAD IN SEGMENTS RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02380 | 2008-05-08 | 20080213 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02382 | 2008-05-08 | 20080212 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02373 | 2008-05-08 | 20080213 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02366 | 2008-05-08 | 20080206 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; PARTIAL LEAD IN SEGMENTS RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02368 | 2008-05-08 | 20080213 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS EXPLANTED AND REPLACED DUE TO OVERSENSING, 1 INAPPROPRIATE SHOCKS AND AN APPARENT LEAD FRACTURE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02369 | 2008-05-08 | 20080213 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02370 | 2008-05-08 | 20080212 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02371 | 2008-05-08 | 20080208 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02372 | 2008-05-08 | 20080201 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02357 | 2008-05-08 | 20080215 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02363 | 2008-05-08 | 20080214 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02364 | 2008-05-08 | 20071024 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02365 | 2008-05-08 | 20080219 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE PATIENT RECEIVED AN INAPPROPRIATE SHOCK DUE TO NOISE. THE PATIENT ALERT SOUNDED, DUE TO UNSTABLE/HIGH IMPEDANCE. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02344 | 2008-05-08 | 20080208 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED PATIENT RECEIVED 32 INAPPROPRIATE THERAPIES IN TWO DAYS, AND IMPEDANCE INCREASED TO GREATER THAN 2000 OHMS. THE LEAD WAS REPLACED. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED ALLEGING THAT THE PATIENT 'EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY', REQUIRING REPLACEMENT OF THE 'DEFECTIVE' LEAD. IT ALSO ALLEGED THAT THE PATIENT HAS 'SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH, ECONOMIC LOSSES AND OTHER DAMAGES. [PATIENT] HAS FURTHER SUFFERED PHYSICAL INJURIES AND VARIOUS PHYSICAL MANIFESTATIONS OF EMOTIONAL DISTRESS'. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02347 | 2008-05-08 | 20071129 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02350 | 2008-05-08 | 20080222 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED |  | Y | Y | LAY USER/PATIENT
2649622-2008-02352 | 2008-05-08 | 20080304 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02353 | 2008-05-08 | 20080211 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02355 | 2008-05-08 | 20080212 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02333 | 2008-05-08 |  | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02336 | 2008-05-08 | 20080213 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02341 | 2008-05-08 | 20080211 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02331 | 2008-05-08 |  | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02343 | 2008-05-08 | 20080211 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | INAPPROPRIATE SHOCKS AND OVERSENSING WERE REPORTED. THE LEAD WAS REPLACED. THE PATIENT FURTHER REPORTED THAT THE DEVICE "SHOCKED" THE PATIENT 4 TIMES. THE PATIENT WAS TAKEN TO THE ER AND THEY COULDN'T FIGURE OUT WHY THE DEVICE "SHOCKED" THE PATIENT. THE PATIENT REPORTED RECEIVING MORE SHOCKS IN THE ER AND WAS FLOWN TO ANOTHER HOSPITAL AND THEY TOLD THE PATIENT THE LEAD "BROKE". NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02322 | 2008-05-08 | 20080208 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02324 | 2008-05-08 | 20080201 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02312 | 2008-05-08 | 20080208 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02313 | 2008-05-08 |  | No answer provided | Manufacturer report | SPRINT FIDELIS | ASKU | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE PATIENT RECEIVED MULTIPLE SHOCKS DUE TO NOISE AND OVERSENSING. THE PLAN IS TO REPLACE THE LEAD. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | 0
2649622-2008-02314 | 2008-05-08 | 20080208 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02316 | 2008-05-08 | 20080205 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE, DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02317 | 2008-05-08 | 20080130 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6930 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02300 | 2008-05-08 | 20080204 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02301 | 2008-05-08 | 20080214 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. OTHER : IT WAS REPORTED THE PATIENT RECEIVED 10 SHOCKS DUE TO OVERSENSING. THE PACING IMPEDANCE HAD JUMPED TO 1500 OHMS. THE LEAD WAS REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONALLY, THE PATIENT'S DAUGHTER SAID THE LEAD WAS REPLACED DUE TO A FRACTURE AND SAID "IT ALMOST KILLED" HER FATHER. IT WAS FURTHER REPORTED THAT THERE WAS NOISE. NO CONCLUSION CAN BE DRAWN. IMPEDANCE, HIGH, INTERFERENCE, LEAD(S), FRACTURE OF, OVERSENSING, SHOCK, INAPPROPRIATE. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02302 | 2008-05-08 | 20080222 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: PROXIMAL CONDUCTOR WAS FRACTURED; PARTIAL LEAD IN SEGMENTS WAS RETURNED FOR ANALYSIS. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02303 | 2008-05-08 | 20080215 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02306 | 2008-05-08 | 20080131 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: DEFIB CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED |  | N | Y | LAY USER/PATIENT
2649622-2008-02307 | 2008-05-08 | 20080210 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02308 | 2008-05-08 | 20080103 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02309 | 2008-05-08 | 20080206 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-02285 | 2008-05-08 | 20080213 | Injury | Manufacturer report | SPRINT FIDELIS | DEFIBRILLATOR, AUTOMATIC IMPLANTABLE CARDIOVERTER | MEDTRONIC PUERTO RICO, INC. | 6949 | 6949 |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02289 | 2008-05-08 | 20080208 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02290 | 2008-05-08 | 20080205 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT OVERSENSING LED TO 30 INAPPROPRIATE THERAPIES. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. FURTHER INFORMATION RECEIVED FROM AN ATTORNEY ALLEGES PATIENT SUFFERED PHYSICAL AND OTHER INJURY AS A RESULT OF THE RECALLED LEAD. PATIENT EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY REQUIRING MEDICAL INTERVENTION INCLUDING BUT NOT LIMITED TO SURGERY, REMOVAL AND/OR REPLACEMENT OF THE DEFECTIVE SPRINT FIDELIS LEAD. ATTORNEY ALSO ALLEGES THE PATIENT "HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH". |  | Y | Y | LAY USER/PATIENT
2649622-2008-02294 | 2008-05-08 | 20080211 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE PATIENT RECEIVED INAPPROPRIATE SHOCKS FROM LEAD FRACTURE. OVERSENSING, WHICH LED TO BOTH INAPPROPRIATE SHOCKS AND PACING INHIBITION, WAS ALSO NOTED. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02295 | 2008-05-08 | 20080207 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02283 | 2008-05-08 | 20080207 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02296 | 2008-05-08 | 20080208 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; PARTIAL LEAD IN SEGMENTS RETURNED AND ANALYZED. OTHER: IT WAS REPORTED THE PATIENT ALERT SOUNDED FOR IMPEDANCE OF 1032 OHMS. OVERSENSING OF NOISE WAS ALSO NOTED. THE LEAD WAS EXPLANTED AND REPLACED DUE TO FRACTURE. UPON REMOVAL, THE HELIX REQUIRED 50 TURNS TO RETRACT. THE PATIENT SUBSEQUENTLY REPORTED LEAD REPLACEMENT, DUE TO "FRAYED WIRE." NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ACTUAL DEVICE INVOLVED IN INCIDENT WAS EVALUATED. ELECTRICAL TESTS PERFORMED; MECHANICAL TESTS PERFORMED. VISUAL EXAMINATION: LEAD CONDUCTOR. COMPONENT/SUBASSEMBLY FAILURE. DEVICE WAS OUT OF SPECIFICATION IN A MANNER THAT RELATES TO EVENT. FRAYED, IMPEDANCE, HIGH, INTERFERENCE, LEAD(S), FRACTURE OF, OVERSENSING. |  | N | Y | LAY USER/PATIENT
2649622-2008-02274 | 2008-05-08 | 20080208 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02278 | 2008-05-08 | 20080208 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02280 | 2008-05-08 | 20080313 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02281 | 2008-05-08 | 20080207 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE PATIENT RECEIVED 3-4 INAPPROPRIATE SHOCKS IN ONE DAY DUE TO OVERSENSING. THE PACE SENSE PORTION OF THE LEAD WAS REPLACED AND THE HIGH VOLTAGE PORTION REMAINS IN USE. IT WAS REPORTED THAT THE LEAD RECENTLY HAD A COIL FRACTURE. HIGH AND LOW IMPEDANCE WAS NOTED. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02282 | 2008-05-08 | 20080204 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02263 | 2008-05-08 | 20080107 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02264 | 2008-05-08 | 20080206 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE LEAD IS PART OF THE SPRINT FIDELIS FIELD ADVISORY, AND TESTED OUT OF SPECIFICATION CONSISTENT WITH THAT ADVISORY. EVALUATION SUMMARY: DISTAL CONDUCTOR WAS FRACTURED; FULL LEAD IN SEGMENTS WAS RETURNED FOR ANALYSIS. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02266 | 2008-05-08 | 20080204 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02268 | 2008-05-08 | 20080124 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED BY THE PATIENT'S WIFE THAT THE PATIENT WAS SHOCKED 36 TIMES. SHE WAS TOLD THE LEAD WAS BROKEN. THE LEAD WAS REPLACED. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ATTORNEY LATER ALLEGES PATIENT "SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH, ECONOMIC LOSSES AND OTHER DAMAGES" AND "SUFFERED PHYSICAL INJURIES AND VARIOUS PHYSICAL MANIFESTATIONS OF EMOTIONAL DISTRESS ASSOCIATED WITH ONE OR MORE OF THE FOLLOWING: THE IMPLANTATION, RECALL, FAILURE, REMOVAL/REPLACEMENT, AND/OR INABILITY TO HAVE THE DEFECTIVE SPRINT FIDELIS LEAD REMOVED OR REPLACED." REVIEW OF MANUFACTURER'S DATABASE VERIFIED PATIENT'S DEATH AND THAT THE SPRINT FIDELIS LEAD WAS NOT IN USE AT THE TIME OF PATIENT'S DEATH. THERE IS NO ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH HAS BEEN RESEARCHED AND NOT RECEIVED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02269 | 2008-05-08 | 20080218 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | REVIEW OF DEVICE DISK DATA SHOWED THE IMPEDANCE HAD INCREASED, AND MEASURED AS HIGH AS 11,000 OHMS. BOTH HV COILS MEASURED OUT OF RANGE. THE LEAD WAS REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED IN A LAWSUIT WHICH FURTHER ALLEGED THAT THE PATIENT "SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, AND ECONOMIC LOSSES AND CONSEQUENTIAL DAMAGES". |  | Y | Y | LAY USER/PATIENT
2649622-2008-02262 | 2008-05-08 |  | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY AND TESTED OUT OF SPECIFICATION CONSISTENT WITH THE ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR WAS FRACTURED; DISTAL SEGMENT WAS RETURNED FOR ANALYSIS. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02272 | 2008-05-08 | 20080213 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES FOUND; FULL LEAD RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02273 | 2008-05-08 | 20080212 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | N | LAY USER/PATIENT
2649622-2008-02255 | 2008-05-08 | 20080207 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT RECEIVED UNEXPECTED SHOCKS. THE LEAD HAD AN APPARENT FRACTURE. THE LEAD WAS CAPPED AND REPLACED. ADDITIONAL INFORMATION RECEIVED INDICATED THAT THE LEAD SHOWED OVERSENSING. THERE WERE NO FURTHER REPORTED PATIENT COMPLICATIONS AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02256 | 2008-05-08 | 20080201 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THAT THE RV LEAD WAS EXPLANTED DUE TO INAPPROPRIATE SHOCKS AND HIGH IMPEDANCE. IT WAS NOTED THAT THE PATIENT ALERT WAS GENERATED FOR HIGH LEAD IMPEDANCE VALUES. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02259 | 2008-05-08 | 20080125 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR WAS FRACTURED; FULL LEAD WAS RETURNED FOR ANALYSIS. |  | N | Y | LAY USER/PATIENT
2649622-2008-02260 | 2008-05-08 | 20080131 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | IT WAS REPORTED THE PATIENT RECEIVED INAPPROPRIATE SHOCKS, DUE TO OVERSENSING OF NOISE. A LEAD FRACTURE WAS SUSPECTED. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT, AND THE PATIENT'S STATUS WAS GOOD. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02261 | 2008-05-08 | 20080131 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02240 | 2008-05-08 | 20080130 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. AS A DIRECT AND PROXIMATE RESULT OF THE DEFENDANTS' CONDUCT,(PT) HAVE SUFFERED PHYSICAL INJURIES AND VARIOUS PHYSICAL MANIFESTATIONS OF EMOTIONAL DISTRESS ASSOCIATED WITH ONE OR MORE OF THE FOLLOWING: THE IMPLANTATION, RECALL, FAILURE, REMOVAL/ REPLACEMENT, AND/OR INABILITY TO HAVE THE DEFECTIVE SPRINT FIDELIS LEAD REMOVED OR REPLACED." PATIENT LATER REQUESTED MEDICAL RECORDS OF THE LEAD REPLACEMENT AND STATED "IT JUST BLEW UP." |  | Y | Y | LAY USER/PATIENT
2649622-2008-02242 | 2008-05-08 | 20080206 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02243 | 2008-05-08 | 20080201 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02244 | 2008-05-08 | 20080205 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02245 | 2008-05-08 | 20080124 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02246 | 2008-05-08 | 20070720 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE LEAD WAS EXPLANTED AND REPLACED DUE TO "INAPPROPRIATE SHOCKS". NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02248 | 2008-05-08 | 20080207 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS EXPLANTED AND REPLACED DUE TO INCREASED IMPEDANCE (>2500 OHMS) AND OVERSENSING. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED REPORTING THAT THE LEAD FAILED, AND THE PATIENT RECEIVED SIX SHOCKS IN A FIVE MINUTE PERIOD. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02249 | 2008-05-08 | 20080208 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE RV LEAD WAS EXPLANTED DUE TO HIGH IMPEDANCE OF GREATER THAN 3000 OHMS AND OVERSENSING. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. AN ATTORNEY FURTHER ALLEGED THE PATIENT EXPERIENCED "INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY REQUIRING MEDICAL INTERVENTION INCLUDING BUT NOT LIMITED TO SURGERY, REMOVAL AND/OR REPLACEMENT OF THE DEFECTIVE" LEAD. ATTORNEY ALSO ALLEGED THE PATIENT HAD "SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES, AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH, ECONOMIC LOSSES AND OTHER DAMAGES". |  | Y | Y | LAY USER/PATIENT
2649622-2008-02250 | 2008-05-08 | 20080204 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02239 | 2008-05-08 | 20080207 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED BY THE PATIENT THAT HE WAS ADMITTED TO THE EMERGENCY DEPARTMENT AFTER BEING SHOCKED SEVERAL TIMES. THE PHYSICIAN FELT "THERE MUST BE A SHORT IN THERE." THE LEAD WAS REPLACED. ADDITIONAL INFORMATION WAS LATER RECEIVED REPORTING A POSSIBLE FRACTURE. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02229 | 2008-05-08 | 20080125 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02230 | 2008-05-08 | 20080201 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE LEAD WAS REPLACED, DUE TO APPARENT LEAD FRACTURE, AND INCREASE/FLUCTUATION IN RV (RIGHT VENTRICULAR) IMPEDANCE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-02231 | 2008-05-08 | 20080205 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02232 | 2008-05-08 | 20080208 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: DISTAL CONDUCTOR FRACTURED; FULL LEAD RETURNED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02235 | 2008-05-08 | 20080201 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE LEAD WAS EXPLANTED DUE TO OVERSENSING. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED FROM AN ATTORNEY ALLEGING THAT DUE TO THE LEAD, THE PATIENT 'HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR INCREASED RISK OF DEATH, LOSS OF COMPANIONSHIP AND SOCIETY, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH, ECONOMIC LOSSES AND OTHER DAMAGES'. NO SPECIFIC PATIENT INJURIES HAVE BEEN REPORTED. |  | Y | N | LAY USER/PATIENT
2649622-2008-02237 | 2008-05-08 | 20080205 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED LEAD WAS EXPLANTED AND REPLACED DUE TO UNDERSENSING, UNSTABLE THRESHOLDS, NO CAPTURE, AND DISLODGEMENT. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | N | LAY USER/PATIENT
2649622-2008-02238 | 2008-05-08 | 20080204 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE PATIENT RECEIVED INAPPROPRIATE THERAPIES, DUE TO OVERSENSING AND APPARENT FRACTURE. HIGH IMPEDANCE, GREATER THAN 3000 OHMS, WAS ALSO REPORTED. THE LEAD WAS FOUND TO BE "WADDED UP IN THE POCKET" AT REPLACEMENT. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02221 | 2008-05-08 | 20080129 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE LEAD HAD HIGH IMPEDANCE, OVERSENSING, AND NO CAPTURE. THERE WAS A VISABLE FRACTURE OF THE LEAD ON X-RAY DISTAL TO THE SUTURE SLEEVE. THE LEAD WAS EXPLANTED AND REPLACED. THERE WERE NO REPORTED PATIENT COMPLICATIONS AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-02222 | 2008-05-08 | 20080131 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02223 | 2008-05-08 | 20080104 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02224 | 2008-05-08 | 20080131 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02225 | 2008-05-08 | 20080207 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY, AND TESTED OUT OF SPECIFICATION CONSISTENT WITH THAT ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR WAS FRACTURED; FULL LEAD WAS RETURNED FOR ANALYSIS. |  | N | Y | LAY USER/PATIENT
2649622-2008-02218 | 2008-05-08 | 20080130 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: DISTAL CONDUCTOR FRACTURED; FULL LEAD ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02227 | 2008-05-08 | 20071229 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02213 | 2008-05-08 | 20080211 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02216 | 2008-05-08 | 20070910 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02217 | 2008-05-08 | 20080207 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. EVALUATION SUMMARY: PROXIMAL CONDUCTOR WAS FRACTURED; FULL LEAD WAS RETURNED FOR ANALYSIS. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02200 | 2008-05-08 | 20080213 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES FOUND; FULL LEAD IN SEGMENTS RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02204 | 2008-05-08 | 20080122 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02206 | 2008-05-08 | 20080125 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02207 | 2008-05-08 | 20080130 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02208 | 2008-05-08 | 20080128 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02189 | 2008-05-08 | 20080123 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02197 | 2008-05-08 | 20080121 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02187 | 2008-05-08 | 20070701 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02198 | 2008-05-08 | 20080129 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES FOUND; FULL LEAD IN SEGMENTS RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02199 | 2008-05-08 | 20080206 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02179 | 2008-05-08 | 20080211 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02181 | 2008-05-08 | 20080125 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02186 | 2008-05-08 | 20080129 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02170 | 2008-05-08 | 20071109 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02171 | 2008-05-08 | 20071107 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02173 | 2008-05-08 | 20071012 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02174 | 2008-05-08 | 20080121 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02169 | 2008-05-08 | 20071108 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02175 | 2008-05-08 | 20071119 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02157 | 2008-05-08 | 20071116 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. DISTAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | Y | N | LAY USER/PATIENT
2649622-2008-02159 | 2008-05-08 | 20080115 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THAT THERE WAS LOSS OF CAPTURE, INCREASED IMPEDANCE TO 2500 OHMS, AND OVERSENSING. THE RIGHT VENTRICULAR LEAD WAS EXPLANTED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED REPORTING THAT THE LEAD WAS FRACTURED. ADDITIONAL INFORMATION WAS RECEIVED REPORTING THAT THE PATIENT ALSO RECEIVED INAPPROPRIATE SHOCKS. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02162 | 2008-05-08 | 20070309 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02163 | 2008-05-08 | 20070308 | Injury | Manufacturer report | SPRINT FIDELIS | DEFIBRILLATOR, AUTOMATIC IMPLANTABLE CARDIOVERTER | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  | THIS DEVICE WAS REPORTED AS INCLUDED IN THE FIELD ACTION. BASED ON THE INFORMATION RECEIVED AND WITHOUT THE RETURN OF THE PRODUCT, IT CANNOT BE CONFIRMED THAT THIS DEVICE PERFORMED AS DESCRIBED IN THE FIELD ACTION. IT IS INCLUDED IN THE FIELD ACTION IN THE ABUNDANCE OF CAUTION. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02165 | 2008-05-08 | 20071003 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02166 | 2008-05-08 | 20080208 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02167 | 2008-05-08 | 20080115 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02156 | 2008-05-08 | 20071113 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY AND TESTED OUT OF SPECIFICATION CONSISTENT WITH THE ADVISORY. PROXIMAL CONDUCTOR WAS FRACTURED; PARTIAL LEAD IN SEGMENTS WAS RETURNED FOR ANALYSIS. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02147 | 2008-05-08 | 20080128 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE DEVICE IS PART OF THE ADVISORY FOR THIS MODEL. |  | N | Y | LAY USER/PATIENT
2649622-2008-02148 | 2008-05-08 | 20080212 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES FOUND; PROXIMAL SEGMENT RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02149 | 2008-05-08 | 20071113 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02150 | 2008-05-08 | 20071211 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02151 | 2008-05-08 | 20080128 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-02152 | 2008-05-08 | 20080204 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02153 | 2008-05-08 | 20080201 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT WAS TAKEN BY AMBULANCE TO THE HOSPITAL AFTER RECEIVING SHOCKS. THERE WAS AN APPARENT LEAD FRACTURE, AND THE DEVICE WAS DISABLED UNTIL THE LEAD COULD BE REPLACED. THE LEAD WAS EXPLANTED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED REPORTING THAT THERE WAS OVERSENSING AND NOISE. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02154 | 2008-05-08 | 20080123 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02146 | 2008-05-08 | 20080211 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02155 | 2008-05-08 | 20071112 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02136 | 2008-05-08 | 20080114 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6930 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02141 | 2008-05-08 | 20071228 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02142 | 2008-05-08 | 20080117 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02144 | 2008-05-08 | 20050621 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02145 | 2008-05-08 | 20080117 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02127 | 2008-05-08 | 20071219 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02129 | 2008-05-08 | 20080108 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE LEAD WAS CAPPED DUE TO APPARENT FRACTURE, NO PACING AND NOISE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. FURTHER INFORMATION FROM AN ATTORNEY ALLEGES THE PLAINTIFF SUFFERED PHYSICAL AND OTHER INJURY RELATED TO THE RECALLED LEAD INCLUDING INAPPROPRIATE SHOCK, FRACTURE, OR OTHER INJURY REQUIRING MEDICAL INTERVENTION. THE ATTORNEY ALSO ALLEGES SEVERE EMOTIONAL STRESS, AND MENTAL ANGUISH. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02132 | 2008-05-08 | 20071201 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | IT WAS REPORTED THAT THE RV LEAD WAS CAPPED DUE TO INAPPROPRIATE THERAPIES. A SIGNIFICANT INCREASE IN RV IMPEDANCE WAS ALSO REPORTED. NO FURTHER COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02133 | 2008-05-08 | 20071115 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02120 | 2008-05-08 | 20071219 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02121 | 2008-05-08 | 20070605 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THE PATIENT'S MOTHER SAID HER SON COLLAPSED AT HOME IN 2007 (SEVEN MONTHS AGO) AND SAID THE DEATH CERTIFICATE STATED CARDIAC ARREST. SHE QUESTIONED WHY THE DEVICE DID NOT SHOCK HIM. THERE IS NO ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT THE DEATH WAS DEVICE RELATED. |  | Y | N | LAY USER/PATIENT
2649622-2008-02122 | 2008-05-08 | 20071107 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02123 | 2008-05-08 | 20071231 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02125 | 2008-05-08 | 20071217 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-02126 | 2008-05-08 | 20071217 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02114 | 2008-05-08 | 20080111 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02115 | 2008-05-08 | 20080111 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THAT THE LEAD WAS REPLACED DUE TO UNSTABLE/HIGH IMPEDANCE AND NOISE ON THE EGM. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED IN A LAWSUIT REPORTING THAT THE PATIENT EXPERIENCED EXCESSIVE SHOCKING DUE TO THE LEAD. |  | N | Y | LAY USER/PATIENT
2649622-2008-02117 | 2008-05-08 | 20071206 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02108 | 2008-05-08 | 20071211 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02109 | 2008-05-08 |  | No answer provided | Manufacturer report | SPRINT FIDELIS | ASKU | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THERE WAS SENSING DIFFICULTY AND A LEAD FRACTURE. THE LEAD WAS EXPLANTED AND REPLACED. A LAWSUIT FURTHER ALLEGED THAT THE PATIENT EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING AND LEAD FRACTURE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | 0
2649622-2008-02110 | 2008-05-08 | 20080228 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES FOUND; THE ANALYST NOTED THAT THERE WERE TWO SETS OF SETSCREW MARKS ON THE IS-1 PIN AND BOTH SETS ARE TOO PROXIMAL, THUS LEAD WAS NOT FULLY INSERTED INTO THE DEVICE; FULL LEAD IN SEGMENTS RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02100 | 2008-05-08 | 20071127 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02101 | 2008-05-08 | 20071120 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS PARTIALLY REMOVED DUE TO HIGH IMPEDANCES AND AN APPARENT FRACTURE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-02095 | 2008-05-08 | 20071108 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-02097 | 2008-05-08 | 20071204 | Injury | Manufacturer report | SPRINT FIDELIS | DEFIBRILLATOR, AUTOMATIC IMPLANTABLE CARDIOVERTER | MEDTRONIC PUERTO RICO, INC. | 694858 | 694858 |  | IT WAS REPORTED BY THE PATIENT'S ATTORNEY THAT THE PATIENT "HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, AND ECONOMIC LOSSES." THE LEAD WAS EXPLANTED AND REPLACED PER PHYSICIAN DECISION. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02093 | 2008-05-08 | 20071207 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02092 | 2008-05-08 | 20080204 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE PATIENT ALERT SOUNDED FOR RV LEAD IMPEDANCE GREATER THAN 1000 OHMS. THE IMPEDANCE WAS NOTED TO FLUCTUATE BETWEEN 425 & 1200 OHMS. IT WAS LATER REPORTED THE PACING IMPEDANCE SPIKED TO 1800 OHMS, THEN BACK TO BASELINE OF 400-600 OHMS. ABOUT TWO MONTHS LATER, IT AGAIN SPIKED TO 1726 OHMS. THE PATIENT RECEIVED 4 SHOCKS, BUT THEY WERE DUE TO ARRHYTHMIAS. ADDITIONAL INFORMATION WAS SUBSEQUENTLY RECEIVED REPORTING HIGH IMPEDANCE & SOME OVERSENSING, BUT NO INAPPROPRIATE SHOCKS. THE PHYSICIAN TURNED OFF DETECTION & THERAPIES FOR VF, FAST VT, & SVT, IN 2008, & KEPT THE VT ZONE ACTIVE, BUT THERAPIES OFF, FOR MONITORING PURPOSES. THE LEAD WILL BE FURTHER EVALUATED IF THE PATIENT DEVELOPES SYMPTOMS. IT WAS FURTHER REPORTED THE LEAD WAS EXTRACTED, DUE TO FLUCTUATING IMPEDANCES. A CURRENT FVT EPISODE CONVINCED THE PATIENT & MD TO HAVE EXTRACTION DONE, SO THE DEVICE CAN BE TURNED BACK ON. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-02085 | 2008-05-08 | 20080205 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. EVALUATION SUMMARY: NO ANOMALIES FOUND; FULL LEAD RETURNED AND ANALYZED. IT WAS REPORTED THE LEAD WAS OVERSENSING. THE LEAD WAS EXPLANTED. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ACTUAL DEVICE INVOLVED IN INCIDENT WAS EVALUATED ELECTRICAL TESTS PERFORMED, MECHANICAL TESTS PERFORMED, VISUAL EXAMINATION DEVICE PERFORMED ACCORDING TO SPECIFICATIONS DEVICE EVALUATED AND ALLEGED FAILURE COULD NOT BE DUPLICATED OVERSENSING. |  | N | Y | LAY USER/PATIENT
2649622-2008-02090 | 2008-05-08 |  | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02083 | 2008-05-08 | 20080220 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  | IT WAS REPORTED THAT THE VENTRICULAR LEAD WAS ALMOST PULLED OUT OF THE DEVICE, DUE TO TWIDDLER'S SYNDROME. THE LEAD WAS CUT, CAPPED, AND REPLACED BECAUSE OF SHORT V-V INTERVALS, INDICATING OVERSENSING, AND NOISE. ADDITIONAL INFORMATION WAS SUBSEQUENTLY RECEIVED FROM AN ATTORNEY ALLEGING THE "PLAINTIFFS HAVE SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH, ECONOMIC LOSSES AND OTHER DAMAGES" ASSOCIATED WITH ONE OR MORE OF THE FOLLOWING: "THE IMPLANTATION, RECALL, FAILURE, REMOVAL/REPLACEMENT, AND/OR INABILITY TO HAVE THE DEFECTIVE [LEAD] REMOVED OR REPLACED." THE PATIENT REPORTED THEY MAY HAVE A SURGERY SCHEDULED TO "FIX DEVICE." THEY STATED THAT "LEAD CAN SHOCK STOMACH AND DEVICE FLOPS IN CHEST." NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02075 | 2008-05-08 | 20071030 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02076 | 2008-05-08 | 20080306 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02077 | 2008-05-08 | 20080219 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE RV LEAD WAS EXPLANTED DUE TO OVERSENSING NOISE AND ONE INAPPROPRIATE SHOCK DELIVERED. AN IMPEDANCE INCREASE FROM 560 TO 1200 OHMS AND DIAPHRAGMATIC STIMULATION WERE ALSO REPORTED. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02078 | 2008-05-08 | 20070930 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-02792 | 2008-05-08 | 20080311 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. PROXIMAL CONDUCTOR FRACTURED; FULL LEAD IN SEGMENTS ANALYZED |  | Y | Y | LAY USER/PATIENT
2649622-2008-02794 | 2008-05-08 | 20070119 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-02795 | 2008-05-08 | 20080312 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS EXPLANTED AND REPLACED DUE TO OVERSENSING AND HIGH IMPEDANCE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-02797 | 2008-05-08 | 20080310 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02799 | 2008-05-08 | 20080310 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02800 | 2008-05-08 | 20080313 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02801 | 2008-05-08 | 20080313 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02802 | 2008-05-08 | 20070831 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02791 | 2008-05-08 | 20070416 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02778 | 2008-05-08 | 20071214 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02779 | 2008-05-08 | 20080228 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02780 | 2008-05-08 | 20080219 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THE LEAD WAS REMOVED AND REPLACED DUE TO OVERSENSING, PATIENT ALERT, INAPPROPRIATE SHOCKS AND APPARENT LEAD FRACTURE. THE PATIENT'S CONDITION WAS REPORTED AS 'GOOD'. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02781 | 2008-05-08 | 20080306 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE, DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02782 | 2008-05-08 | 20080311 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02783 | 2008-05-08 | 20080306 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02784 | 2008-05-08 | 20080313 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-02785 | 2008-05-08 | 20080314 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: DISTAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02786 | 2008-05-08 | 20070310 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS EXPLANTED AND REPLACED DUE TO SENSING DIFFICULTY AND NOISE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION RECEIVED IS THAT THE LAWSUIT ALLEGED THAT THE PATIENT "HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, AND ECONOMIC AND CONSEQUENTIAL DAMAGES." |  | N | Y | LAY USER/PATIENT
2649622-2008-02787 | 2008-05-08 | 20080313 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02789 | 2008-05-08 | 20080220 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | N | LAY USER/PATIENT
2649622-2008-02790 | 2008-05-08 | 20090304 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. DETAILED ANALYSIS OF THE DEVICE WAS NOT PERFORMED AT THIS TIME DUE TO PENDING LITIGATION. THEREFORE, WE ARE UNABLE TO FULLY EVALUATE THE REPORTED EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02777 | 2008-05-08 |  | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02764 | 2008-05-08 | 20071106 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02765 | 2008-05-08 | 20080312 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02766 | 2008-05-08 | 20080310 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES FOUND; FULL LEAD RETURNED AND ANALYZED |  | Y | Y | LAY USER/PATIENT
2649622-2008-02768 | 2008-05-08 | 20080310 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES FOUND; FULL LEAD RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02769 | 2008-05-08 | 20080129 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02771 | 2008-05-08 |  | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: (B) (4) DETAILED ANALYSIS OF THE LEAD WAS NOT PERFORMED AT THIS TIME DUE TO PENDING LITIGATION. THEREFORE, WE ARE UNABLE TO FULLY EVALUATE THE REPORTED EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02772 | 2008-05-08 | 20080303 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-02763 | 2008-05-08 | 20080311 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: NO ANOMALIES FOUND; FULL LEAD RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02752 | 2008-05-08 | 20071112 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02753 | 2008-05-08 | 20071130 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02754 | 2008-05-08 | 20071019 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02755 | 2008-05-08 | 20071207 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS CAPPED AND REPLACED DUE TO FRACTURE. IT WAS FURTHER REPORTED THAT A IMPLANTABLE CARDIOVERTER DEFIBRILLATOR AND A VENTRICULAR LEAD WERE IMPLANTED IN THE PATIENT. THE PATIENT REQUIRED REPLACEMENT SURGERY OF THE LEAD. THE PATIENT WENT TO THE HOSPITAL AFTER SUFFERING MULTIPLE INAPPROPRIATE FIRINGS AND WAS EMOTIONALLY DISTRAUGHT. THE PATIENT HAS SUSTAINED PAIN AND SUFFERING, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH, ECONOMIC LOSSES AND OTHER DAMAGES. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02756 | 2008-05-08 | 20071219 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02757 | 2008-05-08 | 20080305 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE PATIENT RECEIVED MULTIPLE INAPPROPRIATE SHOCKS DUE TO OVERSENSING. THE LEAD WAS REPLACED. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION RECEIVED REPORTED THE LEAD WAS FRACTURED WITH NOISE AND SENSING DIFFICULTY AND WAS CAPPED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02758 | 2008-05-08 | 20080304 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THERE WERE 5 INAPPROPRIATE THERAPIES DELIVERED. IMPEDANCE OF GREATER THAN 1000 OHMS WAS NOTED. THE LEAD AND DEVICE WERE REPLACED. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02759 | 2008-05-08 | 20080313 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-02762 | 2008-05-08 | 20080313 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD IMPEDANCE TREND HAD BEEN INCREASING FROM 790 TO 1150 OHMS AND THAT THERE WAS OVERSENSING. THE LEAD WAS EXPLANTED AND REPLACED. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-02740 | 2008-05-08 | 20080110 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-02741 | 2008-05-08 | 20080310 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THERE WAS SENSING DIFFICULTY, OVERSENSING, INTERMITTENT HIGH IMPEDANCE VALUES, AND THAT THE PATIENT RECEIVED INAPPROPRIATE THERAPY. IT WAS FURTHER REPORTED THAT THERE WAS A SUSPECTED LEAD FRACTURE, AND THAT THE PATIENT HAD EXPERIENCED A FALL. BOTH THE DEVICE AND LEAD WERE EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. THE ATRIAL LEAD WAS RETURNED TO THE MANUFACTURER AFTER BEING EXPLANTED FOR HIGH IMPEDANCE AND OVERSENSING. THE LEAD SUBSEQUENTLY TESTED OUT OF SPECIFICATION DURING MANUFACTURER'S ANALYSIS. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02742 | 2008-05-08 | 20080124 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02743 | 2008-05-08 | 20080107 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES FOUND; FULL LEAD RETURNED AND ANALZYED. THERE WAS ONLY ONE SET OF SETSCREW MARKS ON THE IS-1 PIN AND THEY WERE TOO PROXIMAL, INDICATING THAT THE LEAD WAS NOT FULLY INSERTED INTO THE DEVICE. |  | N | Y | LAY USER/PATIENT
2649622-2008-02744 | 2008-05-08 | 20080311 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02746 | 2008-05-08 | 20080102 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02747 | 2008-05-08 | 20080317 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THERE WAS OVERSENSING AND UNSTABLE LEAD IMPEDANCE. IT WAS FURTHER REPORTED THAT THERE WAS HIGH IMPEDANCE GREATER THAN 3000 OHMS. THE PATIENT SAID THE LEAD "MALFUNCTIONED". IT WAS PARTIALLY REMOVED AND REPLACED. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-02748 | 2008-05-08 | 20071015 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2182208-2008-00130 | 2008-05-08 | 20071102 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC, INC. | 6948 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02749 | 2008-05-08 | 20080107 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02739 | 2008-05-08 | 20060112 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  | A LAWSUITS ALLEGES, "DAYS AFTER THE DEVICE IMPLANTATION, THE PATIENT BEGAN EXPERIENCING HEAVY CHEST PAIN AND SUFFERING SEVERE BREATHING COMPLICATIONS. THE PATIENT WAS RUSHED TO THE EMERGENCY ROOM WHERE IT WAS FOUND THE PATIENT'S LUNGS WERE FILLING WITH BLOOD-THE RESULT OF A [MYOCARDIAL PERFORATION DUE TO THE IMPLANTED LEAD]. CONSEQUENTLY, THE PATIENT UNDERWENT ANOTHER SURGERY TO REPAIR HER HEART AND DRAIN HER LUNGS ON (B)(4) 2006. FURTHERMORE, KNOWLEDGE OF THE LEAD RECALL HAS CAUSED THE PATIENT A CONSIDERABLE AMOUNT OF ANXIETY." IT WAS LATER REPORTED BY THE PATIENT THAT THE DEVICE "MALFUNCTIONED" AND THE PATIENT WAS SHOCKED MULTIPLE TIMES. FOLLOW-UP WAS CONDUCTED TO OBTAIN INFORMATION ON THE PATIENT AND WHETHER THE EVENT WAS LEAD RELATED, BUT THE ATTEMPTS WERE UNSUCCESSFUL. RECORDS INDICATE THE LEAD REMAINS IN USE. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THE EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02726 | 2008-05-08 | 20070806 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02728 | 2008-05-08 |  | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-02729 | 2008-05-08 | 20080310 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-02732 | 2008-05-08 | 20071106 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02733 | 2008-05-08 | 20071123 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02734 | 2008-05-08 | 20080317 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02735 | 2008-05-08 | 20080325 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY, AND TESTED OUT OF SPECIFICATION CONSISTENT WITH THAT ADVISORY. EVALUATION SUMMARY: DISTAL CONDUCTOR WAS FRACTURED: FULL LEAD IN SEGMENTS WAS RETURNED FOR ANALYSIS. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02736 | 2008-05-08 | 20060222 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02737 | 2008-05-08 | 20080317 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02713 | 2008-05-08 | 20071015 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02714 | 2008-05-08 | 20071205 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-02717 | 2008-05-08 | 20080303 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED. FULL LEAD IN SEGMENTS RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02719 | 2008-05-08 | 20080307 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02720 | 2008-05-08 | 20080303 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE DEVICE AND THE RV LEAD WERE REPLACED DUE TO THE PATIENT HAVING RECEIVED 19 SHOCKS DUE TO NOISE. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION RECEIVED REPORTED THE DEVICE HAD SENSING DIFFICULTY AND NO PACING OUTPUT. THE LEAD WAS CAPPED DUE TO HIGH IMPEDANCE OF >3000 OHMS, OVERSENSING, AND NO CAPTURE. THE PATIENT AND PHYSICIAN ELECTED TO REPLACE THE SYSTEM WITH A BRADY SYSTEM AS THE PATIENT HAD BEEN TRAUMATIZED BY THE INAPPROPRIATE SHOCKS, AND HER EJECTION FRACTION HAD IMPROVED BY APPROXIMATELY 50%. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02721 | 2008-05-08 | 20080303 | Injury | Manufacturer report | SPRINT FIDELIS | DEFIBRILLATOR, AUTOMATIC IMPLANTABLE CARDIOVERTER | MPRI | 6949 | 6949 |  | IT WAS REPORTED THE PATIENT RECEIVED 8 SHOCKS DUE TO NOISE. THE LEAD WAS REPLACED. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02722 | 2008-05-08 | 20080304 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02723 | 2008-05-08 |  | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02724 | 2008-05-08 | 20080313 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02701 | 2008-05-08 | 20080303 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE, DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: DISTAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02702 | 2008-05-08 | 20080203 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02703 | 2008-05-08 | 20080229 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02704 | 2008-05-08 | 20080305 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02705 | 2008-05-08 | 20080304 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02706 | 2008-05-08 | 20080305 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02707 | 2008-05-08 | 20080305 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02708 | 2008-05-08 | 20080304 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. NO ANOMALIES FOUND; FULL LEAD IN SEGMENTS ANALYZED |  | N | Y | LAY USER/PATIENT
2649622-2008-02709 | 2008-05-08 | 20080305 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02710 | 2008-05-08 | 20071108 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02711 | 2008-05-08 | 20080303 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-02712 | 2008-05-08 | 20080303 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02700 | 2008-05-08 | 20080303 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT WAS TAKEN TO THE HOSPITAL BY AMBULANCE AND HAD RECEIVED 36 SHOCKS IN ONE HOUR, DUE TO OVERSENSING. IT WAS ALSO REPORTED THAT THERE WAS HIGH AND FLUCTUATING IMPEDANCE. THE LEAD WAS EXPLANTED AND REPLACED. IT WAS LATER REPORTED THAT THE DEVICE WAS DETECTING NOISE, DUE TO A SUSPECTED LEAD FRACTURE, SO INAPPROPRIATE SHOCKS WERE DELIVERED. THE PATIENT REPORTEDLY HAD AN INFECTION WITH THE NEW SYSTEM FOLLOWING THE REPLACEMENT PROCEDURE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT, AND THE PATIENT'S STATUS WAS REPORTED AS "FINE". |  | Y | Y | LAY USER/PATIENT
2649622-2008-02690 | 2008-05-08 | 20080304 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02691 | 2008-05-08 | 20080306 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02693 | 2008-05-08 |  | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02694 | 2008-05-08 | 20071030 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02695 | 2008-05-08 | 20080221 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | IT WAS REPORTED THE PATIENT RECEIVED INAPPROPRIATE SHOCKS DUE TO OVERSENSING OF NOISE. THE PATIENT ALARM WAS SOUNDING. A LEAD FRACTURE WAS SUSPECTED. THE LEAD WAS CAPPED AND REPLACED. IT WAS LATER REPORTED THE PATIENT EXPERIENCED BLEEDING AT HOME IN THE EVENING POST RV LEAD IMPLANT. PRESSURE BANDAGE WAS APPLIED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02696 | 2008-05-08 | 20080222 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02697 | 2008-05-08 | 20080228 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02698 | 2008-05-08 | 20080229 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02689 | 2008-05-08 | 20080307 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. ANALYSIS SUMMARY - NO ANOMALIES FOUND. PROXIMAL SEGMENT RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02678 | 2008-05-08 | 20080312 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE PATIENT REPORTED THAT HE RECEIVED 30 SHOCKS. THE PACE/SENSE PORTION OF THE LEAD WAS CAPPED, AND A NEW PACE/SENSE LEAD WAS IMPLANTED. IT WAS FURTHER REPORTED THAT THERE WAS A COIL FRACTURE AND HIGH SVC AND HVB COIL IMPEDANCES AND OVERSENSING. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | N | LAY USER/PATIENT
2649622-2008-02679 | 2008-05-08 | 20080318 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | OTHER: IT WAS REPORTED BY THE PATIENT THAT HE HAD A "MICROFRACTURE" AND IT WILL BE "REPLACED TODAY." ADDITIONAL INFORMATION RECEIVED FROM THE PHYSICIAN WAS THAT THE LEAD SHOWED OVERSENSING, HIGH IMPEDANCE AND AN APPARENT FRACTURE. THE LEAD WAS CAPPED AND AN NEW LEAD IMPLANTED. THERE WERE NO REPORTED PATIENT COMPLICATIONS AS A RESULT OF THIS EVENT. NO CONCLUSION CAN BE DRAWN. IMPEDANCE, HIGH, LEAD(S), FRACTURE OF, OVERSENSING. |  | N | Y | LAY USER/PATIENT
2649622-2008-02680 | 2008-05-08 | 20080228 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE PATIENT RECEIVED 11 INAPPROPRIATE SHOCKS. THE PACE/SENSE PORTION OF THE HIGH VOLTAGE LEAD WAS REPLACED WITH A PACE/SENSE LEAD. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. THE LEAD WAS RETURNED 11 MONTHS LATER AFTER THE PATIENT DIED. THERE IS NO ALLEGATION THE DEATH WAS DEVICE RELATED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02681 | 2008-05-08 | 20080306 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02683 | 2008-05-08 | 20080306 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02684 | 2008-05-08 | 20080229 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02685 | 2008-05-08 | 20080214 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02687 | 2008-05-08 | 20080229 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02688 | 2008-05-08 | 20080229 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02677 | 2008-05-08 | 20080327 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT HAD A "FAILED LEAD." ADDITIONAL INFORMATION RECEIVED WAS THAT THE LEAD SHOWED OVERSENSING AND UNDERSENSING. THE LEAD WAS REPLACED. THERE WERE NO REPORTED PATIENT COMPLICATIONS AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-02668 | 2008-05-08 | 20080123 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02669 | 2008-05-08 | 20080310 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02670 | 2008-05-08 | 20080107 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-02672 | 2008-05-08 | 20080311 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (NON-CRT) | MEDTRONIC PUERTO RICO VILLALBA | 6949 | 6949 |  | IT WAS REPORTED THAT THE PATIENT HEARD THE PATIENT ALERT. THE PATIENT WENT TO THE HOSPITAL, WAS REPORTED TO HAVE BEEN TOLD HE HAD "1600 FATAL CARDIAC ARRHYTHMIAS." IT WAS ALSO REPORTED THAT THE PATIENT WAS NEVER SHOCKED. THE LEAD WAS REPLACED. ALLEGED THAT THE [PATIENT] "WAS IMPLANTED WITH A [LEAD WIRE SYSTEM], AND SUFFERED PHYSICAL AND OTHER INJURY AS A RESULT OF THE RECALLED LEAD." ALSO ALLEGED WAS THAT THE [PATIENT] EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY." FURTHER ALLEGED WAS THAT THE PATIENT" HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH, ECONOMIC LOSSES AND OTHER DAMAGES." THE [PATIENT] HAS FURTHER "SUFFERED PHYSICAL INJURIES AND VARIOUS PHYSICAL MANIFESTATIONS OF EMOTIONAL DISTRESS ASSOCIATED WITH OF THE FOLLOWING: THE IMPLANTATION, RECALL, FAILURE, REMOVAL/REPLACEMENT, AND/OR INABILITY TO HAVE TO HAVE THE DEFECTIVE [LEAD] REMOVED OR REPLACED." |  | Y | Y | LAY USER/PATIENT
2649622-2008-02673 | 2008-05-08 | 20080215 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED BY THE PATIENT THE LEAD WAS FRACTURED. THE LEAD WAS REPLACED. ADDITIONAL INFORMATION WAS SUBSEQUENTLY RECEIVED FROM AN ATTORNEY REPORTING THE PATIENT HAD EXPERIENCED "INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY REQUIRING MEDICAL INTERVENTION INCLUDING BUT NOT LIMITED TO SURGERY, REMOVAL AND/OR REPLACEMENT" OF THE LEAD. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02674 | 2008-05-08 | 20080103 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02675 | 2008-05-08 | 20080318 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02676 | 2008-05-08 | 20080304 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02667 | 2008-05-08 | 20071127 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE ATTORNEY ALLEGES THAT DUE TO THE 'DESIGN, MANUFACTURE, ASSEMBLY, MARKETING AND SALES' OF THE LEAD, THE PATIENT HAS 'SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, AND ECONOMIC LOSSES AND CONSEQUENTIAL DAMAGES'. A PACE/SENSE LEAD WAS PLACED IN THE RIGHT VENTRICLE, AND THE HIGH VOLTAGE PORTION OF THE LEAD IS STILL IN USE. NO SPECIFIC PATIENT COMPLICATIONS WERE REPORTED. ADDITIONAL INFORMATION WAS RECEIVED FROM THE MEDTRONIC SALES REP REPORTING THAT, DUE TO THE PATIENT'S CONDITION, THE DEVICE WAS UPGRADED FROM A SINGLE CHAMBER TO A DUAL CHAMBER ICD. THE PHYSICIAN CHOSE AT THAT TIME TO PLACE A PACE/SENSE LEAD PROPHYLACTICALLY. INTERROGATION REVEALED THAT LEAD MEASUREMENTS WERE STABLE. THE CASE WENT NORMALLY, AND THERE WERE NO PATIENT COMPLICATIONS. NO CONCLUSION CAN BE DRAWN DEVICE REMAINS ACTIVATED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02658 | 2008-05-08 | 20080201 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02659 | 2008-05-08 | 20080304 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: DISTAL CONDUCTOR FRACTURED; PARTIAL LEAD IN SEGMENTS RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02660 | 2008-05-08 | 20080307 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; PARTIAL LEAD IN SEGMENTS RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02661 | 2008-05-08 | 20080304 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | IT WAS REPORTED THAT THE PATIENT RECEIVED INAPPROPRIATE SHOCKS, DUE TO OVERSENSING OF NOISE. AN APPARENT LEAD FRACTURE WAS ALSO REPORTED. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02662 | 2008-05-08 | 20080304 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02663 | 2008-05-08 | 20080204 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02664 | 2008-05-08 | 20080303 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02665 | 2008-05-08 | 20080306 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: DISTAL CONDUCTOR FRACTURED; DISTAL SEGMENT RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02649 | 2008-05-08 | 20080305 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE RV LEAD WAS EXPLANTED AND REPLACED DUE TO VARYING AND HIGH IMPEDANCES, 768 TO 1520 OHMS. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-02652 | 2008-05-08 | 20080228 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02653 | 2008-05-08 | 20080212 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02654 | 2008-05-08 | 20080220 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | N | LAY USER/PATIENT
2649622-2008-02655 | 2008-05-08 | 20080226 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THAT THE PATIENT ALERT SOUNDED FOR HIGH IMPEDANCE. THE LEAD WAS EXPLANTED AND REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. THE PATIENT'S STATUS WAS REPORTED AS "OK". |  | N | Y | LAY USER/PATIENT
2649622-2008-02656 | 2008-05-08 | 20080228 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02657 | 2008-05-08 | 20080228 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6930 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02647 | 2008-05-08 | 20080221 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02639 | 2008-05-08 | 20070914 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02640 | 2008-05-08 | 20080301 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES FOUND; PROXIMAL SEGMENT RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-02641 | 2008-05-08 | 20080313 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02642 | 2008-05-08 | 20080317 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02643 | 2008-05-08 | 20080228 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02644 | 2008-05-08 | 20080214 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02645 | 2008-05-08 | 20070808 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THIS SUPPLEMENTAL IS BASED SOLELY ON THE RECEIPT OF A 3500A REPORT. ADDITIONAL INFORMATION RECEIVED ON 3500A REPORT INDICATES THE PATIENT HAD A CENTRAL VENOUS PERFORATION LEADING TO INTRAOPERATIVE BLEEDING WHICH REQUIRED A MEDIAL STERNOTOMY AND CARDIOPULMONARY BYPASS. IT WAS REPORTED THAT DURING A LASER LEAD EXTRACTION, THE PATIENT DEVELOPED CARDIAC TAMPONADE AND DIED IN THE CATHETERIZATION LAB. IT WAS FURTHER REPORTED THAT THERE WAS A POSSIBLE HIGH VOLTAGE COIL FRACTURE AND IMPEDANCE GREATER THAN 200 OHMS. THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE DEVICE IS PART OF THE ADVISORY FOR THIS MODEL. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02646 | 2008-05-08 |  | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02625 | 2008-05-08 | 20080227 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE PATIENT RECEIVED 14 INAPPROPRIATE SHOCKS, WITH OVERSENSING INDICATED. THE LEAD WAS REPLACED. AT EXPLANT, IT WAS FOUND THERE WAS A LARGE U-SHAPE OF THE LEAD IN THE POCKET, AND IT WAS SPECULATED THIS COULD HAVE CAUSED A FRACTURE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONALLY, A LAWSUIT ALLEGES THE PATIENT "SUFFERED PHYSICAL AND OTHER INJURY" DUE TO THE LEAD. IT ALSO STATES THE PATIENT "EXPERIENCED INAPPROPRIATE AND OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY REQUIRING MEDICAL INTERVENTION" DUE TO THE "DEFECTIVE" LEAD. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02629 | 2008-05-08 | 20080109 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02633 | 2008-05-08 | 20080303 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY NO ANOMALIES FOUND; PARTIAL LEAD RETURNED IN SEGMENTS IT WAS REPORTED THAT THERE WAS A LEAD FRACTURE, HIGH IMPEDANCE, OVERSENSING, AND NOISE. IT WAS ALSO REPORTED THAT THE PATIENT RECEIVED SEVERAL INAPPROPRIATE SHOCKS. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ACTUAL DEVICE INVOLVED IN INCIDENT WAS EVALUATED ELECTRICAL TESTS PERFORMED MECHANICAL TESTS PERFORMED, VISUAL EXAMINATION DEVICE PERFORMED ACCORDING TO SPECIFICATIONS NO CONCLUSION CAN BE DRAWN IMPEDANCE, HIGH INTERFERENCE LEAD(S), FRACTURE OF OVERSENSING SHOCK, INAPPROPRIATE. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02623 | 2008-05-08 | 20080228 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE PATIENT EXPERIENCED 25 INAPPROPRIATE SHOCKS. A LEAD IMPEDANCE THAT ROSE TO 2600 OHMS AND OVERSENSING WERE ALSO REPORTED. THE PATIENT FURTHER REPORTED THAT HER LAST "MACHINE ALMOST KILLED ME" AND THAT THE ISSUES CAUSED HER PAIN AND FEAR. THE LEAD WAS REPLACED. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02613 | 2008-05-08 | 20080304 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT RECEIVED MORE THAN 20 INAPPROPRIATE SHOCKS. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02614 | 2008-05-08 | 20080228 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. EVALUATION SUMMARY: NO ANOMALIES WERE FOUND; FULL LEAD WAS RETURNED FOR ANALYSIS. ANALYSIS REVEALED THAT THERE WAS ONLY ONE SET OF SETSCREW MARKS ON THE CONNECTOR PIN AND THEY ARE TOO PROXIMAL INDICATING THE LEAD WAS NOT FULLY INSERTED INTO THE THE DEVICE. |  | N | Y | LAY USER/PATIENT
2649622-2008-02617 | 2008-05-08 | 20080312 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02619 | 2008-05-08 | 20080226 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02621 | 2008-05-08 | 20080304 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02622 | 2008-05-08 | 20080225 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: THE ACTUAL DEVICE WAS NOT RECEIVED FOR EVALUATION. WE DID RECEIVE PERFORMANCE DATA COLLECTED FROM THE DEVICE AND HAVE ANALYZED THE DATA. ANALYSIS OF THAT DATA REVEALED NOISE, INCREASED IMPEDANCE FROM 500 OHMS TO 648-768 OHMS, AND AN ELEVATED SENSING INTEGRITY COUNTER. A HIGH VALUE OF THIS COUNTER CAN INDICATE A POSSIBLE INTERMITTENCY IN THE SYSTEM. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02612 | 2008-05-08 | 20080303 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02603 | 2008-05-08 | 20080222 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6930 | ASKU |  | ASKU |  | Y | N | LAY USER/PATIENT
2649622-2008-02604 | 2008-05-08 | 20080225 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-02605 | 2008-05-08 | 20080216 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02607 | 2008-05-08 |  | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-02608 | 2008-05-08 | 20080303 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY DISTAL CONDUCTOR FRACTURED; FULL LEAD RETURNED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02609 | 2008-05-08 | 20080225 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02611 | 2008-05-08 | 20080226 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE, DUE TO CONFIDENTIALITY CONCERNS. EVALUATION SUMMARY: DISTAL CONDUCTOR WAS FRACTURED; FULL LEAD WAS RETURNED FOR ANALYSIS. EVALUATION SUMMARY: NO ANOMALIES FOUND. SUPPLEMENT SENT TO CORRECT THE CONCLUSION CODE FOR THE DEVICE. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02600 | 2008-05-08 | 20080303 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02593 | 2008-05-08 | 20080229 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | OTHER: IT WAS REPORTED THE LEAD WAS EXPLANTED DUE TO FRACTURE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION RECEIVED FROM THE PATIENT'S ATTORNEY ALLEGES THE PATIENT HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND SEVERE EMOTIONAL DISTRESS. NO CONCLUSION CAN BE DRAWN. LEAD(S), FRACTURE OF. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02594 | 2008-05-08 | 20080227 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02596 | 2008-05-08 | 20070723 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE LEAD WAS EXPLANTED DUE TO A FRACTURE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION RECEIVED REPORTING AN ALLEGATION FROM AN ATTORNEY INDICATED THE PATIENT IS DECEASED AND FURTHER INDICATED THAT THE LEAD HAD EXHIBITED A FRACTURE AND/OR WAS EXPLANTED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02597 | 2008-05-08 | 20080228 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS EXPLANTED AND REPLACED DUE TO INAPPROPRIATE SHOCKS, OVERSENSING, HIGH IMPEDANCE, AND LEAD FRACTURE. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02578 | 2008-05-08 | 20080305 | Injury | Manufacturer report | SPRINT FIDELIS | DEFIBRILLATOR, AUTOMATIC IMPLANTABLE CARDIOVERTER | MPRI | 694965 | 694965 |  | THIS DEVICE WAS REPORTED AS INCLUDED IN THE FIELD ACTION. BASED ON THE INFORMATION RECEIVED AND WITHOUT THE RETURN OF THE PRODUCT, IT CANNOT BE CONFIRMED THAT THIS DEVICE PERFORMED AS DESCRIBED IN THE FIELD ACTION. IT IS INCLUDED IN THE FIELD ACTION IN THE ABUNDANCE OF CAUTION. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02579 | 2008-05-08 | 20071002 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02582 | 2008-05-08 | 20080305 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02584 | 2008-05-08 | 20080313 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-02586 | 2008-05-08 | 20080310 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-02587 | 2008-05-08 | 20080310 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-02575 | 2008-05-08 | 20080227 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-03233 | 2008-05-20 |  | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03234 | 2008-05-22 | 20080318 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03235 | 2008-05-22 | 20080430 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-03236 | 2008-05-22 | 20080409 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THAT THE RV LEAD WAS REPLACED FOLLOWING A PATIENT ALERT FOR HIGH IMPEDANCE OF GREATER THAN 3000OHMS. ADDITIONAL INFORMATION REPORTS A SUSPECTED LEAD FRACTURE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. THE PATIENT'S STATUS WAS REPORTED AS "WELL". |  | N | Y | LAY USER/PATIENT
2649622-2008-03238 | 2008-05-22 |  | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | N | Y | LAY USER/PATIENT
1051502 | 2008-05-12 | 20080425 | Malfunction | User Facility report | SPRINT FIDELIS | LEAD, ICD | MEDTRONIC, INC. CARDIAC RHYTHM MANAGEMENT | 694965 | * |  | THE PATIENT WAS ADMITTED AFTER MULTIPLE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR SHOCKS. PATIENT HAD AN IMPLANTABLE CARDIOVERTER DEFIBRILLATOR PLACED TWO YEARS PRIOR. UPON INTERROGATION OF THIS DEVICE, IT WAS CLEAR TO THE CARDIOLOGIST THAT THE PATIENT HAD A LEAD FRACTURE CAUSING OVER-SENSING OF THE ELECTRICAL INTERFERENCE WITH SUBSEQUENT IMPLANTABLE CARDIOVERTER DEFIBRILLATOR SHOCKS. IT WAS THEN DECIDED TO IMPLANT A NEW LEAD TO ENSURE PROPER FUNCTION OF THE DEVICE. THE ORIGINAL LEAD WAS CAPPED OFF AND LEFT IN PLACE. A NEW GENERATOR WAS ALSO PLACED AT THIS TIME DUE TO THE ORIGINAL GENERATOR BEING OVERLY USED BY THE FAULTY LEAD. OF NOTE IS THAT THE PATIENT RECEIVED A NOTIFICATION LETTER FROM MEDTRONIC SIX MONTHS BEFORE THIS EVENT REGARDING THE POSSIBILITY OF A LEAD MALFUNCTION WITH HIS PARTICULAR MODEL. IT IS UNCLEAR WHY THE PATIENT FAILED TO FOLLOW-UP WITH HIS PROVIDER REGARDING THIS POSSIBILITY. THE PATIENT HAS SINCE RECOVERED FROM THIS PROCEDURE AND HAS BEEN DISCHARGED HOME. |  | N | Y | INVALID DATA
1053309 | 2008-05-20 | 20080411 | Malfunction | User Facility report | SPRINT FIDELIS LEAD | LEAD, ICD | MEDTRONIC INC. | 6931 | NA |  | PATIENT MOST LIKELY WAS EXPERIENCING OVERSENSING ON LEAD WIRES, WHICH RESULTED IN UP TO 18 PAINFUL SHOCKS PRIOR TO TRANSFER TO OUR FACILITY FOR REMOVAL OF IMPLANTABLE CARDIOVERTER DEFIBRILLATOR SPRINT FIDELIS LEAD. EXPLANTED AS DEFECTIVE PRODUCT AND ON A RECALL. ====================== MANUFACTURER RESPONSE FOR LEAD WIRE, SPRINT FIDELIS======================NOT KNOWN |  | N | Y | INVALID DATA
1055950 | 2008-05-20 | 20080516 | No answer provided | User Facility report | VIRTUOSO | ICD | MEDTRONIC INC. | DDE-DDDR | * |  | DURING AUTOPSY, A CARDIOVASCULAR IMPLANTABLE DEVICE AND LEAD WERE REMOVED. |  | N | N | INVALID DATA
2649622-2008-03245 | 2008-06-04 | 20070115 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-03246 | 2008-06-04 | 20070101 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-03242 | 2008-05-28 | 20070319 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
1061155 | 2008-06-09 | 20080428 | Injury | User Facility report | SPRINT FIDELIS LEAD | LEAD, ICD | MEDTRONIC INC. | 6959-65 | * |  | ICD MALFUNCTION - LEAD FRACTURE WITH 17 ICD DISCHARGES. DEVICE HAD TO BE TURNED OFF. RECEIVED MULTIPLE SHOCKS. LEAD REPLACED WITHOUT DIFFICULTY.====================== MANUFACTURER RESPONSE FOR RIGHT VENTRICULAR PACEMAKER LEAD, FIDELIS LEAD======================THIS LEAD WAS RECALLED. |  | Y | Y | INVALID DATA
2649622-2008-03249 | 2008-06-11 | 20080221 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. FOLLOWUP DID NOT OFFER ANY ADDITIONAL INFORMATION; THEREFORE, WE ARE UNABLE TO FULLY EVALUATE THE REPORTED EVENT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THERE IS NO ALLEGATION FROM A HEALTHCARE PROFESSIONAL THAT THE DEATH WAS DEVICE RELATED. THE DEVICES HAVE SINCE BEEN RECEIVED BY THE MANUFACTURER. DETAILED ANALYSIS OF THE DEVICE WAS NOT PERFORMED AT THIS TIME DUE TO PENDING LITIGATION. THEREFORE, WE ARE UNABLE TO FULLY EVALUATE THE REPORTED EVENT. |  | N | N | LAY USER/PATIENT
2649622-2008-03254 | 2008-06-17 | 20070924 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  |  |  | Y | Y | LAY USER/PATIENT
1076316 | 2008-05-19 | 20080303 | Injury | User Facility report | SPRINT FIDELIS | LEAD, ICD | MEDTRONIC INC. | MEDTRONIC 6959 | * |  | THE PATIENT HAS A HISTORY OF LEFT BUNDLE CONGESTIVE HEART FAILURE AND AORTIC VALVE REPLACEMENT. HE HAS CHRONIC CONGESTIVE HEART FAILURE AND HAD A BIVENTRICULAR DEFIBRILLATOR INSERTED IN 2006. THIS WAS A MEDTRONIC 7299. THE DEFIBRILLATOR LEAD WAS A MEDTRONIC 6949. THIS LEAD HAS HAD A WARNING CONCERNING POSSIBLE LEAD MALFUNCTION. HE WAS DOING WELL WITH DEFIBRILLATOR CHECK IN JANUARY SHOWING THE LEAD TO BE FUNCTIONING WELL. THE PATIENT SUDDENLY HAD THE ONSET OF 6 DISCHARGES FROM HIS DEFIBRILLATOR WITHOUT ANY PRIOR WARNING. A MAGNET WAS APPLIED TO THE DEFIBRILLATOR AND HE WAS TRANSFERRED TO THE HOSPITAL WHERE INTERROGATION REVEALED THAT THERE WAS APPROPRIATE DISCHARGE SECONDARY TO NOISE IN THE SYSTEM. THERE WAS HIGH IMPEDANCE IN THE RIGHT VENTRICULAR LEAD. THIS WAS A DRAMATIC CHANGE COMPARED TO THE PACEMAKER CHECK ONE MONTH PREVIOUSLY. IT WAS FELT THAT THERE WAS A FRACTURE IN THE LEAD AND, THEREFORE, THE DEFIBRILLATOR WAS TURNED OFF. THE PATIENT HAS A ST. JUDE AORTIC VALVE. THE PATIENT WAS TAKEN TO THE CARDIAC CATH LAB WHERE HE WAS INTUBATED AND PLACED UNDER GENERAL ANESTHESIA. THE DEFIBRILLATOR POCKET WAS THEN OPENED USING SHARP AND BLUNT DISSECTION AND THE PREVIOUS DEFIBRILLATOR BATTERY WAS REMOVED FROM THE POCKET. THIS IS A MEDTRONIC 7299. THE EXISTING DEFIBRILLATOR LEAD IS A MEDTRONIC MODEL #6959. THIS WAS INSERTED IN 2006. ALL THE LEADS WERE FREED FROM SCAR TISSUE USING CAUTERY. THE EXISTING CORONARY SINUS LEAD WAS A MEDTRONIC 4194. ON INSPECTION, IT WAS EVIDENT THAT THERE WERE MULTIPLE AREAS OF BREAKS IN THE INSULATION. THE CORONARY SINUS LEAD WAS CHECKED AND FOUND TO HAVE A VERY HIGH THRESHOLD OF APPROXIMATELY 5 VOLTS AT 1.5 MSEC. PREVIOUS CHECKS HAD SHOWN THAT THE THRESHOLD WAS LESS THAN 3 VOLTS AT 0.5 MSEC. IT WAS FELT THAT THE LEAD HAD EITHER RETRACTED OR THERE WAS A PROBLEM IN THE LEAD ITSELF. THEREFORE, IT WAS DECIDED TO INSERT A NEW CORONARY SINUS LEAD. BECAUSE THE SUBCLAVIAN VEIN WAS VERY TIGHT, THE OLD RIGHT VENTRICULAR LEAD WAS THEN FREED FROM THE FASCIA. A STYLETTE WAS INSERTED AND THE SCREW WAS RETRACTED. THE LEAD WAS THEN REMOVED WITHOUT MUCH DIFFICULTY. ON INSPECTION, THERE WAS A CRACK IN THE LEAD, PROBABLY AT THE SITE OF THE OLD SLEEVE. SINCE THE LEAD DOES HAVE A RECALL, IT WILL BE SENT FOR INSPECTION. |  | Y | Y | INVALID DATA
MW5004811 | 2007-12-26 | 20061217 | Malfunction | Voluntary report | MEDTRONIC SPRINT FIDELIS LEADS | DEFIBRILLATOR-PACEMAKER | MEDTRONIC INC | 7232CXAA |  |  | IN 2005, THE ICD WAS IMPLANTED AND IN 2006, IT WAS SHOCKED OVER 39 TIMES IN ABOUT 5 HRS BEFORE IT WAS SHUT OFF IN THE ER AT THE HOSPITAL. THE DEFECT LEAD WAS REPLACED WITH ANOTHER LEAD ON THE RECALL LIST. DATES OF USE: 2005-2006. DIAGNOSIS OR REASON FOR USE: FAINTING AT CHURCH. EVENT ABATED AFTER USE STOPPED OR DOSE REDUCED? YES. |  | N | Y | 0
MW5004856 | 2007-12-26 | 20071214 | Injury | Voluntary report | MEDTRONIC SPRINT FIDELIS MODEL #6945 | DEFIBRILLATOR LEAD | MEDTRONIC, INC. | 6949 |  |  | IN 2007, I HAD A MEDTRONIC VIRTUOSO DR DEFIBRILLATOR IMPLANTED. ABOUT 7 MONTHS LATER, I WAS NOTIFIED BY MY DOCTOR AND MEDTRONIC THAT THERE WAS A RECALL OF THE SPRINT FIDELIS LEAD -MODEL #6949- USED IN MY IMPLANT. I HAVE BEEN IN CONSULTATION WITH MY PHYSICIAN WHO HAS MADE CHANGES TO MY MEDICATION. APPROX 9 MONTHS POST-IMPLANT, I EXPERIENCED TWO UNEXPLAINED "SHOCKS." I WAS CONTACTED BY MY PHYSICIAN FOR FOLLOW-UP TESTS. IT IS UNCLEAR WHETHER THIS CAN BE A CONSEQUENCE OF THE FAULTY LEADS OR NOT. |  | Y | N | HEALTH PROFESSIONAL
MW5004947 | 2008-01-09 | 20070923 | Other | Voluntary report | SPRINT FIDELIS | UNK | MEDTRONIC INC. | 6949 LEAD |  |  | IN 2007, PT WAS "SHOCKED" OVER SIXTY (60) INAPPROPRIATE TIMES BY THE MALFUNCTIONING DEVICE (A/K/A "SPRINT FIDELIS" LEADS) DUE TO A FRACTURE IN THE LEAD; BODY ROCKING SHOCKS FROM THE DEFECTIVE LEAD(S); HIGH VOLTAGE SHOCKS FROM THE DEFECTIVE LEADS; SCREAMING IN PAIN; FEAR OF IMPENDING DEATH; V-TACH; SYNCOPE; RESPIRATORY ARREST; AGONAL RESPIRATIONS; PULSELESS; APNEIC; DIED AND REQUIRED CPR WITH ASSISTED VENTILATION AND CHEST COMPRESSIONS; INTUBATION; SHOCKED BY 360 J; RESUSCITATED; DAMAGE TO HEART; DAMAGE TO RIBS; BRUISES AND CONTUSIONS; COMPLEX LIFE THREATENING SURGERY TO REPLACE THE PACEMAKER AND LEADS; HOSPITALIZATION; AND EXTREME PHYSICAL AND EMOTIONAL PAIN AND SUFFERING. PLEASE ADVISE ME WHO IS HANDLING THESE CLAIMS AND WHO I SHOULD SEND THE MEDICAL RECORDS AND HIPAA COMPLAINT AUTHORIZATIONS TO. PURSUANT TO THE FREEDOM OF INFORMATION ACT, I KINDLY REQUEST A COPY OF THE FILE CONCERNING MEDTRONIC'S FAULTY DEFIBRILLATOR WIRE (SPRINT FIDELIS # 6949 LEAD) WHICH WAS CAUSING SHOCKS TO PATIENTS. |  | N | N | 0
MW5005052 | 2008-01-16 | 20060915 | Injury | Voluntary report | SPRINT FIDELIS ICD LEADS | DEIBRILLATOR LEADS | MEDTRONIC, INC | 694865ID |  |  | PT HAD A MEDTRONIC DEFIBRILLATOR IMPLANTED IN 2005, WITH THE SPRINT FIDELIS LEADS. IN 2006, PT EXPERIENCED A SHOCK FROM THE DEFIBRILLATOR. PRIOR ADVICE FROM HIS DOCTOR WAS TO SIT DOWN, AVOID EXERTION, AND REPORT THE EVENT IF IT WAS REPEATED A SECOND TIME. ON THE NIGHT OF THE SAME DATE -OR EARLY THE DAY-, HAVING GONE TO BED FOR THE NIGHT, PT EXPERIENCED A LONG SERIES OF SHOCKS FROM THE DEFIBRILLATOR NUMBERING 11 SHOCKS IN A PERIOD OF ABOUT ONE HR. HE WAS TAKEN BY AMBULANCE TO THE HOSPITAL. WHILE AT THE HOSPITAL, HE WAS EXAMINED BY A MEDTRONIC MANUFACTURER'S REP WHO STATED, THAT HE COULD NOT READ THE HISTORY FROM THE DEVICE. PT WAS RELEASED FROM THE HOSPITAL FOUR DAYS LATER AND FLEW HOME ONE DAY LATER. WHILE AT THE HOSPITAL AND AFTERWARD UNDER THE CARE OF HIS OWN CARDIOLOGIST, PT WAS GIVEN AMIODARONE, AND SUBSEQUENTLY SUFFERED SEVERE LUNG DAMAGE THAT NOW REQUIRES HIM TO HAVE 24/7 OXYGEN. |  | Y | Y | 0
2649622-2007-03229 | 2007-12-27 | 20070923 | No answer provided | Manufacturer report | SPRINT FIDELIS | LEAD, ICD | MEDTRONIC INC, CARDIAC RHYTHM MGMT. | 6949-58CM | * |  |  |  | Y | Y | INVALID DATA
983741 | 2008-01-02 | 20071220 | Malfunction | User Facility report | SPRINT FIDELIS | LEAD, ICD | MEDTRONIC INC. | 6944 | * |  | RV LEAD REPLACEMENT PER RECALL. MD HAD DIFFICULTY WITH THE SET SCREW IN THE ICD DEVICE AND ELECTED TO REPLACE THE GENERATOR AS WELL. |  | N | Y | INVALID DATA
983749 | 2008-01-02 | 20071220 | No answer provided | User Facility report | SPRINT FIDELIS | LEAD, ICD | MEDTRONIC INC. | 6949 | * |  | REPLACED AS PER RECALL. |  | N | N | INVALID DATA
983761 | 2008-01-02 | 20071214 | Malfunction | User Facility report | SPRINT FIDELIS | LEAD, ICD | MEDTRONIC, INC. | 6949 | * |  | RECALLED MEDTRONIC ICD LEAD REMOVED AND REPLACED. |  | N | Y | INVALID DATA
983898 | 2007-12-28 | 20071223 | Malfunction | User Facility report | SPRINT FIDELIS | LEAD, ICD | MEDTRONIC INC. | 6931-58 | * |  | THE RECALLED FIDELIS 6931 RV LEAD. PATIENT WAS SEEN IN OFFICE RECENTLY WITH A NORMAL LEAD FUNCTIONING. THE PATIENT WAS SEEN IN ED FIVE DAYS LATER WITH 27 SHOCKS FROM THE ICD. THE MEDTRONIC REPRESENTATIVE WAS CALLED IN AND INTERROGATED THE DEVICE AND THEY FOUND THAT THE 6931 RV LEAD WAS FRACTURED. |  | N | Y | INVALID DATA
986164 | 2008-01-02 | 20071217 | Malfunction | User Facility report | SPRINT FIDELIS | LEAD, ICD | MEDTRONIC, INC. | 6931 | * |  | ICD LEAD REPLACED PER RECALL. |  | N | Y | INVALID DATA
2649622-2008-00824 | 2008-01-15 | 20071114 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES FOUND; THE ANALYST NOTED THAT THERE WAS ONLY ONE SET OF SETSCREW MARKS ON THE IS-1 PIN AND IT IS TOO PROXIMAL, INDICATING THE LEAD WAS NOT FULLY INSERTED INTO THE DEVICE;FULL LEAD RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-00201 | 2008-01-10 | 20070816 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00202 | 2008-01-10 | 20071010 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-00209 | 2008-01-10 | 20061215 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS ALLEGED BY AN ATTORNEY THAT THE PATIENT RECEIVED FIVE OR SIX SHOCKS FOR NO APPARENT REASON. AT THE TIME OF THE INCIDENT, THE PATIENT WAS PREPARING TO LEAVE HIS HOME, GOT DRESSED, PICKED UP HIS PERSONAL ITEMS NEEDED FOR THE DAY, AND REACHED FOR HIS COAT WHEN THE DEVICE WENT OFF. THE STATUS OF THE DEVICE AND LEAD IS UNKNOWN. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. FURTHER INFORMATION FROM AN ATTORNEY ALLEGES THE PLAINTIFF SUFFERED INJURIES AS A RESULT OF THE RECALLED LEAD. IN 2005, THE PLAINTIFF HAD THE LEAD IMPLANTED AND STARTED RECEIVING INAPPROPRIATE SHOCKS IN 2006. TWO MONTHS LATER, PLAINTIFF HAD THE LEAD EXPLANTED. AS A RESULT OF THE LEAD, PLAINTIFF SUFFERED INJURY AND DAMAGES AS SET FORTH IN THE MASTER COMPLAINT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00193 | 2008-01-10 | 20071008 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: (B)(4) DETAILED ANALYSIS OF THE LEADS WAS NOT PERFORMED AT THIS TIME DUE TO PENDING LITIGATION. THEREFORE, WE ARE UNABLE TO FULLY EVALUATE THE REPORTED EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00197 | 2008-01-10 | 20070228 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | OTHER: THE PATIENT REPORTED THAT SHE RECEIVED 33 SHOCKS FROM HER DEVICE. SHE SAID SHE WAS TOLD BY HER PHYSICIAN THAT THE LEAD WAS SENSING INCORRECTLY, SO SHE HAD TO HAVE A NEW LEAD IMPLANTED. AS A RESULT, THE PACE/SENSE PORTION OF THE MODEL 6949 LEAD WAS CAPPED AND REPLACED WITH A NEW LEAD. THE HIGH VOLTAGE PORTION REMAIN IN USE. ADDITIONAL INFORMATION WAS LATER RECEIVED IN A LAWSUIT ALLEGING THE "PLAINTIFF HAS SUFFERED SEVERE EMOTIONAL TRAUMA, PHYSICAL CONSEQUENCES AND LONG CONTINUED EMOTIONAL DISTURBANCE." NO SPECIFIC DETAILS WERE RECEIVED, AND NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. NO CONCLUSION CAN BE DRAWN. DEVICE REMAINS ACTIVATED, SENSITIVITY, SHOCK, INAPPROPRIATE. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00198 | 2008-01-10 | 20071008 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE LEAD WAS EXPLANTED AND REPLACED DUE TO OVERSENSING AND FRACTURE. ADDITIONAL INFORMATION WAS LATER RECEIVED FROM THE PATIENT REPORTING SHE RECEIVED TWO SHOCKS DUE TO THE "DEFECTIVE" LEAD. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00192 | 2008-01-10 | 20070321 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00182 | 2008-01-10 | 20070503 | Injury | Manufacturer report | SPRINT FIDELIS | DEFIBRILLATOR, AUTOMATIC IMPLANTABLE CARDIOVERTER | MEDTRONIC PUERTO RICO, INC. | 6949 | 6949 |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-00183 | 2008-01-10 | 20071002 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00184 | 2008-01-10 | 20071004 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00186 | 2008-01-10 | 20071011 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00187 | 2008-01-10 | 20071008 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00188 | 2008-01-10 | 20071007 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT RECEIVED 16 INAPPROPRIATE SHOCKS. OVERSENSING AND HIGH LEAD IMPEDANCE OF GREATER THAN 2500 OHMS WERE ALSO OBSERVED. THE HIGH VOLTAGE THERAPY HAS BEEN TURNED OFF PENDING THE PATIENT'S DECISION ABOUT A LEAD REVISION. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. FURTHER INFORMATION FROM AN ATTORNEY ALLEGES THE PLAINTIFF EXPERIENCED PHYSICAL, AND OTHER INJURY AS A RESULT OF THE RECALLED LEAD INCLUDING INAPPROPRIATE SHOCKS, FRACTURE OR OTHER INJURY REQUIRING MEDICAL INTERVENTION. THE ATTORNEY FURTHER ALLEGES MENTAL ANGUISH AND SEVERE EMOTIONAL DISTRESS. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00189 | 2008-01-10 | 20071008 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE RV LEAD WAS CAPPED DUE TO LEAD FRACTURE. A LAWSUIT FURTHER ALLEGED THAT THE PATIENT "SUFFERED PHYSICAL AND OTHER INJURY" AND THAT THE PATIENT "EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY REQUIRING MEDICAL INTERVENTION INCLUDING BUT NOT LIMITED TO SURGERY, REMOVAL AND/OR REPLACEMENT OF THE DEFECTIVE [LEAD]". NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00191 | 2008-01-10 | 20050430 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00181 | 2008-01-10 | 20070615 | Injury | Manufacturer report | SPRINT FIDELIS | DEFIBRILLATOR, AUTOMATIC IMPLANTABLE CARDIOVERTER | MPRI | 6949 | 6949 |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. EVALUATION SUMMARY: THE ACTUAL DEVICE WAS NOT RECEIVED FOR EVALUATION. WE DID RECEIVE PERFORMANCE DATA COLLECTED FROM THE DEVICE AND HAVE ANALYZED THE DATA. THE VENTRICULAR SENSING INTEGRITY COUNTER INDICATES A POSSIBLE INTERMITTENCY IN THE SYSTEM. NO IMPEDANCE ANOMALIES WERE FOUND THROUGH THE LAST MEASURED VALUE IN THE RECORD FROM 10 DAYS PRIOR TO EXPLANT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00172 | 2008-01-10 | 20071001 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00173 | 2008-01-10 | 20071001 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00175 | 2008-01-10 | 20071001 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00177 | 2008-01-10 | 20071003 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS FRACTURED AND THE PATIENT RECEIVED MULTIPLE INAPPROPRIATE SHOCKS. THE LEAD WAS EXPLANTED AND REPLACED. THE PATIENT'S STATUS WAS NOTED AS 'OK'. FURTHER INFORMATION RECEIVED INDICATES THAT THE PATIENT PRESENTED TO THE ER WITH CHEST PAIN AND REPEATED SHOCKS. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00178 | 2008-01-10 | 20071001 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00179 | 2008-01-10 | 20071005 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE, DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY : NO ANOMALIES FOUND; PROXIMAL CONDUCTOR FRACTURED; FULL LEAD ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00180 | 2008-01-10 | 20071003 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00171 | 2008-01-10 | 20071005 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00161 | 2008-01-10 | 20071001 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00162 | 2008-01-10 | 20071005 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00165 | 2008-01-10 | 20071002 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00166 | 2008-01-10 | 20071002 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; DISTAL SEGMENT RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00167 | 2008-01-10 | 20071005 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | A CALL WAS RECEIVED THROUGH TECHNICAL SERVICES, REPORTING THE PATIENTRECEIVED 30 INAPPROPRIATE SHOCKS DUE TO OVERSENSING. THE LEAD WAS REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION RECEIVED ON THE EVENT REPORTING LEAD FRACTURE. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00168 | 2008-01-10 | 20071002 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS REMOVED DUE TO FRACTURE VISIBLE ON X-RAY. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. REVIEW OF CARELINK DATA INDICATES HIGH HV IMPEDANCE VALUES AND RV OVERSENSING NOTED. |  | N | Y | LAY USER/PATIENT
2649622-2008-00169 | 2008-01-10 | 20071005 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | A CALL WAS RECEIVED THROUGH TECHNICAL SERVICES REPORTING PT RECEIVED INAPPROPRIATE SHOCKS AND IMPEDANCE VALUES WERE FLUCTUATING. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00170 | 2008-01-10 | 20071001 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00160 | 2008-01-10 | 20071004 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00149 | 2008-01-10 | 20071004 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00150 | 2008-01-10 | 20070901 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00151 | 2008-01-10 | 20071008 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00152 | 2008-01-10 | 20070926 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE PATIENT RECEIVED MULTIPLE INAPPROPRIATE SHOCKS DUE TO NOISE, AND THERE WAS AN APPARENT LEAD FRACTURE. THE LEAD WAS EXPLANTED. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00153 | 2008-01-10 | 20070921 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE PATIENT HAD RECEIVED INAPPROPRIATE SHOCKS DUE TO OVERSENSING. A LEAD FRACTURE WAS SUSPECTED, THOUGH NOT OBSERVED ON FLUOROSCOPY. THE LEAD WAS EXPLANTED AND REPLACED. THE HCP WAS UNABLE TO ADVANCE THE STYLET AND NOTED THE HELIX WAS NOT OPERABLE. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-00154 | 2008-01-10 | 20070907 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE RV HIGH VOLTAGE LEAD IMPEDANCES WERE OUT OF RANGE AT GREATER THAN 200 OHMS. THE LEAD WAS EXPLANTED AND REPLACED. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-00155 | 2008-01-10 | 20070921 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00157 | 2008-01-10 | 20071114 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00159 | 2008-01-10 | 20071003 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: THE ACTUAL DEVICE WAS NOT RECEIVED FOR EVALUATION. WE DID RECEIVE PERFORMANCE DATA COLLECTED FROM THE DEVICE AND HAVE ANALYZED THE DATA, AND THERE WERE NO ANOMALIES FOUND. |  | Y | N | LAY USER/PATIENT
2649622-2008-00144 | 2008-01-10 | 20071004 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00145 | 2008-01-10 | 20071001 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00147 | 2008-01-10 | 20071002 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00148 | 2008-01-10 | 20071004 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: THE FULL LEAD WAS RETURNED FOR ANALYSIS WHICH DETECTED A CONDUCTOR FRACTURE IN THE BIFURCATION OF THE LEAD. THE FRACTURE OF THE RV CONDUCTOR APPEARS TO HAVE OCCURRED AT THE TIME OF LEAD REVISION. |  | N | Y | LAY USER/PATIENT
2649622-2008-00131 | 2008-01-10 | 20071002 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | TWO INAPPROPRIATE SHOCKS DUE TO OVERSENSING REPORTED. IMPEDANCE OF 1900 OHMS ALSO REPORTED. THE LEAD WAS CAPPED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. IT WAS FURTHER REPORTED THAT AT REPLACEMENT, THE RING OF THE LEAD WAS NOTED TO BE GIVING THE NOISE. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00133 | 2008-01-10 |  | No answer provided | Manufacturer report | SPRINT FIDELIS | ASKU | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | SEVEN INAPPROPRIATE SHOCKS WERE REPORTED. IMPEDANCE >3000 OHMS AND OVERSENSING WERE ALSO REPORTED. A PACE/SENSE LEAD WAS IMPLANTED TO REPLACE THE PACE/SENSE PORTION OF THE HIGH VOLTAGE LEAD. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. LAWSUIT ALLEGES THE FIDELIS LEAD CAUSED HIS PULSE GENERATOR TO MALFUNCTION AND SUBJECT THE PATIENT TO RANDOM SHOCKING SENSATIONS. THE RV LEAD WAS FOUND TO HAVE A FRACTURE, AND A SECOND SURGERY WAS CONDUCTED TO REPLACE THE DEFECTIVE LEAD. THE PATIENT HAS SUFFERED DAMAGES, INCLUDING, BUT NOT LIMITED TO, LIFE THREATENING ELECTRICAL SHOCKS TO HIS HEART, LOSS OF CONSCIOUSNESS, A SECOND SURGERY AND THE RECOVERY ASSOCIATED THEREWITH. |  | Y | Y | 0
2649622-2008-00134 | 2008-01-10 | 20071008 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PACE/SENSE PORTION OF THE RV LEAD WAS CAPPED AND REPLACED DUE TO INAPPROPRIATE THERAPIES. OVERSENSING WAS ALSO NOTED. SENSING AND LEAD IMPEDANCE WERE OBSERVED TO BE STABLE. THE HIGH VOLTAGE PORTION OF THE LEAD REMAINS IN USE. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. THE LEAD WAS LATER RETURNED WITH NO INFORMATION. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00137 | 2008-01-10 | 20071002 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00130 | 2008-01-10 | 20071010 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT RECEIVED EXTRA SHOCKS (INAPPROPRIATE THERAPY). AN ABRUPT IMPEDANCE RISE TO GREATER THAN 3000 OHMS AND OVERSENSING WERE ALSO REPORTED. THE LEAD WAS CAPPED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. IT WAS FURTHER REPORTED THE LEAD WAS CAPPED MOST LIKELY DUE TO A FRACTURE. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00119 | 2008-01-10 | 20071011 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00120 | 2008-01-10 | 20070928 | Injury | Manufacturer report | SPRINT FIDELIS | DEFIBRILLATOR, AUTOMATIC IMPLANTABLE CARDIOVERTER | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  | IT WAS REPORTED THAT THE LEAD WAS REPLACED DUE TO HV IMPEDANCE VARIABILITY. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00121 | 2008-01-10 | 20070928 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE LEAD WAS REPLACED DUE TO MULTIPLE INAPPROPRIATE THERAPIES CAUSED BY OVERSENSING. VARYING/HIGH IMPEDANCE WAS ALSO NOTED. THE LEAD WAS REPLACED. ADDITIONAL INFORMATION WAS SUBSEQUENTLY RECEIVED FROM AN ATTORNEY ALLEGING THE PATIENT EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE, OR OTHER INJURY REQUIRING REPLACEMENT OF THE DEFECTIVE LEAD. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00122 | 2008-01-10 | 20071003 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE LEAD WAS EXPLANTED AND REPLACED BECAUSE THE PATIENT RECEIVED INAPPROPRIATE SHOCKS. FRACTURE WAS INDICATED. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00123 | 2008-01-10 | 20070830 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-00124 | 2008-01-10 | 20070925 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00125 | 2008-01-10 | 20070914 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00126 | 2008-01-10 | 20070711 | Injury | Manufacturer report | SPRINT FIDELIS | DEFIBRILLATOR, AUTOMATIC IMPLANTABLE CARDIOVERTER | MPRI | 6931 | 6931 |  | OTHER: IT WAS REPORTED THAT THE LEAD WAS EXPLANTED DUE TO FRACTURE. ADDITIONAL INFORMATION WAS SUBSEQUENTLY RECEIVED REPORTING OVERSENSING. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ACTUAL DEVICE INVOLVED IN INCIDENT WAS EVALUATED. ELECTRICAL TESTS PERFORMED; MECHANICAL TESTS PERFORMED. VISUAL EXAMINATION: LEAD CONDUCTOR, COMPONENT/SUBASSEMBLY FAILURE. DEVICE WAS OUT OF SPECIFICATION IN A MANNER THAT RELATES TO EVENT; LEAD(S), FRACTURE OF, OVERSENSING. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00127 | 2008-01-10 | 20071001 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | INAPPROPRIATE SHOCKS CAUSED BY OVERSENSING AND AN IMPEDANCE INCREASE WERE REPORTED. THE LEAD WAS CAPPED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. APPARENT FRACTURE WAS FURTHER REPORTED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00128 | 2008-01-10 | 20071001 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00129 | 2008-01-10 | 20071001 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00118 | 2008-01-10 | 20071002 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-00113 | 2008-01-10 | 20070705 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00114 | 2008-01-10 | 20071002 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00115 | 2008-01-10 | 20071001 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00116 | 2008-01-10 |  | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00117 | 2008-01-10 | 20071013 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT IS IN A HOSPITAL IN ANOTHER COUNTRY. THE PATIENT RECEIVED SHOCKS. THE ICD AND LEAD ARE SCHEDULED FOR REPLACEMENT. ADDITIONAL INFORMATION WAS LATER RECEIVED REPORTING THE LEAD WAS EXPLANTED AND REPLACED, DUE TO MULTIPLE SHOCKS, SECONDARY TO LEAD FRACTURE AND SUSPECTED NOISE. THE DEVICE WAS ALSO EXPLANTED AND REPLACED, DUE TO EARLY BATTERY DEPLETION, CAUSED BY THE MULTIPLE SHOCKS. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00111 | 2008-01-10 | 20071001 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | N | LAY USER/PATIENT
2649622-2008-00110 | 2008-01-10 | 20071001 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00101 | 2008-01-10 | 20070920 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE PATIENT HAD BEEN RECEIVING MULTIPLE SHOCKS AND FELT "SOMEWHAT TRAUMATIZED." THE LEAD WAS REPLACED DUE TO LEAD FRACTURE. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | N | LAY USER/PATIENT
2649622-2008-00102 | 2008-01-10 | 20070924 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. DETAILED ANALYSIS OF THE DEVICE WAS NOT PERFORMED AT THIS TIME DUE TO PENDING LITIGATION. THEREFORE, WE ARE UNABLE TO FULLY EVALUATE THE REPORTED EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00103 | 2008-01-10 |  | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE PATIENT HEARD THE ALARM SOUND AND WENT TO THE EMERGENCY DEPARTMENT. HE WAS TOLD A LEAD HAD FRACTURED. FURTHER INVESTIGATION REVEALED THE LEAD WAS REPLACED FOLLOWING MULTIPLE INAPPROPRIATE SHOCKS DUE TO NOISE SENSING AND IMPEDANCE ELEVATION TO 4758 OHMS. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION FURTHER REPORTS THERE WAS A VISUALIZED BREAK IN THE LEAD INSULATION AND IT WAS DIFFICULT TO UNSCREW THE LEAD DESPITE HUNDREDS OF TURNS. |  | Y | N | LAY USER/PATIENT
2649622-2008-00107 | 2008-01-10 | 20070920 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00108 | 2008-01-10 | 20071003 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | 6949 |  | THIS MEDICAL DEVICE REPORT IS BEING RESUBMITTED TO THE EMDR SYSTEM DUE TO A CONNECTION ERROR WITHIN THE EMDR SYSTEM. THE EVENT WAS PREVIOUSLY REPORTED TO THE EMDR PORTAL AND AN ACKNOWLEDGEMENT 3 RECEIVED, BUT THE CONNECTION ERROR PREVENTED THE EMDR SYSTEM FROM DOCUMENTING THE REPORT. THE REPORT NUMBER IN THIS REPORT IS THE SAME NUMBER AS THE IMPACTED REPORT THAT HAS THE CONNECTION ISSUE. PRODUCT EVENT SUMMARY: THE FULL LEAD WAS RETURNED IN SEGMENTS, ANALYZED, AND THE RV (RIGHT VENTRICULAR) DEFIBRILLATION COIL DEVELOPED A FRACTURE DUE TO FLEXING WHILE IN VIVO. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00092 | 2008-01-10 | 20070926 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00093 | 2008-01-10 | 20070914 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED BY A COMPETITOR THAT HIGH PACING IMPEDANCE WAS DETECTED IN 2007. THE LEAD SEEMED TO HAVE FRACTURED. REPLACEMENT WAS SCHEDULED THE NEXT DAY. ADDITIONAL INFORMATION WAS SUBSEQUENTLY RECEIVED REPORTING THAT THE PATIENT RECEIVED INAPPROPRIATE SHOCKS CAUSED BY NON-PHYSIOLOGIC OVERSENSING DUE TO LEAD FRACTURE. PACING IMPEDANCE WAS ALSO HIGH AT 3000 OHMS. THE LEAD WAS CAPPED AND REPLACED. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00094 | 2008-01-10 | 20070919 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00095 | 2008-01-10 | 20071001 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00097 | 2008-01-10 | 20071011 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00098 | 2008-01-10 | 20071010 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00091 | 2008-01-10 | 20071008 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00084 | 2008-01-10 | 20070921 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THERE WAS OVERSENSING, AND THE LEAD WAS REPLACED. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED AND AN ALLEGATION FROM AN ATTORNEY INDICATED THE PATIENT IS DECEASED AND FURTHER INDICATED THAT THE LEAD HAD EXHIBITED A FRACTURE AND/OR WAS EXPLANTED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00074 | 2008-01-10 |  | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | PT'S WIFE REPORTS HUSBAND RECEIVED MULTIPLE SHOCKS AND ASKED IF CARE ALERT WAS SENT TO DR'S OFFICE. FOLLOWUP WITH PHYSICIAN'S OFFICE INDICATES PT HAD BEEN ADMITTED TO THE HOSPITAL 13 DAYS LATER. ADMITTING DIAGNOSIS WAS NOT AVAILABLE. NO FURTHER INFORMATION WAS AVAILABLE FROM THE DR'S OFFICE. THE PT LATER REPORTS KIDNEYS AFFECTED AND THAT THE ICD IS NOW TURNED OFF WHILE IN THE HOSPITAL. FOLLOW UP INFORMATION WAS RECEIVED WHICH INDICATES THE DEVICE WAS FUNCTIONING APPROPRIATELY; PT WAS HAVING INCESSANT VT WHICH THE DEVICE WAS TREATING. 30 PLUS APPROPRIATE AND EFFECTIVE THERAPIES WERE DELIVERED. NO CONCLUSION CAN BE DRAWN DEVICE REMAINS ACTIVATED SHOCK, INAPPROPRIATE. |  | Y | N | LAY USER/PATIENT
2649622-2008-00075 | 2008-01-10 | 20071003 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00080 | 2008-01-10 | 20070918 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT RECEIVED MULTIPLE INAPPROPRIATE THERAPIES DUE TO OVERSENSING. THE LEAD IMPEDANCE WAS ALSO NOTED TO FLUCTUATE FROM NORMAL RANGE TO HIGH. ADDITIONAL INFORMATION WAS SUBSEQUENTLY RECEIVED REPORTING THE PATIENT RECEIVED INAPPROPRIATE SHOCKS, WHICH WERE FELT TO BE DUE TO LEAD MIGRATION. THE LEAD WAS CAPPED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2182208-2008-00005 | 2008-01-10 | 20071023 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC, INC. | 6948 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: THE ACTUAL DEVICE WAS NOT RECEIVED FOR EVALUATION. HOWEVER, PERFORMANCE DATA WAS COLLECTED FROM THE DEVICE AND ANALYZED. NOISE AND HIGH IMPEDANCE WAS NOTED. THE LIFETIME VENTRICULAR SIC (SENSING INTEGRITY COUNTER) WAS 315,696. A HIGH VALUE OF THIS COUNTER CAN INDICATE A POSSIBLE INTERMITTENCY IN THE SYSTEM. THE VENTRICULAR PACING IMPEDANCE MEASUREMENT WAS CONSISTENT IN THE RANGE OF 400 - 500 OHMS UNTIL THE WEEK ENDING 2007, WHEN THE MAXIMUM VALUE WAS 648 OHMS. THE VALUE INCREASED FROM THIS POINT DURING THE MEASUREMENT PERIOD, THROUGH THE WEEK ENDING 2007, WHEN THE MINIMUM WAS 2384 OHMS AND MAXIMUM WAS 2544 OHMS. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00082 | 2008-01-10 | 20070914 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT IMPEDANCE WAS OUT OF RANGE. THE LEAD WAS REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION RECEIVED REPORTED LEAD OVERSENSING, HIGH IMPEDANCES, AND DELIVERY OF INAPPROPRIATE SHOCKS. THE LEAD WAS PARTIALLY REMOVED AND CAPPED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00060 | 2008-01-10 | 20070919 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00061 | 2008-01-10 | 20070911 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: THE ACTUAL DEVICE WAS NOT RECEIVED FOR EVALUATION. HOWEVER, PERFORMANCE DATA WAS COLLECTED FROM THE DEVICE AND ANALYZED. NOISE WAS NOTED. THE LIFETIME VENTRICULAR SIC (SENSING INTEGRITY COUNTER) WAS 1760. THESE COUNTS BEGAN IN 2007. A HIGH VALUE OF THIS COUNTER CAN INDICATE A POSSIBLE INTERMITTENCY IN THE SYSTEM. WEEKLY RV (RIGHT VENTRICULAR) IMPEDANCE MEASUREMENTS WERE AROUND 400 OHMS, WITH A VALUE OF 896 OHMS RECORDED AT ABOUT 8 DAYS PRIOR AND 592 OHMS AT ABOUT 4 DAYS AND 3 DAYS LATER |  | Y | Y | LAY USER/PATIENT
2649622-2008-00062 | 2008-01-10 | 20070906 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00063 | 2008-01-10 | 20070209 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00051 | 2008-01-10 | 20070910 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. THIS EVENT INVOLVES A LEGAL CASE IN PROGRESS OR POTENTIAL LITIGATION. THE PROPRIETARY NATURE OF THE EVENT MAY AFFECT FOLLOW UP EFFORTS. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: (B)(4) THE LEAD WAS RETURNED. DETAILED ANALYSIS OF THE DEVICE WAS NOT PERFORMED AT THIS TIME DUE TO PENDING LITIGATION. THEREFORE, WE ARE UNABLE TO FULLY EVALUATE THE REPORTED EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00052 | 2008-01-10 | 20070723 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00053 | 2008-01-10 | 20070905 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00054 | 2008-01-10 | 20070907 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00055 | 2008-01-10 | 20070912 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE RIGHT VENTRICULAR LEAD WAS REPLACED DUE TO A LEAD FRACTURE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION RECEIVED REPORTED THE LEAD HAD ALSO BEEN SHOWING HIGH IMPEDANCES. |  | N | Y | LAY USER/PATIENT
2649622-2008-00056 | 2008-01-10 | 20070612 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00045 | 2008-01-10 | 20070915 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00046 | 2008-01-10 | 20070817 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6930 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00047 | 2008-01-10 | 20070823 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00048 | 2008-01-10 | 20070523 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | A LAWSUIT ALLEGED THAT THE PATIENT HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, AND/OR CONSEQUENTIAL DAMAGES DUE TO THE IMPLANTED LEAD. THE DEVICE WAS RETURNED TO THE MANUFACTURER, ANALYZED AND SUBSEQUENTLY TESTED OUT OF SPECIFICATION. NO SPECIFIC PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-00049 | 2008-01-10 | 20070723 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00043 | 2008-01-10 | 20070813 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00035 | 2008-01-10 | 20071016 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | A CALL WAS RECEIVED THROUGH TECHNICAL SERVICES REPORTING THE RV (RIGHT VENTRICULAR) LEAD IMPEDANCE WAS MEASURED AT GREATER THAN 3000 OHMS TWICE, 1536 OHMS, AND 500 OHMS. THE SHORT INTERVAL COUNTER MEASURED GREATER THAN 300, INDICATIVE OF OVERSENSING. ADDITIONAL INFORMATION SUBSEQUENTLY PROVIDED BY THE MEDTRONIC REPRESENTATIVE REPORTED THE EP MET WITH THE PATIENT, GAVE HER A MAGNET, BUT INSTRUCTED HER NOT TO USE IT UNLESS THE EMT WAS AT HER SIDE. THE DOCTOR MADE NO OTHER CHANGES. THE LEAD WAS EVENTUALLY REMOVED POST VT (VENTRICULAR TACHYCARDIA) ABLATION. IT APPEARED TO BE FRACTURED, AND SUBCLAVIAN CRUSH WAS QUESTIONED. THE DEVICE HAD BEEN PROGRAMMED DOO FOR TWO MONTHS PRIOR TO EXTRACTION, DUE TO OVERSENSING ISSUES. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00036 | 2008-01-10 | 20070802 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: THE ACTUAL DEVICE WAS NOT RECEIVED FOR EVALUATION. HOWEVER, PERFORMANCE DATA WAS COLLECTED FROM THE DEVICE AND ANALYZED. NOISE WAS OBSERVED. THE LIFETIME SIC (SENSING INTEGRITY COUNTER) WAS 19,421. A HIGH VALUE OF THIS COUNTER CAN INDICATE A POTENTIAL INTERMITTENCY IN THE SYSTEM. OTHER: IT WAS REPORTED THAT THERE WERE INCREASED THRESHOLDS ON THE LEAD, ALONG WITH OVERSENSING AND NOISE. THE LEAD WAS EXPLANTED AND REPLACED. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. NO CONCLUSION CAN BE DRAWN. OVERSENSING NOISE HIGH THRESHOLD. |  | N | Y | LAY USER/PATIENT
2649622-2008-00037 | 2008-01-10 | 20071029 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | SENSING ISSUES WERE REPORTED, AND REVIEW OF DEVICE DATA SHOWED THE PACING IMPEDANCE AND SENSING HAVE BEEN ERRACTIC SINCE IMPLANT, INCLUDING NOISY EGMS AND SHORT INTERVALS. BASED ON THE EARLY IMPEDANCE CHANGE, IT WAS BELIEVED TO BE 'MORE OF A SETSCREW PROBLEM'. X-RAY SHOWED QUESTIONABLE PIN PLACEMENT IN THE DEVICE HEADER, HOWEVER, THE PHYSICIAN COULD NOT VISUALLY CONFIRM A POOR PLACEMENT. THE LEAD WAS CAPPED, AND THE ICD WAS EXPLANTED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-00030 | 2008-01-10 | 20070514 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT WENT TO THE HOSPITAL AFTER RECEIVING INAPPROPRIATE SHOCKS. THERE WAS OVERSENSING, AND THE VENTRICULAR PACE/SENSE IMPEDANCE WAS GREATER THAN 3000 OHMS. A NEW LEAD WAS PLACED FOR PACING/SENSING. THE HIGH VOLTAGE PORTION OF THIS LEAD REMAINS IN USE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED REPORTING THAT THE SVC AND DEFIB IMPEDANCES HAVE BEEN GREATER THAN 200 OHMS SINCE JULY. IT HAS BEEN FURTHER REPORTED THAT THERE WAS AN APPARENT FRACTURE OF THE HIGH VOLTAGE COIL. THE HIGH VOLTAGE LEAD WAS EXPLANTED AND REPLACED WITH ANOTHER HIGH VOLTAGE LEAD. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00029 | 2008-01-10 | 20071001 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-00026 | 2008-01-10 | 20071029 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00022 | 2008-01-10 | 20051125 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00023 | 2008-01-10 | 20060725 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00025 | 2008-01-10 | 20060719 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00021 | 2008-01-10 | 20071004 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00014 | 2008-01-10 | 20071109 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: NO ANOMALIES FOUND; PARTIAL LEAD RETURNED IN SEGMENTS. |  | N | Y | LAY USER/PATIENT
2649622-2008-00015 | 2008-01-10 | 20071017 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THAT THE PATIENT RECEIVED INAPPROPRIATE THERAPY DUE TO NOISE, AND THERE WAS OVERSENSING. PACING AND HV (HIGH VOLTAGE) THERAPIES WERE PROGRAMMED OFF UNTIL THE LEAD COULD BE REPLACED. HOWEVER, WITHOUT BI-VENTRICULAR PACING, THE PATIENT DECOMPENSATED AND WENT INTO RENAL FAILURE. THE PATIENT IS NOW UNABLE TO TOLERATE A PROCEDURE TO REPLACE THE LEAD, AND IS CURRENTLY IN THE INTENSIVE CARE UNIT. THE PATIENT WAS REPORTEDLY "NOT DOING WELL," DUE TO NOT RECEIVING BI-VENTRICULAR PACING AND HAS SINCE HAD THE LEAD REPLACED. ADDITIONAL INFORMATION WAS RECEIVED WITH THE RETURNED LEAD REPORTING THAT THERE WAS AN APPARENT LEAD FRACTURE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00017 | 2008-01-10 |  | Injury | Manufacturer report | SPRINT FIDELIS | DEFIBRILLATOR, AUTOMATIC IMPLANTABLE CARDIOVERTER | MPRI | 6949 | 6949 |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES WERE FOUND; FULL LEAD WAS RETURNED FOR ANALYSIS. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00019 | 2008-01-10 | 20060512 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00012 | 2008-01-10 | 20070918 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE LEAD WAS RETURNED TO THE MANUFACTURER, ANALYZED, AND TESTED OUT OF SPECIFICATION. ADDITIONAL INFORMATION WAS LATER RECEIVED REPORTING INCREASED/HIGH PACING IMPEDANCE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-00002 | 2008-01-10 | 20071205 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THE IMPEDANCE WAS NOTED TO BE 500 - 1700 AND FLUCTUATING. ONE EPISODE OF OVERSENSING WAS NOTED WITH NO SHOCK DELIVERED. THE LEAD WAS EXPLANTED AND REPLACED DUE TO HIGH IMPEDANCE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00003 | 2008-01-10 | 20071106 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THERE WAS HIGH PACING IMPEDANCE. EXCESSIVE SHORT V-V INTERVALS WERE NOTED, INDICATING OVERSENSING AND INTERMITTENT LOSS OF CAPTURE WAS ALSO OBSERVED. THERE WAS SOME DISCONTINUITY SEEN IN THE DISTAL SECTION OF THE LEAD, BEYOND THE COILS.THERE WAS A 1CM RADIUS LOOP IN THE LEAD, IN THE POCKET AREA, WITH DIFFICULTY PASSING THE STYLET. THE PACE/SENSE PORTION OF THE LEAD WAS REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED, STATING THE LEAD WAS FRACTURED. IT WAS REPLACED. |  | N | Y | LAY USER/PATIENT
2649622-2008-00005 | 2008-01-10 | 20071221 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED BY THE PATIENT'S FATHER SHE WAS HOSPITALIZED DUE TO 'ISSUE WITH LEAD'. FOLLOW-UP INFORMATION INDICATES LEAD WAS EXPLANTED AND REPLACED DUE TO LEAD FRACTURE. IT WAS FURTHER REPORTED THAT THERE WAS INCREASED/HIGH IMPEDANCE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | N | N | LAY USER/PATIENT
2649622-2008-00006 | 2008-01-10 | 20071109 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-00007 | 2008-01-10 | 20071201 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE RV LEAD WAS EXPLANTED AND REPLACED DUE TO OVERSENSING, HIGH IMPEDANCE OF GREATER THAN 3000 OHMS, AND AN APPARENT LEAD FRACTURE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-00821 | 2008-01-10 | 20071126 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THERE WAS NO CAPTURE, OVERSENSING, AND AN APPARENT LEAD FRACTURE. THE LEAD WAS EXPLANTED AND REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-00822 | 2008-01-10 | 20071116 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00823 | 2008-01-10 | 20070827 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00820 | 2008-01-10 | 20071107 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THAT THERE WAS AN APPARENT LEAD FRACTURE. THE LEAD WAS EXPLANTED AND REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. A LAWSUIT LATER ALLEGED THE LEAD WAS "DEFECTIVE" AND THE PATIENT RECEIVED "UNWANTED SHOCKS", IN ADDITION TO "SCARRING" HER HEART, AND HAVING ADDITIONAL "UNNECESSARY" SURGERIES. IT FURTHER ALLEGES THE PATIENT "SUFFERED SEVERE PHYSICAL AND EMOTIONAL INJURIES". |  | N | Y | LAY USER/PATIENT
2649622-2008-00811 | 2008-01-10 | 20071128 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-00812 | 2008-01-10 | 20071120 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | IT WAS REPORTED THAT THE PATIENT CAME TO THE HOSPITAL, ADVISING SHE HAD HEARD THE DEVICE'S PATIENT ALERT ALARM. WHEN THE DEVICE WAS INTERROGATED, IT WAS DETERMINED THE ALARM SOUNDED DUE TO HIGH IMPEDANACES (> 200 OHMS). THE LEAD WAS EXPLANTED AND REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-00813 | 2008-01-10 | 20071108 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THERE WAS OVERSENSING AND AN APPARENT LEAD FRACTURE. THE LEAD WAS REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONALLY, ALLEGED THAT THE PATIENT HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES, INCLUDING INAPPROPRIATE THERAPIES, SEVERE EMOTIONAL DISTRESS, AND ECONOMIC AND CONSEQUENTIAL DAMAGES DUE TO THE 'DEFECTIVE' AND FRACTURED LEAD. |  | N | Y | LAY USER/PATIENT
2649622-2008-00815 | 2008-01-10 | 20071129 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | LEAD IMPEDANCE TREND VARIATION AND SENSING OF NOISE OBSERVED ON THE EGM WERE REPORTED. THE LEAD WAS EXPLANTED AND REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. THE PATIENT'S STATUS WAS REPORTED AS "GOOD". |  | N | Y | LAY USER/PATIENT
2649622-2008-00817 | 2008-01-10 | 20071101 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00818 | 2008-01-10 | 20071113 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THE LEAD WAS EXPLANTED AND REPLACED DUE TO FRACTURE THAT HAD CAUSED INAPPROPRIATE SHOCKS. AT A PREVIOUS FOLLOW UP VISIT, THE LEAD IMPEDANCE WAS NOTED TO BE GREATER THAN 3000 OHMS. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT AND THE PATIENT WAS DOING WELL. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00819 | 2008-01-10 | 20071130 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THERE WAS HIGH IMPEDANCE AND OVERSENSING. THE LEAD WAS REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. A LAWSUIT ALLEGES THE LEAD HAS CAUSED THE PATIENT TO "SUSTAIN SEVERE PHYSICAL INJURIES, SEVERE EMOTIONAL DISTRESS" THERE WERE FURTHER ALLEGATIONS BY AN ATTORNEY INDICATED THE LEAD HAD EXHIBITED A FRACTURE AND/OR WAS EXPLANTED. ADDITIONALLY IT WAS ALLEGED THE PATIENT IS DECEASED AND "DEATH ASSOCIATED WITH EXPLANT." |  | Y | Y | LAY USER/PATIENT
2649622-2008-00809 | 2008-01-10 | 20071031 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00801 | 2008-01-10 | 20071102 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-00802 | 2008-01-10 | 20071107 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00803 | 2008-01-10 | 20071109 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00805 | 2008-01-10 | 20071025 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | NON-CAPTURE, SENSING DIFFICULTY, AND APPARENT RV PERFORATION WERE REPORTED. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. FURTHER INFORMATION FROM AN ATTORNEY ALLEGES PATIENT SUFFERED PHYSICAL AND OTHER INJURY AS A RESULT OF THE RECALLED LEAD. ATTORNEY ALSO ALLEGES THE PATIENT "HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH". |  | Y | N | LAY USER/PATIENT
2649622-2008-00806 | 2008-01-10 | 20071102 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00807 | 2008-01-10 | 20071029 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00808 | 2008-01-10 | 20071106 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-00800 | 2008-01-10 | 20071106 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00790 | 2008-01-10 | 20071114 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00791 | 2008-01-10 | 20071129 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE LEAD WAS EXPLANTED AND REPLACED DUE TO A DIMINISHING R WAVE AND THE SUBSEQUENT RISK FOR UNDERDETECTION. POSSIBLE LEAD DISLODGEMENT WAS A SUSPECTED CAUSE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | N | N | LAY USER/PATIENT
2649622-2008-00792 | 2008-01-10 | 20071119 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | PATIENT REPORTS DEVICE SHOCKED 10-12 TIMES AND WAS TRANSPORTED TO HOSPITAL VIA AMBULANCE WITH TWO MORE THERAPIES BEING DELIVERED. AT HOSPITAL, MAGNET WAS THEN PLACED OVER THE DEVICE. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION ON THE EVENT INDICATES RV PACING IMPEDANCE WAS > 3000 OHMS AND OVERSENSING WAS NOTED ON DEVICE INTERROGATION. FURTHER INFORMATION WAS RECEIVED IN A LAWSUIT ALLEGING THAT THE LEAD WAS FRACTURED. |  | Y | N | LAY USER/PATIENT
2649622-2008-00794 | 2008-01-10 | 20070918 | Injury | Manufacturer report | SPRINT FIDELIS | DEFIBRILLATOR, AUTOMATIC IMPLANTABLE CARDIOVERTER | MPRI | 6949 | 6949 |  | IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00795 | 2008-01-10 | 20071105 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00796 | 2008-01-10 | 20070918 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00797 | 2008-01-10 | 20071108 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00798 | 2008-01-10 | 20071112 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00799 | 2008-01-10 | 20071106 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00778 | 2008-01-10 |  | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00781 | 2008-01-10 | 20071127 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00785 | 2008-01-10 | 20071129 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: THE ACTUAL DEVICE WAS NOT RECEIVED FOR EVALUATION. PERFORMANCE DATA COLLECTED FROM THE ICD DEVICE INDICATES NOISE AND HIGH IMPEDANCE. |  | N | Y | LAY USER/PATIENT
2649622-2008-00786 | 2008-01-10 | 20070103 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  | IT WAS REPORTED THAT THE PATIENT HAD BEEN FEELING FATIGUED WITH SOME POUNDING IN HIS CHEST. THE PATIENT RECEIVED 10 TO 15 SHOCKS, AND WAS TAKEN TO THE HOSPITAL. IT WAS REPORTED THAT THE LEAD WAS FRACTURED. THE PACE/SENSE PORTION WAS CAPPED AND REPLACED. IT WAS ALSO REPORTED THAT THE ATRIAL PORT WAS UNABLE TO BE TIGHTENED, AND THAT THERE WAS NOISE AND OVERSENSING. THE DEVICE WAS REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. THE PATIENT FURTHER REPORTED THAT THE LEAD FAILED AND THAT THE PATIENT WAS SHOCKED 72 TIMES. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00788 | 2008-01-10 | 20070703 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00777 | 2008-01-10 | 20071113 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | PRODUCT PERFORMANCE ISSUE REPORTED. THE LEAD WAS EXPLANTED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION REPORTS THAT THE LEAD WAS EXPLANTED DUE TO ADVISORY/RECALL/UPGRADE. |  | N | Y | LAY USER/PATIENT
2649622-2008-00768 | 2008-01-10 | 20071120 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE LEAD WAS EXPLANTED AND REPLACED DUE TO HIGH HV AND SVC IMPEDANCE OF GREATER THAN 200 OHMS, AND AN APPARENT LEAD FRACTURE. ADDITIONAL INFORMATION WAS PROVIDED. ALLEGED THAT THE PATIENT HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES, INCLUDING INAPPROPRIATE OR INADEQUATE THERAPIES, SEVERE EMOTIONAL DISTRESS, AND ECONOMIC AND CONSEQUENTIAL DAMAGES DUE TO THE 'DEFECTIVE' AND FRACTURED LEAD. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00769 | 2008-01-10 | 20071126 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT RECEIVED INAPPROPRIATE THERAPY, AND THERE WAS A SUDDEN INCREASE IN IMPEDANCE FROM 500 OHMS TO 1100 OHMS. THE LEAD WAS REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED REPORTING THAT THERE WAS ALSO AN APPARENT LEAD FRACTURE, AND THE LEAD WAS PARTIALLY REMOVED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00773 | 2008-01-10 | 20071105 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00774 | 2008-01-10 | 20071105 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE RIGHT VENTRICULAR LEAD WAS REPLACED DUE TO INCREASED IMPEDANCE AND AN APPARENT LEAD FRACTURE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION RECEIVED REPORTED AN ALLEGATION FROM AN ATTORNEY INDICATED THE PATIENT IS DECEASED AND FURTHER INDICATED THAT THE LEAD HAD EXHIBITED A FRACTURE AND/OR WAS EXPLANTED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00767 | 2008-01-10 | 20071120 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. WE DID RECEIVE PERFORMANCE DATA COLLECTED FROM THE DEVICE AND HAVE ANALYZED THE DATA. THE LIFETIME VENTRICULAR SENSING INTEGRITY COUNTER IS HIGH, INDICATING A POSSIBLE INTERMITTENCY IN THE SYSTEM. HIGH IMPEDANCE VALUES WERE ALSO NOTED. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD IN SEGMENTS RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-00756 | 2008-01-10 | 20071121 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00757 | 2008-01-10 | 20071121 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: THE ACTUAL DEVICE WAS NOT RECEIVED FOR EVALUATION. WE DID RECEIVE PERFORMANCE DATA COLLECTED FROM THE DEVICE. ANALYSIS OF THAT DATA REVEALED NOISE, HIGH VENTRICULAR IMPEDANCE RANGING FROM 1328-1648 OHMS, AND AN ELEVATED SENSING INTEGRITY COUNTER. A HIGH VALUE OF THIS COUNTER CAN INDICATE A POSSIBLE INTERMITTENCY IN THE SYSTEM. |  | N | Y | LAY USER/PATIENT
2649622-2008-00758 | 2008-01-10 | 20071127 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00759 | 2008-01-10 | 20071123 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE RV LEAD WAS REPLACED DUE TO MULTIPLE INAPPROPRIATE SHOCKS CAUSED BY OVERSENSING. A PATIENT ALERT WAS ALSO NOTED FOR HIGH LEAD IMPEDANCE. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. FURTHER INFORMATION INDICATES THE CHRONIC RV LEAD WAS EXPLANTED DUE TO INAPPROPRIATE SHOCKS CAUSED BY OVERSENSING AND A POSSIBLE LEAD FRACTURE. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00761 | 2008-01-10 | 20071126 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00762 | 2008-01-10 | 20071128 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE RV LEAD WAS REPLACED DUE TO HIGH IMPEDANCE OF GREATER THAN 2000 OHMS. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. FURTHER INFORMATION FROM AN ATTORNEY ALLEGES THE PLAINTIFF SUFFERED PHYSICAL AND OTHER INJURY AS A RESULT OF THE RECALLED LEAD INCLUDING INAPPROPRIATE SHOCKS, FRACTURE, SEVERE EMOTIONAL DISTRESS, AND MENTAL ANGUISH. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-00763 | 2008-01-10 | 20071127 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00765 | 2008-01-10 | 20071015 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | Y | N | LAY USER/PATIENT
2649622-2008-00754 | 2008-01-10 | 20071018 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00744 | 2008-01-10 | 20071115 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2182208-2008-00031 | 2008-01-10 | 20071120 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC, INC. | 6948 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00745 | 2008-01-10 | 20070219 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00746 | 2008-01-10 | 20071119 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT RECEIVED 26 INAPPROPRIATE SHOCKS, THAT THE IMPEDANCE WAS VARYING BETWEEN 1300 OHMS AND 564 OHMS, AND THAT THE DEVICE DATA INDICATED OVERSENSING. THE PACE/SENSE PORTION OF THE LEAD WAS CAPPED, AND A NEW PACE/SENSE LEAD WAS IMPLANTED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ALLEGED THAT THE PATIENT "SUFFERED PHYSICAL AND OTHER INJURY", THAT THERE WAS A LEAD FRACTURE, THAT THE PATIENT "EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING OR FAILURE TO RECEIVE THERAPEUTIC SHOCKS, REQUIRING [PATIENT] TO SEEK MEDICAL INTERVENTION RESULTING IN REPLACEMENT OF THE DEFECTIVE [LEAD]", AND THAT THE PATIENT "SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH, ECONOMIC LOSSES AND OTHER DAMAGES. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00748 | 2008-01-10 | 20071120 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: DISTAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-00749 | 2008-01-10 | 20071119 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | Y | N | LAY USER/PATIENT
2649622-2008-00750 | 2008-01-10 | 20071116 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY THE ACTUAL DEVICE WAS NOT RECEIVED FOR EVALUATION. WE DID RECEIVE PERFORMANCE DATA COLLECTED FROM THE DEVICE AND HAVE ANALYZED THE DATA. THE VENTRICULAR SENSING INTEGRITY COUNTER IS HIGH, INDICATING A POSSIBLE INTERMITTENCY IN THE SYSTEM. HIGH IMPEDANCE MEASUREMENTS WERE ALSO NOTED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00751 | 2008-01-10 | 20071116 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY THE ACTUAL DEVICE WAS NOT RECEIVED FOR EVALUATION. WE DID RECEIVE PERFORMANCE DATA COLLECTED FROM THE DEVICE AND HAVE ANALYZED THE DATA, WHICH SHOWS THE SENSING MEASUREMENTS WERE ALL WITHIN ACCEPTABLE RANGE OVER THE ENTIRE TIME THE DEVICE WAS IMPLANTED. NO OVERSENSING WAS SEEN. THERE WERE VARYING HIGH IMPEDANCE MEASUREMENTS NOTED. |  | N | Y | LAY USER/PATIENT
2649622-2008-00752 | 2008-01-10 | 20071126 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00743 | 2008-01-10 | 20071120 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00732 | 2008-01-10 | 20071116 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00733 | 2008-01-10 | 20071007 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00734 | 2008-01-10 | 20071117 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | PATIENT STATES 'SHOCKED SEVERAL TIMES AND LEAD IS FRACTURED. FOLLOWUP WITH THE HCP INDICATES OVERSENSING AND HIGH IMPEDANCES WERE NOTED AS WELL AS INAPPROPRIATE SHOCKS. LEAD FRACTURE WAS SUSPECTED. THE PACE/SENSE PORTION OF THE LEAD WAS CAPPED, AND A NEW PACE/SENSE LEAD WAS IMPLANTED. A LAWSUIT FURTHER ALLEGED THAT THE PATIENT "SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR INCREASED RISK OF DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH, ECONOMIC LOSSES AND OTHER DAMAGES". NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00736 | 2008-01-10 | 20070715 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00737 | 2008-01-10 | 20071114 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT RECEIVED 3 SHOCKS DUE TO NOISE. THE LEAD WAS REPLACED. IT WAS FURTHER REPORTED THAT THERE WAS SENSING DIFFICULTY. THE DEVICE WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00738 | 2008-01-10 | 20070804 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00739 | 2008-01-10 | 20071130 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE RV LEAD WAS EXPLANTED DUE TO MULTIPLE INAPPROPRIATE SHOCKS CAUSED BY OVERSENSING, AND AN APPARENT LEAD FRACTURE. A VARYING IMPEDANCE TREND WAS ALSO NOTED. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. LEAD WAS RETURNED WITH ADDITIONAL INFORMATION NOTING HIGH IMPEDANCE. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00740 | 2008-01-10 | 20071108 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00741 | 2008-01-10 | 20070908 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-00721 | 2008-01-10 | 20070518 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE RV LEAD WAS EXPLANTED DUE TO OVERSENSING. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ATTORNEY ALLEGES PATIENT EXPERIENCED INJURY, INCLUDING INAPPROPRIATE SHOCKS, AND FRACTURE RELATED TO THE DEFECTIVE LEAD. ALLEGATION FURTHER INDICATES THE PATIENT HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, AND MENTAL ANGUISH. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-00730 | 2008-01-10 | 20071126 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00709 | 2008-01-10 | 20071113 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS EXPLANTED AND REPLACED DUE TO OVERSENSING THAT LED TO 2 INAPPROPRIATE SHOCKS. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION REPORTS THE PHYSICIAN FELT THE OVERSENSING MAY HAVE BEEN RELATED TO CROSSTALK. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00710 | 2008-01-10 | 20071113 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00711 | 2008-01-10 | 20071113 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS EXPLANTED AND REPLACED BECAUSE OF INAPPROPRIATE SHOCKS DUE TO OVERSENSING AND A HIGH IMPEDANCE OF 2500 OHMS. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00713 | 2008-01-10 | 20071113 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | Y | N | LAY USER/PATIENT
2649622-2008-00714 | 2008-01-10 | 20071112 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THAT THE LEAD WOULD BE REPLACED DUE TO FRACTURE, AND THAT THE LOWER PATIENT ALERT LIMITS TRIGGERED THE ALARM. THE LEAD WAS REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED FROM THE PATIENT'S WIFE REPORTED THAT "IT STARTED BEEPING AND LEAD WAS REPLACED RIGHT AWAY, HE DIDN'T HAVE ANY SHOCKS." IT WAS ALSO REPORTED THAT THERE WAS SENSING DIFFICULTY, OVERSENSING, NOISE, HIGH IMPEDANCE, AND A LEAD FAILURE. THE LEAD WAS CAPPED AND REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-00717 | 2008-01-10 | 20071120 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-00718 | 2008-01-10 | 20071127 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00701 | 2008-01-10 | 20070912 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00702 | 2008-01-10 | 20071108 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00704 | 2008-01-10 | 20071114 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00705 | 2008-01-10 | 20071112 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT WAS TAKEN BY AMBULANCE TO THE EMERGENCY ROOM AFTER EXPERIENCING MULTIPLE INAPPROPRIATE SHOCKS DUE TO A LEAD FRACTURE. THE PATIENT REQUIRED EMERGENCY SURGERY TO REPLACE THE LEAD BECAUSE HIS HEART RATE WAS DANGEROUSLY LOW. IT WAS ALSO REPORTED THAT BECAUSE OF THE SHOCKS THE PATIENT RECEIVED, HIS HEART IS NOW WEAKER THAN WHEN THE DEVICE WAS FIRST IMPLANTED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONALLY, ALLEGED THAT THE PATIENT HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES, INCLUDING INAPPROPRIATE THERAPIES, SEVERE EMOTIONAL DISTRESS, AND ECONOMIC AND CONSEQUENTIAL DAMAGES DUE TO THE 'DEFECTIVE' AND FRACTURED LEAD. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00706 | 2008-01-10 | 20071127 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00708 | 2008-01-10 | 20071113 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00698 | 2008-01-10 | 20051221 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00691 | 2008-01-10 | 20071024 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE IMPEDANCE WAS > 3000 OHMS AND THERE WAS NOISE ON THE LEAD RESULTING IN 6 INAPPROPRIATE SHOCKS. THE LEAD WAS CAPPED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION REPORTED BY PATIENT'S FAMILY MEMBER REPORTED THAT THE LEADS WERE DEFECTIVE SOME TIME AGO AND THAT IN 2007 THEY SHORTED OUT AND THE PATIENT RECEIVED 3 SHOCKS. ADDITIONAL INFORMATION REPORTED OVERSENSING. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00695 | 2008-01-10 |  | No answer provided | Manufacturer report | SPRINT FIDELIS | ASKU | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE PATIENT RECEIVED INAPPROPRIATE SHOCKS DUE TO OVERSENSING. THE PACE/SENSE PORTION OF THE LEAD WAS CAPPED, AND THE HIGH VOLTAGE PORTION REMAINS IN USE. A PREVIOUSLY ABANDONED LEAD WAS REUSED FOR PACING/SENSING. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. LAWSUIT ALLEGES PATIENT SUFFERED PHYSICAL AND OTHER INJURY AS A RESULT OF THE RECALLED LEAD. PATIENT EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY REQUIRING MEDICAL INTERVENTION INCLUDING, BUT NOT LIMITED TO SURGERY, REMOVAL AND/OR REPLACEMENT OF THE DEFECTIVE LEAD. PATIENT SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH. PATIENT HAS SUFFERED PHYSICAL INJURIES AND VARIOUS PHYSICAL MANIFESTATIONS OF EMOTIONAL DISTRESS ASSOCIATED WITH ONE OR MORE OF THE FOLLOWING: THE IMPLANTATION, RECALL, FAILURE, REMOVAL/REPLACEMENT, AND/OR INABILITY TO HAVE THE DEFECTIVE LEAD REMOVED OR REPLACED. |  | Y | Y | 0
2649622-2008-00697 | 2008-01-10 | 20071116 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-00690 | 2008-01-10 | 20071129 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00679 | 2008-01-10 | 20071115 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00681 | 2008-01-10 | 20070731 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00682 | 2008-01-10 | 20071024 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00685 | 2008-01-10 | 20070604 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00686 | 2008-01-10 | 20070705 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THERE WAS A 'PROBLEM' WITH THE LEAD. THE LEAD WAS REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. IT WAS FURTHER REPORTED THAT THE PATIENT RECEIVED INAPPROPRIATE SHOCKS. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00688 | 2008-01-10 | 20071115 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00689 | 2008-01-10 | 20071112 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS EXPLANTED AND REPLACED DUE TO ACUTE INCREASE IN IMPEDANCE AND SUSPECTED LEAD FRACTURE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-00669 | 2008-01-10 | 20060308 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00670 | 2008-01-10 | 20071116 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT ALERT SOUNDED AND THE PATIENT WENT TO HIS MD. THE LEAD WAS 'STARTING TO FRACTURE', AND WAS REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS PROVIDED BY THE PATIENT'S ATTORNEY. ALLEGED THAT THE PATIENT HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES, INCLUDING INAPPROPRIATE THERAPIES, SEVERE EMOTIONAL DISTRESS, AND ECONOMIC AND CONSEQUENTIAL DAMAGES DUE TO THE 'DEFECTIVE' AND FRACTURED LEAD. |  | N | Y | LAY USER/PATIENT
2649622-2008-00672 | 2008-01-10 | 20071101 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00667 | 2008-01-10 | 20071101 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00657 | 2008-01-10 | 20071112 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PACE/SENSE PORTION OF THE LEAD WAS CAPPED BECAUSE OF LEAD FRACTURE. A NEW PACE/SENSE LEAD WAS IMPLANTED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION RECEIVED REPORTED THERE WAS ALSO OVERSENSING AND VARIABLE IMPEDANCES NOTED ON THE RV LEAD. |  | N | N | LAY USER/PATIENT
2649622-2008-00659 | 2008-01-10 | 20071107 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00660 | 2008-01-10 | 20071107 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-00662 | 2008-01-10 | 20071028 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00663 | 2008-01-10 | 20071023 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | Y | N | LAY USER/PATIENT
2649622-2008-00664 | 2008-01-10 | 20071025 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00665 | 2008-01-10 | 20071107 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. EVALUATION SUMMARY: DEFIB CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | N | N | LAY USER/PATIENT
2649622-2008-00666 | 2008-01-10 | 20071031 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THAT THE PATIENT RECEIVED INAPPROPRIATE THERAPY, AND LEAD IMPEDANCE HAD INCREASED. IT WAS FURTHER REPORTED THAT THERE WAS OVERSENSING AND AN APPARENT LEAD FRACTURE. THE LEAD WAS EXPLANTED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00646 | 2008-01-10 | 20071029 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00647 | 2008-01-10 | 20071030 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6930 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00648 | 2008-01-10 | 20071101 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS EXPLANTED AND REPLACED WITH A COMPETITOR LEAD DUE TO OVERSENSING, HIGH IMPEDANCE, HIGH THRESHOLD, AND FRACTURE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-00650 | 2008-01-10 | 20070120 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT WENT TO THE ER AFTER 6 SHOCKS OCCURRED DUE TO OVERSENSING. THE LEAD WAS LATER REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | N | LAY USER/PATIENT
2649622-2008-00654 | 2008-01-10 | 20050331 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00655 | 2008-01-10 |  | No answer provided | Manufacturer report | SPRINT FIDELIS | ASKU | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | THE PATIENT'S COUSIN REPORTED THAT THE PATIENT WAS ADMITTED TO THE ER DUE TO FOUR CARDIAC ARRESTS IN THE DRIVEWAY. PATIENT'S COUSIN ALLEGED THAT THE DEVICE WAS DEFECTIVE AND THAT IT WAS NOT FUNCTIONING THE WAY IT WAS SUPPOSED TO. THE PACE/SENSE PORTION OF THE HIGH VOLTAGE LEAD WAS REPLACED WITH A NEW PACE/SENSE LEAD. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. LAWSUIT ALLEGES THE PATIENT SUFFERED PHYSICAL AND OTHER INJURY AS A RESULT OF THE RECALLED LEAD. PATIENT EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY REQUIRING MEDICAL INTERVENTION INCLUDING BUT NOT LIMITED TO SURGERY, REMOVAL AND/OR REPLACEMENT OF THE DEFECTIVE SPRINT FIDELIS LEAD. PATIENT HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH. |  | Y | Y | 0
2649622-2008-00634 | 2008-01-10 | 20071109 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PACE/SENSE PORTION OF THE RV LEAD WAS CAPPED AND REPLACED WITH A NEW PACE/SENSE LEAD DUE TO FLUCTUATING HIGH LEAD IMPEDANCE FROM 400 TO 1700 OHMS. THE HIGH VOLTAGE PORTION OF THE LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. FURTHER INFORMATION FROM AN ATTORNEY ALLEGES THE PLAINTIFF SUFFERED PHYSICAL AND OTHER INJURY AS A RESULT OF THE RECALLED LEAD INCLUDING LOSS OF CONSORTIUM, SEVERE EMOTIONAL STRESS, AND MENTAL ANGUISH. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-00636 | 2008-01-10 | 20071108 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00637 | 2008-01-10 | 20071107 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE RV LEAD WAS EXPLANTED DUE TO T-WAVE OVERSENSING AND NOISE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION RECEIVED FROM THE PATIENT'S ATTORNEY REPORTED THE PATIENT SUSTAINED INJURY FROM THE DELIVERY OF INAPPROPRIATE SHOCKS. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00638 | 2008-01-10 | 20071105 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | N | N | LAY USER/PATIENT
2649622-2008-00639 | 2008-01-10 | 20071108 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT WAS BROUGHT TO THE ER FOLLOWING MULTIPLE INAPPROPRIATE SHOCKS. DEEP BREATHING AND COUGHING ELICITED NOISE AS WELL AS ONE HIGH IMPEDANCE MEASUREMENT OF GREATER THAN 2500 OHMS. THE RV LEAD WAS REPLACED AND NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. FURTHER INFORMATION RECIEVED STATES OVERSENSING AND HIGH IMPEDANCE WERE OBSERVED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00640 | 2008-01-10 | 20071108 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES FOUND; FULL LEAD IN SEGMENTS RETURNED AND ANALYZED. OTHER: IT WAS REPORTED THAT THE LEAD IMPEDANCE INCREASED TO 2768 OHMS. THE LEAD WAS REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION RECEIVED WITH THE DEVICE REPORTED THAT THERE WAS A LEAD FRACTURE. ACTUAL DEVICE INVOLVED IN INCIDENT WAS EVALUATED. ELECTRICAL TESTS PERFORMED; MECHANICAL TESTS PERFORMED. VISUAL EXAMINATION: DEVICE PERFORMED ACCORDING TO SPECIFICATIONS. DEVICE EVALUATED AND ALLEGED FAILURE COULD NOT BE DUPLICATED, IMPEDANCE, HIGH, LEAD(S), FRACTURE OF. |  | N | Y | LAY USER/PATIENT
2649622-2008-00641 | 2008-01-10 | 20071105 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | N | LAY USER/PATIENT
2649622-2008-00642 | 2008-01-10 | 20071026 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00643 | 2008-01-10 | 20071106 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-00644 | 2008-01-10 | 20071115 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY AND TESTED OUT OF SPECIFICATION CONSISTENT WITH THE ADVISORY. EVALUATION SUMMARY: DISTAL CONDUCTOR WAS FRACTURED; FULL LEAD IN SEGMENTS WAS RETURNED FOR ANALYSIS. |  | Y | N | LAY USER/PATIENT
2649622-2008-00623 | 2008-01-10 | 20071105 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT, WHO HAD A HISTORY OF CHF, WAS FOUND DEAD. THE EXACT CAUSE OF DEATH WAS NOTED TO BE UNDETERMINED, BUT CARDIAC. THE PATIENT HAD A CARELINK TRANSMISSION AFTER THE FIRST EPISODE OF DEVICE-DEFINED VF (3 SHOCKS). UNABLE TO DETERMINE IF TRUE VF OR IF RELATED TO MODEL 6949 LEAD FAILURE. ADDITIONAL INFORMATION WAS RECEIVED REPORTING ABNORMAL EGMS; QUESTION IF TRUE VF OR NOT. ATRIAL EGM WAS NON-EXISTENT, WITH MINIMAL SENSING; PATIENT HAD A HISTORY OF AF. REVIEW OF SAVE-TO-DISK DATA NOTED THERE WERE NO ATRIAL EGM DEFLECTIONS AND AN IMPEDANCE INCREASE ON THE RV (RIGHT VENTRICULAR) AND SVC (SUPERIOR VENA CAVA) COILS, AS WELL AS AN INCREASE IN PACING LEAD IMPEDANCES FOR ATRIAL, RV, AND LV (LEFT VENTRICULAR) ON THE REPORTED DATE OF DEATH. THE PATIENT ALSO HAD MULTIPLE VF EPISODES ON THAT DAY. VF EPISODES SHOW SHORT INTERVALS, WITH A CORRELATING FIRST SHORT INTERVAL COUNT DATE, AT THE TIME OF THE VF EPISODES. LEAD IMPEDANCES WERE NOTED TO HAVE BEEN NORMAL UNTIL THE DAY BEFORE THE PATIENT'S DEATH. INFORMATION CONTINUED IN NARRATIVE. IT WAS REPORTED THAT THE PT, WHO HAD A HISTORY OF CHF, WAS FOUND DEAD. THE EXACT CAUSE OF DEATH WAS NOTED TO BE UNDETERMINED, BUT CARDIAC. THE PT HAD A CARELINK TRANSMISSION AFTER THE FIRST EPISODE OF DEVICE-DEFINED VF (3 SHOCKS). UNABLE TO DETERMINE IF TRUE VF OR IF RELATED TO MODEL 6949 LEAD FAILURE. ADDT'L INFO WAS RECEIVED THROUGH TECH SERV REPORTING ABNORMAL EGMS. ? IF TRUE VF OR NOT. ATRIAL EGM WAS NON-EXISTENT WITH MINIMAL SENSING; PT HAD A HISTORY OF AF. REVIEW OF SAVE-TO-DISK DATA NOTED THERE WERE NO ATRIAL EGM DEFLECTIONS AND AN IMPEDANCE INCREASE ON THE RV AND SVC COILS, AS WELL AS AN INCREASE IN PACING LEAD IMPEDANCES FOR ATRIAL, RV, AND LV, ON THE REPORTED DATE OF DEATH. THE PT ALSO HAD MULTIPLE VF EPISODES ON THAT DAY. VF EPISODES SHOW SHORT INTERVALS, WITH A CORRELATING FIRST SHORT INTERVAL COUNT DATE, AT THE TIME OF VF EPISODES. LEAD IMPEDANCES WERE NOTED TO HAVE BEEN NORMAL UNTIL THE DAY BEFORE THE PT'S DEATH |  | Y | Y | LAY USER/PATIENT
2649622-2008-00625 | 2008-01-10 | 20071112 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS EXPLANTED AND REPLACED DUE TO NO CAPTURE, HIGH IMPEDANCE AND AN APPARENT LEAD FRACTURE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONALLY, A LAWSUIT ALLEGES THE PT "SUFFERED PHYSICAL AND OTHER INJURY" DUE TO THE LEAD. IT ALSO STATES THE PT "EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY REQUIRING MEDICAL INTERVENTION". |  | N | Y | LAY USER/PATIENT
2649622-2008-00626 | 2008-01-10 | 20071113 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. WE DID RECEIVE PERFORMANCE DATA COLLECTED FROM THE DEVICE AND HAVE ANALYZED THE DATA, WHICH SHOWED VARYING IMPEDANCE VALUES. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. IMPEDANCE INCREASE WAS REPORTED. THE LEAD WAS EXPLANTED AND REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ACTUAL DEVICE INVOLVED IN INCIDENT WAS EVALUATED ELECTRICAL TESTS PERFORMED, MECHANICAL TESTS PERFORMED, VISUAL EXAMINATION LEAD CONDUCTOR COMPONENT/SUBASSEMBLY FAILURE DEVICE WAS OUT OF SPECIFICATION IN A MANNER THAT RELATES TO EVENT IMPEDANCE, HIGH. |  | N | Y | LAY USER/PATIENT
2649622-2008-00627 | 2008-01-10 | 20071108 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED |  | Y | Y | LAY USER/PATIENT
2649622-2008-00628 | 2008-01-10 | 20071112 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT WAS SHOCKED MULTIPLE TIMES DUE TO OVERSENSING AND LEAD FRACTURE. THE LEAD WAS REPLACED. IT WAS FURTHER REPORTED BY THE PATIENT THAT HER EJECTION FRACTION DROPPED SIGNIFICANTLY AFTER ALL THE SHOCKS FROM THE BROKEN LEAD. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00629 | 2008-01-10 | 20071018 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | NON SUSTAINED EPISODES WERE REPORTED. POSSIBLE LEAD FRACTURE. THE LEAD WAS REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION REPORTED THAT THE PATIENT HAD SUDDEN ONSET OF 4 ICD SHOCKS AT HOME AND 3 MORE SHOCKS IN THE ER. IT WAS FOUND IN THE OR THAT CLAVICULAR CRUSH MAY HAVE BEEN THE REASON FOR THE RV LEAD FRACTURE. THE RV LEAD WAS CAPPED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00631 | 2008-01-10 | 20071026 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | Y | N | LAY USER/PATIENT
2649622-2008-00632 | 2008-01-10 | 20071106 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-00633 | 2008-01-10 | 20071107 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY - DEFIB CONDUCTOR FRACTURED. FULL LEAD RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-00622 | 2008-01-10 | 20071005 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00610 | 2008-01-10 | 20071030 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. IT WAS REPORTED THAT THE PATIENT RECEIVED INAPPROPRIATE THERAPY DUE TO OVERSENSING, AND THERE WAS A SUSPECTED LEAD FRACTURE. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ACTUAL DEVICE INVOLVED IN INCIDENT WAS EVALUATED ELECTRICAL TESTS PERFORMED, MECHANICAL TESTS PERFORMED, VISUAL EXAMINATION LEAD CONDUCTOR COMPONENT/SUBASSEMBLY FAILURE DEVICE WAS OUT OF SPECIFICATION IN A MANNER THAT RELATES TO EVENT LEAD(S), FRACTURE OF OVERSENSING SHOCK, INAPPROPRIATE. |  | Y | N | LAY USER/PATIENT
2649622-2008-00611 | 2008-01-10 | 20071030 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00612 | 2008-01-10 | 20071023 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE, DUE TO CONFIDENTIALITY CONCERNS. EVALUATION SUMMARY - DISTAL CONDUCTOR FRACTURED. FULL LEAD RETURNED AND ANALYZED. ADDITIONAL ANALYST COMMENT - LEAD APPEARS TO HAVE BEEN IMPLANTED WITH A BEND IN IT AT THE FRACTURE SITE. |  | Y | N | LAY USER/PATIENT
2649622-2008-00613 | 2008-01-10 | 20071107 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE RIGHT VENTRICULAR LEAD IMPEDANCE WAS FLUCTUATING BETWEEN 600 OHMS AND 1300 OHMS. THE LEAD WAS EXPLANTED AND REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-00614 | 2008-01-10 | 20071030 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00615 | 2008-01-10 | 20071106 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00616 | 2008-01-10 | 20070524 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00617 | 2008-01-10 | 20071106 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS EXPLANTED AND REPLACED DUE TO INAPPROPRIATE SHOCKS. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION RECEIVED REPORTED THE IMPEDANCE HAD BEEN GREATER THAN 3000 OHMS AND A LEAD FRACTURE WAS SUSPECTED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00618 | 2008-01-10 | 20071014 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00619 | 2008-01-10 | 20071030 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00620 | 2008-01-10 | 20071030 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00621 | 2008-01-10 | 20071029 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00609 | 2008-01-10 | 20071029 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00598 | 2008-01-10 | 20071106 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. DETAILED ANALYSIS OF THE DEVICE WAS NOT PERFORMED AT THIS TIME DUE TO PENDING LITIGATION. THEREFORE, WE ARE UNABLE TO FULLY EVALUATE THE REPORTED EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00599 | 2008-01-10 | 20071106 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE LEAD WAS EXPLANTED AND REPLACED DUE TO OVERSENSING AND HIGH IMPEDANCE OF 2800 OHMS. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-00603 | 2008-01-10 | 20071005 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT SUFFERED A CARDIAC ARREST, AND THAT THE DEVICE (A CLINICAL DEVICE) HAD FAILED TO SHOCK THE PATIENT'S VENTRICULAR FIBRILLATION. AFTER INITIAL INTERROGATION DURING THE ARREST, THE SENSITIVITY WAS INCREASED, AND THE PATIENT RECEIVED SIX FULL ENERGY SHOCKS FROM THE DEVICE WHICH WERE UNSUCCESSFUL. EARLIER THAT SAME DAY, THE DEVICE DELIVERED SEVERAL SUCCESSFUL SHOCKS, AND IT APPEARED THAT ATP MAY HAVE ACCELERATED THE PATIENT'S RHYTHM TO VF. DEVICE INTERROGATION DURING THE ARREST WITH CPR IN PROGRESS SHOWED EVIDENCE OF INTERMITTENT T-WAVE OVERSENSING, NOISE, AND UNDERSENSING. THE SHORT INTERVAL COUNT WAS 402, AND HAD INCREASED STARTING AROUND THE TIME OF THE FIRST SHOCK. SAVE TO DISK REVIEW SHOWED INITIAL APPROPRIATE DETECTION OF VT AND APPROPRIATE THERAPY, BUT LATER EPISODES DURING CPR SHOW OVERSENSING. IT WAS QUESTIONED WHETHER THE LEAD MAY HAVE FRACTURED DURING CPR. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00604 | 2008-01-10 | 20071102 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00605 | 2008-01-10 | 20070824 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00596 | 2008-01-10 | 20071109 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00584 | 2008-01-10 | 20071019 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00586 | 2008-01-10 | 20070403 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00587 | 2008-01-10 | 20070118 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00588 | 2008-01-10 | 20050526 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | PATIENT'S WIFE ALLEGES 'LEAD BROKE'. ALSO STATES PT CANNOT LIFT ARM AND CANNOT GO BACK TO WORK'. THE LEAD WAS REPLACED. ADDITIONAL INFORMATION RECEIVED FROM ATTORNEY ALLEGES THE ATRIAL LEAD HAD DISLODGED (IN 2005) AND IT WAS AT THAT TIME THAT THE DR DISCOVERED THE SPRINT FIDELIS LEAD WAS FRACTURED. ATTORNEY FURTHER ALLEGES ' [PATIENT] HAS SUFFERED SEVERE PHYSICAL AND EMOTIONAL INJURIES AND ECONOMIC LOSS IN THE PAST FOR MEDICAL CARE NECESSITATED BY THE UNWARRANTED SHOCKS, AND THE ADDITIONAL SURGERIES'. |  | Y | N | LAY USER/PATIENT
2649622-2008-00589 | 2008-01-10 | 20070329 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | PATIENT REPORTS LEAD WAS REPLACED BECAUSE 'FAULTY'. F/U WITH CLINIC INDICATES PT HAD PHANTOM ICD FIRINGS PRESENTING WITH ATYPICAL CHEST PAIN. PT'S CHART STATES DEVICE FAILED TO CONVERT VF X 3 AND PT RESCUED EXTERNALLY IN CATH LAB. THE HIGH VOLTAGE COIL WAS CAPPED AND A NEW DEFIB LEAD IMPLANTED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION RECEIVED REPORTED AN ALLEGATION FROM AN ATTORNEY INDICATED THE PATIENT IS DECEASED AND FURTHER INDICATED THAT THE LEAD HAD EXHIBITED A FRACTURE AND/OR WAS EXPLANTED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00590 | 2008-01-10 | 20071106 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00594 | 2008-01-10 | 20071116 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE LEAD WAS CAPPED DUE TO HIGH SVC IMPEDANCE OF GREATER THAN 200 OHMS. REPEATED MEASUREMENTS BROUGHT THE IMPEDANCE DOWN TO 70 TO 80 OHMS. IT WAS NOTED THE PATIENT HAS A HISTORY OF FALLING. FURTHER INFORMATION RECEIVED INDICATES THAT THE LEAD FAILURE WAS DUE TO 1ST RIB/CLAVICLE CRUSH. WORN INSULATION WAS ALSO REPORTED WITH THE LEAD RETURN. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-00595 | 2008-01-10 | 20070708 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00574 | 2008-01-10 | 20071019 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE RV LEAD WAS REPLACED DUE TO MULTIPLE INAPPROPRIATE SHOCKS, OVERSENSING, AND A LEAD IMPEDANCE SPIKE. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. FURTHER INFORMATION PROVIDED INDICATES NOISE ON LEAD. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00575 | 2008-01-10 | 20071105 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00576 | 2008-01-10 | 20071105 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-00577 | 2008-01-10 | 20071101 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  | THIS MEDICAL DEVICE REPORT IS BEING RESUBMITTED TO THE EMDR SYSTEM DUE TO A CONNECTION ERROR WITHIN THE EMDR SYSTEM. THE EVENT WAS PREVIOUSLY REPORTED TO THE EMDR PORTAL AND AN ACKNOWLEDGEMENT 3 RECEIVED, BUT THE CONNECTION ERROR PREVENTED THE EMDR SYSTEM FROM DOCUMENTING THE REPORT. THE REPORT NUMBER IN THIS REPORT IS THE SAME NUMBER AS THE IMPACTED REPORT THAT HAS THE CONNECTION ISSUE. PRODUCT EVENT SUMMARY: THE PROXIMAL SEGMENT OF THE LEAD WAS RETURNED AND ANALYZED. NO ANOMALIES WERE FOUND. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00578 | 2008-01-10 | 20071108 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT RECEIVED INAPPROPRIATE THERAPY DUE TO OVERSENSING OF NOISE. THE LEAD IS TO BE REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED REPORTING THAT THE LEAD WAS FRACTURED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00580 | 2008-01-10 | 20071105 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | IT WAS REPORTED THAT THE SVC IMPEDANCE INCREASED FROM 62 OHMS TO 160 OHMS, THE HVB IMPEDANCE DECREASED FROM 532 OHMS TO 144 OHMS, AND A COIL FRACTURE WAS SUSPECTED. THE RIGHT VENTRICULAR LEAD WAS REPLACED. ADDITIONAL INFORMATION WAS PROVIDED BY THE PATIENT'S ATTORNEY. ALLEGED THAT THE PATIENT HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES, INCLUDING INAPPROPRIATE THERAPIES, SEVERE EMOTIONAL DISTRESS, AND ECONOMIC AND CONSEQUENTIAL DAMAGES DUE TO THE 'DEFECTIVE' AND FRACTURED LEAD. IN ADDITION, INDICATES THE "ENHANCED MONITORING SYSTEM" FAILED. |  | N | Y | LAY USER/PATIENT
2649622-2008-00581 | 2008-01-10 | 20070501 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00582 | 2008-01-10 | 20071101 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT EXPERIENCED INAPPROPRIATE SHOCKS, AND THERE WAS A LEAD FRACTURE. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00562 | 2008-01-10 | 20071031 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00563 | 2008-01-10 | 20071105 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00564 | 2008-01-10 | 20070112 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00567 | 2008-01-10 | 20071108 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE PATIENT RECEIVED INAPPROPRIATE SHOCKS DUE TO OVERSENSING. HIGH IMPEDANCE WAS ALSO REPORTED, AT GREATER THAN 3000 OHMS. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00568 | 2008-01-10 | 20071102 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00571 | 2008-01-10 | 20071031 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PACE/SENSE PORTION OF THE RV (RIGHT VENTRICULAR) LEAD WAS REPLACED DUE TO AN APPARENT LEAD FRACTURE. THE HIGH VOLTAGE PORTION REMAINS IN USE. IT WAS LATER REPORTED BY THE PATIENT'S ATTORNEY THAT THE PATIENT RECEIVED INJURIES FROM THE SPRINT FIDELIS LEAD. INAPPROPRIATE SHOCKS DUE TO MALFUNCTION OF DEFECTIVE PRODUCT WERE NOTED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00552 | 2008-01-10 | 20071112 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE PATIENT REPORTED HER LEAD IS FRACTURED. IT WAS EXPLANTED AND REPLACED, WITH A REPORT OF A FRACTURE AND HIGH IMPEDANCE OF GREATER THAN 3000 OHMS. OVERSENSING WAS ALSO REPORTED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. A LAWSUIT ALLEGES THE SURGERY THAT WAS REQUIRED DUE TO THE FRACTURED LEAD CAUSED THE PATIENT TO EXPERIENCE INJURY. IT ALSO ALLEGES THE PATIENT HAS AND WILL CONTINUE TO SUSTAIN "SEVERE PHYSICAL INJURIES, SEVERE EMOTIONAL DISTRESS" DUE TO THE LEAD. ACTUAL DEVICE INVOLVED IN INCIDENT WAS EVALUATED ELECTRICAL TESTS PERFORMED, MECHANICAL TESTS PERFORMED, VISUAL EXAMINATION IMPEDANCE, HIGH LEAD(S), FRACTURE OF OVERSENSING SHOCK, INAPPROPRIATE. |  | N | Y | LAY USER/PATIENT
2649622-2008-00553 | 2008-01-10 | 20071009 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE PATIENT REPORTED RECEIVING THREE SHOCKS. SHE STATED THE LEAD BROKE. ADDITIONAL INFORMATION WAS SUBSEQUENTLY RECEIVED FROM THE PHYSICIAN'S OFFICE REPORTING LEAD FRACTURE. THE PACE/SENSE PORTION OF THE LEAD WAS CAPPED AND REPLACED WITH A NEW LEAD. THE HIGH VOLTAGE PORTION REMAINS IN USE. SUBSEQUENT INFORMATION WAS RECEIVED REPORTING HIGH DEFIB COIL IMPEDANCES OF GREATER THAN 200 OHMS. THE PACE/SENSE PORTION OF THE LEAD HAD PREVIOUSLY BEEN REPLACED DUE TO HIGH IMPEDANCE AND OVERSENSING. NOW, THE ENTIRE LEAD WAS REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00555 | 2008-01-10 | 20070915 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00556 | 2008-01-10 | 20071015 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT RECEIVED INAPPROPRIATE SHOCKS. IT WAS ALSO NOTED THAT THE PATIENT'S EJECTION FRACTION WAS LOW, AND BRADY AND TACHY FEATURES WERE PROGRAMMED OFF. IT IS NOT KNOWN IF THE LEAD HAS BEEN EXPLANTED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED REPORTING THAT THERE WAS OVERSENSING, AND AN APPARENT LEAD FRACTURE. THE LEAD WAS EXPLANTED. |  | Y | Y | LAY USER/PATIENT
2182208-2008-00024 | 2008-01-10 | 20071114 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC, INC. | 6948 | ASKU |  | OTHER: IT WAS REPORTED THAT THE RIGHT VENTRICULAR LEAD WAS REPLACED, DUE TO VARYING IMPEDANCE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED REPORTING THAT THERE WAS HIGH IMPEDANCE AND A LEAD FRACTURE. NO CONCLUSION CAN BE DRAWN. IMPEDANCE, HIGH, LEAD(S), FRACTURE OF. |  | N | Y | LAY USER/PATIENT
2649622-2008-00557 | 2008-01-10 | 20071031 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT PRESENTED TO THE EMERGENCY ROOM AFTER RECEIVING SEVERAL INAPPROPRIATE SHOCKS DUE TO OVERSENSING, AND THERE WAS HIGH IMPEDANCE OF GREATER THAN 1500 OHMS. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. AN ALLEGATION FROM AN ATTORNEY INDICATED THE LEAD HAD EXHIBITED A FRACTURE AND/OR WAS EXPLANTED. ADDITIONALLY IT WAS ALLEGED THE PATIENT IS DECEASED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00558 | 2008-01-10 | 20071027 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED BY THE PATIENT THAT HE HAD EXPERIENCED 2 SHOCKS AND WAS TOLD BY THE MANUFACTURER'S REPRESENTATIVE THAT THE LEAD HAD FRACTURED. THE LEAD WAS EXPLANTED AND REPLACED THE NEXT MORNING. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00559 | 2008-01-10 | 20071101 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE PATIENT WAS SEEN IN THE EMERGENCY ROOM AFTER EXPERIENCING MULTIPLE SHOCKS. OVERSENSING WAS OBSERVED. THE PACE/SENSE PORTION OF THE LEAD WAS CAPPED AND REPLACED WITH A NEW LEAD. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION RECEIVED REPORTED THE PACE/SENSE PORTION OF THE LEAD WAS FRACTURED AND THERE HAD BEEN A SLIGHT INCREASE IN IMPEDANCE OF 50 - 100 OHMS. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00560 | 2008-01-10 | 20070910 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED BY THE PATIENT THAT SHE WAS SHOCKED 6 TIMES IN ONE MINUTE AND WAS TAKEN BY AMBULANCE TO THE EMERGENCY ROOM, WHERE THE DEVICE WAS TURNED OFF. SHE WAS TOLD THE LEAD WAS "A BAD LEAD." THE PATIENT REPORTED BEING WEAK, SCARED, AND EMOTIONAL AFTER THE SHOCKS. THE LEAD WAS REPLACED. IT WAS SUBSEQUENTLY REPORTED BY AN ATTORNEY THAT THE PATIENT "EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY REQUIRING MEDICAL INTERVENTION INCLUDING BUT NOT LIMITED TO SURGERY, REMOVAL AND/OR REPLACEMENT" OF THE DEFECTIVE LEAD. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00551 | 2008-01-10 | 20071030 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00540 | 2008-01-10 | 20071025 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00541 | 2008-01-10 | 20071031 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00542 | 2008-01-10 | 20071031 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-00543 | 2008-01-10 | 20071029 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00544 | 2008-01-10 | 20071106 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-00545 | 2008-01-10 | 20071023 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00546 | 2008-01-10 | 20071025 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE RIGHT VENTRICULAR LEAD WAS EXPLANTED DUE TO T-WAVE OVERSENSING. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED REPORTING THAT THE PATIENT RECEIVED INAPPROPRIATE SHOCKS. ACTUAL DEVICE INVOLVED IN INCIDENT WAS EVALUATED ELECTRICAL TESTS PERFORMED, MECHANICAL TESTS PERFORMED, VISUAL EXAMINATION DEVICE PERFORMED ACCORDING TO SPECIFICATIONS DEVICE EVALUATED AND ALLEGED FAILURE COULD NOT BE DUPLICATED INTERFERENCE OVERSENSING SHOCK, INAPPROPRIATE. |  | Y | N | LAY USER/PATIENT
2649622-2008-00547 | 2008-01-10 | 20071101 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE RIGHT VENTRICULAR LEAD WAS EXPLANTED DUE TO AN IMPEDANCE INCREASE UP TO 2500 OHMS, AND OVERSENSING. IT WAS NOTED THAT IMPEDANCE VARIED BY AS MUCH AS 1000 OHMS EACH TIME THE DEVICE WAS CHECKED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED ALLEGING, THE PATIENT 'SUFFERED PHYSICAL AND OTHER INJURY AS A RESULT OF THE RECALLED LEAD'. IT ALSO ALLEGED THAT THE PATIENT 'EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY', RESULTING IN REMOVAL OF THE 'DEFECTIVE' LEAD. THE PATIENT 'SUFFERED REPEATED SHOCKS AND FRACTURES, AFTER THE ENHANCED MONITORING SYSTEM WAS IN PLACE, AND THE MONITORING FAILED'. NO MONITOR ALARM SOUNDED. AND THE PATIENT 'HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH, ECONOMIC LOSSES AND OTHER DAMAGES.' |  | Y | N | LAY USER/PATIENT
2649622-2008-00548 | 2008-01-10 | 20071024 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00549 | 2008-01-10 | 20071031 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00550 | 2008-01-10 | 20071029 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00531 | 2008-01-10 | 20071112 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00532 | 2008-01-10 | 20071108 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE LEAD WAS EXPLANTED DUE TO INAPPROPRIATE SHOCKS CAUSED BY OVERSENSING OF NOISE. TWO STRESS TESTS PERFORMED ON THE IMPLANT SITE YIELDED NO OVERSENSING, HOWEVER, WHEN THE PATIENT TURNED SIDEWAYS OVERSENSING WAS SEEN. VARYING LEAD IMPEDANCE WAS ALSO NOTED. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. LEAD WAS RETURNED WITH LEAD FRACTURE NOTED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00536 | 2008-01-10 | 20071120 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-00539 | 2008-01-10 | 20071018 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00519 | 2008-01-10 | 20071027 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00520 | 2008-01-10 | 20071026 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00521 | 2008-01-10 | 20071026 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00522 | 2008-01-10 | 20070713 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00523 | 2008-01-10 | 20071015 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00525 | 2008-01-10 | 20070928 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00526 | 2008-01-10 | 20071031 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE PATIENT RECEIVED MULTIPLE INAPPROPRIATE SHOCKS DUE TO OVERSENSING OF NOISE. THE LEAD HAD BEEN HAVING UNSTABLE PACING IMPEDANCE VALUES. IT WAS EXPLANTED AND REPLACED WITH A COMPETITOR'S LEAD. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00527 | 2008-01-10 | 20071026 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-00528 | 2008-01-10 | 20071024 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00529 | 2008-01-10 | 20070809 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THAT AT ONE MONTH FOLLOW-UP, OVERSENSING, CAUSED BY NOISE, WAS DETECTED. THE LEAD WAS EXPLANTED AND REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-00507 | 2008-01-10 | 20071024 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. EVALUATION SUMMARY: DISTAL CONDUCTOR FRACTURED; FULL LEAD IN SEGMENTS RETURNED AND ANALYZED |  | Y | Y | LAY USER/PATIENT
2649622-2008-00508 | 2008-01-10 | 20070608 | Injury | Manufacturer report | SPRINT FIDELIS | DEFIBRILLATOR, AUTOMATIC IMPLANTABLE CARDIOVERTER | MEDTRONIC PUERTO RICO, INC. | 6949 | 6949 |  | IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00509 | 2008-01-10 | 20071024 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE, DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY:PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00510 | 2008-01-10 | 20071024 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00513 | 2008-01-10 | 20071025 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00514 | 2008-01-10 | 20071025 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6930 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00506 | 2008-01-10 | 20071010 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00496 | 2008-01-10 | 20071023 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR WAS FRACTURED; FULL LEAD WAS RETURNED FOR ANALYSIS. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00497 | 2008-01-10 | 20071012 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. EVALUATION SUMMARY: DISTAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. IT WAS REPORTED THAT THE PATIENT RECEIVED INAPPROPRIATE THERAPY DUE TO OVERSENSING. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT ACTUAL DEVICE INVOLVED IN INCIDENT WAS EVALUATED ELECTRICAL TESTS PERFORMED, MECHANICAL TESTS PERFORMED, VISUAL EXAMINATION LEAD CONDUCTOR COMPONENT/SUBASSEMBLY FAILURE DEVICE WAS OUT OF SPECIFICATION IN A MANNER THAT RELATES TO EVENT OVERSENSING SHOCK, INAPPROPRIATE. |  | Y | N | LAY USER/PATIENT
2649622-2008-00498 | 2008-01-10 | 20071022 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00499 | 2008-01-10 | 20071029 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-00500 | 2008-01-10 | 20071022 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED FOR ANALYSIS. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00501 | 2008-01-10 | 20071022 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-00502 | 2008-01-10 | 20071015 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: NO ANOMALIES WERE FOUND; FULL LEAD IN SEGMENTS WAS RETURNED FOR ANALYSIS. |  | N | Y | LAY USER/PATIENT
2649622-2008-00503 | 2008-01-10 | 20071102 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE PATIENT PRESENTED TO THE EMERGENCY DEPARTMENT AFTER RECEIVING SEVERAL INAPPROPRIATE SHOCKS. THE PACING IMPEDANCE WAS DETERMINED TO BE UNSTABLE. A LEAD FRACTURE WAS SUSPECTED. THE LEAD WAS EXPLANTED AND REPLACED WITH A COMPETITOR LEAD. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00504 | 2008-01-10 | 20071101 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00494 | 2008-01-10 | 20071012 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00484 | 2008-01-10 | 20070822 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO SEPSIS FOLLOWING A DEVICE CHANGEOUT. THE PATIENT HAD A DEVICE SYSTEM UPGRADE IN 2007. ON THE FOLLOWING MONTH, THAT SYSTEM WAS REMOVED DUE TO POCKET INFECTION. THE PATIENT WAS PLACED ON ANTIBIOTICS AND WHEN HIS CULTURES WERE NEGATIVE THIS SYSTEM WAS IMPLANTED. THE PATIENT'S INFECTION REAPPEARED AND HE WAS AGAIN PUT ON ANTIBIOTICS UNTIL HIS DEATH. THERE IS NO ALLEGATION THAT THE DEATH WAS DEVICE RELATED. |  | Y | N | LAY USER/PATIENT
2649622-2008-00487 | 2008-01-10 | 20071019 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00488 | 2008-01-10 | 20071022 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6930 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-00489 | 2008-01-10 | 20071022 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00490 | 2008-01-10 | 20071019 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD IN SEGMENTS RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-00491 | 2008-01-10 | 20071030 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00492 | 2008-01-10 | 20071024 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: DISTAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-00483 | 2008-01-10 | 20071011 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS EXPLANTED AND REPLACED DUE TO OVERSENSING. MORE THAN 59000 SHORT INTERVAL COUNTS WERE MEASURED IN 3 DAYS. A LEAD FRACTURE WAS SUSPECTED. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-00474 | 2008-01-10 | 20071030 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PACE/SENSE PORTION OF THE RV LEAD WAS CAPPED AND REPLACED BY A NEW PACE/SENSE DUE TO AN INAPPROPRIATE SHOCK CAUSED BY OVERSENSING. LEAD IMPEDANCE ALSO JUMPED TO GREATER THAN 3000 OHMS. THE LEAD WAS EXPLANTED AND REPLACED. IT WAS ALSO REPORTED THAT A LAWSUIT ALLEGED THAT THE PATIENT 'SUFFERED PHYSICAL AND OTHER INJURY AS A RESULT OF THE RECALLED LEAD', AND THE PATIENT 'EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY REQUIRING MEDICAL INTERVENTION INCLUDING BUT NOT LIMITED TO SURGERY, REMOVAL AND/OR REPLACEMENT OF THE DEFECTIVE' LEAD. IT ALSO ALLEGED THAT THE PATIENT HAS 'SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH, ECONOMIC LOSSES AND OTHER DAMAGES.' THE PATIENT HAS FURTHER SUFFERED 'VARIOUS PHYSICAL MANIFESTATIONS OF EMOTIONAL DISTRESS'. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00475 | 2008-01-10 | 20071113 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THAT THE RV LEAD WAS REPLACED DUE TO A SPIKE SEEN IN THE HIGH VOLTAGE COIL IMPEDANCE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. THE LEAD WAS RETURNED WITH A REPORT OF LEAD FRACTURE. |  | N | Y | LAY USER/PATIENT
2649622-2008-00476 | 2008-01-10 | 20071029 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PACE/SENSE PORTION OF THE RV LEAD WAS CAPPED AND REPLACED WITH A NEW PACE/SENSE LEAD DUE TO INAPPROPRIATE SHOCKS CAUSED BY OVERSENSING OF NOISE. THE HIGH VOLTAGE PORTION OF THE LEAD REMAINS IN USE. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. THE PACE/SENSE PORTION OF THE LEAD WAS RETURNED FOR ANALYSIS DUE TO AN APPARENT LEAD FRACTURE. THE HIGH VOLTAGE PORTION OF THE LEAD WAS LATER TAKEN OUT OF USE. |  | Y | Y | LAY USER/PATIENT
2182208-2008-00022 | 2008-01-10 | 20071023 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC, INC. | 6948 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00478 | 2008-01-10 | 20071016 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-00479 | 2008-01-10 | 20071010 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6930 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00480 | 2008-01-10 | 20071018 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-00481 | 2008-01-10 | 20071016 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00462 | 2008-01-10 | 20071025 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00463 | 2008-01-10 | 20071031 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00464 | 2008-01-10 | 20071022 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00465 | 2008-01-10 | 20071029 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00466 | 2008-01-10 | 20071031 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00468 | 2008-01-10 | 20071107 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00469 | 2008-01-10 | 20071025 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD CAUSED A SHOCK TO BE DELIVERED, DUE TO NOISE. THERE WAS ALSO A BRIEF IMPEDANCE SPIKE OF 3000 OHMS. A LEAD FRACTURE WAS SUSPECTED. THE LEAD WAS CAPPED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00470 | 2008-01-10 | 20071020 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT BOTH SVC AND HVB LEAD IMPEDANCE WERE GREATER THAN 200 OHMS. THE LEAD WAS REPLACED. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION RECEIVED REPORTED THERE WAS NO PACING OR DEFIBRILLATION OUTPUT AND THE LEAD WAS FRACTURED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00471 | 2008-01-10 | 20071012 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: THE ACTUAL DEVICE WAS NOT RECEIVED FOR EVALUATION. WE DID RECEIVE PERFORMANCE DATA COLLECTED FROM THE DEVICE AND HAVE ANALYZED THE DATA. ANALYSIS OF THAT DATA REVEALED NOISE, HIGH VENTRICULAR PACING IMPEDANCE OF 928-2224 OHMS, AND AN ELEVATED VENTRICULAR SENSING INTEGRITY COUNTER. A HIGH VALUE OF THIS COUNTER CAN INDICATE A POSSIBLE INTERMITTENCY IN THE SYSTEM. OTHER: IT WAS REPORTED THE PACE/SENSE PORTION OF THE LEAD WAS CAPPED AND REPLACED WITH A NEW PACE/SENSE LEAD DUE TO NOISE ON THE LEAD. OVERSENSING AND MULTIPLE INAPPROPRIATE SHOCKS WERE ALSO REPORTED. THE HIGH VOLTAGE PORTION OF THE LEAD REMAINS IN USE. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. NO CONCLUSION CAN BE DRAWN. INTERFERENCE, OVERSENSING; DEVICE REMAINS ACTIVATED, SHOCK, INAPPROPRIATE. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00461 | 2008-01-10 | 20061009 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00455 | 2008-01-10 | 20071102 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES FOUND; FULL LEAD RETURNED AND ANALYZED. THE ANALYST ALSO NOTED THERE WAS ONLY ONE SET OF SETSCREW MARKS ON THE IS-1 PIN. THESE SETSCREW MARKS WERE TOO PROXIMAL, INDICATING THE LEAD WAS NOT FULLY INSERTED INTO THE DEVICE. |  | N | Y | LAY USER/PATIENT
2649622-2008-00457 | 2008-01-10 | 20071001 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00458 | 2008-01-10 | 20071012 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | A CALL WAS RECEIVED THROUGH TECHNICAL SERVICES REPORTING LEAD FRACTURE AND OVERSENSING. THE LEAD WAS EXPLANTED AND RETURNED WITH A REPORT OF FRACTURE AND 'SEVERAL' INAPPROPRIATE SHOCKS. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00459 | 2008-01-10 | 20071025 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-00452 | 2008-01-10 | 20071013 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00445 | 2008-01-10 | 20071101 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00446 | 2008-01-10 | 20071101 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00431 | 2008-01-10 | 20071024 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00432 | 2008-01-10 | 20071023 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00433 | 2008-01-10 | 20071025 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00434 | 2008-01-10 | 20071026 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE LEAD WAS EXPERIENCING OVERSENSING, AND THAT THE IMPEDANCE WAS GREATER THAN 3000 OHMS. THE LEAD WAS EXPLANTED AND REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED IN A LAWSUIT ALLEGING THE PATIENT "EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, OR FAILURE TO RECEIVE THERAPEUTIC SHOCKS. EACH OF WHICH HAVE REQUIRED THEM TO SEEK MEDICAL INTERVENTION RESULTING IN REPLACEMENT OF THE DEFECTIVE" LEAD. IT FURTHER ALLEGED THAT THE PATIENT "SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH, ECONOMIC LOSSES AND OTHER DAMAGES" AND THAT THE "[PATIENT HAS] FURTHER SUFFERED PHYSICAL INJURIES AND VARIOUS PHYSICAL MANIFESTATIONS OF EMOTIONAL DISTRESS". LEAD "FAILURE" WAS ALSO NOTED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00435 | 2008-01-10 | 20071024 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: NO ANOMALIES FOUND; FULL LEAD IN SEGMENTS RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-00436 | 2008-01-10 | 20071023 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00437 | 2008-01-10 | 20071023 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00438 | 2008-01-10 | 20071031 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY DISTAL CONDUCTOR FRACTURED; FULL LEAD RETURNED IN SEGMENTS. |  | N | Y | LAY USER/PATIENT
2649622-2008-00439 | 2008-01-10 | 20070125 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00440 | 2008-01-10 | 20071025 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00442 | 2008-01-10 | 20071025 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00421 | 2008-01-10 | 20060607 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00423 | 2008-01-10 | 20071015 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00425 | 2008-01-10 | 20050616 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00426 | 2008-01-10 | 20070920 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00420 | 2008-01-10 | 20070324 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00411 | 2008-01-10 | 20071019 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00412 | 2008-01-10 | 20071024 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00415 | 2008-01-10 | 20071025 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS DELIVERING INAPPROPRIATE THERAPY DUE TO OVERSENSING. THE IMPEDANCES HAD BEEN INCREASING OVER THE PREVIOUS COUPLE WEEKS, AND THERE HAD BEEN A GRADUAL REDUCTION IN SENSING OF R WAVES. FURTHER ALLEGED THAT THERE WAS A LEAD FRACTURE AND THAT THE PATIENT "EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING". THE LEAD WAS CAPPED AND REPLACED. IT WAS NOTED THAT THE HELIX WOULD NOT RETRACT. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00416 | 2008-01-10 | 20071022 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00417 | 2008-01-10 | 20071018 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS EXPLANTED AND REPLACED DUE TO MULTIPLE SHOCKS DUE TO OVERSENSING. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00418 | 2008-01-10 | 20070907 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00419 | 2008-01-10 | 20060602 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00410 | 2008-01-10 | 20071023 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE LEAD WAS EXPLANTED DUE TO INAPPROPRIATE SHOCKS DUE TO NOISE ON RV LEAD. ALSO, VENTRICULAR IMPEDANCE WAS HIGH, THREHOLDS WERE UNACCEPTABLE, AND APPARENT LEAD FRACTURE WAS NOTED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00402 | 2008-01-10 | 20071022 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00403 | 2008-01-10 | 20071023 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-00404 | 2008-01-10 | 20071022 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT RECEIVED INAPPROPRIATE SHOCKS, THERE WAS HIGH IMPEDANCE AND OVERSENSING. THE PATIENT WAS ADMITTED TO THE HOSPITAL FOR LEAD REPLACEMENT AS A FRACTURE WAS SUSPECTED. HVB AND SVC IMPEDANCES WERE NORMAL, SO THE HIGH VOLTAGE COIL REMAINS IN USE. THE PACE/SENSE PORTION OF THE LEAD WAS CAPPED, AND A NEW LEAD WAS PLACED FOR PACING AND SENSING. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED FROM AN ATTORNEY ALLEGING THE PATIENT SUSTAINED INJURIES. NO SPECIFIC PATIENT INJURIES WERE REPORTED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00405 | 2008-01-10 | 20071128 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE RV LEAD WAS CAPPED DUE TO OVERSENSING AND VARYING IMPEDANCE FROM 400S TO 1200S OHMS. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. FURTHER INFORMATION FRON AN ATTORNEY ALLEGES THE PLAINTIFF HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, AND MENTAL ANGUISH AS A RESULT OF THE LEAD. |  | N | Y | LAY USER/PATIENT
2649622-2008-00406 | 2008-01-10 | 20071107 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00407 | 2008-01-10 | 20071031 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: DISTAL CONDUCTOR FRACTURED; FULL LEAD IN SEGMENTS RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-00408 | 2008-01-10 | 20071101 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-00409 | 2008-01-10 | 20071001 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE DEFIBRILLATION COIL IMPEDANCE WAS HIGH AND THE LEAD WAS COMPLETELY FLAT. IT WAS ALSO REPORTED THAT THE PATIENT IS A ROOFER AND CARRIES SHINGLES ON THAT SIDE OF HIS BODY. THE LEAD WAS EXPLANTED AND RETURNED TO MEDTRONIC FOR ANALYSIS. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION REPORTED LEAD FRACTURE. |  | N | Y | LAY USER/PATIENT
2649622-2008-00400 | 2008-01-10 | 20071027 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD EXHIBITED HIGH PACING IMPEDANCE OF GREATER THAN 3000 OHMS, OVERSENSING, AND NOISE. AN APPARENT FRACTURE WAS SUSPECTED, SO THE LEAD WAS EXPLANTED AND REPLACED. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-00390 | 2008-01-10 | 20071026 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00391 | 2008-01-10 | 20071018 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00392 | 2008-01-10 | 20071107 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-00393 | 2008-01-10 | 20070912 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00394 | 2008-01-10 | 20060428 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT EXPIRED, AND THE PATHOLOGIST CONFIRMED A LEAD FRACTURE. THERE IS NO ALLEGATION THAT THE LEAD FRACTURE CONTRIBUTED TO THE PATIENT'S DEATH. ADDITIONAL INFORMATION WAS RECEIVED REPORTING THAT THE PATIENT HAD BEEN ADMITTED TO THE HOSPITAL FOUR DAYS PRIOR TO DEATH DUE TO SYMPTOMS OF EXERTIONAL CHEST DISCOMFORT AND SHORTNESS OF BREATH. DEVICE INTERROGATION REVEALED NORMAL FUNCTION WITH NO VF/VT OR BRADYCARDIC EPISODES. THE PATIENT WAS IN THE HOSPITAL AND ON TELEMETRY WHEN HE STOPPED BREATHING AND BECAME PULSELESS. RESCUE ATTEMPTS WERE UNSUCCESSFUL AND THE PATIENT DIED. THE PATIENT'S HEART RYTHYM AT THE TIME OF DEATH WAS SINUS RYTHYM WITH PULSELESS ELECTRICAL ACTIVITY. THE CAUSE OF DEATH WAS REPORTED TO BE DILATED CARDIOMYOPATHY. INFORMATION FROM THE HCP REPORTED THAT THE PATIENT'S DEATH WAS NOT SYSTEM RELATED. THERE IS NO NOTE IN THE AUTOPSY REPORT OF A LEAD FRACTURE. THE LEAD WAS EXPLANTED, BUT HAS NOT BEEN RETURNED TO MEDTRONIC. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00395 | 2008-01-10 | 20071019 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-00396 | 2008-01-10 | 20071023 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD HAD DELIVERED MULTIPLE INAPPROPRIATE THERAPIES DUE TO OVERSENSING, AS INDICATED BY THE SIC (SHORT INTERVAL COUNTER), WHICH MEASURED 80170 COUNTS. LEAD IMPEDANCE HAD FLUCTUATED BETWEEN NORMAL RANGE TO GREATER THAN 3000 OHMS. THE LEAD WAS REPLACED. ADDITIONAL INFORMATION WAS LATER RECEIVED FROM AN ATTORNEY REPORTING THE PATIENT "EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY REQUIRING MEDICAL INTERVENTION INCLUDING BUT NOT LIMITED TO SURGERY, REMOVAL AND/OR REPLACEMENT OF THE DEFECTIVE [LEAD]." NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00397 | 2008-01-10 | 20071031 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00398 | 2008-01-10 | 20071016 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00399 | 2008-01-10 | 20071031 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00378 | 2008-01-10 | 20071025 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE LEAD WAS EXPLANTED AND REPLACED DUE TO INAPPROPRIATE SHOCKS CAUSED BY OVERSENSING FROM A FRACTURED LEAD. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00380 | 2008-01-10 | 20071022 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00382 | 2008-01-10 | 20071022 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT RECEIVED NUMEROUS SHOCKS AND WAS BROUGHT TO THE ER, WHERE A MAGNET WAS APPLIED, AND THE SHOCKS STOPPED. REVIEW OF SAVE-TO-DISK DATA REVEALED RV (RIGHT VENTRICULAR) IMPEDANCE OF 1440 OHMS, FROM PREVIOUSLY STABLE IMPEDANCES IN THE MID 400S, VARYING IMPEDANCE, AND NOISE. IT WAS ALSO NOTED THAT THE PATIENT FELL THE WEEK PRIOR AND HIT HIS HEAD. IT WAS REPORTED THAT THE FALL WAS NOT RELATED TO THE DEVICE FUNCTION, BUT IT WAS QUESTIONABLE IF IT MAY HAVE CONTRIBUTED TO A POSSIBLE LEAD FRACTURE. ADDITIONAL INFORMATION WAS SUBSEQUENTLY RECEIVED FROM THE PATIENT'S SON REPORTING "AGONIZING PAIN" AND SUFFERING FOR HIS FATHER DURING THE "UNWARRANTED" SHOCKS. THE LEAD WAS REPORTEDLY DIAGNOSED AS MALFUNCTIONING IN THE ER. ALLEGATIONS WERE MADE BY THE PATIENT'S SON THAT THE PATIENT'S QUALITY OF LIFE HAS BEEN DIMINISHED BY THE LEAD FAILURE. AS A RESULT, HE IS NOW ON OXYGEN 100% OF THE TIME, CANNOT WALK MORE THAN 20 FEET WITHOUT LOSING BREATH, AND CAN NO LONGER DRIVE A CAR. THE LEAD WAS CAPPED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00385 | 2008-01-10 | 20071030 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00377 | 2008-01-10 | 20071017 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00367 | 2008-01-10 | 20071015 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00368 | 2008-01-10 | 20071018 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | IT WAS REPORTED THAT THE PATIENT RECEIVED INAPPROPRIATE SHOCKS DUE TO OVERSENSING, AND RIGHT VENTRICULAR PACING IMPEDANCE HAD INCREASED TO 1400 OHMS. THE LEFT VENTRICULAR LEAD HAD DISLODGED. BOTH LEADS WERE EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00370 | 2008-01-10 | 20071018 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00371 | 2008-01-10 | 20071016 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00373 | 2008-01-10 | 20061106 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00374 | 2008-01-10 | 20070720 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-00375 | 2008-01-10 | 20071024 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00376 | 2008-01-10 | 20071015 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00366 | 2008-01-10 | 20071002 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6930 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-00357 | 2008-01-10 | 20071015 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. EVALUATION SUMMARY - NO ANOMALIES FOUND. FULL LEAD RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-00358 | 2008-01-10 | 20071017 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-00359 | 2008-01-10 | 20071016 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00360 | 2008-01-10 | 20071016 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00361 | 2008-01-10 | 20071016 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE, DUE TO CONFIDENTIALITY CONCERNS. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00362 | 2008-01-10 | 20071013 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00363 | 2008-01-10 | 20071016 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. EVALUATION SUMMARY: NO ANOMALIES FOUND; FULL LEAD RETURNED IN SEGMENTS AND ANALYZED. AT THE TIME OF ANALYSIS, IT WAS ALSO NOTED, THERE IS ONLY ONE SET OF SETSCREW MARKS ON THE IS-1 PIN AND IT IS TOO PROXIMAL - LEAD WAS NOT FULLY INSERTED INTO THE DEVICE. |  | N | Y | LAY USER/PATIENT
2649622-2008-00364 | 2008-01-10 | 20071009 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00365 | 2008-01-10 | 20071017 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-00356 | 2008-01-10 | 20071015 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00345 | 2008-01-10 | 20071018 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00348 | 2008-01-10 | 20071019 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00349 | 2008-01-10 | 20071019 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY - NO ANOMALIES FOUND. FULL LEAD RETURNED AND ANALYZED. THERE IS ONLY ONE SET OF SETSCREW MARKS ON THE IS-1 PIN AND IT IS TOO PROXIMAL, WHICH INDICATES THAT THE LEAD WAS NOT FULLY INSERTED INTO THE DEVICE HEADER. |  | N | Y | LAY USER/PATIENT
2649622-2008-00350 | 2008-01-10 | 20071018 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00352 | 2008-01-10 | 20071004 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00353 | 2008-01-10 | 20071012 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00355 | 2008-01-10 | 20071116 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-00344 | 2008-01-10 | 20071022 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00336 | 2008-01-10 | 20071019 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | FRACTURE OF THE PACE/SENSE PORTION OF THE HIGH VOLTAGE LEAD WAS REPORTED. THE LEAD WAS EXPLANTED. ATTORNEY ALLEGED THE PATIENT SUFFERED PHYSICAL AND OTHER INJURY DUE TO THE LEAD. IT WAS ALSO ALLEGED THE PATIENT EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY REQUIRING MEDICAL INTERVENTION. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-00337 | 2008-01-10 | 20071022 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE DEVICE WAS EXPLANTED AND REPLACED DUE TO PREMATURE BATTERY DEPLETION. THE LEAD WAS EXPLANTED AND REPLACED DUE TO INAPPROPRIATE THERAPY, LOSS OF CAPTURE, SENSING DIFFICULTY, FRACTURE, AND THRESHOLD INCREASE. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | N | LAY USER/PATIENT
2649622-2008-00338 | 2008-01-10 |  | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00342 | 2008-01-10 | 20071022 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00343 | 2008-01-10 | 20071022 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00328 | 2008-01-10 | 20071019 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00329 | 2008-01-10 | 20071016 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00330 | 2008-01-10 | 20071011 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | HIGH PACING IMPEDANCE WAS REPORTED. THE LEAD WAS EXPLANTED AND REPLACED. AN ATTORNEY ALLEGED THE PATIENT SUFFERED PHYSICAL AND OTHER INJURY DUE TO THE LEAD. A LAWSUIT STATES THE PATIENT EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE, OR OTHER INJURY REQUIRING MEDICAL INTERVENTION. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-00332 | 2008-01-10 | 20071126 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-00324 | 2008-01-10 | 20071023 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE DEVICE BATTERY VOLTAGE HAD DROPPED FROM 3.20V, ONE MONTH POST IMPLANT, TO 2.59V, TWO MONTHS LATER. ONE MONTH LATER, THE DEVICE COULD NOT BE INTERROGATED DURING A CLINIC VISIT. THERE WAS NO EVIDENCE OF PACING OR OTHER ACTIVITY FROM THE DEVICE. IT WAS REPLACED. AN ATTORNEY LATER ALLEGED THE PRODUCTS WERE 'DEFECTIVE AND FAILED'. HIS CLAIM ALSO ALLEGED PATIENT INJURY. ADDITIONAL INFORMATION STATES A LAWSUIT ALLEGES THE DEVICE HAD EARLY BATTERY DEPLETION, DUE TO THE LEAD. IT ALSO STATES THE REPLACEMENT OF THE ICD AND LEAD CAUSED 'SCARRING ' TO THE PATIENT'S 'ALREADY FRAGILE HEART' AND FORCED HIM TO UNDERGO 'ADDITIONAL AND UNNECESSARY COMPLICATED SURGERIES'. IT FURTHER STATES AS A RESULT, THE PATIENT 'SUFFERED SEVERE PHYSICAL AND EMOTIONAL INJURIES' AND 'SEVERE POST-TRAUMATIC EMOTIONAL DISTRESS AS A RESULT OF THE DEFECTIVE LEAD'. NO CONCLUSION CAN BE DRAWN; DEFECTIVE DEVICE. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00314 | 2008-01-10 | 20071009 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00319 | 2008-01-10 | 20070720 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00320 | 2008-01-10 | 20070921 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00321 | 2008-01-10 | 20071012 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT RECEIVED INAPPROPRIATE THERAPY. THE SHORT INTERVAL COUNTER WAS GREATER THAN 300, INDICATING OVERSENSING, AND THERE WAS AN APPARENT LEAD FRACTURE. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED IN A LAWSUIT WHICH ALLEGES THAT AS A RESULT OF THE LEAD, THE PATIENT 'HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS AND ECONOMIC AND CONSEQUENTIAL DAMAGES'. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00322 | 2008-01-10 | 20071018 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00323 | 2008-01-10 | 20071019 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | UNDERSENSING WAS REPORTED. THE LEAD WAS EXPLANTED AND REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. FURTHER INFORMATION RECEIVED FROM AN ATTORNEY ALLEGES PATIENT SUFFERED PHYSICAL AND OTHER INJURY AS A RESULT OF THE RECALLED LEAD. PATIENT EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY REQUIRING MEDICAL INTERVENTION INCLUDING BUT NOT LIMITED TO SURGERY, REMOVAL AND/OR REPLACEMENT OF THE DEFECTIVE SPRINT FIDELIS LEAD. PATIENT HAS SUFFERED REPEATED SHOCKS AND FRACTURES, AFTER THE ENHANCED MONITORING SYSTEM WAS IN PLACE, AND THE MONITORING FAILED. PATIENTS SUFFERED THESE SHOCKS AND FRACTURES WITHOUT ANY ALARM SOUNDING FROM THIS PURPORTEDLY "ENHANCED" MONITOR. ATTORNEY ALSO ALLEGES THE PATIENT "HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH". |  | N | N | LAY USER/PATIENT
2649622-2008-00305 | 2008-01-10 | 20070918 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: THE ACTUAL DEVICE WAS NOT RECEIVED FOR EVALUATION. WE DID RECEIVE PERFORMANCE DATA COLLECTED FROM THE DEVICE AND HAVE ANALYZED THE DATA, WHICH SHOWED HIGH IMPEDANCE MEASUREMENTS.. |  | N | Y | LAY USER/PATIENT
2649622-2008-00310 | 2008-01-10 | 20071015 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | INAPPROPRIATE THERAPY WAS REPORTED AND THE LEAD WAS REPLACED. ADDITIONAL INFORMATION STATES AN ATTORNEY ALLEGES THE PATIENT IS 'SUFFERING COMPLICATIONS' DUE TO THE LEAD. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00311 | 2008-01-10 | 20071016 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00303 | 2008-01-10 | 20071113 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | HIGH IMPEDANCE WAS REPORTED. THE LEAD WAS EXPLANTED AND REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION RECEIVED REPORTED LEAD IMPEDANCE WAS GREATER THAN 2500 OHMS AND THERE WAS AN APPARENT FRACTURE. FURTHER INFORMATION RECEIVED FROM AN ATTORNEY ALLEGES PATIENT SUFFERED PHYSICAL AND OTHER INJURY AS A RESULT OF THE RECALLED LEAD. PATIENT EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY REQUIRING MEDICAL INTERVENTION INCLUDING BUT NOT LIMITED TO SURGERY, REMOVAL AND/OR REPLACEMENT OF THE DEFECTIVE SPRINT FIDELIS LEAD. ATTORNEY ALSO ALLEGES THE PATIENT "HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH". |  | Y | Y | LAY USER/PATIENT
2649622-2008-00296 | 2008-01-10 | 20071017 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT RECEIVED MULTIPLE INAPPROPRIATE SHOCKS. THE IMPEDANCE WAS FOUND TO HAVE INCREASED ABOUT THREE DAYS PRIOR TO THE SHOCKS BEING DELIVERED. A FRACTURE WAS ALSO REPORTED. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00297 | 2008-01-10 | 20070227 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00298 | 2008-01-10 | 20071017 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE DEVICE AND LEAD WERE EXPLANTED AND REPLACED. EARLY DEPLETION WAS REPORTED ON THE DEVICE. FRACTURE, HIGH IMPEDANCE, AND INAPPROPRIATE SHOCKS, DUE TO OVERSENSING, WERE REPORTED ON THE LEAD. ADDITIONAL INFORMATION WAS SUBSEQUENTLY RECEIVED REPORTING THE PATIENT PRESENTED TO THE EMERGENCY ROOM DUE TO RECEIVING SHOCKS. THE DEVICE BATTERY DEPLETED FROM 2.9V TO 2.62V ERI (ELECTIVE REPLACEMENT INDICATORS) IN 7 DAYS, SECONDARY TO DELIVERY OF MULTIPLE THERAPIES. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | N | LAY USER/PATIENT
2649622-2008-00299 | 2008-01-10 | 20061106 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00300 | 2008-01-10 | 20071017 | Injury | Manufacturer report | SPRINT FIDELIS | DEFIBRILLATOR, AUTOMATIC IMPLANTABLE CARDIOVERTER | MEDTRONIC PUERTO RICO, INC. | 6949 | 6949 |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00301 | 2008-01-10 | 20071018 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-00302 | 2008-01-10 | 20071016 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00285 | 2008-01-10 | 20071011 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00287 | 2008-01-10 | 20070723 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | N | Y | LAY USER/PATIENT
2649622-2008-00288 | 2008-01-10 | 20071012 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY: NO ANOMALIES FOUND; FULL LEAD RETURNED AND ANALYZED. OTHER: IT WAS REPORTED THE LEAD WAS EXPLANTED AND REPLACED DUE TO A THRESHOLD RISE AND AN INCREASE IN IMPEDANCE TO GREATER THAN 200 OHMS. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION RECEIVED FROM THE PATIENT'S ATTORNEY ALLEGES THE PATIENT "EXPERIENCED FRACTURE OR OTHER MALFUNCTION OF HER LEAD WIRE REQUIRING MEDICAL INTERVENTION." THE ATTORNEY FURTHER ALLEGES THE PATIENT "HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES" AND "MENTAL ANGUISH." ACTUAL DEVICE INVOLVED IN INCIDENT WAS EVALUATED ELECTRICAL TESTS PERFORMED MECHANICAL TESTS PERFORMED VISUAL EXAMINATION DEVICE PERFORMED ACCORDING TO SPECIFICATIONS DEVICE EVALUATED AND ALLEGED FAILURE COULD NOT BE DUPLICATED IMPEDANCE, HIGH LEAD(S), FRACTURE OF HIGH THRESHOLD. |  | Y | N | LAY USER/PATIENT
2649622-2008-00289 | 2008-01-10 | 20071023 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | IT WAS REPORTED THAT THE RIGHT VENTRICULAR LEAD WAS REPLACED BECAUSE IT WAS NOT SENSING APPROPRIATELY. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION RECEIVED REPORTS THERE WAS INTERMITTENT CAPTURE THRESHOLD OF GREATER THAN 6.0V WITH QUESTIONABLE DISLODGEMENT PRIOR TO REPLACEMENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-00291 | 2008-01-10 | 20071008 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | IT WAS REPORTED THAT THE RIGHT VENTRICULAR LEAD WAS EXPLANTED DUE TO A FRACTURE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION RECEIVED REPORTED THE RV LEAD HAD HIGH IMPEDANCES AND OVERSENSING WITH DELIVERY OF INAPPROPRIATE SHOCKS. |  | N | Y | LAY USER/PATIENT
2649622-2008-00292 | 2008-01-10 | 20071012 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-00293 | 2008-01-10 | 20071016 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00294 | 2008-01-10 | 20071014 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00295 | 2008-01-10 | 20071012 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | ASKU |  | Y | N | LAY USER/PATIENT
2649622-2008-00277 | 2008-01-10 | 20071010 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PACE/SENSE PORTION OF THE RV (RIGHT VENTRICULAR) LEAD WAS CAPPED AND REPLACED WITH A NEW PACE/SENSE LEAD, DUE TO A SUDDEN INCREASE IN LEAD IMPEDANCE AND AN APPARENT LEAD FRACTURE. THE HIGH VOLTAGE PORTION OF THE RV LEAD REMAINS IN USE. ADDITIONAL INFORMATION WAS RECEIVED FROM THE PATIENT'S ATTORNEY ALLEGING THE PATIENT EXPERIENCED INAPPROPRIATE SHOCKS, AS A RESULT OF A DEFECTIVE LEAD, WHICH CAUSED THE PATIENT PAIN AND ANXIETY. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00278 | 2008-01-10 | 20071017 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00279 | 2008-01-10 | 20071017 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE LEAD FRACTURED, RESULTING IN 13 INAPPROPRIATE SHOCKS. THE LEAD WAS CAPPED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. INFORMATION RECEIVED FROM AN ATTORNEY ALLEGES PATIENT SUFFERED PHYSICAL AND OTHER INJURY AS A RESULT OF THE RECALLED LEAD. PATIENT EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY REQUIRING MEDICAL INTERVENTION INCLUDING BUT NOT LIMITED TO SURGERY, REMOVAL AND/OR REPLACEMENT OF THE DEFECTIVE SPRINT FIDELIS LEAD. PATIENT HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, AND MENTAL ANGUISH. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00280 | 2008-01-10 | 20070917 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00282 | 2008-01-10 | 20071019 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00265 | 2008-01-10 | 20071016 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00266 | 2008-01-10 | 20071011 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00267 | 2008-01-10 | 20071016 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00268 | 2008-01-10 |  | No answer provided | Manufacturer report | SPRINT FIDELIS | ASKU | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: (B)(4) DISTAL CONDUCTOR WAS FRACTURED; FULL LEAD WAS RETURNED FOR ANALYSIS. |  | N | Y | 0
2649622-2008-00264 | 2008-01-10 | 20071108 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE HIGH VOLTAGE LEAD IMPEDANCE WAS GREATER THAN 200 OHMS. THE LEAD WAS REPLACED. ADDITIONAL INFORMATION PROVIDED REPORTED THE LEAD WAS FRACTURED. ADDITIONAL INFORMATION WAS PROVIDED BY THE PATIENT'S ATTORNEY. ALLEGED THAT THE PATIENT HAS SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES, INCLUDING INAPPROPRIATE THERAPIES, SEVERE EMOTIONAL DISTRESS, AND ECONOMIC AND CONSEQUENTIAL DAMAGES DUE TO THE 'DEFECTIVE' AND FRACTURED LEAD. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00253 | 2008-01-10 | 20071011 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. EVALUATION SUMMARY NO ANOMALIES FOUND; PARTIAL LEAD RETURNED IN SEGMENTS. THERE IS ONLY ONE SET OF SETSCREW MARKS ON THE IS-1 PIN AND THEY APPEAR TO BE TOO PROXIMAL, INDICATING THE LEAD WAS NOT FULLY INSERTED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00254 | 2008-01-10 | 20071010 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE, DUE TO CONFIDENTIALITY CONCERNS. EVALUATION SUMMARY: DISTAL CONDUCTOR FRACTURED; PARTIAL LEAD RETURNED IN SEGMENTS AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-00255 | 2008-01-10 | 20071010 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | N | Y | LAY USER/PATIENT
2649622-2008-00256 | 2008-01-10 | 20070208 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT RECEIVED 18 SHOCKS WHEN HER LEAD FRACTURED. IT WAS FURTHER REPORTED THAT THE PATIENT RECEIVED SEVERAL INAPPROPRIATE SHOCKS FOLLOWING SIGNIFICANT MANIPULATION OF THE DEVICE POCKET DURING ROUTINE MAMMOGRAPHY. VARYING IMPEDANCES WERE NOTED, ALONG WITH NOISE, OVERSENSING, AND LEAD FRACTURE WAS CONFIRMED VIA X-RAY. THE FRACTURE OCCURRED IN THE PACE/SENSE PORTION JUST PROXIMAL TO THE TRIFURCATION. THE LEAD WAS REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00259 | 2008-01-10 | 20071015 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE PATIENT PRESENTED WITH VT STORM AND NUMEROUS APPROPRIATE SHOCKS. HOWEVER, OVERSENSING, V PACING IMPEDANCE > 3000 OHMS, HIGH CAPTURE THRESHOLD, AND SENSING DECREASE WERE REPORTED. THE LEAD WAS REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. APPARENT CONDUCTOR FRACTURE WAS FURTHER REPORTED. THE LEAD WAS CAPPED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00260 | 2008-01-10 | 20071101 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT ALERT SOUNDED DUE TO HIGH IMPEDANCE. ADDITIONAL INFORMATION WAS SUBSEQUENTLY RECEIVED REPORTING FRACTURE. THE LEAD WAS REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2008-00261 | 2008-01-10 | 20071011 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MPRI | 6949 | 6949 |  | THIS MEDICAL DEVICE REPORT IS BEING RESUBMITTED TO THE EMDR SYSTEM DUE TO A CONNECTION ERROR WITHIN THE EMDR SYSTEM. THE EVENT WAS PREVIOUSLY REPORTED TO THE EMDR PORTAL AND AN ACKNOWLEDGEMENT 3 RECEIVED, BUT THE CONNECTION ERROR PREVENTED THE EMDR SYSTEM FROM DOCUMENTING THE REPORT. THE REPORT NUMBER IN THIS REPORT IS THE SAME NUMBER AS THE IMPACTED REPORT THAT HAS THE CONNECTION ISSUE. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00262 | 2008-01-10 | 20071018 | Injury | Manufacturer report | SPRINT FIDELIS | DEFIBRILLATOR, AUTOMATIC IMPLANTABLE CARDIOVERTER | MPRI | 694965 | 694965 |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-00243 | 2008-01-10 | 20071004 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00246 | 2008-01-10 | 20071010 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00247 | 2008-01-10 | 20071010 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00248 | 2008-01-10 | 20071025 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | A CALL WAS RECEIVED THROUGH TECHNICAL SERVICES REPORTING VARYING SVC IMPEDANCE VALUES (50-200 OHMS) AND A CONNECTION ISSUE WAS SUSPECTED. THE LEAD WAS LATER REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION RECEIVED INDICATES LEAD FRACTURE WAS SUSPECTED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00249 | 2008-01-10 | 20071010 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00251 | 2008-01-10 | 20071011 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MPRI | 6949 | 6949 |  | THIS DEVICE WAS REPORTED AS INCLUDED IN THE FIELD ACTION. BASED ON THE INFORMATION RECEIVED AND WITHOUT THE RETURN OF THE PRODUCT, IT CANNOT BE CONFIRMED THAT THIS DEVICE PERFORMED AS DESCRIBED IN THE FIELD ACTION. IT IS INCLUDED IN THE FIELD ACTION IN THE ABUNDANCE OF CAUTION. PRODUCT EVENT SUMMARY: THE LEAD WAS NOT RETURNED FOR ANALYSIS. HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DEVICE MEMORY INDICATED THE IMPEDANCE ON THE RV (RIGHT VENTRICULAR) DEFIBRILLATION COIL WAS BEYOND THE EXPECTED UPPER RANGE. THE RV DEFIB IMPEDANCE WAS STEAD AT APPROXIMATELY 49 OHMS THROUGH (B)(4) 2015 THEN RISES OUT OF RANGE STARTING ON (B)(4) 2015. ANALYSIS OF THE DEVICE MEMORY INDICATED THE IMPEDANCE ON THE SVC (SUPERIOR VENA CAVA) DEFIBRILLATION COIL WAS BEYOND THE EXPECTED UPPER RANGE. THE SVC DEFIB IMPEDANCE WAS STEAD AT APPROXIMATELY 65 OHMS THROUGH (B)(4) 2015 THEN RISES OUT OF RANGE STARTING ON (B)(4) 2015. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00252 | 2008-01-10 | 20071108 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00242 | 2008-01-10 | 20071005 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THIS EVENT OCCURRED OUTSIDE THE US WHERE THE SAME MODEL IS DISTRIBUTED. ALL INFORMATION PROVIDED IS INCLUDED IN THIS REPORT. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PATIENT INFORMATION IS NOT GENERALLY AVAILABLE DUE TO CONFIDENTIALITY CONCERNS. EVALUATION SUMMARY: PROXIMAL CONDUCTOR FRACTURED; FULL LEAD RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00233 | 2008-01-10 | 20071016 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE RV LEAD WAS EXPLANTED DUE TO INAPPROPRIATE SHOCKS, OVERSENSING, AND HIGH IMPEDANCE. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00234 | 2008-01-10 | 20071003 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00236 | 2008-01-10 | 20070730 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00237 | 2008-01-10 | 20061101 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00238 | 2008-01-10 | 20070514 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00239 | 2008-01-10 | 20071017 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00241 | 2008-01-10 | 20070621 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00232 | 2008-01-10 | 20071008 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00227 | 2008-01-10 | 20071005 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00228 | 2008-01-10 | 20071016 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | INCREASED DEFIBRILATOR LEAD IMPEDANCE WITH 2 SHOCKS DELIVERED SECONDARY TO OVERSENSING OF ARTIFACT IN THE LEAD WAS REPORTED. RV LEAD FRACTURE WAS REPORTED. THE LEAD WAS EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | Y | Y | LAY USER/PATIENT
2649622-2008-00229 | 2008-01-10 | 20071005 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2008-00230 | 2008-01-10 | 20071009 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT RECEIVED NUMEROUS INAPPROPRIATE SHOCKS DUE TO 'DOUBLE COUNTING'. THE LEAD WAS EXPLANTED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED IN A LAWSUIT ALLEGING THE PATIENT 'SUFFERED PHYSICAL AND OTHER INJURIES', AND THAT THE PATIENT 'EXPERIENCED INPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE, OR OTHER INJURY REQUIRING MEDICAL INTERVENTION. |  | Y | N | LAY USER/PATIENT
2649622-2008-00231 | 2008-01-10 | 20071002 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00214 | 2008-01-10 | 20071012 | Death | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE RV (RIGHT VENTRICULAR) LEAD WAS EXPLANTED AND REPLACED DUE TO INAPPROPRIATE SHOCKS, HIGH PACING IMPEDANCE OF GREATER THAN 2000 OHMS, OVERSENSING, AND LOSS OF CAPTURE. FURTHER INFORMATION FROM AN ATTORNEY ALLEGES THE PATIENT SUSTAINED INJURY AS A RESULT OF A DEFECTIVE LEAD. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. FURTHER ALLEGED THAT THE PATIENT "SUFFERED PHYSICAL AND OTHER INJURY", THAT THE PATIENT "EXPERIENCED INAPPROPRIATE AND/OR EXCESSIVE SHOCKING, FRACTURE OR OTHER INJURY REQUIRING MEDICAL INTERVENTION INCLUDING BUT NOT LIMITED TO SURGERY, REMOVAL AND/OR REPLACEMENT OF THE DEFECTIVE [LEAD]", AND THAT THE PATIENT "SUSTAINED AND WILL CONTINUE TO SUSTAIN SEVERE PHYSICAL INJURIES AND/OR DEATH, SEVERE EMOTIONAL DISTRESS, MENTAL ANGUISH, ECONOMIC LOSSES AND OTHER DAMAGES." |  | Y | Y | LAY USER/PATIENT
2649622-2008-00215 | 2008-01-10 | 20071011 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2008-00216 | 2008-01-10 | 20071014 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00217 | 2008-01-10 | 20070130 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00218 | 2008-01-10 | 20071015 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6948 | ASKU |  | IT WAS REPORTED THAT THE RIGHT VENTRICULAR LEAD WAS EXPLANTED DUE TO T-WAVE OVERSENSING. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED FROM AN ATTORNEY ALLEGING THAT 'DEFECTIVE' LEAD CAUSED THE PLAINTIFF TO 'SUSTAIN SEVERE, PERMANENT, AND POTENTIALLY LIFE-THREATENING PHYSICAL INJURY'. |  | Y | N | LAY USER/PATIENT
2649622-2008-00219 | 2008-01-10 | 20071004 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE SPRINT FIDELIS LEAD MODEL IS INCLUDED IN A FIELD ADVISORY. EVALUATION SUMMARY: DISTAL CONDUCTOR FRACTURED; DISTAL SEGMENT RETURNED AND ANALYZED. |  | Y | Y | LAY USER/PATIENT
2649622-2007-03237 | 2008-01-23 | 20071106 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
2649622-2008-00009 | 2008-01-23 | 20071015 | Injury | Manufacturer report | SPRINT FIDELIS | ASKU | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  |  |  | Y | Y | 0
2649622-2008-00010 | 2008-01-23 | 20071011 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | NA |  |  |  | Y | N | LAY USER/PATIENT
2649622-2008-00011 | 2008-01-23 | 20071107 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE RV LEAD WAS EXPLANTED DUE TO OVERSENSING, HIGH IMPEDANCE OF 2620 OHMS, AND AN APPARENT LEAD FRACTURE. A MEDWATCH WAS ALSO RECEIVED AND REPORTS THE APPARENT LEAD FRACTURE, HIGH IMPEDANCE, OVERSENSING, AND NOISE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2007-03236 | 2008-01-23 | 20071123 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE PATIENT ALERT SOUNDED. A CARELINK TRANSMITTAL INDICATED HIGH/INCREASED RV (RIGHT VENTRICULAR) LEAD IMPEDANCE FROM 450 OHMS TO 1500. OVERSENSING WAS ALSO NOTED. THE PATIENT DID NOT RECEIVE ANY INAPPROPRIATE SHOCKS. THE LEAD WAS EXPLANTED AND REPLACED. A 3500A WAS SUBSEQUENTLY RECEIVED REPORTING JUMP IN IMPEDANCE AND DEVICE REPROGRAMMING TO LONGER INTERVALS. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2007-03230 | 2008-01-23 | 20070831 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
2649622-2007-03231 | 2008-01-23 | 20071030 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | ADDITIONAL INFORMATION REPORTED THERE WAS NOISE ON THE LEAD WITH MULTIPLE INAPPROPRIATE SHOCKS. LEAD FRACTURE WAS SUSPECTED. THE LEAD WAS EXPLANTED AND REPLACED. |  | Y | Y | LAY USER/PATIENT
2649622-2007-03232 | 2008-01-23 | 20071101 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THAT THE PATIENT PRESENTED TO THE HOSPITAL AFTER RECEIVING MULTIPLE INAPPROPRIATE SHOCKS DUE TO OVERSENSING OF NOISE, TYPICAL OF LEAD FRACTURE. PACING IMPEDANCE WAS ALSO NOTED TO BE HIGH (>2000 OHMS). THE LEAD WAS REPLACED. |  | Y | Y | LAY USER/PATIENT
2649622-2007-03233 | 2008-01-23 | 20071012 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | Y | Y | LAY USER/PATIENT
2649622-2007-03234 | 2008-01-23 | 20071110 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | ASKU |  | IT WAS REPORTED THE PT PRESENTED TO THE ER WHERE THE ICD DEMONSTRATED INAPPROPRIATE SENSING OF ICD LEAD AND ICD FIRING. WE HAVE NO INFORMATION TO SUGGEST THE LEAD WAS REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | N | Y | LAY USER/PATIENT
2649622-2007-03235 | 2008-01-23 | 20070810 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | Y | LAY USER/PATIENT
988805 | 2008-01-25 | 20080114 | Injury | User Facility report | SPRINT FIDELIS LEAD | IMPLANTABLE CARDIAC LEAD | MEDTRONIC INC. |  |  |  | A MALE ADMITTED WITH DIAGNOSIS OF POSSIBLE LEAD FRACTURE. THE PATIENT HAD A HISTORY OF ISCHEMIC CARDIOMYOPATHY S/P PLACEMENT OF IMPLANTABLE CARDIOVERTER DEFIBRILLATOR FOR PRIMARY PREVENTION OF SUDDEN DEATH. PATIENT WENT FOR AN ELECTROPHYSIOLOGIC EVALUATION, PER CARDIOLOGIST, IT WAS NOTED THAT THE PATIENT HAD P TO P INTERVALS OF CIRCA 110 MSEC. HOWEVER, MORE IMPORTANTLY THE IMPEDENCE ON HIS RIGHT VENTRICULAR LEAD HAD CLIMBED FROM A BASELINE UP 484 TO 500 OHMS TO HIS CURRENT LEVEL OF 1600 OHMS. THE PATIENT ALSO REVEALED EPISODES OF SOB (SHORTNESS OF BREATH). THE PATIENT UNDERWENT LEAD REVISION AND REPLACEMENT IN 2008. NOTE: 10/15/2007 MEDTRONIC INC. RELEASED AN URGENT MEDICAL DEVICE INFORMATION LETTER REGARDING SPRINT FIDELIS LEADS MODELS 6930, 6931, 6948 AND 6949 WHICH THIS PATIENT HAD. THE LETTER WAS FOR VOLUNTARY RECALL OF ABOVE MODELS DUE TO REPORTS OF LEAD CHRONIC CONDUCTOR FRACTURES. |  | Y | N | HEALTH PROFESSIONAL
2649622-2008-00828 | 2008-01-30 | 20080117 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6949 | NA |  |  |  | N | * | LAY USER/PATIENT
2649622-2008-00827 | 2008-01-30 | 20071221 | Injury | Manufacturer report | SPRINT FIDELIS | IMPLANTABLE TACHY LEAD | MEDTRONIC PUERTO RICO, INC. | 6931 | ASKU |  | ASKU |  | Y | Y | LAY USER/PATIENT
MW5005168 | 2008-01-30 | 20080128 | Injury | Voluntary report | MEDTRONICS SPRINT FIDELIS LEAD | NONE | MEDTRONIC WORLD HEADQUARTERS | 694965ID |  |  | I HAD ONE OF THE MEDTRONIC RECALLED DEFIBRILLATOR LEADS. MY LEAD WAS SHOWING A HIGH RESISTANCE LEVEL IN THE COIL AREA OF THE LEAD. AT MEDTRONIC'S ADVICE, MY ELECTROPHYSIOLOGIST PERFORMED THE LEAD REPLACEMENT SURGERY IN 2008. I WAS HOSPITALIZED OVERNIGHT AND MISSED A TOTAL OF 5 DAYS OF WORK. DATES OF USE: 11 MONTHS 28 DAYS, 2007 - 2008. DIAGNOSIS OR REASON FOR USE: LEFT BUNDLE BRANCH BLOCK, LOW EJECTION FRACTION. CARDIOMYOPATHY, CARDIOMEGALY. |  | Y | Y | HEALTH PROFESSIONAL
MW5005207 | 2008-01-30 | 20080125 | Other | Voluntary report | SPRINT FIDELIS | LEAD | MEDTRONIC | 6949 |  |  | SPRINT FIDELIS LEAD MALFUNCTION AND IT WAS REPLACED. |  | N | N | HEALTH PROFESSIONAL
997138 | 2008-01-02 | 20071113 | Malfunction | User Facility report | SPRINT FIDELIS | LEAD, ICD | MEDTRONIC, INC | 6949 | * |  | ICD LEAD REPLACEMENT PER THE RECALL. |  | N | Y | INVALID DATA
999489 | 2008-01-07 | 20071218 | Malfunction | User Facility report | SPRINT FIDELIS | LEAD, ICD | MEDTRONIC INC. | 694965 | * |  | ELDERLY MALE WITH ISCHEMIC CARDIOMYOPATHY, SEVERE END-STAGE CORONARY ARTERY DISEASE, ANGINA PECTORIS, CONGESTIVE HEART FAILURE, AND AN INDWELLING BIVENTRICULAR PACEMAKER THAT WAS SHOWN TO HAVE A PACEMAKER MALFUNCTION WITH LIFE-THREATENING BRADYDYSRHYTHMIAS AND NUMEROUS INAPPROPRIATE SHOCKS. HE UNDERWENT EMERGENCY PROGRAMMING CHANGES, WHICH TEMPORIZED THE SITUATION TO A POINT THAT HE COULD GET TO THE OPERATING SUITE IN STABLE CONDITION. THE SURGICAL PROCEDURE WAS EXTREMELY COMPLEX AND RESULTED IN THE SURGEON NOT BEING ABLE TO REMOVE THE LEAD ENTIRELY DUE TO RIGHT VENTRICULAR LEAD BEING ACTIVELY FIXATED TO THE VENTRICULAR SEPTUM. THE GENERATOR AND LEAD FRAGMENTS WERE REMOVED FROM THE LEFT DUE TO THE INABILITY TO FULLY CANNULATE TO THE RIGHT. THE GENERATOR WAS THEREFORE REPLACED ALSO. |  | N | Y | INVALID DATA
999497 | 2008-01-07 | 20071219 | Malfunction | User Facility report | SPRINT FIDELIS | LEAD, ICD | MEDTRONIC INC. | 6949 | * |  | ELDERLY MALE HAD A BIVENTRICULAR ICD IMPLANTED FOR CONGESTIVE HEART FAILURE AND DILATED ISCHEMIC CARDIOMYOPATHY. HE REQUIRED LEFT ANTERIOR THORACOTOMY FOR PLACEMENT OF LEFT VENTRICULAR PACING ELECTRODE. THE PATIENT DID WELL AND HAS BEEN MAINTAINED ON A CONGESTIVE FAILURE MEDICAL REGIMEN. HOWEVER, 8 MONTHS LATER HE HAD 4 ICD DISCHARGES. THESE WERE INAPPROPRIATE WHEN ASSESSED IN THE HOSPITAL EMERGENCY ROOM. PER THE SURGEON THE PATIENT HAD A RECALLED RV DUAL-COILED DEFIBRILLATOR ELECTRODE. HE WOULD APPEAR TO HAVE A LEAD FRACTURE IN THIS RECALL LEAD AND THEREFORE HE REQUIRED READMISSION TO THE HOSPITAL FOR LEAD EXTRACTION AND REINSERTION. THE DEFIBRILLATOR WAS DEACTIVATED BY A COMPANY REPRESENTATIVE. THE NEXT DAY, THE RIGHT VENTRICULAR LEAD WAS DISSECTED UP TO ITS INSERTION POINT INTO THE SUBCLAVIAN VEIN. THROUGH THE EXISTING WOUND, THE SURGEON CANNULATED THE LEFT SUBCLAVIAN VEIN WITHOUT DIFFICULTY AND PASSED THE GUIDEWIRE UNDER FLUOROSCOPY INTO THE RIGHT ATRIUM. A STYLET WAS THEN PASSED DOWN THE CHRONIC, RECALLED, DUAL COIL RV LEAD. THE ACTIVE FIXATION DEVICE WAS WITHDRAWN. USING CONSTANT GENERAL TRACTION AND SOME ROTATIONAL TORSION, THE LEAD WAS REMOVED IN ITS ENTIRETY WITHOUT DISRUPTION OF THE RIGHT ATRIAL CHRONIC ELECTRODE AND WITHOUT DIFFICULTY WHATSOEVER. |  | N | Y | INVALID DATA
1010520 | 2008-02-07 | 20080117 | Malfunction | User Facility report | MAXIMO | IMPLANTED CARDIAC DEFIBRILLATOR (ICD) | MEDTRONIC, INC. | VR7232 CX | * |  | THE PATIENT WAS EXPERIENCING POOR LEAD SENSING WITH THE MAXIMO SINGLE CHAMBER PACEMAKER. NO SYMPTOMS AT THIS TIME, HOWEVER, DUE TO THE SEVERITY OF THE PATIENT'S PAST V-TACH, THE PHYSICIAN FELT THE NEED TO REPLACE THE DEVICE. |  | N | Y | INVALID DATA
1011229 | 2008-01-25 | 20071217 | Injury | User Facility report | FIDELIS LEAD | LEAD, ICD | MEDTRONIC INC. CARDIAC RHYTHM MANAGEMENT | 6949 | * |  | PATIENT EXPERIENCING A PRODUCTIVE COUGH, LOWER EXTREMITY EDEMA, AND DYSPNEA WITH EXERTION. IT WAS DETERMINED THAT LEAD ON INTERNAL DEFIBRILLATOR HAD NON-PHYSIOLOGIC SENSING. UPON REPLACEMENT, LEAD WAS FOUND FRACTURED. |  | Y | Y | INVALID DATA
MW5007091 | 2008-05-31 | 20080512 | Injury | Voluntary report | SPRING FIDELIS | ICD LEAD | METRONIC, INC. | 6949 |  |  | PATIENT WITH A RIGHT MEDTRONIC ICD WITH THE SPRINT FIDELIS LEAD. THE LEAD FRACTURED PREMATURELY, TRIGGERING INAPPROPRIATE ICD SHOCKS. PATIENT RECEIVED LEAD REVISION AND HAD POST OP BLEEDING WHICH REQUIRED A POCKET REVISION. |  | N | Y | HEALTH PROFESSIONAL
991334 | 2008-01-31 | 20080116 | Injury | User Facility report | FIDELIS LEAD | IMPLANTABLE VENTRICULAR LEAD | MEDTRONIC INC. | 694965 |  |  | FEMALE WITH HISTORY OF ISCHEMIC CARDIOMYOPATHY, CHRONIC ARTERIAL FIBRILLATION STATUS POST PLACEMENT OF AN AUTOMATIC IMPLANTABLE CARDIOVERTER DEFIBRILLATOR WHO PRESENTED FOR AN ELECTROPHYSIOLOGIC EVALUATION. INTERROGATION OF HER FIDELIS LEAD REVEALED IMPEDANCE OF 0.88. THIS THEN ROSE TO 6.88 AND THEN TO OVER 1000 WITH THE TREND SUGGESTING POSSIBLE IMPEDING LEAD FRACTURE. THE PT UNDERWENT DEFIB LEAD REPLACEMENT IN 2008. |  | Y | Y | HEALTH PROFESSIONAL
998986 | 2008-02-07 | 20080122 | Malfunction | User Facility report | FIDELIS | LEAD, PACEMAKER | MEDTRONIC, INC. | 6947 | * |  | THE PATIENT HAD A HISTORY OF HYPERTROPHIC CARDIOMYOPATHY. AFTER PLACEMENT OF THE FIDELIS RIGHT VENTRICULAR LEAD, THE PATIENT REPORTED EXPERIENCING MULTIPLE INAPPROPRIATE SHOCKS. THE LEAD WAS REPLACED. THIS DEVICE IS PART OF A WELL-KNOWN RECALL. |  | N | Y | INVALID DATA

## Monthly Summary
year_month | count
2007-12 | 4
2008-01 | 567
2008-02 | 11
2008-03 | 87
2008-04 | 36
2008-05 | 756
2008-06 | 5

## Event Types
Event Type | Count
Injury | 1356
Death | 78
Malfunction | 17
No answer provided | 13
Other | 2

## Brand Names
Brand Name | Count
SPRINT FIDELIS | 1447
SPRINT FIDELIS LEAD | 4
FIDELIS LEAD | 3
MEDTRONIC SPRINT FIDELIS | 2
MEDTRONIC/SPRINT FIDELIS | 1
SPRINT FIDELIS DEFIBRILLATOR LEAD | 1
VIRTUOSO | 1
MEDTRONIC SPRINT FIDELIS LEADS | 1
MEDTRONIC SPRINT FIDELIS MODEL #6945 | 1
SPRINT FIDELIS ICD LEADS | 1
MEDTRONICS SPRINT FIDELIS LEAD | 1
MAXIMO | 1
SPRING FIDELIS | 1
FIDELIS | 1

## Case Metadata
Case Name | MDL Number | Filing Date | Court | Status | Settlement | Data Window Start | Data Window End | Total Records | Search Queries Used
Medtronic Sprint Fidelis | 1905 | 2008-06-26 | D. Minnesota | Settled | $268 million (2010-2012) | 2007-12-26 | 2008-06-26 | 1466 | 3