## All Reports
report_number | date_received | date_of_event | event_type | report_source_code | brand_name | generic_name | manufacturer_d_name | model_number | catalog_number | device_class | event_description | manufacturer_narrative | adverse_event_flag | product_problem_flag | device_operator | manufacturer_g1_name
1043534-2011-00736 | 2011-11-03 | 20110815 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4248 |  | ALLEGEDLY REVISED DUE TO A LOOSE COMPONENT. |  | Y | Y | PHYSICIAN
1043534-2011-00737 | 2011-11-03 | 20110815 | Injury | Manufacturer report | CONSERVE(R) TOTAL HEAD W/BFH | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3801-4204 |  | ALLEGEDLY REMOVED DURING THE REVISION OF ANOTHER COMPONENT. |  | Y | Y | PHYSICIAN
1043534-2011-00656 | 2011-11-04 | 20100528 | Injury | Manufacturer report | CONSERVE(R) PLUS HA CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HA-5056 |  | ALLEGEDLY, AS REPORTED IN (B)(4) REGISTRY, REVISED DUE TO ASEPTIC LOOSENING SOCKET/STEM. |  | Y | Y | PHYSICIAN
1043534-2011-00658 | 2011-11-04 | 20101115 | Injury | Manufacturer report | CONSERVE(R) PLUS HA CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HA-4450 |  | CONCLUSION: NO CONCLUSION CAN BE DRAWN. THE COMPLAINT WAS REVIEWED. THE PRODUCT WAS NOT RETURNED. EVIDENCE THAT PRODUCT IN SPECIFICATION WHEN USED. |  | Y | Y | PHYSICIAN
1043534-2011-00654 | 2011-11-04 | 20101007 | Injury | Manufacturer report | CONSERVE(R) PLUS HA CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HA-4652 |  | INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. THE EVENT CODE IS ADDRESSED IN THE PACKAGE INSERT. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS IS THE SAME EVENT AS 1043534-2011-00655. THIS EVENT OCCURRED IN (B)(6). |  | Y | Y | PHYSICIAN
1043534-2011-00667 | 2011-11-04 | 20100820 | Injury | Manufacturer report | CONSERVE(R) SUPER FINISH FEMORAL | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3803-1046 |  | ALLEGEDLY, AS REPORTED IN (B)(4) REGISTRY, REVISED DUE TO PAIN. |  | Y | Y | PHYSICIAN
1043534-2011-00657 | 2011-11-04 | 20100528 | Injury | Manufacturer report | CONSERVE(R) SUPER FINISH FEMORAL | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3803-1050 |  | ALLEGEDLY, AS REPORTED IN (B)(4) REGISTRY, REVISED DUE TO ASEPTIC LOOSENING SOCKET/STEM. |  | Y | Y | PHYSICIAN
1043534-2011-00655 | 2011-11-04 | 20101007 | Injury | Manufacturer report | CONSERVE(R) SUPER FINISH FEMORAL | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3803-1046 |  | ALLEGEDLY, AS REPORTED IN (B)(4) REGISTRY, REVISED DUE TO PAIN; ADVERSE SOFT TISSUE REACTION TO PARTICLE DEBRIS. |  | Y | Y | PHYSICIAN
1043534-2011-00660 | 2011-11-04 | 20100322 | Injury | Manufacturer report | CONSERVE(R) PLUS HA CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HA-5258 |  | INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. THE EVENT CODE IS ADDRESSED IN THE PACKAGE INSERT. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS IS THE SAME EVENT AS 1043534-2011-00661. THIS EVENT OCCURRED IN (B)(6). |  | Y | Y | PHYSICIAN
1043534-2011-00659 | 2011-11-04 | 20101115 | Injury | Manufacturer report | CONSERVE(R) SUPER FINISH FEMORAL | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3803-1044 |  | ALLEGEDLY, AS REPORTED IN (B)(4) REGISTRY, REVISED DUE TO UNKNOWN REASONS. |  | Y | Y | PHYSICIAN
1043534-2011-00661 | 2011-11-04 | 20100322 | Injury | Manufacturer report | CONSERVE(R) A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-1052 |  | ALLEGEDLY, AS REPORTED IN (B)(4) REGISTRY, REVISED DUE TO PERIPROSTHETIC FRACTURE STEM; ADVERSE SOFT TISSUE REACTION TO PARTICLE DEBRIS. |  | Y | Y | PHYSICIAN
1043534-2011-00663 | 2011-11-04 | 20100518 | Injury | Manufacturer report | CONSERVE(R) SUPER FINISH FEMORAL | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3803-1052 |  | ALLEGEDLY, AS REPORTED IN (B)(4) REGISTRY, REVISED DUE TO UNKNOWN REASONS. |  | Y | Y | PHYSICIAN
1043534-2011-00668 | 2011-11-04 | 20100601 | Injury | Manufacturer report | CONSERVE(R) PLUS HA CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HA-5460 |  | INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. THE EVENT CODE IS ADDRESSED IN THE PACKAGE INSERT. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS IS THE SAME EVENT AS 1043534-2011-00669. THIS EVENT OCCURRED IN THE (B)(6). |  | Y | Y | PHYSICIAN
1043534-2011-00669 | 2011-11-04 | 20100601 | Injury | Manufacturer report | CONSERVE(R) SUPER FINISH FEMORAL | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3803-1054 |  | DEVICE #2: INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. THE EVENT CODE IS ADDRESSED IN THE PACKAGE INSERT. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS IS THE SAME EVENT AS 1043534-2011-00668. THIS EVENT OCCURRED IN THE (B)(6). |  | Y | Y | PHYSICIAN
1043534-2011-00670 | 2011-11-04 | 20100622 | Injury | Manufacturer report | CONSERVE(R) PLUS HA CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HA-5056 |  | CONCLUSION: NO DEVICE FAILURE. THE COMPLAINT WAS REVIEWED AND ANALYSIS SHOWED NO TREND. THE PACKAGE INSERT WAS ALSO REVIEWED. THE PRODUCT WAS NOT RETURNED. EVIDENCE THAT PRODUCT IN SPECIFICATION WHEN USED. |  | Y | Y | PHYSICIAN
1043534-2011-00675 | 2011-11-04 | 20101108 | Injury | Manufacturer report | CONSERVE(R) SUPER FINISH FEMORAL | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3803-1042 |  | ALLEGEDLY, AS REPORTED IN THE NATIONAL JOINT REGISTRY, REVISED DUE TO ASEPTIC LOOSENING SOCKET. |  | Y | Y | PHYSICIAN
1043534-2011-00665 | 2011-11-04 | 20100806 | Injury | Manufacturer report | CONSERVE(R) SUPER FINISH FEMORAL | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3803-1050 |  | ALLEGEDLY, AS REPORTED IN (B)(4) REGISTRY, REVISED DUE TO ASEPTIC LOOSENING STEM; DISLOCATION SUBLUXATION. |  | Y | Y | PHYSICIAN
1043534-2011-00666 | 2011-11-04 | 20100820 | Injury | Manufacturer report | CONSERVE(R) PLUS HA CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4656 |  | ALLEGEDLY, AS REPORTED IN (B)(4) REGISTRY, REVISED DUE TO PAIN. |  | Y | Y | PHYSICIAN
1043534-2011-00664 | 2011-11-04 | 20100806 | Injury | Manufacturer report | CONSERVE(R) PLUS HA CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HA-5056 |  | ALLEGEDLY, AS REPORTED IN (B)(4) REGISTRY, REVISED DUE TO ASEPTIC LOOSENING STEM; DISLOCATION SUBLUXATION. |  | Y | Y | PHYSICIAN
1043534-2011-00673 | 2011-11-04 | 20100622 | Injury | Manufacturer report | CONSERVE(R) SUPER FINISH FEMORAL | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3803-1054 |  | ALLEGEDLY, AS REPORTED IN THE NATIONAL JOINT REGISTRY, REVISED DUE TO ASEPTIC LOOSENING SOCKET. |  | Y | Y | PHYSICIAN
1043534-2011-00671 | 2011-11-04 | 20100622 | Injury | Manufacturer report | CONSERVE(R) SUPER FINISH FEMORAL | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3803-1050 |  | ALLEGEDLY, AS REPORTED IN THE NATIONAL JOINT REGISTRY, REVISED DUE TO ASEPTIC LOOSENING STEM; PAIN. |  | Y | Y | PHYSICIAN
1043534-2011-00674 | 2011-11-04 | 20101108 | Injury | Manufacturer report | CONSERVE(R) PLUS HA CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HA-4248 |  | INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. THE EVENT CODE IS ADDRESSED IN THE PACKAGE INSERT. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS IS THE SAME EVENT AS 1043534-2011-00675. THIS EVENT OCCURRED IN (B)(6). |  | Y | Y | PHYSICIAN
1043534-2011-00672 | 2011-11-04 | 20100622 | Injury | Manufacturer report | CONSERVE(R) PLUS HA CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HA-5460 |  | ALLEGEDLY, AS REPORTED IN (B)(4) REGISTRY, REVISED DUE TO ASEPTIC LOOSENING SOCKET. |  | Y | Y | PHYSICIAN
1043534-2011-00662 | 2011-11-05 | 20100518 | Injury | Manufacturer report | CONSERVE(R) PLUS HA CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HA-5258 |  | INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS IS THE SAME EVENT AS 1043534-2011-00663. THIS EVENT OCCURRED IN (B)(6). |  | Y | Y | PHYSICIAN
1043534-2011-00678 | 2011-11-08 | 20101221 | Injury | Manufacturer report | CONSERVE(R) PLUS HA CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HA-4046 |  | ALLEGEDLY, AS REPORTED IN THE NATIONAL JOINT REGISTRY, REVISED DUE TO ASEPTIC LOOSENING STEM; PAIN. |  | Y | Y | PHYSICIAN
1043534-2011-00681 | 2011-11-08 | 20100610 | Injury | Manufacturer report | CONSERVE(R) SUPER FINISH FEMORAL | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3803-1050 |  | ALLEGEDLY, AS REPORTED IN THE (B)(6) REGISTRY, REVISED DUE TO UNKNOWN REASONS. |  | Y | Y | PHYSICIAN
1043534-2011-00680 | 2011-11-08 | 20100610 | Injury | Manufacturer report | CONSERVE(R) PLUS HA CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HA-5056 |  | CONCLUSION: NO CONCLUSION CAN BE DRAWN. THE COMPLAINT WAS REVIEWED. THE PRODUCT WAS NOT RETURNED. |  | Y | Y | PHYSICIAN
1043534-2011-00682 | 2011-11-08 | 20100211 | Injury | Manufacturer report | CONSERVE(R) PLUS HA CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HA-5258 |  | ALLEGEDLY, AS REPORTED IN THE (B)(6) REGISTRY, REVISED DUE TO PAIN. |  | Y | Y | PHYSICIAN
1043534-2011-00686 | 2011-11-08 | 20110628 | Injury | Manufacturer report | CONSERVE(R) PLUS HA CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HA-5056 |  | CONCLUSION: NO DEVICE FAILURE. THE COMPLAINT WAS REVIEWED AND ANALYSIS SHOWED NO TREND. THE PACKAGE INSERT WAS REVIEWED. THE PRODUCT WAS NOT RETURNED. |  | Y | Y | PHYSICIAN
1043534-2011-00676 | 2011-11-08 | 20101202 | Injury | Manufacturer report | CONSERVE(R) PLUS HA CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HA-4450 |  | CONCLUSION: NO CONCLUSION CAN BE DRAWN. THE COMPLAINT AND PACKAGE INSERT WERE REVIEWED. THE PRODUCT WAS NOT RETURNED. |  | Y | Y | PHYSICIAN
1043534-2011-00679 | 2011-11-08 | 20101221 | Injury | Manufacturer report | CONSERVE(R) SUPER FINISH FEMORAL | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3803-1040 |  | ALLEGEDLY, AS REPORTED IN THE NATIONAL JOINT REGISTRY, REVISED DUE TO ASEPTIC LOOSENING STEM; PAIN. |  | Y | Y | PHYSICIAN
1043534-2011-00683 | 2011-11-08 | 20100211 | Injury | Manufacturer report | CONSERVE(R) PLUS SUPER FINISH FEMORAL | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-1052 |  | DEVICE #2:INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS IS THE SAME EVENT AS 1043534-2011-00682. THIS EVENT OCCURRED IN THE (B)(6). |  | Y | Y | PHYSICIAN
1043534-2011-00684 | 2011-11-08 | 20110427 | Injury | Manufacturer report | CONSERVE(R) PLUS HA CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HA-4046 |  | ALLEGEDLY, AS REPORTED IN THE (B)(6) REGISTRY, REVISED DUE TO PAIN; ADVERSE TISSUE REACTION TO PARTICLE DEBRIS. |  | Y | Y | PHYSICIAN
1043534-2011-00685 | 2011-11-08 | 20110427 | Injury | Manufacturer report | CONSERVE(R) PLUS SUPER FINISH FEMORAL | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-1040 |  | ALLEGEDLY, AS REPORTED IN THE (B)(6) REGISTRY, REVISED DUE TO PAIN; ADVERSE TISSUE REACTION TO PARTICLE DEBRIS. |  | Y | Y | PHYSICIAN
1043534-2011-00687 | 2011-11-08 | 20110628 | Injury | Manufacturer report | CONSERVE(R) SUPER FINISH FEMORAL | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3803-1050 |  | ALLEGEDLY, AS REPORTED IN THE (B)(6) REGISTRY, REVISED DUE TO ASEPTIC LOOSENING STEM. |  | Y | Y | PHYSICIAN
1043534-2011-00677 | 2011-11-08 | 20101202 | Injury | Manufacturer report | CONSERVE(R) SUPER FINISH FEMORAL | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3803-1044 |  | ALLEGEDLY, AS REPORTED IN THE NATIONAL JOINT REGISTRY, REVISED DUE TO ADVERSE SOFT TISSUE REACTION TO PARTICLE DEBRIS. |  | Y | Y | PHYSICIAN
1043534-2011-00689 | 2011-11-08 | 20110630 | Injury | Manufacturer report | CONSERVE(R) SUPER FINISH FEMORAL | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3803-1046 |  | DEVICE #2: INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. THE EVENT CODE IS ADDRESSED IN THE PACKAGE INSERT. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS IS THE SAME EVENT AS 1043534-2011-00688. THIS EVENT OCCURRED IN (B)(6). |  | Y | Y | PHYSICIAN
1043534-2011-00690 | 2011-11-08 | 20110602 | Injury | Manufacturer report | CONSERVE(R) PLUS HA CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HA-4450 |  | ALLEGEDLY, AS REPORTED IN THE (B)(4) REGISTRY, REVISED DUE TO ADVERSE SOFT TISSUE REACTION TO PARTICLE DEBRIS. |  | Y | Y | PHYSICIAN
1043534-2011-00695 | 2011-11-08 | 20110112 | Injury | Manufacturer report | CONSERVE(R) SUPER FINISH FEMORAL | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3803-1046 |  | ALLEGEDLY, AS REPORTED IN THE (B)(4) REGISTRY, REVISED DUE TO PAIN. |  | Y | Y | PHYSICIAN
1043534-2011-00699 | 2011-11-08 | 20110503 | Injury | Manufacturer report | CONSERVE(R) SUPER FINISH FEMORAL | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3803-1040 |  | CONCLUSION: NO CONCLUSION CAN BE DRAWN. THE COMPLAINT AND PACKAGE INSERT WERE REVIEWED. THE PRODUCT WAS NOT RETURNED. |  | Y | Y | PHYSICIAN
1043534-2011-00698 | 2011-11-08 | 20110503 | Injury | Manufacturer report | CONSERVE(R) PLUS HA CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HA-4046 |  | ALLEGEDLY, AS REPORTED IN THE (B)(4) REGISTRY, REVISED DUE TO ASEPTIC LOOSENING SOCKET; LYSIS SOCKET; PAIN. |  | Y | Y | PHYSICIAN
1043534-2011-00701 | 2011-11-08 | 20110225 | Injury | Manufacturer report | CONSERVE(R) SUPER FINISH FEMORAL | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3803-1042 |  | ALLEGEDLY, AS REPORTED IN THE (B)(4) REGISTRY, REVISED DUE TO ASEPTIC LOOSENING STEM. |  | Y | Y | PHYSICIAN
1043534-2011-00688 | 2011-11-08 | 20110630 | Injury | Manufacturer report | CONSERVE(R) PLUS HA CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HA-4652 |  | ALLEGEDLY, AS REPORTED IN THE (B)(4) REGISTRY, REVISED DUE TO ADVERSE SOFT TISSUE REACTION TO PARTICLE DEBRIS. |  | Y | Y | PHYSICIAN
1043534-2011-00691 | 2011-11-08 | 20110602 | Injury | Manufacturer report | CONSERVE(R) SUPER FINISH FEMORAL | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3803-1044 |  | ALLEGEDLY, AS REPORTED IN THE (B)(4) REGISTRY, REVISED DUE TO ADVERSE SOFT TISSUE REACTION TO PARTICLE DEBRIS. |  | Y | Y | PHYSICIAN
1043534-2011-00693 | 2011-11-08 | 20110127 | Injury | Manufacturer report | CONSERVE(R) SUPER FINISH FEMORAL | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3803-1048 |  | ALLEGEDLY, AS REPORTED IN THE (B)(4) REGISTRY, REVISED DUE TO ADVERSE SOFT TISSUE REACTION TO PARTICLE DEBRIS. |  | Y | Y | PHYSICIAN
1043534-2011-00694 | 2011-11-08 | 20110112 | Injury | Manufacturer report | CONSERVE(R) PLUS HA CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HA-4652 |  | ALLEGEDLY, AS REPORTED IN THE (B)(4) REGISTRY, REVISED DUE TO PAIN. |  | Y | Y | PHYSICIAN
1043534-2011-00692 | 2011-11-08 | 20110127 | Injury | Manufacturer report | CONSERVE(R) PLUS HA CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HA-4854 |  | ALLEGEDLY, AS REPORTED IN THE (B)(4) REGISTRY, REVISED DUE TO ADVERSE SOFT TISSUE REACTION TO PARTICLE DEBRIS. |  | Y | Y | PHYSICIAN
1043534-2011-00696 | 2011-11-08 | 20110208 | Injury | Manufacturer report | CONSERVE(R) PLUS HA CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HA-4652 |  | ALLEGEDLY, AS REPORTED IN THE (B)(4) REGISTRY, REVISED DUE TO UNKNOWN REASONS. |  | Y | Y | PHYSICIAN
1043534-2011-00700 | 2011-11-08 | 20110225 | Injury | Manufacturer report | CONSERVE(R) PLUS HA CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HA-4248 |  | ALLEGEDLY, AS REPORTED IN THE (B)(4) REGISTRY, REVISED DUE TO ASEPTIC LOOSENING STEM. |  | Y | Y | PHYSICIAN
1043534-2011-00697 | 2011-11-08 | 20110208 | Injury | Manufacturer report | CONSERVE(R) SUPER FINISH FEMORAL | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3803-1046 |  | ALLEGEDLY, AS REPORTED IN THE (B)(4) REGISTRY, REVISED DUE TO UNKNOWN REASONS. |  | Y | Y | PHYSICIAN
1043534-2011-00703 | 2011-11-08 | 20110322 | Injury | Manufacturer report | CONSERVE(R) SUPER FINISH FEMORAL | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3803-1050 |  | ALLEGEDLY, AS REPORTED IN (B)(4) REGISTRY, REVISED DUE TO ADVERSE SOFT TISSUE REACTION TO PARTICLE DEBRIS. |  | Y | Y | PHYSICIAN
1043534-2011-00702 | 2011-11-08 | 20110322 | Injury | Manufacturer report | CONSERVE(R) PLUS HA CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HA-5056 |  | CONCLUSION: NO CONCLUSION CAN BE DRAWN. THE COMPLAINT AND PACKAGE INSERT WERE REVIEWED. THE PRODUCT WAS NOT RETURNED. EVIDENCE THAT PRODUCT IN SPECIFICATION WHEN USED. |  | Y | Y | PHYSICIAN
1043534-2011-00767 | 2011-11-17 | 20111011 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-5664 |  | ALLEGEDLY REVISED DUE TO INFECTION. |  | Y | Y | PHYSICIAN
1043534-2011-00768 | 2011-11-17 | 20111011 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-5635 |  | ALLEGEDLY REVISED DUE TO INFECTION. |  | Y | Y | PHYSICIAN
1043534-2011-00772 | 2011-11-22 | 20110124 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-4800 |  | INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. THE EVENT CODE IS ADDRESSED IN THE PACKAGE INSERT. ALTHOUGH SEVERAL ATTEMPTS HAVE BEEN MADE, NO MEDWATCH 3500 HAS BEEN RECEIVED FROM THE USER FACILITY. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS IS THE SAME EVENT AS 1043534-2011-00773. |  | Y | Y | PHYSICIAN
1043534-2011-00773 | 2011-11-22 | 20110124 | Injury | Manufacturer report | CONSERVE(R) PLUS SPIKED CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38SP-4854 |  | CONCLUSION: NO CONCLUSION CAN BE DRAWN. THE COMPLAINT WAS REVIEWED. THE PRODUCT WAS NOT RETURNED. |  | Y | Y | PHYSICIAN
1043534-2011-00785 | 2011-11-23 | 20111024 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-5200 |  | ADDITIONAL INFORMATION RECEIVED 12/12/2011:PLEASE SEE THE MEDWATCH 3500A RECEIVED FROM THE USER FACILITY. THIS IS THE SAME EVENT AS 1043534-2011-00783, 00784, 00786. |  | Y | Y | PHYSICIAN
1043534-2011-00786 | 2011-11-23 | 20111024 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-5258 |  | ALLEGEDLY REMOVED DURING THE REVISION OF ANOTHER COMPONENT. |  | Y | Y | PHYSICIAN
1043534-2011-00788 | 2011-11-23 | 20100204 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-4435 |  | ALLEGEDLY REMOVED DURING THE REVISION OF ANOTHER COMPONENT. |  | Y | Y | PHYSICIAN
1043534-2011-00794 | 2011-11-28 | 20111026 | Injury | Manufacturer report | CONSERVE(R) TOTAL HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38CH-4200 |  | ALLEGEDLY THE PATIENT COMPLAINED OF PAIN. THIS COMPONENT WAS REMOVED DURING THE REVISION OF ANOTHER COMPONENT. |  | Y | Y | PHYSICIAN
1043534-2011-00797 | 2011-11-28 | 20111025 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-4600 |  | ALLEGEDLY REVISED DUE TO PAIN. |  | Y | Y | PHYSICIAN
1043534-2011-00796 | 2011-11-28 | 20111025 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4652 |  | ALLEGEDLY REVISED DUE TO PAIN. |  | Y | Y | PHYSICIAN
1043534-2011-00804 | 2011-11-29 | 20100319 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-5235 |  | CONCLUSION: NO DEVICE FAILURE. COMPLAINT HISTORY REVIEWED AND ANALYSIS SHOWS NO TREND FOR ITEM/LOT NUMBER. DEVICE HISTORY RECORD REVIEWED AND INDICATES THE PRODUCT MET SPECIFICATIONS WHEN MANUFACTURED. PRODUCT NOT RETURNED FOR EVALUATION. THERE WAS NO SPECIFIC COMPLAINT STATED AGAINST THIS COMPONENT AND REMOVAL OF THIS COMPONENT WOULD BE TYPICAL DURING THIS TYPE OF REVISION SURGERY. PRODUCT DID NOT CONTRIBUTE TO THIS EVENT. |  | Y | N | PHYSICIAN
2364060 | 2011-12-04 | 20111024 | Malfunction | User Facility report | CONSERVE TOTAL FEMORAL HEAD | BALL WITH BROKEN PEICE OF FEMORAL NECK | WRIGHT MEDICAL TECHNOLOGY, INC. | 52MM MEDIUM NECK | 38AM5200 |  | MALE PT WITH MULTIPLE MEDICAL COMORBIDITIES STATUS POST CARDIAC STENT PLACEMENT TIMES 2 FOUR WEEKS PRIOR TO PRESENTATION, WAS BENDING OVER TODAY TO PICK UP A COVER FOR HIS SNOWMOBILE WHEN HE FELT HIS LEFT HIP GIVE WAY. HE IS STATUS POST A LEFT TOTAL HIP ARTHROPLASTY IN 2006 AND HAD BEEN DOING WELL UNTIL YESTERDAY. HE DID, HOWEVER, ENDORSE A BURNING SENSATION IN HIS LEFT HIP WITH WALKING OVER THE PAST FEW MONTHS. HE DENIES ANY FEVERS, CHILLS OR RECENT ILLNESS. HE WAS EVALUATED IN AN URGENT CARE AND THEN DECIDED TO DRIVE TO THE HOSPITAL FOR FURTHER CARE. AFTER HIS HIP GAVE WAY HE NOTICED SIGNIFICANT HIP PAIN AND INABILITY TO STAND UPRIGHT. |  | N | Y | INVALID DATA
1043534-2011-00822 | 2011-12-08 | 20110919 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-5035 |  | CORRECTED DATA RECEIVED (B)(6) 2012: REVISION DATE (B)(6) 2009. AGE OF DEVICE 3.5 YEARS. ADDITIONAL DATA RECEIVED (B)(6) 2012: IN REVIEW OF THE FILES, THIS WAS DETERMINED TO BE A DUPLICATE OF MEDICAL DEVICE REPORT 1043534-2009-00332 PREVIOUSLY SUBMITTED AND THEREFORE NOT REQUIRED. |  | Y | Y | PHYSICIAN
1043534-2011-00823 | 2011-12-08 | 20110919 | Injury | Manufacturer report | CONSERVE(R) PLUS SPIKED CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38SP-5056 |  | ALLEGEDLY REVISED DUE TO PATIENT REACTION TO MOM. |  | Y | Y | PHYSICIAN
1043534-2011-00833 | 2011-12-09 | 20111102 | Injury | Manufacturer report | CONSERVE(R) TOTAL BCH FEMORAL HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38CH-3800 |  | ALLEGEDLY REMOVED DURING THE REVISION OF ANOTHER COMPONENT. |  | Y | Y | PHYSICIAN
1043534-2011-00832 | 2011-12-09 | 20111102 | Injury | Manufacturer report | CONSERVE(R) TOTAL NECK SLEEVE | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38NS-0035 |  | ALLEGEDLY REMOVED DURING THE REVISION OF ANOTHER COMPONENT. |  | Y | Y | PHYSICIAN
1043534-2011-00836 | 2011-12-10 | 20111107 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4652 |  | ALLEGEDLY REVISED DUE TO INFECTION AND SLIPPAGE. |  | Y | Y | PHYSICIAN
1043534-2011-00839 | 2011-12-10 | 20111107 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4854 |  | ALLEGEDLY REVISED DUE TO PAIN AND DISLOCATION. |  | Y | Y | PHYSICIAN
1043534-2011-00840 | 2011-12-10 | 20111107 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS(R) FEMORAL HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AC-4800 |  | ALLEGEDLY REMOVED DURING THE REVISION OF ANOTHER COMPONENT. |  | Y | Y | PHYSICIAN
1043534-2011-00842 | 2011-12-10 | 20111107 | Injury | Manufacturer report | CONSERVE(R) TOTAL SHORT NECK SLEEVE | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38NS-0N35 |  | DEVICE #4:INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. NO COMPLAINT WAS STATED AGAINST THIS DEVICE. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS IS THE SAME EVENT AS 1043534-2011-00839, 00840, 00841. |  | Y | Y | PHYSICIAN
1043534-2011-00837 | 2011-12-12 | 20111107 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-4604 |  | ALLEGEDLY REVISED DUE TO INFECTION. |  | Y | Y | PHYSICIAN
1043534-2011-00846 | 2011-12-12 | 20111031 | Injury | Manufacturer report | CONSERVE(R) PLUS SPIKED CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38SP-5460 |  | INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. THE EVENT CODE IS ADDRESSED IN THE PACKAGE INSERT. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS IS THE SAME EVENT AS 1043534-2011-00847, 00848. |  | Y | Y | PHYSICIAN
1043534-2011-00847 | 2011-12-12 | 20111031 | Injury | Manufacturer report | CONSERVE(R) TOTAL HEAD W/BFH | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3801-5404 |  | ALLEGEDLY REVISED DUE TO PAIN. |  | Y | Y | PHYSICIAN
1043534-2011-00852 | 2011-12-13 | 20111107 | Injury | Manufacturer report | CONSERVE(R) TOTAL HEAD W/BFH | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3801-5635 |  | ALLEGEDLY REMOVED DURING THE REVISION OF ANOTHER COMPONENT. |  | Y | Y | PHYSICIAN
1043534-2011-00855 | 2011-12-14 | 20111121 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-5204 |  | DEVICE #2: INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. THE EVENT CODE IS ADDRESSED IN THE PACKAGE INSERT. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS IS THE SAME EVENT AS 1043534-2011-00854. |  | Y | Y | PHYSICIAN
1043534-2011-00860 | 2011-12-15 | 20111009 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-5200 |  | ALLEGEDLY ON (B)(6) 2011, PATIENT WAS SEEN IN THE CLINIC AND FOUND TO BE DOING WELL. ON (B)(6) 2011, HE HAD A SUDDEN ONSET OF PAIN IN HIS HIP AND WAS SEEN IN THE ER. HE WAS FOUND TO HAVE AN IRRITABLE HIP. AN ASPIRATION REVEALED WHITE BLOOD CELLS AND GRAM-POSITIVE COCCI. WORKUP ON THE MEDICINE SERVICE SUGGESTED THAT HE HAD POSITIVE BLOOD CULTURES. REVISED DUE TO INFECTION. |  | Y | Y | PHYSICIAN
1043534-2011-00859 | 2011-12-15 | 20111009 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-5258 |  | ALLEGEDLY ON (B)(6) 2011, PATIENT WAS SEEN IN THE CLINIC AND FOUND TO BE DOING WELL. ON (B)(6) 2011, HE HAD A SUDDEN ONSET OF PAIN IN HIS HIP AND WAS SEEN IN THE ER. HE WAS FOUND TO HAVE AN IRRITABLE HIP. AN ASPIRATION REVEALED WHITE BLOOD CELLS AND GRAM-POSITIVE COCCI. WORKUP ON THE MEDICINE SERVICE SUGGESTED THAT HE HAD POSITIVE BLOOD CULTURES. REVISED DUE TO INFECTION. |  | Y | Y | PHYSICIAN
1043534-2011-00863 | 2011-12-15 | 20091215 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-4435 |  | ALLEGEDLY HIP NEVER FULLY HEALED. ADVISES MUSCLE TISSUE DAMAGE AND DISCOMFORT STILL. SEEKING SETTLEMENT. |  | Y | Y | PHYSICIAN
1043534-2011-00862 | 2011-12-15 | 20091215 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4450 |  | ALLEGEDLY HIP CUP WAS LOOSE AND NEVER FULLY HEALED. ADVISES MUSCLE TISSUE DAMAGE AND DISCOMFORT STILL. SEEKING SETTLEMENT. |  | Y | Y | PHYSICIAN
1043534-2011-00866 | 2011-12-15 | 20110725 | Injury | Manufacturer report | CONSERVE(R) TOTAL HEAD W/BFH | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3801-5200 |  | ALLEGEDLY COMPONENT FAILED AND PATIENT HAD INFECTION AND REQUIRED A REVISION SURGERY. |  | Y | Y | PHYSICIAN
1043534-2011-00868 | 2011-12-15 | 20110718 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-5258 |  | ALLEGEDLY CUP WAS REVISED DUE TO INFECTION. |  | Y | Y | PHYSICIAN
1043534-2011-00880 | 2011-12-21 | 20091021 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4652 |  | ALLEGEDLY REVISED DUE TO MOM COMPLICATIONS. |  | Y | Y | PHYSICIAN
1043534-2011-00882 | 2011-12-21 | 20091021 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-4600 |  | ALLEGEDLY REVISED DUE TO MOM COMPLICATIONS. |  | Y | Y | PHYSICIAN
1043534-2011-00885 | 2011-12-22 | 20111115 | Injury | Manufacturer report | CONSERVE(R) TOTAL HEAD W/BFH | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3801-5000 |  | ALLEGEDLY REMOVED DURING THE REVISION OF ANOTHER COMPONENT. |  | Y | Y | PHYSICIAN
1043534-2011-00893 | 2011-12-22 | 20110922 | Injury | Manufacturer report | CONSERVE(R) PLUS HA CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HA-4652 |  | ALLEGEDLY PROBLEMATIC SINCE INSTALLATION. PER MEDWATCH REPORT # (B)(4). |  | Y | Y | PHYSICIAN
1043534-2011-00894 | 2011-12-22 | 20110922 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-4635 |  | CONCLUSION: NO CONCLUSION CAN BE DRAWN. THE COMPLAINT AND PACKAGE INSERT WERE REVIEWED. THE PRODUCT WAS NOT RETURNED. |  | Y | Y | PHYSICIAN
1043534-2011-00895 | 2011-12-22 | 20111013 | Injury | Manufacturer report | CONSERVE(R) TOTAL HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3801-4400 |  | ADDITIONAL INFORMATION RECEIVED 12/23/2011: THE PATIENT SENT THE CATALOG/LOT NUMBER. THIS IS THE SAME EVENT AS 1043534-2011-00896, 00898. ALSO, THE USER FACILITY ADVISES NO MEDWATCH WAS FILED FOR THIS EVENT. |  | Y | Y | PHYSICIAN
1043534-2011-00898 | 2012-01-03 | 20111013 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-1154 |  | ALLEGEDLY REVISED DUE TO PATIENT REACTION. PER MEDWATCH REPORT (B)(4). |  | Y | Y | PHYSICIAN
1043534-2011-00902 | 2012-01-04 | 20111011 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS(R) FEMORAL HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AC-3800 |  | ALLEGEDLY REMOVED DURING THE REVISION OF ANOTHER COMPONENT. |  | Y | Y | PHYSICIAN
1043534-2011-00903 | 2012-01-04 | 20111011 | Injury | Manufacturer report | CONSERVE(R) TOTAL NECK SLEEVE | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38NS-0035 |  | ALLEGEDLY REMOVED DURING THE REVISION OF ANOTHER COMPONENT. |  | Y | Y | PHYSICIAN
1043534-2011-00905 | 2012-01-04 | 20110912 | Injury | Manufacturer report | CONSERVE(R) TOTAL HEAD W/BFH | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3801-4604 |  | DEVICE #2:INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. NO COMPLAINT HAS BEEN STATED AGAINST THIS DEVICE. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS IS THE SAME EVENT AS 1043534-2011-00904, 00906. THIS EVENT OCCURRED IN (B)(6). |  | Y | Y | PHYSICIAN
1043534-2011-00906 | 2012-01-04 | 20110912 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4652 |  | ALLEGEDLY REMOVED DURING THE REVISION OF ANOTHER COMPONENT. |  | Y | Y | PHYSICIAN
1043534-2012-00003 | 2012-01-06 | 20091116 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4450 |  | CONCLUSION: NO CONCLUSION CAN BE DRAWN. THE COMPLAINT AND PACKAGE INSERT WERE REVIEWED. THE PRODUCT WAS NOT RETURNED. |  | Y | Y | PHYSICIAN
1043534-2012-00004 | 2012-01-06 | 20091116 | Injury | Manufacturer report | CONSERVE(R) TOTAL HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3801-4400 |  | ALLEGEDLY REMOVED DURING THE REVISION OF ANOTHER COMPONENT. |  | Y | Y | PHYSICIAN
1043534-2012-00014 | 2012-01-09 | 20111208 | Injury | Manufacturer report | CONSERVE(R) TOTAL HEAD W/BFH | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3801-5035 |  | ALLEGEDLY REMOVED DURING THE REVISION OF ANOTHER COMPONENT. |  | Y | Y | PHYSICIAN
1043534-2012-00022 | 2012-01-10 |  | Injury | Manufacturer report | CONSERVE(R) BFH HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38XX-XXXX |  | ALLEGEDLY REVISED DUE TO SUPECTED TISSUE REACTION. |  | Y | Y | PHYSICIAN
1043534-2012-00023 | 2012-01-10 |  | Injury | Manufacturer report | CONSERVE(R) CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38XX-XXXX |  | CONCLUSION: NO CONCLUSION CAN BE DRAWN. THE PRODUCT WAS NOT RETURNED. |  | Y | Y | PHYSICIAN
1043534-2012-00024 | 2012-01-10 |  | Injury | Manufacturer report | CONSERVE(R) FEMORAL COMPONENT | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38PF-XXXX |  | ALLEGEDLY REVISED DUE TO BONE FRACTURE. |  | Y | Y | PHYSICIAN
1043534-2012-00029 | 2012-01-12 | 20100906 | Injury | Manufacturer report | CONSERVE(R) TOTAL HEAD W/BFH | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3801-5204 |  | CONCLUSION: NO CONCLUSION CAN BE DRAWN. COMPLAINT REVIEWED AND ANALYSIS SHOWED NO TREND FOR (B)(4), LOT NO: 016306330. DEVICE HISTORY RECORD REVIEWED. PRODUCT NOT RETURNED. EVIDENCE THAT PRODUCT IN SPECIFICATION WHEN USED. |  | Y | Y | PHYSICIAN
1043534-2012-00030 | 2012-01-12 | 20100906 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-5258 |  | DEVICE #4: INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. NO COMPLAINT WAS STATED AGAINST THIS DEVICE. ALTHOUGH SEVERAL ATTEMPTS HAVE BEEN MADE, NO MEDWATCH 3500A HAS BEEN RECEIVED FROM THE USER FACILITY. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS IS THE SAME EVENT AS 1043534-2012-00027, 00028, 00029. |  | Y | Y | PHYSICIAN
1043534-2012-00033 | 2012-01-13 | 20111212 | Injury | Manufacturer report | CONSERVE(R) TOTAL HEAD W/BFH | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3801-5200 |  | ALLEGEDLY REMOVED DURING THE REVISION OF ANOTHER COMPONENT. |  | Y | Y | PHYSICIAN
1043534-2012-00044 | 2012-01-18 | 20070510 | Injury | Manufacturer report | CONSERVE(R) PLUSS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-5258 |  | ALLEGEDLY REVISED DUE TO LOOSENESS. |  | Y | Y | PHYSICIAN
1043534-2012-00045 | 2012-01-18 | 20091019 | Injury | Manufacturer report | CONSERVE(R) PLUS SPIKED CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38SP-5258 |  | INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. THE EVENT CODE IS ADDRESSED IN THE PACKAGE INSERT. ALTHOUGH SEVERAL ATTEMPTS HAVE BEEN MADE, NO MEDWATCH 3500 HAS BEEN RECEIVED FROM THE USER FACILITY. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS IS THE SAME EVENT AS 1043534-2012-00046. |  | Y | Y | PHYSICIAN
1043534-2012-00046 | 2012-01-18 | 20091019 | Injury | Manufacturer report | CONSERVE(R) PLUS SUPER FINISH FEMORAL | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-1052 |  | DEVICE #2:INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. THE EVENT CODE IS ADDRESSED IN THE PACKAGE INSERT. ALTHOUGH SEVERAL ATTEMPTS HAVE BEEN MADE, NO MEDWATCH 3500 HAS BEEN RECEIVED FROM THE USER FACILITY. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS IS THE SAME EVENT AS 1043534-2012-00045. |  | Y | Y | PHYSICIAN
1043534-2012-00057 | 2012-01-19 |  | Injury | Manufacturer report | CONSERVE(R) PLUS HIP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  |  |  | INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. ALTHOUGH SEVERAL ATTEMPTS HAVE BEEN MADE, NO MEDWATCH 3500 OR ADDITIONAL INFORMATION HAS BEEN RECEIVED FROM THE REPORTER. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. |  | Y | Y | PHYSICIAN
1043534-2012-00064 | 2012-01-23 | 20111116 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4854 |  | INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. THE EVENT CODE IS ADDRESSED IN THE PACKAGE INSERT. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS IS THE SAME EVENT AS 1043534-2012-00065, 00066, 00067. |  | Y | Y | PHYSICIAN
1043534-2012-00065 | 2012-01-23 | 20111116 | Injury | Manufacturer report | CONSERVE(R) TOTAL HEAD W/BFH | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3801-4804 |  | DEVICE #2:INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. THE EVENT CODE IS ADDRESSED IN THE PACKAGE INSERT. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS IS THE SAME EVENT AS 1043534-2012-00064, 00066, 00067. |  | Y | Y | PHYSICIAN
1043534-2012-00072 | 2012-01-24 | 20111215 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-3604 |  | ALLEGEDLY REVISED DUE TO INFECTION. |  | Y | Y | PHYSICIAN
1043534-2012-00084 | 2012-01-25 | 20111207 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4248 |  | ALLEGEDLY, THE CUP TORE LOOSE FROM THE ACETABULUM. THE CUP'S ANGLE WAS ABOUT 90 DEGREES. REVISION SURGERY WAS PERFORMED AND THE SURGEON COULDN'T FIND ANY METAL DEBRIS. |  | Y | Y | PHYSICIAN
1043534-2012-00085 | 2012-01-25 | 20111207 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-4235 |  | CONCLUSION: NO DEVICE FAILURE. VISUALLY EXAMINED. DIMENSIONAL INSPECTION. PHOTOGRAPHIC IMAGES MADE. COMPLAINT REVIEWED AND ANALYSIS SHOWED NO TREND FOR 38AM-4235 LOT NO: 088690237. DEVICE HISTORY RECORD REVIEWED. UNDETERMINED. |  | Y | Y | PHYSICIAN
1043534-2012-00091 | 2012-01-26 | 20110203 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-5000 |  | ALLEGEDLY REVISED DUE TO SUSPECTED TISSUE REACTION. |  | Y | Y | PHYSICIAN
1043534-2012-00088 | 2012-01-26 | 20101209 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-4635 |  | INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. THE EVENT CODE IS ADDRESSED IN THE PACKAGE INSERT. ALTHOUGH SEVERAL ATTEMPTS HAVE BEEN MADE, NO MEDWATCH 3500 HAS BEEN RECEIVED FROM THE USER FACILITY. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS IS THE SAME EVENT AS 1043534-2012-00087. |  | Y | Y | PHYSICIAN
1043534-2012-00089 | 2012-01-26 | 20080101 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4652 |  | ALLEGEDLY REVISED DUE TO LOOSENESS. |  | Y | Y | PHYSICIAN
1043534-2012-00092 | 2012-01-26 | 20110203 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3208-5056 |  | CONCLUSION: NO DEVICE FAILURE COMPLAINT REVIEWED. DEVICE HISTORY RECORD REVIEWED. PACKAGE INSERT REVIEWED. PRODUCT NOT RETURNED. (B)(4) - EVIDENCE THAT PRODUCT IN SPECIFICATION WHEN USED. |  | Y | Y | PHYSICIAN
1043534-2012-00093 | 2012-01-26 | 20101201 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-5056 |  | ALLEGEDLY REVISED DUE TO LOOSENESS. |  | Y | Y | PHYSICIAN
1043534-2012-00087 | 2012-01-26 | 20101209 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4652 |  | ALLEGEDLY REVISED DUE TO SUSPECTED TISSUE REACTION. |  | Y | Y | PHYSICIAN
1043534-2012-00097 | 2012-01-26 | 20100629 | Injury | Manufacturer report | CONSERVE(R) PLUS SPIKED CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38SP-4652 |  | CONCLUSION: NO CONCLUSION CAN BE DRAWN.COMPLAINT REVIEWED. DEVICE HISTORY RECORD REVIEWED. PACKAGE INSERT REVIEWED. PRODUCT NOT RETURNED.(B)(4). |  | Y | Y | PHYSICIAN
1043534-2012-00098 | 2012-01-26 | 20100629 | Injury | Manufacturer report | CONSERVE(R) TOTAL HEAD W/BFH | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3801-4600 |  | ALLEGEDLY REMOVED DUE TO TISSUE REACTION. |  | Y | Y | PHYSICIAN
1043534-2012-00094 | 2012-01-26 | 20101201 | Injury | Manufacturer report | CONSERVE(R) TOTAL HEAD W/BFH | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3801-5035 |  | CONCLUSION: NO CONCLUSION CAN BE DRAWN. COMPLAINT REVIEWED AND ANALYSIS SHOWED NO TREND. DEVICE HISTORY RECORD REVIEWED. PRODUCT NOT RETURNED. EVIDENCE THAT PRODUCT IN SPECIFICATION WHEN USED. |  | Y | Y | PHYSICIAN
1043534-2012-00100 | 2012-01-26 | 20101102 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4450 |  | ALLEGEDLY REVISED DUE TO LOOSENESS. |  | Y | Y | PHYSICIAN
1043534-2012-00101 | 2012-01-26 | 20101102 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-4400 |  | ALLEGEDLY REVISED DUE TO REACTION. |  | Y | Y | PHYSICIAN
1043534-2012-00104 | 2012-01-26 | 20100121 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4248 |  | ALLEGEDLY REVISED DUE TO TUMOR. |  | Y | Y | PHYSICIAN
1043534-2012-00103 | 2012-01-26 | 20090420 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-5000 |  | ALLEGEDLY REMOVED DURING THE REVISION OF ANOTHER COMPONENT. |  | Y | Y | PHYSICIAN
1043534-2012-00102 | 2012-01-26 | 20090420 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-5056 |  | ALLEGEDLY REVISED DUE TO LOOSENESS. |  | Y | Y | PHYSICIAN
1043534-2012-00105 | 2012-01-26 | 20100121 | Injury | Manufacturer report | CONSERVE(R) TOTAL HEAD W/BFH | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3801-4204 |  | ALLEGEDLY REVISED DUE TO TUMOR. |  | Y | Y | PHYSICIAN
1043534-2012-00110 | 2012-01-26 | 20100519 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4248 |  | CONCLUSION: NO DEVICE FAILURE. COMPLAINT REVIEWED AND ANALYSIS SHOWED NO TREND. DEVICE HISTORY RECORD REVIEWED. PACKAGE INSERT REVIEWED. PRODUCT NOT RETURNED. EVIDENCE THAT PRODUCT IN SPECIFICATION WHEN USED. |  | Y | Y | PHYSICIAN
1043534-2012-00112 | 2012-01-26 | 20101220 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4854 |  | CONCLUSION: NO CONCLUSION CAN BE DRAWN. THE COMPLAINT WAS REVIEWED AND ANALYSIS SHOWED NO TREND. THE PACKAGE INSERT WAS ALSO REVIEWED. THE PRODUCT WAS NOT RETURNED. |  | Y | Y | PHYSICIAN
1043534-2012-00114 | 2012-01-26 | 20101220 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-4804 |  | ADDITIONAL INFORMATION/CORRECTED DATA RECEIVED (B)(4) 2012: ADDITIONAL INFORMATION RECEIVED THAT THIS PATIENT IS A BILATERAL. WHAT WAS ORIGINALLY REPORTED AS THE REVISION DATE IS ACTUALLY THE ORIGINAL SURGERY FOR THE OPPOSITE HIP. ALTHOUGH THE MONTH/YEAR ARE KNOWN, THE DAY IS NOT. |  | Y | Y | PHYSICIAN
1043534-2012-00116 | 2012-01-26 | 20091203 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4652 |  | ALLEGEDLY REVISED DUE TO PATIENT REACTION. |  | Y | Y | PHYSICIAN
1043534-2012-00117 | 2012-01-26 | 20091203 | Injury | Manufacturer report | CONSERVE(R) TOTAL HEAD W/BFH | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3801-4604 |  | ALLEGEDLY REVISED DUE TO PATIENT REACTION. |  | Y | Y | PHYSICIAN
1043534-2012-00119 | 2012-01-26 | 20111220 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-3600 |  | ALLEGEDLY REVISED DUE TO DISLOCATION. |  | Y | Y | PHYSICIAN
1043534-2012-00128 | 2012-01-27 | 20110127 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-3800 |  | ALLEGEDLY REVISED DUE TO SUSPECTED TISSUE REACTION. |  | Y | Y | PHYSICIAN
1043534-2012-00132 | 2012-01-31 | 20110329 | Injury | Manufacturer report | CONSERVE(R) TOTAL BCH FEMORAL HEAD IDE | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38CH-3800 |  | ALLEGEDLY REVISED DUE TO DISLOCATION. |  | Y | Y | PHYSICIAN
1043534-2012-00135 | 2012-01-31 | 20110329 | Injury | Manufacturer report | CONSERVE(R) TOTAL NECK SLEEVE | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38NS-0000 |  | ALLEGEDLY REMOVED DURING THE REVISION OF ANOTHER COMPONENT. |  | Y | Y | PHYSICIAN
1043534-2012-00137 | 2012-01-31 | 20090701 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-4835 |  | DEVICE #2:INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. THE EVENT CODE IS ADDRESSED IN THE PACKAGE INSERT. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS IS THE SAME EVENT AS 1043534-2012-00136, 00138, 00139. |  | Y | Y | PHYSICIAN
1043534-2012-00136 | 2012-01-31 | 20090701 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4854 |  | ALLEGEDLY REVISED DUE TO SUSPECTED TISSUE REACTION. |  | Y | Y | PHYSICIAN
1043534-2012-00141 | 2012-02-01 | 20110809 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4854 |  | ALLEGEDLY REVISED DUE TO SUSPECTED TISSUE REACTION. |  | Y | Y | PHYSICIAN
1043534-2012-00140 | 2012-02-01 | 20110809 | Injury | Manufacturer report | CONSERVE(R) SF W/HM TECH. FEMORAL COMPONENT | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HM-1048 |  | ALLEGEDLY REVISED DUE TO SUSPECTED TISSUE REACTION. |  | Y | Y | PHYSICIAN
1043534-2012-00151 | 2012-02-01 | 20111110 | Injury | Manufacturer report | CONSERVE(R) TOTAL NECK SLEEVE | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38NS-3500 |  | ALLEGEDLY REVISED DUE TO REPEATED DISLOCATIONS. |  | Y | Y | PHYSICIAN
1043534-2012-00153 | 2012-02-01 | 20111020 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-3600 |  | CONCLUSION: NO CONCLUSION CAN BE DRAWN. COMPLAINT HISTORY REVIEWED AND ANALYSIS SHOWS NO TREND FOR ITEM (B)(4), LOT NO: 016268864. PRODUCT NOT RETURNED. DEVICE HISTORY REVIEWED. PACKAGE INSERT REVIEWED. |  | Y | Y | PHYSICIAN
1043534-2012-00157 | 2012-02-03 | 20100222 | Injury | Manufacturer report | CONSERVE(R) PLUS SPIKED CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38SP-5258 |  | ALLEGEDLY REVISED DUE TO PATIENT REACTION AND LOOSENESS. |  | Y | Y | PHYSICIAN
1043534-2012-00158 | 2012-02-03 | 20100222 | Injury | Manufacturer report | CONSERVE(R) TOTAL HEAD W/BFH | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3801-5200 |  | ALLEGEDLY REMOVED DURING THE REVISION OF ANOTHER COMPONENT. |  | Y | Y | PHYSICIAN
1043534-2012-00168 | 2012-02-09 | 20120103 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-5600 |  | ALLEGEDLY REMOVED DURING THE REVISION OF ANOTHER COMPONENT. |  | Y | Y | PHYSICIAN
1043534-2012-00167 | 2012-02-09 | 20120103 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-5662 |  | INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. THE EVENT CODE IS ADDRESSED IN THE PACKAGE INSERT. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS IS THE SAME EVENT AS 1043534-2012-00168, 00169. |  | Y | Y | PHYSICIAN
1043534-2012-00177 | 2012-02-10 | 20110521 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4652 |  | INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. THE EVENT CODE IS ADDRESSED IN THE PACKAGE INSERT. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS IS THE SAME EVENT AS 1043534-2012-00178, 00179. THIS EVENT OCCURRED IN (B)(6). |  | Y | Y | PHYSICIAN
1043534-2012-00178 | 2012-02-10 | 20110521 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-4604 |  | DEVICE #2: INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. NO COMPLAINT WAS STATED AGAINST THIS DEVICE. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS IS THE SAME EVENT AS 1043534-2012-00177, 00179. THIS EVENT OCCURRED IN (B)(6). |  | Y | Y | PHYSICIAN
1043534-2012-00195 | 2012-02-15 | 20120119 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-4835 |  | ALLEGEDLY REVISED DUE TO PAIN. |  | Y | N | PHYSICIAN
1043534-2012-00197 | 2012-02-16 | 20110811 | Injury | Manufacturer report | CONSERVE(R) TOTAL HEAD W/BFH | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3801-4600 |  | ALLEGEDLY REVISED DUE TO SUSPECTED TISSUE REACTION AND LOOSENING OF CUP. |  | Y | Y | PHYSICIAN
1043534-2012-00196 | 2012-02-16 | 20110811 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4652 |  | ALLEGEDLY REVISED DUE TO SUSPECTED TISSUE REACTION AND LOOSENING. |  | Y | Y | PHYSICIAN
1043534-2012-00214 | 2012-03-01 | 20110329 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4854 |  | CONCLUSION: NO CONCLUSION CAN BE DRAWN. COMPLAINT REVIEWED. DEVICE HISTORY RECORD REVIEWED. PACKAGE INSERT REVIEWED. PRODUCT NOT RETURNED. EVIDENCE THAT PRODUCT IN SPECIFICATION WHEN USED, UNDETERMINED. |  | Y | Y | PHYSICIAN
1043534-2012-00213 | 2012-03-01 | 20110329 | Injury | Manufacturer report | CONSERVE(R) HM FEMORAL COMPONENT | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HM-1048 |  | ALLEGEDLY REVISED DUE TO PATIENT REACTION OF PAIN AND SUFFERING. |  | Y | Y | PHYSICIAN
1043534-2012-00216 | 2012-03-01 | 20120206 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-5400 |  | DEVICE #2: INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. NO COMPLAINT WAS STATED AGAINST THIS DEVICE. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS IS THE SAME EVENT AS 1043534-2012-00215. |  | Y | Y | PHYSICIAN
1043534-2012-00222 | 2012-03-05 | 20081103 | Injury | Manufacturer report | CONSERVE(R) TOTAL HEAD W/BFH | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3801-4204 |  | ALLEGEDLY REVISED DUE TO TISSUE REACTION. |  | Y | Y | PHYSICIAN
1043534-2012-00221 | 2012-03-05 | 20081103 | Injury | Manufacturer report | CONSERVE(R) PLUS SPIKED CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38SP-4248 |  | ALLEGEDLY REVISED DUE TO LOOSENESS AND TISSUE REACTION. |  | Y | Y | PHYSICIAN
MW5024525 | 2012-03-01 | 20071020 | Injury | Voluntary report | WRIGHT MEDICAL CONSERVE PLUS DEVICES | NONE | WRIGHT MEDICAL |  |  |  | UNREMITTING CHRONIC PAIN AND STIFFNESS FOLLOWING BILATERAL HIP REPLACEMENT SURGERIES IN (B)(6) 2007. THE PROSTHESES ARE WRIGHT MEDICAL CONSERVE PLUS DEVICES. ALTHOUGH I AM NO LONGER CANE-DEPENDENT, THE SURGERIES NEVER ALLEVIATED ALL OF THE HIP-RELATED PAIN, WHICH HAS WORSENED OVER TIME. POTENTIAL RELATED PROBLEMS ARE SHORTNESS OF BREATH AND PROSTATE CANCER DIAGNOSED IN (B)(6) 2011. THERE IS NO HISTORY OF PROSTATE IN MY FAMILY. |  | Y | Y | 0
1043534-2012-00225 | 2012-03-08 |  | Injury | Manufacturer report | CONSERVE(R) HIP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  |  |  | INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. THE EVENT CODE IS ADDRESSED IN THE PACKAGE INSERT. ALTHOUGH SEVERAL ATTEMPTS HAVE BEEN MADE, NO ADDITIONAL INFORMATION HAS BEEN RECEIVED AT THIS TIME. THE REPORTER DID INDICATE THAT NO REVISION HAS OCCURRED TO DATE. NO PATIENT, CATALOG NUMBER, OR EVENT DATES HAVE BEEN ADVISED. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS EVENT OCCURRED IN (B)(6). |  | Y | Y | PHYSICIAN
1043534-2012-00229 | 2012-03-09 | 20120206 | Injury | Manufacturer report | CONSERVE(R) TOTAL FEMORAL HEAD W/BFH | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3801-5600 |  | ALLEGEDLY REMOVED AT THE CONTRA-LATERAL SURGERY. NO COMPLAINT STATED AGAINST THIS COMPONENT. |  | Y | Y | PHYSICIAN
1043534-2012-00231 | 2012-03-09 |  | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4450 |  | ALLEGEDLY REVISED DUE TO DISLOCATION. |  | Y | Y | PHYSICIAN
1043534-2012-00236 | 2012-03-14 | 20100924 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4652 |  | ALLEGEDLY, PER PATIENT MEDWATCH REFERENCE NUMBER (B)(4), THIS REVISION WAS DUE TO PAIN AND INFECTION. |  | Y | Y | PHYSICIAN
1043534-2012-00239 | 2012-03-15 | 20120215 | Injury | Manufacturer report | CONSERVE(R) PLUS SPIKED CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38SP-4652 |  | ALLEGEDLY REVISED DUE TO SUSPECTED REACTION. |  | Y | Y | PHYSICIAN
1043534-2012-00238 | 2012-03-15 | 20120215 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS(R) FEMORAL HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AC-4600 |  | CONCLUSION: NO CONCLUSION CAN BE DRAWN. COMPLAINT REVIEWED AND ANALYSIS SHOWED NO TREND FOR (B)(4) LOT NO: 02339444. DEVICE HISTORY RECORD REVIEWED. PACKAGE INSERT. PRODUCT NOT RETURNED. EVIDENCE THAT PRODUCT IN SPECIFICATION WHEN USED, UNDETERMINED. |  | Y | Y | PHYSICIAN
1043534-2012-00240 | 2012-03-15 | 20120215 | Injury | Manufacturer report | CONSERVE(R) TOTAL NECK SLEEVE | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38NS-0035 |  | ALLEGEDLY REVISED DUE TO SUSPECTED REACTION. |  | Y | Y | PHYSICIAN
1043534-2012-00246 | 2012-03-15 | 20120301 | Injury | Manufacturer report | CONSERVE(R) TOTAL FEMORAL HEAD W/BFH | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3801-5600 |  | DEVICE #2: INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. NO COMPLAINT WAS STATED AGAINST THIS DEVICE. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS IS THE SAME EVENT AS 1043534-2012-00245, 00247. |  | Y | Y | PHYSICIAN
1043534-2012-00251 | 2012-03-15 | 20120215 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS(R) FEMORAL HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-4835 |  | ALLEGEDLY REVISED DUE TO PAIN AND CORROSION. |  | Y | Y | PHYSICIAN
1043534-2012-00253 | 2012-03-15 | 20120215 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4854 |  | ALLEGEDLY REMOVED DURING THE REVISION OF ANOTHER DEVICE. |  | Y | Y | PHYSICIAN
1043534-2012-00262 | 2012-03-28 |  | Injury | Manufacturer report | CONSERVE(R) PLUS SPIKED CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38SP-5056 |  | INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. THE EVENT CODE IS ADDRESSED IN THE PACKAGE INSERT. ALTHOUGH SEVERAL ATTEMPTS HAVE BEEN MADE, NO MEDWATCH 3500 NOR DATE OF REVISION HAS BEEN RECEIVED FROM THE USER FACILITY. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS IS THE SAME EVENT AS 1043534-2012-00263, 00264. |  | Y | Y | PHYSICIAN
1043534-2012-00266 | 2012-03-28 |  | Injury | Manufacturer report | CONSERVE(R) TOTAL HEAD W/BFH | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3801-4800 |  | ALLEGEDLY REVISED DUE TO MOM ISSUES. |  | Y | Y | PHYSICIAN
1043534-2012-00265 | 2012-03-28 |  | Injury | Manufacturer report | CONSERVE(R) PLUS SPIKED CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38SP-4854 |  | INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. THE EVENT IS ADDRESSED IN THE PACKAGE INSERT. ALTHOUGH SEVERAL ATTEMPTS HAVE BEEN MADE, NO MEDWATCH 3500 NOR DATE OF REVISION HAS BEEN RECEIVED FROM THE USER FACILITY. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS IS THE SAME EVENT AS 1043534-2012-00266, 00267. |  | Y | Y | PHYSICIAN
1043534-2012-00263 | 2012-03-29 |  | Injury | Manufacturer report | CONSERVE(R) TOTAL HEAD W/BFH | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3801-5000 |  | ALLEGEDLY REVISED DUE TO MOM ISSUES. |  | Y | Y | PHYSICIAN
1043534-2012-00269 | 2012-03-29 | 20100225 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-5400 |  | ALLEGEDLY REVISED DUE TO MOM ISSUES. |  | Y | Y | PHYSICIAN
1043534-2012-00268 | 2012-03-29 | 20100225 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-5460 |  | ALLEGEDLY REVISED DUE TO MOM ISSUES. |  | Y | Y | PHYSICIAN
1043534-2012-00274 | 2012-04-03 | 20120229 | Injury | Manufacturer report | CONSERVE(R) PLUS SPIKED CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38SP-5056 |  | CONCLUSION: NO CONCLUSION CAN BE DRAWN. COMPLAINT REVIEWED. DEVICE HISTORY RECORD REVIEWED. PACKAGE INSERT REVIEWED. PRODUCT NOT RETURNED. EVIDENCE THAT PRODUCT IN SPECIFICATION WHEN USED; UNDETERMINED. |  | Y | Y | PHYSICIAN
1043534-2012-00273 | 2012-04-03 | 20120229 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-5035 |  | ALLEGEDLY REVISED DUE TO METAL ON METAL ISSUES. |  | Y | Y | PHYSICIAN
1043534-2012-00281 | 2012-04-04 | 20120217 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-3804 |  | ALLEGEDLY REVISED DUE TO A FEMUR FRACTURE. |  | Y | Y | PHYSICIAN
1043534-2012-00277 | 2012-04-04 | 20120221 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4450 |  | ALLEGEDLY REVISED DUE TO A LOOSE COMPONENT. |  | Y | Y | PHYSICIAN
1043534-2012-00278 | 2012-04-04 | 20120221 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-4435 |  | ALLEGEDLY REMOVED DURING THE REVISION OF ANOTHER COMPONENT. |  | Y | Y | PHYSICIAN
1043534-2012-00283 | 2012-04-04 | 20080307 | Injury | Manufacturer report | CONSERVE(R) SUPER FINISH FEMORAL | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3803-1048 |  | ALLEGEDLY REVISED DUE TO MIDCERVICAL RIGHT FEMORAL NECK FRACTURE. |  | Y | Y | PHYSICIAN
1043534-2012-00285 | 2012-04-04 | 20110404 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-4404 |  | ALLEGEDLY REMOVED DURING THE REVISION OF ANOTHER COMPONENT. |  | Y | Y | PHYSICIAN
1043534-2012-00284 | 2012-04-04 | 20110404 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4450 |  | ALLEGEDLY REVISED DUE TO ASEPTIC LOOSENING. |  | Y | Y | PHYSICIAN
1043534-2012-00286 | 2012-04-04 | 20100608 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4450 |  | ALLEGEDLY REVISED DUE TO A COMPONENT SHIFTED VERTICALLY, ASEPTIC LOOSENING, RIGHT HIP INSTABILITY AND POSSIBLE METALLOSIS. |  | Y | Y | PHYSICIAN
1043534-2012-00287 | 2012-04-04 | 20100608 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-4404 |  | ALLEGEDLY REVISED DUE TO RIGHT HIP INSTABILITY AND POSSIBLE METALLOSIS. |  | Y | Y | PHYSICIAN
1043534-2012-00292 | 2012-04-05 | 20120227 | Injury | Manufacturer report | CONSERVE(R) PLUS HA CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HA-4854 |  | ALLEGEDLY REVISED DUE PAIN AND A PSEUDOTUMOR. |  | Y | Y | HEALTH PROFESSIONAL
1043534-2012-00293 | 2012-04-05 | 20120227 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-4804 |  | ALLEGEDLY REVISED DUE PAIN AND A PSEUDOTUMOR. |  | Y | Y | HEALTH PROFESSIONAL
1043534-2012-00295 | 2012-04-05 | 20120308 | Injury | Manufacturer report | CONSERVE(R) PLUS HA CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HA-5258 |  | ALLEGEDLY REVISED DUE PAIN. |  | Y | Y | PHYSICIAN
1043534-2012-00296 | 2012-04-05 | 20120307 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4858 |  | ALLEGEDLY REVISED DUE TO PAIN. |  | Y | Y | PHYSICIAN
1043534-2012-00297 | 2012-04-05 | 20120307 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-4835 |  | CONCLUSION: NO CONCLUSION CAN BE DRAWN. THE COMPLAINT, DEVICE HISTORY RECORD, AND PACKAGE INSERT WERE REVIEWED. THE PRODUCT WAS NOT RETURNED.. |  | Y | Y | PHYSICIAN
1043534-2012-00302 | 2012-04-10 |  | Injury | Manufacturer report | CONSERVE(R) CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  |  |  | ALLEGEDLY REVISED DUE PAIN. A PSEUDOTUMOR WAS REVEALED ON MRI; HOWEVER, METAL IONS WERE RELATIVELY LOW. |  | Y | Y | HEALTH PROFESSIONAL
1043534-2012-00303 | 2012-04-10 |  | Injury | Manufacturer report | CONSERVE(R) BFH | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  |  |  | ALLEGEDLY REVISED DUE TO ASYMPTOMATIC WITH MILD SCIATIC SYMPTOMS AROUND LOWER BACK. |  | Y | Y | PHYSICIAN
1043534-2012-00306 | 2012-04-10 | 20120307 | Injury | Manufacturer report | CONSERVE(R) PLUS HA CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HA-5056 |  | ALLEGEDLY REVISED DUE TO PAIN AND ABNORMALLY HIGH LEVELS OF CHROME. |  | Y | Y | PHYSICIAN
1043534-2012-00307 | 2012-04-10 | 20120307 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-5000 |  | ALLEGEDLY REVISED DUE TO PAIN AND ABNORMALLY HIGH LEVELS OF CHROME. |  | Y | Y | PHYSICIAN
1043534-2012-00311 | 2012-04-10 | 20120227 | Injury | Manufacturer report | CONSERVE(R) PLUS HA CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HA-5056 |  | ALLEGEDLY REMOVED DURING THE REVISION OF ANOTHER COMPONENT. |  | Y | Y | PHYSICIAN
1043534-2012-00310 | 2012-04-11 | 20120227 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-5000 |  | ALLEGEDLY REMOVED DURING THE REVISION OF ANOTHER COMPONENT. |  | Y | Y | PHYSICIAN
1043534-2012-00319 | 2012-04-12 | 20120312 | Malfunction | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-5404 |  | ALLEGEDLY REVISED DUE TO PATIENT NEVER FELT GREAT AND CONCERN OVER METAL IONS. REVISION OCCURRED AND THERE WERE NO VISABLE SIGNS OF METAL WEAR. |  | Y | N | OTHER
1043534-2012-00320 | 2012-04-12 | 20120312 | Malfunction | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-5460 |  | CONCLUSION: NO CONCLUSION CAN BE DRAWN THE COMPLAINT WAS REVIEWED. THE PRODUCT WAS NOT RETURNED. |  | Y | N | OTHER
1043534-2012-00321 | 2012-04-13 | 20110101 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-5000 |  | ALLEGEDLY REVISED DUE TO HIGH METAL ION LEVEL IN BLOOD. |  | Y | N | PHYSICIAN
1043534-2012-00322 | 2012-04-13 | 20110101 | Injury | Manufacturer report | CONSERVE(R) PLUS HA CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38HA-5056 |  | INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. THE EVENT CODE IS ADDRESSED IN THE PACKAGE INSERT. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS THE SAME EVENT AS 1043534-2012-00321, 00323, 00324. THIS EVENT OCCURRED IN (B)(6). |  | Y | N | PHYSICIAN
1043534-2012-00326 | 2012-04-16 | 20120319 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-5600 |  | ALLEGEDLY REMOVED DURING THE REVISION OF ANOTHER COMPONENT. |  | N | N | PHYSICIAN
1043534-2012-00332 | 2012-04-16 | 20120319 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-4000 |  | ALLEGEDLY REVISED DUE TO SUSPECTED TISSUE REACTION AND MISMATCH COMPONENTS. |  | Y | Y | PHYSICIAN
1043534-2012-00331 | 2012-04-16 | 20120319 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4450 |  | INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. THE EVENT CODE IS ADDRESSED IN THE PACKAGE INSERT. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS IS THE SAME EVENT AS 1043534-2012-00332, 00333, 00334. |  | Y | Y | PHYSICIAN
1043534-2012-00339 | 2012-04-19 | 20120308 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-5235 |  | ADDITIONAL INFORMATION: PRODUCT RETURNED FOR EVALUATION. INCIDENT INVESTIGATION REOPENED. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS IS THE SAME EVENT AS 1043534-2012-00338 AND 1043534-2013-01230, 01231. |  | N | N | OTHER
1043534-2012-00341 | 2012-04-19 | 20120322 | Injury | Manufacturer report | CONSERVE(R) TOTAL HEAD W/BFH | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3801-5200 |  | ALLEGEDLY REMOVED DURING THE REVISION OF ANOTHER COMPONENT. |  | N | N | PHYSICIAN
1043534-2012-00342 | 2012-04-20 | 20120312 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4652 |  | INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. THE EVENT CODE IS ADDRESSED IN THE PACKAGE INSERT. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS IS THE SAME EVENT AS 1043534-2012-00343. THIS EVENT OCCURRED IN (B)(6). |  | Y | Y | PHYSICIAN
1043534-2012-00343 | 2012-04-20 | 20120312 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-4635 |  | ALLEGEDLY REVISED DUE TO LOOSENESS. |  | Y | Y | PHYSICIAN
1043534-2012-00345 | 2012-04-20 | 20111121 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-5200 |  | ALLEGEDLY REVISED DUE TO REACTION. |  | Y | Y | PHYSICIAN
1043534-2012-00344 | 2012-04-20 | 20111121 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-5258 |  | ALLEGEDLY REVISED DUE TO LOOSENESS. |  | Y | Y | PHYSICIAN
1043534-2012-00348 | 2012-04-20 | 20120111 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-5056 |  | INVESTIGATION IS NOT COMPLETE. TRENDS WILL BE EVALUATED. THE EVENT IS ADDRESSED IN THE PACKAGE INSERT. THIS REPORT WILL BE UPDATED WHEN THE INVESTIGATION IS COMPLETE. THIS IS THE SAME EVENT AS 1043534-2012-00349, 00350, 00351. |  | Y | Y | PHYSICIAN
1043534-2012-00349 | 2012-04-20 | 20120111 | Injury | Manufacturer report | CONSERVE(R) TOTAL AM HEAD W/BFH | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | PL38HM5035 |  | ALLEGEDLY REVISED DUE TO REACTION. |  | Y | Y | PHYSICIAN
1043534-2012-00358 | 2012-04-25 | 20111010 | No answer provided | Manufacturer report | CONSERVE(R) TOTAL A-CLASS FEMORAL HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-4800 |  | ALLEGEDLY REMOVED DURING THE REVISION OF ANOTHER COMPONENT. |  | N | N | PHYSICIAN
1043534-2012-00360 | 2012-04-25 | 20101210 | Injury | Manufacturer report | CONSERVE(R) PLUS SPIKED CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38SP-4854 |  | ALLEGEDLY REVISED DUE TO LOOSENESS AND REACTION. |  | Y | Y | PHYSICIAN
1043534-2012-00361 | 2012-04-25 | 20101210 | Injury | Manufacturer report | CONSERVE(R) TOTAL A-CLASS HEAD | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 38AM-4804 |  | ALLEGEDLY REVISED DUE TO REACTION. |  | Y | N | PHYSICIAN
1043534-2012-00371 | 2012-05-04 | 20100830 | Injury | Manufacturer report | CONSERVE(R) PLUS CUP | HIP COMPONENT | WRIGHT MEDICAL TECHNOLOGY, INC. |  | 3802-4854 |  | ALLEGEDLY REVISED DUE TO METALLOSIS. |  | Y | N | OTHER

## Monthly Summary
year_month | count
2011-11 | 63
2011-12 | 26
2012-01 | 49
2012-02 | 13
2012-03 | 22
2012-04 | 39
2012-05 | 1

## Event Types
Event Type | Count
Injury | 209
Malfunction | 3
No answer provided | 1

## Brand Names
Brand Name | Count
CONSERVE(R) PLUS CUP | 46
CONSERVE(R) TOTAL A-CLASS HEAD | 44
CONSERVE(R) PLUS HA CUP | 31
CONSERVE(R) SUPER FINISH FEMORAL | 23
CONSERVE(R) TOTAL HEAD W/BFH | 20
CONSERVE(R) PLUS SPIKED CUP | 12
CONSERVE(R) TOTAL NECK SLEEVE | 5
CONSERVE(R) TOTAL A-CLASS(R) FEMORAL HEAD | 4
CONSERVE(R) PLUS SUPER FINISH FEMORAL | 3
CONSERVE(R) TOTAL HEAD | 3
CONSERVE(R) TOTAL FEMORAL HEAD W/BFH | 2
CONSERVE(R) CUP | 2
CONSERVE(R) TOTAL A-CLASS | 2
CONSERVE(R) TOTAL SHORT NECK SLEEVE | 1
CONSERVE(R) TOTAL BCH FEMORAL HEAD | 1
CONSERVE(R) BFH HEAD | 1
CONSERVE TOTAL FEMORAL HEAD | 1
CONSERVE(R) FEMORAL COMPONENT | 1
CONSERVE(R) PLUSS CUP | 1
CONSERVE(R) PLUS HIP | 1

## Case Metadata
Case Name | MDL Number | Filing Date | Court | Status | Settlement | Data Window Start | Data Window End | Total Records | Search Queries Used
Wright Conserve Hip | 2329 | 2012-05-04 | N.D. Georgia | Settled | $240 million (2015) | 2011-11-03 | 2012-05-04 | 213 | 2