## All Reports
report_number | date_received | date_of_event | event_type | report_source_code | brand_name | generic_name | manufacturer_d_name | model_number | catalog_number | device_class | event_description | manufacturer_narrative | adverse_event_flag | product_problem_flag | device_operator | manufacturer_g1_name
1822565-2015-02075 | 2015-10-06 | 20091110 | Injury | Manufacturer report | UNKNOWN ZIMMER M/L TAPER STEM | HIP PROSTHESIS | ZIMMER INC |  |  |  | NO DEVICE OR PHOTOS WERE RECEIVED; THEREFORE THE CONDITION OF THE DEVICE IS UNKNOWN. THE LOT NUMBER OF THE STEM IS UNKNOWN; THEREFORE THE DEVICE HISTORY RECORDS, COMPLAINT HISTORY COULD NOT BE REVIEWED. THE REPORTED DEVICE IS USED FOR TREATMENT. IT COULD NOT BE CONFIRMED IF THE DEVICE WAS USED IN AN APPROVED AND COMPATIBLE COMBINATION. PRIMARY OPERATIVE SURGICAL NOTES WERE NOT PROVIDED. IT IS UNKNOWN WHETHER THE COMPONENTS WERE IMPLANTED WITH THE CORRECT FIT AND ORIENTATION AS PER THE SURGICAL TECHNIQUE. RELEVANT MEDICAL HISTORY AND ADHERENCE TO REHABILITATION PROTOCOL ARE UNKNOWN. OPERATIVE NOTES FOR REVISION SURGERY PERFORMED ON (B)(6) 2009 WERE REVIEWED. MINOR AMOUNT OF FLUID WAS NOTED IN THE CAPSULE. THE STEM WAS NOTED TO HAVE SUBSIDED INTO RETROVERSION. THE STEM WAS REMOVED WITH MINIMUM DIFFICULTY. THE ACETABULAR COMPONENT WAS NOTED TO HAVE GOOD INGROWTH. A DEFINITIVE ROOT CAUSE CANNOT BE DETERMINED WITH THE INFORMATION PROVIDED. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02090 | 2015-10-08 |  | Injury | Manufacturer report | UNKNOWN M/L TAPER STANDARD OFFSET STEM | KWZ | ZIMMER INC |  |  |  | THE DEVICES WERE NOT RETURNED FOR REVIEW, THEREFORE THEIR CONDITIONS CANNOT BE DESCRIBED. THE DEVICE HISTORY RECORDS COULD NOT BE REVIEWED AS THE ITEM/LOT NUMBERS ARE UNKNOWN. THE DEVICES WERE USED IN TREATMENT. NO APPLICABLE PATIENT HISTORY IS AVAILABLE FOR REVIEW. COMPATIBILITY OF THE DEVICES IS UNKNOWN. A COMPLAINT HISTORY REVIEW COULD NOT BE PERFORMED WITH THE LACK OF IDENTIFYING PRODUCT INFORMATION. WITH THE INFORMATION PROVIDED, A DEFINITIVE ROOT CAUSE CANNOT BE STATED. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02091 | 2015-10-08 |  | Injury | Manufacturer report | UNKNOWN M/L TAPER STANDARD OFFSET STEM | KWZ | ZIMMER INC |  |  |  | THE DEVICES WERE NOT RETURNED FOR REVIEW, THEREFORE THEIR CONDITIONS CANNOT BE DESCRIBED. THE DEVICE HISTORY RECORDS COULD NOT BE REVIEWED AS THE ITEM/LOT NUMBERS ARE UNKNOWN. THE DEVICES WERE USED IN TREATMENT. NO APPLICABLE PATIENT HISTORY IS AVAILABLE FOR REVIEW. COMPATIBILITY OF THE DEVICES IS UNKNOWN. A COMPLAINT HISTORY REVIEW COULD NOT BE PERFORMED WITH THE LACK OF IDENTIFYING PRODUCT INFORMATION. WITH THE INFORMATION PROVIDED, A DEFINITIVE ROOT CAUSE CANNOT BE STATED. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02098 | 2015-10-08 |  | Injury | Manufacturer report | UNKNOWN M/L TAPER EXTENDED OFFSET STEM | KWZ | ZIMMER INC |  |  |  | THE DEVICES WERE NOT RETURNED FOR REVIEW, THEREFORE THEIR CONDITIONS CANNOT BE DESCRIBED. THE DEVICE HISTORY RECORDS COULD NOT BE REVIEWED AS THE ITEM/LOT NUMBERS ARE UNKNOWN. THE DEVICES WERE USED IN TREATMENT. NO APPLICABLE PATIENT HISTORY IS AVAILABLE FOR REVIEW. COMPATIBILITY OF THE DEVICES IS UNKNOWN. A COMPLAINT HISTORY REVIEW COULD NOT BE PERFORMED WITH THE LACK OF IDENTIFYING PRODUCT INFORMATION. WITH THE INFORMATION PROVIDED, A DEFINITIVE ROOT CAUSE CANNOT BE STATED. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02100 | 2015-10-08 |  | Injury | Manufacturer report | UNKNOWN M/L TAPER STANDARD OFFSET STEM | KWZ | ZIMMER INC |  |  |  | THE DEVICES WERE NOT RETURNED FOR REVIEW, THEREFORE THEIR CONDITIONS CANNOT BE DESCRIBED. THE DEVICE HIS COULD NOT BE REVIEWED AS THE ITEM/LOT NUMBERS ARE UNKNOWN. THE DEVICES WERE USED IN TREATMENT. NO APPLICABLE PATIENT HISTORY IS AVAILABLE FOR REVIEW. COMPATIBILITY OF THE DEVICES IS UNKNOWN. A COMPLAINT HISTORY REVIEW COULD NOT BE PERFORMED WITH THE LACK OF IDENTIFYING PRODUCT INFORMATION. WITH THE INFORMATION PROVIDED, A DEFINITIVE ROOT CAUSE CANNOT BE STATED. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02092 | 2015-10-08 |  | Injury | Manufacturer report | UNKNOWN M/L TAPER STANDARD OFFSET STEM | KWZ | ZIMMER INC |  |  |  | THE DEVICES WERE NOT RETURNED FOR REVIEW, THEREFORE THEIR CONDITIONS CANNOT BE DESCRIBED. THE DEVICE HISTORY RECORDS COULD NOT BE REVIEWED AS THE ITEM/LOT NUMBERS ARE UNKNOWN. THE DEVICES WERE USED IN TREATMENT. NO APPLICABLE PATIENT HISTORY IS AVAILABLE FOR REVIEW. COMPATIBILITY OF THE DEVICES IS UNKNOWN. A COMPLAINT HISTORY REVIEW COULD NOT BE PERFORMED WITH THE LACK OF IDENTIFYING PRODUCT INFORMATION. WITH THE INFORMATION PROVIDED, A DEFINITIVE ROOT CAUSE CANNOT BE STATED. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02094 | 2015-10-08 |  | Injury | Manufacturer report | UNKNOWN M/L TAPER EXTENDED OFFSET STEM | KWZ | ZIMMER INC |  |  |  | THE DEVICES WERE NOT RETURNED FOR REVIEW, THEREFORE THEIR CONDITIONS CANNOT BE DESCRIBED. THE DEVICE HISTORY RECORDS COULD NOT BE REVIEWED AS THE ITEM/LOT NUMBERS ARE UNKNOWN. THE DEVICES WERE USED IN TREATMENT. NO APPLICABLE PATIENT HISTORY IS AVAILABLE FOR REVIEW. COMPATIBILITY OF THE DEVICES IS UNKNOWN. A COMPLAINT HISTORY REVIEW COULD NOT BE PERFORMED WITH THE LACK OF IDENTIFYING PRODUCT INFORMATION. WITH THE INFORMATION PROVIDED, A DEFINITIVE ROOT CAUSE CANNOT BE STATED. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02095 | 2015-10-08 |  | Injury | Manufacturer report | UNKNOWN M/L TAPER STANDARD OFFSET STEM | KWZ | ZIMMER INC |  |  |  | THE DEVICES WERE NOT RETURNED FOR REVIEW, THEREFORE THEIR CONDITIONS CANNOT BE DESCRIBED. THE DEVICE HISTORY RECORDS COULD NOT BE REVIEWED AS THE ITEM/LOT NUMBERS ARE UNKNOWN. THE DEVICES WERE USED IN TREATMENT. NO APPLICABLE PATIENT HISTORY IS AVAILABLE FOR REVIEW. COMPATIBILITY OF THE DEVICES IS UNKNOWN. A COMPLAINT HISTORY REVIEW COULD NOT BE PERFORMED WITH THE LACK OF IDENTIFYING PRODUCT INFORMATION. WITH THE INFORMATION PROVIDED, A DEFINITIVE ROOT CAUSE CANNOT BE STATED. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02093 | 2015-10-08 |  | Injury | Manufacturer report | UNKNOWN M/L TAPER STANDARD OFFSET STEM | KWZ | ZIMMER INC |  |  |  | THE DEVICES WERE NOT RETURNED FOR REVIEW, THEREFORE THEIR CONDITIONS CANNOT BE DESCRIBED. THE DEVICE HISTORY RECORDS COULD NOT BE REVIEWED AS THE ITEM/LOT NUMBERS ARE UNKNOWN. THE DEVICES WERE USED IN TREATMENT. NO APPLICABLE PATIENT HISTORY IS AVAILABLE FOR REVIEW. COMPATIBILITY OF THE DEVICES IS UNKNOWN. A COMPLAINT HISTORY REVIEW COULD NOT BE PERFORMED WITH THE LACK OF IDENTIFYING PRODUCT INFORMATION. WITH THE INFORMATION PROVIDED, A DEFINITIVE ROOT CAUSE CANNOT BE STATED. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02096 | 2015-10-09 |  | Injury | Manufacturer report | UNKNOWN M/L TAPER EXTENDED OFFSET STEM | KWZ | ZIMMER INC |  |  |  | THE DEVICES WERE NOT RETURNED FOR REVIEW, THEREFORE THEIR CONDITIONS CANNOT BE DESCRIBED. THE DEVICE HISTORY RECORDS COULD NOT BE REVIEWED AS THE ITEM/LOT NUMBERS ARE UNKNOWN. THE DEVICES WERE USED IN TREATMENT. NO APPLICABLE PATIENT HISTORY IS AVAILABLE FOR REVIEW. COMPATIBILITY OF THE DEVICES IS UNKNOWN. A COMPLAINT HISTORY REVIEW COULD NOT BE PERFORMED WITH THE LACK OF IDENTIFYING PRODUCT INFORMATION. WITH THE INFORMATION PROVIDED, A DEFINITIVE ROOT CAUSE CANNOT BE STATED. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02101 | 2015-10-09 |  | Injury | Manufacturer report | UNKNOWN M/L TAPER STANDARD OFFSET STEM | KWZ | ZIMMER INC |  |  |  | THE DEVICES WERE NOT RETURNED FOR REVIEW, THEREFORE THEIR CONDITIONS CANNOT BE DESCRIBED. THE DEVICE HISTORY RECORDS COULD NOT BE REVIEWED AS THE ITEM/LOT NUMBERS ARE UNKNOWN. THE DEVICES WERE USED IN TREATMENT. NO APPLICABLE PATIENT HISTORY IS AVAILABLE FOR REVIEW. COMPATIBILITY OF THE DEVICES IS UNKNOWN. A COMPLAINT HISTORY REVIEW COULD NOT BE PERFORMED WITH THE LACK OF IDENTIFYING PRODUCT INFORMATION. WITH THE INFORMATION PROVIDED, A DEFINITIVE ROOT CAUSE CANNOT BE STATED. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02097 | 2015-10-09 |  | Injury | Manufacturer report | UNKNOWN M/L TAPER STANDARD OFFSET STEM | KWZ | ZIMMER INC |  |  |  | THE DEVICES WERE NOT RETURNED FOR REVIEW, THEREFORE THEIR CONDITIONS CANNOT BE DESCRIBED. THE DEVICE HISTORY RECORDS COULD NOT BE REVIEWED AS THE ITEM/LOT NUMBERS ARE UNKNOWN. THE DEVICES WERE USED IN TREATMENT. NO APPLICABLE PATIENT HISTORY IS AVAILABLE FOR REVIEW. COMPATIBILITY OF THE DEVICES IS UNKNOWN. A COMPLAINT HISTORY REVIEW COULD NOT BE PERFORMED WITH THE LACK OF IDENTIFYING PRODUCT INFORMATION. WITH THE INFORMATION PROVIDED, A DEFINITIVE ROOT CAUSE CANNOT BE STATED. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02169 | 2015-10-16 |  | Injury | Manufacturer report | ZIMMER M/L TAPER FEMORAL STEM | LZO | ZIMMER INC |  | 00771101520 |  | (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02354 | 2015-11-12 |  | Injury | Manufacturer report | ZIMMER M/L TAPER FEMORAL STEM | LZO | ZIMMER INC. |  | 65771101520 |  | NO DEVICES OR PHOTOGRAPHS WERE RETURNED FOR REVIEW; THEREFORE THE CONDITION OF THE DEVICES IS UNKNOWN. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. REVIEW OF THE OPERATIVE NOTES CONFIRMS THAT THE PATIENT UNDERWENT LEFT TOTAL HIP ARTHROPLASTY DUE TO LEFT HIP DEGENERATIVE JOINT DISEASE. THE TRIAL COMPONENTS GAVE GOOD STABILITY. THE TRIALS WERE REMOVED AND FINAL COMPONENTS WERE IMPLANTED. REVIEW OF THE OFFICE VISIT NOTES CONFIRMS THAT THE PATIENT WAS SUFFERING FROM INTERMITTENT THIGH PAIN. IT WAS NOTED FROM THE NOTES THAT THE PATIENT HAD EXCELLENT MOTION OF THE LEFT HIP. IT WAS ALSO NOTED IN THE NOTES THAT THE X RAYS DEMONSTRATE SOME MILD HETEROTOPIC OSSIFICATION, ROUGHLY A GRADE 2 AND THE PATIENT DOES HAVE SOME SCLEROTIC LINES AROUND THE SMOOTH PART OF HIS PROSTHESIS BUT NO LUCENCIES AROUND THE POROUS PART OF THE PROSTHESIS. THE SURGEON NOTED THAT THE INTERMITTENT THIGH PAIN COULD BE FROM DIFFERENTIAL MOTION BETWEEN THE STEM AND FEMUR. IT WAS REPORTED THAT THE PATIENT DOES LOT OF HEAVY LIFTING AND REGULARLY PARTICIPATES IN VERY ACTIVE EVENTS SUCH AS BOWLING OR GOLF. PRODUCT COMPATIBILITY WAS REVIEWED WITH NO ISSUES NOTED. A DEFINITE ROOT CAUSE FOR THE PATIENT'S PAIN CANNOT BE DETERMINED WITH THE INFORMATION PROVIDED. THESE DEVICES ARE USED FOR TREATMENT. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02356 | 2015-11-12 | 20151019 | Malfunction | Manufacturer report | ZIMMER M/L TAPER STEM | LZO | ZIMMER INC. |  | 00771100920 |  | THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | N | Y | HEALTH PROFESSIONAL
1822565-2015-02457 | 2015-11-24 | 20151027 | Malfunction | Manufacturer report | ZIMMER M/L TAPER HIP FEMORAL STEM | KWA | ZIMMER INC |  | 00771301600 |  | THIS REPORT IS BEING AMENDED TO REFLECT CHANGES IN SECTIONS. THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | N | Y | HEALTH PROFESSIONAL
1822565-2015-02596 | 2015-12-09 | 20151112 | Injury | Manufacturer report | ZIMMER M/L TAPER KINECTIV STEM | KWA | ZIMMER INC |  | 00771300600 |  | NO DEVICE OR PHOTOS WERE RECEIVED; THEREFORE THE CONDITION OF THE COMPONENTS IS UNKNOWN. DEVICE HISTORY RECORDS CANNOT BE REVIEWED SINCE THE LOT NUMBER IS UNKNOWN. THIS DEVICE IS USED FOR TREATMENT. SURGICAL NOTES WERE NOT PROVIDED; IT IS UNKNOWN WHETHER THE COMPONENTS WERE IMPLANTED WITH THE CORRECT FIT AND ORIENTATION AS PER THE SURGICAL TECHNIQUE. A DEFINITE ROOT CAUSE CANNOT BE DETERMINED WITH THE INFORMATION PROVIDED. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02657 | 2015-12-15 |  | Injury | Manufacturer report | ZIMMER M/L TAPER HIP PROSTHESIS FEMORAL STEM | HIP PROSTHESIS | ZIMMER INC |  | 6577110090 |  | NO DEVICE OR PHOTOS WERE RECEIVED; THEREFORE THE CONDITION OF THE COMPONENTS IS UNKNOWN. DEVICE HISTORY RECORDS CANNOT BE REVIEWED SINCE THE PART AND LOT NUMBER IS UNKNOWN. THIS DEVICE IS USED FOR TREATMENT. SURGICAL NOTES WERE NOT PROVIDED; IT IS UNKNOWN WHETHER THE COMPONENTS WERE IMPLANTED WITH THE CORRECT FIT AND ORIENTATION AS PER THE SURGICAL TECHNIQUE. A DEFINITE ROOT CAUSE CANNOT BE DETERMINED WITH THE INFORMATION PROVIDED. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02658 | 2015-12-16 |  | Injury | Manufacturer report | UNKNOWN ZIMMER M/L TAPER | HIP PROSTHESIS | ZIMMER INC |  |  |  | NO DEVICE OR PHOTOS WERE RECEIVED; THEREFORE THE CONDITION OF THE DEVICE IS UNKNOWN. DEVICE HISTORY RECORDS CANNOT BE REVIEWED SINCE THE PART AND LOT NUMBER IS UNKNOWN. SURGICAL NOTES WERE NOT PROVIDED; IT IS UNKNOWN WHETHER THE COMPONENTS WERE IMPLANTED WITH THE CORRECT FIT AND ORIENTATION AS PER THE SURGICAL TECHNIQUE. A DEFINITE ROOT CAUSE CANNOT BE DETERMINED WITH THE INFORMATION PROVIDED. THIS DEVICE IS USED FOR TREATMENT. |  | N | Y | HEALTH PROFESSIONAL
1822565-2015-02723 | 2015-12-22 | 20151112 | Injury | Manufacturer report | M/L TAPER FEMORAL STEM | PROSTHESIS, HIP | ZIMMER INC |  | 00771100540 |  | (B)(4) THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
1822565-2016-00020 | 2016-01-06 | 20151209 | Injury | Manufacturer report | ZIMMER M/L TAPER KINECTIV NECK | KWA | ZIMMER INC |  | 00784801200 |  | NO DEVICE WAS RETURNED FOR REVIEW. PHOTOGRAPHS OF THE EXPLANTED HEAD AND NECK WERE PROVIDED. NO NOTABLE DAMAGES WERE NOTED ON THE NECK FROM THE PHOTOGRAPHS PROVIDED. THE STEM REMAINS IMPLANTED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. THESE DEVICES ARE USED FOR TREATMENT. REVIEW OF THE SURGICAL NOTES DATED (B)(6) 2015 CONFIRMS THAT THE PATIENT UNDERWENT LEFT CONVERSION FROM PREVIOUS SURGERY TO THR DUE TO OSTEOARTHRITIS. THE TRIAL COMPONENTS GAVE EXCELLENT STABILITY. THE TRIALS WERE REMOVED AND THE FEMORAL NECK WAS CHECKED BOTH ANTERIORLY AND POSTERIORLY FOR ANY FRACTURES AND NO FRACTURES WERE IDENTIFIED. THE FINAL COMPONENTS ALSO GAVE EXCELLENT STABILITY. SURGICAL HISTORY FROM THE FOLLOW UP NOTES DATED (B)(6) 2015 NOTES THAT THE PATIENT HAD CANNULATED SCREW FIXATION LEFT HIP WITH ZIMMER MAGNA-FX SCREWS ON (B)(6) 2013. ON (B)(6) 2014 IRRIGATION AND DEBRIDEMENT OF THE LEFT HIP BURSA WITH SUBCUTANEOUS FLUID COLLECTION AND HARDWARE REMOVAL WAS PERFORMED. THE NOTES ALSO SHOW THAT AN INCISION AND DRAINAGE OF LEFT HIP WOUND WAS PERFORMED ON (B)(6) 2015. X RAYS NOTES CONFIRM THAT THE PROSTHESIS HAD GOOD ALIGNMENT WITH NO EVIDENCE OF LOOSENING. IT WAS ALSO NOTED THAT THE PATIENT¿S LEFT HIP IS DOING FINE AND THE WOUND HEALED NICELY. REVISION SURGERY NOTES ARE NOT PROVIDED. PRODUCT HISTORY SEARCH REVEALED NO ADDITIONAL COMPLAINTS AGAINST THE RELATED PART AND LOT COMBINATIONS. REVIEW OF THE COMPATIBILITY OF THE DEVICES CONFIRMED THAT THESE ARE AN APPROVED COMPATIBLE COMBINATION. A DEFINITE ROOT CAUSE FOR THE REPORTED ISSUE CANNOT BE DETERMINED WITH THE INFORMATION PROVIDED. |  | Y | N | HEALTH PROFESSIONAL
1822565-2016-00061 | 2016-01-12 | 20151214 | Injury | Manufacturer report | ZIMMER M/L TAPER HIP FEMORAL STEM | KWA | ZIMMER INC |  | 00771301600 |  | INFORMATION WAS RECEIVED FROM A DISTRIBUTOR WHO IS NOT REQUIRED TO COMPLETE FORM 3500A. (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | Y | HEALTH PROFESSIONAL
1822565-2016-00134 | 2016-01-25 | 20150902 | Injury | Manufacturer report | M/L TAPER PRESS-FIT FEMORAL STEM | LPH | ZIMMER INC | N/A | 00771101600 |  | (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
1822565-2016-00137 | 2016-01-25 | 20151230 | Malfunction | Manufacturer report | ZIMMER M/L TAPER PRESS-FIT FEMORAL STEM | LZO | ZIMMER INC |  | 00771101220 |  | (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | N | Y | HEALTH PROFESSIONAL
1822565-2016-00146 | 2016-01-26 | 20151229 | Injury | Manufacturer report | ZIMMER M/L TAPER PRESS-FIT STEM | LZO | ZIMMER INC |  | 00771101220 |  | (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
1822565-2016-00161 | 2016-02-01 | 20151218 | Injury | Manufacturer report | ZIMMER M/L TAPER HIP STEM | LZO | ZIMMER INC |  | 00771101200 |  | (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
1822565-2016-00187 | 2016-02-03 |  | Injury | Manufacturer report | M/L TAPER FEMORAL STEM | LZO | ZIMMER INC |  | 00771100740 |  | THIS REPORT IS BEING SUBMITTED AS THE PREVIOUS REPORT WAS MISTAKENLY FILED UNDER AN MFR NUMBER WITH THREE LEADING ZEROS. NO DEVICES OR PHOTOS WERE RECEIVED SINCE THE DEVICES STILL REMAIN IMPLANTED; THEREFORE THE CONDITION OF THE COMPONENTS IS UNKNOWN. THESE DEVICES ARE USED FOR TREATMENT. REVIEW OF THE OPERATIVE NOTES CONFIRMS THAT THE PATIENT UNDERWENT LEFT TOTAL HIP REPLACEMENT DUE TO DEGENERATIVE ARTHRITIS. THE TRIAL COMPONENTS GAVE STABLE RANGE OF MOTION. THE TRIALS WERE REMOVED AND FINAL COMPONENTS WERE IMPLANTED. THE FINAL COMPONENTS ALSO GAVE STABLE RANGE OF MOTION. LEG LENGTHS AND OFFSET SEEMED TO BE EQUALIZED. FOLLOW UP NOTES IDENTIFIED THAT THE LEFT HIP HAS GOOD RANGE OF MOTION BUT WHILE DOING STRAIGHT-LEG LIFT OR DURING FORCED FLEXION AND INTERNAL ROTATION THE PATIENT HAS MILD GROIN DISCOMFORT. THE X-RAY EVALUATION NOTES CONFIRM NO SIGNS OF LOOSENING OR FAILURE. IT HAS ALSO BEEN NOTED IN THE FOLLOW UP NOTES THAT THERE IS AN IMPRESSION OF PSOAS TENDINITIS BUT NOT SURE WHAT¿S CAUSING IT. THE NOTES ALSO STATE THAT THE PATIENT MAY BE GETTING SOME IMPINGEMENT IRRITATION. LABORATORY RESULTS CONFIRMS NO INFECTION. MRI NOTES THAT NO ABNORMALITIES WERE IDENTIFIED. PRODUCT HISTORY SEARCH REVEALED NO ADDITIONAL COMPLAINTS AGAINST THE RELATED PART AND LOT COMBINATIONS. REVIEW OF THE COMPATIBILITY OF THE DEVICES CONFIRMED THAT THESE ARE AN APPROVED COMPATIBLE COMBINATION. IT WAS REPORTED THAT THE PATIENT REGULARLY PARTICIPATES IN VERY ACTIVE EVENTS SUCH AS BOWLING OR GOLF. THE ADHERENCE TO REHABILITATION PROTOCOL IS UNKNOWN. A DEFINITE ROOT CAUSE FOR THE REPORTED ISSUE CANNOT BE DETERMINED WITH THE INFORMATION PROVIDED. |  | Y | N | HEALTH PROFESSIONAL
1822565-2016-00290 | 2016-02-16 | 20151118 | Injury | Manufacturer report | ZIMMER M/L TAPER STEM | JDI | ZIMMER INC | N/A | 00875201032 |  | INFORMATION WAS RECEIVED FROM A HEALTH PROFESSIONAL WHO IS NOT REQUIRED TO COMPLETE FORM 3500A. (B)(4). OTHER DEVICE USED: CATALOG #00875705201, CONTINUUM SHELL, LOT #61629478. CATALOG #00875201032, LONGEVITY CROSSLINKED LINER, LOT #61505133. CATALOG #00877503202, BIOLOX DELTA HEAD, LOT #2559495 - MANUFACTURED BY ZIMMER (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
1822565-2016-00306 | 2016-02-17 |  | Injury | Manufacturer report | ZIMMER M/L TAPER STEM | LZO | ZIMMER INC |  | 00771100640 |  | (B)(4). OTHER DEVICES USED: CATALOG #00875705001, CONTINUUM TM SHELL WITH CLUSTER HOLES, LOT #62898790; CATALOG #00885100932 , CONTINUUM VIVACIT-E POLY NEUTRAL LINER, LOT #62579604; CATALOG #00877503202 , BIOLOX DELTA HEAD, LOT #2766843 MANUFACTURED BY ZIMMER (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
1822565-2016-00315 | 2016-02-17 | 20160125 | Injury | Manufacturer report | ZIMMER M/L TAPER FEMORAL STEM | LZO | ZIMMER INC |  | 00771301500 |  | INFORMATION WAS RECEIVED FROM A DISTRIBUTOR WHO IS NOT REQUIRED TO COMPLETE FORM 3500A. (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
1822565-2016-00450 | 2016-03-02 |  | Injury | Manufacturer report | UNKNOWN ZIMMER M/L TAPER STEM | HIP PROSTHESIS | ZIMMER INC |  |  |  | INFORMATION WAS RECEIVED VIA JOURNAL ARTICLE. PLEASE REFERENCE LITERATURE AT THE FOLLOWING LOCATION: HTTP://WWW.SCIENCEDIRECT.COM/SCIENCE/ARTICLE/PII/S088354031501102X. (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
1822565-2016-00749 | 2016-03-22 | 20160620 | Injury | Manufacturer report | ZIMMER M/L TAPER KINECTIV FEMORAL STEM | LPH | ZIMMER INC |  | 00771301000 |  | THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
2648920-2015-00347 | 2015-09-24 | 20150817 | Malfunction | Manufacturer report | VERSYS FEMORAL HEAD | LPH | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801802801 |  | (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | N | Y | HEALTH PROFESSIONAL
2648920-2015-00350 | 2015-09-25 | 20150831 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | LPH | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801804002 |  | OTHER DEVICE USED: CATALOG #00630506440, TRILOGY LONGEVITY CROSSLINKED POLYETHYLENE ACETABULAR LINER, LOT #62496997 - MANUFACTURED BY ZIMMER (B)(4). VISUAL INSPECTION OF THE RETURNED FEMORAL HEAD SHOWS DARK DEBRIS ON THE TAPER. IT HAS ALSO BEEN NOTED THAT THERE ARE FEW SCRATCHES AND SOME FOREIGN PARTICLE ON THE OUTSIDE SURFACE OF THE HEAD. DIMENSIONS WERE FOUND CONFORMING TO PRINT SPECIFICATIONS WHERE MEASURED. VISUAL INSPECTION OF THE LINER INDICATE THIRD BODY WEAR AND SOME FOREIGN PARTICLE ON THE INNER SURFACE OF THE LINER. IT WAS ALSO NOTED THAT THERE ARE GOUGES ON THE RIM OF THE LINER. DIMENSIONAL ANALYSIS CANNOT BE PERFORMED ON THE LINER BECAUSE OF THE DAMAGES. THE STEM HAS NOT BEEN RETURNED FOR REVIEW AND IT IS UNKNOWN IF CORROSION IS PRESENT ON THE STEM NECK TAPER. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. SCANNING ELECTRON MICROSCOPY IMAGES CONFIRM THE PRESENCE OF DARK DEBRIS ON THE HEAD TAPER. ENERGY DISPERSIVE SPECTROSCOPY ANALYSIS OF THE DARK DEBRIS ON THE HEAD TAPER INDICATED THE ELEMENTS C, O, NA, AL, SI, P, CL, TI, CR, CO, AND MO. THIS HAS BEEN A COMMON ELEMENT BREAKDOWN OF THE BLACK DEBRIS FOUND IN HIP FEMORAL CORROSION EVENTS. THESE DEVICES ARE USED FOR TREATMENT. REVIEW OF THE COMPATIBILITY OF THE DEVICES CONFIRMED THAT THESE ARE AN APPROVED COMPATIBLE COMBINATION. SURGICAL NOTES WERE NOT PROVIDED. A DEFINITIVE ROOT CAUSE FOR THE CORROSION CANNOT BE DETERMINED WITH THE INFORMATION PROVIDED. |  | Y | N | HEALTH PROFESSIONAL
2648920-2015-00352 | 2015-09-30 | 20150902 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801804001 |  | (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | Y | HEALTH PROFESSIONAL
1822565-2015-02017 | 2015-10-01 |  | Injury | Manufacturer report | UNKNOWN VERSYS ET STEM | HIP PROSTHESIS | ZIMMER INC |  |  |  | (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. PLEASE REFERENCE LITERATURE AT THE FOLLOWING LOCATION: HTTP://WWW.RESEARCHGATE.NET/PUBLICATION/42110135_SEVEN_TO_TWELVE_YEAR_RESULTS_WITH_VERSYS_ET_CEMENTLESS_STEM._A_RETROSPECTIVE_STUDY_OF_225_CASES |  | Y | Y | HEALTH PROFESSIONAL
1822565-2015-02005 | 2015-10-01 |  | Injury | Manufacturer report | UNKNOWN VERSYS ET STEM | HIP PROSTHESIS | ZIMMER INC |  |  |  | NO DEVICE OR PHOTOS WERE RECEIVED; THEREFORE THE CONDITION OF THE DEVICE IS UNKNOWN. THE PART AND LOT NUMBER OF THE PRODUCT IS UNKNOWN; THEREFORE THE DEVICE HISTORY RECORDS, COMPLAINT HISTORY COULD NOT BE REVIEWED. THE REPORTED DEVICE IS USED FOR TREATMENT. IT COULD NOT BE CONFIRMED IF THE DEVICE WAS USED IN AN APPROVED AND COMPATIBLE COMBINATION. SURGICAL NOTES WERE NOT PROVIDED. IT IS UNKNOWN WHETHER THE COMPONENTS WERE IMPLANTED WITH THE CORRECT FIT AND ORIENTATION AS PER THE SURGICAL TECHNIQUE. RELEVANT MEDICAL HISTORY AND ADHERENCE TO REHABILITATION PROTOCOL ARE UNKNOWN. AS REPORTED IN THE JOURNAL ARTICLE, THERE WAS EXCESSIVE ANTEVERSION OF THE FEMUR DUE TO HIP DYSPLASIA AND THE FEMORAL COMPONENT WAS REVISED WITH A ROUNDED SECTION STEM TO MODIFY THE ANTEVERSION. A LIKELY CAUSE FOR THE REPORTED ISSUE IS EXCESSIVE ANTEVERSION OF THE FEMUR. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02007 | 2015-10-01 |  | Injury | Manufacturer report | UNKNOWN VERSYS ET STEM | HIP PROSTHESIS | ZIMMER INC |  |  |  | INFORMATION WAS RECEIVED VIA PUBLISHED LITERATURE. (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. PLEASE REFERENCE LITERATURE AT THE FOLLOWING LOCATION: HTTP://WWW.RESEARCHGATE.NET/PUBLICATION/42110135._SEVEN_TO_TWELVE_YEAR_RESULTS_WITH_VERSYS_ET_CEMENTLESS_STEM._A_RETROSPECTIVE_STUDY_OF_225_CASES. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02004 | 2015-10-01 |  | Injury | Manufacturer report | UNKNOWN VERSYS ET STEM | HIP PROSTHESIS | ZIMMER INC |  |  |  | NO DEVICE OR PHOTOS WERE RECEIVED; THEREFORE THE CONDITION OF THE DEVICE IS UNKNOWN. THE PART AND LOT NUMBER OF THE PRODUCT IS UNKNOWN; THEREFORE THE DEVICE HISTORY RECORDS, COMPLAINT HISTORY COULD NOT BE REVIEWED. THE REPORTED DEVICE IS USED FOR TREATMENT. IT COULD NOT BE CONFIRMED IF THE DEVICE WAS USED IN AN APPROVED AND COMPATIBLE COMBINATION. SURGICAL NOTES WERE NOT PROVIDED. IT IS UNKNOWN WHETHER THE COMPONENTS WERE IMPLANTED WITH THE CORRECT FIT AND ORIENTATION AS PER THE SURGICAL TECHNIQUE. RELEVANT MEDICAL HISTORY AND ADHERENCE TO REHABILITATION PROTOCOL ARE UNKNOWN. A DEFINITIVE ROOT CAUSE CANNOT BE DETERMINED WITH THE INFORMATION PROVIDED. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02018 | 2015-10-01 |  | Injury | Manufacturer report | UNKNOWN VERSYS ET STEM | HIP PROSTHESIS | ZIMMER INC |  |  |  | (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. PLEASE REFERENCE LITERATURE AT THE FOLLOWING LOCATION: HTTP://WWW.RESEARCHGATE.NET/PUBLICATION/42110135_SEVEN_TO_TWELVE_YEAR_RESULTS_WITH_VERSYS_ET_CEMENTLESS_STEM._A_RETROSPECTIVE_STUDY_OF_225_CASES |  | Y | Y | HEALTH PROFESSIONAL
1822565-2015-02006 | 2015-10-01 |  | Injury | Manufacturer report | UNKNOWN VERSYS ET STEM | HIP PROSTHESIS | ZIMMER INC |  |  |  | (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. PLEASE REFERENCE LITERATURE AT THE FOLLOWING LOCATION: HTTP://WWW.RESEARCHGATE.NET/PUBLICATION/42110135_SEVEN_TO_TWELVE_YEAR_RESULTS_WITH_VERSYS_ET_CEMENTLESS_STEM._A_RETROSPECTIVE_STUDY_OF_225_CASES |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02002 | 2015-10-01 |  | Injury | Manufacturer report | UNKNOWN VERSYS ET STEM | HIP PROSTHESIS | ZIMMER INC |  |  |  | INFORMATION WAS RECEIVED VIA PUBLISHED LITERATURE. (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. PLEASE REFERENCE LITERATURE AT THE FOLLOWING LOCATION: HTTP://WWW.RESEARCHGATE.NET/PUBLICATION/42110135_SEVEN_TO_TWELVE_YEAR_RESULTS_WITH_VERSYS_ET_CEMENTLESS_STEM._A_RETROSPECTIVE_STUDY_OF_225_CASES. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02019 | 2015-10-01 |  | Injury | Manufacturer report | UNKNOWN VERSYS ET STEM | HIP PROSTHESIS | ZIMMER INC |  |  |  | INFORMATION WAS RECEIVED VIA PUBLISHED LITERATURE. (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. PLEASE REFERENCE LITERATURE AT THE FOLLOWING LOCATION: HTTP://WWW.RESEARCHGATE.NET/PUBLICATION/42110135_SEVEN_TO_TWELVE_YEAR_RESULTS_WITH_VERSYS_ET_CEMENTLESS_STEM._A_RETROSPECTIVE_STUDY_OF_225_CASES. |  | Y | Y | HEALTH PROFESSIONAL
1822565-2015-02016 | 2015-10-02 |  | Injury | Manufacturer report | UNKNOWN VERSYS ET STEM | HIP PROSTHESIS | ZIMMER INC |  |  |  | INFORMATION WAS RECEIVED VIA PUBLISHED LITERATURE. (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. PLEASE REFERENCE LITERATURE AT THE FOLLOWING LOCATION: HTTP://WWW.RESEARCHGATE.NET/PUBLICATION/42110135_SEVEN_TO_TWELVE_YEAR_RESULTS_WITH_VERSYS_ET_CEMENTLESS_STEM._A_RETROSPECTIVE_STUDY_OF_225_CASES. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02008 | 2015-10-02 |  | Injury | Manufacturer report | UNKNOWN VERSYS ET STEM | HIP PROSTHESIS | ZIMMER INC |  |  |  | NO DEVICE OR IMAGES WERE RECEIVED. THE CONDITION OF THE DEVICE IS UNKNOWN. THE DEVICE HISTORY RECORD REVIEW AND COMPLAINT HISTORY REVIEW COULD NOT BE PERFORMED AS PART AND LOT INFORMATION OF THE REPORTED DEVICE ARE UNKNOWN. THE REPORTED DEVICE IS USED FOR TREATMENT. BASED ON THE INFORMATION PROVIDED IN THE JOURNAL ARTICLE, A LIKELY CAUSE FOR THE REPORTED ISSUE IS THE TRAUMATIC EVENT OF PATIENT FALL, THAT LIKELY CAUSED OR CONTRIBUTED TO THE STEM FRACTURE. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02015 | 2015-10-02 |  | Injury | Manufacturer report | UNKNOWN VERSYS ET STEM | HIP PROSTHESIS | ZIMMER INC |  |  |  | INFORMATION WAS RECEIVED VIA PUBLISHED LITERATURE. (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. PLEASE REFERENCE LITERATURE AT THE FOLLOWING LOCATION: HTTP://WWW.RESEARCHGATE.NET/PUBLICATION/42110135_SEVEN_TO_TWELVE_YEAR_RESULTS_WITH_VERSYS_ET_CEMENTLESS_STEM._A_RETROSPECTIVE_STUDY_OF_225_CASES. |  | Y | Y | HEALTH PROFESSIONAL
1822565-2015-02009 | 2015-10-02 |  | Injury | Manufacturer report | UNKNOWN VERSYS ET STEM | HIP PROSTHESIS | ZIMMER INC |  |  |  | INFORMATION WAS RECEIVED VIA PUBLISHED LITERATURE. (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. PLEASE REFERENCE LITERATURE AT THE FOLLOWING LOCATION: HTTP://WWW.RESEARCHGATE.NET/PUBLICATION/42110135_SEVEN_TO_TWELVE_YEAR_RESULTS_WITH_VERSYS_ET_CEMENTLESS_STEM._A_RETROSPECTIVE_STUDY_OF_225_CASES . |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02014 | 2015-10-02 |  | Injury | Manufacturer report | UNKNOWN VERSYS ET STEM | HIP PROSTHESIS | ZIMMER INC |  |  |  | INFORMATION WAS RECEIVED VIA PUBLISHED LITERATURE. (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. PLEASE REFERENCE LITERATURE AT THE FOLLOWING LOCATION: HTTP://WWW.RESEARCHGATE.NET/PUBLICATION/42110135_SEVEN_TO_TWELVE_YEAR_RESULTS_WITH_VERSYS_ET_CEMENTLESS_STEM._A_RETROSPECTIVE_STUDY_OF_225_CASES. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02013 | 2015-10-02 |  | Injury | Manufacturer report | UNKNOWN VERSYS ET STEM | HIP PROSTHESIS | ZIMMER INC |  |  |  | INFORMATION WAS RECEIVED VIA PUBLISHED LITERATURE. (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. PLEASE REFERENCE LITERATURE AT THE FOLLOWING LOCATION: HTTP://WWW.RESEARCHGATE.NET/PUBLICATION/42110135_SEVEN_TO_TWELVE_YEAR_RESULTS_WITH_VERSYS_ET_CEMENTLESS_STEM._A_RETROSPECTIVE_STUDY_OF_225_CASES . |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02012 | 2015-10-02 |  | Injury | Manufacturer report | UNKNOWN VERSYS ET STEM | HIP PROSTHESIS | ZIMMER INC |  |  |  | INFORMATION WAS RECEIVED VIA PUBLISHED LITERATURE. (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. PLEASE REFERENCE LITERATURE AT THE FOLLOWING LOCATION: HTTP://WWW.RESEARCHGATE.NET/PUBLICATION/42110135_SEVEN_TO_TWELVE_YEAR_RESULTS_WITH_VERSYS_ET_CEMENTLESS_STEM._A_RETROSPECTIVE_STUDY_OF_225_CASES. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02011 | 2015-10-02 |  | Injury | Manufacturer report | UNKNOWN VERSYS ET STEM | HIP PROSTHESIS | ZIMMER INC |  |  |  | NO DEVICES OR PHOTOS WERE RECEIVED; THEREFORE THE CONDITION OF THE DEVICES IS UNKNOWN. THE PART AND LOT NUMBERS OF THE PRODUCTS ARE UNKNOWN; THEREFORE THE DEVICE HISTORY RECORDS, COMPLAINT HISTORY COULD NOT BE REVIEWED. THE REPORTED DEVICES ARE USED FOR TREATMENT. IT COULD NOT BE CONFIRMED IF THE DEVICES WERE USED IN AN APPROVED AND COMPATIBLE COMBINATION. SURGICAL NOTES WERE NOT PROVIDED. IT IS UNKNOWN WHETHER THE COMPONENTS WERE IMPLANTED WITH THE CORRECT FIT AND ORIENTATION AS PER THE SURGICAL TECHNIQUE. RELEVANT MEDICAL HISTORY AND ADHERENCE TO REHABILITATION PROTOCOL ARE UNKNOWN. A DEFINITIVE ROOT CAUSE CANNOT BE DETERMINED WITH THE INFORMATION PROVIDED. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02010 | 2015-10-02 |  | Injury | Manufacturer report | UNKNOWN VERSYS ET STEM | HIP PROSTHESIS | ZIMMER INC |  |  |  | NO DEVICE OR PHOTOS WERE RECEIVED; THEREFORE THE CONDITION OF THE DEVICE IS UNKNOWN. THE PART AND LOT NUMBER OF THE PRODUCT IS UNKNOWN; THEREFORE THE DEVICE HISTORY RECORDS, COMPLAINT HISTORY COULD NOT BE REVIEWED. THE REPORTED DEVICE IS USED FOR TREATMENT. STERILIZED DEVICES MANUFACTURED OR DISTRIBUTED BY ZIMMER ARE STERILIZED IN ACCORDANCE WITH FDA REGULATIONS AND ISO STANDARDS TO A STERILITY ASSURANCE LEVEL OF 1.0 X 10-6 OR BETTER. THEREFORE, IT IS HIGHLY UNLIKELY THAT THE SPECIFIED DEVICE CAUSED OR CONTRIBUTED TO PATIENT INFECTION. |  | Y | N | HEALTH PROFESSIONAL
2648920-2015-00358 | 2015-10-02 | 20150905 | Death | Manufacturer report | VERSYS FEMORAL HEAD | LPH | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801802801 |  | (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02072 | 2015-10-03 |  | Injury | Manufacturer report | UNKNOWN VERSYS ET STEM | HIP PROSTHESIS | ZIMMER, INC. |  |  |  | NO DEVICE OR PHOTOS WERE RECEIVED; THEREFORE THE CONDITION OF THE DEVICE IS UNKNOWN. THE PART AND LOT NUMBER OF THE PRODUCT IS UNKNOWN; THEREFORE THE DEVICE HISTORY RECORDS, COMPLAINT HISTORY COULD NOT BE REVIEWED. THE REPORTED DEVICE IS USED FOR TREATMENT. IT COULD NOT BE CONFIRMED IF THE DEVICE WAS USED IN AN APPROVED AND COMPATIBLE COMBINATION. SURGICAL NOTES WERE NOT PROVIDED. IT IS UNKNOWN WHETHER THE COMPONENTS WERE IMPLANTED WITH THE CORRECT FIT AND ORIENTATION AS PER THE SURGICAL TECHNIQUE. RELEVANT MEDICAL HISTORY AND ADHERENCE TO REHABILITATION PROTOCOL ARE UNKNOWN. A DEFINITIVE ROOT CAUSE CANNOT BE DETERMINED WITH THE INFORMATION PROVIDED. |  | Y | N | HEALTH PROFESSIONAL
2648920-2015-00361 | 2015-10-07 | 20150913 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803602 |  | (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
2648920-2015-00362 | 2015-10-07 | 20150909 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803603 |  | (B)(4). OTHER DEVICE USED: CATALOG #UNK, UNKNOWN ZIMMER EPOCH STEM, LOT #UNK- REMAINS IMPLANTED; CATALOG #UNK, UNKNOWN ZIMMER TRILOGY SHELL, LOT #UNK. THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
2648920-2015-00363 | 2015-10-07 | 20150909 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801804003 |  | INFORMATION WAS RECEIVED FROM A DISTRIBUTOR WHO IS NOT REQUIRED TO COMPLETE FORM 3500A. (B)(4). OTHER DEVICE USED: CATALOG #65771101120, ZIMMER M/L TAPER FEMORAL STEM, LOT #62465818 MANUFACTURED BY ZIMMER INC. (B)(4) REMAINS IMPLANTED. THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | Y | HEALTH PROFESSIONAL
1822565-2015-02120 | 2015-10-13 |  | Injury | Manufacturer report | UNKNOWN VERSYS BEADED FULLCOAT PLUS STEM | HIP PROSTHESIS | ZIMMER INC |  |  |  | OPERATIVE NOTES WERE REQUESTED; HOWEVER, NONE WERE PROVIDED. RELEVANT MEDICAL HISTORY AND ADHERENCE TO THE REHABILITATION PROTOCOL ARE UNKNOWN. NO DEVICES OR PHOTOS WERE RECEIVED; THEREFORE, THE CONDITION OF THE COMPONENT IS UNKNOWN. THE PART AND LOT NUMBERS OF THE PRODUCT ARE UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORDS COULD NOT BE REVIEWED. THE PRODUCT WAS USED FOR TREATMENT. THE COMPLAINT HISTORY FOR THIS PRODUCT COULD NOT BE REVIEWED DUE TO THE LACK OF LOT NUMBERS. IT COULD NOT BE CONFIRMED IF THE DEVICES USED ARE AN APPROVED AND COMPATIBLE COMBINATION. A DEFINITIVE ROOT CAUSE CANNOT BE DETERMINED WITH THE INFORMATION PROVIDED. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE AVAILABLE INFORMATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02119 | 2015-10-13 |  | Injury | Manufacturer report | UNKNOWN VERSYS BEADED FULLCOAT PLUS STEM | HIP PROSTHESIS | ZIMMER INC |  |  |  | OPERATIVE NOTES WERE REQUESTED; HOWEVER, NONE WERE PROVIDED. RELEVANT MEDICAL HISTORY AND ADHERENCE TO THE REHABILITATION PROTOCOL ARE UNKNOWN. NO DEVICES OR PHOTOS WERE RECEIVED; THEREFORE, THE CONDITION OF THE COMPONENT IS UNKNOWN. THE PART AND LOT NUMBERS OF THE PRODUCT ARE UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORDS COULD NOT BE REVIEWED. THE PRODUCT WAS USED FOR TREATMENT. THE COMPLAINT HISTORY FOR THIS PRODUCT COULD NOT BE REVIEWED DUE TO THE LACK OF LOT NUMBERS. IT COULD NOT BE CONFIRMED IF THE DEVICES USED ARE AN APPROVED AND COMPATIBLE COMBINATION. A DEFINITIVE ROOT CAUSE CANNOT BE DETERMINED WITH THE INFORMATION PROVIDED. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE AVAILABLE INFORMATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02118 | 2015-10-13 |  | Injury | Manufacturer report | UNKNOWN VERSYS BEADED FULLCOAT PLUS STEM | HIP PROSTHESIS |  |  |  |  | INFORMATION WAS RECEIVED VIA LITERATURE. PLEASE REFERENCE LITERATURE AT THE FOLLOWING LOCATION: HTTP://WWW.NCBI.NLM.NIH.GOV/PUBMED/25556779. (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02115 | 2015-10-13 |  | Injury | Manufacturer report | UNKNOWN VERSYS BEADED FULLCOAT PLUS STEM | HIP PROSTHESIS | ZIMMER INC |  |  |  | OPERATIVE NOTES WERE REQUESTED; HOWEVER, NONE WERE PROVIDED. RELEVANT MEDICAL HISTORY AND ADHERENCE TO THE REHABILITATION PROTOCOL ARE UNKNOWN. NO DEVICES OR PHOTOS WERE RECEIVED; THEREFORE, THE CONDITION OF THE COMPONENT IS UNKNOWN. THE PART AND LOT NUMBERS OF THE PRODUCT ARE UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORDS COULD NOT BE REVIEWED. THE PRODUCT WAS USED FOR TREATMENT. THE COMPLAINT HISTORY FOR THIS PRODUCT COULD NOT BE REVIEWED DUE TO THE LACK OF LOT NUMBERS. IT COULD NOT BE CONFIRMED IF THE DEVICES USED ARE AN APPROVED AND COMPATIBLE COMBINATION. A DEFINITIVE ROOT CAUSE CANNOT BE DETERMINED WITH THE INFORMATION PROVIDED. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE AVAILABLE INFORMATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02117 | 2015-10-13 |  | Injury | Manufacturer report | UNKNOWN VERSYS BEADED FULLCOAT PLUS STEM | HIP PROSTHESIS | ZIMMER INC |  |  |  | OPERATIVE NOTES WERE REQUESTED; HOWEVER, NONE WERE PROVIDED. RELEVANT MEDICAL HISTORY AND ADHERENCE TO THE REHABILITATION PROTOCOL ARE UNKNOWN. NO DEVICES OR PHOTOS WERE RECEIVED; THEREFORE, THE CONDITION OF THE COMPONENT IS UNKNOWN. THE PART AND LOT NUMBERS OF THE PRODUCT ARE UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORDS COULD NOT BE REVIEWED. THE PRODUCT WAS USED FOR TREATMENT. THE COMPLAINT HISTORY FOR THIS PRODUCT COULD NOT BE REVIEWED DUE TO THE LACK OF LOT NUMBERS. IT COULD NOT BE CONFIRMED IF THE DEVICES USED ARE AN APPROVED AND COMPATIBLE COMBINATION. A DEFINITIVE ROOT CAUSE CANNOT BE DETERMINED WITH THE INFORMATION PROVIDED. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE AVAILABLE INFORMATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02116 | 2015-10-13 |  | Injury | Manufacturer report | UNKNOWN VERSYS BEADED FULLCOAT PLUS STEM | HIP PROSTHESIS | ZIMMER INC |  |  |  | OPERATIVE NOTES WERE REQUESTED; HOWEVER, NONE WERE PROVIDED. RELEVANT MEDICAL HISTORY AND ADHERENCE TO REHABILITATION PROTOCOL ARE UNKNOWN. NO DEVICES OR PHOTOS WERE RECEIVED; THEREFORE, THE CONDITION OF THE COMPONENT IS UNKNOWN. THE PART AND LOT NUMBERS OF THE PRODUCT ARE UNKNOWN; THEREFORE, THE DEVICE HISTORY RECORDS COULD NOT BE REVIEWED. THE PRODUCT WAS USED FOR TREATMENT. THE COMPLAINT HISTORY FOR THIS PRODUCT COULD NOT BE REVIEWED DUE TO THE LACK OF LOT NUMBERS. IT COULD NOT BE CONFIRMED IF THE DEVICES USED ARE AN APPROVED AND COMPATIBLE COMBINATION. A DEFINITIVE ROOT CAUSE CANNOT BE DETERMINED WITH THE INFORMATION PROVIDED. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE AVAILABLE INFORMATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02113 | 2015-10-13 |  | Injury | Manufacturer report | UNKNOWN ZIMMER VERSYS FIBER METAL STEM | HIP PROSTHESIS | ZIMMER INC |  |  |  | INFORMATION WAS RECEIVED VIA PUBLISHED LITERATURE. PLEASE REFERENCE LITERATURE AT THE FOLLOWING LOCATION: HTTP://WWW.NCBI.NLM.NIH.GOV/PMC/ARTICLES/PMC4232831/. (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
2648920-2015-00365 | 2015-10-13 | 20150914 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00902603607 |  | VISUAL INSPECTION OF THE RETURNED VERSYS FEMORAL HEAD SHOWS DARK DEBRIS ON THE HEAD TAPER. IT HAS ALSO BEEN NOTED THAT THERE ARE SOME WEAR MARKS ON THE SURFACE OF THE HEAD. THE RETURNED VERSYS STEM HAS DARK DEBRIS ON THE STEM NECK TAPER. DIMENSIONS OF THE HEAD AND THE STEM WERE FOUND CONFORMING TO PRINT SPECIFICATIONS WHERE MEASURED. THE RETURNED LINER HAS A PLATINUM COLORED COATING. IT HAS ALSO BEEN NOTED THAT THERE IS PITTING ON THE ARTICULAR SURFACE INDICATIVE OF THIRD BODY WEAR AND ALSO THE LINER HAS GOUGES AND CUTS. DIMENSIONAL ANALYSIS CANNOT BE PERFORMED DUE TO THE CONDITION OF THE LINER. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. SCANNING ELECTRON MICROSCOPE IMAGES CONFIRM THE PRESENCE OF DARK DEBRIS ON THE HEAD TAPER AND SHOWS BASE METAL AND CU-DISPERSIVE X-RAY SPECTROSCOPY ANALYSIS OF THE DARK DEBRIS ON THE HEAD TAPER INDICATED THE ELEMENTS C, O, MG, SI, P, CA, CR, MN, CO, AND MO. THIS HAS BEEN A COMMON ELEMENT BREAKDOWN OF THE BLACK DEBRIS FOUND IN HIP FEMORAL CORROSION EVENTS. THESE DEVICES ARE USED FOR TREATMENT. REVIEW OF THE COMPATIBILITY OF THE DEVICES CONFIRMED THAT THESE ARE AN APPROVED COMPATIBLE COMBINATION. SURGICAL NOTES WERE NOT PROVIDED; IT IS UNKNOWN WHETHER THE COMPONENTS WERE IMPLANTED WITH THE CORRECT FIT AND ORIENTATION AS PER THE SURGICAL TECHNIQUE. A DEFINITIVE ROOT CAUSE FOR THE CORROSION CANNOT BE DETERMINED WITH THE INFORMATION PROVIDED. |  | Y | N | HEALTH PROFESSIONAL
2648920-2015-00366 | 2015-10-14 | 20150916 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | HIP PROSTHESIS | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803203 |  | THIS INFORMATION WAS PREVIOUSLY SUBMITTED UNDER THE INCORRECT REPORT #0002648920-2015-00366-1. INFORMATION REPORTED WAS PREVIOUSLY UNKNOWN. CATALOG #00408801206, BEADED FULLCOAT FEMORAL STEM, LOT #61974962. |  | Y | Y | HEALTH PROFESSIONAL
2648920-2015-00367 | 2015-10-14 | 20150915 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803202 |  | PHOTOS OF THE EXPLANTED DEVICES WERE RETURNED FOR REVIEW. THE FEMORAL HEAD TAPER APPEARS TO SHOW A DARK RESIDUE OF AN UNKNOWN SUBSTANCE. THE LINER'S CONDITION IS DIFFICULT TO DISCERN FROM THE PHOTOS DUE TO BLOOD REMAINING ON THE SURFACE, HOWEVER THERE IS A SCREW HOLE PRESENT FROM THE REMOVAL EFFORT. THE DEVICE HISTORY RECORDS WERE REVIEWED FOR THE FEMORAL HEAD AND STEM AND BOTH LOTS WERE FOUND CONFORMING. THESE DEVICES WERE USED FOR TREATMENT. A WRITTEN REPORT ANALYZING THE PATIENT'S (B)(6) 2015 MRI AND RADIOGRAPHIC SERIES STATES THAT THERE IS SYNOVIAL THICKENING AND NODULARITY WITH A SOFT TISSUE COMPONENT THAT EXTENDS POSTERIOR TO THE JOINT TOWARDS THE SCIATIC NOTCH. ATROPHY OF THE MUSCLES IS ALSO STATED WITH A COMPLETE TEAR OF THE GLUTEUS MINIMUS TENDON AND A PARTIAL TEAR OF THE GLUTEUS MEDIUS TENDON. THE PATIENT HAS BILATERAL THAS AND THE LEFT HIP WAS REVISED IN (B)(6) 2015 FOR ADVERSE LOCAL TISSUE REACTION. PRIMARY OPERATIVE NOTES WERE REVIEWED AND WERE UNREMARKABLE. COMPONENT COMPATIBILITY WAS REVIEWED WITH NO ISSUES NOTED. NO ADDITIONAL COMPLAINTS HAVE BEEN RECEIVED AGAINST THE FEMORAL HEAD OR STEM MANUFACTURING LOTS. WITH THE INFORMATION PROVIDED, A DEFINITIVE ROOT CAUSE CANNOT BE STATED WITH CERTAINTY. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02134 | 2015-10-14 | 20150914 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER INC |  | 00801803605 |  | OTHER DEVICE USED: CATALOG #00620006022, TRILOGY POROUS ACETABULAR SHELL, LOT #13363400. THROUGH INSPECTION OF PROVIDED PHOTOS, THE INNER TAPER OF THE FEMORAL HEAD APPEARS TO HAVE A UNIFORM DARK COATING INSIDE. THE FEMORAL STEM TAPER IS NOT IN GOOD LIGHTING AND THEREFORE NO DETAILS CAN BE GATHERED ABOUT ITS CONDITION EXCEPT FOR A POTENTIAL KNICK ON THE TAPER SURFACE. THE DEVICES WERE USED IN TREATMENT. THE EXACT PRIMARY SURGERY DATE IS UNKNOWN, HOWEVER THE PATIENT INDICATED IT OCCURRED IN 1997, MAKING A POTENTIAL IN-VIVO TIME OF UP TO 18 YEARS. SINGLE-USE, STERILIZED DEVICES MANUFACTURED OR DISTRIBUTED BY ZIMMER ARE STERILIZED IN ACCORDANCE WITH FDA REGULATIONS AND (B)(4) STANDARDS TO A STERILITY ASSURANCE LEVEL OF (B)(4) OR BETTER, THEREFORE IT IS NOT SUSPECTED THAT THE DEVICES CAUSED OR CONTRIBUTED TO THE STAPH INFECTION. A DEFINITIVE ROOT CAUSE FOR THE TAPER CORROSION CANNOT BE DETERMINED. |  | Y | N | HEALTH PROFESSIONAL
2648920-2015-00375 | 2015-10-16 | 20150723 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803602 |  | (B)(4). CONCOMITANT MEDICAL PRODUCTS: MANUFACTURED BY (B)(4),CATALOG #00875201036, CONTINUUM, TRILOGY IT, ALLOFIT IT LINER, LOT #62770417. THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
2648920-2015-00374 | 2015-10-16 | 20150923 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803603 |  | (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | Y | HEALTH PROFESSIONAL
2648920-2015-00376 | 2015-10-16 | 20150803 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803602 |  | (B)(4). OTHER DEVICE USED: MANUFACTURED BY (B)(4), CATALOG #00875201036, CONTINUUM, TRILOGY IT, ALLOFIT IT LINER, LOT #62770417. THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
2648920-2015-00371 | 2015-10-19 | 20140715 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803603 |  | THIS REPORT WAS PREVIOUSLY SUBMITTED UNDER MEDWATCH 1822565-2015-01156. (B)(4). OTHER DEVICE USED: CATALOG #00875101236, CONTINUUM, TRILOGY IT, ALLOFIT IT POLY LINER, LOT #62387631 MANUFACTURED BY ZIMMER (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
2648920-2015-00382 | 2015-10-27 |  | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803602 |  | (B)(4). OTHER DEVICES USED: CATALOG #00625006530, TRILOGY BONE SCREW, LOT #62032209. THE FOLLOWING PRODUCTS WERE MANUFACTURED BY ZIMMER INC. (B)(4): CATALOG #UNK, UNKNOWN ZIMMER STEM, LOT #UNK- IMPLANTED ON (B)(6) 2014; CATALOG #00875201036, CONTINUUM, TRILOGY IT, ALL OF IT IT LINER, LOT #62071813; CATALOG #00875705202, CONTINUUM TM SHELL WITH MULTI HOLES, LOT #62070524. THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
2648920-2015-00381 | 2015-10-27 | 20150928 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801804002 |  | (B)(4). OTHER DEVICE USED: CATALOG #00784801300, ZIMMER KINECTIV FEMORAL NECK, LOT #61710088, MANUFACTURED BY ZIMMER INC. (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02225 | 2015-10-28 |  | Injury | Manufacturer report | UNKNOWN ZIMMER VERSYS STEM | HIP PROSTHESIS | ZIMMER INC |  |  |  | INFORMATION WAS RECEIVED VIA LITERATURE. PLEASE REFERENCE LITERATURE AT THE FOLLOWING LOCATION: HTTP://WWW.NCBI.NLM.NIH.GOV/PUBMED/19056213. (B)(4). PRODUCT WAS NOT RETURNED FOR EVALUATION THE DEVICES WERE NOT RETURNED FOR REVIEW, THEREFORE THEIR CONDITIONS CANNOT BE DESCRIBED. THE DEVICE HISTORY RECORDS COULD NOT BE REVIEWED AS THE ITEM/LOT NUMBERS ARE UNKNOWN. THE DEVICES WERE USED IN TREATMENT. NO APPLICABLE PATIENT HISTORY IS AVAILABLE FOR REVIEW. COMPATIBILITY OF THE DEVICES IS UNKNOWN. A COMPLAINT HISTORY REVIEW COULD NOT BE PERFORMED WITH THE LACK OF IDENTIFYING PRODUCT INFORMATION. WITH THE INFORMATION PROVIDED, A DEFINITIVE ROOT CAUSE CANNOT BE STATED. |  | Y | N | HEALTH PROFESSIONAL
2648920-2015-00383 | 2015-10-28 | 20151001 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803602 |  | (B)(4). OTHER DEVICE USED: CATALOG #00771101200, M/L TAPER SIZE 12.5, LOT #61765947 - MANUFACTURED BY ZIMMER (B)(4), IN REMAINS IMPLANTED. SURGICAL AND EXPLANT PHOTOS WERE PROVIDED. THERE APPEARS TO BE SOME DARK DEBRIS AROUND AND INSIDE THE OPENING OF THE FEMORAL HEAD TAPER; HOWEVER, IT CANNOT BE DEFINITIVELY DETERMINED WHAT ITS MAKE-UP IS. THERE ALSO APPEARS BE SOME SCUFFING DAMAGE JUST INSIDE THE FEMORAL HEAD TAPER, POTENTIALLY DOWN THE TAPER WALLS. NO CLEAR DEBRIS IS SEEN ON THE FEMORAL STEM TAPER FROM THE PHOTOS, WHICH WAS NOT REVISED. A MRI OF THE HIP IS NOTED TO DEMONSTRATE COMPLETE TEARING OF THE ABDUCTOR TENDONS AS WELL AS A COMPLEX FLUID COLLECTION CONSISTENT WITH PARTICLE DISEASE. NOTES STATE THAT THE PATIENT HAS SUFFERED MULTIPLE DISLOCATIONS, WHICH LIKELY CONTRIBUTED TO BY HER DEFICIENT ABDUCTORS, BUT WAS ALSO CALLED NONCOMPLIANT BY HER SURGEON. LAB NOTES ALSO INDICATE HIGH WHOLE BLOOD COBALT LEVELS. A COMPLAINT HISTORY SEARCH OF THE LOTS OF THE FEMORAL HEAD AND STEM RETURNED NO ADDITIONAL COMPLAINTS AGAINST EITHER MANUFACTURING LOT. PRODUCT COMPATIBILITY WAS CONFIRMED. THE INFORMATION PROVIDED, IT APPEARS THAT TAPER CORROSION HAS LED TO COMPLICATIONS INCLUDING ADVERSE TISSUE REACTION AND DISLOCATION; HOWEVER THE CAUSE FOR THIS TAPER CORROSION CANNOT BE DETERMINED. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02226 | 2015-10-28 |  | Injury | Manufacturer report | UNKNOWN ZIMMER VERSYS STEM | HIP PROSTHESIS | ZIMMER INC |  |  |  | PLEASE REFERENCE LITERATURE AT THE FOLLOWING LOCATION: HTTP://WWW.NCBI.NLM.NIH.GOV/PUBMED/19056213 (B)(4). PRODUCT WAS NOT RETURNED FOR EVALUATION. THE DEVICE WAS NOT RETURNED FOR REVIEW, THEREFORE ITS CONDITION CANNOT BE DESCRIBED. THE DHR COULD NOT BE REVIEWED AS THE ITEM/LOT NUMBERS ARE UNKNOWN. THE DEVICE WAS USED IN TREATMENT. NO APPLICABLE PATIENT HISTORY IS AVAILABLE FOR REVIEW. COMPATIBILITY OF THE DEVICES IS UNKNOWN. A COMPLAINT HISTORY REVIEW COULD NOT BE PERFORMED WITH THE LACK OF IDENTIFYING PRODUCT INFORMATION. WITH THE INFORMATION PROVIDED, A DEFINITIVE ROOT CAUSE CANNOT BE STATED. |  | Y | N | HEALTH PROFESSIONAL
2648920-2015-00389 | 2015-11-04 | 20151005 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801804002 |  | INFORMATION WAS RECEIVED FROM A DISTRIBUTOR WHO IS NOT REQUIRED TO COMPLETE FORM 3500A. (B)(4). OTHER DEVICE USED: CATALOG #65771101100, ZIMMER M/L TAPER STEM, LOT #60876615 MANUFACTURED BY ZIMMER INC. (B)(4) - REMAINS IMPLANTED. THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02339 | 2015-11-10 | 20130312 | Injury | Manufacturer report | VERSYS EPOCH STEM | LPH | ZIMMER INC |  | 00408801326 |  | THIS REPORT IS BEING AMENDED TO REFLECT CHANGES IN SECTIONS. NO DEVICES OR PHOTOS WERE RECEIVED; THEREFORE THE CONDITION OF THE COMPONENTS IS UNKNOWN. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. THIS DEVICE IS USED FOR TREATMENT. PRIMARY OPERATIVE NOTES WERE NOT PROVIDED. REVIEW OF THE REVISION OPERATIVE NOTES CONFIRMS THAT THE PATIENT UNDERWENT REVISION RIGHT TOTAL HIP ARTHROPLASTY. IT WAS NOTED IN THE NOTES THAT THE FEMORAL HEAD WAS REMOVED. THE ACETABULAR COMPONENT WAS DISSECTED OUT AND WAS CHECKED FOR LOOSENING, AND THERE WAS NO EVIDENCE OF LOOSENING. THE ACETABULAR COMPONENT WAS SOLIDLY FIXED. IT WAS ALSO NOTED THAT THE PROXIMAL PORTION OF THE STEM WAS PROMINENT AND WAS LOOSE. THE PROXIMAL PORTION OF THE PROSTHESIS WAS RELATIVELY EASILY REMOVED. PRODUCT HISTORY SEARCH REVEALED NO ADDITIONAL COMPLAINTS AGAINST THE RELATED PART AND LOT COMBINATION. A DEFINITE ROOT CAUSE CANNOT BE DETERMINED WITH THE INFORMATION PROVIDED. |  | Y | N | HEALTH PROFESSIONAL
2648920-2015-00396 | 2015-11-12 | 20151019 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803604 |  | VISUAL INSPECTION OF THE RETURNED FEMORAL HEAD SHOWS DARK DEBRIS ON THE HEAD TAPER. DIMENSIONS WERE FOUND CONFIRMING TO PRINT SPECIFICATIONS WHERE MEASURED. THE STEM HAS NOT BEEN RETURNED FOR REVIEW SINCE IT STILL REMAINS IMPLANTED; THEREFORE THE CONDITION OF THE STEM IS UNKNOWN. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. INSPECTION DOCUMENTATION SHOWS THE TAPER ANGLE AND DIAMETER IS 100% INSPECTED AT THE TIME OF MANUFACTURE. SCANNING ELECTRON MICROSCOPE IMAGES CONFIRM THE PRESENCE OF DARK DEBRIS ON THE HEAD TAPER. ENERGY-DISPERSIVE X-RAY SPECTROSCOPY ANALYSIS OF THE DARK DEBRIS ON THE HEAD TAPER INDICATED THE ELEMENTS C, O, SI, P, CL, CR, CO, AND MO. THIS HAS BEEN A COMMON ELEMENT BREAKDOWN OF THE BLACK DEBRIS FOUND IN HIP FEMORAL CORROSION EVENTS. THESE DEVICES ARE USED FOR TREATMENT. REVIEW OF THE COMPATIBILITY OF THE DEVICES CONFIRMED THAT THESE ARE AN APPROVED COMPATIBLE COMBINATION. THE IMPLANTS WERE IN-VIVO FOR APPROXIMATELY 3 YEARS. PRODUCT HISTORY SEARCH REVEALED NO ADDITIONAL COMPLAINTS AGAINST THE RELATED PART AND LOT COMBINATION. PATIENT'S ADHERENCE TO REHABILITATION PROTOCOL IS UNKNOWN. A DEFINITIVE ROOT CAUSE FOR THE CORROSION CANNOT BE DETERMINED WITH THE INFORMATION PROVIDED. |  | Y | N | HEALTH PROFESSIONAL
2648920-2015-00398 | 2015-11-17 | 20151020 | Malfunction | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801802230 |  | (B)(4). THE DEVICES REMAIN IMPLANTED. THE IMPACT AND CONDITION OF THE DEVICES AS A RESULT OF THE IMPLANTATION OF THIS COMBINATION IS UNKNOWN. THE DEVICE HISTORY RECORD REVIEW, FOR THE REPORTED HEAD LOT, IDENTIFIED NO DEVIATIONS OR ANOMALIES IN THE MANUFACTURING PROCESS. REVIEW OF COMPLAINT HISTORY IDENTIFIED NO PREVIOUS COMPLAINTS FOR THE PART-LOT COMBINATION ASSOCIATED WITH THE REPORTED HEAD. THE REPORTED DEVICE IS USED FOR TREATMENT. A PRODUCT COMPATIBILITY CHECK DETERMINED THAT THE BIOMET G7 ACETABULAR SYSTEM DUAL MOBILITY ACETABULAR LINER, THE ACTIVE ARTICULATION DUAL MOBILITY BEARING (AR COM XL) ARE NOT COMPATIBLE WITH THE 22MM DIAMETER VERSYS FEMORAL HEAD. THE AR COM XL IS RECOMMENDED FOR USE WITH A 22.2MM SIZED HEAD. ZIMMER-BIOMET HAS NOT ASSESSED OR CONFIRMED THE COMPATIBILITY OF THIS COMBINATION OF DEVICES, AND THIS WOULD BE CONSIDERED AN OFF-LABEL USE OF THIS PRODUCT AS INDICATED ON THE PACKAGING INSERT. |  | N | Y | HEALTH PROFESSIONAL
2648920-2015-00402 | 2015-11-20 | 20151026 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803203 |  | THE POLY LINER AND FEMORAL HEAD WERE RETURNED FOR REVIEW. THE OUTER SPHERE OF THE HEAD WAS IN GOOD CONDITION. DIMENSIONAL MEASUREMENTS OF THE DEVICE CONFORM TO PRINT SPECIFICATIONS, WHERE MEASURED. IMAGE PROVIDED OF THE STEM CONFIRMED TAPER CORROSION. DEVICE HISTORY RECORDS FOR THE LOT CODE ASSOCIATED WITH THE HEAD WERE REVIEWED. NO DEVIATIONS OR ANOMALIES WERE NOTED IN THE MANUFACTURING PROCESS. THE PART AND LOT CODE ASSOCIATED WITH THE STEM WAS NOT PROVIDED. THE STEM WAS NOT RETURNED FOR FURTHER EVALUATION AS THIS REMAINS IMPLANTED. THE REPORTED DEVICE IS USED FOR TREATMENT. SCANNING ELECTRON MICROSCOPY IMAGES CONFIRMED THE PRESENCE OF DARK DEBRIS ON THE FEMORAL HEAD TAPER. ENERGY-DISPERSIVE SPECTROSCOPY ELEMENTAL ANALYSIS OF THE DEBRIS ON THE HEAD TAPER SURFACE WAS PERFORMED. THE DEBRIS ON THE HEAD TAPER PRIMARILY CONSISTED OF ELEMENTS CARBON, NITROGEN, OXYGEN, SILICON, PHOSPHOROUS, CHROMIUM, COBALT AND MOLYBDENUM. THIS HAS BEEN A COMMON ELEMENT BREAKDOWN OF THE BLACK DEBRIS FOUND IN HIP FEMORAL CORROSION EVENTS. THE HEAD AND STEM ARE A COMPATIBLE COMBINATION; HOWEVER IT COULD NOT BE VERIFIED IF THE SHELL IS COMPATIBLE AS PART-LOT INFORMATION OF THE SHELL ARE UNKNOWN. IT IS UNKNOWN WHETHER THE COMPONENTS WERE IMPLANTED WITH THE CORRECT FIT AND ORIENTATION AS PER THE SURGICAL TECHNIQUE. A DEFINITIVE ROOT CAUSE CANNOT BE DETERMINED WITH THE INFORMATION PROVIDED. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02432 | 2015-11-20 |  | Injury | Manufacturer report | VERSYS FEMORAL STEM BEADED FULLCOAT | LZO | ZIMMER INC |  | 00784301326 |  | OTHER DEVICE USED: CATALOG #00875801128, TRILOGY CROSSLINKED LINER, LOT #61609237. CATALOG #00875705402, CONTINUUM SHELL, LOT #62343832. CATALOG #00784301326, VERSYS FULL COAT STEM, LOT #07888693. CATALOG #0080180814, VERSYS FEMORAL HEAD, LOT #62435554 - MANUFACTURED BY ZIMMER (B)(4). NO DEVICES OR PHOTOS WERE RECEIVED SINCE THE DEVICES STILL REMAIN IMPLANTED; THEREFORE THE CONDITION OF THE COMPONENTS IS UNKNOWN. THESE DEVICES ARE USED FOR TREATMENT. SURGICAL NOTES WERE NOT PROVIDED; IT IS UNKNOWN WHETHER THE COMPONENTS WERE IMPLANTED WITH THE CORRECT FIT AND ORIENTATION AS PER THE SURGICAL TECHNIQUE. PRODUCT HISTORY SEARCH REVEALED NO ADDITIONAL COMPLAINTS AGAINST THE RELATED PART AND LOT COMBINATIONS. PATIENT ACTIVITY LEVEL AND ADHERENCE TO REHABILITATION PROTOCOL IS UNKNOWN. A DEFINITE ROOT CAUSE CANNOT BE DETERMINED WITH THE INFORMATION PROVIDED. |  | Y | N | HEALTH PROFESSIONAL
2648920-2015-00403 | 2015-11-20 | 20151019 | Injury | Manufacturer report | VERSYS HIP SYSTEM FEMORAL HEAD | LPH | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803203 |  | NO DEVICES OR PHOTOS WERE RECEIVED; THEREFORE THE CONDITION OF THE COMPONENTS IS UNKNOWN. THE STEM REMAINED IMPLANTED. DEVICE HISTORY RECORDS FOR THE LOT CODE ASSOCIATED WITH THE HEAD WERE REVIEWED. NO DEVIATIONS OR ANOMALIES WERE NOTED IN THE MANUFACTURING PROCESS FOR THIS LOT. THE DEVICE HISTORY RECORD AND COMPLAINT HISTORY REVIEW COULD NOT BE PERFORMED FOR THE FEMORAL STEM AS PART AND LOT INFORMATION WAS NOT PROVIDED. THE REPORTED DEVICE IS USED FOR TREATMENT. IT COULD NOT BE CONFIRMED IF THE DEVICE WAS USED IN AN APPROVED AND COMPATIBLE COMBINATION. REVIEW OF COMPLAINT HISTORY IDENTIFIED NO PREVIOUS COMPLAINTS FOR THE PART-LOT COMBINATION ASSOCIATED WITH THE REPORTED HEAD. SURGICAL NOTES WERE NOT PROVIDED. IT IS UNKNOWN WHETHER THE COMPONENTS WERE IMPLANTED WITH THE CORRECT FIT AND ORIENTATION AS PER THE SURGICAL TECHNIQUE. RELEVANT MEDICAL HISTORY AND ADHERENCE TO REHABILITATION PROTOCOL ARE UNKNOWN. A DEFINITIVE ROOT CAUSE CANNOT BE DETERMINED WITH THE INFORMATION PROVIDED. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02450 | 2015-11-24 |  | Injury | Manufacturer report | UNKNOWN VERSYS BEADED FULLCOAT STEM | HIP PROSTHESIS | ZIMMER INC |  |  |  | INFORMATION WAS RECEIVED VIA PUBLISHED LITERATURE. PLEASE REFERENCE LITERATURE AT THE FOLLOWING LOCATION: HTTP://WWW.NCBI.NLM.NIH.GOV/PUBMED/26410635 (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
2648920-2015-00409 | 2015-11-25 | 20151027 | Malfunction | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803202 |  | THE FEMORAL HEAD WAS RETURNED FOR REVIEW. VISUAL INSPECTION CONFIRMS THAT THERE ARE NO NOTABLE DAMAGES ON THE HEAD. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. DEVICE HISTORY RECORDS REVIEW INDICATES TAPERS WERE 100% INSPECTED ON THE DIAMETER AND ANGLE AND ALL PARTS WERE CONFORMING AT TIME OF MANUFACTURE. DIMENSIONS OF THE RETURNED HEAD WERE FOUND CONFORMING TO THE PRINT SPECIFICATIONS WHERE MEASURED. THE INNER DIAMETER, THE TAPER ANGLE AND DEPTH OF THE TAPER WERE MEASURED AND FOUND TO BE CONFORMING. THIS DEVICE IS USED FOR TREATMENT. A WAGNER STEM WAS USED WITH THE REPORTED VERSYS HEAD AND A REVIEW OF THE COMPATIBILITY OF THE DEVICES CONFIRMED THAT THESE ARE AN APPROVED COMPATIBLE COMBINATION. SINCE THE FEMORAL HEAD WAS FOUND TO BE WITHIN SPECIFICATION WHEN MEASURED IT CAN BE CONFIRMED THAT THE FAILURE MODE OF THE REPORTED ISSUE COULD NOT BE CORRELATED TO THE MANUFACTURING PROCESS. A DEFINITE ROOT CAUSE FOR THE REPORTED ISSUE CANNOT BE DETERMINED WITH THE INFORMATION PROVIDED. |  | N | Y | HEALTH PROFESSIONAL
2648920-2015-00411 | 2015-11-25 | 20140415 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801802801 |  | INFORMATION WAS RECEIVED FROM A CONSUMER WHO IS NOT REQUIRED TO COMPLETE FORM 3500A. (B)(4) OTHER DEVICE USED: CATALOG #00784301807, VERSYS FEMORAL STEM, LOT #61102544 MANUFACTURED BY ZIMMER INC. (B)(4). THE REPORTED DEVICES ARE USED FOR TREATMENT. DEVICE HISTORY RECORD REVIEW IDENTIFIED NO DEVIATIONS OR ANOMALIES IN THE MANUFACTURING PROCESS. IT IS REPORTED THAT A STRYKER ADM MOBILE BEARING ACETABULAR SYSTEM INSERT 28/50 WAS USED WITH LEGACY ZIMMER 12/14, VERSYS 28MM TAPERED COCA FEMORAL HEAD AND VERSYS HIP SYSTEM FEMORAL BEADED FULLCOAT COLLARED STEM. ZIMMER BIOMET HAS NOT TESTED THE COMPATIBILITY OF THIS COMBINATION OF DEVICES, AND THIS WOULD BE CONSIDERED AN OFF-LABEL USE OF THIS PRODUCT AS INDICATED ON THE PACKAGING INSERT. REVIEW OF REVISION OPERATIVE NOTES DETERMINED THAT PATIENT HAD SIGNS OF TROCHANTERIC BURSITIS WITH THE TROCHANTERIC BURSA INFLAMED AND SWOLLEN. AS NOTED ONE OF THE CABLES FROM THE PLATE USED FOR A PREVIOUS TROCHANTER/FEMUR FRACTURE WAS BROKEN. A LARGE AREA OF HETEROTOPIC OSSIFICATION WAS NOTED ON REMOVAL OF THE PLATE. AS STATED BY THE SURGEON THIS COULD BE POTENTIALLY FROM THE REMNANTS OF THE TROCHANTER. WITH THE INFORMATION PROVIDED A DEFINITIVE CAUSE FOR THE REPORTED CONDITION COULD NOT BE DETERMINED. |  | Y | N | HEALTH PROFESSIONAL
2648920-2015-00414 | 2015-12-02 | 20151105 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803601 |  | INFORMATION WAS RECEIVED FROM A DISTRIBUTOR WHO IS NOT REQUIRED TO COMPLETE FORM 3500A. (B)(4). OTHER DEVICE USED: CATALOG #00784301236, VERSYS POROUS FULL COAT FEMORAL STEM, LOT #60143491 MANUFACTURED BY ZIMMER (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | Y | HEALTH PROFESSIONAL
2648920-2015-00413 | 2015-12-02 | 20151104 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803204 |  | INFORMATION WAS RECEIVED FROM A DISTRIBUTOR WHO IS NOT REQUIRED TO COMPLETE FORM 3500A.(B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | Y | HEALTH PROFESSIONAL
1822565-2015-02498 | 2015-12-02 | 20140318 | Injury | Manufacturer report | VERSYS HIP SYSTEM FEMORAL HEAD | JDI | ZIMMER INC |  | 00801803604 |  | NO DEVICE OR PHOTOS WERE RECEIVED; THEREFORE THE CONDITION OF THE DEVICE IS UNKNOWN. THE LOT NUMBER OF THE PRODUCT IS UNKNOWN; THEREFORE THE DEVICE HISTORY RECORDS, COMPLAINT HISTORY COULD NOT BE REVIEWED. IT COULD NOT BE CONFIRMED IF THE DEVICE WAS USED IN AN APPROVED AND COMPATIBLE COMBINATION. SURGICAL NOTES WERE NOT PROVIDED. IT COULD NOT BE CONFIRMED IF THE COMPONENTS WERE IMPLANTED WITH THE CORRECT FIT AND ORIENTATION AS PER THE SURGICAL TECHNIQUE. RELEVANT MEDICAL HISTORY AND ADHERENCE TO REHABILITATION PROTOCOL ARE UNKNOWN. A DEFINITIVE ROOT CAUSE CANNOT BE DETERMINED WITH THE INFORMATION PROVIDED. THE REPORTED DEVICE IS USED FOR TREATMENT. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02518 | 2015-12-02 | 20151104 | Malfunction | Manufacturer report | VERSYS PROVISIONAL FEMORAL HEAD | HWT | ZIMMER INC |  | 00789503202 |  | THE FEMORAL HEAD PROVISIONAL WAS RETURNED FOR REVIEW. FEW MARKS ARE NOTED ON THE HEAD TAPER WHICH LIKELY HAPPENED DUE TO THE ASSEMBLY AND REMOVAL OF THE PROVISIONAL HEAD DURING TRIAL REDUCTION. DIMENSIONS WERE FOUND CONFORMING TO THE PRINT SPECIFICATIONS WHERE MEASURED. THE PROVISIONAL HEAD WAS MANUFACTURED ON FEB 3, 2015 AND THEREFORE HAVE A POTENTIAL FIELD AGE OF 1 YEAR 9 MONTHS. PRODUCT HISTORY SEARCH FOR THE PROVISIONAL HEAD REVEALED NO ADDITIONAL COMPLAINTS AGAINST THE RELATED PART AND LOT COMBINATION. A DEFINITE ROOT CAUSE FOR THE REPORTED ISSUE CANNOT BE DETERMINED WITH THE INFORMATION PROVIDED. THIS DEVICE IS USED FOR TREATMENT. |  | N | Y | HEALTH PROFESSIONAL
2648920-2015-00416 | 2015-12-03 | 20151106 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803202 |  | (B)(4). OTHER DEVICE USED: CATALOG #UNK, UNKNOWN EPOCH STEM, LOT #UNK MANUFACTURED BY ZIMMER INC. (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
2648920-2015-00419 | 2015-12-03 | 20151102 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801804005 |  | VISUAL INSPECTION OF THE RETURNED FEMORAL HEAD SHOWS DARK DEBRIS ON THE HEAD TAPER. DIMENSIONS WERE FOUND CONFIRMING TO PRINT SPECIFICATIONS WHERE MEASURED. REVIEW OF THE DEVICE HISTORY RECORDS FOR THE HEAD DID NOT FIND ANY DEVIATIONS OR ANOMALIES. REVIEW OF THE DEVICE HISTORY RECORDS FOR THE STEM CANNOT BE PERFORMED SINCE THE PART AND LOT NUMBERS ARE UNKNOWN. SCANNING ELECTRON MICROSCOPE IMAGES CONFIRM THE PRESENCE OF DARK DEBRIS ON THE HEAD TAPER. EDS ANALYSIS OF THE DARK DEBRIS ON THE HEAD TAPER INDICATED THE ELEMENTS C, O, AL, P, CL, TI, V, CR, CO, AND MO. THIS HAS BEEN A COMMON ELEMENT BREAKDOWN OF THE BLACK DEBRIS FOUND IN HIP FEMORAL CORROSION EVENTS. THE ENERGY-DISPERSIVE X-RAY SPECTROSCOPY ANALYSIS OF THE BASE MATERIAL IS CONSISTENT WITH COCRMO. THESE DEVICES ARE USED FOR TREATMENT. SURGICAL NOTES WERE NOT PROVIDED; IT IS UNKNOWN WHETHER THE COMPONENTS WERE IMPLANTED WITH THE CORRECT FIT AND ORIENTATION AS PER THE SURGICAL TECHNIQUE. A DEFINITIVE ROOT CAUSE FOR THE CORROSION CANNOT BE DETERMINED WITH THE INFORMATION PROVIDED. |  | Y | N | HEALTH PROFESSIONAL
2648920-2015-00418 | 2015-12-03 | 20151109 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803214 |  | INFORMATION WAS RECEIVED FROM A DISTRIBUTOR WHO IS NOT REQUIRED TO COMPLETE FORM 3500A. (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
1822565-2015-02595 | 2015-12-09 | 20140925 | Injury | Manufacturer report | VERSYS FEMORAL STEM | LPH | ZIMMER INC |  | 00786401300 |  | NO DEVICE OR PHOTOS WERE RECEIVED; THEREFORE THE CONDITION OF THE COMPONENTS IS UNKNOWN. REVIEW OF THE DEVICE HISTORY RECORD DID NOT FIND ANY DEVIATIONS OR ANOMALIES. THIS DEVICE IS USED FOR TREATMENT. REVIEW OF THE IMPLANTATION OPERATIVE NOTES CONFIRM THAT THE PATIENT UNDERWENT LEFT TOTAL HIP ARTHROPLASTY DUE TO LEFT FEMORAL NECK FRACTURE. IT WAS NOTED THAT THE TRIALS GAVE EXCELLENT STABILITY. TRIALS WERE REMOVED AND FINAL COMPONENTS WERE IMPLANTS WHICH ALSO GAVE EXCELLENT STABILITY. NO COMPLICATIONS NOTED. REVIEW OF THE REVISION OPERATIVE NOTES CONFIRM THAT THE PATIENT UNDERWENT REVISION LEFT TOTAL HIP ARTHROPLASTY DUE TO SUBJECTIVE INSTABILITY AND ELEVATED COBALT LEVELS. IT WAS NOTED THAT THERE WAS NO EVIDENCE OF CLEAR INFECTION. THERE WAS MODERATE AMOUNT OF METALLOSIS NOTED PARTICULARLY IN THE INFERIOR PORTION OF THE ACETABULUM. THE FEMORAL COMPONENT WAS WELL FIXED AND WAS VERY STABLE. IT WAS NOTED THAT THE FEMORAL COMPONENT WAS REMOVED DUE TO STABILITY REASONS AND AN OSTEOTOMY WAS PERFORMED TO REMOVE THE STEM. IT WAS NOTED THAT THE STEM HAD GREAT INGROWTH TO IT. THE PATIENT ACTIVITY LEVEL AND ADHERENCE TO REHABILITATION PROTOCOL IS UNKNOWN. REVIEW OF THE COMPATIBILITY OF THE DEVICES CONFIRMED THAT THESE ARE AN APPROVED COMPATIBLE COMBINATION. A DEFINITE ROOT CAUSE FOR THE REPORTED ISSUE CANNOT BE DETERMINED WITH THE INFORMATION PROVIDED. |  | Y | N | HEALTH PROFESSIONAL
2648920-2015-00422 | 2015-12-10 | 20151111 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00630505032 |  | (B)(4). OTHER DEVICE USED: CATALOG #00630505032, TRILOGY LONGEVITY CROSSLINKED POLY LINER, LOT #62265633 MANUFACTURED BY ZIMMER (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
2648920-2015-00433 | 2015-12-16 |  | Injury | Manufacturer report | VERSYS FEMORAL HEAD | LPH | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803201 |  | NO DEVICES OR PHOTOS WERE RECEIVED SINCE THE DEVICES STILL REMAIN IMPLANTED; THEREFORE THE CONDITION OF THE COMPONENTS IS UNKNOWN. SURGICAL NOTES WERE NOT PROVIDED; IT IS UNKNOWN WHETHER THE COMPONENTS WERE IMPLANTED WITH THE CORRECT FIT AND ORIENTATION AS PER THE SURGICAL TECHNIQUE. PATIENT ACTIVITY LEVEL AND ADHERENCE TO REHABILITATION PROTOCOL IS UNKNOWN. FDA RECALL Z-1031-2014 CONTAINS THE RELATED FEMORAL HEAD LOT NUMBER 62175268. RECALL WAS DONE SINCE IT WAS DETERMINED THAT THE PACKAGING OPERATIONS WERE NOT PROPERLY VALIDATED; HOWEVER THE SEAL INTEGRITY TESTING CONDUCTED BY ZIMMER CONFIRMED THAT THE PACKAGING TRAY CONFIGURATIONS MET THE ACCEPTANCE CRITERIA FOR STERILE BARRIER INTEGRITY OF BOTH THE INNER AND OUTER TRAYS. A DEFINITE ROOT CAUSE CANNOT BE DETERMINED WITH THE INFORMATION PROVIDED. THESE DEVICES ARE USED FOR TREATMENT. |  | Y | N | HEALTH PROFESSIONAL
2648920-2015-00437 | 2015-12-17 | 20151202 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | LPH | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803203 |  | (B)(4). OTHER DEVICE USED: CATALOG #UNK, ZIMMER EPOCH STEM, LOT #UNK- REMAINS IMPLANTED.THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | Y | HEALTH PROFESSIONAL
2648920-2015-00434 | 2015-12-18 | 20151130 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | LPH | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803202 |  | (B)(4). OTHER DEVICE USED: CATALOG #00630505632, TRILOGY LONGEVITY CROSSLINKED POLY LINER, LOT #60486081 ,MANUFACTURED BY ZIMMER INC. (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | Y | HEALTH PROFESSIONAL
2648920-2015-00439 | 2015-12-18 | 20151125 | Injury | Manufacturer report | VERSYS ENDO FEMORAL HEAD | LWJ | ZIMMER, TURPEAUX INDUSTRIAL PARK | 00781803800 |  |  | INFORMATION WAS RECEIVED FROM A DISTRIBUTOR WHO IS NOT REQUIRED TO COMPLETE FORM 3500A. (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
2648920-2015-00438 | 2015-12-18 | 20151202 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | LPH | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803203 |  | (B)(4).OTHER DEVICE USED:CATALOG #00408801106, ZIMMER EPOCH STANDARD OFFSET STEM, LOT #UNKTHIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
1822565-2016-00012 | 2016-01-05 | 20151208 | Malfunction | Manufacturer report | VERSYS PROVISIONAL FEMORAL HEAD | HWT | ZIMMER INC |  | 00789503602 |  | INFORMATION WAS RECEIVED FROM A FOREIGN SOURCE WHO IS NOT REQUIRED TO COMPLETE FORM 3500A. (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | N | Y | PHYSICIAN
2648920-2016-00002 | 2016-01-06 | 20151210 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | LPH | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801802802 |  | NO DEVICES OR PHOTOS WERE RECEIVED; THEREFORE THE CONDITION OF THE COMPONENTS IS UNKNOWN. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. THESE DEVICES ARE USED FOR TREATMENT. SURGICAL NOTES WERE NOT PROVIDED; IT IS UNKNOWN WHETHER THE COMPONENTS WERE IMPLANTED WITH THE CORRECT FIT AND ORIENTATION AS PER THE SURGICAL TECHNIQUE. THE PATIENT ACTIVITY LEVEL AND ADHERENCE TO REHABILITATION PROTOCOL IS UNKNOWN. PRODUCT HISTORY SEARCH REVEALED NO ADDITIONAL COMPLAINTS AGAINST THE RELATED PART AND LOT COMBINATIONS. A DEFINITE ROOT CAUSE CANNOT BE DETERMINED WITH THE INFORMATION PROVIDED. |  | Y | N | HEALTH PROFESSIONAL
1822565-2016-00055 | 2016-01-08 | 20151207 | Malfunction | Manufacturer report | VERSYS HIP SYSTEM STEM IMPACTOR LONG NOSE | HWA | ZIMMER INC |  | 00789605200 |  | (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | N | Y | HEALTH PROFESSIONAL
2648920-2016-00010 | 2016-01-13 | 20151215 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803602 |  | (B)(4). OTHER DEVICE USED: MANUFACTURED BY ZIMMER (B)(4). CATALOG #00631006436, TRILOGY ACETABULAR LONGEVITY CROSSLINKED LINER LOT #61559823. CATALOG#65771101300, ZIMMER M/L TAPER STEM, LOT NUMBER # 61585414 - REMAINS IMPLANTED. THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
2648920-2016-00011 | 2016-01-13 | 20151216 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | LPH | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803202 |  | (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
2648920-2016-00012 | 2016-01-14 | 20151221 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803602 |  | (B)(4). OTHER DEVICES USED: MANUFACTURED BY ZIMMER INC. (B)(4): CATALOG #00630506036 , TRILOGY LONGEVITY CROSSLINKED LINER, LOT #62224381; CATALOG #00784801200, ZIMMER M/L KINECTIV NECK, LOT #62238279. THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
2648920-2016-00022 | 2016-02-02 | 20151221 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, POROUS UNCEMENTED | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803602 |  | (B)(4). OTHER DEVICE USED: CATALOG #UNK, VERSYS FITMORE STEM, LOT #UNK - MANUFACTURED BY ZIMMER (B)(4) REMAINS IMPLANTED. |  | Y | Y | HEALTH PROFESSIONAL
2648920-2016-00023 | 2016-02-03 | 20160104 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803201 |  | (B)(4). OTHER DEVICE USED: CATALOG #00631005032, TRILOGY CROSSLINKED LINER, LOT 61399506. CATALOG #0771100640, M/L TAPER STEM, LOT #60688539 - REMAINS IMPLANTED, MANUFACTURED BY ZIMMER (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
2648920-2016-00025 | 2016-02-04 | 20160106 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | LPH | ZIMMER, TURPEAUX INDUSTRIAL PARK | N/A | 00801803203 |  | DEVICE HISTORY RECORDS FOR THE PART-LOT COMBINATION OF THE HEAD WAS REVIEWED. NO DEVIATIONS OR ANOMALIES WERE NOTED IN THE MANUFACTURING PROCESS. THE REPORTED DEVICE IS USED FOR TREATMENT. COMPATIBILITY COULD NOT BE ASSESSED AS PART AND LOT INFORMATION OF THE STEM, SHELL USED IN THE CONSTRUCT ARE UNKNOWN. REVIEW OF THE COMPLAINT HISTORY INDICATED NO PRIOR COMPLAINTS FOR THE PART-LOT COMBINATIONS ASSOCIATED WITH THE REPORTED HEAD. A DEFINITIVE ROOT CAUSE CANNOT BE DETERMINED WITH THE INFORMATION PROVIDED. |  | Y | Y | HEALTH PROFESSIONAL
2648920-2016-00027 | 2016-02-09 | 20150113 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801804003 |  | INFORMATION WAS RECEIVED FROM A DISTRIBUTOR WHO IS NOT REQUIRED TO COMPLETE FORM 3500A. (B)(4). OTHER DEVICE USED: CATALOG #UNK, UNKNOWN ZIMMER STEM, LOT #UNK- REMAINS IMPLANTED. THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | Y | HEALTH PROFESSIONAL
2648920-2016-00034 | 2016-02-10 | 20151222 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | PROSTHESIS, HIP | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803602 |  | (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | Y | HEALTH PROFESSIONAL
1822565-2016-00289 | 2016-02-16 |  | Injury | Manufacturer report | VERSYS FEMORAL STEM BEADED FULL COAT | LZO | ZIMMER INC |  | 00784301526 |  | INFORMATION WAS RECEIVED FROM A CONSUMER WHO IS NOT REQUIRED TO COMPLETE FORM 3500A. (B)(4). OTHER DEVICE USED: THE FOLLOWING WERE MANUFACTURED BY MANUFACTURED BY ZIMMER (B)(4); CATALOG #00620006622, TRILOGY ACETABULAR CLUSTER SHELL , LOT #23902000 CATALOG #00630506640, TRILOGY CROSSLINKED LINER , LOT #60611720; CATALOG #00801804003, VERSYS FEMORAL HEAD , LOT #60774059. THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
2648920-2016-00039 | 2016-02-16 | 20160118 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801804003 |  | (B)(4).OTHER DEVICE USED:CATALOG #UNK, UNKNOWN ZIMMER BOWED EPOCH STEM, LOT #UNK- REMAINS IMPLANTEDTHIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
2648920-2016-00042 | 2016-02-16 | 20160114 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | PROSTHESIS, HIP | ZIMMER B.V. | N/A | 00801803203 |  | (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | PHYSICIAN
2648920-2016-00050 | 2016-02-18 | 20160112 | Malfunction | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK | N/A | 00801803203 |  | THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | N | Y | HEALTH PROFESSIONAL
1822565-2016-00381 | 2016-02-23 | 20160108 | Injury | Manufacturer report | VERSYS FEMORAL STEM | MEH | ZIMMER INC |  | 65764501400 |  | (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
2648920-2016-00054 | 2016-02-24 | 20150903 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803201 |  | (B)(4). OTHER DEVICE USED: CATALOG #00771100700, ZIMMER M/L TAPER FEMORAL STEM, LOT #61694540 MANUFACTURED BY ZIMMER INC. (B)(4). REMAINS IMPLANTED. THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
1822565-2016-00382 | 2016-02-25 |  | Injury | Manufacturer report | VERSYS FEMORAL STEM | LPH | ZIMMER INC |  | 00784301708 |  | (B)(4). OTHER DEVICES USED: CATALOG #00875101640, TRILOGY LONGEVITY POLY LINER, LOT #62998324- IMPLANTED (B)(6) 2015; CATALOG #00875706401, CONTINUUM SHELL WITH CLUSTER HOLES, LOT #61442591; THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
2648920-2016-00059 | 2016-02-25 | 20160127 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803602 |  | OTHER DEVICE USED: CATALOG #00630505836, TRILOGY CROSSLINKED LINER, LOT #61543444; CATALOG #00784801200, M/L KINECTIV NECK, LOT #61513128 - MANUFACTURED BY ZIMMER (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
1822565-2016-00404 | 2016-02-26 | 20151022 | Injury | Manufacturer report | VERSYS PROVISIONAL FEMORAL HEAD | HWT | ZIMMER INC |  | 00789503202 |  | INFORMATION WAS RECEIVED VIA MW5058876. (B)(4) OTHER DEVICE USED: CATALOG #00789503202, VERSYS PROVISIONAL FEMORAL HEAD, LOT #UNK. THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
2648920-2016-00068 | 2016-02-26 |  | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801802802 |  | INFORMATION WAS RECEIVED FROM A CONSUMER WHO IS NOT REQUIRED TO COMPLETE FORM 3500A.(B)(4). OTHER DEVICES USED: CATALOG #00783401500, VERSYS FEMORAL STEM, LOT #61878471; CATALOG #00500104800, BIPOLAR METAL SHELL, LOT #62158173; CATALOG #00500104728, MULTIPOLAR BIPOLAR CUP LINER, LOT #62203207. THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
2648920-2016-00067 | 2016-02-26 | 20140620 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | PROSTHESIS, HIP | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803203 |  | INFORMATION WAS RECEIVED FROM A CONSUMER WHO IS NOT REQUIRED TO COMPLETE FORM 3500A. (B)(4). OTHER DEVICE USED: CATALOG #00771100920, ZIMMER M/L TAPER STEM, LOT #60447574 MANUFACTURED BY ZIMMER INC. (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | Y | HEALTH PROFESSIONAL
2648920-2016-00066 | 2016-02-26 | 20160201 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801804003 |  | INFORMATION WAS RECEIVED FROM A DISTRIBUTOR WHO IS NOT REQUIRED TO COMPLETE FORM 3500A.(B)(4). OTHER DEVICE USED: CATALOG #00630505840, ZIMMER TRILOGY ACETABULAR LINER, LOT #62815057 MANUFACTURED BY ZIMMER INC. (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. H3 OTHER TEXT : 81 |  | Y | N | HEALTH PROFESSIONAL
1822565-2016-00459 | 2016-03-02 |  | Injury | Manufacturer report | UNKNOWN ZIMMER VERSYS FEMORAL STEM | HIP PROSTHESIS | ZIMMER INC |  |  |  | INFORMATION WAS RECEIVED VIA PUBLISHED LITERATURE. PLEASE REFERENCE LITERATURE AT THE FOLLOWING LOCATION: HTTP://WWW.NCBI.NLM.NIH.GOV/PUBMED/26724809. (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
2648920-2016-00070 | 2016-03-02 | 20160203 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803201 |  | (B)(4). OTHER DEVICE USED: CATALOG #00784501600, VERSYS FEMORAL STEM, LOT #60300060 REMAINS IMPLANTED, MANUFACTURED BY ZIMMER INC. (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
1822565-2016-00562 | 2016-03-03 |  | Injury | Manufacturer report | UNKNOWN ZIMMER VERSYS FEMORAL STEM | HIP PROSTHESIS | ZIMMER INC |  |  |  | (B)(4).THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
1822565-2016-00561 | 2016-03-03 |  | Injury | Manufacturer report | UNKNOWN ZIMMER VERSYS FEMORAL STEM | HIP PROSTHESIS | ZIMMER INC |  |  |  | INFORMATION WAS RECEIVED VIA PUBLISHED LITERATURE. PLEASE REFERENCE LITERATURE AT THE FOLLOWING LOCATION: HTTP://WWW.NCBI.NLM.NIH.GOV/PUBMED/26709566. (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
2648920-2016-00074 | 2016-03-04 | 20160203 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | LPH | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803202 |  | INFORMATION WAS RECEIVED FROM A DISTRIBUTOR WHO IS NOT REQUIRED TO COMPLETE FORM 3500A. (B)(4). OTHER DEVICE USED: CATALOG #00784801300, M/L TAPER KINECTIV NECK, LOT #61548982. THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
2648920-2016-00075 | 2016-03-08 | 20160129 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801804002 |  | THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. RECEIVED, BUT NOT YET EVALUATED. |  | Y | N | HEALTH PROFESSIONAL
1822565-2016-00717 | 2016-03-18 | 20150729 | Injury | Manufacturer report | VERSYS FULLCOAT COLLARED FEMORAL STEM | LPH | ZIMMER INC |  | 00784301456 |  | INFORMATION WAS RECEIVED FROM A CONSUMER WHO IS NOT REQUIRED TO COMPLETE FORM 3500A. (B)(4). OTHER DEVICES USED: CATALOG #00620205020, TRABECULAR METAL SHELL WITH MULTI HOLES, LOT #62960497; CATALOG #00630505036, TRILOGY LONGEVITY CROSSLINKED LINER, LOT #62902909; CATALOG #0801803602, VERSYS FEMORAL HEAD, LOT #62694696 MANUFACTURED BY (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
1822565-2016-00728 | 2016-03-18 |  | Injury | Manufacturer report | VERSYS EPOCH FULLCOAT FEMORAL STEM | HIP PROSTHESIS | ZIMMER INC | N/A | 00408801206 |  | INFORMATION WAS RECEIVED FROM A CONSUMER WHO IS NOT REQUIRED TO COMPLETE FORM 3500A. (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | Y | HEALTH PROFESSIONAL
2648920-2016-00020 | 2016-03-21 | 20150630 | Injury | Manufacturer report | ZIMMER VERSYS FEMORAL HEAD | LPH | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803202 |  | (B)(4).THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
2648920-2016-00084 | 2016-03-22 | 20160217 | Injury | Manufacturer report | VERSYS FEMORAL HEAD | LPH | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00801803602 |  | INFORMATION WAS RECEIVED FROM A DISTRIBUTOR WHO IS NOT REQUIRED TO COMPLETE FORM 3500A. (B)(4). OTHER DEVICES USED: CATALOG # (B)(4), TRILOGY LONGEVITY CROSSLINKED POLY LINER, LOT #61149077; CATALOG # (B)(4), ZIMMER M/L KINECTIV NECK, LOT #60969523 MANUFACTURED BY ZIMMER INC. (B)(6). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL
1822565-2016-00754 | 2016-03-23 | 20160226 | Malfunction | Manufacturer report | VERSYS FEMORAL STEM BEADED FULLCOAT | LPH | ZIMMER INC |  | 00784301506 |  | (B)(4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | N | Y | HEALTH PROFESSIONAL
2648920-2016-00087 | 2016-03-24 |  | Injury | Manufacturer report | VERSYS FEMORAL HEAD | JDI | ZIMMER, TURPEAUX INDUSTRIAL PARK |  | 00902602800 |  | (B)(4). OTHER DEVICES USED: CATALOG #00632004828, TRILOGY LONGEVITY CROSSLINKED POLY LINER, LOT #60468570; MANUFACTURED BY ZIMMER INC. (B)(4). CATALOG #00620004820, TRILOGY SHELL WITH POROUS HOLES, LOT #60586280; CATALOG #UNK, UNKNOWN ZIMMER FEMORAL STEM, LOT #UNK. THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. |  | Y | N | HEALTH PROFESSIONAL

## Monthly Summary
year_month | count
2015-09 | 3
2015-10 | 55
2015-11 | 13
2015-12 | 18
2016-01 | 11
2016-02 | 22
2016-03 | 14

## Event Types
Event Type | Count
Injury | 124
Malfunction | 11
Death | 1

## Brand Names
Brand Name | Count
VERSYS FEMORAL HEAD | 56
UNKNOWN VERSYS ET STEM | 18
UNKNOWN M/L TAPER STANDARD OFFSET STEM | 8
UNKNOWN VERSYS BEADED FULLCOAT PLUS STEM | 6
UNKNOWN M/L TAPER EXTENDED OFFSET STEM | 3
ZIMMER M/L TAPER FEMORAL STEM | 3
ZIMMER M/L TAPER STEM | 3
UNKNOWN ZIMMER VERSYS FEMORAL STEM | 3
VERSYS FEMORAL STEM | 3
VERSYS PROVISIONAL FEMORAL HEAD | 3
VERSYS HIP SYSTEM FEMORAL HEAD | 2
VERSYS FEMORAL STEM BEADED FULLCOAT | 2
UNKNOWN ZIMMER VERSYS STEM | 2
UNKNOWN ZIMMER M/L TAPER STEM | 2
ZIMMER M/L TAPER HIP FEMORAL STEM | 2
M/L TAPER FEMORAL STEM | 2
ZIMMER M/L TAPER HIP STEM | 1
ZIMMER M/L TAPER HIP PROSTHESIS FEMORAL STEM | 1
VERSYS EPOCH FULLCOAT FEMORAL STEM | 1
VERSYS FULLCOAT COLLARED FEMORAL STEM | 1

## Case Metadata
Case Name | MDL Number | Filing Date | Court | Status | Settlement | Data Window Start | Data Window End | Total Records | Search Queries Used
Zimmer M/L Taper Hip | 2716 | 2016-03-24 | N.D. Indiana | Active | N/A | 2015-09-23 | 2016-03-24 | 136 | 3