## All Reports
report_number | date_received | date_of_event | event_type | report_source_code | brand_name | generic_name | manufacturer_d_name | model_number | catalog_number | device_class | event_description | manufacturer_narrative | adverse_event_flag | product_problem_flag | device_operator | manufacturer_g1_name | year_month
9616680-2012-00478 | 2012-06-27 | 2012-06-05 | I, n, j, u, r, y | Manufacturer report | UNK REJUVENATE MODULAR NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT, "THE SURGEON NOTED CORROSION AND METALLOSIS. THE PT WAS REVISED." | AN EVAL OF THE DEVICES CANNOT BE PERFORMED AS THE DEVICES ARE CURRENTLY WITH HOSP PATHOLOGY AND WERE NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICES REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD DEVICES OR ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS MFR #2249697-2012-00813. | Y | N | HEALTH PROFESSIONAL |  | 2012-06
2249697-2012-00813 | 2012-06-27 | 2012-06-05 | I, n, j, u, r, y | Manufacturer report | UNK REJUVENATE STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, "THE SURGEON NOTED CORROSION AND METALLOSIS. THE PT WAS REVISED." | AN EVAL OF THE DEVICES CANNOT BE PERFORMED AS THE DEVICES ARE CURRENTLY WITH HOSPITAL PATHOLOGY AND WERE NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICES REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD DEVICES OR ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS MFR #9616680-2012-00478. | Y | N | HEALTH PROFESSIONAL |  | 2012-06
2249697-2012-00831 | 2012-06-27 | 2012-06-06 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT THE PT WAS HAVING PAIN SO THE SURGEON REVISED. | IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. WHEN COMPLETED, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-06
2249697-2012-00817 | 2012-06-27 | 2012-06-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | REJUVENATE STEM INSERTER | INSTRUMENT | STRYKER ORTHOPAEDICS MAHWAH | NA | 1601-1600 |  | IT WAS REPORTED THAT, "IMPACTOR DOESN'T FIT STEM." | SUMMARY OF EVAL: VISUAL INSPECTION CONFIRMED THE REPORTED EVENT. THE DEVICE RETURNED IN USED CONDITION, IMPACTION MARKS WERE NOTED ON THE STRIKE PLATE CONSISTENT WITH NORMAL USE. THE SPLIT COLLET HAS BEEN DEFORMED AND FRACTURED AS A RESULT OF IMPROPER SEATING. THE PRODUCT EXPERIENCING REPORTING DATABASE SHOWS THIS EVENT IS RELATED TO A TREND OF SIMILAR EVENTS WHERE THE INSERTER FAILS TO FULLY ENGAGE WITH THE STEM. THE ROOT CAUSE WAS PRESUMED TO BE DEFORMATION OF THE SPLIT COLLAR OF THE STEM INSERTER DUE TO IMPROPER SEATING OF THE STEM IN THE INSERTION FEATURE. IT WAS ALSO DETERMINED THAT THE REQUIREMENTS OF THE GAGE DRAWING WERE NOT MET. | N | Y | HEALTH PROFESSIONAL |  | 2012-06
2249697-2012-00836 | 2012-06-27 | 2011-09-01 | I, n, j, u, r, y | Manufacturer report | C-TAPER COCR LFIT HEAD 32MM/0 | HIP JOINT FEMORAL (HEMI-HIP) METALLIC CEMENTED OR UNCEMENTED PROSTHESIS. | STRYKER ORTHOPAEDICS-MAHWAH | NA | 06-3200 |  | IT WAS REPORTED THAT, METALLOSIS, PAIN, HIP REVISION. | AN EVALUATION OF THE DEVICES CANNOT BE PERFORMED AS THE DEVICES WERE RETAINED BY THE PT AND WERE NOT RETURNED TO THE MANUFACTURER. DEVICE HISTORY REVIEW DETERMINED ALL DEVICES IN THE REPORTED LOTS WERE ACCEPTED INTO FINAL STOCK WITH NO REPORTED DISCREPANCIES. COMPLAINT HISTORY REVIEW FOUND NO PREVIOUSLY REPORTED SIMILAR EVENTS FOR THE PRODUCT LOT, CATALOG NUMBER, OR FAMILY. ADDITIONAL INFO PERTAINING TO THE DEVICES REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME EVENT AS MFR #2249697-2012-00835. | Y | N | HEALTH PROFESSIONAL |  | 2012-06
2249697-2012-00835 | 2012-06-27 | 2011-09-01 | I, n, j, u, r, y | Manufacturer report | OMNIFIT EON 132 | HIP JOINT METAL/POLYMER SEMI-CONSTRAINED CEMENTED PROSTHESIS. | STRYKER ORTHOPAEDICS-MAHWAH | NA | 6098-0735 |  | IT WAS REPORTED THAT, METALLOSIS, PAIN, HIP REVISION. | AN EVALUATION OF THE DEVICES CANNOT BE PERFORMED AS THE DEVICES WERE RETAINED BY THE PT AND WERE NOT RETURNED TO THE MANUFACTURER. DEVICE HISTORY REVIEW DETERMINED ALL DEVICES IN THE REPORTED LOTS WERE ACCEPTED INTO FINAL STOCK WITH NO REPORTED DISCREPANCIES. COMPLAINT HISTORY REVIEW FOUND NO PREVIOUSLY REPORTED SIMILAR EVENTS FOR THE PRODUCT LOT, CATALOG NUMBER, OR FAMILY. ADDITIONAL INFO PERTAINING TO THE DEVICES REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME EVENT AS MFR #2249697-2012-00836. | Y | N | HEALTH PROFESSIONAL |  | 2012-06
2249697-2012-00840 | 2012-06-29 | 2012-06-11 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT, PAIN AND SWELLING. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: DELTA V-40 CERAMIC HEAD 36/-2, 5, CAT# 6570-0-436, LOT# 32286202. TRIDENT 10 X3 INSERT 36MM ID, CAT# 623-10-36E, LOT# MJNXH0. TRIDENT PSL HA SOLID BACK 52MM, CAT# 540-11-52E, LOT# 33421302. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. WHEN COMPLETED, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-06
2249697-2012-00845 | 2012-07-03 | 2012-06-12 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRIGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT, THE PT WAS REVISED FOR A PAINFUL HIP. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0DEG 30MM, CAT# NLS-300000B, LOT# 35683501; PRIMARY TRITANIUM HEMI CLUSTER HOLE CUP, 58MM, CAT# 502-03-58F, LOT# MJM78A; TRIDENT 0 DEG INSERT 40MM, CAT# 623-00-40F, LOT# MKDT7P; V40 COCR LFIT HEAD 40MM/+0, CAT# 6260-9-140, LOT# MHN2K; 6.5 CANCELLOUS BONE SCREW 25MM, CAT# 2030-6525-1, LOT# MKK1EV; 6.5 CANCELLOUS BONE SCREW 35MM, CAT# 2030-6535-1, LOT# MKA0Y7. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVAL OF THE DEVICES CANNOT BE PERFORMED AS THE DEVICES WERE RETAINED BY THE SURGEON AND WERE NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFO PERTAINING TO THE DEVICES REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED, BUT NOT MADE AVAILABLE. SHOULD ADDITIONAL INFO BECOMES AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-00856 | 2012-07-03 | 2012-06-13 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT, PT PRESENTED WITH HIP PAIN APPROX 10 MONTHS POST-OP X-RAY AND SHOWED A LESION ON LATERAL SIDE OF CUP. BLOOD WORK-UP WAS POSITIVE FOR METAL LEVEL. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD DEVICE OR ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME EVENT AS: MFR# 9616680-2012-00480, 2249697-2012-00857. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-00867 | 2012-07-05 | 2012-02-12 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE LEFT HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | . | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-00897 | 2012-07-10 | 2012-06-20 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT, PT CAME TO DR. (B)(6) WITH PAIN AND SWELLING SO REVISING FOR ALTR. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 8DEG 34MM, CAT# NLV-340800G, LOT# 32987801. DELTA V-40 CERAMIC HEAD 36/0, CAT #6570-0-136, LOT# 35973101. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVALUATION OF THE DEVICES CANNOT BE PERFORMED AS THE DEVICE WERE RETAINED BY THE DOCTOR AND WERE NOT RETURNED TO THE MANUFACTURER. REVIEW OF THE DEVICE HISTORY RECORDS INDICATES ALL DEVICE WERE ACCEPTED INTO FINAL STOCK MET SPECIFICATIONS. A REVIEW OF PRODUCT EXPERIENCE REPORTING DATABASE INDICATES THAT THERE HAVE BEEN NO OTHER REPORTED EVENTS FOR THE REPORTED PRODUCT LOTS. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED, BUT NOT MADE AVAILABLE. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-00880 | 2012-07-10 | 2012-06-15 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 10 | IMPLANT | STRYKER ORTHOPAEDICS | NA | SPT-100000S |  |  |  | Y | N | HEALTH PROFESSIONAL |  | 2012-07
9616680-2012-00516 | 2012-07-12 | 2012-05-03 | I, n, j, u, r, y | Manufacturer report | REJUVENATE NECK 30MM | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | NLS-300000B |  | USER FACILITY REPORT # (B)(4), RECEIVED (B)(4) 2012, INDICATED: (B)(6) FEMALE HAD LTHR REVISED (B)(6) 2012 FOR METAL REACTION S/P LTHR (B)(6) 2011, FOR OA (STRYKER REJUVENATE IMPLANTS). LTHR REVISION (B)(6) 2012, W/O COMPLICATION. HISTOPATHOLOGY FINDINGS CONSISTENT W/ IMMUNOLOGICALLY MEDIATED IMPLANT DEBRIS REACTION. ALVAL SCORE IS 10/10 (SYNOVIAL LING 3 INFLAMMATORY INFILTRATE 4 TISSUE ORGANIZATION 3). THERE IS MARKED LYMPHOCYTIC INFILTRATE ASSOCIATED W/ NECROTIC MACROPHAGES IN MARROW W/ FORMATION ACTIVE GERMINAL CENTERS. | THE REVISION WAS ORIGINALLY REPORTED UNDER MFR REPORT #9616680-2012-00384. THE UF REPORT WAS RECEIVED ON (B)(4) 2012, ALONG WITH UF REPORT # (B)(4). A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. ADDITIONAL INFORMATION FROM THE USER FACILITY REPORT: COMMON DEVICE NAME: MODULAR NECK, FACILITY: STRYKER ORTHOPAEDICS, (B)(4), MODEL NUMBER: NLS300000B, OTHER NUMBER: 127/132 0 30 MM, DEVICE AVAILABLE FOR EVALUATION: YES. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-00898 | 2012-07-12 | 2012-05-03 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | SPT070000S | SPT-070000S |  | USER FACILITY REPORT # (B)(4), RECEIVED (B)(4) 2012, INDICATED: (B)(6) FEMALE HAD LTHR REVISED (B)(4) 2012, FOR METAL REACTION S/P LTHR (B)(4) 2011, FOR OA (STRYKER REJUVENATE IMPLANTS). LTHR REVISION (B)(6) 2012, W/O COMPLICATION. HISTOPATHOLOGY FINDINGS CONSISTENT W/ IMMUNOLOGICALLY MEDIATED IMPLANT DEBRIS REACTION. ALVAL SCORE IS 10/10 (SYNOVIAL LING 3 INFLAMMATORY INFILTRATE 4 TISSUE ORGANIZATION 3). THERE IS MARKED LYMPHOCYTIC INFILTRATE ASSOCIATED W/ NECROTIC MACROPHAGES IN MARROW W/ FORMATION ACTIVE GERMINAL CENTERS. | THE REVISION WAS ORIGINALLY REPORTED UNDER MFR REPORT # 9616680-2012-00384. THE ATTACHED UF REPORT WAS RECEIVED ON (B)(4) 2012, ALONG WITH UF REPORT # (B)(4). A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. (B)(6). | Y | Y | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-00902 | 2012-07-12 | 2012-06-20 | I, n, j, u, r, y | Manufacturer report | UNK SIZE 7 REJUVENATE STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, "PRE-OPERATIVE MRI REVEALED A LARGE PSEUDO TUMOR. SURGEON REVISED. LAB: CR=0.6, TI=3.4" | AN EVAL OF THE DEVICES CANNOT BE PERFORMED AS THE DEVICES WERE RETAINED BY THE HOSPITAL AND WERE NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICES REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME EVENT AS: MFR #9616680-2012-00517. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
9616680-2012-00517 | 2012-07-12 | 2012-06-20 | I, n, j, u, r, y | Manufacturer report | UNK REJUVENATE 34 NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT, "PRE-OPERATIVE MRI REVEALED A LARGE PSEUDOTUMOR. SURGEON REVISED. LABS: CR=0.6, TI=3.4". | AN EVAL OF THE DEVICES CANNOT BE PERFORMED AS THE DEVICES WERE RETAINED BY THE HOSPITAL AND WERE NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICES REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME EVENT AS: MFR #2249697-2012-00902. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-00909 | 2012-07-13 | 2012-06-20 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT, "THE PT WAS REVISED DUE TO LOOSE, SUBSIDED IMPLANTS." | ADDITIONAL INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. WHEN COMPLETED, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-00917 | 2012-07-17 | 2012-06-24 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRIGTH PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS ¿ MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT, "THE PT HAD AN INFECTION SO THE IMPLANTS WERE REMOVED. NO IMPLANTS WERE REIMPLANTED AT THIS TIME. UNK SCREWS ALSO REMOVED." | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0DEG 38 MM, CAT# NLS-380000B, LOT# 23327203. UNK CUP, UNK HEAD, UNK LINER, UNK SCREWS. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S INFECTION. AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-00933 | 2012-07-17 | 2012-06-27 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS ¿ MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT, "PT EXPERIENCING GROIN PAIN." | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0DEG 34MM, CAT# NLS-340000B, LOT#35956002; 28MM STD LFIT V40 HEAD, CAT# 6260-9-128, LOT# 36464503; RESTORATION ADM CUP W/HA, CAT# 1235-2-481, LOT# G3044872; RESTORATION ADM X 3 INS 28/48, CAT# 1236-2-848, LOT# 35332201. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR DUE TO HOSPITAL POLICY. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-00922 | 2012-07-17 | 2011-03-08 | I, n, j, u, r, y | Manufacturer report | UNK REJUVENATE STEM - HIP RIGHT | IMPLANT | STRYKER ORTHOPAEDICS ¿ MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PT STATED HE HAD A RIGHT HIP REPLACEMENT AND AT FIRST HE FELT IT WAS A SUCCESS. THEN HE STARTED HAVING TROUBLE SLEEPING, HE COULDN'T ROLL ONTO HIS RIGHT HIP IT WAS VERY PAINFUL. HE WENT THE DR AND HAD AN MRI -THERE WAS FLUID ON THE HIP. THE DR ASPIRATED 35 CC OF GREENISH FLUID. DR STATED THAT THERE WAS TROUBLE WITH THE REJUVENATE STEM. PT CONTINUES TO HAVE SWELLING, HE FEELS WEAK AT TIMES AND THE FLUID CONTINUES TO BUILD UP. PT HAD MORE TESTING AND DR FOUND COBALT METAL IN HIS BLOOD STREAM. PT STATES HIS BODY IS REJECTING THE IMPLANT AND NOW THE DR WILL MONITOR THE FLUID FOR AWHILE BUT A REVISION IS EMINENT. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-00943 | 2012-07-17 | 2012-06-29 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT, THE PATIENT WAS REVISED DUE TO PAIN. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0 DEG 34MM, CAT# NLS-340000B, LOT# NLS-340000B; DELTA V-40 CERAMIC HEAD 36/+2,5, CAT# 6570-0-536, LOT# 34887001; TRIDENT 10 DEG X3 INSERT 36MM ID, CAT# 623-10-36E, LOT# MJT49H. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S PAIN. WHEN COMPLETED, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
MW5026252 | 2012-07-18 | 2011-09-23 | I, n, j, u, r, y | Voluntary report | STRYKER REJUVENATE MODULAR FEMORAL STEM | REJUVENATE MODULAR FEMORAL STEM | STRYKER |  |  |  | RIGHT TOTAL HIP REPLACEMENT ON ABOVE DATE WITH STRYKER MOBILE BEARING HIP INCLUDING REJUVENATE MODULAR FEMORAL STEM. HOSPITALIZATION (B)(6) 2011 FOR RIGHT BUTTOCK CELLULITIS WITH HIGH FEVER, VOMITING AND RASH OVER RIGHT BUTTOCK STARTING AT THR INCISION. SINCE THEN PAIN IN RIGHT HIP OF VARIABLE DEGREES, LOCALIZED REDNESS AND SWELLING ALONG INCISION. SURGERY SCHEDULED (B)(6) 2012 TO REMOVE THR WAS CANCELLED DUE TO PROBLEMS WITH RIGHT SHOULDER (PREVIOUSLY REPLACED 2007 COMPLICATED BY P ACNES INFECTION), ANTIBIOTIC SPACER FROM 2008 STILL IN PLACE. TREATMENT PLAN STILL UNDECIDED AT THIS TIME. RIGHT HIP OA. |  | Y | N | 0 |  | 2012-07
2249697-2012-00958 | 2012-07-19 | 2012-07-02 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT, REVISION OF THE RIGHT TOTAL HIP. PATIENT WAS EXPERIENCING PAIN. THEREFORE, SURGEON DECIDED TO REVISE. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0DEG 30MM, CAT# NLS-300000B, LOT# 32904301; TRIDENT 0DEG X3 INSERT 36MM ID, CAT# 623-00-36E, LOT# MKN7MY; DELTA V-40 CERAMIC HEAD 36/+5, CAT# 6570-0-236, LOT# 31808901. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S PAIN. WHEN COMPLETED, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-00955 | 2012-07-19 | 2012-07-02 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT, "PAINFUL, SWELLING HIP." | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: REJUVENATE MOD NECK OV 127/132 BLK 34 MM, CAT# NLS-340000B, LOT# 33271202; TRIDENT 0 DEG INSERT 40MM, CAT# 623-00-40E, LOT# MJMPAH; V40 COCR LIFT HEAD 40MM/+4, CAT# 6260-9-240, LOT# METKMX. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S PAIN. WHEN COMPLETED, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-00959 | 2012-07-19 | 2012-07-03 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT, "THE PATIENT COMPLAINED OF GROIN PAIN SO THE SURGEON REVISED THE LEFT HIP. SCREW ALSO REMOVED." | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0DEG 30MM, CAT# NLS-300000B, LOT# 37106301; TRIDENT PSL HA CLUSTER 48MM, CAT# 542-11-48D, LOT# MKH57D; TRIDENT 0 DEG X3 INSERT 32MM HEAD, CAT# 623-00-32D, LOT# MKD5XX; 32MMSTED LFIT V40 HEAD, CAT# 6260-9-132, LOT# 36317303; 6.5 CANCELLOUS BONE SCREW 30MM, CAT# 2030-6530-1, LOT# MKKEPK. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE PATIENTS MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S PAIN. WHEN COMPLETED, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-00963 | 2012-07-20 | 2012-07-03 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT, "ADVERSE LOCAL TISSUE REACTION." | AN EVALUATION OF THE DEVICES CANNOT BE PERFORMED AS THE DEVICES WERE RETAINED BY THE DOCTOR AND WERE NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD DEVICE OR ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS: MFR # 9616680-2012-00552. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-00964 | 2012-07-20 | 2012-07-03 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT, "PATIENT HAS GROIN AND HIP PAIN." | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0DEG 38MM: CAT# NLS380000B, LOT# 29438402. RESTORATION (TM) ADM. CUP W/HA: CAT# 1235-2-581, LOT# G2943827. RESTORATION ADM X3 INS 28/58: CAT# 1236-2-858, LOT# 34238401. 28MM STD LFIT V40 HEAD, 6: CAT# 260-9-128, LOT# 36858404. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER DUE TO HOSPITAL POLICY. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-01014 | 2012-07-27 | 2012-07-09 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT, PT HAD ONSET OF HIP PAIN ON (B)(6) 2010. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 8DEG 30MM, CAT# NLV-300800G, LOT# 31617901. TRIDENT PSL HA CLUSTER 54MM, CAT# 542-11-54F, LOT# MJELM8. TRIDENT 0 DEG INSERT 40MM, CAT# 623-00-40F, LOT# MJHPKJ. UNI ADAPTOR SLEEVE V40 TI, CAT# 6519-T-100, LOT# 3463081. DELTA UN FEM HD 40MM, CAT# 6519-1-040, LOT# 33669601. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-01006 | 2012-07-27 | 2011-04-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN PRODUCT REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PT IS A QUADRIPLEGIC AND HAS REDUCED SENSITIVITY FROM THE NECK DOWN. HE HAD A HIP REPLACEMENT IN 2011 AND HE STATED HIS RECOVERY WAS VERY DIFFICULT. HE ALSO STATED THAT HE QUESTIONED HIS SURGEON THOROUGHLY ABOUT THE CORROSION ISSUES BECAUSE HE KNEW ABOUT IT FROM OUR COMPETITORS. WHEN HE FEELS PAINS IT IS EXTREME BUT ON THE OTHER HAND AS AN EXAMPLE IF YOU STICK A FORK IN HIS THIGH HE WILL NOT FEEL IT. PT HAD METAL SENSITIVITY TEST PRIOR TO SURGERY BUT HE IS NOW VERY CONCERNED ABOUT THIS RECALL AND HOW IT WILL AFFECT HIM. PT WILL SUPPLY MEDICAL INFO TO STRYKER. PT WILL BE SPEAKING TO LEGAL COUNSEL. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-01010 | 2012-07-27 | 2011-03-19 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT HIP REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, SEVERAL MONTHS AFTER SURGERY PT STARTED HAVING PROBLEMS. PT CLAIMS THAT PAIN HAS ALWAYS BEEN PRESENT FROM HIP TO KNEE. HOWEVER, HE STATES THAT SITTING AND ELEVATION OF LEGS IS "OK." PT NEEDS REVISION AND ITS SCHEDULED FOR AN UPCOMING REVISION ON (B)(6) 2012 WITH DR BONENBERGER. HE OBTAINED A SECOND OPINION FROM ANOTHER DOCTOR, WHO FOUND LOOSENING. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-01015 | 2012-07-27 | 2012-07-09 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT, "DOCTOR STATES PT STARTED HAVING HIP PAIN (B)(6) 2011." | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 8DEG 42MM, CAT# NLV-420800G, LOT# 32440702. TRIDENT PSL HA CLUSTER 56MM, CAT# 542-11-56F, LOT# MJM2HK. TRIDENT 0 DEG INSERT 40MM, CAT# 623-00-40F, LOT# MJNEJX. UNI ADAPTOR SLEEVE V40 TI, CAT# 6519-T-100, LOT# 34630807. DELTA UN FEM HD 40MM, CAT# 6519-1-040, LOT# 35197001. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-01009 | 2012-07-27 | 2012-01-05 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, SINCE PT'S SURGERY ON (B)(6) 2012 HE HAS NEVER HEALED AND HAS PAIN. MRI SHOWED FLUID IN IMPLANT AND INFLAMED TISSUE. COBALT LEVEL IS 27.2, CHROMIUM LEVEL IS 2. PT IS SCHEDULED FOR A REVISION ON (B)(6) 2012. PER PT, SURGEON WANTS TO REMOVE ALL IMPLANTS BUT PLANS ON KEEPING THE SHELL. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-01017 | 2012-07-27 | 2012-07-09 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT, "DR (B)(6) STATES THE PT HAD ONSET OF HIP PAIN STARTING (B)(6) 2011." | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0DEG 38MM, CAT# NLS-380000B, LOT# 2494501. TRIDENT PSL HA CLUSTER 56MM, CAT# 542-11-56F, LOT# MHN8KA. TRIDENT 0X3 INSERT 36MM ID, CAT# 623-00-36F, LOT# MHMM9M. DELTA V-40 CERAMIC HEAD 36/0, CAT# 6570-0-136, LOT# 30758501. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-01018 | 2012-07-27 | 2012-07-09 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED: "DR (B)(6) STATED THE PT HAD ONSET OF PAIN STARTING (B)(6) 2011." | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0DEG 34MM, CAT# NLS-340000B, LOT# 29380602. TRIDENT PSL HA CLUSTER 56MM, CAT# 542-11-56F, LOT# MJA2K9. TRIDENT 0 X 3 INSERT 36MM ID, CAT# 623-00-360F, LOT# MJATV8. DELTA V-40 CERAMIC HEAD 36/0, CAT# 6570-0-136, LOT# 30593201. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-01013 | 2012-07-27 | 2012-07-09 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 10 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-100000S |  | IT WAS REPORTED THAT, PT STARTED EXPERIENCING HIP PAIN IN (B)(6) 2012. | THE FOLLOWING OTHER HIP DEVICE WAS ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 16DEG 42MM, CAT# NLS-421600P, LOT# 26409201. IT CANNOT BE DETERMINED WHICH, IF ANY OF THE REPORTED DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
MW5026332 | 2012-07-29 | 2012-07-24 | I, n, j, u, r, y | Voluntary report | STRYKER REJUVENATE | FEMORAL IMPLANT-BIPOLAR HEAD | STRYKER CORP |  |  |  | I HAD A STRYKER REJUVENATE HIP IMPLANT PLACE IN ME ON (B)(6) 2011. AFTER (B)(6) MONTHS, MY PAIN GOT WORSE AND WORSE. I WENT BACK TO THE TREATING DOCTOR FOUR TIMES OVER THE LAST (B)(6) MONTHS. HE REPEATEDLY SAID NOTHING WAS WRONG. I FINALLY WAS REFERRED TO SOME (B)(6) DOCTORS THAT ADVISED ME (B)(6) WEEKS AGO THAT THIS STRYKER IMPLANT HAS TO COME OUT. THEY ARE NOT SURE JUST HOW THEY ARE GOING TO GET THIS THING OUT, AS IT WASN'T INTENDED TO COME OUT. AS OF THIS WRITING I AM (B)(6) AND BACK WALKING WITH A CANE DUE TO THIS THING. ON CONSTANT PAIN MEDICATION FOR THE LAST (B)(6) MONTHS. THEY PLAN ON TAKING THIS IMPLANT OUT THE LAST WEEK OF (B)(6). BLOOD WORK AND X-RAYS HAVE BEEN DONE AND A BONE SCAN TOMORROW. THANKS, (B)(6). |  | N | Y | LAY USER/PATIENT |  | 2012-07
2249697-2012-01029 | 2012-07-30 | 2010-12-06 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR SIZE 8 STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, "PATIENT IS EXPERIENCED INCREASING PAIN. HE STATED THAT HIS PAIN IS AROUND THE HIP AND THAT IT STARTED OUT AS A MUSCULAR PAIN AND NOW THE PAIN IS CONSTANTLY THERE. PATIENT STATED THAT HE CAN BARELY WALK NOW AND CAN'T BEND OVER. HIS LEGS ARE EVEN. PATIENT STATED THAT HE HAD AN AUTO IMMUNE ISSUE DEVELOP AFTER SURGERY. HE HAS HAD FOLLOW UPS AND X-RAYS AND REPORTED THAT HIS X-RAYS ARE NORMAL. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PATIENT/EVENT AS MFR# 9616680-2012-00588. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-01052 | 2012-07-31 | 2012-07-10 | I, n, j, u, r, y | Manufacturer report | UNKNOWN #8 REJUVENATE STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: "PT WAS COMPLAINING OF A PAINFUL HIP AND WAS TAKEN TO THE OPERATING ROOM FOR REVISION. COMPONENTS WERE EXPLANTED WITHOUT INCIDENT AND REVISION COMPONENTS WERE REIMPLANTED." | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: UNK 36MM +0 BIOLOX DELTA HEAD. UNK 30MM TRUNNION NEUTRAL VERSION. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS RETAINED BY THE HOSP AND WAS NOT RETURNED TO THE MFR. ADDITIONAL INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD DEVICE OR ADDITIONAL INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-01051 | 2012-07-31 | 2012-06-27 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT, METALLOSIS AT THE JUNCTURE OF THE MODULAR NECK CONNECTED TO THE HIP STEM. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS RETAINED BY THE PT AND WAS NOT RETURNED TO THE MFR. ADDITIONAL INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS: MFR # 9616680-2012-00600. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-01030 | 2012-07-31 | 2012-07-09 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT: DR. (B)(6) STATED THAT THE PT. HAD ONSET OF PAIN STARTING (B)(6) 2010. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0DEG 30MM, CAT# NLS-300000B, LOT# 32662102. TRIDENT PSL HA CLUSTER 56 MM, CAT# 542-11-56F, LOT# MJE7H8. TRIDENT 0 DEG INSERT 44 MM, CAT# 623-00-44F, LOT# MJD8T9. V40 COCR LFIT HEAD 44 MM/+0, CAT# 6260-9-144, LOT# MRH7A1. A 6.5 CANCELLOUS BONE SCREW 25 MM, CAT# 2030-6525-1, LOT# MJA2NX. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-01034 | 2012-07-31 | 2012-07-09 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE RIGHT HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PT WAS TOLD BY DR. (B)(6) TO CALL STRYKER BECAUSE THEY HAVE A REIMBURSEMENT PLAN SET UP. DR. (B)(6)'S EVAL DETERMINED PT NEEDS MRI AND BLOOD TESTS. PT STATED SHE HAS AN "ACHE" BUT NOT PAIN. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-01033 | 2012-07-31 | 2012-07-09 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT: DR. (B)(6) STATED THAT THE PT. HAD INITIAL ONSET OF PAIN STARTING (B)(6) 2012. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0DEG 42MM, CAT# NLS-420000B, LOT# 25613301. TRITANIUM REVISION ACETABULAR, CAT# 509-02-54E, LOT# 1JK0N0. TRIDENT X3 ECCENTRIC 0 36MM ID, CAT# 663-0036E, LOT# 15LJE2. DELTA V-40 CERAMIC HEAD 36/0, CAT# 6570-0-136, LOT# 35714002. GAP PLATE SCREWS, CAT# 2080-0050, LOT# LRDMRD. GAP PLATE SCREWS, CAT# 2080-0025, LOT# MJH01Y. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-01031 | 2012-07-31 | 2018-11-05 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | HIP JOINT METAL/CERAMIC/POLYMER SEMI-CONSTRAINED CEMENTED OR NONPOROUS UNCEMENTE | STRYKER ORTHOPAEDICS-MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT: DR. (B)(6) STATED THAT THE PT. HAD ONSET OF PAIN STARTING (B)(6) 2011. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 8DEG 42MM, CAT# NLV-420800Y, LOT# 31618701. TRIDENT PSL HA CLUSTER 58MM, CAT# 542-11-58G, LOT# MJMTYN. TRIDENT 0 DEG INSERT 44MM, CAT# 623-00-44G, LOT# MJN94K. DELTA UN.FEM HD 44MM, CAT# 6519-1-044, LOT# 4518502. UNI ADAPTOR SLEEVE V40 TI, CAT# 6519-T-204, LOT# 34661405. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-01025 | 2012-07-31 | 2012-07-09 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT: DR. (B)(6) STATED THE PT HAD ONSET OF PAIN STARTING (B)(6) 2011. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0DEG 30MM, CAT# NLS-300000B, LOT# 33133501. TRIDENT PSL HA CLUSTER 54MM, CAT# 542-11-54F, LOT# MJP9H8. TRIDENT 0 DEG INSERT 40MM, CAT# 623-00-40F, LOT# MJEE2P. DELTA UN.FEM HD 40MM, CAT# 6519-1-040, LOT# 31070602. UNI ADAPTER SLEEVE V40 TI, CAT# 6519-T-100, LOT# UNK. A 6.5 CANCELLOUS BONE SCREW 30MM, CAT# 2030-6530-1, LOT# MJN99E. A 6.5 CANCELLOUS BONE SCREW 35MM, CAT# 2030-6535-1, LOT# MJM28H. A 6.5 CANCELLOUS BONE SCREW 40MM, CAT# 2030-6540-1, LOT# MJD5PP. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A F/U REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-01027 | 2012-07-31 | 2012-07-12 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT, PATIENT ADMITTED FOR REVISION FOR PAINFUL RIGHT HIP. | CATALOG NUMBER AND LOT CODE OF THE OTHER DEVICE LISTED IN THIS REPORT: CAT # NLS-300000B, LOT # 37569001, DESCRIPTION: 30 MM MODULAR NECK. IT CANNOT BE DETERMINED WHICH DEVICE, IF ANY, MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S EXPERIENCE. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-07
2249697-2012-01035 | 2012-07-31 | 2012-07-09 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT: DR. (B)(6) STATED THE PT. HAD INITIAL ONSET OF PAIN STARTING (B)(6) 2011." | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 8 DEG 30MM, CAT# NLV-300800Y, LOT# 26850801. TRITANIUM REVISION ACETABULAR, CAT# 509-02-58F, LOT# MKHJAE. TRIDENT 0 DEG INSERT 40MM, CAT# 623-00-40F, LOT# MKEMTW. V40 COCR LFIT HEAD 40MM/+0, CAT# 6260-9-140, LOT# MJAVJE. GAP PLATE SCREWS, CAT# 2080-0050, LOT 4HWMME. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-00979 | 2012-07-31 | 2012-07-05 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP SYSTEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, THE PT IS EXPERIENCING PAIN AND HAS NOT BEEN PAIN FREE SINCE THE SURGERY. STEROIDS MADE THE PAIN A LITTLE BETTER, BUT WITHOUT STEROID HE IS IN SIGNIFICANT CONSTANT PAIN. STRENUOUS ACTIVITY MAKES THE PAIN MUCH WORSE. HE IS ALLERGIC TO OTC MEDS SO HE DOES NOT TAKE MANY PAIN MEDICATIONS. A BLOOD TEST WAS PERFORMED AND THE RESULTS SHOWED CR WAS NORMAL, BUT THE OTHER METAL TESTED HIGH. HE IS BEING SENT FOR AN MRI. THE DOCTOR HAS STATED THAT HE BELIEVES THIS PT WILL NEED A REVISION. HE IS PAYING FOR HIMSELF AND CANNOT AFFORD TO PAY FOR TESTS AND SURGERY. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-00970 | 2012-07-31 | 2012-07-03 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRIGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS - MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT, SURGEON SAID THAT THE PT EXPERIENCED AN INEXPLICABLE PAIN AND SWELLING. SURGEON SAID COBALT LEVELS WERE ELEVATED. THEREFORE, SURGEON EXPLANTED THE LISTED ITEMS AND REPLANTED THE FOLLOWING ITEMS: A RESTORATION STEM, ADM POLYETHYLENE AND 28 MM COBALT CHROME -3 HEAD. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED FOR THE MFR. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS: MFR# 9616680-2012-00555. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-00966 | 2012-07-31 | 2012-07-03 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT, PT WAS COMPLAINING OF PAIN AND HAD ELEVATED COBALT-CHROMIUM LEVELS. REJUVENATE STEM TRUNNION AND HEAD WERE EXPLANTED. LINER AND CUP WERE THEN REMOVED UNEVENTFULLY. REIMPLANTATION WAS THEN INITIATED AND COMPLETED IN A TIMELY MANNER WITH REVISION COMPONENTS. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS RETAINED BY THE HOSPITAL AND WAS NOT RETURNED TO THE MFR. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS: MFR# 9616680-2012-00553. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-01022 | 2012-07-31 | 2012-07-09 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED: "DR. (B)(6) STATED THE PT. HAD ONSET OF PAIN STARTING (B)(6) 2011." | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0DEG 42MM, CAT# NLS-420000B, LOT# 25705902. TRIDENT PSL HA CLUSTER 56MM, CAT# 542-11-56F, LOT# MJL03M. TRIDENT 0 DEG INSERT 40MM, CAT# 623-00-40F, LOT# MJKY2P. UNI ADAPTOR SLEEVE V40 TI, CAT# 6510-T-204, LOT# 3331006. A 6.5 CANCELLOUS BONE SCREW 50MM, CAT# 2030-6550-1, LOT# MPH995. A 6.5 CANCELLOUS BONE SCREW 16MM, CAT# 2030-6516-1, LOT# MHNAKW. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESES DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-01028 | 2012-07-31 | 2012-07-09 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 11 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-110000S |  | IT WAS REPORTED THAT: "DR. (B)(6) STATED, THE PT HAD ONSET OF PAIN STARTING (B)(6) 2011." | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 8DEG 38MM, CAT# NLV-380800Y, LOT# 29589401. RESTORATION (TM) ADM. CUP W/HA, CAT# 1235-2-621, LOT# G2848547. RESTORATION ADM X3 INS 28/62, CAT# 1236-2-862, LOT# 33079901. DELTA V-40 CERAMIC HEAD 28/+4, CAT# 6570-0-228, LOT# 35066702. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-01026 | 2012-07-31 | 2012-07-09 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED: DR. (B)(6) STATED THAT THE PT. HAD ONSET OF PAIN STARTING (B)(6) 2012. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: REJUVENATE MODULAR NECK, CAT# NLV-300800G, LOT# 32662201. TRIDENT PSL HA CLUSTER 50 MM, CAT# 542-11-50E, LOT# MJH98Y. TRIDENT 0 X3 INSERT 36MM ID, CAT# 623-00-36E, LOT# MJK6DJ. DELTA V-40 CERAMIC HEAD 36/0, CAT# 6570-0-136, LOT# 33430202. A 6.5 CANCELLOUS BONE SCREW 60MM, CAT# 2030-6560-1, LOT# MJH4HY. A 6.5 CANCELLOUS BONE SCREW 35MM, CAT# 2030-6535-1, LOT# MJM929. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-01048 | 2012-07-31 | 2011-03-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT HIP REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, FIRST FEW MONTHS AFTER SURGERY, HE WAS FINE. THEN HE WAS IN SO MUCH PAIN HE WOULD CRY. METAL TESTS ON (B)(6) 2012 SHOWED METAL LEAKING INTO BLOOD STREAM. COBALT LEVEL WAS HIGH. REVISION IS SCHEDULED FOR (B)(6) 2012. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-01047 | 2012-07-31 | 2012-07-09 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT: "FLUID AND SWELLING CAUSED PAIN. MOST PAINFUL WHEN SITTING. PULLED THE RECALLED REJUVENATE STEM AND NOTICED FRETTING ON THE NECK - STEM JUNCTION. TISSUES WERE BLACK AND SHOWED SIGNS OF METALLOSIS. STEM WAS WELL FIXED, FLUID WAS DRAINED AND THE STEM WAS PULLED AND REPLACED WITH RESTORATION MODULAR." | WHEN COMPLETED, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS: MFR # 9616680-2012-00598. | Y | N | HEALTH PROFESSIONAL |  | 2012-07
2249697-2012-01064 | 2012-08-01 | 2012-07-11 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT: DR. (B)(6) STATED THE PT HAD ONSET OF PAIN STARTING (B)(6) 2012. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0 DEG 30MM, CAT # NLS-300000B, LOT # 34095901. UNI ADAPTOR SLEEVE V40 TI, CAT # 6519-T-100, LOT# 34630802. TRIDENT 0 DEG INSERT 40MM, CAT# 623-00-40F, LOT # MKEJKE. TRIDENT PSL HA CLUSTER 54MM, CAT # 542-11-54F, LOT # MKD9H3. 5.5 CANCELLOUS BONE SCREW 50MM, CAT # 2029-5550-1, LOT # UNK. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01039 | 2012-08-01 | 2012-07-17 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | UNKNOWN | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT, STAGE ONE REVISION DUE TO INFECTION. IMPLANTS WERE REMOVED AND REPLACED WITH AN ANTIBIOTIC SPACER. REJUVENATE STEM WAS WELL FIXED AND HAD TO BE REMOVED VIA AN OSTEOTOMY, PER SURGEON. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. CATALOG NUMBERS AND LOT CODES FOR OTHER DEVICES LISTED IN THIS REPORT: CAT # 2030-6535-1, LOT # MJTY70 DESCRIPTION: 6.5 CANCELLOUS BONE SCREW 35MM. CAT # 2030-6530-1, LOT # MJRJA6 DESCRIPTION: 6.5 CANCELLOUS BONE SCREW 30MM. CAT # 623-00-36E, LOT #MKDEWL DESCRIPTION: TRIDENT 0 X3 INSERT 36MM ID. CAT # NLS-380000B, LOT # 32143301 DESCRIPTION: 38 REJUVENATE NECK. CAT # 6570-0-136, LOT # 35238301 DESCRIPTION: DELTA V-40 CERAMIC HEAD 36/0. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S INFECTION. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01038 | 2012-08-01 | 2012-07-18 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | PATIENT'S SON REPORTED THAT PATIENT HAD HIP PAIN NEAR INCISION. HE STATED XRAYS ARE "OK" BUT PATIENT WILL BE HAVING BONE SCAN. NO CONTACT INFORMATION OR ADDITIONAL INFORMATION WAS PROVIDED BY THE SON. HE STATED THAT THEY WILL CALL BACK IF THEY NEED TO. HE HAD QUESTIONS ABOUT THE RECALL AND QUESTIONS ABOUT HIS FATHER'S SYMPTOMS. HE WAS ADVISED TO SEE HIS SURGEON WITH ANY MEDICAL CONCERNS. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER. IF ADDITIONAL INFORMATION BECOMES AVAILABLE, IT WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01046 | 2012-08-01 | 2009-12-14 | I, n, j, u, r, y | Manufacturer report | UNKNOWN HIP REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PATIENT HAD HIS PRIMARY SURGERY ON (B)(6) 2007. HE BEGAN EXPERIENCING PAIN AND HAD A REVISION ON (B)(6) 2009. PATIENT IS CURRENTLY IN PAIN AND HAS BEEN INFORMED BY HIS SURGEON THAT DUE TO HIS BLOOD WORK AND TESTS HE SEES THE NEED FOR A REVISION DUE TO THE RECALL IMPLANTS. PATIENT STATED THAT HE HAS EXTENSIVE DAMAGE TO THE BONE AND MUSCLE TISSUE. HE IS UNSURE OF HIS IMPLANT BUT STATES THAT HIS SURGEON SAID IT WAS ONE OF THE RECALL. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01054 | 2012-08-01 | 2010-09-21 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE LEFT HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PATIENT IS EXPERIENCING PAIN SINCE SURGERY IN BOTH HIPS. SHE REPORTS TO WADDLE WHEN SHE WALKS WITH A WALKER. SHE CAN'T WALK BY HERSELF. HER HIPS FEEL LIKE THEY ARE WOBBLING. X-RAYS TAKEN SHOWED NOTHING WRONG. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01062 | 2012-08-01 | 2011-05-02 | I, n, j, u, r, y | Manufacturer report | UNKNOWN PRODUCT REJUVENATE LEFT HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, THIS PATIENT STARTED HAVING PAIN LITTLE BY LITTLE BUT WAS GETTING PROGRESSIVELY WORSE. APPROXIMATELY 2 MONTHS AGO, THE PAIN STARTED TO GET MORE SEVERE AND THE PATIENT STARTED TAKING OTC PAIN MEDICINE. THE PATIENT WAS WALKING WITH A LIMP AND MADE AN APPOINTMENT WITH THE SURGEON. AN MRI WAS DONE AND INFLAMMATION WAS DETECTED. THE SURGEON WANTS THE PATIENT TO REVISE THESE IMPLANTS AS SOON AS POSSIBLE TO AVOID ANY FURTHER TISSUE DAMAGE. THE REVISION HAS NOT BE SCHEDULED TO DATE. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01061 | 2012-08-01 | 2010-09-27 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: "THE PT WAS HAVING PAIN SINCE THE SURGERY. HE WAS PRESCRIBED 24/7 CARE IN HIS HOME FOR APPROXIMATELY 13 WEEKS AS WELL AS REHABILITATION THERAPY IN AN OFFICE. THEN THE PT WENT FOR 30 WEEKS OF THERAPY. THE HIP DOCTOR SENT HIM TO A SPINAL DOCTOR THINKING THAT THE ISSUES WERE STEMMING FROM THE BACK. THE SPINAL DOCTOR SAID IT WAS NOT THE BACK THE PAIN WAS COMING FROM THE HIP AND HE WAS SENT FOR ANOTHER 18 WEEKS THERAPY. APPROXIMATELY 1 YEAR AGO, THE PT WAS HEARING NOISES AND COULD FEEL GRINDING IN HIS HIP. ANOTHER VISIT WAS MADE TO DR. (B)(6) BECAUSE OF PAIN. HE IS WAITING FOR BLOOD WORK AND A MARS MRI. THE PT HAD 2 VISITS TO THE ER SINCE HIS SURGERY, ONE FOR THE NOISE AND THE OTHER FOR INSTABILITY. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01071 | 2012-08-01 | 2018-02-19 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | HIP JOINT METAL/CERAMIC/POLYMER SEMI-CONSTRAINED CEMENTED OR NONPOROUS UNCEMENTE | STRYKER ORTHOPAEDICS-MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT: DR. (B)(6) STATED THAT "THE PT HAD ONSET OF PAIN STARTING (B)(6) 2011." | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 8 DEG 38MM, CAT# NLV-380800Y, LOT# 27502301. TRIDENT PSL HA CLUSTER 56MM, CAT# 542-11-56F, LOT# MJK6R6. TRIDENT 0 DEG INSERT 40MM, CAT# 623-00-40F, LOT# MJPW11. DELTA UN.FEM HD 40MM, CAT# 6519-1-040, LOT# 34337401. UNI ADAPTOR SLEEVE V40 TI, CAT# 6519-T-100, LOT# 31766204. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01055 | 2012-08-01 | 2011-02-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE RIGHT HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PATIENT IS EXPERIENCING PAIN SINCE SURGERY IN BOTH HIPS. SHE REPORTS TO WADDLE WHEN SHE WALKS WITH A WALKER. SHE CAN'T WALK BY HERSELF. HER HIPS FEEL LIKE THEY ARE WOBBLING. X-RAYS TAKEN SHOWED NOTHING WRONG. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01074 | 2012-08-01 | 2012-07-11 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT: DR. (B)(6) STATED, THE PT STARTED HAVING PAIN ON (B)(6) 2010. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 16 DEG 24MM, CAT# NLS-341600P, LOT# 25613002. TRITANIUM REVISION ACETABULAR, CAT# 509-02-64G, LOT# MHP2H4. TRIDENT 0 DEG INSERT 44MM, CAT# 623-00-44G, LOT# MJJNOX. DELTA UN.FEM HD 44MM, CAT# 6519-1-044, LOT# 33811403. UNI ADAPTOR SLEEVE V40 TI, CAT# 6519-T-100, LOT# 34435602. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01072 | 2012-08-01 | 2012-07-11 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT: DR. (B)(6) STATED THE PT STARTED HAVING PAIN ON (B)(6) 2012. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 8 DEG 38MM, CAT# NLV-380800Y, LOT# 27502801. TRIDENT 0 DEG INSERT 44MM, CAT# 623-00-44G, LOT# MJMTPA. UNI ADAPTOR SLEEVE V40 TI, CAT# 6519-T-204, LOT# 31988807. DELTA UN.FEM HD 44MM, CAT# 6519-1-044, 3 LOT# 1056502. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01068 | 2012-08-01 | 2012-07-11 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT: DR. (B)(6) STATED, THE PT STARTED HAVING PAIN (B)(6) 2011. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0 DEG 38MM, CAT# NLS-380000B, LOT# 33810001. TRITANIUM REVISION ACETABULAR, CAT# 509-02-58F, LOT# MJAX7J. TRIDENT 0 DEG INSERT 40MM, CAT# 623-00-40F, LOT# MJP385. V40 COCR LFIT HEAD 40MM/+0, CAT# 6260-9-140, LOT# 3V8MLE. GAP PLATE SCREWS, CAT# 2080-0040, LOT# MJE372. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01065 | 2012-08-01 | 2012-07-11 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT: DR. (B)(6) STATED THE PT HAD INITIAL ONSET OF PAIN STARTING (B)(6) 2012. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0 DEG 42MM, CAT# NLS-420000B, LOT# 31842202. TRIDENT PSL WITH PUREFIX HA 60MM, CAT# 542-11-60G, LOT# 32179801. TRIDENT 0 DEG INSERT 44MM, 6 CAT# 623-00-44G, LOT# MJHTDN. DELTA UN.FEM HD 44MM, CAT# 6519-1-044, LOT# 35690901. UNI ADAPTOR SLEEVE V40 TI, CAT# 6519-T-100, LOT# 35643505. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01069 | 2012-08-01 | 2012-07-11 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT: DR. (B)(6) STATED THE PT STARTED HAVING PAIN ON (B)(6) 2011. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0 DEG 42MM, CAT# NLS-420000B, LOT# 32277401. TRIDENT PSL WITH PUREFIX HA 60MM, CAT# 542-11-60G, LOT# 35525202. V40 COCR LFIT HEAD 44MM/+0, CAT# 6260-9-144, LOT# A94MEE. TRIDENT 0 DEG INSERT 44MM, CAT# 623-00-44G, LOT# MJAEHV. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01067 | 2012-08-01 | 2012-07-11 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT: DR. (B)(6) STATED THAT THE PT STARTED HAVING PAIN (B)(6) 2010. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0 DEG 42MM, CAT# NLS-420000B, LOT# 31842301. TRIDENT PSL WITH PUREFIX HA 60MM, CAT# 542-11-60G, LOT# 33152202. TRIDENT 0 DEG INSERT 44MM, CAT# 623-00-44G, LOT# MJDE28. UNI ADAPTOR SLEEVE V40 TI, CAT# 6519-T-100, LOT# 30443406. DELTA UN.FEM HD 44MM, CAT# 6519-1-044, LOT# 30042201. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01073 | 2012-08-01 | 2012-07-11 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT: DR. (B)(6) STATED THE PT HAD PAIN STARTING (B)(6) 2012. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 8 DEG 38MM, CAT# NLV-380800G, LOT# 31714901. TRIDENT PSL WITH PUREFIX HA 64 MM, CAT# 542-11-64H, LOT# 32026701. TRIDENT 0 DEG INSERT 44 MM, CAT# 623-00-44H, LOT# MHPETP. UNI ADAPTOR SLEEVE V40 TI, CAT# 6519-7-100, LOT# 34630806. DELTA UN.FEM HD 44 MM, CAT# 6519-1-044, LOT# 34967401. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01042 | 2012-08-01 | 2010-07-19 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT HIP REJUVENATE SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PATIENT DID NOT EXPERIENCE ANY PROBLEMS AT FIRST. HOWEVER, PATIENT BEGAN EXPERIENCING PAIN IN (B)(6) OF 2011. IT BOTHERED HIM SO MUCH THAT HE HAD XRAYS TAKEN BUT THEY WERE NORMAL. ALSO, HE CONTINUES WITH PHYSICAL THERAPY BUT THE PAIN IS GETTING WORSE AND THE THERAPY IS NOT HELPING. HE IS CURRENTLY ON PERCOCET TO CONTROL HIS PAIN. HE HAD A FOLLOW UP VISIT WITH HIS DOCTOR YESTERDAY AND IT WAS DETERMINED THAT HE WILL NEED A BONE SCAN AND FURTHER BLOOD TESTS. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01086 | 2012-08-02 | 2011-12-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN SIZE 12 REJUVENATE STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT SOON AFTER THE IMPLANTS WERE PUT IN, THIS PATIENT STARTED EXPERIENCING PROBLEMS. THE PATIENT IS FEELING PAIN IN THE JOINT THAT IS GETTING PROGRESSIVELY WORSE. HE ALSO FEELS THAT THERE IS A REDUCTION IN ROM. AS HIS SYMPTOMS PROGRESSED, HE WAS TAKEN TO (B)(6) APPROXIMATELY (B)(6) MONTHS AGO, WHICH RESULTED IN COCR = 10. NEXT WEEK (B)(6) FOR MARS MRI. THE SURGEON IS STATING THAT A REVISION IS LIKELY, WHICH HE WILL ASSESS AFTER TESTING IS COMPLETE. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PATIENT/EVENT AS: MFR # 9616680-2012-00615. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
9616680-2012-00615 | 2012-08-02 | 2011-12-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN 34 MM REJUVENATE NECK 132 DEG ANGLE NEUTRAL VERSIO | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT SOON AFTER THE IMPLANTS WERE PUT IN THIS PATIENT STARTED EXPERIENCING PROBLEMS. THE PATIENT IS FEELING PAIN IN THE JOINT THAT IS GETTING PROGRESSIVELY WORSE. HE ALSO FEELS THAT THERE IS A REDUCTION IN ROM. AS HIS SYMPTOMS PROGRESSED, HE WAS TAKEN TO (B)(6) HOSPITAL APPROXIMATELY 1 1/2 MONTHS AGO WHICH RESULTED IN COCR = 10. NEXT WEEK (B)(6) HOSPITAL FOR MARS MRI. THE SURGEON IS STATING THAT A REVISION IS LIKELY WHICH HE WILL ASSESS AFTER TESTING IS COMPLETE. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PATIENT/EVENT AS: MFR # 2249697-2012-01086. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01090 | 2012-08-02 | 2012-07-12 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT: "DR. (B)(6) REVISED PATIENTS HIP DUE TO ADVERSE LOCAL TISSUE REACTION." | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER DUE TO HOSPITAL POLICY. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PATIENT/EVENT AS: MFR # 9616680-2012-00617. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01092 | 2012-08-02 | 2012-07-12 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 10 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-100000S |  | IT WAS REPORTED THAT: "DOCTOR (B)(6) REVISED PATIENT'S HIP DUE TO SUSPECTED ADVERSE LOCAL TISSUE REACTION." | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REQUESTED FROM HOSPITAL BY PATIENTS ATTORNEY AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PATIENT'/EVENT AS: MFR# 9616680-2012-00618. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01094 | 2012-08-02 | 2012-07-12 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRAIGHT PRFIT TMZF MOD STEM SIZE 10 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-100000S |  | IT WAS REPORTED THAT: "DR (B)(6) REVISED PT'S HIP DUE TO ADVERSE LOCAL TISSUE REACTION." | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR DUE TO HOSPITAL POLICY. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS: MFR NUMBER 9616680-2012-00619. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01079 | 2012-08-02 | 2011-10-24 | I, n, j, u, r, y | Manufacturer report | REJUVENATE RIGHT HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT WHEN PATIENT WENT INTO THE REVISION SURGERY THE IMPLANT COULDN'T BE REPLACED WITHOUT BREAKING THE BONE. THE PATIENT WAS SITTING IN AN OFFICE ON A COMFORTABLE CHAIR AND WHILE ADJUSTING HERSELF IN HER SEAT SHE HEARD A POP AND THE HIP DISLOCATED. SHE THEN HAD NO CONTROL OVER THAT LEG. ADDITIONAL FRACTURES HAVE NOW BEEN DETECTED IN THAT AREA AND A SURGERY IS PLANNED (B)(6) 2012 TO TRY FIX THOSE FRACTURES BY DR (B)(6). | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01088 | 2012-08-02 | 2012-07-12 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT: "I MET WITH MY DR. - (B)(6) OF (B)(6) ORTHOPEDICS, (B)(6) TODAY, WHERE WE REVIEWED MY X-RAYS, BLOOD WORK AND MRI. I HAD MY REPLACEMENT ON (B)(6) 2010 AND ESSENTIALLY HAVE BEEN IN PAIN AND LIMPING EVER SINCE. I WAS (B)(6) WHEN REPLACED. DR. (B)(6) SAID, MY COBALT LEVELS WERE THE WORST HE HAS SEEN THUS FAR, AND THE FLUID BUILD UP AND BONE LOSS IS VERY CONCERNING TO ME. I AM SCHEDULED FOR SURGERY ON (B)(6) TO REPLACE THE STEM. UP UNTIL I CAME BACK IN ON MY OWN ON (B)(6), I HAD NO KNOWLEDGE OF THE PROBLEMS WITH THE STEM. AT MY (B)(6) 2011 CHECK UP I WAS TOLD MY PAIN AND SWELLING WAS COMMON. NO RECOMMENDATIONS FOR TREATMENT WERE MADE AT THAT TIME AND I SPENT THE PAST YEAR TREATING AT OTHER CLINICS, ETC. TO NO AVAIL." | ON (B)(4) 2012: SALES REP REPORTED: "DOCTOR REVISED PATIENT'S HIP DUE TO ADVERSE LOCAL TISSUE REACTION." AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER DUE TO HOSPITAL POLICY. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PATIENT/EVENT AS: MFR # 9616680-2012-00616. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01077 | 2012-08-03 | 2012-05-18 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT: "THE SURGEON STATED THAT THE PT HAD INITIAL ONSET OF PAIN STARTING (B)(6) 2012." | CATALOG NUMBERS AND LOT CODES OF OTHER DEVICES LISTED IN THIS REPORT: DESCRIPTION: REJUVENATE MODULAR NECK: CAT # NLS-300000B, LOT # 34887301. DESCRIPTION: PSL HA CLUSTER 54MM: CAT # 542-11-54F, LOT # MJPREN. DESCRIPTION: TRIDENT O DEG INSERT 40MM: CAT # 623-00-40F, LOT # MKAXX4. DESCRIPTION: DELTA UN. FEM HD 40MM. DESCRIPTION: UNI ADAPTOR SLEEVE V40 TI: CAT # 6519-5-100, LOT #37094101. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S EXPERIENCE. ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01076 | 2012-08-03 | 2012-07-16 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT: "THE SURGEON STATED THE PT STARTED HAVING PAIN ON (B)(6) 2011." | CATALOG NUMBERS AND LOT CODES OF OTHER DEVICES LISTED IN THIS REPORT. DESCRIPTION: REJUVENATE MODULAR NECK: CAT # NLS-300000B, LOT # 35793502. DESCRIPTION: TRIDENT PSL WITH PUREFIX HA 60MM: CAT # 542-11-60G, LOT # 34081002. DESCRIPTION TRIDENT 0 DEG INSERT 44MM: CAT # 623-00-44G, LOT # 6TWMHD. DESCRIPTION: DELTA UN. FEM HD 44MM: CAT # 6519-1-044, LOT # 35418301. DESCRIPTION: UNI ADAPTOR SLEEVE V40 TI: CAT # 6519-T-100, LOT # 35050001. IT CANNOT BE DETERMINE WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S EXPERIENCE. ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01103 | 2012-08-03 | 2012-07-18 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, THE PT COMPLAINED OF PAIN SO THE SURGEON REVISED. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: UNK REJUVENATE NECK. UNK HEAD. UNK LINER. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. WHEN COMPLETED, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01095 | 2012-08-03 | 2011-06-20 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT: ¿(B)(6) 2012 THE PT STARTED EXPERIENCING EXTREME PAIN AND ACHING AT NIGHT WHICH WAS PREVENTING HER FROM BEING ABLE TO SLEEP. IT IS VERY TENDER/SENSITIVE TO THE TOUCH. IT IS NOTICEABLE WHEN SHE HAS TO CARRY ANYTHING HEAVY. SHE HAS BLOOD WORK COCR OF 5.9. MRI WAS TAKEN ON (B)(6) 2012. A LESION AND PSEUDO TUMOR WAS NOTED IN HER HIP IN WHICH DR. (B)(6) REPORTED ALTR.¿ | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD¿L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD¿L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS: MFR #: 9616680-2012-00620. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01087 | 2012-08-03 | 2020-02-25 | I, n, j, u, r, y | Manufacturer report | ABGII. MODULAR STEM | PROSTHESIS, HIP, SEMI-CONSTRAINED, UNCEMENTED, METAL/POLYMER, NON-POROUS, CALICU | STRYKER ORTHOPAEDICS-MAHWAH | NA | 4845-4-205 |  | PT REPORTS THAT IN THE MORNINGS, HE'S ALRIGHT BUT AS THE DAY PROGRESSES, THE PAIN GETS WORSE. HE'S EXPERIENCING A SHOOTING PAIN AND IT'S DIFFICULT FOR HIM TO WALK. HE IS UNABLE TO PUT PRESSURE ON HIS LEFT HIP. PT SAW DOCTOR AND HAD X-RAYS TAKEN AS WELL AS BLOOD WORK. HE IS SCHEDULED TO HAVE A BONE SCAN. | AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER, BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | 0 |  | 2012-08
2249697-2012-01127 | 2012-08-07 | 2012-04-18 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT IS IN PAIN AND IN A FOLLOW UP VISIT THAT SHE HAD WITH HER DOCTOR, HER SURGEON ADVISED HER TO CALL STRYKER SINCE SHE HAS THE REJUVENATE IMPLANT. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01126 | 2012-08-07 | 2009-11-23 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE | IMPLANT | STRYKE ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT WAS IN PAIN. SHE HAD HER INITIAL SURGERY ON (B)(6) 2009 AND WAS REVISED ON (B)(6) 2012. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01141 | 2012-08-07 | 2011-09-23 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PATIENT HAS BEEN HAVING HIP PROBLEMS. HE IS IN PAIN AND SINCE HIS SURGERY HE HAS HAD FIVE ASPIRATIONS OF JOINTS/BURSA PERFORMED. THE MOST RECENT ASPIRATIONS DATED ON (B)(6) 2012, (B)(6) 2012 AND (B)(6) 2012. AS PER WIFE, THE PATIENT IS EXPERIENCING REDNESS AT INCISION SITE. PATIENT HAS ALSO HAD PHLEBITIS AT THE RIGHT BUTTOCK. PATIENT HAD BEEN SCHEDULED FOR A REVISION SURGERY YESTERDAY BUT IT WAS CANCELLED DUE TO OTHER MEDICAL PROBLEMS. PATIENT HAS ALSO HAD A BIOPSY DONE ON (B)(6) 2012. HE WILL HAVE HIS REVISION RESCHEDULED. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01109 | 2012-08-07 | 2012-07-12 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT, PT RETURNED TO OFFICE ON (B)(6) /2012 COMPLAINING OF THIGH PAIN. X-RAY WAS TAKEN AND REVEALED A BROKEN PLATE AT THE PREVIOUS FX SITE DUE TO A NON UNION OF THE PERI-PROSTHETIC FX. REJUVENATE STEM AND NECK AND V40 HEAD WERE REMOVED. SMITH NEPHEW REVISION STEM WAS IMPLANTED. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 8 DEG 34MM, CAT# NLV-340800G, LOT# 35231101. V40 COCR LFIT HEAD 36MM/+5, CAT# 6260-9-236, LOT# MKEN4L. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S NON UNION OF THE PERI PROSTHETIC FRACTURE. AN EVALUATION OF THE REPORTED DEVICES CANNOT BE PERFORMED AS THE DEVICES WERE SENT TO THE HOSPITAL'S PATHOLOGY LAB PER THEIR PROTOCOL AND WILL BE DESTROYED. ADDITIONAL INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
9616680-2012-00627 | 2012-08-07 | 2011-01-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE 34MM NECK 127 DEGREE X 8 ANTEVE | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT, LEFT TOTAL HIP. EXTREME GROIN PAIN. (B)(6) 2011 SHE ATTEMPTED TO TAKE HER WALKER OUT OF THE CAR AND SHE BELIEVES THAT IS WHEN THE PAIN STARTED. THROUGH (B)(6) 2011 SHE CONTINUED TO SEE HER ORTHOPEDIC SURGEON. THE DOCTOR WAS TRYING TO TREAT HER WITH CORTISONE SHOTS IN THE ILIOPSOAS. SHE WAS THEN REFERRED TO DR. (B)(6) ON (B)(6) 2011. (B)(6) 2011 OR (B)(6) 2012 AN MRI WAS DONE. A CORTISONE SHOT WAS DONE AND SHE WAS PAIN FREE FOR SOME TIME AND THEN THE PAIN STARTED AGAIN WHILE PUTTING HER WALKER INTO HER CAR. SITTING AND WALKING IS THE MOST EXCRUCIATING PAIN. SHE IS CONCERNED ABOUT ALL OF THE PAIN MEDICATION THAT SHE HAS TO TAKE TO RELIEVE SOME OF THE PAIN AND COSTS ASSOCIATED WITH HER HEALTH CARE. A BLOOD TEST HAS NOW REVEALED A HIGH LEVEL OF COCR. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PATIENT/EVENT AS: MFR # 2249697-2012-01112. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01113 | 2012-08-07 | 2012-01-13 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PT HAD SURGERY SIX MONTHS AGO. HE IS CURRENTLY COMPLAINING OF PAIN, LIMPING AND SWELLING IN UPPER THIGH. HE CAN'T GO UP THE STAIRS. HE HAS A SCHEDULED APPOINTMENT TO SEE HIS DOCTOR ON (B)(6) 2012. ONE MONTH AGO HE HAD COBALT AND CHROMIUM LEVEL TESTING WHICH WERE NORMAL. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01108 | 2012-08-07 | 2012-01-23 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT, PRIMARY SURGERY WAS (B)(6) 2011 WITH REJUVENATE STEM AND PRIMARY TRITANIUM CUP. ON (B)(6) 2012 PT DEVELOPED A PERI PROSTHETIC FX AND 13H AXSOS 5.0 COMPRESSION WAS IMPLANTED WITH 4 CABLES AND 4 4.5 CORTICAL SCREWS. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 8 DEG 34MM, CAT# NLV-340800G, LOT# 35231101. V40 COCR LFIT HEAD 36MM/+5, CAT# 6260-9-236, LOT# MKEN4L. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PERI PROSTHETIC FRACTURE. AN EVALUATION OF THE REPORTED DEVICES CANNOT BE PERFORMED AS THE DEVICES REMAINED IMPLANTED IN THE PT AND WERE NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01119 | 2012-08-07 | 2012-07-16 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PT HAS SYMPTOMS AND WILL HAVE A REVISION. HE HAD QUESTIONS ABOUT THE RECALL AND REIMBURSEMENT AND JUST WANTED INFORMATION FOR NOW. HE STATED HE WILL TALK TO HIS LAWYER FIRST AND DID NOT WANT TO PROVIDE HIS INFORMATION. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION WAS REQUESTED, BUT NOT MADE AVAILABLE. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01121 | 2012-08-07 | 2012-02-29 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, THE PT IS HAVING PAIN, SWELLING AND CANNOT EXERT TOO MUCH WEIGHT OR EXERCISE. INITIALLY PT WAS DOING FINE BUT AT THE END OF MAY HE STARTED TO FEEL PAIN. HE HAD AN MRI TODAY AND BLOOD TESTS AND WILL FOLLOW UP WITH SURGEON. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01112 | 2012-08-07 | 2011-01-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE #9 STEM LEFT | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, LEFT TOTAL HIP. EXTREME GROIN PAIN. (B)(6), 2011 SHE ATTEMPTED TO TAKE HER WALKER OUT OF THE CAR AND SHE BELIEVES THAT IS WHEN THE PAIN STARTED. THROUGH (B)(6) OF 2011 SHE CONTINUED TO SEE HER ORTHOPEDIC SURGEON. THE DOCTOR WAS TRYING TO TREAT HER WITH CORTISONE SHOTS IN THE ILIOPSOAS. SHE WAS THEN REFERRED TO DR. (B)(6) ON (B)(6) 2011. (B)(6) OF 2011 OR (B)(6) OF 2012 AN MRI WAS DONE. A CORTISONE SHOT WAS DONE AND SHE WAS PAIN FREE FOR SOME TIME AND THEN THE PAIN STARTED AGAIN WHILE PUTTING HER WALKER INTO HER CAR. SITTING AND WALKING IS THE MOST EXCRUCIATING PAIN. SHE IS CONCERNED ABOUT ALL OF THE PAIN MEDICATION THAT SHE HAS TO TAKE TO RELIEVE SOME OF THE PAIN AND COSTS ASSOCIATED WITH HER HEALTH CARE. A BLOOD TEST HAS NOW REVEALED A HIGH LEVEL OF COCR. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PATIENT/EVENT AS: MFR # 9616680-2012-00627. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01118 | 2012-08-07 | 2011-10-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE RIGHT HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, THE PT STARTED HAVING SYMPTOMS IMMEDIATELY FOLLOWING THE SURGERY WHICH WAS NOT THE CASE WHEN SHE HAD HER LEFT HIP REPLACEMENT. SHE HAD PAIN IN THE GROIN AREA. SHE STARTED EXPERIENCING DIZZINESS, RINGING IN HER EAR, VERY FATIGUED WHERE SHE WOULD SLEEP 10-12 HOURS A DAY. TEST REVEALED LOW IRON AND HI COCR OF 21. A TUMOR WAS DETECTED AS WELL AS ALTR. THE AREA WAS YELLOW AND DETERIORATED. WHEN SHE WENT INTO THE REVISION SURGERY THE IMPLANT COULDN'T BE REPLACED WITHOUT BREAKING THE BONE. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01120 | 2012-08-07 | 2011-10-31 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE | IMPLANT | STRYKER ORHTOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PT RECEIVED A LETTER FROM HER SURGEON AND HAD QUESTIONS ABOUT THE INFLAMMATION PROCESS AND HER PAIN. SHE WAS REFERRED TO HER SURGEON TO HAVE THESE QUESTIONS ANSWERED. SHE STATED THAT SHE IS CURRENTLY EXPERIENCING PAIN THAT BEGAN ABOUT 5 WEEKS AGO. ON (B)(6) 2012, SHE SAW HER SURGEON AND HAD AN X-RAY PERFORMED. SURGEON STATED THAT IT MAY JUST BE PAIN FROM TIRED MUSCLES. SHE IS DOING MUCH BETTER AND JUST HAD MAINLY QUESTIONS REGARDING HER SYMPTOMS. SHE SAID, SHE WILL SEE HER DOCTOR IN A FEW WEEKS AGAIN AND MAY CALL BACK. UPDATE: PT SAW SURGEON ON (B)(6) 2012 AND SHE STATED THAT HER BLOOD LEVELS WERE HIGH. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE(S) REFERENCE IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01133 | 2012-08-07 | 2011-11-15 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PT HAD INITIAL SURGERY ON (B)(6) 2010. IT WAS DETERMINED THAT PT HAD WEAR AND DEBRIS. SHE ALREADY HAD FRETTING AND CORROSION AND WAS REVISED ON (B)(6) 2011. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01134 | 2012-08-07 | 2011-11-15 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, SHE IS NOW IN A LOT OF PAIN. SHE STATES THAT SHE WADDLES AND LIMPS. SHE STATED THAT SHE WILL BE HAVING A REVISION IN THE NEAR FUTURE. HER SURGEON WAS COMMUNICATED TO HER THAT THE STEM IS SOLID AND THEREFORE NEEDS TO BE CHIP AWAY AT THE BONE. THEREFORE, BONE WILL HAVE TO HEAL BEFORE THE REVISION IS DONE. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOMES AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01138 | 2012-08-07 | 2012-07-17 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT, PT HAS HIP PAIN. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MODE NCK 8 DEG 30 MM, CAT# NLV-300800G, LOT# 32795201; 28MM STD LFIT V40 HEAD, CAT# 6260-9-128, LOT# 34894801; RESTORATION (TM) ADM. CUP W/HA, CAT# 1235-2-502, LOT# G2958387; RESTORATION ADM X3 INS 28/50, CAT# 1236-2-850, LOT# 34169701. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER DUE TO HOSPITAL POLICY. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS ) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOMES AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01116 | 2012-08-07 | 2010-07-14 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 10 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-100000S |  | IT WAS REPORTED THAT, APPROX 8 MONTHS AGO, AN MRI REVEALED A TORN ABDUCTOR TENDON. BLOOD TESTS WERE INCONCLUSIVE AT THAT TIME. THE PT HAD BEEN EXPERIENCING PAIN AND SICKNESS. ON (B)(6) 2012 AT (B)(6) MEDICAL CENTER (B)(6) FOLLOW UP PROCEDURE WAS TO REPAIR THE TORN TENDON. WHEN THE SURGEON OPENED THE PT, HE FOUND SEVERE ALTR. A SIGNIFICANT AMOUNT OF TISSUE HAD TO BE REMOVED. THE SURGEON RESURFACED THE PELVIS, REPLACED THE BALL AND REPAIRED THE TENDON. THE SURGEON STATED HE NEEDED TO REPAIR THE TENDON TO TRY TO RETURN PROPER BLOOD SUPPLY TO THE AREA SO ANOTHER REVISION COULD BE PERFORMED. A FOLLOW UP APPOINTMENT WITH THE SURGEON IS SCHEDULED (B)(6) 2012 AND AN UPCOMING REVISION WILL NEED TO BE PERFORMED TO REMOVE THE FEMUR AND PERFORM A PROPER REVISION. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. WHEN COMPLETED, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PATIENT/EVENT AS: MFR # 9616680-2012-00628. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01149 | 2012-08-08 | 2010-06-18 | I, n, j, u, r, y | Manufacturer report | LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED PER THE PATIENT'S HUSBAND, PATIENT AT THE BEGINNING WAS "OK" BUT THEN THE PAIN BEGAN A FEW MONTHS AFTER THE SURGERY. THIS IS AN ON AND OFF BURNING PAIN FEELING. X-RAYS WERE NORMAL AND SHE WAS SENT TO PHYSICAL THERAPY BUT IS NOT HELPING HER. X-RAYS AND BLOOD TESTS WERE TAKEN TODAY. PATIENT IS ALSO LIMPING. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01142 | 2012-08-08 | 2011-09-27 | I, n, j, u, r, y | Manufacturer report | RIGHT UNKNOWN REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT COMPLAINS OF PAIN AND SORENESS. SHE STATES THAT THE PROBLEM HAS BEEN THERE SINCE THE SURGERY BUT RECENTLY IT'S MORE NOTICABLE. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01152 | 2012-08-08 | 2012-05-14 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT, SURGEON TOOK PATIENT TO SURGERY ON (B)(6) 2012 FOR ABDUCTOR REPAIR. SURGEON REPLACED MODULAR NECK AND HEAD AT THAT TIME. SURGEON NOTED SOFT TISSUES, DESCRIBED AS "ABNORMAL". | ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. WHEN COMPLETED, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PATIENT EVENT AS MFR # 9616680-2012-00635. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01151 | 2012-08-08 | 2012-07-19 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT, REJUVENATE STEM WAS EXPLANTED DUE TO ADVERSE LOCAL TISSUE REACTION. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER DUE TO HOSPITAL POLICY. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED, BUT MADE AVAILABLE. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PATIENT/EVENT AS MFR # 9616680-2012-00634. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01146 | 2012-08-08 | 2012-04-03 | I, n, j, u, r, y | Manufacturer report | UNKNOWN PRODUCT LEFT HIP REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT IS EXPERIENCING PAIN AND DISCOMFORT. HE STATES THAT HE IS IN CHRONIC PAIN. HE HAS BEEN FOLLOWING UP WITH HIS SURGEON AND HAS HAD X-RAYS AND MRI. HE HAS HAD CORTISONE SHOTS FOR THE PAIN AND SURGEON HAS COMMUNICATED TO PATIENT THAT HE WILL NEED A REVISION. HE HAS ALSO BEEN ON PAIN MEDICATIONS FOR THE LAST YEAR. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | 0 |  | 2012-08
2249697-2012-01148 | 2012-08-08 | 2012-05-07 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT IMPLANTS WERE REMOVED FOR SUSPECTED INFECTION. | CATALOG NUMBERS AND LOT CODES OF OTHER DEVICES LISTED IN THIS REPORT: DESCRIPTION: REJUVENATE MODULAR NECK, CAT #NLS-300000B, LOT #35368001; DESCRIPTION: TRIDENT PSL WITH PUREFIX HA 46MM, CAT #542-11-46D, LOT #34081102; DESCRIPTION: TRIDENT 0 DEG X3 INSERT 32MM, CAT #623-00-32D, LOT #MKJMV4; DESCRIPTION: DELTA V-40 CERAMIC HEAD 32/0, CAT #6570-0-132, LOT #36132401; DESCRIPTION: 6.5 CANCELLOUS BONE SCREW 30MM, CAT #2030-6530-1, LOT #MKHR88. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S EXPERIENCE. A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-08
2249697-2012-01172 | 2012-08-13 | 2011-10-26 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PT REPORTED TO HAVE NUMBNESS, PAIN AND BURNING SENSATION ON THE HIP/THIGH. SHE REPORTS THAT THESE SYMPTOMS HAVE ALWAYS BEEN PRESENT, IT NEVER GOES AWAY. PT HAD A F/U VISIT WITH HER SURGEON ON (B)(6) 2012. X-RAYS WERE TAKEN THE SAME DAY. SHE WILL F/U WITH HER SURGEON. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORD) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01163 | 2012-08-13 | 2012-07-05 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | THE SURGEON STATED THE PT TESTED POSITIVE FOR A (B)(6). SURGEON REMOVED MODULAR NECK, FEMORAL HEAD, MDM LINER, MDM INSERT, WASHED PT OUT. REPLACEMENT LINER, INSERT, AND FEMORAL HEAD WERE IMPLANTED. AS MODULAR HEAD WAS ON A PRODUCT HOLD, THE SURGEON CHOSE TO AUTOCLAVE MODULAR NECK AND RE-IMPLANTS. | CATALOG NUMBERS AND LOT CODES OF OTHER DEVICES LISTED IN THIS REPORT. CAT # NLS-300000B, LOT # 3870002, DESCRIPTION: REJUVENATE MODULAR NECK. CAT # 502-11-46C, LOT # 34896701, DESCRIPTION: TRIDENT HEMISPHERICAL CLUSTER HOLE SHELL. CAT # 1236-2-242, LOT # 39073001, DESCRIPTION: RESTORATION ADM X3 INS. CAT # 626-00-36C, LOT # 38150001, DESCRIPTION: MODULAR DUAL MOBILITY INSERT. CAT # 2030-6525-1, LOT # MLDHNH, DESCRIPTION: 6.5 CANCELLOUS BONE SCREW 25MM. CAT # 6260-4-122, LOT # 38730002, DESCRIPTION: 22.2MM STD V40 HEAD. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S EXPERIENCE. A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-08
2249697-2012-01155 | 2012-08-13 | 2010-11-11 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT THE SURGEON REMOVED IMPLANTS DUE TO STREP INFECTION. | CATALOG NUMBERS AND LOT CODES OF OTHER DEVICES LISTED IN THIS REPORT: CAT # NLS-380000B, LOT # 26582101, DESCRIPTION: LRG TAP PRI MOD NCK 0 DEG 38MM; CAT # 502-11-52E, LOT # 31346601, DESCRIPTION: TRIDENT HEMISPHERICAL CLUSTER HOLE SHELL; CAT # 623-00-36E, LOT # MJHEWE, DESCRIPTION: TRIDENT 0 DEG X3 INSERT 36MM ID; CAT # 6260-9-036, LOT # MJA560, DESCRIPTION: V40 COCR LFIT HEAD 36MM/-5; CAT # 2030-6520, LOT # MHTJRY, DESCRIPTION: CANCELLOUS BONE SCREW 6.5MM. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S EXPERIENCE. A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | 0 |  | 2012-08
2249697-2012-01159 | 2012-08-13 | 2012-01-18 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT IS UNABLE TO RAISE HIS LEG, ADN REPORTS PAIN AROUND THE GROIN AREA. PT IS FOLLOWING UP WITH HIS SURGEON AND WILL HAVE MRI AND BLOOD WORK FOR FURTHER EVALUATION. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01156 | 2012-08-13 | 2011-11-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT COMPLAINS OF PAIN AND SWELLING SINCE THE SURGERY. SURGEON HAS SCHEDULED A REVISION SURGERY FOR (B)(6) 2012. PT HAS HAD MRIS, X-RAYS, AND BLOOD WORK. BLOOD ION TESTS HAVE SHOWED SOME INCREASE IN METAL BUT PT IS UNSURE OF THE EXACT RESULTS. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01168 | 2012-08-13 | 2011-08-09 | I, n, j, u, r, y | Manufacturer report | UNKNOWN PRODUCT LEFT STRYKER REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT COMPLAINS OF PAIN, BURNING SENSATION AND SORENESS AROUND THE INCISION AREA. PT STATES THAT IT IS DIFFICULT FOR HIM TO EXERCISE AND WALK. HE HAS DAILY DISCOMFORT. PT WILL HAVE A JOINT ASPIRATION IN THE NEXT FEW DAYS AND WILL SCHEDULED A REVISION SURGERY WITH HIS SURGEON THIS MONTH. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01166 | 2012-08-13 | 2012-01-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN PRODUCT LEFT STRYKER REJUVENATE MODULAR HI | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PT COMPLAINS OF PAIN, SHORTNESS AND RESTRICTED MOVEMENT IN HIP AREA. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01160 | 2012-08-13 | 2011-10-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE MODULAR HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT COMPLAINS OF STIFFNESS AND PAIN TOWARDS THE GROIN AREA AND OUTSIDE OF THE HIP. SHE WILL NEED FURTHER TESTS. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01161 | 2012-08-13 | 2011-07-20 | I, n, j, u, r, y | Manufacturer report | UNKNOWN PRODUCT RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT COMPLAINS OF PAIN SINCE PRIMARY SURGERY AND SWELLING THAT LED TO THE REVISION. PT RECEIVED CORTISONE SHOTS BUT NOTHING EVER HELPED THE PAIN. PT WILL BE FOLLOWING UP WITH HIS SURGEON IN ABOUT 3 WEEKS. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01158 | 2012-08-13 | 2009-02-02 | I, n, j, u, r, y | Manufacturer report | UNKNOWN PRODUCT LEFT HIP REJUVENATE MODULAR HIP ST | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PATIENT COMPLAINS OF PAIN AND STIFFNESS. PATIENT WILL FOLLOW UP WITH SURGEON AND ADDITIONAL TESTS AS HE WAS ADVISED BY THE DOCTORS OFFICE TO GO AHEAD AND FILE A CLAIM. PATIENT WILL LIKE TO KNOW WHO PAYS FOR ALL THESE COSTS. (CLAIMS EVALUATION PROCESS.) | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01157 | 2012-08-13 | 2011-10-31 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT IS EXPERIENCING PAIN AND SWELLING. HE WALKS WITH A LIMP. PT HAS CALLED HIS DOCTOR AND HE HAS SCHEDULED A VISIT TO HAVE FURTHER EVALUATION. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01192 | 2012-08-13 | 2011-01-11 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT HIP REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT COMPLAINS HIP/THIGH PAIN AND BURNING SENSATION DOWN THE LEG AND KNEE. PATIENT STATES THAT IT IS DIFFICULT TO WALK. THERE IS SOME OCCASIONAL SWELLING. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01154 | 2012-08-13 | 2010-09-20 | I, n, j, u, r, y | Manufacturer report | UNKNOWN PRODUCT RIGHT HIP REJUVENATE/ABG II STRYKE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PT COMPLAINS OF PAIN AND DIFFICULTY WALKING SINCE SURGERY. A MICROSCOPIC SURGERY WAS PERFORM ON (B)(6) 2011 TO SEE WHAT WAS CAUSING THE PROBLEM. PT STATES THAT THEY REMOVED SOME TISSUE BUT NO IMPLANTS WERE REMOVED. SHE HAS HAD MRI AND BLOOD WORK. PT HAS BEEN UNABLE TO GO BACK TO WORK SINCE HER SECOND MICROSCOPIC INTERVENTION AND WILL LIKE TO KNOW WHAT WILL STRYKER COVER. PT WILL BE HAVING A REVISION DUE TO THE CONTINOUS PAIN AND INABILITY TO BEAR WEIGHT ON THAT HIP. (CLAIMS EVALUATION PROCESS). | AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01208 | 2012-08-13 | 2012-08-06 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT'S DAUGHTER CALLED ON BEHALF OF HER DAD WHO HAS A REJUVENATE HIP AND SINCE ONSET OF SURGERY HAS NEVER BEEN OUT OF PAIN. SINCE SURGERY HE HAS BEEN IN THERAPY, PAIN NEVER SUBSIDING. HE IS HAVING A REVISION THIS MONDAY. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01209 | 2012-08-13 | 2012-08-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT THE PT WAS REVISED DUE TO PAIN. | CATALOG NUMBERS AND LOT CODES OF OTHER DEVICES LISTED IN THIS REPORT: CAT#NLS-340000B, LOT# 38893902, DESCRIPTION: ERG TAP PRI MOD NCK 0DEG 34MM; CAT # 6570-0-436, LOT# 39149901, DESCRIPTION: DELTA V-40 CERAMIC HEAD 36/-2,5; CAT# 540-11-52E, LOT# 38325501, DESCRIPTION: TRIDENT PSL HA SOLID BACK 52MM; CAT# 623-00-36E, LOT# MLA51Y, DESCRIPTION: TRIDENT 0 X3 INSERT 36MM ID. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S EXPERIENCE. ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01229 | 2012-08-13 | 2011-10-15 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT BEGAN EXPERIENCING PAIN ABOUT 2 MONTHS AFTER THE SURGERY. DUE TO THE PAIN PT HAS DIFFICULTY SLEEPING, IS UNABLE TO WALK OR DO PHYSICAL EXERCISE. PT WENT TO A ROUTINE CHECK UP LAST WEEK AND TESTS WERE ORDERED. HE STATES THAT HIS BLOOD RESULTS SHOW EVIDENCE OF COBALT AND CHROMIUM METALS. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01222 | 2012-08-13 | 2011-04-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT COMPLAINS OF PAIN SINCE SURGERY. IN THE PAST, TESTS HAVE BEEN PERFORMED BUT NOTHING HAS BEEN FOUND. HOWEVER, PT CONTINUES TO HAVE PAIN AND IS UNABLE TO BEAR WEIGHT ON THAT HIP. HE HAS A LIMP. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01226 | 2012-08-13 | 2011-08-04 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT COMPLAINS OF PAIN. HE HAS HAD STEROID INJECTIONS. AN MRI, BONE SCAN AND HIP ASPIRATION HAVE BEEN SCHEDULED FOR THE PT. HE IS SCHEDULED FOR REVISION SURGERY ON (B)(6) 2012. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01207 | 2012-08-13 | 2012-06-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT IS EXPERIENCING PAIN SINCE (B)(6) OF THIS YEAR. SHE ALSO REPORTS SOME CLICKING NOISE. SHE WILL FOLLOW UP WITH HER SURGEON. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01200 | 2012-08-13 | 2012-07-23 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT EARLIER THIS SPRING, THE PATIENT HAD BEEN GARDENING, BENDING AND STANDING AND HAD SOME ANTERIOR HIP PAIN. THEN ABOUT 3-4 WEEKS AGO, SHE COMPLAINED OF INCREASING HIP PAIN AND WAS SEEN IN THE DOCTOR'S OFFICE ON (B)(6) 2012. SHE CAME TO THE EMERGENCY ROOM ON SATURDAY WITH PAIN AND HIP ASPIRATE CAME BACK POSITIVE FOR STREP. SHE WAS BROUGHT TO THE OPERATING ROOM TODAY FOR EXPLANT OF HER PROSTHESIS AND INTRODUCTION OF AN ARTICULATING ANTIBIOTIC SPACER. | CATALOG NUMBERS AND LOT CODES OF OTHER DEVICES LISTED IN THIS REPORT: CAT # NLV-340800Y, LOT # 25520001, DESCRIPTION: REJUVENATE NECK. CAT # 502-03-56E, LOT # MJPV31, DESCRIPTION: PRIMARY TRITANIUM HEMI CLUSTER HOLE CUP 56MM, CAT # 6260-9-140, LOT # MJL5H0, DESCRIPTION: V40 COCR LFIT HEAD 40MM/+0. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S EXPERIENCE. A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-08
2249697-2012-01227 | 2012-08-13 | 2012-08-06 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT COMPLAINED OF PAIN IN GROIN, AND CAN'T MOVE IN AND OUT OF CAR WITHOUT REACHING FOR HER PANT LEG TO PULL IT UP. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01203 | 2012-08-13 | 2011-08-31 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT WAS "OK" FOR THE FIRST FEW MONTHS. HOWEVER, AFTER THE FOURTH MONTH, PATIENT BEGAN EXPERIENCING PAIN. THE PATIENT REPORTED MILD PAIN INITIALLY BUT THE PAIN INCREASED GRADUALLY THAT IT NEEDS TO BE REPLACED. HE IS SCHEDULED FOR A REVISION ON (B)(6) 2012. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01179 | 2012-08-13 | 2010-05-03 | I, n, j, u, r, y | Manufacturer report | UNKNOWN PRODUCT RIGHT HIP REJUVENATE/ABGII | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT COMPLAINS OF TENDERNESS AND OCCASIONAL PAIN. RECEIVED LETTER FROM HIS SURGEON ABOUT THE RECALL STATING THAT HE NEEDS AN MRI. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01195 | 2012-08-13 | 2011-08-02 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE MODULAR NECK HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT COMPLAINS OF SWELLING, PAIN AND DIFFICULTY TO WALK AND EXERCISE. PATIENT STATES SHE ALSO HAS A METAL TASTE IN HER MOUTH. PATIENT SPOKE TO HER DOCTOR AND STATES THAT SHE WILL NEED FURTHER TESTING. MRI AND BLOOD TESTS. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01253 | 2012-08-14 | 2011-07-27 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT HIP REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT HAS BEEN EXPERIENCING PAIN SINCE SURGERY. | THE PT IS (B)(6). ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01239 | 2012-08-14 | 2010-07-18 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT REPORTS PAIN IN GROIN, LEGS, AND LOWER BACK. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01254 | 2012-08-14 | 2011-12-05 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT HIP REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PAIN, SWELLING, STIFFNESS, RASH FOR 3-4 DAYS, UNABLE TO SLEEP. REVISION SURGERY IS SCHEDULED FOR (B)(6) 2012. | THE PT IS (B)(6). ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01248 | 2012-08-14 | 2011-08-03 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE MODULAR NECK HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT HAS EXPERIENCED MINIMAL SORENESS. | THE PATIENT IS (B)(6). ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01243 | 2012-08-14 | 2012-01-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT EXPERIENCED PAIN AROUND TOP OF LEG AND HIP, HAVING TROUBLE GETTING UP. METAL LEVELS ARE 4. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01259 | 2012-08-14 | 2010-05-25 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE LEFT HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT IS EXPERIENCING PAIN IN LEFT HIP AND LEGS, AND IS NOT SURE IF THIS PAIN IS DUE TO THE SURGERY OR HER AGE. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01244 | 2012-08-14 | 2010-03-25 | I, n, j, u, r, y | Manufacturer report | LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT COMPLAINS OF PAIN AND DISCOMFORT SINCE THE SURGERY. PATIENT HAS HAD ASPIRATION, BLOOD TEST AND MRI. SURGEON STATED TO PATIENT THAT THE MRI SHOWED A PSEUDOTUMOR AND THAT THIS CAN CAUSE PERMANENT MUSCLE DAMAGE. THEREFORE, HE WILL NEED TO HAVE A REVISION. | THE PATIENT IS (B)(6). ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01271 | 2012-08-15 | 2012-08-06 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 6 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT, PT WAS EXPERIENCING PAIN AND HAD ELEVATED COBALT-CHROMIUM LEVELS. PT WAS TAKEN TO THE OPERATING ROOM AND HAD ALL COMPONENTS REMOVED SUCCESSFULLY. REVISION COMPONENT PREPARATION AS CARRIED OUT AND REIMPLANTATION WAS SUCCESSFULLY CARRIED OUT. | THIS IS THE SAME PT/EVENT AS: MFR # 9616680-2012-00657. AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE HOSPITAL COULD NOT RELEASE DEVICE SO IT WAS NOT RETURNED TO THE MANUFACTURER. X-RAYS WERE NOT MADE AVAILABLE EITHER. WHEN COMPLETED, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01284 | 2012-08-15 | 2006-04-03 | I, n, j, u, r, y | Manufacturer report | ACCOLADE PLUS TMZF HIP STEM #3 | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, POROUS UNCEMENTED | STRYKER ORTHOPAEDICS-MAHWAH | NA | 6021-0335 |  | THE PATIENT REPORTED THAT SHE NEEDED A REVISION FOR HER RIGHT HIP BECAUSE AT THE VERY END OF THE STEM THERE WAS A GAP. THE SURGEON PUT IN A BONE GRAFT AT THE VERY TIP OF THE STEM TO CLOSE THE GAP. SURGEON STATED THAT THERE WERE "HOT SPOTS" ON THE BONE. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01281 | 2012-08-15 | 2012-04-29 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT THE PATIENT SLIPPED AND FELL ON (B)(6) 2012. THE SURGEON DECIDED TO REVISE ON (B)(6) 2012. WHEN THE SURGEON WAS REVISING THE PATIENT, A MILKY FLUID WAS NOTED SO THE SURGEON COMPLETELY REVISED THE PATIENT. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PATIENT/EVENT AS MFR # 9616680-2012-00659. | Y | N | OTHER |  | 2012-08
2249697-2012-01267 | 2012-08-15 | 2012-07-26 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT THE PT HAS A REJUVENATE MODULAR HIP IMPLANTED AND HAS BEEN EXPERIENCING PAIN, SWELLING, ELEVATED METAL LEVELS IN HIS BLOOD AND A HOT BONE SCAN. THE DOCTOR HAS RECOMMENDED THAT THE PT HAVE A REVISION SURGERY. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS MFR # 9616680-2012-00656. | Y | N | OTHER |  | 2012-08
2249697-2012-01272 | 2012-08-15 | 2012-04-03 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT HIP REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT COMPLAINS OF PAIN, SORENESS, HIGH METAL LEVELS. | THE PATIENT IS (B)(6). ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01268 | 2012-08-15 | 2010-07-08 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT HIP REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT HAS PAIN AS WELL AS RHEUMATOID ARTHRITIS. | THE PATIENT IS (B)(6). ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01288 | 2012-08-15 | 2012-08-07 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT THE PT HAD HIP PAIN. | THE REPORTED LOT ID IS NOT VALID FOR THE REPORTED DEVICE. CAT NUMBERS AND LOT CODES OF OTHER DEVICES LISTED IN THIS REPORT: DESC: RESTORATION ADM X3 INS 28/54: CAT# 1236-2-854, LOT# 36297801. DESCP: RESTORATION (TM) ADM. CUP W/HA: CAT # 1235-2-541, LOT# G3057945. DESCP: LRG TAP PRI MOD NCK 0DEG 30MM: CAT # NLS-300000B, LOT# 37106301. DESC: 28MM-4 LFIT V40 HEAD: CAT # 6260-9-028, LOT# 36580602. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S EXPERIENCE. ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01283 | 2012-08-15 | 2012-08-10 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | PROSTHESIS, HIP, SEMI-CONSTRAINED, UNCEMENTED, METAL/POLYMER, NON-POROUS, CALICU | STRYKER ORTHOPAEDICS-MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT THE PT IS COMPLAINING OF PAIN AND SWELLING. THE PT REPORTS HE PLAYS GOLF AND RAQUETBALL. HE FURTHER REPORTS IN THE BEGINNING HE HAD NO PAIN BUT HAS EXPERIENCED PAIN IN THE LAST 2 MONTHS. | THE PT IS (B)(6). AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER, BELIEVES THEM TO BE EITHER REJUVENATE OR A BG II. ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01285 | 2012-08-15 | 2011-07-07 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 10 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPOT-100000S |  | THE PATIENT REPORTED THAT HE IS EXPERIENCING STIFFNESS AND PAIN SINCE THE PRIMARY SURGERY. PATIENT FEELS IT DID NOT HEAL PROPERLY. SCHEDULED FOR A REVISION AT THE END OF (B)(6) 2012. | CATALOG NUMBERS AND LOT CODES OF OTHER DEVICES LISTED IN THIS REPORT: CAT # NLS-380000B, LOT # 33808802, DESCRIPTION: REJUVENATE MODULAR NECK - 127/132 DEG NECK ANGLE. CAT # 502-03-56E, LOT # MKE8HD, DESCRIPTION: PRIMARY TRITANIUM HEMI CLUSTER HOLE CUP 56MM. CAT # 623-0036E, LOT # MKE90D, DESCRIPTION: TRIDENT 0X3 INSERT 36MM ID. CAT # 6570-0-136, LOT # 36955903, DESCRIPTION: DELTA V-40 CERAMIC HEAD 36/0. CAT # 2030-6525-1, LOT #MKE4DX, DESCRIPTION: 6.5 CANCELLOUS BONE SCREW 25MM. CAT # 2030-6535-1, LOT # MKEY2Y, DESCRIPTION: 6.5 CANCELLOUS BONE SCREW 35MM. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S EXPERIENCE. ADDITIONAL INFORMATION HAS BEEN REQUESTED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-08
2249697-2012-01295 | 2012-08-16 | 2010-11-23 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE RIGHT HIP | IMPLANT | STRYKER ORTHOPAEDICS | NA | UNK |  | IT WAS REPORTED THAT THE PT HAS BEEN EXPERIENCING PAIN, STIFFNESS, SORENESS, AND SWELLING SINCE DATE OF IMPLANT. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01296 | 2012-08-16 | 2012-07-30 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPANT | STRYKER ORTHOPAEDICS | NA | SPT-080000S |  | IT WAS REPORTED THAT THE PT COMPLAINED OF PAIN AND THE HIP WAS REVISED. ACCORDING TO THE SURGEON, SOFT TISSUE INFLAMMATION WAS FOUND IN AND AROUND THE JOINT CAPSULE. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. THIS IS THE SAME PT/EVENT AS MFR # 9616680-2012-00664. | Y | N | OTHER |  | 2012-08
2249697-2012-01292 | 2012-08-16 | 2012-07-28 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANTED | STRYKER ORTHOPAEDICS | NA | SPT-070000S |  | IT WAS REPORTED THAT, PT WAS EXPERIENCING HIP PAIN. | CATALOG NUMBERS AND LOT CODES OF OTHER DEVICES LISTED IN THIS REPORT: CAT #NLV-300800Y, LOT # 26794102, DESCRIPTION: REJUVENATE MODULAR NECK. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S EXPERIENCE. A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-08
2249697-2012-01324 | 2012-08-21 | 2011-12-22 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: THE PATIENT COMPLAINED OF A BURNING SENSATION - NOT SURE IF ASSOCIATED WITH IMPLANT. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINS IMPLANTED IN THE PATIENT IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WILL BE REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01322 | 2012-08-21 | 2011-03-03 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT PATIENT HAD HIP PAIN. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0DEG 30MM, CAT# NLS-300000B, LOT# 34565102; RESTORATION ADM. CUP W/HA, CAT# 1235-2-461, LOT# G1843866; RESTORATION ADM X 3 INS 28/46, CAT# 1236-2-846, LOT# 34568201; 28MM STD LFIT V40 HEAD, CAT# 6260-9-128, LOT# 35522005. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S REVISION PAIN. AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WILL BE REQUESTED. SHOULD DEVICE OR ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01327 | 2012-08-21 | 2012-02-14 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: THE PATIENT IS COMPLAINING OF PAIN IN HIS MID-THIGH INCISION, CLAIMING THAT IT IS INFLAMMED AND REDDISH. HE FURTHER REPORTS HE ALSO HAD SURGERY IN HIS RIGHT HIP WHERE A REJUVENATE HIP WAS USED IN (B)(6) 2011 WITH DR. (B)(6) NO SYMPTOMS WERE REPORTED WITH THE RIGHT HIP. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01314 | 2012-08-21 | 2011-04-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT HIP REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PATIENT COMPLAINS OF PAIN AND STIFFNESS SINCE SURGERY. PATIENT RECEIVED RECALL LETTER AND WILL FOLLOW UP WITH HER SURGEON FOR FURTHER TESTING AND EVALUATION. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WILL BE REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01318 | 2012-08-21 | 2012-08-08 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT, THE PATIENT WAS HAVING PAIN SO THE SURGEON REVISED. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0DEG 34MM, CAT# NLS-340000B, LOT# 38369301; TRIDENT 0 DEG X3 INSERT 36MM ID, CAT# 623-00-36F, LOT# MKPV3D; TRIDENT PSL HA SOLID BACK 56MM, CAT# 540-11-56F, LOT# 37809701; DELTA V-40 CERAMIC HEAD 36/2,5, CAT# 6570-0-436, LOT# 38309401. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S PAIN. AN EVALUATION OF THE DEVICES CANNOT BE PERFORMED AS THE REVISED DEVICES WERE RETAINED BY THE HOSPITAL AND WERE NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01320 | 2012-08-21 | 2011-12-20 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT HIP REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PATIENT'S DAUGHTER CALLING ON BEHALF OF HER FATHER ABOUT COMPENSATION FOR HIS PAIN. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WILL BE REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01334 | 2012-08-21 | 2010-07-14 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: THE PATIENT COMPLAINS OF SOME SWELLING, BUT NO PAIN. SHE HAS RECEIVED A RECALL LETTER AND EXPECTS TO BE PAID FOR ALL THESE COSTS. THE PATIENT IS SCHEDULED TO SEE HER SURGEON ON (B)(6) 2012. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WILL BE REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01319 | 2012-08-21 | 2012-08-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THE PT'S LEFT HIP WAS REVISED DUE TO PAIN. | CATALOG NUMBERS AND LOT CODES OF OTHER DEVICES LISTED IN THIS REPORT. CAT # NLS-340000B, LOT # 37413202, DESCRIPTION: LRG TAP PRI MOD NCK 0DEG 34MM. CAT # 6570-0-128, LOT # 37785005, DESCRIPTION: DELTA V-40 CERAMIC HEAD 28/0. CAT # 1235-2-542, LOT # G3039401, DESCRIPTION: RESTORATION AMD. CUP W/HA. CAT # 1236-2-854, LOT # 35041001, DESCRIPTION: RESTORATION ADM X3 INS 28/54. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S EXPERIENCE. A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-08
2249697-2012-01342 | 2012-08-21 | 2010-01-10 | I, n, j, u, r, y | Manufacturer report | REJUVENATE LEFT HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: THE PATIENT COMPLAINS OF OCCASIONAL PAIN IN LEFT LEG AROUND THIGH AREA. PATIENT WANTS TO KNOW WHO WILL PAY FOR TESTING. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WILL BE REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01315 | 2012-08-21 | 2012-08-03 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT THE PT WAS REVISED DUE TO PAIN. | CATALOG NUMBERS AND LOT CODES OF OTHER DEVICES LISTED IN THIS REPORT. CAT # NLS-340000B, LOT # 35286701, DESCRIPTION: LRG TAP PRI MOD NCK 0DEG 34MM. CAT # 6570-0-536, LOT # 37314605, DESCRIPTION: DELTA V-40 CERAMIC HEAD 36/+2.5. CAT # 623-00-36F, LOT # MKL43P, DESCRIPTION: TRIDENT 0X3 INSERT 36MM ID. CAT # 540-11-54F, LOT # 31773801, DESCRIPTION: TRIDENT PSL HA SOLID BACK 54 MM. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S EXPERIENCE. A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-08
2249697-2012-01336 | 2012-08-21 | 2010-04-14 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: THE PATIENT COMPLAINS OF SWELLING, PAIN, BURNING SENSATION. PATIENT REPORTS THAT HIS LEFT HIP IS EXTREMELY SENSITIVE TO THE TOUCH. THE PATIENT IS CURRENTLY WITH NO INSURANCE AND WOULD LIKE FOR STRYKER TO COVER ALL THE ADDITIONAL TESTING HE WILL REQUIRE. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WILL BE REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01310 | 2012-08-21 | 2012-07-30 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT THE PT WAS REVISED DUE TO PAIN. | CATALOG NUMBERS AND LOT CODES OF OTHER DEVICES LISTED IN THIS REPORT. CAT # NLS-300000B, LOT # 38307501, DESCRIPTION: LRG TAP PRI MOD NCK 0DEG 30MM. CAT # 540-11-46D, LOT # 38119001, DESCRIPTION: TRIDENT PSL HA SOLID BACK 46MM. CAT # 6570-0-132, LOT # 39131001, DESCRIPTION: DELTA V-40 CERAMIC HEAD 32/0. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S EXPERIENCE. A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-08
2249697-2012-01308 | 2012-08-21 | 2012-08-03 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT THE PT PRESENTED TO THE OPERATING ROOM WITH PAIN. EXPLORATION AND EXPLANTATION OF TOTAL HIP WAS PERFORMED AND REVISION COMPONENTS WERE IMPLANTED AFTER PREPARATIONS. | CATALOG NUMBERS AND LOT CODES OF OTHER DEVICES LISTED IN THIS REPORT. CAT # NLS-34000B, LOT # 36080401, DESCRIPTION: REJUVENATE MODULAR NECK. CAT # 502-03-58F, LOT # MKAH11, DESCRIPTION: PRIMARY TRITANIUM HEMI CLUSTER HOLE CUP 58MM. CAT # 623-00-36F, LOT # MKJ2EM, DESCRIPTION: TRIDENT 0X3 INSERT 36MM ID. CAT # 6570-0-136, LOT # 36591901, DESCRIPTION: DELTA V-40 CERAMIC HEAD 36/0. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S EXPERIENCE. A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. IT WAS NOTED THAT THE HOSPITAL IS NOT WILLING TO PROVIDE THE DEVICES OR ADD'L INFO. | Y | N | OTHER |  | 2012-08
2249697-2012-01333 | 2012-08-22 | 2012-07-31 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 11 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-110000S |  | IT WAS REPORTED THAT THE DOCTOR REVISED THE PT'S LEFT HIP DUE TO ADVERSE LOCAL TISSUE REACTION. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. THIS IS THE SAME PT/EVENT AS MFR # 9616680-2012-00671. | Y | N | OTHER |  | 2012-08
2249697-2012-01330 | 2012-08-22 | 2011-09-12 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | THE PT COMPLAINS OF PAIN ALONG THE THIGH AND MUSCLES SINCE THE SURGERY. PT IS FOLLOWING UP WITH HER SURGEON AND WILL BE HAVING BLOOD WORK TO DETERMINE THE CAUSE OF HER SYMPTOMS. | CATALOG NUMBERS AND LOT CODES OF OTHER DEVICES LISTED IN THIS REPORT: CAT # NLS-300000B, LOT # 35368002, DESCRIPTION: LRG TAP PRI MOD NECK 0DEG 30MM. CAT # 6570-0-228, LOT # 35728702, DESCRIPTION: DELTA V-40 CERAMIC HEAD 28/+4. CAT # 1235-2-481, LOT # G3040424, DESCRIPTION: RESTORATION (TM) ADM. CUP W/HA. CAT # 1236-2-848, LOT # 35101401, DESCRIPTION: RESTORATION ADM X3 INS 28/48. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S EXPERIENCE. A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-08
2249697-2012-01323 | 2012-08-22 | 2010-07-14 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT HIP REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | THE PT REPORTED THAT HER PRIMARY LEFT THA WAS (B)(6) 2010. SHE WILL NEED TO HAVE THIS REPLACED DUE TO COBALT LEVELS BEING HIGH AND PAIN. THE REVISION SURGERY WILL BE IN APPROXIMATELY 2 MONTHS. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01335 | 2012-08-22 | 2012-07-31 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT THE SURGEON PERFORMED A REVISION ON THE PT. NO ADDITIONAL INFO WAS AVAILABLE. | THE REASON FOR THE PT'S REVISION WAS NOT PROVIDED. THIS CASE HAS BEEN TURNED OVER TO AN ATTORNEY. ADDITIONAL INFO WILL BE REQUESTED THROUGH STRYKER ORTHOPAEDICS LEGAL AFFAIRS. IF ADDITIONAL INFO BECOMES AVAILABLE, IT WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01331 | 2012-08-22 | 2011-11-29 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT HIP REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | THE PT REPORTED THAT SHE HAD HER PRIMARY SURGERY ON THE RIGHT HIP (B)(6) 2011 DUE TO AVASCULAR NECROSIS. ON (B)(6) 2011, IT HAD TO BE REPLACED BECAUSE THE HIP NEVER SETTLED IN. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01345 | 2012-08-22 | 2012-08-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT THE PT HAD A REVISION SURGERY DUE TO PAIN AND PSEUDO TUMOR. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. THIS IS THE SAME PT/EVENT AS MFR # 9616680-2012-00672. | Y | N | OTHER |  | 2012-08
9616680-2012-00674 | 2012-08-23 | 2012-07-13 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | NLS-380000B |  | IT WAS REPORTED THAT THE DOCTOR REVISED PT'S LEFT HIP DUE TO ADVERSE LOCAL TISSUE REACTION. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. THIS IS THE SAME PT/EVENT AS MFR #2249697-2012-01349. | Y | N | OTHER |  | 2012-08
2249697-2012-01355 | 2012-08-23 | 2012-08-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT THE SURGEON REVISED THE PT'S LEFT HIP DUE TO ADVERSE LOCAL TISSUE REACTION. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. THIS IS THE SAME PT/EVENT AS MFR # 9616680-2012-00679. | Y | N | OTHER |  | 2012-08
2249697-2012-01371 | 2012-08-23 | 2011-02-02 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE/ABGII HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PATIENT IS EXPERIENCING PAIN IN GROIN, DISCOMFORT AND SWELLING. DR. SENT LETTER AND CALLED PATIENT AND STATED FURTHER TESTING, BLOOD WORK AND MRI IS NECESSARY. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OF ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01351 | 2012-08-23 | 2012-07-27 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT THE PT CAME IN TO THE DOCTOR'S OFFICE WITH ACUTE LATERAL THIGH REDNESS. X-RAY REVEALED WHAT APPEARED TO BE LUCENCY ON LATERAL CORTEX. SUSPECTED INFECTION. | CATALOG NUMBERS AND LOT CODES OF OTHER DEVICES LISTED IN THIS REPORT: CAT# NLS-380000B, LOT# 37412501, DESCRIPTION: LRG TAP PRI MOD NCK 0 DEG 38MM; CAT# 542-11-52E, LOT# MKPRJT, DESCRIPTION: TRIDENT PSL HA CLUSTER 52 MM; CAT# 6260-9-136, LOT# MKNL36, DESCRIPTION: V40 COCR LFIT HEAD 36MM/0; CAT# 623-00-36E, LOT# MKN69K, DESCRIPTION: TRIDENT 0 DEG X3 INSERT 36MM ID; CAT# 2030-6530-1, LOT# MKNPWV, DESCRIPTION: 6.5 CANCELLOUS BONE SCREW 30MM; CAT# 2030-6540-1, LOT# MKNAP7, DESCRIPTION: 6.5 CANCELLOUS BONE SCREW 40MM. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S EXPERIENCE. A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-08
2249697-2012-01349 | 2012-08-23 | 2012-07-13 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT THE DOCTOR REVISED PT'S LEFT HIP DUE TO ADVERSE LOCAL TISSUE REACTION. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. THIS IS THE SAME PT/EVENT AS MFR # 9616680-2012-00674. | Y | N | OTHER |  | 2012-08
2249697-2012-01353 | 2012-08-23 | 2012-08-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 10 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-100000S |  | IT WAS REPORTED THAT THE SURGEON EXPLANTED THE PT'S HIP ON (B)(6) 2012 AND IMPLANTED CEMENT HIP SPACERS. PT RECEIVED HIP. ON (B)(6) 2012, DOCTOR REIMPLANTED FINAL COMPONENTS. DURING THIS TIME THE STEM HAD BEEN RECALLED. DOCTOR STATED INFECTION WAS POSSIBLY SECONDARY TO ADVERSE LOCAL TISSUE REACTION. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. THIS IS THE SAME PT/EVENT AS MFR #9616680-2012-00678. CATALOG NUMBERS AND LOT CODES OF OTHER DEVICES LISTED IN THIS REPORT: CAT# 1235-2-561, LOT# G2899438, DESCRIPTION: RESTORATION ADM. CUP W/HA; CAT# 1236-2-856, LOT# 33564801, DESCRIPTION: RESTORATION ADM X3 INS 28/56; CAT #6570-0-228, LOT# 32790801, DESCRIPTION: 28MM BIOLOX HEAD. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S EXPERIENCE. | Y | N | OTHER |  | 2012-08
2249697-2012-01352 | 2012-08-23 | 2012-07-31 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT080000S |  | IT WAS REPORTED THAT THE SURGEON REPLACED COMPONENTS LISTED DUE TO SKIN RASH, PAIN AND PSEUDO TUMOR. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. THIS IS THE SAME PATIENT/EVENT AS MFR # 9616680-2012-00677. | Y | N | OTHER |  | 2012-08
2249697-2012-01366 | 2012-08-23 | 2010-07-23 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT HIP REJUVENATE/ABG II | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PATIENT IS EXPERIENCING LITTLE PAIN, NOT SURE IF IT IS FROM THE HIP. HE WORKS OUT ALOT SO HE ISN'T SURE ABOUT THE PAIN FACTOR. HE RECEIVED A LETTER FROM DR. (B)(6) AND HE WANTS HIM TO GO FOR TESTING, MRI AND BLOOD TEST. | AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01369 | 2012-08-23 | 2010-02-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN PRODUCT RIGHT REJUVENATE/ABG II (ONE OF TH | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PATIENT COMPLAINS OF NUMBNESS, PAIN AND DIFFICULTY LYING DOWN SINCE SURGERY. PATIENT UNABLE TO BEAR WEIGHT ON THE RIGHT HIP. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01362 | 2012-08-23 | 2012-08-02 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STEM #10 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT WAS EXPERIENCING HIP PAIN 3 MONTHS AFTER INITIAL SURGERY. | CATALOG NUMBER AND LOT CODE OF OTHER DEVICE LISTED IN THIS REPORT: CAT # NLS-380000B, LOT # 37413001, DESCRIPTION: LRG TAP PRI MOD NCK 0DEG 38MM. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S EXPERIENCE. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01372 | 2012-08-27 | 2012-08-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT THERE WAS METALLOSIS AT THE JUNCTURE OF THE MODULAR NECK CONNECTED TO THE HIP STEM. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. THIS IS THE SAME PT/EVENT AS MFR # 9616680-2012-00687. | Y | N | OTHER |  | 2012-08
2249697-2012-01385 | 2012-08-28 | 2010-10-07 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT HIP REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT COMPLAINTS OF GROIN PAIN. PT HAS SCHEDULED A VISIT WITH HER SURGEON. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01380 | 2012-08-28 | 2011-06-20 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT THE PT HAD HIGH COBALT/CHROMIUM LEVELS FROM BLOOD TEST REPORTED ON (B)(6) 2012. PT ALSO EXPERIENCING SORENESS. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. THIS IS THE SAME PT/EVENT AS MFR # 9616680-2012-00691. | Y | N | OTHER |  | 2012-08
2249697-2012-01394 | 2012-08-28 | 2012-08-06 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | THE SURGEON REPORTED THAT THE PATIENT HAS BEEN NOTICING CONTINUOUS HIP PAIN SINCE ORIGINAL SURGERY. PATIENT CAME INTO OFFICE AND STILL HAS PAIN AT INCREASING LEVEL. HE HAS BEEN EVALUATED FOR INFECTION AND LOOSENING WITH NEGATIVE RESULTS. | CATALOG NUMBER/LOT CODE OF OTHER DEVICE LISTED IN THIS REPORT: CAT # NLS-301600P, LOT # 26852701. DESCRIPTION: LRG TAP PRI MOD NECK 16DEG 30MM. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S EXPERIENCE. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01384 | 2012-08-28 | 2010-04-06 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | THE PT'S WIFE CALLED ON BEHALF OF HER HUSBAND WHO HAS BEEN IN PAIN SINCE HIS SURGERY 2 YEARS AGO. | ADD'L INFO HAS BEEN REQUETED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01379 | 2012-08-28 | 2009-01-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | THE PT COMPLAINS OF PAIN IN GROIN AREA. IN (B)(6) 2012, PT REPORTS THAT HIS CHROMIUM/COBALT LEVELS WERE HIGH. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01398 | 2012-08-29 | 2012-05-31 | I, n, j, u, r, y | Manufacturer report | UNK LEFT HIP REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS ¿ MAHWAH | NA | UNK |  | THE PT REPORTED NO PAIN OR SWELLING, BUT NOTED THAT HE WAS ABOUT TO FINISH ANTIBIOTICS FOR CELLULITIS. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01420 | 2012-08-31 | 2012-03-01 | I, n, j, u, r, y | Manufacturer report | UNK LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: THE PT IS COMPLAINING OF REAL NAUSEOUS. COBALT LEVELS ARE A LITTLE ELEVATED. SORE TO THE TOUCH. PT WAS IN THE KITCHEN, LEFT LEG BECAME LIKE IT WAS PARALYZED. NEXT DAY SAME FEELING OF NUMBNESS CAME AGAIN BUT SLOWER. WILL BE GOING FOR FURTHER TESTING. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-08
2249697-2012-01416 | 2012-08-31 | 2012-08-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT THE PT'S HIP WAS INFECTED. THE SURGEON REMOVED REJUVENATE STEM, 34MM NECK, 40 CERAMIC HEAD 40MM LINER, BONE SCREW. | CATALOG NUMBERS AND LOT CODES OF OTHER DEVICES LISTED IN THIS REPORT: DESCRIPTION: REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7, CAT #NLS-341600P, LOT #34297201; DESCRIPTION: UNI ADAPTOR SLEEVE V40 TI, CAT #6519-T-100, LOT #37100001; DESCRIPTION: DELTA UN FEM HD 40MM, CAT #6519-1-040, LOT #37862001; DESCRIPTION: CANCELLOUS BONE SCREW 6.5MM, CAT # 2030-6535, LOT # MKE5VW; DESCRIPTION: PRIMARY TRITANIUM HEMI CLUSTER HOLE CUP 58MM, CAT # 502-03-58F, LOT # MKHVW8; DESCRIPTION: TRIDENT 0 DEG INSERT 40MM, CAT # 623-00-40F, LOT # MKEYLR; DESCRIPTION: NECK TRIAL ADAPTER, CAT # 1601-2000, LOT # MHPN4Y. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S EXPERIENCE. A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-08
2249697-2012-01427 | 2012-08-31 | 2010-11-15 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE RIGHT HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT COMPLAINS OF TENDERNESS AROUND THE HIP AREA. THE PATIENT REPORTS DISCOMFORT UPON WALKING. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-08
2249697-2012-01478 | 2012-09-05 | 2012-08-09 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | A REVIEW OF REJUVENATE STEM WAS REPORTED. | THE REASON FOR THE REVISION WAS NOT PROVIDED. THE CATALOG NUMBERS AND LOT CODES OF OTHER DEVICES LISTED IN THIS REPORT: CAT # NLV-340800Y, LOT # 31618301, DESCRIPTION: LRG TAP PRI MOD NECK 8DEG 34MM, CAT # 6260-5-128, LOT # 33490103, DESCRIPTION: 28MM STD V40 TAPER VIT HEAD, CAT # 1235-2-850, LOT # G2139065, DESCRIPTION: RESTORATION (TM) ADM. INSERT, CAT # 6570-0-136, LOT # 36591901, DESCRIPTION: DELTA V-40 CERAMIC HEAD 36/0. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S EXPERIENCE. ADD'L INFO HAS BEEN REQUESTED AND IF REC'D, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-09
2249697-2012-01451 | 2012-09-05 | 2011-03-07 | I, n, j, u, r, y | Manufacturer report | UNKNOWN_CORK_PRODUCT | UNK HIP IMPLANT | STRYKER ORTHOPAEDICS-MAHWAH | NA | UNK_SHC |  | IT WAS REPORTED THAT THE PT COMPLAINS OF PAIN, SWELLING. PT IS UNABLE TO BEAR WEIGHT ON RIGHT HIP. PT COMPLAINS OF LOSS OF BALANCE. HER LEG "GIVES OUT." PT IS UNSURE AS TO WHAT STRYKER HIP SHE HAS AND SHE WILL REQUEST HER IMPLANT SHEET. | THE PT IS (B)(6) WITH A BMI OF (B)(6), WHICH IS OBESE. AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED. THE EVENT COULD NOT BE CONFIRMED, AS NO ITEMS ASSOCIATED WITH THE EVENT WERE RETURNED OR MADE AVAILABLE FOR IDENTIFICATION OR EVAL. | Y | N | 0 |  | 2012-09
2249697-2012-01444 | 2012-09-05 | 2010-05-04 | I, n, j, u, r, y | Manufacturer report | REJUVENATE TOTAL HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | NOTICE WAS REC'D ON (B)(6) 2012, FROM (B)(6) ALLEGING THAT PT HAD SUSTAINED INJURIES ON (B)(6) 2010, FROM THE USE OF THE STRYKER REJUVENATE TOTAL HIP SYSTEM. | THE INFO IN THIS REPORT WAS PROVIDED BY STRYKER ORTHOPAEDICS LEGAL AFFAIRS DEPARTMENT. NO ADD'L INFO IS AVAILABLE AT THIS TIME. THE DEVICE IS NOT AVAILABLE DUE TO THE ONGOING LITIGATION. SHOULD DEVICE OR ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01443 | 2012-09-05 | 2010-02-17 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT HIP REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: THE PT COMPLAINS OF PAIN IN HIS RIGHT HIP THAT STARTED APPROX 8 WEEKS AGO AND IS GETTING WORSE. HE CLAIMS IT FEELS LIKE GLASS AND HE EXPERIENCES SPASMS. THE PT REPORTS HIS SURGEON TOLD HIM THAT HE WILL BE TAKING IT OUT. THE PT ALSO MENTIONED THAT A REJUVENATE WAS IMPLANTED IN THE LEFT HIP, BUT WAS ASYMPTOMATIC. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE(S) REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01447 | 2012-09-05 | 2011-07-14 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWH | NA | UNK |  | IT WAS REPORTED BY THE PT'S HUSBAND THAT THE PT WAS EXPERIENCING STIFFNESS, SWELLING, AND PAIN AND WANTED TO KNOW HOW THIS IS GOING TO BE PAID. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE(S) REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01471 | 2012-09-05 | 2012-08-14 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT THE PT EXPERIENCED PAIN. NO SIGN OF INFECTION. | CATALOG NUMBER AND LOT CODE OF THE OTHER DEVICE LISTED IN THIS REPORT: CAT # NLS-380000B, LOT # 33537202, DESCRIPTION: LRG TAP PRI MOD NECK 0DEG 38MM. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S EXPERIENCE. ADD'L INFO HAS BEEN REQUESTED AND IF REC'D, WILL BE PROVIDED IN THE SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-09
2249697-2012-01439 | 2012-09-05 | 2010-09-08 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE RIGHT HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PT COMPLAINING ABOUT PAIN AND LEG GROIN. SHE HAD X-RAYS AND BLOOD WORK DONE. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE(S) REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
9616680-2012-00712 | 2012-09-05 | 2012-08-13 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT, METALLOSIS AT JUNCTURE OF MODULAR NECK SEGMENT ATTACHED TO STEM. | AN EVAL OF THE REPORTED DEVICES CANNOT BE PERFORMED AS THEY WERE RETAINED BY THE PT AND WERE NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS MFR # 2249697-2012-01431. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01445 | 2012-09-05 | 2012-08-16 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 10 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-100000S |  | IT WAS REPORTED THAT A REJUVENATE STEM WAS REMOVED DUE TO PAIN FELT BY THE PT. NO INFECTION WAS FOUND. STEM, NECK AND HEAD WERE REMOVED - HAD CONSIDERABLE WEAR. CUP AND LINER REMAIN IN THE PT W/O PROBLEM. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0DEG 34MM, CAT# NLS-340000B, LOT# 35367901, 32MM STD LFIT V40 HEAD, CAT# 6260-9-132, LOT# 36587805. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. WHEN COMPLETED, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01470 | 2012-09-05 | 2012-08-20 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS ¿ MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT, POST OPERATIVE DEVELOPMENT OF MECHANICALLY ASSISTED CREVICE CORROSION WITH ADVERSE LOCAL SOFT ISSUE REACTION WITH TISSUE NECROSIS RIGHT HIP. | AN EVAL OF THE REPORTED DEVICES CANNOT BE PERFORMED AS THEY WERE RETAINED BY THE PT AND WERE NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS MR# 9616680-2012-00722. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01479 | 2012-09-05 | 2012-08-20 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 10 | IMPLANT | STRYKER ORTHOPAEDICS ¿ MAHWAH | NA | SPT-100000S |  | IT WAS REPORTED THAT, DR. (B)(6) REVISED PT'S LEFT HIP DUE TO ADVERSE LOCAL TISSUE REACTION. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR DUE TO HOSPITAL POLICY. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS MR# 9616680-2012-00725. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01481 | 2012-09-05 | 2012-08-20 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS ¿ MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT, ADVERSE LOCAL TISSUE REACTION RESULTING IN REVISION SURGERY. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE KEPT BY THE HOSPITAL AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS MR# 9616680-2012-00726. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01482 | 2012-09-05 | 2012-08-20 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS ¿ MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT A REVISION HIP SURGERY OF A REJUVENATE HIP WAS DONE ON THE PT DUE TO PAIN AND A PSEUDO TUMOR. THE SURGEON TOOK OUT THE REJUVENATE HIP AND IMPLANTED A RESTORATION MODULAR HIP. DR. (B)(6). NOTED OBSERVATION OF THICKENED NECROTIC SOFT TISSUE. | ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. WHEN COMPLETED, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS MFR# 9616680-2012-00727. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01431 | 2012-09-05 | 2012-08-13 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, METALLOSIS AT JUNCTURE OF MODULAR NECK SEGMENT ATTACHED TO STEM. | AN EVAL OF THE REPORTED DEVICES CANNOT BE PERFORMED AS THEY WERE RETAINED BY THE PT AND WERE NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS MFR # 9616680-2012-00712. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01442 | 2012-09-05 | 2010-11-17 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE LEFT HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: (B)(6) DAUGHTER OF THE PT, CALLED ON HER MOTHER'S BEHALF. SHE REPORTS PAIN. SHE FURTHER REPORTS THAT SHE HAD AN INFECTION AND WAS IMPLANTED ON (B)(6). CURRENTLY, SHE DOES NOT HAVE ANYTHING IN PLACE OF THE EXPLANTED PARTS. SHE IS WAITING FOR THE INFECTION TO BE COMPLETELY REMOVED BEFORE IMPLANTING AGAIN IN ABOUT A MONTH. THE PT HAS ALSO BEEN DR. (B)(6), DR.(B)(6), DR. (B)(6), AND DR. (B)(6) AT (B)(6) REHAB. SHE IS NOW IN (B)(6). SHE HAS ALSO BEEN TO (B)(6) HOSPITAL. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01430 | 2012-09-05 | 2012-07-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE FEMORAL IMPLANT | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | . | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01448 | 2012-09-05 | 2011-07-14 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED BY THE PT'S HUSBAND THAT THE PT WAS EXPERIENCING STIFFNESS, SWELLING, AND PAIN. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE(S) WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01449 | 2012-09-05 | 2012-03-07 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT HAS HAD PAIN AND SWELLING SINCE THE SURGERY. SHE WILL F/U WITH HER SURGEON. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE(S) REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01494 | 2012-09-06 | 2011-06-27 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT COMPLAINS OF PAIN IN HIP AND ALONG SCAR AREA. PT HAS HAD DIFFICULTY WALKING AND SITTING SINCE THE SURGERY. | ADD'L INFO HAS BEEN REQUESTED AND IF REC'D, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-09
2249697-2012-01495 | 2012-09-06 | 2011-04-01 | I, n, j, u, r, y | Manufacturer report | UNK LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT COMPLAINS OF PAIN AND LOSS OF FLEXIBILITY. PT CAN'T SLEEP. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-09
2249697-2012-01490 | 2012-09-06 | 2011-11-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE LEFT HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT COMPLAINS HE IS A LITTLE STIFF AND SORE, BUT EXPECTS SORENESS. SAYS HE IS STILL YOUNG AND ACTIVE, PLAYS GOLF. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | 0 |  | 2012-09
2249697-2012-01489 | 2012-09-06 | 2011-12-05 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT EXPERIENCED PAIN, STIFFNESS, AND SWELLING. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-09
2249697-2012-01485 | 2012-09-06 | 2011-04-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE RIGHT HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT COMPLAINS OF PAIN IN GROIN, THIGH, AND KNEE THAT IS SOMETIMES SEVERE. | ADD'L INFO HAS BEEN REQUESTED AND IF REC'D, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-09
2249697-2012-01492 | 2012-09-06 | 2011-10-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE RIGHT HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT HAS NO MOBILITY AND PAIN ON INSIDE UPPER HIP AND OUTER HIP. DOCTOR HAS DETERMINED HE NEEDS A REVISION SURGERY. | ADD'L INFO HAS BEEN REQUESTED AND IF REC'D, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-09
2249697-2012-01503 | 2012-09-07 | 2010-12-09 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT THE PT COMPLAINS OF PAIN IN HIS HIP AND GROIN AREA. | CATALOG NUMBERS AND LOT CODES OF OTHER DEVICES LISTED IN THIS REPORT: DESCRIPTION: REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8: CAT#NLS-300000B, LOT# UNK. DESCRIPTION: TRIDENT 0 X3 INSERT 36MM ID: CAT# 623-00-36F, LOT# UNK. DESCRIPTION: DELTA V-40 CERAMIC HEAD 36/0: CAT # 6570-0-136, LOT# UNK. DESCRIPTION: TRIDENT PSL HA CLUSTER 56MM: CAT# 542-11-56F, LOT# UNK. DESCRIPTION: ACETABULAR DOME HOLE PLUG: CAT# 2060-0000-1, LOT# UNK. DESCRIPTION: CANCELLOUS BONE SCREW 25MM: CAT# 2025-6530-1, LOT# UNK. DESCRIPTION: CANCELLOUS BONE SCREW 30MM: CAT# 2030-6530-1, LOT# UNK. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S EXPERIENCE. A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-09
2249697-2012-01506 | 2012-09-07 | 2012-08-21 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT THE SURGEON DID A REVISION HIP SURGERY ON THE PT'S RIGHT HIP AND REMOVED A STRYKER REJUVENATE HIP STEM AND NECK DUE TO A PSEUDO TUMOR AND PAIN. UPON OPENING OF PSEUDO TUMOR CAPSULE, 40 CC'S OF BROWN FLUID CAME OUT OF THE CELL. | ADDITIONAL INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. WHEN COMPLETED, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS MFR # 9616680-2012-00731. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
9616680-2012-00730 | 2012-09-07 | 2012-08-20 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT METALLOSIS AT JUNCTURE OF MODULAR NECK SEGMENT ATTACHED TO STEM. | AN EVALUATION OF THE REPORTED DEVICES CANNOT BE PERFORMED AS THEY WERE RETAINED BY THE PT AND WERE NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS: MFR # 2249697-2012-01502. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01502 | 2012-09-07 | 2012-08-20 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT METALOSIS AT JUNCTURE OF MODULAR NECK SEGMENT ATTACHED TO STEM. | AN EVALUATION OF THE REPORTED DEVICES CANNOT BE PERFORMED AS THEY WERE RETAINED BY THE PT AND WERE NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS MFR # 9616680-2012-00730. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01500 | 2012-09-07 | 2011-10-04 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT HIP REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT COMPLAINS OF SORENESS. | THE PT IS (B)(6), WHICH IS OBESE. ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-09
2249697-2012-01508 | 2012-09-10 | 2011-12-08 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE RIGHT HIP | IMPLANT | STRYKER ORTHOPAEDICS | NA | UNK |  | IT WAS REPORTED THAT THE PT COMPLAINS OF PAIN. PT IS SCHEDULED FOR AN UPCOMING REVISION. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-09
2249697-2012-01509 | 2012-09-10 | 2012-08-14 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 10 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-34000B |  | AN INFECTION WAS REPORTED. | CATALOG NUMBERS AND LOT CODES OF OTHER DEVICES LISTED IN THIS REPORT: CAT # NLS-340000B, LOT # 35170502, DESCRIPTION: LRG TAP PRI MOD NCK 0DEG 34MM. CAT# 509-0254E, LOT # MJTEWH, DESCRIPTION: TRITANIUM REVISION ACETABULAR. CAT # 2080-0025, LOT # MJRNKN, DESCRIPTION: GAP PLATE SCREWS. CAT # 623-00-40E, LOT # MKD62A, DESCRIPTION: TRIDENT 0 DEG INSERT 40 MM. CAT # 6519-T-100, LOT # 36356902, DESCRIPTION: UNI ADAPTOR SLEEVE V40 TI. CAT # 6519-1-040, LOT # 36359503, DESCRIPTION: DELTA UN. FEM HD 40MM. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S EXPERIENCE. THE ROOT CAUSE OF THE EVENT WAS NOT DETERMINED. A REVIEW OF THE DHR INDICATES THAT THE REPORTED DEVICES WERE MANUFACTURED AND ACCEPTED INTO FINAL STOCK WITH NO REPORTED DISCREPANCIES RELATED TO THE EVENT. THE DEVICES WERE WITHIN STERILIZATION SPECIFICATION. THE COMPLAINT REVIEW INDICATED THAT THERE WERE NO SIMILAR EVENTS FOR THE REPORTED LOTS. | Y | N | OTHER |  | 2012-09
2249697-2012-01517 | 2012-09-10 | 2011-02-02 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 10 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-1000000S |  | IT WAS REPORTED THAT THE PT HAS PAIN AND SWELLING. PT HAD BILATERAL HIP DONE AND IS NOT SURE OF THE IMPLANTATION DATE. | THE PT IS (B)(6). CATALOG NUMBERS AND LOT CODES OF OTHER DEVICES LISTED IN THIS REPORT: DESCRIPTION: LRG TAP PRI MOD NCK 0DEG 334MM: CAT # NLS-340000B, LOT # 34887202. DESCRIPTION: DELTA V-40 CERAMIC HEAD 30/0: CAT # 6570-0-136, LOT # 35424102. DESCRIPTION: PRIMARY TRITANIUM HEMI CLUSTER HOLE CUP 58MM: CAT # 502-03-58F, LOT # MJND46. DESCRIPTION: TRIDENT 0 DEG X3 INSERT 36MM ID: CAT # 623-00-36F, LOT # MJN84P. DESCRIPTION: 6.5 CANCELLOUS BONE SCREW 35MM: CAT # 2030-6535-1, LOT # MJPNJE. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S EXPERIENCED. ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-09
2820293 | 2012-09-10 | 2012-09-05 | N, o, , a, n, s, w, e, r, , p, r, o, v, i, d, e, d | User Facility report | REJUVENATE SPT | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER, CE | STRYKER ORTHOPAEDICS | SPT-080000S | * |  |  |  | N | * | INVALID DATA |  | 2012-09
2249697-2012-01542 | 2012-09-11 | 2012-08-23 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT PT EXPERIENCING HIP PAIN. | THE FOLLOWING DEVICES WERE ALSO LISTED IN THE REPORT: LRG TAP PRI MOD NCK 0DEG 30 MM, CAT# NLS-30000B, LOT# 38653302; DELTA V-40 CERAMIC HEAD 32/+4, CAT# 6570-0-232, LOT# 722227. IT CANNOT BE DETERMINED WHICH, IF ANY, OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED, AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER DUE TO HOSPITAL POLICY. ADDITIONAL INFO HAS BEEN REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01534 | 2012-09-11 | 2012-08-22 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT PT PRESENTED WITH HIP PAIN. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0DEG 30MM: CAT# NLS-30000B, LOT# 32794001; TRIDENT 0 DEG INSERT 40MM: CAT# 623-00-40H, LOT# MHJ41X; V40 COCR LFIT HEAD 40MM/-4: CAT# 6260-9-040, LOT# MJM5R7. IT CANNOT BE DETERMINED WHICH, IF ANY, OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED, AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER DUE TO HOSPITAL POLICY. ADDITIONAL INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WILL BE REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01543 | 2012-09-11 | 2011-10-10 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE MODULAR HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PT COMPLAINS OF PAIN AND SWELLING. PT WILL BE FOLLOWING UP WITH HIS SURGEON. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01516 | 2012-09-11 | 2012-08-21 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT, INFLAMMATION AND PAIN IN HIP. NECROSIS FOUND IN HIP BIOPSY. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER DUE TO HOSPITAL POLICY. HOWEVER, SOME X-RAYS AND PHOTOS HAVE BEEN RECEIVED. WHEN COMPLETED, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS MFR # 9616680-2012-00737. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01522 | 2012-09-11 | 2012-08-21 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT UNEXPLAINED PAIN, LEFT HIP. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0DEG 30MM, CAT# NLS-300000B, LOT# 32356201; DELTA V-40 CERAMIC HEAD 36/+2.5, CAT# 6570-0-536, LOT# 32905801. IT CANNOT BE DETERMINED WHICH, IF ANY, OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVALUATION OF THE REPORTED DEVICES CANNOT BE PERFORMED AS THE DEVICES WERE KEPT BY THE HOSPITAL AND WERE NOT RETURNED TO THE MANUFACTURER. HOWEVER, SOME X-RAYS AND MEDICAL RECORDS HAVE BEEN RECEIVED. WHEN COMPLETED, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01523 | 2012-09-11 | 2011-10-10 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 10 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-100000S |  | IT WAS REPORTED THAT: SAW PT IN (B)(6) 2012, AT 3.5 MONTHS POST-OP. WAS FEELING PRETTY WELL, BUT SINCE THAT TIME, HE HAS DEVELOPED PAIN FROM THE LATERAL SIDE. COBALT-CHROMIUM LEVELS HAVE BEEN ORDERED AS WELL AS A MASS MRI. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0DEG 34MM: CAT# NLS-340000B, LOT# 37633601; PRIMARY TRITANIUM HEMI CLUSTER HOLE CUP 60MM: CAT# 502-03-60F, LOT# MJT763; TRIDENT 0DEG X3 INSERT 36MM ID: CAT# 623-00-36F, LOT# MKK3V9; 6.5 CANCELLOUS BONE SCREW 30 MM: CAT# 2030-6530-1, LOT# MKLV92, MKMYPY; DELTA V-40 CERAMIC HEAD 36+0: CAT# 6570-0-136, LOT# 7530303. IT CANNOT BE DETERMINED WHICH, IF ANY, OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVALUATION OF THE DEVICES CANNOT BE PERFORMED AS THE DEVICES REMAINED IMPLANTED IN THE PT AND WERE NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01526 | 2012-09-11 | 2012-08-21 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT, PT HIP REVISED DUE TO PAIN. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0DEG 34MM, CAT # NLS-340000B, LOT # 35231001. DELTA V-40 CERAMIC HEAD 36/0, CAT # 6570-0-136, LOT # 37298504. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01511 | 2012-09-11 | 2012-08-21 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT, PT COMPLAINED OF PAIN. SURGEON DID BONE SCAN AND THE FEMUR LITE UP AND SCHEDULED FEMORAL REVISION. FEMUR WAS REMOVED WITHOUT TOO MUCH DIFFICULTY AND A RESTORATION MODULAR STEM WAS UTILIZED. THE REJUVENATE FEMORAL COMPONENT WAS RETRIEVED AS WELL AS THE NECK AND HEAD. | ADDITIONAL INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. WHEN COMPLETED, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS MFR # 9616680-2012-00734. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01529 | 2012-09-11 | 2012-03-29 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT HAD PAIN BETWEEN HIP AND KNEE AND IN THE JOINT. THE PATIENT MENTIONED HIS PRIMARY SURGERY WAS IN (B)(6) 2005 AND WAS NOT A STRYKER IMPLANT. REVISION USING STRYKER IMPLANTS ON (B)(6) 2012. AFTER SURGERY, THE PATIENT CLAIMS IT TOOK AWHILE TO SLEEP ON HIS LEFT. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01530 | 2012-09-11 | 2012-04-09 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT HAD PAIN. HE CAN'T WALK MORE THAN HALF A BLOCK. THERE IS A CLICKING NOISE. MRI SHOWED FLUID AROUND THE JOINT. A BLOOD CLOT DEVELOPED AND HAS REACHED THE LUNGS; PATIENT IS ON BLOOD THINNERS. DR. WANTS TO HAVE A REVISION SURGERY. PATIENT WANTS TO KNOW, WHO WILL PAY FOR THE REVISION. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01531 | 2012-09-11 | 2011-08-19 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE MODULAR STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT COMPLAINS OF PAIN, NOT ABLE TO WALK, AND INABILITY TO BEAR WEIGHT ON HIP. PATIENT WILL FOLLOW UP WITH A NEW SURGEON IN (B)(6). | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01532 | 2012-09-11 | 2011-04-25 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PATIENT HAS LACK OF MOTION, WHICH STARTED SHORTLY AFTER SURGERY. PATIENT HAD PAIN, NAUSEA, AND A PSEUDOTUMOR IN THE HIP. PATIENT COMPLAINS OF SWELLING IN THE RIGHT LEG, ANKLE, FOOT, AND MOVING TO HANDS. DR. DIAGNOSED HER WITH ALTR. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01521 | 2012-09-11 | 2012-08-21 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 10 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-100000S |  | IT WAS REPORTED THAT PT EXPERIENCING PAIN IN LEFT HIP. THE PT HAS ELEVATED COBALT LEVEL. A REJUVENATE STEM WAS REMOVED AND REPLACED WITH RESTORATION HA STEM, 56 ADM POLY LINER, 28 CERAMIC HEAD. SURGEON SAID IT DID NOT APPEAR LIKE PAST CORROSIVE CASES. | AN EVALUATION OF THE REPORTED DEVICES CANNOT BE PERFORMED, AS THEY WERE RETAINED BY THE PT AND WERE NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS MFR # 9616680-2012-00739. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01547 | 2012-09-13 | 2010-08-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT THE PT COMPLAINS OF PAIN IN THE RIGHT HIP AND CAN'T WALK. | CATALOG NUMBERS AND LOT CODES OF OTHER DEVICES LISTED IN THIS REPORT: UNK, 30MM 132 DEG REJUVENATE NECK. UNK, 52MM TRITANIUM ACETABULUM. UNK, 10 DEG X3 LINER. UNK, 32MM +0 HEAD. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S EXPERIENCE. A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | N | N | OTHER |  | 2012-09
2249697-2012-01548 | 2012-09-13 | 2012-12-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT THE PT COMPLAINS OF PAIN IN LEFT HIP AND CAN'T WALK. | CATALOG NUMBERS AND LOT CODES OF OTHER DEVICES LISTED IN THIS REPORT: UNK, 30MM 132DEG REJUVENATE NECK. UNK 54MM TRITANIUM ACETABULAR. UNK, 32MM-4 HEAD. UNK, 32MM 10DEG X3 LINER. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S EXPERIENCED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-09
2249697-2012-01549 | 2012-09-13 | 2012-06-25 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | HOWMEDICA OSTEONICS CORP | SRT08000S | SPT-08000S |  | USER FACILITY MEDWATCH #(B)(4) REPORTED THAT, (B)(6) HAD LTHR REVISED (B)(6) 2012 FOR METAL REACTION S/P BTHR (B)(6) 2010 FOR OA (STRYKER REJUVENATE IMPLANTS). RIGHT HIP WILL ALSO BE REVISED (ALSO REJUVENATE MODULAR METAL COMPONENTS) AFTER 3-MONTH RECOVERY. PT COMPLAINED TO SURGEON OF PAIN AND FEVERS IN (B)(6) 2011; WAS TO FOLLOW UP WITH INTERNIST FOR FEVER. RETURNED (B)(6) 2012 TO SURGEON STILL WITH PAIN; FEVERS RESOLVED. (B)(6) 2012 SERUM COBALT ELEVATED AT 22.B UG/L (WNL >1.0). JULY 5, 2012 MRI: DEHISCENCE OF LATERAL ATTACHMENT OF POSTERIOR CAPSULAR REPAIR BILATERALLY; ON RIGHT HIP RESULTS IN EXTENSIVE DEBRIS IN GREATER TROCHANTERIC BURSA. ALSO ON RIGHT ARE LOW SIGNAL INTENSITY DEPOSITS, STRONGLY SUSPICIOUS FOR METALLIC DEPOSITION IN SURROUNDING SOFT TISSUE ENVELOPE. ADD¿L INFO FROM USER FACILITY REPORT: A (B)(6) FEMALE TO HAVE RTHR REVISED FOR METAL REACTION S/P BTHR (B)(6) 2010 FOR OA (STRYKER REJUVENATE IMPLANTS). LEFT. HIP REVISED (B)(6) 2012 (ALSO REJUVENATE MODULAR METAL COMPONENTS). AWAITING 3-MONTH RECOVERY S/P RIGHT REVISION. PT COMPLAINED TO SURGEON OF PAIN AND FEVERS IN (B)(6) 2011; WAS TO FOLLOW UP WITH INTERNIST FOR FEVER IN (B)(6) 2011; WAS TO FOLLOW UP WITH INTERNIST FOR FEVER. RETURNED (B)(6) 2012 TO SURGEON STILL WITH PAIN; FEVERS RESOLVED. (B)(6) 2012, SERUM COBALT ELEVATED AT 22.8 UG/L (WNL >1.0). (B)(6) 2012 MRI: DEHISCENCE OF LATERAL ATTACHMENT OF POSTERIOR CAPSULAR REPAIR BILATERALLY; ON RIGHT HIP RESULTS IN EXTENSIVE DEBRIS IN GREATER TROCHANTERIC BURSA. ALSO ON RIGHT ARE LOW SIGNAL INTENSITY DEPOSITS, STRONGLY SUSPICIOUS FOR METALLIC DEPOSITION IN SURROUNDING SOFT TISSUE ENVELOPE. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD¿L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD¿L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS MFR # 9616680-2012-00752. (B)(4). ADD'L INFO FROM USER FACILITY REPORT: REJUVENTATE. MODULAR STEM. # TMZF SIZE 8. | Y | Y | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01557 | 2012-09-14 | 2012-08-27 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT PATIENT COMPLAINED OF HIP PAIN AND AS SOON AS HE AND HIS LAWYERS SAW THE RECALL COMMERCIALS AND HE WANTED IT OUT. PULLED THE HEAD AND NECK AND SOME FRETTING AND METALLOSIS WAS OBSERVED UPON EXTRACTING THE NECK. ALSO FOUND SOME PUS THAT LEAD DR. (B)(6) TO BELIEVE HE MIGHT HAVE SOME INFECTION. LABS CAME BACK NEGATIVE SO A RESTORATION MODULAR WAS IMPLANTED. | ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. WHEN COMPLETED, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PATIENT/EVENT AS MFR # 9616680-2012-00769. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01556 | 2012-09-14 | 2012-08-27 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT A REJUVENATE STEM AND NECK ALONG WITH A BIOLOX HEAD WERE EXPLANTED DUE TO ADVERSE LOCAL TISSUE REACTION. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER DUE TO HOSPITAL POLICY. ADDITIONAL INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS MFR # 9616680-2012-00768. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01552 | 2012-09-14 | 2012-04-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PT COMPLAINS OF SORENESS AND TIGHTNESS IN HIP AREA. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01555 | 2012-09-14 | 2012-07-27 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PT HAS EXPERIENCED CONSTANT PAIN FOR AT LEAST A MONTH, ACHING, AND DIFFICULTY SLEEPING. WANTS TO BE REIMBURSED FOR TESTS. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORD(S) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01550 | 2012-09-14 | 2010-12-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT THE PATIENT IS EXPERIENCING PAIN. PATIENT IS CURRENTLY FOLLOWING UP WITH HIS SURGEON. MRI SHOWED A "LITTLE" FLUID BUILD UP IN JOINT. SURGEON WILL FOLLOW UP WITH PATIENT IN 6 MONTHS. | CATALOG NUMBERS AND LOT CODES OF OTHER DEVICES LISTED IN THIS REPORT: CAT # NLS-380000B, LOT # 32143201, DESCRIPTION: REJUVENATE MODULAR NECK - 127/132 NECK ANGLE 38MM V40. CAT # 6570-0-128, LOT # 24217701, DESCRIPTION: DELTA V-40 CERAMIC HEAD 28/0. CAT # 1235-2-522, LOT # G2947455, DESCRIPTION: RESTORATION ADM. CUP W/HA CAT # 1236-2-852, LOT # 34788001, DESCRIPTION: RESTORATION ADM X3 INS 28/52. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S EXPERIENCE. A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-09
2249697-2012-01553 | 2012-09-14 | 2011-02-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT HIP REJUVENATE MODULAR STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PT COMPLAINS OF PAIN AND "POPPING AND CLICKING" NOISE. PT CALLING TO FIND OUT IF HIS HIP PARTS HAVE BEEN PART OF A RECALL. THEY HAVE CALLED THEIR DOCTOR'S OFFICE AND THEY WERE GIVEN THE NAME OF HIS IMPLANT LISTED. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
9616680-2012-00762 | 2012-09-14 | 2009-10-05 | I, n, j, u, r, y | Manufacturer report | TRIDENT ALUMINA INSERT | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/CERAMIC/METAL, CEMENTED OR UNCE | STRYKER ORTHOPAEDICS-MAHWAH | NA | 625-0T-32F |  | IT WAS REPORTED THROUGH THE ATTORNEY FOR THE PATIENT, AS A RESULT OF A LEGAL CLAIM, THAT ALLEGEDLY THE PATIENT RECEIVED A TRIDENT ACETABULAR HIP STEM. IT WAS FURTHER ALLEGED THAT THE PATIENT "BEGAN EXPERIENCING AUDIBLE SOUNDS DURING MOTION EMANATING FROM LOCATION OF THE DEVICE." | THE INFORMATION IN THIS REPORT WAS PROVIDED BY STRYKER ORTHOPAEDICS LEGAL AFFAIRS. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. THEREFORE, THE EXACT CAUSE OF THE REPORTED EVENT CANNOT BE DETERMINED. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
MW5027002 | 2012-09-17 | 2011-03-08 | I, n, j, u, r, y | Voluntary report | ABG II MODULAR V 40 SHORT NECK | HIP PROSTHESIS | STRYKER ORTHOPAEDICS |  | 4845-4-411 |  | COBALT POISONING DUE TO FRETTING FROM STRYKER ABGII HIP PROSTHESIS IMPLANTED (B)(6) 2011 - RIGHT LEG. TOOK 15 MONTHS TO DIAGNOSE. (B)(6) 2011 - HIP SWELLED UP AND WAS VERY PAINFUL. DIFFICULTY WALKING, COULD NOT HIKE. (B)(6) 2012 - BAD PAIN IN LOWER RIGHT LEG (NIGHTS ONLY) THAT WOKE ME UP AT 3 AM. COULD NOT GO BACK TO SLEEP. PAINFUL AND DANGEROUSLY TIRED. (B)(6) 2012, CANNOT LIFT RIGHT LEG WHILE SITTING OR STRAIGHT UP WHILE LYING DOWN. HARD TO TIE SHOE, ENTER CARS, CANNOT RIDE A BIKE. (B)(6) 2012, SLOW MOVING, TIRED, LOST SENSE OF BALANCE. (B)(6) 2012, WHEN SNEEZING OR COUGHING, INTENSE PAIN IN GROIN. (B)(6) 2012 - CONTINUOUS ALLERGIC HACKING COUGH. COUGH STOPPED 2 DAYS AFTER SURGERY. (B)(6) 2012, HIP REVISION SURGERY TO REMOVE PROSTHESIS CAUSING COBALT POISONING. TAN PUS-LIKE DEBRIS WAS FOUND ALL OVER THE AREA AND SCARRING ON MUSCLES SURROUNDING PROSTHESIS. |  | Y | Y | LAY USER/PATIENT |  | 2012-09
2249697-2012-01560 | 2012-09-18 | 2011-08-22 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PT HAS DISCOMFORT, SHARP PAIN, AND CAN BARELY WALK. PT WILL BE GOING FOR BLOODWORK AND POSSIBLE MRI. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01561 | 2012-09-18 | 2011-12-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE MODULAR HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PT SCHEDULED FOR REVISION SURGERY ON (B)(6) 2012. PT HAD MRI AND BLOOD WORK TAKEN. HIS COBALT LEVELS ARE NEARLY SIX TIMES THE NORMAL. HE HAS ELEVATED CHROMIUM LEVELS. MRI SHOWED FLUID ACCUMULATION IN JOINT, POSSIBLE PSEUDOTUMOR AND OTHER FINDINGS. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01562 | 2012-09-18 | 2011-02-17 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, WIFE AND PT CALLED TOGETHER. PT HAS BEEN HAVING PAIN ON AND OFF ON THE SIDE AND BACK OF HIS HIP. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01564 | 2012-09-18 | 2010-10-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE MODULAR HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PT COMPLAINS OF SORENESS IN HIP AREA AND BACK. BLOOD TESTS SHOWED COBALT LEVELS WERE HIGH. CHROMIUM LEVELS ARE NORMAL. MRI IS SCHEDULED AND PT WILL HAVE FLUID ASPIRATED FROM JOINT. SHE WILL FOLLOW UP WITH SURGEON. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01567 | 2012-09-18 | 2011-10-03 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PT CAN'T SLEEP ON EITHER SIDE; PT CAN'T TOUCH AREA WHERE BURSA SACS USED TO BE. SEVERE GROIN PAIN SINCE JUNE: CAN'T CLIMB STAIR. FOUND INFLAMMATION ON MRI AND CYST IS GROWING. BLOOD WORK INDICATES INFLAMMATION DOCTOR IS RUNNING TESTS. COBALT LEVELS ELEVATED. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01568 | 2012-09-18 | 2012-04-23 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT HIP REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, NO INFECTION FROM SURGERY. LEFT HIP FEELS DIFFERENT FROM RIGHT HIP, DOESN'T FEEL RIGHT AND DOESN'T MOVE LIKE RIGHT HIP. HAS DIFFERENT SENSATIONS. STILL SOME SWELLING AT SITE. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01583 | 2012-09-18 | 2011-04-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT HAS DEEP THIGH, GROIN, AND HIP PAIN. PT CAN ONLY, BARELY RAISE LEFT LEG ENOUGH TO GO UP STAIRS. REVISION PLANNED FOR (B)(6) 2012. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-09
2249697-2012-01566 | 2012-09-18 | 2011-08-03 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PT COMPLAINS OF GROIN PAIN AND INABILITY TO SLEEP ON EITHER SIDE. PT CANNOT TOUCH WHERE THE BURSA SACS WERE. BLOOD WORK INDICATES POSSIBLE INFLAMMATION; DOCTOR IS RUNNING TESTS, COBALT LEVELS ELEVATED. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01587 | 2012-09-18 | 2012-08-29 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 10 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-100000S |  | IT WAS REPORTED THAT, DOCTOR REVISED THE PT'S RIGHT HIP DUE TO ADVERSE LOCAL TISSUE REACTION. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR DUE TO HOSPITAL POLICY. ADDITIONAL INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED, BUT HAS NOT BEEN MADE AVAILABLE. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS MFR # 9616680-2012-00774. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01581 | 2012-09-18 | 2012-03-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT HAD A REVISION DUE TO THE STEM BREAKING. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01580 | 2012-09-18 | 2012-02-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT A REVISION WAS PERFORMED DUE TO IMPLANT BREAKING OFF AT STEM. PELVIS WAS PUNCTURED. WAS IN WHEELCHAIR FOR 5 MONTHS. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | Y | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01582 | 2012-09-18 | 2012-04-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT HAD A THIRD REVISION. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01585 | 2012-09-18 | 2012-05-10 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT IS IN PAIN AND HAS NUMBNESS IN THE UPPER PART OF HIP. IT IS HARD TO SLEEP AND WALK. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01584 | 2012-09-18 | 2012-05-10 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT HAD A FOURTH REVISION. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01586 | 2012-09-18 | 2010-05-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PATIENT HAS PAIN AND IT IS HARD TO WALK AND SLEEP. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01570 | 2012-09-18 | 2012-09-04 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 10 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT - 100000S |  | IT WAS REPORTED THAT THE PATIENT PRESENTED WITH PAIN AND OSTEOLYSIS. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PATIENT/EVENT AS MFR # 9616680-2012-00770. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01589 | 2012-09-18 | 2012-08-29 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT A REJUVENATE STEM, NECK, AND BIOLOX HEAD WERE EXPLANTED DUE TO ADVERSE LOCAL TISSUE REACTION. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PATIENT/EVENT AS MFR # 9616680-2012-00775. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01598 | 2012-09-19 | 2007-09-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE/ABG II HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PT IS EXPERIENCING SORENESS IN HIP. ULTRASOUND AND OTHER TESTS HAVE SHOWED CORROSION. PT IS NOT A CANDIDATE FOR REPLACEMENT OF HIP AND WILL REQUIRE OTHER TREATMENTS. SHE WILL UNDERGO ASPIRATION OF HIP TO REMOVE FLUID. SHE HAS HAD BLOOD WORK FOR CHROMIUM AND COBALT LEVELS. | AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDE DIN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01599 | 2012-09-19 | 2010-10-20 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE/ABG II HIP | IMPLANT | STRYKER ORHTOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PT COMPLAINS OF PAIN AND LOSS OF MUSCLE TONE. HE THINKS HE MAY NEED A REVISION. PT INQUIRED ABOUT HOW STRYKER WILL HELP HIM AS HE NEEDS TO HAVE THE MRI AND BLOOD TEST. PT WAS EXPLAINED THE PROCESS, AND HE DID NOT WANT TO PROVIDE ANY ADDITIONAL INFORMATION TO PROCEED WITH THE INVESTIGATION AT THIS TIME. HE STATED THE HE WILL CONTACT HIS SURGEON FIRST. | AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01597 | 2012-09-19 | 2012-08-29 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT THE PT WAS HAVING PAIN IN THEIR LEFT HIP. THE PT'S LEFT HIP WAS REVISED ON (B)(6) 2012 BY SURGEON. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0DEG 30MM, CAT# NLS-300000B, LOT# 36968302. TRIDENT 0DEG X3 INSERT 36MM ID, CAT# 623-00-36F, LOT# MKPXL2. DELTA V-40 CERAMIC HEAD 36/0, CAT# 6570-0-136, LOT# 38567701. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS RETAINED BY THE HOSPITAL FOR PATHOLOGY AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD DEVICE OR ADDITIONAL INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01607 | 2012-09-20 | 2012-08-30 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE RIGHT HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PT HAS BILATERAL HIPS IMPLANTED BY DR (B)(6). HE HAS STARTED TO EXPERIENCE SOME DISCOMFORT IN HIS RIGHT HIP. HE HAS HAD SOME TESTING DONE AND HIS LAWYER HAS CONTACTED THE STRYKER HOTLINE. | THE REPORTED EVENT BECAME LEGAL AND WILL BE HANDLED BY STRYKER ORTHOPAEDICS LEGAL AFFAIRS DEPARTMENT. NO ADD'L INFO IS AVAILABLE AT THIS TIME DUE TO THE ONGOING LITIGATION. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
9616680-2012-00786 | 2012-09-20 | 2012-03-26 | I, n, j, u, r, y | Manufacturer report | UNKNOWN TRIDENT HEMISPHERICAL HA CLUSTER SHELL | HIP IMPLANT | STRYKER ORTHOPAEDICS-MAHWAH | NA | UNK_JR |  | IT WAS REPORTED THROUGH THE ATTORNEY FOR THE PT, AS A RESULT OF A LEGAL CLAIM, THAT ALLEGEDLY, THE PT RECEIVED A TRIDENT ACETABULAR SYSTEM. IT WAS FURTHER ALLEGED THAT THE PT "BEGAN EXPERIENCING AUDIBLE SOUNDS DURING MOTION EMANATING FROM THE LOCATION OF THE DEVICE." | THE INFO IN THIS REPORT WAS PROVIDED BY STRYKER ORTHOPAEDICS LEGAL AFFAIRS. NO ADD'L INFO IS AVAILABLE AT THIS TIME. THEREFORE, THE EXACT CAUSE OF THE REPORTED EVENT CANNOT BE DETERMINED. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01611 | 2012-09-20 | 2012-08-31 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT THE PT HAD AN INFECTED TOTAL HIP. ALL IMPLANTS REMOVED AND REPLACED WITH AN ANTIBIOTIC SPACER. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP MOD NCK 8DEG 30MM, CAT# NLV-300800G, LOT# 32987601; DELTA V40 HEAD 36MM+0, CAT# UNK, LOT# UNK; TRIDENT 10 DEG X3 INSERT 36MM ID, CAT# 623-10-36F, LOT# MJRNLL; PRIMARY TRITANIUM HEMI CLUSTER HOLE CUP 60MM, CAT# 502-03-60F, LOT# MJPJ96. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S INFECTION. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. WHEN COMPLETED, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01605 | 2012-09-20 | 2011-07-26 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS | NA | UNK |  | IT WAS REPORTED THAT PT MOVEMENT ON RIGHT LEG IS GETTING LESS. PT HAS DEEP THIGH, GROIN, AND HIP PAIN, BUT IT IS LESS THAN LEFT LEG. | THE PT IS (B)(6). IT WAS INDICATED THAT THE DEVICE IS NOT AVAILABLE FOR EVAL, AS IT IS STILL IMPLANTED. ADD¿L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-09
2249697-2012-01604 | 2012-09-20 | 2011-02-09 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT HAS PAIN AND SWELLING IN GROIN. HARD TO WALK WITH THE EXTREME PAIN FROM LOWER BACK TO FOOT. BURNING SENSATION OCCURS WITH JUST SITTING OR TOO MUCH EFFORT. PT ON ANTI-INFLAMMATORY PILLS, AND HAS WENT TO EMERGENCY ROOM AND RECEIVED MORPHINE SHOTS. | THE PT IS (B)(6). ADD¿L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-09
2249697-2012-01610 | 2012-09-20 | 2011-08-11 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PT COMPLAINS OF PAIN IN LEG, GROIN, AND LOWER BACK. PT COMPLAINS OF HAVING HEADACHES AND ITCHING, SWELLING, AND TENDERNESS TO THE TOUCH. PT MENTIONED HAVING A REVISION SURGERY IN THE NEXT WEEKS. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD¿L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD¿L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01608 | 2012-09-20 | 2012-08-30 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT THE PT HAD EXTREME PAIN, SURGEON DID ALLERGY TESTS AND INFECTION TESTS AND ALL TESTS CAME BACK NEGATIVE. THE PT DID HAVE VERY HIGH CHROMIUM LEVELS. SURGEON REMOVED ALL IMPLANTS - REPLACED WITH A SPACER. | ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. WHEN COMPLETED, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS: MFR # 9616680-2012-00791. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01643 | 2012-09-24 | 2009-12-11 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STYRYKER ORTHOPAEDICS | NA | UNK |  | IT WAS REPORTED THAT THE PT RECEIVED A LETTER FROM HER SURGEON ABOUT THE RECALL. PT STATED THAT SHE HAS PAIN WHEN SHE IS WALKING. SHE STATED THAT WHATEVER IS IN HER BLOOD TEST IS HIGH. HER MRI SHOWED CYST AROUND THE IMPLANT. ON (B)(6) 2012 PT HAD ASPIRATION OF HER CYST. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01628 | 2012-09-24 | 2011-04-07 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE MODULAR HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE BLOOD TESTS SHOWED METAL LEVELS ELEVATED. SURGEON ADVISED PT THAT HE NEEDS A REVISION SURGERY. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-09
2249697-2012-01630 | 2012-09-24 | 2011-06-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT STARTED TO HAVE PAIN IN (B)(6) 2011. SHE HAD TWO MRIS THAT SHOWED ACCUMULATION OF FLUID IN THE HIP JOINT. HER CHROMIUM OR COBALT LEVEL (PT NOT SURE WHICH ONE) IS 6.3. THE TP ALSO NOTED THAT HER DOCTOR IS CONSIDERING A REVISION SURGERY. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-09
2249697-2012-01635 | 2012-09-24 | 2012-04-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT STATED THAT HE HAS EXTREME PAIN RADIATING FROM HIS THIGH TO FOOT SINCE (B)(6) 2012. HE HAD FLUID BUILD UP. THE DOCTOR ASPIRATED 15CC ONE TIME FROM HIS HIP. HE HAS LOCAL TISSUE REACTION. HE HAS PAIN IN THE BUTTOCKS AND GROIN AREAS. HE HAS PAIN SITTING DOWN, FEELS LIKE "SITTING ON BUNCH OF ROCKS". HE HAS ELEVATED COBALT LEVEL. DOCTOR TOLD THE PT THAT HE NEEDS A REVISION. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01649 | 2012-09-25 | 2011-10-31 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PT COMPLAINS OF PAIN IN RIGHT HIP. PT IS IN SO MUCH PAIN THAT SHE IS UNABLE TO MOVE LEG FOR 20 MINUTES, UNTIL SHE SLOWLY REGAINS MOVEMENT. PT STATED, SHE HAS NO MORE MEDICAL COVERAGE AND NOW REQUIRES A BONE SCAN AND BLOOD TESTS THAT SHE IS UNABLE TO COVER. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01648 | 2012-09-25 | 2011-02-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PT HAS SOME STIFFNESS AND SORENESS WITH WEATHER CHANGES AND PAIN AT SCAR SITE. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01658 | 2012-09-25 | 2012-09-05 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT PT, THROUGH THE FILING OF A LAWSUIT, IS REPORTING AND ALLEGING THAT A REJUVENATE MODULAR FEMORAL STEM REQUIRED REVISION DUE TO PAIN AND THAT AT REVISION SURGERY ALLEGEDLY THERE WERE FINDINGS OF "THE PRESENCE OF MILKY, TURBID FLUID; LARGE PSEUDOTUMOR FORMATION; SOFT TISSUE NECROSIS; MUSCLE LOSS AND/OR BONY NECROSIS AT THE PROXIMAL FEMUR". (B)(6) 2012 ADD'L INFO: SURGEON REVISED PTS RIGHT HIP DUE TO ADVERSE LOCAL TISSUE REACTION. | THE INFO IN THIS REPORT WAS PROVIDED BY STRYKER ORTHOPAEDICS LEGAL AFFAIRS DEPARTMENT. NO ADD'L INFO IS AVAILABLE AT THIS TIME. THE DEVICE IS NOT AVAILABLE DUE TO THE ONGOING LITIGATION. SHOULD DEVICE OR ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS MFR # 9616680-2012-00806. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01676 | 2012-09-25 | 2012-09-06 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT REJUVENATE STEM, NECK, BIOLOX HEAD, PSL SHELL, LINER, DOME HOLE PLUG, AND SCREW WERE EXPLANTED DUE TO ADVERSE LOCAL TISSUE REACTION. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR DUE TO HOSPITAL POLICY. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED, BUT NOT MADE AVAILABLE. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS MFR # 9616680-2012-00810. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01637 | 2012-09-25 | 2012-07-13 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PT HAD FLUID BUILD UP. PT HAS PAIN, AND COULDN'T WALK BECAUSE OF THE PAIN. DR TOOK SOME TISSUE AND MUSCLE OUT. DR DIAGNOSED PT WITH ALTR. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01633 | 2012-09-25 | 2011-07-19 | I, n, j, u, r, y | Manufacturer report | UNK RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS ¿ MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PT COMPLAINS OF CONSTANT PAIN AND SORENESS. PT HAS DIFFICULTY SLEEPING AND FEELS THAT HER HIP HAS NOT HEALED. PT IS SCHEDULED FOR AN MRI AND BLOOD TESTS. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01657 | 2012-09-25 | 2012-06-14 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PT, THROUGH THE FILING OF A LAWSUIT, IS REPORTING AND ALLEGING THAT A REJUVENATE MODULAR FEMORAL STEM REQUIRED REVISION DUE TO PAIN AND THAT AT REVISION SURGERY ALLEGEDLY THERE WERE FINDINGS OF "THE PRESENCE OF MILKY, TURBID FLUID; LARGE PSEUDOTUMOR FORMATION; SOFT TISSUE NECROSIS; MUSCLE LOSS AND/OR BNY NECROSIS AT THE PROXIMAL FEMUR". | THE INFO IN THIS REPORT WAS PROVIDED BY STRYKER ORTHOPAEDICS LEGAL AFFAIRS DEPT. NO ADD'L INFO IS AVAILABLE AT THIS TIME. THE DEVICE IS NOT AVAILABLE DUE TO THE ONGOING LITIGATION. SHOULD DEVICE OR ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01646 | 2012-09-25 | 2012-09-07 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS - MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT THE PT HAS HAD FLUID LEACHING. ABOUT FOUR MONTHS AGO, STARTED HAVING PAIN AND DISCOMFORT. MRI AND BLOODWORK WERE DONE. PT HAS HIGH LEVELS OF COBALT. PT HAS BEEN HOBBLING. REVISION IS SCHEDULED FOR (B)(6) 2012. | IT WAS FURTHER REPORTED THAT PT HAD A REJUVENATE PRIMARY RIGHT HIP REPLACEMENT DONE (B)(6) 2011 AND EXPERIENCED PAIN. PT WAS REVISED ON (B)(6) 2012. S SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. THIS IS THE SAME PT/EVENT AS MFR# 9616680-2012-00803. | Y | N | OTHER |  | 2012-09
2249697-2012-01655 | 2012-09-25 | 2010-10-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PT HAS HAD PAIN AND HAS BEEN GETTING CORTISONE SHOTS. MRI WAS DONE AND THERE WAS FLUID AROUND THE JOINT. REVISION IS PLANNED FOR (B)(6) 2012. ADD'L INFO RECEIVED (B)(6) 2012: PT STATED THAT SHE WAS RELEASED FROM THE HOSP TODAY AFTER HER REVISION SURGERY ON (B)(6) 2012. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD DEVICE OR ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01651 | 2012-09-25 | 2012-05-07 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PT HAD SEVERE PAIN SINCE DAY ONE. DOCTOR DID MANY TESTS FOR INFECTION, WHICH CAME BACK NEGATIVE. PT STARTED SEEING DIFFERENT DOCTOR (B)(6). HE ALSO DID TESTS WHICH CAME BACK NEGATIVE ALSO, BUT HE SUSPECTED BONE INFECTION. IT WAS DISCOVERED THAT IMPLANT CORRODED IN HER. IT WAS REMOVED ON (B)(6) AND SPACER WAS PUT IN UNTIL (B)(6) 2012, WHEN ANOTHER IMPLANT WAS PUT IN. HER LEFT LEG IS STILL NUMB. HAS SUFFERED FROM C-DIFF (SEVERE INTESTINAL INFECTION) AND IS STILL SUFFERING FROM YEAST INFECTION ALL OVER HER BODY. | AN EVAL OF THE REPORTED DEVICE CANNOT BE PERFORMED AS IT WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01647 | 2012-09-25 | 2010-02-22 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT, RECALL. REVISION IS SCHEDULED FOR (B)(6) 2012. | THE LRG TAP PRI MOD NCK 0DEG 30MM, CAT # NLS-300000B, LOT# 29682801 WAS ALSO LISTED IN THIS REPORT. IT CANNOT BE DETERMINED WHICH, IF ANY OF THE REPORTED DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S FUTURE REVISION. AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. THE REPORTED EVENT BECAME LEGAL AND WILL BE HANDLED BY STRYKER ORTHOPAEDICS LEGAL AFFAIRS DEPARTMENT. NO ADD'L INFO IS AVAILABLE AT THIS TIME DUE TO THE ONGOING LITIGATION. SHOULD DEVICE OR ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01654 | 2012-09-25 | 2012-04-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PT HAS HAD PAIN IN GROIN, LEG, AND HIP AREA, SINCE SURGERY. PT HAD SURGERY TO RELEASE TENDON IN (B)(6) TO TRY AND RELIEVE PAIN. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | LAY USER/PATIENT |  | 2012-09
2249697-2012-01656 | 2012-09-25 | 2012-04-23 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PT, THROUGH THE FILING OF A LAWSUIT, IS REPORTING AND ALLEGING THAT A REJUVENATE MODULAR FEMORAL STEM REQUIRED REVISION DUE TO PAIN AND THAT AT REVISION SURGERY ALLEGEDLY THERE WERE FINDINGS OF "THE PRESENCE OF MILKY, TURBID FLUID; LARGE PSEUDOTUMOR FORMATION; SOFT TISSUE NECROSIS; MUSCLE LOSS AND/OR BONY NECROSIS AT THE PROXIMAL FEMUR". | THE INFO IN THIS REPORT WAS PROVIDED BY STRYKER ORTHOPAEDICS LEGAL AFFAIRS DEPARTMENT. NO ADD'L INFO IS AVAILABLE AT THIS TIME. THE DEVICE IS NOT AVAILABLE DUE TO THE ONGOING LITIGATION. SHOULD DEVICE OR ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01659 | 2012-09-26 | 2011-06-08 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT STATED THAT SHE HAS HAD PAIN SINCE SURGERY. SHE HAS PAIN IN HER LEFT HIP, LEG, AND EXTREME PAIN IN HER LOWER BACK. SHE SAID THAT EXTREME AMOUNT OF FLUID WAS EXTRACTED FROM HER HIP JOINT. HER DOCTOR ORDERED HER PAIN SHOTS. HER BLOOD LEVELS ARE HIGHER THAN NORMAL. SHE HAS HARD TIME STANDING OR SITTING FOR LONG TIME BECAUSE, HER LEG WOULD GET NUMB. SHE HAS HARD TIME SLEEPING ON HER LEFT SIDE. PT'S REVISION SURGERY IS SCHEDULED FOR (B)(6) 2012. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01677 | 2012-09-26 | 2012-09-04 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE #10 STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT WILL HAVE A REVISION SURGERY IN RIGHT HIP. MRI SHOWED MASSIVE TISSUE DESTRUCTION AND FLUID BUILD UP IN JOINT. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. THIS IS THE SAME PT/EVENT AS MFR # 9616680-2012-00811. | Y | N | OTHER |  | 2012-09
9616680-2012-00811 | 2012-09-26 | 2012-09-04 | I, n, j, u, r, y | Manufacturer report | UNK REJUVENATE 127, 38 NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT THE PT WILL HAVE A REVISION SURGERY IN RIGHT HIP. MRI SHOWED MASSIVE TISSUE DESTRUCTION AND FLUID BUILD UP IN JOINT. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. THIS IS THE SAME PT/EVENT AS MFR #2249697-2012-01677. | Y | N | OTHER |  | 2012-09
2249697-2012-01684 | 2012-09-26 | 2012-09-19 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED BY THE PT THAT THE SURGEON WAS AMAZED AT THE AMOUNT OF SCAR TISSUE. PAIN FOR A YEAR AND A HALF (2009) IN PELVIC AND GROIN AREA, INCLUDING SWELLING STARTING FROM (B)(6). STARTING FROM LAST (B)(6), SHE WAS WALKING WITH RIGHT HAND ON RIGHT LEG. HAS BEEN IN A WHEELCHAIR SINCE LAST (B)(6). SCAR TISSUE WAS CLEANED ON (B)(6) AND WAS RELIEVED TEMPORARILY. SIX WEEKS AGO THE PAIN RETURNED. ALLERGIC TO COBALT; IMPLANT IS CORRODING. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. THIS IS THE SAME PT/EVENT AS MFR # 9616680-2012-00813. | Y | N | OTHER |  | 2012-09
2249697-2012-01660 | 2012-09-26 | 2011-08-15 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPACT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT HAD SIX TO EIGHT INCHES DISCOMFORT AREA INTO THE THIGH. HIS RIGHT LEG IS WEAKER THAN THE LEFT LEG. HIS RIGHT LEG FEELS SORE AND STIFF. HE HAD X-RAY ON FRIDAY (B)(6) 2012 AND WILL DO AN MRI AND BLOOD TESTS. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01688 | 2012-09-26 | 2020-01-02 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | PROSTHESIS, HIP, SEMI-CONSTRAINED, UNCEMENTED, METAL/POLYMER, NON-POROUS, CALICU | STRYKER ORTHOPAEDICS-MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT THE PT'S HUSBAND CALLED BECAUSE PT IS DEAF. HE SAID THAT THE PT SAW HER DOCTOR ON (B)(6) 2012. HE STATED THAT THE PT HAS PAIN SLEEPING ON HER SIDE. HE ALSO SAID HER COBALT LEVELS ARE HIGH. HE SAID THAT PT CANNOT HAVE AN MRI SO HER DOCTOR ORDERED A CAT SCAN INSTEAD, THE SCAT SCAN IS NOT YET DONE AS OF (B)(6) 2012. | THE PT IS (B)(6). IT WAS INDICATED THAT THE DEVICE IS NOT AVAILABLE FOR EVALUATION, AS IT IS STILL IMPLANTED. ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | 0 |  | 2012-09
2249697-2012-01707 | 2012-09-27 | 2012-09-07 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT PATIENT HAS PAIN AND A LARGE PSEUDOTUMOR. DR (B)(6) REMOVED ALL COMPONENTS AND REPLACED WITH A RESTORATION HA STEM, TITANIUM SHELL AND A 40 MM HEAD AND LINER. | AN EVAL OF THE REPORTED DEVICES CANNOT BE PERFORMED AS THEY WERE RETAINED BY THE PT AND WERE NOT RETURNED TO THE MFR. ADD¿L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD¿L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS MFR NUMBER: 9616680-2012-00827. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01701 | 2012-09-27 | 2012-09-06 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT, THE PT HAD COMPLAINED OF ONSET OF PAIN WITHIN THE LAST FEW MONTHS. PT WAS TREATED BY DR (B)(6), AND TESTS WERE DONE THAT INDICATED THE PT WAS SCHEDULED FOR REMOVAL OF THE REJUVENATE STEM/NECK AND HEAD AS WELL AS ADM PLASTIC PORTION. UPON REVISION A RESPONSIVE HEMATOMA TYPE FLUID WAS EVIDENT. A SLURRY FLUID AND RESPONSIVE TISSUE WAS REMOVED AND RESECTED. THE PT WAS REIMPLANTED WITH SECUR-FIT. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR. ADD¿L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD DEVICE OR ADD¿L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS MFR NUMBER: 9616680-2012-00825. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01705 | 2012-09-27 | 2012-09-06 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT THE PT PRESENTED WITH PAINFUL HIP. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0DEG 34MM, CAT# NLS-34000B, LOT # 36169001. RESTORATION ADM. CUP W/HA, CAT# 1235-2-501, LOT # G3075322. RESTORATION ADM X3 INS 28/50, CAT # 1236-2-850, LOT # 34923201. 28MM STD LFIT V40 HEAD, CAT# 6260-9-128, LOT# 36679306. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR DUE TO HOSPITAL POLICY. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01700 | 2012-09-27 | 2012-09-06 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS | NA | SPT-090000S |  | IT WAS REPORTED THAT, PT HAD HIP COMPONENTS EXPLANTED DUE TO ADVERSE LOCAL TISSUE REACTION. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR DUE TO HOSPITAL POLICY. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED, BUT NOT MADE AVAILABLE. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PATIENT/EVENT AS MFR # 9616680-2012-00823. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01712 | 2012-09-28 | 2011-11-30 | I, n, j, u, r, y | Manufacturer report | REJUVENATE HIP COMPONENTS | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT NOTICE WAS RECEIVED FROM (B)(6) ALLEGING THAT PATIENT HAD SUSTAINED PERSONAL INJURIES ON (B)(6) 2011 FROM THE IMPLANTATION OF THE STRYKER REJUVENATE HIP COMPONENTS. | THE INFORMATION IN THIS REPORT WAS PROVIDED BY STRYKER ORTHOPAEDICS LEGAL AFFAIRS DEPARTMENT. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DUE TO ONGOING LITIGATION. SHOULD DEVICE OR ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01725 | 2012-09-28 | 2012-06-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: HUSBAND ((B)(6)) CALLED FOR HIS WIFE. HE CLAIMED THAT HIS WIFE HAS HAD PAIN IN HER RIGHT LEG FOR TWO TO THREE MONTHS. HER MRI SHOWED A TUMOR AROUND THE IMPLANT. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01714 | 2012-09-28 | 2009-11-18 | I, n, j, u, r, y | Manufacturer report | REJUVENATE HIP COMPONENTS | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT NOTICE WAS RECEIVED FROM (B)(4), ALLEGING THAT PATIENT HAD SUSTAINED PERSONAL INJURIES ON (B)(6) 2009 FROM THE IMPLANTATION OF THE STRYKER REJUVENATE HIP COMPONENTS. | THE INFORMATION IN THIS REPORT WAS PROVIDED BY STRYKER ORTHOPAEDICS LEGAL AFFAIRS DEPARTMENT. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DUE TO THE ONGOING LITIGATION. SHOULD DEVICE OR ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-09
2249697-2012-01762 | 2012-10-03 | 2011-02-14 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT HIP REJUVENATE/ABGII | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS STATED THAT THE PT HAS GROIN PAIN SINCE THE SURGERY AND HIGH LEVELS OF METAL IN HIS BLOOD. PT'S WIFE STATED THAT PT MAY NEED A REVISION SURGERY. PT'S WIFE STATED SHE HAS NO MONEY FOR THIS SURGERY, MRIS AND RELATED EXPENSES. PT'S WIFE STATED THAT THE SURGERY IS CAUSING TIME, MONEY, ANGUISH, AND HER INSURANCE IS NOT COVERING A LOT OF THESE COSTS. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01751 | 2012-10-03 | 2010-03-03 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT HIP REJUVENATE/ ABGII | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS STATED THAT THE PT EXPERIENCED PAIN IN THE GROIN AREA. PT HAD BLOOD WORK AND THE RESULTS WERE ELEVATED METAL LEVELS. PT¿S WIFE REPORTED THAT PT IS SCHEDULED FOR MRI. PT¿S WIFE STATED THAT PT MAY NEED A REVISION SURGERY. PT¿S WIFE STATED SHE HAS NO MONEY FOR SURGERY, MRIS AND RELATED EXPENSES. PT¿S WIFE STATED THAT THE SURGERY IS CAUSING TIME, MONEY, ANGUISH AND HER INSURANCE IS NOT COVERING A LOT OF THESE COSTS. | ADD¿L INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADD¿L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
MW5027197 | 2012-10-03 | 2012-08-15 | I, n, j, u, r, y | Voluntary report | STRYKER | PROSTHETIC HIP | STRYKER |  | UNK |  | PATIENT WAS INFORMED IN A LETTER MAILED BY SURGEON THAT PROSTHETIC HIP REPLACEMENT WAS DEFECTIVE AND RECALLED. LETTER RECOMMENDED PATIENT HAVE CONSULTATION WITH SURGEON. DURING THIS TIME FRAME PATIENT NEEDED TWO ROOT CANAL THERAPIES ALL OF A SUDDEN AND WAS CONCERNED THAT CHROME/COBALT MATERIAL LEACHING FROM DEFECTIVE HIP PROSTHESIS WAS CONTRIBUTORY TO THESE SUDDEN CHANGES IN HEALTH. |  | N | Y | 0 |  | 2012-10
2249697-2012-01763 | 2012-10-03 | 2011-06-29 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT HIP REJUVENATE /ABGII | IMPLANT | STRYKER ORTHOPAEDIC MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT IS EXPERIENCING PAIN, LIMITED MOVEMENT OF THE LEFT LEG WHERE THE IMPLANT IS AND DIFFICULTY SLEEPING. PT STATED THAT SHE CANNOT WALK AND CANNOT PUT WEIGHT ON THE LEFT LEG. AN MRI SHOWS FLUID IN THE HIP. PT HAD BLOOD WORK THAT SHOWS ELEVATED METAL LEVELS. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED. HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01738 | 2012-10-03 | 2011-09-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PATIENT COMPLAINS OF PAIN SINCE SURGERY. HE IS UNABLE TO WALK DUE TO THE PAIN. HE IS ON PAIN MEDICATION, BUT THIS IS NOT HELPING HIS PAIN. HIS SURGEON HAS COMMUNICATED TO HIM THAT HE WILL NEED A REVISION SURGERY. PATIENT IS CURRENTLY ON DISABILITY. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01733 | 2012-10-03 | 2012-09-10 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT PT COMPLAINED OF PAINFUL HIP. POSSIBLE PUSS ENCOUNTERED DURING DISSECTION. CULTURES WERE TAKEN AND AN ANTIBIOTIC SPACER WAS PLACED. PT MAY SEEK COMPENSATION. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0DEG 38MM, CAT# NLS-380000B, LOT# 32987901R; PRIMARY TRITANIUM HEMI CLUSTER HOLE CUP 58MM, CAT# 502-03-58F, LOT# MKLYHH; 6.5 CANCELLOUS BONE SCREW 25MM, CAT# 2030-6525-1, LOT# MKMJP1; TRIDENT 10DEG X3 INSERT 36MM ID, CAT #623-10-36F, LOT# MKP93L; DELTA V-40 CERAMIC HEAD 36/0, CAT# 6570-0-136, LOT# 38522301. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S INFECTION. AN EVAL OF THE REPORTED DEVICES CANNOT BE PERFORMED AS THEY WERE RETAINED BY THE PT AND WERE NOT RETURNED TO THE MFR. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01764 | 2012-10-03 | 2010-06-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE/ABG II MODULAR HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT COMPLAINS OF PAIN. PT WILL REQUIRE A REVISION SURGERY BECAUSE TESTS HAVE SHOWN FLUID BUILD UP IN THREE AREAS OF THE LEFT HIP JOINT. THIS HAS DESTROYED THE MUSCLE IN HIS LEG AND HE HAS LOST MUSCLE STRENGTH. TESTS HAVE ALSO SHOWN TITANIUM IN HIS BLOOD. HIS REVISION SURGERY IS SCHEDULED FOR (B)(6) 2012. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01730 | 2012-10-03 | 2011-05-31 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, WIFE CALLED ON BEHALF OF HER HUSBAND. SHE SAID THAT HER HUSBAND HAS DISCOMFORT IN HIS LEFT HIP. HE HAD A F/U ON (B)(6) 2012. SHE NOTED THAT PT WILL HAVE MRI AND BLOOD WORK ON (B)(6) 2012. SHE MENTIONED THAT TESTING WILL BE DONE IN AN OUT-OF-NETWORK HOSPITAL AND PT WILL TAKE A DAY OFF WORK. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01743 | 2012-10-03 | 2012-02-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE/ABG II HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PT IS EXPERIENCING PAIN IN THE BACK OF THE RIGHT HIP IMPLANT, SORENESS DOWN THE LEG OF IMPLANT, PAIN IN THE GROIN AREA AND CANNOT WALK SINCE THE SURGERY. PT STATED THAT SHE WILL BE NEEDING REVISION SURGERY. PT STATED THAT SHE WILL BE HAVING AN MRI AROUND (B)(6), 2012. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01747 | 2012-10-03 | 2012-02-10 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT HIP REJUVENATE/ABG II HIP | IMPLANT | STRYKER ORTHOPAEDICS MEHWAH | NA | UNK |  | IT WAS REPORTED THAT PT HAD IMPLANT SURGERY TO HIS LEFT HIP IN (B)(6) 2012. HE UNDERWENT PHYSICAL THERAPY FOR A FEW WEEKS AND CONTINUED HIS EXERCISES AT HOME. PT STATES THAT HE WAS GETTING BETTER; HOWEVER, IN (B)(6) HE NOTICED THAT HE WAS NOT IMPROVING IN HIS MOBILITY AND WAS STILL EXPERIENCING PAIN. PT STATES HIS PAIN IS LOCALIZED DIRECTLY IN THE AREA OF THE HIP IMPLANT. HE HAS BEEN TAKING OVER THE COUNTER PAIN MEDICATION TO ALLEVIATE THE PAIN. HE CALLED THE SURGEON TO REPORT THE PAIN AND WAS ADVISED TO COME IN. HE WAS ADVISED OF THE RECALL DURING THIS APPOINTMENT. THE PT HAD AN X-RAY AND HIS SURGEON STATED THAT IT DID NOT SHOW ANY ABNORMALITIES. NO OTHER TESTS HAVE BEEN ORDERED AT THIS TIME. THE DOCTOR ADVISED HIM TO TAKE TWO WEEKS OFF FROM WORK TO RESET HIS HIP IN THE EXPECTATION THAT THE PAIN WOULD SUBSIDE. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01759 | 2012-10-03 | 2010-11-22 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PT HAS THE RECALLED HIP. HE HAS HAD AN MRI AND BLOOD TESTS. HIS COBALT LEVEL IS 17 AND WILL BE REQUIRED TO HAVE FOLLOW UP TESTS EVERY 2-3 MONTHS AS THE FDA DOES NOT HAVE A SET LEVEL OF COBALT. PT HAS ALSO LOST SOME HEARING AND IS CONSIDERING THAT IT MAY BE A SIDE EFFECT OF HIGH LEVELS OF COBALT. HE WILL FOLLOW UP WITH HIS SURGEON TO ADDRESS HIS CLINICAL QUESTIONS. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADDITIONAL INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01760 | 2012-10-03 | 2011-12-12 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PT RECEIVED LETTER FROM DR. AND THE DR. MENTIONED THAT SHE WILL PROBABLY NEED A HIP REVISION SURGERY. SHE HAS AN MRI SCHEDULED FOR (B)(6) 2012. SHE HAS HAD BLOODWORK AND XRAYS COMPLETED ON (B)(6) 2012. SHE HAS HAD TENDERNESS SINCE THE SURGERY AND HER BLOOD LEVELS ARE ELEVATED. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADDITIONAL INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01755 | 2012-10-03 | 2012-09-12 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT A REJUVENATE REMOVAL AND CONVERSION TO SECUR-FIT PLUS. PT HAD AN ELEVATED ION LEVEL AND COMPLAINED OF PAIN. STEM WAS REMOVED AS WELL AS MODULAR NECK AND HEAD/MDM IMPLANT. SECUR-FIT PLUS WAS IMPLANTED WITH MDM (NEW) AND APPROPRIATE LENGTH HEAD. | AN EVAL OF THE REPORTED DEVICES CANNOT BE PERFORMED AS THEY WERE RETAINED BY THE HOSPITAL AND WERE NOT RETURNED TO THE MFR. ADDITIONAL INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS MFR # 9616680-2012-00841. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01732 | 2012-10-03 | 2011-10-25 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE #9 HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PT SAID THAT HE HAD NO SYMPTOMS OTHER THAN LITTLE DISCOMFORT WALKING. HIS MRI SHOWED FLUID ACCUMULATION AT THE JOINT. HIS BLOOD TEST RESULTS INDICATED HIGH LEVEL OF CHROMIUM AND COBALT. PT SAID THAT HE HAS NO SWELLING. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01753 | 2012-10-03 | 2010-04-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PT COMPLAINS OF PAIN AND STIFFNESS. BLOOD TESTS HAVE SHOWN 8.9 LEVEL OF SOME METAL. AN MRI HAS SHOWN A "3 INCH HOT DOG" SIZED PSEUDOTUMOR. SURGEON SEES SUITABLE TO HAVE PT GO THROUGH A REVISION SURGERY. PT CALLED TO FILE A CLAIM AS HE WANTS TO KNOW WHO IS RESPONSIBLE FOR ALL THESE EXPANSES. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01737 | 2012-10-03 | 2011-01-19 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PT COMPLAINS OF PAIN IN GROIN AREA AND OUTER THIGH. HE HAS RECEIVED THE RECALL LETTER FROM HIS SURGEONS AND IT'S BEING REQUIRED TO HAVE AN MRI AND OTHER STUDIES, BUT STATES THAT HE CANNOT AFFORD THEM. THE ANXIETY OF ALL THIS IS REALLY AFFECTING HIM AND HE WANTS STRYKER TO HELP HIM. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01746 | 2012-10-03 | 2011-11-02 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE MODULAR HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED PT COMPLAINS OF MINOR PAIN IN JOINT AREA AND MORE PRONOUNCED PAIN IN TROCHANTER AREA. PT HAS NOTICED LACK OF STRENGTH AND HIS SURGEON HAS HAD AN MRI AND BLOOD WORK. HE HAS ELEVATED COBALT LEVELS. THE MRI SHOWS SOME IRREGULARITIES. PT WILL F/U WITH SURGEON. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01740 | 2012-10-03 | 2012-07-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN SIZE 7 REJUVENATE STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PT'S PSL CUP, SIZE 54, LOOSENED AND DR. L. PLANNED TO REVISE CUP. IN LIEU OF THE REJUVENATE RECALL, WE ALSO PLANNED TO REVISE THE SIZE 7 REJUVENATE STERM WITH A 34 NEUTRAL NECK. THE PRIMARY HEAD WAS A 40MM +0. UPON MAKING HIS INCISION, DR. L. NOTICED A MODERATELY SIZED GRAYISH TUMOR THAT, WHEN EXCISED, DISCHARGED FLUID. THERE WERE SEVERAL OTHER MUCH SMALLER FLUID FILED SACS NOTED WITHIN THE WOUND. BONE STOCK WAS NOT AFFECTED AND THE FLUID FILED TUMORS WERE CONTAINED IN THE REGION OF THE HIP. THE REVISION SURGERY WAS COMPLETED WITHOUT COMPLICATION. | AN EVAL OF THE REPORTED DEVICES CANNOT BE PERFORMED AS THEY WERE RETAINED BY THE HOSPITAL AND WERE NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS MFR # 9616680-2012-00836 AND 2249697-2012-01741. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01739 | 2012-10-03 | 2010-03-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PATIENT COMPLAINS OF SORENESS. PATIENT HAS RECEIVED RECALL LETTER AND HE WILL REQUIRE AN MRI, TEST, AND FURTHER EVALUATIONS. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01769 | 2012-10-04 | 2010-04-13 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | MPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT PATIENT IS HAVING REVISION SURGERY ON (B)(6) 2012. SHE HAS BEEN EXPERIENCING PAIN (B)(6) 2012. PATIENT STATES THAT SHE CANNOT LIFT HER LEG MORE THAN A FEW INCHES, STAND OR SIT FOR LONG PERIODS OF TIME WITHOUT PAIN. SHE ALSO STATES SHE CANNOT BEND OVER OR CROUCH WITHOUT SIGNIFICANT DISCOMFORT. SHE HAD AN MRI TAKEN APPROXIMATELY 2 WEEKS AGO WHICH SHOWED SIGNIFICANT FLUID ACCUMULATION IN THE HIP AREA. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 8DEG 34MM, CAT #NLV-340800G, LOT # 26580402; 28MM STD LFIT V40 HEAD, CAT# 6260-9-128, LOT # 31073802. RESTORATION ADM X3 INS 28/52, CAT# 1236-2-852, LOT # 32805801. MD VIT SLV AND 2.0MM HOMOG CBL, CAT# 6704-0-510, LOT # 296575503. RESTORATION ADM, CUP W/HA, CAT# 1235-2-521, LOT # G2941841. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S PAIN. AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01775 | 2012-10-04 | 2011-04-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PT REC'D A LETTER FROM DOCTOR (B)(6) REGARDING THE RECALL. THE PT HAS BEEN EXPERIENCING SIGNIFICANT PAIN SINCE THE (B)(6) 2012. HE STATES THAT HE HAS LIMITED FUNCTIONALITY AND MOBILITY IN HIS HIP. HE VISITED HIS DOCTOR ON (B)(6) 2012. MRI AND X-RAYS WERE TAKEN. BLOOD WORK SHOWED ELEVATED BLOOD LEVELS. HE HAD 15CC'S OF FLUID ASPIRATED FROM HIS HIP. REVISION SURGERY IS SCHEDULED FOR (B)(6) 2012. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01770 | 2012-10-04 | 2012-09-14 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT HIP/REJUVENATE IMPLANT | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PT REPORTED THAT HER HIP HAS ALWAYS BEEN SORE. SHE STATES THAT THE SORENESS INCREASES WITH CERTAIN ACTIVITIES BUT NEVER GOES AWAY. THE RIGHT HIP FEELS HEAVY AND PAIN IS IN FRONT OF HER THIGH. THE PT HAS BEEN TO HER DR. ON (B)(6) FOR TESTING I.E, BLOOD WORK, BONE SCAN AND MRI. DOCTOR HAS DIAGNOSED HER WITH METALLOSIS AND STATES THAT SHE WILL NEED REVISION SURGERY. THE SURGERY HAS NOT YET BEEN SCHEDULED. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE(S) REFERENCE IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01776 | 2012-10-04 | 2010-05-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PT VISITED HER SURGEON FOR HER TWO YEAR CHECKUP IN (B)(6) 2012. AT THE TIME SHE REPORTED HAVING SOME PAIN AND TENDERNESS IN THE RIGHT SIDE OF HER HIP. SHE DOES STATE THAT SHE HAS ALWAYS HAD SOME PAIN SINCE THE SURGERY BUT ATTRIBUTED IT TO A NORMAL SIDE EFFECT OF THE IMPLANT PROCESS. AN MRI WAS TAKEN IN AUGUST WHICH SHOWED PSEUDO-TUMORS AROUND THE IMPLANT SIGHT. SHE IS SCHEDULED TO HAVE REVISION SURGERY ON (B)(6) 2012. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCE IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01768 | 2012-10-04 | 2011-12-12 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORHTOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT THE PT COMPLAINS OF PAIN SINCE THE SURGERY. AN MRI AND TESTS HAVE SHOWN THAT HER HIP NEEDS TO BE REMOVED AND REPLACED. SHE IS SCHEDULED TO HAVE A REVISION ON (B)(6) 2012. IT WAS FURTHER REPORTED THAT MECHANICALLY ASSISTED CREVICE CORROSION WITH SECONDARY ADVERSE LOCAL LYMPHOCYTIC SOFT TISSUE REACTION - PSEUDOTUMOR-SECONDARY TO MODULAR REJUVENATE FEMORAL COMPONENT- LOCAL TISSUE NECROSIS. | THE PT IS (B)(6). THE PT REQUESTED TO KEEP THE REPORTED DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-10
2249697-2012-01783 | 2012-10-05 | 2012-09-17 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT ON (B)(6) 2011, PT HAD A TOTAL LEFT HIP DONE AND WAS HAVING PAIN. SURGEON REVISED REJUVENATE STEM AND ALL OTHER PARTS - HEAD, NECK, CUP AND LINER, AND CEMENTED IN AN EXETER STEM, NEW LINER, NEW CUP, NEW NECK AND NEW HEAD ON (B)(6) 2012. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0DEG 34MM, CAT #NLS-340000B, LOT #38368902; TRIDENT PSL HA SOLID BACK 54MM, CAT #540-11-54F, LOT #38118601; TRIDENT 0 DEG X3 INSERT 36MM ID, CAT #623-00-36F, LOT #MKL2LW; DELTA V-40 CERAMIC HEAD 36/+2,5, CAT #6570-0-536, LOT #37902903. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVAL OF THE DEVICES CANNOT BE PERFORMED AS THE REVISED DEVICES WERE RETAINED BY THE HOSP AND WERE NOT RETURNED TO THE MFR. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01782 | 2012-10-05 | 2012-09-14 | I, n, j, u, r, y | Manufacturer report | UNK REJUVENATE MODULAR HIP STEM (LEFT) | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PT COMPLAINS OF PAIN SINCE THE SURGERY. PAIN IS IN THE JOINT AND GROIN AREA. SHE REPORTS ELEVATED LEVELS OF CHROMIUM. SHE HAS HAD LAB WORK, MRI AND WILL HAVE OTHER TESTS TO TEST FOR INFECTION IN JOINT AREA. SHE IS SCHEDULED FOR A REVISION SURGERY ON (B)(6) 2012. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WILL BE REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS: MFR #9616680-2012-00854. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
9616680-2012-00854 | 2012-10-05 | 2012-09-14 | I, n, j, u, r, y | Manufacturer report | UNK REJUVENATE MODULAR NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT, PT COMPLAINS OF PAIN SINCE THE SURGERY. PAIN IS IN THE JOINT AND GROIN AREA. SHE REPORTS ELEVATED LEVELS OF CHROMIUM. SHE HAS HAD LAB WORK, MRI AND WILL HAVE OTHER TESTS TO TEST FOR INFECTION IN JOINT AREA. SHE IS SCHEDULED FOR A REVISION SURGERY ON (B)(6) 2012. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WILL BE REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS: MFR #2249697-2012-01782. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01778 | 2012-10-05 | 2012-09-19 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE LEFT HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT COMPLAINTS OF PAIN THAT STARTED A FEW MONTHS AFTER THE SURGERY AND HAS LACK OF MOBILITY. PT HAD AN MRI AND BLOOD TESTS WHICH SHOWED INCREASED CHROMIUM LEVELS, FLUID ACCUMULATION AND INFLAMMATION. PT IS SCHEDULED TO HAVE A REVISION SURGERY ON (B)(6) 2012. | ADD¿L INFO HAS BEEN REQUESTED AND IF REC¿D, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-10
2249697-2012-01828 | 2012-10-10 | 2009-07-29 | I, n, j, u, r, y | Manufacturer report | TRIDENT 0° X3 INSERT 36MM ID | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, POROUS UNCEMENTED | STRYKER ORTHOPAEDICS-MAHWAH | NA | 623-00-36F |  | IT WAS REPORTED THROUGH THE ATTORNEY FOR THE PATIENT, AS A RESULT OF A LEGAL CLAIM, THAT ALLEGEDLY THE PATIENT UNDERWENT RIGHT TOTAL HIP ARTHROPLASTY ON (B)(6) 2006. IT IS FURTHER ALLEGED THAT "ON (B)(6) 2009, HE SUFFERED A DISLOCATION OF HIS RIGHT HIP IMPLANT WHICH REQUIRED A CLOSED REDUCTION. A FEW DAYS LATER, HE SUFFERED A SECOND DISLOCATION WHICH ALSO NECESSITATED A CLOSED REDUCTION." | THE INFORMATION IN THIS REPORT WAS PROVIDED BY STRYKER ORTHOPAEDICS LEGAL AFFAIRS DEPARTMENT. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. THE DEVICE IS NOT AVAILABLE AS IT REMAINED IMPLANTED. DEVICE HISTORY REVIEW INDICATED THE REPORTED DEVICES WERE MANUFACTURED AND ACCEPTED INTO FINAL STOCK WITH NO REPORTED DISCREPANCIES. THE PRODUCT EXPERIENCE REPORTING DATABASE WAS REVIEWED FOR SIMILAR REPORTED EVENTS REGARDING DISLOCATION. (B)(4) EVENT IS REGARDING THE SAME PATIENT. SHOULD DEVICE OR ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01829 | 2012-10-10 | 2011-01-28 | I, n, j, u, r, y | Manufacturer report | V40 COCR LFIT HEAD 36MM/+5 | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | STRYKER ORTHOPAEDICS-MAHWAH | NA | 6260-9-236 |  | IT WAS REPORTED THROUGH THE ATTORNEY FOR THE PATIENT, AS A RESULT OF A LEGAL CLAIM, THAT ALLEGEDLY THE PATIENT UNDERWENT RIGHT TOTAL HIP ARTHROPLASTY ON (B)(6) 2006. IT IS FURTHER ALLEGED, "IN (B)(6) 2011, HE LOST CONTROL OF HIS RIGHT LEG WHILE GETTING UP FROM HIS OFFICE CHAIR. RADIOGRAPHS SHOWED A COMPLETE DISASSOCIATION OF THE FEMORAL HEAD FROM THE STEM WITH THE TRUNNION EXTENDING INTO THE ACETABULAR CUP." THE PATIENT UNDERWENT REVISION SURGERY ON (B)(6) 2011. | THE INFORMATION IN THIS REPORT WAS PROVIDED BY STRYKER ORTHOPAEDICS LEGAL AFFAIRS DEPARTMENT. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. THE DEVICE IS NOT AVAILABLE DUE TO THE ONGOING LITIGATION. SHOULD DEVICE OR ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01799 | 2012-10-10 | 2011-09-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT HIP - REJUVENATE/ABGII | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT IS SCHEDULED FOR REVISION SURGERY ON (B)(6) 2012 ON HIS RIGHT HIP. PT STATES THAT HIS DR ORDERED BLOOD WORK WHICH SHOWED ELEVATED CHROMIUM AND COBALT LEVELS. PT COMPLAINS OF HAVING DULL PAIN ALMOST ALL THE TIME THAT SOMETIMES BECOMES SHARP. HE EXPERIENCES SOME STIFFNESS AND DIFFICULTY WALKING ESPECIALLY AFTER PERIODS OF ACTIVITY. HE HAS DIFFICULTY SLEEPING ESPECIALLY WITHIN THE LAST YEAR DUE TO PAIN IN HIS HIPS. PT STATES HE HAS PERIODIC WEAKNESS IN BOTH LEGS BUT MORE EXTREME IN THE RIGHT. PAIN IN BOTH HIPS WILL SOMETIMES EXTEND DOWN INTO THE FEMUR. PT HAD AN MRI AND ADDITIONAL BLOOD LABS ON (B)(6) 2012. THE PT STATES THAT HE HAS THE SAME SYMPTOMS IN BOTH HIPS BUT THOSE IN HIS LEFT HIP ARE LESS INTENSE AND FREQUENT THAN THOSE IN HIS RIGHT HIP. HE LAST SAW HIS DOCTOR ON (B)(6) 2012. THE PT ASKED TO FILE A CLAIM FOR THE REVISION SURGERY ON HIS RIGHT HIP. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01802 | 2012-10-10 | 2011-08-31 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT HIP REJUVENATE/ABGII | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | PT STATES THAT SHE HAS DISCOMFORT IN THE LEFT HIP. SHE WENT TO THE DOCTOR ON (B)(6) 2012. SHE HAD BLOOD LABS DONE ON (B)(6) 2012. SHE RECEIVED THE RESULTS FROM HER DOCTOR'S OFFICE ON (B)(6) 2012 SAYING THAT HER LEVELS WERE ELEVATED. SHE IS BEING SCHEDULED FOR AN MRI THROUGH THE DOCTOR'S OFFICE. THE PT'S STATES THAT SHE IS HAVING TROUBLE MANEUVERING STAIRS. SHE WAS GIVEN A PRESCRIPTION FOR ANTI-INFLAMMATORY MEDICATION ON (B)(6) 2012. SHE BEGAN HAVING TENDERNESS AND PERIODIC NERVE TINGLING IN THE HIP (B)(6) 2012. SHE ALSO HAS OCCASIONAL DISCOMFORT IN THE GROIN AREA. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01800 | 2012-10-10 | 2008-11-20 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT HIP-REJUVENATE/ABGII | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT IS SCHEDULED FOR REVISION SURGERY ON (B)(6) 2012. PT STATES THAT HIS DR. ORDERED BLOOD WORK THAT SHOWED ELEVATED CHROMIUM AND COBALT LEVELS. HAS DULL PAIN ALMOST ALL THE TIME, SOMETIMES BECOMES SHARP. FREQUENT STIFFNESS AND DIFFICULTY WALKING ESPECIALLY AFTER PERIODS OF ACTIVITY. DIFFICULTY SLEEPING ESPECIALLY IN THE LAST YEAR. FREQUENTLY WAKES FROM PAIN IN HIS HIP. PERIODIC WEAKNESS IN BOTH LEGS BUT MORE EXTREME IN THE RIGHT. OFTEN FEELS THE LEG IS GOING TO GIVE AWAY. PAIN IN BOTH HIPS WILL SOMETIMES EXTEND DOWN INTO THE FEMUR. PT HAD AN MRI AND ADDITIONAL BLOOD LABS ON (B)(6) 2012. HE SAW HIS SURGEON ON (B)(6) 2012. THE MRI SEEMED TO INDICATE FLUID AT THE HIP SITE. THE PT WAS SENT TO THE HOSPITAL FOR ASPIRATION. NO FLUID WAS EXTRACTED. THE PT ASKED TO FILE A CLAIM FOR THE REVISION SURGERY. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01805 | 2012-10-10 | 2011-01-20 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE/ABGII HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT IS EXPERIENCING BLEEDING, SWELLING AND PAIN AT THE JOINT. PT STATED THAT THE JOINT WAS DRAINED. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01801 | 2012-10-10 | 2012-02-02 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT HIP - REJUVENATE/ABGII | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT STATES THAT SHE HAS BEEN HAVING CONSTANT PAIN SINCE THE IMPLANT SURGERY WHICH SHE NEVER HAD PRIOR TO THE SURGERY. THE PAIN OCCURS WHEN SHE TRIES TO STAND FROM A SITTING POSITION. PAIN IS DOWN THE SIDE OF HER THIGH AND IN HER BUTTOCKS. SHE CANNOT LAY ON HER RIGHT SIDE. THE PAIN HAS DIMINISHED SOME OVER THE PAST MONTHS. SHE ALSO STATES THAT SHE HAD INTENSE GROIN AND THIGH PAIN 8 WEEKS AFTER THE SURGERY. SHE HAD X-RAYS TAKEN. AFTER REVIEWING THE X-RAYS, HER DR. TOLD HER THAT THE IMPLANT CUP MIGHT BE LOOSE AND SHE MAY NEED REVISION SURGERY TO STABILIZE IT IF SHE DID NOT STOP BEING SO ACTIVE. HE FURTHER ADVISED HER TO LIMIT OR STOP MOST OF HER PHYSICAL ACTIVITIES. PT HAD PHYSICAL THERAPY AND DID NOT HAVE PROBLEMS WITH MOST OF THE EXERCISES. THE PT IS DISSATISFIED WITH HER ORIGINAL SURGEON. SHE HAD A FOLLOW-UP X-RAY AT THE 3 MONTH MARK AT WHICH TIME THE SURGEON SAID THE POSITIONING OF THE CUP HAD IMPROVED. SHE CHANGED DOCTORS. DR. P. INFORMED HER OF THE REJUVENATE/ABGII RECALL. PT TOOK NEW X-RAYS. DR. CONCLUDED THAT THE IMPLANT IS FINE. HE ALSO REVIEWED RECENT BLOOD TEST RESULTS FROM (B)(6) 2012 AND CONCLUDED THAT HER RESULTS WERE FINE. HE ALSO STATED THAT THE PAIN SHE IS EXPERIENCING MIGHT BE RELATED TO A BACK CONDITION WHICH SHOWED UP ON THE X-RAYS. PT STATES THAT SHE HAS DEGENERATE JOINT DISEASE IN SEVERAL VERTEBRAE. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01834 | 2012-10-10 | 2012-05-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT; PATIENT DISLOCATED HIP 4 TIMES AFTER THE FIRST SURGERY. THE FIRST DISLOCATION HAPPENED ON (B)(6) 2012. PATIENT STATED THAT A REVISION SURGERY WAS PERFORMED ON (B)(6) 2012. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01814 | 2012-10-10 | 2010-05-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN - LEFT REJUVENATE/ABG II HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT HAS BEEN EXPERIENCING PAIN SINCE THE SURGERY IN THE LEFT LEG GOING UP THE THIGH TO THE HIP. PATIENT STATED THAT SHE HAS TROUBLE WALKING ON HARD SURFACES, WALKING UP AND DOWN THE STAIRS, AND NEEDS TO HOLD ON TO SOMETHING WHILE DOING SO. PATIENT IS ALSO EXPERIENCING PAIN WHILE LYING DOWN AND TENDERNESS IN THE LEG. PATIENT STATED THAT HER COBALT BLOOD LEVEL IS 23, AND THE CHROMIUM LEVELS ARE 6.2. PATIENT STATED THAT THE RIGHT KNEE IS SWOLLEN AND ASPIRATION ON THE RIGHT HIP WAS DONE, FLUID WAS TAKEN OUT OF IT. PATIENT STATED THAT SHE IS SCHEDULED FOR REVISION SURGERY ON THE LEFT HIP (B)(6) 2012. PATIENT STATED THAT HER SURGEON WILL PERFORM A REVISION SURGERY ON THE RIGHT HIP THREE MONTHS AFTER THE REVISION ON THE LEFT HIP. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01822 | 2012-10-10 | 2010-02-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT HIP REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT; PT IS BEEN EXPERIENCING PAIN IN LEG, UP THE HIP. PT REPORTED THAT MRI SHOWED A TUMOR "THE SIZE OF A MEDIUM POTATO". PT STATED THAT HER COBALT BLOOD LEVELS WERE ELEVATED (6.4). TRIPLE PHASE BONE SCAN WAS PERFORMED AND SHE HAD 17 CC OF FLUID; FLUID WAS TAKEN OUT OF JOINT. DR. (B)(6) SCHEDULED A REVISION SURGERY. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCE IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOMES AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01813 | 2012-10-10 | 2012-04-02 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORHTOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT; PT COMPLAINS OF PAIN AND SWELLING AT JOIN SITE. HE HAS BEEN HAVING HEADACHES, GI UPSET, LOSS OF BALANCE AND INTERMITTENT INABILITY TO BEAR WEIGHT ON HIS LEFT HIP FOR 2-3 MONTHS. HE HAS EXTREME FATIGUE AND HIS SURGEON HAS SCHEDULED HIM FOR AN UPCOMING REVISION SURGERY. MRI AND BLOOD TESTS HAVE SHOWN INCREASED COBALT LEVELS AND PT HAS BEEN TOLD THAT HE HAS "COBALT TOXICITY". | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCE IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOMES AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01817 | 2012-10-10 | 2010-02-07 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP SYSTEM | IMPLANT | STRYKER ORHTOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: PT IS STILL EXPERIENCING DIFFICULTY LIFTING AND SORENESS AFTERWARDS FOR MORE THAN THREE DAYS. PT STATED THAT PAIN COMES AND GOES. HE IS HAVING DIFFICULTY CLIMBING STAIRS, RIDING A FOUR WHEEL, HIKING AND RIDDING HORSES. PT STATED THAT HE IS CURRENTLY ON MELOXICAM. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCE IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOMES AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01819 | 2012-10-10 | 2012-03-27 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP SYSTEM | IMPLANT | STRYKER ORHTOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: PT IS EXPERIENCING PAIN, SWELLING, DIFFICULTY STANDING AND "FEELS LIKE SOMETHING IS PUSHING INTO THE PELVIC AREA, LIKE IT IS GOING TO DISLOCATE." PT STATED THAT HE HAS HAD MANY INFECTIONS IN THE IMPLANT SINCE ONE MONTH FROM THE SURGERY, INCLUDING A STAPH INFECTION. PT STATED THAT SINCE LAST MONTH THE PAIN HAS INCREASED AND THAT "THE AREA OF INCISION IS NOT NORMAL." PT STATED THAT HE HAD BLOOD WORK DOWN AND THE LEVELS FOR COBALT ARE HIGH; HOWEVER, HE IS STILL WAITING FOR THE OTHER BLOOD TESTS RESULTS. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCE IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOMES AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01812 | 2012-10-10 | 2012-09-19 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 11 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-110000S |  | IT WAS REPORTED THAT, ADVERSE LOCAL TISSUE REACTION. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS: MFR # 9616680-2012-00862. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01797 | 2012-10-10 | 2012-09-18 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP | IMPLANT | STRYKER ORHTOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT PRESENTED SURGEON WITH HIP PAIN. SURGEON TOOK XRAYS AND NOTICED THAT SOME SORT OF METALLOSIS MAY BE OCCURRING. BECAUSE OF THE AREA OF THE PT'S PAIN, SURGEON SUSPECTS THAT METALLOSIS IS VERY PROBABLE. PT WAS DOING FINE SINCE IMPLANTATION UP UNTIL 3-4 WEEKS AGO WHEN PAIN STARTED. SURGEON IS PLANNING ON REVISING. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCE IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOMES AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01809 | 2012-10-10 | 2011-02-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP COMPONENTS | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PATIENT ALLEGEDLY SUSTAINED PERSONAL INJURIES ON (B)(6) 2011 FROM THE IMPLANTATION OF STRYKER REJUVENATE HIP COMPONENTS. | THE INFORMATION IN THIS REPORT WAS PROVIDED BY STRYKER ORTHOPAEDICS LEGAL AFFAIRS. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. THEREFORE, THE EXACT CAUSE OF THE REPORTED EVENT CANNOT BE DETERMINED. | Y | N | OTHER |  | 2012-10
2249697-2012-01831 | 2012-10-10 | 2012-09-20 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT, PATIENT DEVELOPED PAINFUL JOINT. | THE LRG TAP PRI MOD NCK 8DEG 34MM, CAT# NLV-340800Y, LOT# 35231202 WAS ALSO LISTED IN THIS REPORT. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED PATIENT'S PAIN. AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01816 | 2012-10-10 | 2010-11-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE/ABG II HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT IS REPORTED THAT THE PATIENT COMPLAINED OF HAVING CONTINUOUS DULL PAIN IN HIS LEFT HIP SINCE THE SURGERY. HE WAS CONTACTED BY HIS SURGEON REGARDING THE RECALL. HE THEN SCHEDULED AN APPOINTMENT FOR (B)(6) 2012. HE HAD AN MRI AND BLOOD WORK ON (B)(6) 2012. HE STATES, HE WAS ADVISED BY HIS DOCTOR THAT HIS BLOOD LEVELS ARE ELEVATED AND THERE IS SOME INDICATION OF INFLAMMATION. HE STATES HIS DOCTOR SAYS HE WILL CONTINUE TO MONITOR HIS CONDITION. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01810 | 2012-10-10 | 2010-10-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN - RIGHT REJUVENATE/ABG II HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT HAS BEEN EXPERIENCING PAIN SINCE THE SURGERY IN THE LEFT LEG GOING UP THE THIGH TO THE HIP. PT STATED THAT SHE HAS TROUBLE WALKING ON HARD SURFACES, WALKING UP AND DOWN THE STAIRS, AND NEEDS TO HOLD ON TO SOMETHING WHILE DOING SO. PT IS ALSO EXPERIENCING PAIN WHILE LYING DOWN AND TENDERNESS IN THE LEG. PT STATED THAT HER COBALT BLOOD LEVEL IS 23, AND THE CHROMIUM LEVELS ARE 6.2. PT STATED THAT THE RIGHT KNEE IS SWOLLEN AND ASPIRATION ON THE RIGHT HIP WAS DONE, FLUID WAS TAKEN OUT. PT STATED THAT SHE IS SCHEDULED FOR REVISION SURGERY ON THE LEFT HIP (B)(6) 2012. PT STATED THAT HER SURGEON WILL PERFORM A REVISION SURGERY ON THE RIGHT HIP THREE MONTHS AFTER THE REVISION ON THE LEFT HIP. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01804 | 2012-10-10 | 2011-04-20 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE/ABGII RIGHT HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | PT STATES THAT SHE HAS MINOR DISCOMFORT IN THE RIGHT HIP. SHE WENT TO THE DOCTOR ON (B)(6) 2012. SHE HAD BLOOD LABS DONE ON (B)(6) 2012. SHE RECEIVED THE RESULTS FROM HER DOCTOR'S OFFICE ON (B)(6) 2012, SAYING THAT HER LEVELS WERE ELEVATED. SHE IS BEING SCHEDULED FOR AN MRI ON HER RIGHT HIP THROUGH THE DOCTOR'S OFFICE. THE PT STATES THAT SHE IS HAVING TROUBLE MANEUVERING STAIRS. SHE WAS GIVEN A PRESCRIPTION FOR AN ANTI-INFLAMMATORY MEDICATION ON (B)(6) 2012. SHE HAS BEEN EXPERIENCING OCCASIONAL NERVE TINGLING IN THE HIP SINCE (B)(6) 2012. PT STATES THAT THE ISSUES IN HER RIGHT HIP ARE MINOR AND INFREQUENT. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01815 | 2012-10-10 | 2010-02-02 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP SYSTEM | IMPLANT | STRYKER ORHTOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT; PT EXPERIENCED PAIN, SWELLING, DIFFICULTY WALKING, DIFFICULTY CLIMBING STIRS AND DIFFICULTY LIFTING. PT STATED THAT THE REASON FOR REVISION SURGERY WAS "THE THING THAT GOES IN THE FEMUR WAS LOOSE" PT HAD REVISION SURGERY ON (B)(6) 2010. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01795 | 2012-10-10 | 2012-09-18 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT PT HAS HIP PAIN. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0DEG 34MM, CAT# NLS-340000B, LOT# 33809301; 28MM STD LFIT V40 HEAD, CAT# 6260-9-128, LOT# 35367003; RESTORATION ADM X3 INS 28/48, CAT# 1236-2-848, LOT# 34374001; RESTORATION (TM) ADM. CUP W/HA, CAT# 1235-2-482, LOT# G2949823. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR DUE TO HOSP POLICY. ADDITIONAL INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01835 | 2012-10-10 | 2012-09-18 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT, PATIENT WAS REVISED FOR INFECTION, ALSO TESTING (B)(6). SURGEON STATE THE PATIENT COBALT LEVEL WERE VERY HIGH AND SURGEON STATED HE WAS CONCERN ABOUT AN ADVERSE REACTION TO THE REJUVENATE CONSTRUCT BECAUSE OF THE STAPH POSITIVE, ANTIBIOTIC SPACER WERE PLACED. | THE FOLLOWING OTHER DEVICES WERE ALSO LISTED IN THIS REPORT: RESTORATION ADM X3 INS 28/48, CAT# 1236-2-848, LOT# 34022301. DELTA V-40 CERAMIC HEAD 28/0, CAT#6570-0-128, LOT# 36665503. TRIDENT HEMISPHERICAL CLUSTER HOLE SHELL, CAT#502-11-52E, LOT# 3689601. MODULAR DUAL MOBILITY INSERT, CAT# 626-00-42E, LOT#37249201, 6.5 CANCELLOUS BONE SCREW 25MM, CAT# 2030-6525-1, LOT# MKHM6M, 6.5 CANCELLOUS BONE SCREW 25MM, CAT# 2030-6525-1, LOT# MKHRPH. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S INFECTION. AN EVALUATION OF THE DEVICES CANNOT BE PERFORMED AS THE REVISED DEVICES WERE RETAINED BY THE HOSPITAL AND WERE NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PATIENT/EVENT AS: MFR# 9616680-2012-00878. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01847 | 2012-10-12 | 2012-07-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE/ABG II HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PT STATES THAT SHE RECEIVED THE RECALL NOTIFICATION FROM HER SURGEON. SHE STATES THAT SHE RECENTLY HAS BEGUN EXPERIENCING MILD TWINGES OF PAIN IN HER RIGHT HIP WHEN SHE STANDS FROM SITTING. SHE STATES THAT SHE CURRENTLY HAS NO OTHER SYMPTOMS. THE PT HAD X-RAYS TAKEN IN (B)(6) 2012 DURING HER STANDARD F/U VISIT AT WHICH TIME HER PHYSICIAN TOLD HER THE X-RAYS WERE FINE. SHE WAS NOT EXPERIENCING PAIN AT THE TIME. THE PT IS SCHEDULED TO SEE HER SURGEON ON (B)(6) 2012 FOR FURTHER EXAMINATION. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01916 | 2012-10-17 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER DR (B)(6) OFFICE, THIS PATIENT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01885 | 2012-10-17 | 2009-12-14 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE/ABGII HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT HAD BILATERAL HIP SURGERIES ON (B)(6) 2009. SINCE THEN, PATIENT COMPLAINS OF ON AND OFF PAIN IN BOTH HIPS. MRI AND TESTING HAVE SHOWN FLUID BUILD UP IN BOTH HIP JOINTS. THE TOXIC LIQUID HAS CAUSED THE BONE TO DETERIORATE. PATIENT WILL UNDERGO ASPIRATION OF BOTH HIPS NEXT WEEK. HIS SURGEON HAS SCHEDULED HIM FOR REVISION SURGERY ON (B)(6) 2012 FOR BOTH HIPS. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01887 | 2012-10-17 | 2012-06-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT HAS A HARD TIME GOING UP STEPS AND WALKING DUE TO PAIN. SHE ALSO STATES THAT SHE CANNOT PUT PRESSURE ON THE LEFT LEG OR LIE ON IT WHEN SLEEPING. SHE HAS BEEN HAVING THE PROBLEM FOR QUITE SOME TIME, BUT CANNOT PINPOINT THE SPECIFIC ONSET. THE PATIENT COMPLETED BLOOD WORK AND HAD X-RAYS DURING HER DOCTOR'S APPOINTMENT ON (B)(6) 2012. SHE IS SCHEDULED TO HAVE AN MRI NEXT WEEK. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01924 | 2012-10-17 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER DR. (B)(6) OFFICE, THIS PATIENT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01923 | 2012-10-17 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER DR (B)(6) OFFICE, THIS PATIENT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS AND MRI. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01902 | 2012-10-17 | 2010-10-25 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE/ABG II HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT HAS CONSISTENTLY EXPERIENCED PAIN IN HIS HIP SINCE THE SURGERY. HE HAS UNDERGONE PHYSICAL THERAPY FOR MONTHS WHICH HAS NOT RELIEVED HIS PAIN. THE PATIENT STATES THAT HE CANNOT LIE ON HIS LEFT SIDE WITHOUT PAIN. THE PATIENT ALSO STATES THAT HE HAS BEEN COMPLAINING TO HIS SURGEON FOR A LONG TIME ABOUT THE PAIN AND DISCOMFORT. THE PATIENT'S SURGEON IS SCHEDULING HIS APPOINTMENTS FOR MRI, BLOOD WORK, AND OFFICE VISIT WITHIN THE NEXT FEW WEEKS. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01914 | 2012-10-17 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER DR (B)(6) OFFICE, THIS PATIENT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01892 | 2012-10-17 | 2012-06-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 10 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-100000S |  | ADDITIONAL INFORMATION RECEIVED: PLAINTIFF ALLEGES THE LEFT REJUVENATE MODULAR HIP STEM IMPLANTED ON OR ABOUT (B)(6) 2011 FAILED, CAUSING EXCESSIVE LEVELS OF CHROMIUM AND COBALT IN HIS BLOOD. PLAINTIFF HAS NOT YET SCHEDULED A REVISION SURGERY OF THE LEFT HIP STEM. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01891 | 2012-10-17 | 2012-06-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE/ABG II HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT IS REPORTED THAT FOR THE PAST FEW MONTHS THE PATIENT HAS EXPERIENCED DISCOMFORT THAT EXTENDS FROM HIS UPPER THIGH INTO THE HIP. THE PATIENT HAS NOT UNDERGONE ANY TESTING AT THIS TIME. HE IS SCHEDULED FOR AN APPOINTMENT WITH HIS SURGEON ON (B)(6) 2012. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01918 | 2012-10-17 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER DR (B)(6) OFFICE, THIS PATIENT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS AND MRI. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01913 | 2012-10-17 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER DR (B)(6) OFFICE, THE PATIENT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01917 | 2012-10-17 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER DR (B)(6) OFFICE, THIS PATIENT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01879 | 2012-10-17 | 2012-07-31 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE/ABG II HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT COMPLAINS OF PAIN AND SWELLING IN THE JOINT SITE. ULTRASOUND REVEALED PROSTHESIS WAS DETERIORATED AND PATIENT REQUIRED A REVISION SURGERY. PATIENT REPORTED THAT COBALT LEVELS WERE LEAKING INTO HER SYSTEM AND BLOOD STREAM. PATIENT HAD REVISION SURGERY ON (B)(6) 2012. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01921 | 2012-10-17 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER DR (B)(6) OFFICE, THIS PATIENT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01908 | 2012-10-17 | 2012-04-12 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE/ABG II HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PATIENT IS SUFFERING MINOR PAIN AND STIFFNESS. PATIENT FEELS THAT SHE DOESN'T HAVE COMPLETE MOTION AS SHE DID BEFORE. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01911 | 2012-10-17 | 2011-09-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE/ABG II HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT STATES THAT HER RIGHT HIP ACHES BADLY GIVING HER A PRONOUNCED LIMP. SHE STATES THAT SINE (B)(6) 2012, HER HIP IS ALWAYS PAINFUL AND WARM TO THE TOUCH. THE PATIENT HAD A DOCTORS APPOINTMENT ON (B)(6) 2012. THE PATIENT IS SCHEDULED TO COMPLETE BLOOD WORK AND AN MRI ON (B)(6) 2012. THE DOCTOR HAS STATED THAT HE SUSPECTS SHE WILL NEED A REVISION BUT IT CANNOT BE CONFIRMED UNTIL THE RESULTS OF HER TEST ARE IN. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01909 | 2012-10-17 | 2010-11-29 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE/ABG II HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT STATES THAT HE HAS STIFFNESS AND A PULLING PAIN IN THE FRONT OF THE RIGHT THIGH, THE GROIN, AND PAIN ON THE OUTER PART OF HIS WHICH EXTENDS INTO HIS BACK. HE ALSO HAS SOME SWELLING IN THE HIP AND THIGH. HE HAS A LOT OF PAIN WHEN HE STANDS FROM A SITTING POSITION. HE OCCASIONALLY FEELS SHARP PAIN IN THE GROIN AREA WHEN HE LIES DOWN. HE STATES THAT HE CAN'T WALK MORE THAN TWO BLOCKS WITHOUT INTENSE PAIN. IN (B)(6), HIS SURGEON ORDERED X-RAYS AND A CT SCAN. HE HAS HAD BLOOD DRAWN BUT IS NOT AWARE OF THE RESULTS. HIS DOCTOR DIAGNOSED THAT A REVISION SURGEON IS NECESSARY. HIS SURGERY IS SCHEDULED FOR (B)(6) 2012. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01880 | 2012-10-17 | 2012-06-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE/ABG II HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT IS REPORTED THAT THE PATIENT HAS BEEN HAVING A BURNING PAIN IN THE RIGHT HIP AT THE SITE OF THE IMPLANT SINCE (B)(6) 2012. IN (B)(6) 2012, SHE WAS DIAGNOSED AS HAVING INFLAMMATION IN THE RIGHT HIP. SHE RECEIVED A STEROID SHOT, A PRESCRIPTION FOR MOBIC WHICH SHE TAKES FOR INFLAMMATION AND PAIN, AND HAD 4 WEEKS OF PHYSICAL THERAPY. THE PATIENT STATES THAT THESE THERAPIES HAVE HELPED SOMEWHAT TO RELIEVE THE PAIN AND SWELLING. SHE WAS TOLD THAT HER CHROMIUM LEVEL AND COBALT LEVELS ARE ELEVATED. SHE HAS SEEN BOTH HER SURGEON AND PRIMARY CARE PHYSICIANS. RECENT X-RAYS INDICATED THAT THE HIP IS STILL INTACT. THE PATIENT STATES THAT HER SURGEON WILL MONITOR HER CONDITION WITH REPEAT TESTING EVERY 6 MONTHS. THE PATIENT STATES THAT SHE IS VERY ACTIVE AND EXERCISES REGULARLY 1 HOUR, 3 TIMES A WEEK. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01881 | 2012-10-17 | 2011-12-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE/ABG II HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT IS EXPERIENCING OCCASIONAL PAIN WHILE WALKING. PATIENT STATED THAT HER BLOOD WORK SHOWS HIGH COBALT LEVELS AT 8.5 AND THE MRI SHOWS THAT THERE ARE METAL FRAGMENTS. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01882 | 2012-10-17 | 2012-06-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE/ABG II HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT STATES THAT SHE BEGAN TO EXPERIENCE PAIN AND DISCOMFORT IN HER RIGHT HIP ABOUT 2-3 MONTHS AGO. THE PAIN BEGAN SUBTLY BUT HAS INCREASED IN INTENSITY AND FREQUENCY OVER THE PAST FEW MONTHS. THE PATIENT STATES THAT SHE CANNOT STAND FOR MORE THAN 20 MINUTES WITHOUT EXPERIENCING PAIN IN THE FRONT THIGH MUSCLE IN HER RIGHT LEG, ON THE SIDE OF HER RIGHT HIP AND IN THE BUTTOCKS. SHE STATES THAT SHE SEES SOME VISIBLE SWELLING AT THE TOP OF THE IMPLANT INCISION. SHE HAS TROUBLE GETTING UP FROM A CHAIR AND OUT OFF A CAR. SHE CAN'T LIE ON HER RIGHT SIDE FOR LONG WITHOUT PAIN. WHEN THE PAIN GETS INTENSE SHE MUST LIE DOWN AND TAKE ASPIRIN. SHE TRIES TO CONTINUE LIGHT EXERCISE BY DAILY GOING FOR SLOW WALKS. THE PATIENT IS CURRENTLY WAITING FOR THE DOCTOR'S STAFF TO SCHEDULE HER FOR AN APPOINTMENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. |  | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01922 | 2012-10-17 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER DR (B)(6)'S OFFICE, THIS PATIENT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01930 | 2012-10-17 | 2012-09-26 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 10 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-100000S |  | IT WAS REPORTED THAT, NO SPECIFIC EVENT OCCURRED. PATIENT HAD ONGOING PAIN SINCE PRIMARY IMPLANT. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 16DEG 42MM, CAT# NLS-421600P, LOT# 29438701. DELTA V-40 CERAMIC HEAD 36/+2,5, CAT# 6570-0-536, LOT# 33072401. TRIDENT 10DEG X3 INSERT 36MM ID, CAT# 623-10-36G, LOT# MHPX9N. AT THIS TIME, IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S PAIN. AN EVALUATION OF THE REPORTED DEVICE(S) CANNOT BE PERFORMED AS THE DEVICE(S) WAS RETAINED BY THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01860 | 2012-10-17 | 2011-09-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | UNKNOWN | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT THE PT BEGAN EXPERIENCING PAIN AND DISCOMFORT APPROX ONE YR AGO. SHE STATES THAT SHE HAS PAIN IN HER RIGHT HIP AFTER WALKING MORE THAN THREE BLOCKS. SHE SAW HER DOCTOR IN (B)(6) 2011 AND X-RAYS WERE TAKEN. THE PT STATED THAT HER DOCTOR SAID THERE WAS A POSSIBILITY THAT THE IMPLANT MIGHT BE LOOSE OR THAT HER PAIN WAS CAUSED BY ARTHRITIS. SHE SAYS, THE DOCTOR TOLD HER TO MONITOR HER HIP PAIN AND RETURN TO HIM IF THE PAIN PERSISTED. NO ADD'L TREATMENT WAS PRESCRIBED. SHE SAYS THAT SHE IS ALREADY TAKING PRESCRIPTION PAIN MEDICATION FOR OTHER CONDITIONS. SHE STATES THAT HER PAIN MEDICATION LESSENS BUT DOES NOT ELIMINATE THE PAIN IN HER HIP. THE PT HAS NOT YET SCHEDULED AN APPOINTMENT WITH HER SURGEON BUT SAYS SHE WILL DO SO IN THE NEAR FUTURE. | ADD'L INFO HAS BEEN REQUESTED AND IF REC'D WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. CATALOG NUMBERS AND LOT CODES OF OTHER DEVICES LISTED IN THIS REPORT: CAT # NLV-300800Y LOT # 24945202 DESCRIPTION: REJUVENATE MODULAR NECK, CAT # UH1-48-26 LOT # MHRREM DESCRIPTION: UHR BIPOLAR 26X48MM, CAT # 6260-9-126 LOT # 32933402 DESCRIPTION: 26MM STD LFIT V40 HEAD. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01915 | 2012-10-17 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER DR (B)(6) OFFICE, THIS PATIENT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS AND MRI. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01920 | 2012-10-17 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER DR. OFFICE, THIS PATIENT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01895 | 2012-10-17 | 2012-02-03 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE/ABG II HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT STATES THAT SHE VISITED HER SURGEON ON (B)(6) 2012 AND X-RAYS OF HER RIGHT HIP WERE TAKEN. SHE STATES THAT ACCORDING TO HER DOCTOR, THE IMPLANT APPEARS FINE IN X-RAY. SHE STATES THAT SHE HAS TENDERNESS IN THE RIGHT HIP AND EXPERIENCES LIMITED RANGE OF MOTION AND PAIN WHEN LIFTING HER LEG. SHE ALSO HAS SOME SLIGHT BRUISING AND DISCOLORATION AROUND THE IMPLANT SITE. SHE CAN'T LIE ON HER RIGHT SIDE WITHOUT HER LEG GOING TO SLEEP. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01896 | 2012-10-17 | 2012-01-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE/ABG II HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT IS EXPERIENCING PAIN IN THE FRONT OF THE THIGH ON THE LEFT SIDE. THE PAIN BEGAN AROUND (B)(6) 2012. THE PATIENT IS SCHEDULED FOR AN MRI AND BLOOD TESTS ON (B)(6) 2012. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01886 | 2012-10-17 | 2009-12-14 | I, n, j, u, r, y | Manufacturer report | UNK RIGHT REJUVENATE/ABG II HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT HAD BILATERAL HIP SURGERIES ON (B)(6) 2009. SINCE THEN, PT COMPLAINS OF ON AND OFF PAIN IN BOTH HIPS. MRI AND TESTING HAVE SHOWN FLUID BUILD UP IN BOTH HIP JOINTS. THE TOXIC LIQUID HAS CAUSED THE BONE TO DETERIORATE. PT WILL UNDERGO ASPIRATION OF BOTH HIPS NEXT WEEK. HIS SURGEON HAS SCHEDULED HIM FOR REVISION SURGERY ON (B)(6) 2012 FOR BOTH HIPS. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER, BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01866 | 2012-10-17 | 2010-12-21 | I, n, j, u, r, y | Manufacturer report | UNK LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT; PT COMPLAINS OF PAIN IN LEFT HIP. PAIN INCREASES WITH WALKING. PT HAS TROUBLE SLEEPING AND SORENESS IN LEFT HIP ON AND OFF. PT STATES THAT IT'S PAINFUL TO STAND. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01856 | 2012-10-17 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER THE SURGEON'S OFFICE, THIS PT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-10
2249697-2012-01901 | 2012-10-17 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER THE SURGEON'S OFFICE, THIS PT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-10
2249697-2012-01869 | 2012-10-17 | 2012-04-17 | I, n, j, u, r, y | Manufacturer report | UNK LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT; PT IS EXPERIENCING PAIN IN THE LEFT HIP AREA. PT STATED THAT SHE CANNOT STAND UP WITHOUT USING A CANE AND USES A WALKER TO GET AROUND. PT STATED THAT SHE FALLS DOWN EASILY. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01865 | 2012-10-17 | 2012-09-25 | I, n, j, u, r, y | Manufacturer report | UNK REJUVENATE STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE SURGEON REVISED A REJUVENATE STEM AND ADM CUP ON PT AND REPLACED WITH DEPUY. NO REASON GIVEN FOR THE REVISION. | THE UNK ADM CUP WAS ALSO LISTED IN THIS REPORT. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S REVISION. AN EVAL OF THE DEVICE(S) CANNOT BE PERFORMED AS THE DEVICE(S) WAS NOT RETURNED TO THE MFR DUE TO HOSPITAL POLICY. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01863 | 2012-10-17 | 2010-09-12 | I, n, j, u, r, y | Manufacturer report | UNK RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT; PT COMPLAINS OF PAIN IN RIGHT HIP. PAIN INCREASES WITH WALKING. PT HAS TROUBLE SLEEPING AND SORENESS IN RIGHT HIP ON AND OFF. PT STATES THAT IT'S PAINFUL TO STAND. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
9616680-2012-00885 | 2012-10-17 | 2011-05-17 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK 34MM | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | NLS-340000B |  | IT WAS REPORTED THAT PT HAD PAIN SINCE SURGERY. PT STATED THAT A REVISION WAS PERFORMED ON (B)(6) 2012 AND A NON STRYKER HIP WAS PUT IN PLACE. PT STATED THAT THE STRYKER HIP WAS "CORRODED." | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD DEVICE OR ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS: MFR # 2249697-2012-01871. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01905 | 2012-10-17 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER THE SURGEON'S OFFICE, THIS PT IS EXPERIENCING DIFFICULTIES WITH HER HIP IMPLANT. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-10
2249697-2012-01893 | 2012-10-17 | 2012-05-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE/ABG II HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT'S RIGHT LEG AND HIP HURTS WHEN HE WALKS FOR MORE THAN TEN MINUTES. HE ALSO STATES THAT HE HAS PAIN WHENEVER HE ATTEMPTS ANY WEIGHT BEARING ACTIVITIES AND HAS TO TAKE TYLENOL IMMEDIATELY AFTERWARDS. THE PAIN BEGAN APPROXIMATELY FIVE MONTHS AGO AND HAS INCREASED IN SEVERITY OVER TIME. THE PATIENT SAW HIS SURGEON LAST WEEK. HE HAS UNDERGONE BLOOD TEST FOR CHROMIUM AND COBALT LEVELS, TAKEN X-RAYS AND AN MRI. HIS DOCTOR STATED THAT THE MRI SHOWED INFLAMMATION IN THE IMPLANT AREA AND A PSEUDO-TUMOR. THE SURGEON ADVISES THAT HE HAVE A REVISION SURGERY AS SOON AS POSSIBLE. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01873 | 2012-10-17 | 2011-04-01 | I, n, j, u, r, y | Manufacturer report | UNK RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT; PT IS EXPERIENCING DISCOMFORT IN THE HIP AREA. PT'S BLOOD COBALT WAS ELEVATED; METAL ION LEVELS WERE ALSO ELEVATED. PT STATED THAT SHE WENT TO DOCTORS OFFICE ON (B)(6) 2012 X-RAYS WERE DONE. PT STATED THAT SHE IS GOING FOR A MRI IN THE NEAR FUTURE. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01872 | 2012-10-17 | 2012-09-01 | I, n, j, u, r, y | Manufacturer report | UNK LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT; PT IS EXPERIENCING PAIN SITTING AND STANDING. PT STATED THAT THE PAIN STARTED LAST WEEK. PT IS F/U WITH HER DOCTOR. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01877 | 2012-10-17 | 2009-12-28 | I, n, j, u, r, y | Manufacturer report | UNK LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT; PT IS EXPERIENCING PAIN AT THE LEFT HIP ONCE IN AWHILE; AND PAIN IN THE LEFT KNEE. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01868 | 2012-10-17 | 2011-09-06 | I, n, j, u, r, y | Manufacturer report | UNK LFT HIP REJUVENATE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT; PT IS EXPERIENCING PAIN WHILE SITTING AND WALKING. PT STATED THAT SHE HAD TWO BLOOD TESTS AND A HIP ASPIRATION. PT STATED THAT SHE HAD A BONE SCAN. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01878 | 2012-10-17 | 2012-08-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT; PT IS EXPERIENCING PAIN IN HIP AND LIMPING SINCE LAST MONTH. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: REJUVENATE MODULAR NECK 34MM V40, CAT #NLS0340000B, LOT# 3603301; TRIDENT HEMISPHERICAL CLUSTER HOLE SHELL, CAT #502-11-52E, LOT #36380402; TRIDENT 0DEG X3 INSERT 36MM ID, CAT #623-00-36E, LOT # MKE0X8; DELTA V-40 CERAMIC HEAD 36/-5, CAT #6570-0-036, LOT #36674302; 6.5 CANCELLOUS BONE SCREW 25MM, CAT #2030-6525-1, LOT #MKHRPH. AT THIS TIME, IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
9616680-2012-00892 | 2012-10-17 | 2010-12-16 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK 30MM V40 | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | NLS-300000B |  | IT WAS REPORTED THAT: THE PT CLAIMS SHE HURTS, HAS SOFTLY SWELLING, AND HER BLOOD WORK IS ELEVATED (PROBABLY COBALT LEVELS). | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. HOWEVER, SOME X-RAYS AND MEDICAL RECORDS WERE RECEIVED. WHEN COMPLETED, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS: 2249697-2012-01432. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01870 | 2012-10-17 | 2011-05-17 | I, n, j, u, r, y | Manufacturer report | UNK RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT; PT IS EXPERIENCING PAIN IN THE RIGHT HIP AREA DOWN TO THE KNEE. PT STATED THAT SHE HAS DIFFICULTY BENDING HER RIGHT KNEE; SHE CANNOT STAND UP WITHOUT USING A CANE AND USES A WALKER. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01871 | 2012-10-17 | 2011-05-17 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 10 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-100000S |  | IT WAS REPORTED THAT; PT HAD PAIN SINCE SURGERY. PT STATED THAT A REVISION WAS PERFORMED ON (B)(6) 2012 AND A NON STRYKER HIP WAS PUT IN PLACE. PT STATED THAT THE STRYKER HIP WAS "CORRODED". | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD DEVICE OR ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS: MFR# 9616680-2012-00885. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01857 | 2012-10-17 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER THE SURGEON'S OFFICE, THIS PT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS AND MRI. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-10
2249697-2012-01855 | 2012-10-17 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER THE SURGEON'S OFFICE, THIS PT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-10
2249697-2012-01858 | 2012-10-17 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER THE SURGEON'S OFFICE, THIS PT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-10
2249697-2012-01890 | 2012-10-17 | 2010-09-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE/ABG II HIP | IMPLANT | STRYKER ORTHOPAEDICS | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT BEGAN COMPLAINING OF STIFFNESS AND OCCASIONAL PAIN IN THE RIGHT HIP APPROXIMATELY ONE YEAR AFTER HER IMPLANT SURGERY. SHE HAS HAD PHYSICAL THERAPY AND TRIES YOGA AT HER DOCTOR'S RECOMMENDATION, BUT SAYS IT IS PAINFUL AND SHE TAKES IBUPROFEN FOR PAIN RELIEF AFTER EACH SESSION. SHE RECENTLY HAD BLOOD TEST FOR FOLLOW-UP TESTING IN THREE MONTHS. SHE STATES HER SURGEON SAID THAT SHE WILL EVENTUALLY NEED REVISION SURGERY. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01889 | 2012-10-17 | 2011-02-03 | I, n, j, u, r, y | Manufacturer report | UNK LEFT REJUVENATE/ABG II HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT STATES, HE HAS SWELLING IN THE LEFT HIP. HE HAS A CONSTANT DULL ACHE IN THE LEFT HIP AND THIGH. HE SOMETIMES EXPERIENCES SHARP PAINS DOWN THE LEG BETWEEN THE HIP AND THE KNEE. HE HAS A SHARP PAIN WHEN TRYING TO NEGOTIATE STEPS OR GETTING DOWN ON HIS KNEES. ANYTHING THAT PUTS PRESSURE ON THE LEG CAUSES HIS PAINT TO INCREASE. HE ALSO HAS DIFFICULTY GETTING IN AND OUT OF HIS PICKUP TRUCK AND CANNOT BEND OVER TO PUT ON HIS SOCKS. HE IS TAKING PAIN MEDICATION WHEN PAIN INTENSIFIES. HE IS SCHEDULED TO SEE HIS DOCTOR ON (B)(6) 2012. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER, BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01867 | 2012-10-17 | 2010-01-17 | I, n, j, u, r, y | Manufacturer report | UNK LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT; PT COMPLAINS OF PAIN THAT STARTED IN THE PAST OF COUPLE OF WEEKS. PT HAS SWELLING SINCE LAST MONTH AND REPORTS THAT WHEN LAYING DOWN HE CAN FEEL THE SWELLING. MRI SHOWS FLUID BUILD UP. DOCTOR ASPIRATED ON (B)(6) 2012. X-RAY SHOWED BONE DETERIORATION. REVISION SURGERY IS SCHEDULED FOR (B)(6) 2012. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01888 | 2012-10-17 | 2012-03-01 | I, n, j, u, r, y | Manufacturer report | UNK RIGHT REJUVENATE/ABG II HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THE PT STATES THAT SHE HAD BEEN EXPERIENCING PAIN SINCE THE END OF (B)(6) 2012. SHE HAS A SEARING PAIN WHICH SOMETIMES WAKES HER IN THE MIDDLE OF THE NIGHT. THE PAIN BEGINS IN HER HIP AND BEGINS TO RADIATE INTO HER LOWER BACK. WALKING ON HARD SURFACES AND PROLONGED STANDING CAUSE THE PAIN TO INCREASE. THE PAIN IS NOT CONSTANT, HOWEVER, IT IS FREQUENT AND THE FREQUENCY IS INCREASING. SHE HAS LOST AGILITY IN THE RIGHT FOOT. SHE HAS DIFFICULTY LIFTING HER FOOT AND FINDS HERSELF DRAGGING IT. SHE HAS TO LIFT HER FOOT WITH HER HAND TO GET INTO A CAR. SHE WAS GIVEN CORTISONE SHOTS FOR PAIN RELIEF BY HER FAMILY PHYSICIAN IN (B)(6). THE PT'S SURGEON HAS AN APPOINTMENT FOR AN MRI AND BLOOD WORK ON (B)(6) 2012. SHE WILL SEE THE SURGEON ONCE THE TEST RESULTS ARE IN. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER, BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01884 | 2012-10-17 | 2011-12-01 | I, n, j, u, r, y | Manufacturer report | UNK RIGHT REJUVENATE/ABG II HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT STATES THAT SHE HAS BEEN EXPERIENCING EXCRUCIATING PAIN FOR THE PAST 10 MONTHS. ONE DAY, SHE BENT OVER AND EXPERIENCED A SHARP MAJOR PAIN IN HER HIP AND HAS BEEN IN PAIN EVER SINCE. SHE SAYS SHE CONSISTENTLY HAS PAIN WHEN SHE ATTEMPTS TO RISE FROM SITTING AND IS UNABLE TO GO UP AND DOWN STEPS. SHE NOW HAS AN UNUSUAL GAIT WHEN WALKING DUE TO THE PAIN AND WEAKNESS IN HER LEG. PREVIOUS X-RAYS OF THE RIGHT HIP HAVE SHOWN THE IMPLANT TO BE INTACT ACCORDING TO HER DOCTOR'S REPORTING. SHE WAS REFERRED TO THE LOCAL PAIN CLINIC AND WAS TWICE GIVEN CORTISONE INJECTIONS. SHE SAYS THE INJECTIONS HAVE NOT RELIEVED HER HIP PAIN. SHE HAS A PRESCRIPTION FOR HYDROCODONE. THE SURGEON'S OFFICE IS SCHEDULING HER FOR TESTING AND AN OFFICE VISIT NEXT WEEK. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER, BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01883 | 2012-10-17 | 2011-12-16 | I, n, j, u, r, y | Manufacturer report | UNK LEFT REJUVENATE/ABGII HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT IS EXPERIENCING PAIN LAYING DOWN, STANDING, STIFFNESS AND WALKS WITH CANE. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER, BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01900 | 2012-10-17 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS STATED THAT THIS PT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | 0 |  | 2012-10
2249697-2012-01854 | 2012-10-17 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER THE SURGEON'S OFFICE, THIS PT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-10
2249697-2012-01903 | 2012-10-17 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THIS PT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-10
2249697-2012-01906 | 2012-10-17 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THIS PT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-10
2249697-2012-01898 | 2012-10-17 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THIS PT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-10
2249697-2012-01862 | 2012-10-17 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER THE SURGEON¿S OFFICE, THIS PT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | PRODUCT ID, 377775 LOT# N0032985 SERIAL#, IMPLANTED: 2005 (B)(6), EXPLANTED: PRODUCT TYPE LEAD PRODUCT ID, 377775 LOT# N0032985 SERIAL#, IMPLANTED: 2005 (B)(6), EXPLANTED: PRODUCT TYPE LEAD PRODUCT ID, 37742 LOT# SERIAL# (B)(4), IMPLANTED: 2005 (B)(6), EXPLANTED: PRODUCT TYPE PROGRAMMER, PATIENT. (B)(4). | Y | N | OTHER |  | 2012-10
2249697-2012-01864 | 2012-10-17 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER THE SURGEON¿S OFFICE, THIS PT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-10
2249697-2012-01851 | 2012-10-17 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER THE SURGEON'S OFFICE, THIS PT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-10
2249697-2012-01852 | 2012-10-17 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER THE SURGEON'S OFFICE, THIS PT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-10
2249697-2012-01853 | 2012-10-17 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER THE SURGEON'S OFFICE, THIS PT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-10
2249697-2012-01894 | 2012-10-17 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THIS PT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-10
2249697-2012-01849 | 2012-10-17 | 2011-09-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT HIP REJUVENATE/ABG II HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | THE PT REPORTS THAT SHE HAS BEEN HAVING PAIN IN HER LEFT HIP FOR THE PAST YEAR. SOMETIMES THE PAIN SHOOTS FROM THE AREA OF THE IMPLANT DOWN THE LEFT LEG INTO THE CALF. THE IMPLANT AREA OCCASIONALLY FEELS WARM. THE PT STATES THAT SHE CANNOT LIE ON HER LEFT SIDE MORE THAN ONE HALF HR. ON (B)(6) 2012, THE PT HAD BLOOD WORK COMPLETED. SHE STATED THAT HER DOCTOR ADVISED HER THAT NO METAL WAS DETECTED. HER NEXT DOCTOR¿S APPOINTMENT IS SCHEDULED FOR (B)(6) 2012. | AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED BUT BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-10
2249697-2012-01937 | 2012-10-18 | 2012-09-26 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT, COMPONENTS WERE EXPLANTED DUE TO ADVERSE LOCAL TISSUE REACTION. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR DUE TO HOSPITAL POLICY. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED, BUT NOT MADE AVAILABLE. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS: 9616680-2012-00898. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01942 | 2012-10-18 | 2012-09-22 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT PT'S LEFT HIP WILL BE REVISED ON (B)(6) 2012 DUE TO PAIN AND PT HAS INCREASED METAL ION LEVEL/COBALT LEVELS. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD'L DEVICES OR ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS: 9616680-2012-00899. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02022 | 2012-10-23 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER DOCTOR (B)(6)'S OFFICE, THIS PATIENT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WILL BE REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02029 | 2012-10-23 | 2012-09-29 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT, COMPONENTS WERE EXPLANTED DUE TO ADVERSE LOCAL TISSUE REACTION. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR DUE TO HOSPITAL POLICY. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED, BUT NOT MADE AVAILABLE. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS: 9616680-2012-00923. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02031 | 2012-10-23 | 2012-10-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT SURGEON DID A REVISION ON A PATIENT'S RIGHT HIP TO PAIN. SURGEON REMOVED REJUVENATE STEM, NECK, CUP AND LINER. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0DEG 34MM, CAT# NLS-340000B, LOT# 3769001. TRIDENT PSL HA SOLID BACK 52MM, CAT# 540-11-52E, LOT# 36724101. TRIDENT 0/DEG X3 INSERT 36MM ID, AT# 623-00-36E, LOT# MKMKXJ. DELTA V-40 CERAMIC HEAD 36/+2,5, CAT# 6570-0-536, LOT# 37251101. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S PAIN. AN EVALUATION OF THE REPORTED DEVICES CANNOT BE PERFORMED AS THE REVISED DEVICES WERE RETAINED BY THE HOSPITAL AND WERE NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02013 | 2012-10-23 | 2012-01-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER DR (B)(6) OFFICE, THIS PATIENT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS, MRI AND REVISION. | AN EVALUATION OF THE REPORTED DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02015 | 2012-10-23 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | AS PER DOCTOR (B)(6)'S OFFICE, THIS PATIENT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WILL BE REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02019 | 2012-10-23 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER DOCTOR (B)(6)'S OFFICE, THIS PATIENT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WILL BE REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02024 | 2012-10-23 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER DR (B)(6)'S OFFICE, THIS PATIENT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS, MRI AND REVISION. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01960 | 2012-10-23 | 2012-03-08 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE MODULAR HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THE PT COMPLAINS OF DIFFICULTY WALKING, PAIN IN JOINT AREA, AND SORENESS. PT CAN'T BEAR WEIGHT ON HIP JOINT. SHE SAW HER SURGEON TODAY AND WAS ADVISED TO CALL STRYKER. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02020 | 2012-10-23 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER DOCTOR (B)(6) 'S OFFICE, THIS PATIENT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WILL BE REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01955 | 2012-10-23 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNK REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER DOCTOR THE SURGEON'S OFFICE, THIS PT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS AND MRI. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-10
2249697-2012-01968 | 2012-10-23 | 2012-02-01 | I, n, j, u, r, y | Manufacturer report | UNK RIGHT REJUVENATE/ABGII HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT COMPLAINS OF SWELLING AND CONSISTENT PAIN. PT STATES THAT THE PAIN HAS BEEN CONSTANT SINCE THE SURGERY. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER, BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01958 | 2012-10-23 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNK REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER DOCTOR'S OFFICE, THIS PT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-10
2249697-2012-01945 | 2012-10-23 | 2012-09-27 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT THE SURGEON PERFORMED A TOTAL HIP REVISION ON PATIENT'S LEFT HIP DUE TO PAIN. PATIENT HAD A REJUVENATE HIP IMPLANTED (B)(6) 2011 AND EXPERIENCED PAIN. SURGEON NOTICED SEVERE TISSUE DAMAGE AND THERE WAS A SEVERE AMOUNT OF NECROTIC TISSUE TAKEN OUT OF THE HIP CAPSULE. SURGEON ALSO NOTED A LARGE AMOUNT OF METAL DEBRIS AT THE STEM/NECK JUNCTION. SURGEON REVISED WITH RESTORATION MODULAR CONE CONICAL. HOSPITAL POLICY IS THAT NO MANUFACTURER GETS ANY EXPLANTS BUT SALES REP WILL TRY AND FOLLOW UP AS NEEDED TO ATTEMPT TO GET FURTHER INFORMATION, X-RAYS ETC. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS RETAINED BY THE SURGEON AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED, BUT NOT MADE AVAILABLE. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PATIENT/EVENT AS: 9616680-2012-00901. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01944 | 2012-10-23 | 2012-09-27 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT THE SURGEON REVISED THE PATIENT'S LET HIP DUE TO PAIN. MRI SHOWED A PSEUDO TUMOR. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS RETAINED BY THE SURGEON AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED, BUT NOT MADE AVAILABLE. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PATIENT/EVENT AS: 9616680-2012-00900. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01963 | 2012-10-23 | 2010-08-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT COMPLAINS OF PAIN AT JOINT SITE, NUMBNESS IN RIGHT LEG DOWN TO THE FOOT. HE RATES HIS PAIN A 5 OUT OF 10. HE WILL FOLLOW UP WITH SURGEON. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02018 | 2012-10-23 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PER DOCTOR (B)(6)'S OFFICE, THIS PATIENT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS AND MRI. UPDATE: PATIENT STATED THAT BOTH HIPS ARE EXPERIENCING "CLICKING AND POPPING" AND GROIN PAIN. PATIENT IS EXPERIENCING DIFFICULTY LIFTING LEGS GOING UPSTAIRS. PATIENT IS HAVING DIFFICULTY STANDING TO GET OUT OF THE CAR AND HAS TO POP HIS HIPS BACK INTO PLACE. PATIENT IS SCHEDULED FOR REVISION SURGERY ON (B)(6) 2012 FOR RIGHT HIP. LEFT HIP WILL BE REVISED APPROXIMATELY 8 WEEKS AFTER RIGHT HIP. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02027 | 2012-10-23 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER DOCTOR (B)(6)'S OFFICE, THIS PATIENT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WILL BE REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01962 | 2012-10-23 | 2011-10-27 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PT COMPLAINS OF DIFFICULTY WALKING AND PAIN AT JOINT SITE SINCE THE SURGERY. PT'S CT-SCAN AND TESTS ARE PENDING. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02001 | 2012-10-23 | 2012-01-20 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT; PATIENT COMPLAINS OF MINOR PAIN, PAIN WHILE LYING DOWN. MRI AND BLOOD TEST ARE PENDING. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: REJUVENATE MODULAR NECK 30MM V40, CAT# NLV-300800G, LOT #38248301. TRITANIUM REVISION ACETABULAR, CAT #509-02-54E, LOT# MKP69V. TRIDENT 10DEG CROSSFIRE INSERT 36MM ID, CAT# 621-10-36E, LOT# MKP4LL. OSTEOLOCK BONE SCREW 35MM, CAT# 5260-5-035, LOT# 38323402. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S PAIN. AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. HOWEVER, SOME X-RAYS AND MEDICAL RECORDS HAVE BEEN RECEIVED. WHEN COMPLETED, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01995 | 2012-10-23 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT SURGEON REVISED A REJUVENATE STEM, NECK AND HEAD ON PATIENT DUE TO ADVERSE LOCAL TISSUE REACTION. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER. INFORMATION RECEIVED INDICATED THAT THE HOSPITAL IS SENDING THE EXPLANTED DEVICES TO OUTSIDE LAB FOR TESTING. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PATIENT/EVENT AS: 9616680-2012-00919. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02028 | 2012-10-23 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER DOCTOR (B)(6)'S OFFICE, THIS PATIENT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WILL BE REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02025 | 2012-10-23 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER DR (B)(6)'S OFFICE, THIS PATIENT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS AND MRI. UPDATE: PATIENT CALLED TO REPORT THAT HE IS SCHEDULED TO HAVE REVISION SURGERY ON (B)(6) 2012. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02023 | 2012-10-23 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER DOCTOR (B)(6)'S OFFICE, THIS PATIENT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WILL BE REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02021 | 2012-10-23 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER DOCTOR (B)(6)'S OFFICE, THIS PATIENT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WILL BE REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02026 | 2012-10-23 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP DEVICE | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, AS PER DOCTOR (B)(6)'S OFFICE, THIS PATIENT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WILL BE REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02050 | 2012-10-24 | 2012-01-30 | I, n, j, u, r, y | Manufacturer report | UNK LEFT REJUVENATE/ABG II HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT IS COMPLAINING OF HAVING PAIN EVER SINCE HE RECEIVED THE IMPLANT. HE HAS PAIN DIRECTLY IN THE JOINT. HE CANNOT WALK ON CONCRETE OR HARD SURFACES OR WALK OR STAND FOR EXTENDED PERIODS OF TIME (20 MINUTES) WITHOUT EXPERIENCING SEVERE PAIN. HE IS CURRENTLY TAKING TYLENOL TWICE A DAY FOR PAIN. THE PAIN DISSIPATES WITH REST AND THE MEDICATION. HE JUST RECEIVED THE RECALL LETTER FROM HIS PHYSICIAN TODAY. HE IS GOING TO SCHEDULE AN APPOINTMENT WITH HIS DOCTOR TODAY (B)(6) 2012. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02034 | 2012-10-24 | 2019-11-22 | I, n, j, u, r, y | Manufacturer report | ABGII. MODULAR STEM | PROSTHESIS, HIP, SEMI-CONSTRAINED, UNCEMENTED, METAL/POLYMER, NON-POROUS, CALICU | STRYKER ORTHOPAEDICS-MAHWAH | NA | 4845-4-102 |  | IT WAS REPORTED THAT PATIENT IS CURRENTLY SUFFERING NO PAIN. PATIENT STATES THAT SHE HAS HAD AN MRI, AN X-RAY, AND BLOOD WORK DONE TO TEST FOR CHROMIUM AND COBALT IN (B)(6) 2012. THE PATIENT IS REPORTING THAT THE X-RAY IS SHOWING FLUID BUILD UP. PATIENT IS ALSO REPORTING THAT DOCTOR IS RECOMMENDING THAT SHE BE EVALUATED IN SIX MONTHS. | AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED BUT BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02039 | 2012-10-24 | 2011-07-29 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE/AGBII HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT IS EXPERIENCING PAIN AND SWELLING. PATIENT IS REPORTING THAT SWELLING IS ON THE OUTSIDE OF THE THIGH UP TO THE LEFT HIP. PATIENT IS CURRENTLY ON DISABILITY. PATIENT HAS NOT SEEN HIS SURGEON SINCE HE HAS NO INSURANCE. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEN TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02059 | 2012-10-24 | 2012-10-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT THE PT'S REJUVENATE HEMI HIP (LEFT) WAS REVISED DUE TO PAIN. SURGEON REPLACED WITH AN EXETER TOTAL HIP. SURGEON WANTS IT STATED THAT "UPON ENTERING THE CAPSULE, SURGEON FOUND A PSEUDO TUMOR IN THE CAPSULE AND SUCTIONED OFF 150CC OF FLUID. THE SHORT ROTATORS WERE NECROTIC AND BROWN TO THE POINT OF SPREADING TOWARDS THE LESSER TROCHANTER." ORIGINAL SURGERY ON (B)(6) 2012 WAS DONE FOR A HIP FRACTURE OF PT'S LEFT HIP. SURGEON ALSO SAID THAT THERE WERE BLACK MARKINGS ON NECK WHERE THE NECK ENTERS THE HEAD. PT WILL BE SEEKING LEGAL REPRESENTATION. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR AS THE PT IS SENDING THE RECEIVED DEVICES TO INDEPENDENT LAB FOR TESTING. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD DEVICE OR ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS: 9616680-2012-00933. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02041 | 2012-10-24 | 2011-01-10 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT THE PT HAS PERIODIC PAIN IN THE RIGHT HIP. HE STATES THAT HIS PAIN IS MINIMAL AT THIS TIME. HE IS A HAIRDRESSER AND STANDS ON HIS FEET FOR LONG PERIODS. HE CAN'T TAKE OFF FROM WORK BECAUSE HE WON'T GET PAID. HE JUST RECEIVED THE LETTER AND HAS NOT YET SCHEDULED AN APPOINTMENT WITH HIS DOCTOR. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: REJUVENATE MODULAR NECK 34MM V40, CAT #NLV-340800G, LOT #31618201; TRIDENT 0 DEG INSERT 40MM, CAT #623-00-40F, LOT #MJTM19; 6.5 CANCELLOUS BONE SCREW 30MM, CAT #2030-6530-1, LOT #MJT25T; V40 COCR LFIT HEAD 40MM/+0, CAT #6260-9-140, LOT #MJN714. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02060 | 2012-10-24 | 2012-10-01 | I, n, j, u, r, y | Manufacturer report | UNK LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, I AM WRITING IN REGARD TO THE STRYKER REJUVENATE HIP SYSTEM. I HAD THIS DEVICE IMPLANTED IN MY LEFT HIP ON (B)(6) 2010, BY DR (B)(6) AT HOSPITAL IN (B)(6). I WAS ADVISED THAT THERE WAS A RECALL ON THIS DEVICE THIS SUMMER. I AM WRITING TO REGISTER THIS INFO WITH YOU. IT WAS ALSO REPORTED AS PER DOCTOR (B)(6)'S OFFICE, THIS PT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02035 | 2012-10-24 | 2022-03-25 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 10 | PROSTHESIS, HIP, SEMI-CONSTRAINED, UNCEMENTED, METAL/POLYMER, NON-POROUS, CALICU | STRYKER ORTHOPAEDICS-MAHWAH | NA | SPT-100000S |  | IT WAS REPORTED THAT PT HAS SWELLING THAT STARTED LAST WEEK. HIS SURGEON HAS ADVISED HER TO CALL STRYKER AND REGISTER. SHE WILL HAVE AN X-RAY TAKEN. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WILL BE REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02053 | 2012-10-24 | 2010-12-15 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT PT COMPLAINS OF PAIN SINCE HER SURGERY. SHE IS SCHEDULED FOR AN MRI NEXT WEEK. HER BLOOD TEST RESULTS AND X-RAYS ARE PENDING. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: REJUVENATE MODULAR NECK 34MM V40, CAT #NLS-340000B, LOT #34293302; TRIDENT 0 DEG X3 INSERT 32MM HEAD, CAT #623-00-32D, LOT# MJPNLM; TRIDENT HEMISPHERICAL CLUSTER HOLE SHELL, CAT #502-11-50D, LOT# 34139402; 32MM -4 LFIT V40 HEAD, CAT #6260-9-032, LOT #33695401; 6.5 CANCELLOUS BONE SCREW 25MM, CAT #2030-6525-1, LOT #MJM7NY. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01981 | 2012-10-24 | 2012-04-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE/ABGII HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT IS EXPERIENCING SOME PAIN IN RIGHT HIP. PT ALSO STATES THAT HE HAS TO LEAN SIDEWAYS WHEN DRIVING. PT HAS A PENDING APPOINTMENT TO HAVE BLOOD WORK AND X-RAYS DONE ON (B)(6) 2012 AT THE HOSPITAL WHERE HE HAD THE SURGERY. PT HAS ALSO HAD A LEFT HIP IMPLANTED AT THE SAME TIME AS THE RIGHT HIP, BUT LEFT HIP IS CURRENTLY ASYMPTOMATIC. | AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER, BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02049 | 2012-10-24 | 2010-01-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE/ABGII HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT COMPLAINED OF PAIN IN GROIN AREA. PT WENT FOR X-RAY AND REVEALED A WOBBLE TRACT. PT IS CURRENTLY FOLLOWING UP WITH THE SURGEON. IT IS FURTHER REPORTED THAT ADDITIONAL INFO WAS RECEIVED THAT APPROXIMATELY ONE MONTH AGO, THE PT BEGAN HAVING A SHOOTING PAIN AT THE IMPLANT SITE AND ON THE RIGHT SIDE OF HIP. HE CAN ONLY WALK ABOUT FIVE MINUTES BEFORE THE PAIN SETS IN. THE PAIN OCCURS WHEN HE ENGAGES IN ANY ACTIVITY OR ACTION THAT PUTS PRESSURE ON THE LEG OR HIP. HE RECENTLY HAD AN X-RAY TAKEN. HIS ORTHOPEDIST DIAGNOSED HIM WITH "MODERATE CHRONIC POST SURGICAL HETEROTYPIC BONE FORMATION." | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER, BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01992 | 2012-10-24 | 2011-03-04 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE/ABGII HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT IS REPORTED THAT THE PT HAD BLOOD WORK DONE. RESULTS SHOWED ELEVATED LEVELS OF COBALT. HAD CHROMIUM AND COBALT TESTING. NO PAIN OR DISCOMFORT. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER, BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01991 | 2012-10-24 | 2011-09-26 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE/ABGII HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT IS EXPERIENCING PAIN IN HIS RIGHT THIGH AND PAIN AT THE INCISION AREA. PT IS ALSO REPORTING THAT HE IS EXPERIENCING MINOR PAIN IN HIS GROIN AREA. PT HAS HAD BLOOD WORK FOR CHROMIUM TESTING (RESULTS PENDING) AND MRI IS SCHEDULED FOR (B)(6) 2012. | AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01989 | 2012-10-24 | 2010-10-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE/ABGII HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PT HAS BEEN EXPERIENCING PAIN IN HIP AREA FOR THE LAST YEAR. PT STATES THAT, WITHIN THE LAST SIX MONTHS, HE HAS DEVELOPED PAIN IN HIS FEMUR AREA. PT STATES THE PAIN IS NOT CONSTANT, BUT ON AND OFF. | AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02014 | 2012-10-24 | 2010-03-02 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE/AGBII HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | THE PATIENT REPORTS THAT SHE HAS HAD PAIN PERIODICALLY SINCE THE SURGERY. THE PAIN IS NOT OVERWHELMING BUT IT DOES BOTHER HER. SHE HAS BEEN HAVING BACK PAIN ON THE RIGHT SIDE AND BELIEVES IT MAY BE ATTRIBUTABLE TO HER HIP IMPLANT. SHE CAN'T LAY ON HER RIGHT SIDE VERY LONG WITHOUT EXPERIENCING PAIN. SHE HAS PAIN AND DIFFICULTY WHEN GETTING INTO HER VEHICLE. SHE ALSO NOTICES THAT IT HURTS WHEN SHE STANDS FROM A SITTING POSITION OR JUST WHEN SHE MOVES A CERTAIN WAY. SHE IS UNDERGOING PHYSICAL THERAPY FOR HER HIP AND BACK WHICH IS VERY PAINFUL. SHE STATES THAT SHE WAS UNABLE TO FOLLOW THROUGH WITH THE PHYSICAL THERAPY ONE DAY BECAUSE OF THE PAIN IN HER HIP. SHE HAS NOT YET SEEN HER SURGEON AND WILL SCHEDULE AN APPOINTMENT SOON. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02007 | 2012-10-24 | 2010-04-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE/AGBII HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT; PATIENT IS COMPLAINING OF PAIN. PATIENT IS EXPERIENCING LIMPING WHEN WALKING. PATIENT IS REPORTING THAT THE DOCTOR DID BLOOD WORK, X-RAYS AND CAT SCAN. PATIENT IS REPORTING THAT HIS COBALT LEVEL IS HIGH AND THAT THE CAT SCAN IS SHOWING MUSCULAR SWELLING. PATIENT STATES THAT THE DOCTOR WANTS TO PERFORM REVISION SURGERY. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02057 | 2012-10-24 | 2012-10-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT THE REJUVINATE REVISION STEM WAS REMOVED WITH SOME DIFFICULTY A SECURE FIT PLUS STEM WAS IMPLANTED AND A DALL MILES CABLE WAS SECURED AROUND PROXIMAL ASPECT. ADDITIONAL INFO: PT WAS REVISED FOR PAIN AND ELEVATED ION BLOOD LEVELS. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR. ADDITIONAL INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD DEVICE OR ADDITIONAL INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS: 9616680-2012-00932. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02003 | 2012-10-24 | 2011-11-14 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE/AGBII HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT; PATIENT HAS BEEN SUFFERING PAIN SINCE HAVING THE SURGERY. IT WAS ALSO REPORTED THAT THE PATIENT FELL IN (B)(6) 2012 AND HAD REHABILITATION. PATIENT IS SCHEDULED TO SEE DR (B)(6) ON (B)(6) 2012. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01988 | 2012-10-24 | 2010-12-09 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE/ABG II HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PT HAS DISCOMFORT PERIODICALLY, SHE IS UNSURE OF WHEN THE PROBLEMS STARTED EXACTLY. SHE REPORTS THAT SHE FEELS DISCOMFORT WHEN SHE SITS FOR TOO LONG. SHE STATES THAT HER BLOOD WORK SHOW ELEVATED LEVELS. HER MRI WAS DONE ONE WEEK AGO. SHE HAS MET WITH SURGEON AND SHE REPORTS TO HAVE FLUID BUILD UP. SHE WILL NEED A REVISION SURGERY TO REMOVE STEM ONLY. | AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01987 | 2012-10-24 | 2011-11-14 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE/ABGII HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PT IS NOT CURRENTLY EXPERIENCING PAIN OR DISCOMFORT. PT STATES THAT HE RECEIVED A LETTER FROM DOCTOR ON (B)(6) 2012. PT ALSO STATES THAT HE HAD HIS BLOOD WORK DONE ON (B)(6) 2012. PT'S COBALT LEVEL WAS 4.8. PT UNSURE OF HIS CHROMIUM LEVEL. PT WAS ADVISED TO HAVE BLOOD WORK DONE IN SIX MONTHS. | AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02011 | 2012-10-24 | 2010-12-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE/ABGII HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT STATES THAT SHE HAS BEEN HAVING PAIN AND DISCOMFORT IN HER LOWER BACK AND HIP FOR A LONG TIME. HER PAIN BEGAN APPROXIMATELY SEVEN MONTHS AFTER THE SURGERY IN 2010 AND CONTINUES TO BE PRESENT. SHE HAD AN MRI OF HER BACK IN (B)(6) 2012, WHICH INDICATED SOME ARTHRITIS IN HER BACK. THE PATIENT RECEIVED CORTISONE INJECTIONS IN HER SPINE FOR THE PAIN ON (B)(6) 2012. AT THAT TIME, THE DOCTOR STATED THAT HE BELIEVED THE SOURCE OF HER PAIN IS HER HIP AND NOT HER BACK. THE IMPLANT SITE IS VERY SORE TO THE TOUCH, BUT IS NOT SWOLLEN. SHE STATES THAT HER PAIN IS NOT DEBILITATING. SHE IS MORE AWARE OF THE PAIN WHEN SHE ATTEMPTS TO LAY ON IT. SOMETIMES SHE HAS TROUBLE SLEEPING BECAUSE OF THE DISCOMFORT. SHE STATES THAT APPLYING HEAT TO HER HIP HELPS THE PAIN. SHE ALSO STATES THAT SHE HAS HAD BOTH KNEES REPLACED. THE PATIENT IS SCHEDULING AN APPOINTMENT WITH HER SURGEON FOR BLOOD WORK AND FOLLOW-UP WITHIN THE NEXT TWO WEEKS WHEN SHE WILL BE IN (B)(6). | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02006 | 2012-10-24 | 2011-08-30 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT IS REPORTED THAT: THE PATIENT STATES THAT HE DEVELOPED A TERRIBLE PAIN TWO MONTHS AFTER HIS SURGERY. HIS RECENT BLOOD WORK INDICATES ELEVATED COBALT LEVELS. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01993 | 2012-10-24 | 2009-10-04 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE/ABGII HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT IS REPORTED THAT THE PT'S HIP "SQUEAKS AND POPS" AND IS PAINFUL SINCE PRIMARY SURGERY. PT STATES THAT THERAPY IS NOT WORKING. X-RAYS TAKEN. PT STATES THAT DOCTOR TOLD HER THAT SHE MAY NEED TO REPLACE THE MECHANISM OF HER IMPLANTS. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER, BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01990 | 2012-10-24 | 2012-05-15 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE/ABGII HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT IS EXPERIENCING EXTREME PAIN, ESPECIALLY WHILE DRIVING. PT IS CURRENTLY EXPERIENCING PAIN FROM THIS BACK, ON THE RIGHT SIDE DOWN TO HIS KNEE. PT STATED THAT, INITIALLY, AFTER HE HAD THE SURGERY, HE HAD NO PAIN BUT AS TIME PASSED, THE PAIN GRADUALLY INCREASED. PT REPORTED THAT HE WENT TO HIS FAMILY DOCTOR THEN THE FAMILY PHYSICIAN TOLD HIM TO GO TO THE CHIROPRACTOR. PT ALSO REPORTED THAT HE IS CURRENTLY TAKING PAIN MEDICATION (VICODIN/TYLENOL AND ALSO TRAMADOL). | AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02038 | 2012-10-24 | 2012-03-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE/AGBII HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT IS REPORTED THAT: THE PT RECEIVED A LETTER REGARDING THE RECALL. HE HAS HAD NO POST OPERATIVE VISITS SINCE THE IMPLANT BECAUSE HE LIVES IN A DIFFERENT STATE. THE PT HAD X-RAYS TAKEN IN (B)(6) 2012, WHICH HE SENT TO HIS SURGEON. THE PT STATES THAT HE OCCASIONALLY EXPERIENCES A BURNING SENSATION ON THE SIDE OF HIS RIGHT HIP AND A PULLING SENSATION LIKE A MUSCLE PULL IN THE GROIN AREA AND THE FRONT OF THE HIP. HE IS TRAVELING FROM (B)(6) TO (B)(6) FOR TESTING AND FOLLOW-UP WITH HIS SURGEON ON (B)(6) 2012. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
9616680-2012-00929 | 2012-10-24 | 2012-02-15 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK34MM V40 | UNKNOWN | STRYKER ORTHOPAEDICS CORK | NA | NLS-340000B |  | IT WAS REPORTED THAT THE PT HAS EXPERIENCED CONSISTENT PAIN SINCE THE SURGERY IN FEBRUARY. HE HAS UNDERGONE PHYSICAL THERAPY FOR HIS HIP BUT IT HAS NOT HELPED. WITHIN THE LAST MONTH, HIS PAINS HAVE CHANGED. HE HAS BEGUN TO EXPERIENCE MID GLUTEAL AND INGUINAL TENDERNESS IN THE GROIN. HE ALSO HAS OTHER CONDITION IN THAT LEG UNRELATED TO THE HIP IMPLANT. A RECENT MRI HAS SHOWN FLUID ON THE LEFT HIP AND HIS BLOOD WORK SHOWS ELEVATED METAL CONTENT. HE HAS HAD TWO CONSECUTIVE BLOOD TESTS AND HE HEAVY METAL LEVELS HAVE SHOWN MINOR ELEVATIONS. HE WILL HAVE ADDITIONAL TESTING AROUND THANKSGIVING. HIS SURGEON ADVISES THAT IF THE METAL LEVELS ARE INCREASED AT THAT TIME, A REVISION SURGERY IS RECOMMENDED. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS MFR # 2249697-2012-02043. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02043 | 2012-10-24 | 2012-02-15 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 10 | UNKNOWN | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-100000S |  | IT WAS REPORTED THAT THE PT HAS EXPERIENCED CONSISTENT PAIN SINCE THE SURGERY IN FEBRUARY. HE HAS UNDERGONE PHYSICAL THERAPY FOR HIS HIP, BUT IT HAS NOT HELPED. WITHIN THE LAST MONTH, HIS PAINS HAVE CHANGED. HE HAS BEGUN TO EXPERIENCE MID GLUTEAL AND INGUINAL TENDERNESS IN THE GROIN. HE ALSO HAS OTHER CONDITIONS IN THAT LEG UNRELATED TO THE HIP IMPLANT. A RECENT MRI HAS SHOWN FLUID ON THE LEFT HIP AND HIS BLOOD WORK SHOWS ELEVATED METAL CONTENT. HE HAS TWO CONSECUTIVE BLOOD TESTS AND HIS HEAVY METAL LEVELS HAVE SHOWN MINOR ELEVATIONS. HE WILL HAVE ADDITIONAL TESTING AROUND (B)(6). HIS SURGEON ADVISES THAT IF THE METAL LEVELS ARE INCREASED AT THAT TIME, A REVISION SURGERY IS RECOMMENDED. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS MFR # 9616680-2012-00929. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-01979 | 2012-10-24 | 2010-08-03 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT HIP - REJUVENATE/ABGII | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT IS REPORTED THAT THE PT COMPLAINS OF HAVING PAIN IN HIS UPPER HIGH AND HIP FOR THE PAST SIX MONTHS TO A YEAR. THE PT HAS NOTICED A SIGNIFICANT CHANGE IN HIS MOBILITY WITHIN THE LAST YEAR. THE PT STATES THAT HE HAS TROUBLE WALKING DISTANCES AND FREQUENTLY BEGINS TO LIMP FROM THE PAIN. HE RECENTLY HAD AN APPOINTMENT WITH HIS SURGEON AND COMPLETED MEDICAL TESTING. THE PT IS SCHEDULED FOR A REVISION SURGERY ON (B)(6) 2012. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02051 | 2012-10-24 | 2012-02-10 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT HIP - REJUVENATE/ABGII | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT IS REPORTED, "THE PT HAD A TOTAL HIP REPLACEMENT DONE ON (B)(6) 2012. AROUND (B)(6) 2012, HE NOTICED HE WAS NOT IMPROVING IN RELATIONSHIP TO PAIN. HE CONTACTED HIS SURGEON WHO SAID THAT THIS WAS NORMAL. IN (B)(6) 2012, HE STILL HAD PAIN AND SAW HIS SURGEON. DURING HIS VISIT HE WAS INFORMED THAT HIS HIP HAD BEEN RECALLED BY STRYKER. THE PT REPORTS FEELING HELPLESS AND STATES THAT THE THOUGHT OF REVISION SURGERY SENDS HIM INTO MANIC DEPRESSION. HIS SURGEON DOES NOT THINK HE HAS THE FRETTING PROBLEM THAT INITIATED THE RECALL. HE WAS ADVISED BY HIS SURGEON TO TAKE A TWO-WEEK REST FROM HIS EXERCISE REGIMEN, WHICH HE DID. HE STILL HAS THE PAIN. HE HAS ALREADY HAD AN X-RAY. HE VISITED HIS DOCTOR ON (B)(6) 2012 AND IS SCHEDULED TO HAVE BLOOD WORK AND AN MRI THE WEEK OF (B)(6). THE PT STATES THAT HIS PAIN IS AT THE IMPLANT SITE AND SOMETIMES EXTENDS TO THE FRONT OF THE HIP. HE HAS PAIN SITTING, STANDING AND WALKING FOR LONG PERIODS OF TIME; THE MORE INTENSE HIS ACTIVITY, THE MORE INTENSE THE PAIN. HE TAKES AN OVER THE COUNTER ANTI-INFLAMMATORY WHICH PROVIDES SOME RELIEF." | AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER, BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02088 | 2012-10-25 | 2010-12-02 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT; PATIENT IS CURRENTLY IN PAIN. PATIENT WAS INFORMED BY HER DOCTOR ABOUT THE RECALL DURING HER VISIT (B)(6) 2012. PATIENT COMPLAINS THAT, FOR THE PAST YEAR, THE PAIN HAS GRADUALLY GOTTEN WORSE. PATIENT ALSO HAS BEEN HAVING HEART ISSUES AND HIS HEART IS OUT OF SYNC. DOCTOR IS REQUESTING THAT PATIENT HAVE MRI, COBALT, AND CHROMIUM TESTING DONE. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02069 | 2012-10-25 | 2019-06-24 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | PROSTHESIS, HIP, SEMI-CONSTRAINED, UNCEMENTED, METAL / POLYMER, NON-POROUS, CALC | STRYKER ORTHOPAEDICS-MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT THE PATIENT BEGAN EXPERIENCING EXTREME SWELLING (B)(6) 2012, IN THE LOWER PART OF HER LEFT LEG. SHE MOSTLY EXPERIENCES THE SWELLING WHEN THE WEATHER IS HUMID. SHE HAS RECENTLY NOTICED A RASH OF SMALL RED SPOTS UNDER THE SURFACE OF THE SKIN IN HER LEFT AND RIGHT LEGS EXTENDING FROM HER KNEE TO HER FOOT. THE PATIENT HAS AN APPOINTMENT FOR BLOOD TEST ON (B)(6) 2012. SHE IS SCHEDULED TO SEE HER PHYSICIAN ON (B)(6) 2012, FOR TEST RESULTS AND A FOLLOW-UP EXAMINATION. | AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, BUT BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02071 | 2012-10-25 | 2010-09-21 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT HIP REJUVENATE/ABG II | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PATIENT HAS BEEN EXPERIENCING PAIN IN THE RIGHT HIP. PATIENT IS REPORTING THAT SHE HAD AN MRI AND BLOOD WORK DONE AND IS AWAITING RESULTS. PATIENT STATES THAT THE PATIENT STARTED IN (B)(6) 2012 AND IT CONSISTS OF A STING IN HER BUTTOCK AREA. PATIENT ALSO STATES THAT SHE HAD IMPLANT SURGERY ON HER LEFT HIP WHICH IS ALSO A REJUVENATE HIP IMPLANT. PATIENT STATES THAT IS WAS DONE ON (B)(6) 2011 AND IS CURRENTLY ASYMPTOMATIC. | AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02070 | 2012-10-25 | 2010-09-01 | I, n, j, u, r, y | Manufacturer report | UNK LEFT REJUVENATE/ABG II HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PT HAS TWO HIP IMPLANTS. THE RIGHT HIP IS ASYMPTOMATIC. PT IS HAVING PAIN IN HIS LEFT HIP. PT IS EXPERIENCING SWELLING AND SHARP PAIN. PT IS REPORTING THAT THE DOCTOR WANTS TO SCHEDULE AN MRI FOR BOTH OF HIS HIPS. THE RIGHT HIP SURGERY WAS DONE IN (B)(6) 2010, APPROXIMATELY. | AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER, BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02064 | 2012-10-25 | 2012-01-01 | I, n, j, u, r, y | Manufacturer report | UNK LEFT REJUVENATE/AGBII HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT IS REPORTED THAT: THE PT HAD A REVISION SURGERY IN (B)(6) 2012 DUE TO AN INFECTION. THE PT STATES THAT SINCE (B)(6) 2012, HIS HIP HAS DISLOCATED FOUR TIMES. EACH DISLOCATION REQUIRED A SURGICAL ADJUSTMENT. THE LAST DISLOCATION WAS IN (B)(6) 2012. THE PT IS SCHEDULED FOR AN MRI AND BLOOD WORK ON (B)(6) 2012. THE PT STATES THAT HE IS GOING TO HAVE A SECOND REVISION SURGERY. | AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER, BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02062 | 2012-10-25 | 2011-10-05 | I, n, j, u, r, y | Manufacturer report | UNK - RIGHT HIP REJUVENATE/ABGII | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT IS REPORTED THAT THE PT IS HAVING PAIN IN HIS HIP THAT HURTS VERY BADLY WHEN HE WALKS. HE HAD TO PURCHASE A FOUR WHEEL WALKER AND CANNOT GET AROUND WITHOUT IT. THE PAIN STARTED REALLY GETTING BAD AROUND (B)(6) 2012. HIS PHYSICIAN SAID HIS X-RAY WAS FINE AND HE WAS ON SCHEDULE WITH HIS RECOVERY. HE RECEIVED A LETTER REGARDING THE RECALL ON (B)(6) 2012. | AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER, BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02083 | 2012-10-25 | 2011-01-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: PATIENT IS EXPERIENCING STIFFNESS AFTER SITTING FOR 20 MINUTES. PATIENT HAS ALSO EXPERIENCED PAIN AND SWELLING FOR THE LAST SIX MONTHS. PATIENT SAYS THAT HIS HIP FEELS LIKE IT IS SWOLLEN. PATIENT STATES THAT HE HAS HAD TESTING OR CHROMIUM AND THE TESTS ARE CURRENTLY NEGATIVE. PATIENT HAS ALSO HAD AN MRI, BUT HAS YET TO RECEIVE RESULTS. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENT REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02082 | 2012-10-25 | 2011-02-24 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKE ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT PATIENT COMPLAINS OF PAIN IN GROIN AREA. SHE HAS HAD TROUBLE WALKING AND SWIMMING. SHE HAS HAD A NUCLEAR BONE SCAN AND STATES THAT SHE WILL NEED A REVISION SURGERY. A REVISION DATE HAS NOT BEEN SCHEDULED AS PATIENT AND DOCTOR HAVE DECIDED TO WAIT UNTIL AFTER SHE GETS HER LEFT HIP REPLACED. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02106 | 2012-10-26 | 2012-10-03 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT THE DOCTOR REVISED PT'S RIGHT HIP DUE TO ALTR. ADDITIONAL INFO FROM THE PT: THE PT REPORTED EXTENSIVE SWELLING AND BRUISING IN THE HIP AREA. HE RECENTLY HAD AN MRI AND BLOOD WORK COMPLETED. HIS BLOOD TESTS INDICATED SEVERELY ELEVATED COBALT AND CHROMIUM LEVELS. HE STATES THAT HE HAS NEVER BEEN PAIN FREE SINCE THE INITIAL IMPLANT AND OFTEN HAS TO TAKE PAIN PILLS FOR RELIEF. THE PT JUST HAD A REVISION SURGERY ON (B)(6) 2012. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR DUE TO HOSP POLICY. ADDITIONAL INFO INCLUDING SOME X-RAYS AND MEDICAL RECORDS HAS BEEN RECEIVED. WHEN COMPLETED, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS: MFR # 9616680-2012-00954. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02095 | 2012-10-26 | 2012-10-03 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STEM EXTRACTOR | INSTRUMENT | STRYKER ORTHOPAEDICS MAHWAH | NA | 1601-1650 |  | IT WAS REPORTED THAT: THE SURGEON WAS MALLETING THE EXTRACTOR AND THE EXTRACTOR CONTINUALLY DISENGAGED FROM THE WELL FIXED STEM. AN OSTEOTOMY WAS PERFORMED TO REMOVE THE STEM. NO ABNORMAL WEAR OF THE INSTRUMENT WAS NOTICED. | WHEN COMPLETED, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | Y | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02104 | 2012-10-26 | 2012-10-03 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE STEM | IMPLANT | STRYKER ORTHOPAEDICS | NA | UNK |  | IT WAS REPORTED THAT, PT THA DONE APPROX 2 YRS AGO, PRESENTED WITH HIP PAIN, SURGEON SAID ACETABULAR CUP LOOKED LOOSE, BOOKED PT FOR REVISION OF ACETABULAR IMPLANT. DR. NOTED PT HAD A REJUVENATE STEM IN, I EXPLAINED STEM WAS RECALLED, INTRAOPERATIVELY DOC DETERMINED THE JOINT WAS INFECTED, DOC PUT IN ANTIBIOTIC SPACER IN TO ADDRESS INFECTION BEFORE REIMPLANTATION. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: UNK REJUVENATE NECK, UNK HEAD, UNK CUP, UNK LINER, UNK SCREW. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S REVISION INFECTION. AN EVAL OF THE REPORTED DEVICE(S) CANNOT BE PERFORMED AS THE DEVICE(S) WERE RETAINED BY THE PT AND WERE NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICES REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS: MFR # 2249697-2012-02105. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02097 | 2012-10-26 | 2012-10-03 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 10 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-100000S |  | THE PT WAS HAVING PAIN SO THE SURGEON REVISED. AN OSTEOTOMY HAD TO BE PERFORMED TO REMOVE THE STEM. THREE DALL-MILES CABLES WERE ALSO USED. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0DEG 38MM, CAT# NLS-380000B, LOT# 37061801, TRIDENT PSL HA SOLID BACK 62MM, CAT# 540-11-62H, LOT# R5PMLE, TRIDENT 0DEG X3 INSERT 36MM ID, CAT# 623-00-36H, LOT# 50TMMD, DELTA V-40 CERAMIC HEAD 36/0, CAT# 6570-0-136, LOT# 38244008. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVAL OF THE REPORTED DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS RETAINED BY THE HOSPITAL AND WAS NOT RETURNED TO THE MFR. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02107 | 2012-10-26 | 2012-10-03 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT PT'S RIGHT REJUVENATE HIP HAS BEEN CAUSING HER PAIN AND MUST BE REVISED. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0DEG 34MM, CAT#NLS-340000B, LOT# UNK. PRIMARY TRITANIUM HEMI CLUSTER HOLE CUP 54MM, CAT# 502-03-54E, LOT# MJE2NL. TRIDENT 10DEG X3 INSERT 36MM ID, CAT# 623-10-36E, LOT# MJD640. DELTA V-40 CERAMIC HEAD 36/+2.5, CAT# 6570-0-536, LOT# 30650801. D-M 2.0MM BEADED CABLE SET VIT, CAT# 6704-0-520, LOT# 31319004. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICES REMAINED IMPLANTED IN THE PT AND WERE NOT RETURNED TO THE MFR. ADDITIONAL INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02111 | 2012-10-26 | 2012-10-04 | I, n, j, u, r, y | Manufacturer report | REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, IT WAS REPORTED THROUGH THE ATTORNEY FOR THE PT, AS A RESULT OF A LEGAL CLAIM, THAT ALLEGEDLY THE PT HAD A TOTAL HIP REPLACEMENT ON (B)(6) 2011 UTILIZING THE REJUVENATE HIP STEM. IT IS FURTHER ALLEGED THAT AS A RESULT OF THE INSERTION OF THAT DEVICE, THE PT WILL REQUIRE ANOTHER SURGERY TO HAVE THE DEVICE REMOVED AND REPLACED. | THE INFO IN THIS REPORT WAS PROVIDED BY STRYKER ORTHOPAEDICS LEGAL AFFAIRS DEPT. NO ADDITIONAL INFO IS AVAILABLE AT THIS TIME. THE DEVICE IS NOT AVAILABLE AS THE DEVICE REMAINED IMPLANTED IN THE PT. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02102 | 2012-10-26 | 2011-10-04 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT STARTED EXPERIENCING PAIN IN HIS HIP IN (B)(6) 2012. PT ALSO REPORTED THAT HE IS EXPERIENCING SWELLING THAT IS NOT RECEDING OR GOING DOWN. PT STATED THAT HE HAD BLOOD WORK DONE IN (B)(6) FOR THE COBALT AND CHROMIUM AND THE COBALT LEVEL IS 4.7. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-10
9616680-2012-00940 | 2012-10-26 | 2011-12-28 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK 42MM V40 | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | NLS-420000B |  | IT WAS REPORTED THAT, PT IS EXPERIENCING OCCASIONAL PAIN WHILE WALKING. PT STATED THAT HER BLOOD WORK SHOWS HIGH COBALT LEVELS AT 8.5 AND THE MRI SHOWS THAT THERE ARE METAL FRAGMENTS. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS MFR # 2249697-2012-01881. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
9616680-2012-00945 | 2012-10-26 | 2011-09-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK 38MM V40 | UNKNOWN | STRYKER ORTHOPAEDICS CORK | NA | NLS-380000B |  | IT WAS REPORTED THAT, PT IS SCHEDULED FOR REVISION SURGERY ON (B)(6) 2012 ON HIS RIGHT HIP. PT STATES THAT HIS DR ORDERED BLOOD WORK WHICH SHOWED ELEVATED CHROMIUM AND COBALT LEVELS. PT COMPLAINS OF HAVING DULL PAIN ALMOST ALL THE TIME THAT SOMETIMES BECOMES SHARP. HE EXPERIENCES SOME STIFFNESS AND DIFFICULTY WALKING ESPECIALLY AFTER PERIODS OF ACTIVITY. HE HAS DIFFICULTY SLEEPING ESPECIALLY WITHIN THE LAST YEAR DUE TO PAIN IN HIS HIPS. PT STATES HE HAS PERIODIC WEAKNESS IN BOTH LEGS BUT MORE EXTREME IN THE RIGHT. PAIN IN BOTH HIPS WILL SOMETIMES EXTEND DOWN INTO THE FEMUR. PT HAD AN MRI AND ADDITIONAL BLOOD LABS ON (B)(6) 2012. THE PT STATES THAT HE HAS THE SAME SYMPTOMS IN BOTH HIPS BUT THOSE IN HIS LEFT HIP ARE LESS INTENSE AND FREQUENT THAN THOSE IN HIS RIGHT HIP. HE LAST SAW HIS DOCTOR ON (B)(6) 2012. | THIS IS THE SAME PT/'EVENT AS MFR # 2249697-2012-01799. AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFO (X-RAYS, MEDICAL RECORDS, OPERATIVE NOTES) WAS REQUESTED. IF IT BECOMES AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02092 | 2012-10-26 | 2012-10-03 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT HIP COMPONENTS WERE EXPLANTED DUE TO ADVERSE LOCAL TISSUE REACTION. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR DUE TO HOSPITAL POLICY. ADD'L INFO INCLUDING X-RAYS AND MEDICAL RECORDS WAS NOT MADE AVAILABLE EITHER. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS: 9616680-2012-00943. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2814194 | 2012-10-29 | 2012-10-22 | N, o, , a, n, s, w, e, r, , p, r, o, v, i, d, e, d | User Facility report | REVOLUTION CEMENT MIXING SYSTEM WITH FEMORAL | PROSTHESIS, HIP, FEMORAL COMPONENT | STRYKER INSTRUMENTS | * | 606563000 |  | EXPIRED CEMENT MIXER FOUND INSIDE CUSTOM TOTAL HIP PACK CONTAINING MULTIPLE ITEMS WITH DIFFERENT EXPIRATION DATE FROM REST OF PACK. |  | N | N | PHYSICIAN |  | 2012-10
9616680-2012-00957 | 2012-10-29 | 2011-04-20 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK 34MM V40 | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | NLS-340000B |  | PT STATES THAT SHE HAS MINOR DISCOMFORT IN THE RIGHT HIP. SHE WENT TO THE DOCTOR ON (B)(6) 2012. SHE HAD BLOOD LABS DONE ON (B)(6) 2012. SHE RECEIVED THE RESULTS FROM HER DOCTOR'S OFFICE ON (B)(6) 2012 SAYING THAT HER LEVELS WERE ELEVATED. SHE IS BEING SCHEDULED FOR AN MRI ON HER RIGHT HIP THROUGH THE DOCTOR'S OFFICE. THE PT STATES THAT SHE IS HAVING TROUBLE MANEUVERING STAIRS. SHE WAS GIVEN A PRESCRIPTION FOR AN ANTI-INFLAMMATORY MEDICATION ON (B)(6) 2012. SHE HAS BEEN EXPERIENCING OCCASIONAL NERVE TINGLING IN THE HIP SINCE (B)(6) 2012. PT STATES THAT THE ISSUES IN HER RIGHT HIP ARE MINOR AND INFREQUENT. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS MFR #2249697-2012-01804. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02124 | 2012-10-30 | 2011-06-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT IS REPORTED THAT: PATIENT COMPLAINS OF HIP, THIGH, AND GROIN PAIN SINCE THE SURGERY. PT IS CALLING TO FILE A CLAIM AS ADVISED BY HER SURGEON FOR THE EVALUATIONS THAT SHE WILL REQUIRE AS PART OF THE RECALL. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD¿L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD¿L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02120 | 2012-10-30 | 2012-10-05 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT THE PATIENT COMPLAINED OF PAIN IN RIGHT HIP. MRI SHOWED FLUID BUILD-UP AND EXTENSIVE TISSUE DAMAGE. SURGEON NOTICED EXCESSIVE FRETTING AT THE HEAD NECK JUNCTION. REPLACED WITH A RESTORATION MODULAR CONE CONICAL HIP. | AN EVALUATION OF THE REPORTED DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS RETAINED BY THE PATHOLOGY IN THE HOSPITAL AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION INCLUDING X-RAYS AND MEDICAL RECORDS WAS NOT MADE AVAILABLE EITHER. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PATIENT/EVENT AS: 9616680-2012-00965. | Y | N | HEALTH PROFESSIONAL |  | 2012-10
2249697-2012-02171 | 2012-11-02 | 2012-10-17 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT THE COMPONENTS WERE EXPLANTED DUE TO ADVERSE LOCAL TISSUE REACTION. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-11
2249697-2012-02164 | 2012-11-02 | 2019-05-02 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE MODULAR STEM | HIP IMPLANT | STRYKER ORTHOPAEDICS-MAHWAH | NA | UNK_JR |  | IT WAS REPORTED THAT: PATIENT STATES THAT HER HIP IS SWOLLEN; THAT SHE EXPERIENCES SHARP PAIN WHEN WALKING OR ENTERING THE CAR. PATIENT IS ALSO STATING THAT HER HIP FEELS TENDER AND THAT THE PAIN HAS BEEN ONGOING SINCE MAY. PATIENT REPORTS THAT SHE HAS BEEN TO SEE HER SURGEON AND THAT HE HAS DONE THE BLOOD WORK AND HAD X-RAYS TAKEN AND THE RESULTS ARE PENDING. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. DEVICE INFORMATION HAS NOT BEEN PROVIDED AT THIS TIME. INFORMATION RECEIVED FROM THE PATIENT INDICATED THAT REPORTED DEVICE MAY BE EITHER REJUVENATE OR ABGII. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02160 | 2012-11-02 | 2012-10-08 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT: PATIENT'S RIGHT REJUVENATE HIP WAS CAUSING PATIENT PAIN. SURGEON REVISED PATIENT'S HIP ON (B)(6) 2012. ADDITIONAL INFORMATION: BASED ON THE PROVIDED PHOTOS OF THE REVISED DEVICES IT APPEARS THAT THERE IS ADVERSE LOCAL TISSUE REACTION PRESENT. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER DUE TO HOSPITAL POLICY. HOWEVER, PHOTOS OF THE REVISED DEVICES WERE PROVIDED. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PATIENT/EVENT AS: MFR # 9616680-2012-00977. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02167 | 2012-11-02 | 2012-10-08 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT THERE WAS CORROSION AND FRETTING AT NECK AND STEM. | WHEN COMPLETED, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PATIENT/EVENT AS: MFR # 9616680-2012-00980. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02162 | 2012-11-02 | 2012-10-08 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT DOCTOR REVISED PATIENT'S LEFT HIP DUE TO ADVERSE LOCAL TISSUE REACTION (ALTR). | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER DUE TO HOSPITAL POLICY. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED, BUT NOT MADE AVAILABLE. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PATIENT/EVENT AS: MFR # 9616680-2012-00979. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02161 | 2012-11-02 | 2012-10-08 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | UNKNOWN | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT DOCTOR REVISED PATIENT'S RIGHT HIP DUE TO ALTR. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER DUE TO HOSPITAL POLICY. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED, BUT NOT MADE AVAILABLE. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PATIENT/EVENT AS: MFR # 9616680-2012-00978. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-00975 | 2012-11-02 | 2011-10-12 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK 34MM V40 | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | NLS-340000B |  | IT WAS REPORTED THAT THE PT IS REPORTING THAT HER HIP STARTED HAVING PAIN A MONTH AND A HALF AGO. PT IS EXPERIENCING PAIN WHEN SHE TWISTS OR TURNS. PT IS ALSO REPORTING THAT SHE IS EXPERIENCING A SHARP PAIN WHEN SHE STANDS UP. PT IS EXPERIENCING DIFFICULTY GOING UP AND DOWN STAIRS. PT STATES THAT SHE CANNOT SLEEP ON THE LEFT SIDE OF HER HIP. SHE IS ALSO STATING THAT SHE IS RESTRICTED IN HER EXERCISE AND IS COMPLAINING OF INFLAMMATION IN HER LEFT HIP AND ALSO BOTH HER ANKLES. PT IS REPORTING THAT WHEN DRIVING SHE HAS TO USE A CUSHION BECAUSE THE SEATING CAUSES PAIN AND DISCOMFORT. SHE IS COMPLAINING OF MORE STRAIN ON HER HIP WITHIN THE LAST MONTH. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-11
2249697-2012-02191 | 2012-11-02 | 2012-10-10 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT, CASE SCHEDULED AS REVISION OF LOOSE CUP AS PT WAS COMPLAINING OF PAIN AFTER A FALL. WHILE IN SURGERY, THE DR CHECKED TAPER JUNCTION OF NECK AND STEM AND SAW DARK DISCOLORATION. DECISION WAS MADE TO REVISE ALL COMPONENTS. | AN EVALUATION OF THE REPORTED DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS RETAINED BY THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS: MFR #9616680-2012-00986 AND MFR # 9616680-2012-00987. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02189 | 2012-11-02 | 2012-10-09 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 10 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-100000S |  | IT WAS REPORTED THAT, PT REPORTED PAIN IN RIGHT HIP, DR (B)(6) PERFORMED ROUTINE FOLLOW-UP AND X-RAYS. A YEAR AFTER HIS ORIGINAL SURGERY AND FORMER SURGEON FOUND ALL FLUID IN HIS HIP AND THEN HAD A HEAD, NECK, AND LINER SWAPOUT. BONE SCAN LIT UP AROUND THE STEM AND CUP. | AN EVALUATION OF THE REPORTED DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS RETAINED BY THE PT FOR HIS LAW FIRM AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD DEVICE OR ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS: MFR #9616680-2012-00984 AND MFR # 9616680-2012-00985. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02186 | 2012-11-02 | 2009-12-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT: PT IS EXPERIENCING "CLICKING AND POPPING" AND PAIN IN GROIN AREA. PT IS ALSO EXPERIENCING ADDITIONAL PAIN ON TOP OF THIGHS. PT IS REPORTING DIFFICULTY LIFTING LEG GOING UPSTAIRS FOR BOTH HIPS. PT IS ALSO EXPERIENCING DIFFICULTY STANDING TO GET OUT THE CARE AND HAS TO POP HIS HIPS BACK INTO PLACE. PT IS SCHEDULED FOR REVISION SURGERY ON (B)(6) 2012 FOR LEFT HIP. RIGHT HIP WILL BE REVISED APPROXIMATELY 8 WEEKS AFTER LEFT HIP. ADDITIONAL INFO REPORTED VIA SALES REP: IT WAS REPORTED THAT, REJUVENATE REVISION. BLOOD SERUM ION LEVEL WAS ELEVATED. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR. ADDITIONAL INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD DEVICE OR ADDITIONAL INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS: MFR #9616680-2012-00983. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02174 | 2012-11-02 | 2010-11-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT IS REPORTED THAT: PT IS SCHEDULED FOR REVISION SURGERY ON (B)(6) 2012. MRI SHOWS PT HAS A TUMOR AND REQUIRES A REVISION SURGERY. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFO PERTAINING TO THE DEVICE REFERENCE IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02173 | 2012-11-02 | 2011-10-11 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT IS REPORTED THAT: PT IS EXPERIENCING SWELLING AND PAIN IN HIP AREA SINCE THE SURGERY, BACK, AND BUTTOCKS. PT EXPERIENCES PAIN WHILE WALKING. PT IS SCHEDULED FOR FOLLOW UP VISION WITH DOCTOR ON (B)(6) 2012. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADDITIONAL INFO PERTAINING TO THE DEVICE REFERENCE IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02141 | 2012-11-02 | 2011-10-12 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT THE PT IS REPORTING THAT HER HIP STARTED HAVING PAIN A MONTH AND A HALF AGO. PT IS EXPERIENCING PAIN WHEN SHE TWISTS OR TURNS. PT IS ALSO REPORTING THAT SHE IS EXPERIENCING A SHARP PAIN WHEN SHE STANDS UP. PT IS EXPERIENCING DIFFICULTY GOING UP AND DOWN STAIRS. PT STATES THAT SHE CANNOT SLEEP ON THE LEFT SIDE OF HER HIP. SHE IS ALSO STATING THAT SHE IS RESTRICTED IN HER EXERCISE AND IS COMPLAINING OF INFLAMMATION IN HER LEFT HIP AND ALSO BOTH HER ANKLES. PT IS REPORTING THAT WHEN DRIVING SHE HAS TO USE A CUSHION BECAUSE THE SEATING CAUSES PAIN AND DISCOMFORT. SHE IS COMPLAINING OF MORE STRAIN ON HER HIP WITHIN THE LAST MONTH. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-11
2249697-2012-02212 | 2012-11-02 | 2012-10-11 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT, PT PRESENTED WITH HIP PAIN. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0DEG 30MM, CAT # NLS-340000B, LOT # 33810101. RESTORATION ADM X3 INS 28/56, CAT # 1236-2-856, LOT # 33571001. RESTORATION (TM) ADM. CUP W/HA, CAT# 1235-2-561, LOT # G285217K6Q01, 28MM STD LFIT V40 HEAD, CAT # 6260-9-128, LOT # 34590101. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR DUE TO HOSPITAL POLICY. HOWEVER, X-RAYS AND PHOTOS WILL BE PROVIDED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02175 | 2012-11-02 | 2012-10-16 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT COMPONENTS WERE EXPLANTED DUE TO ADVERSE LOCAL TISSUE REACTION. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-11
2249697-2012-02165 | 2012-11-02 | 2011-05-18 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT: PATIENT IS EXPERIENCING PAIN IN HIS LEFT HIP ALONG WITH INFLAMMATION. PATIENT STATES THAT THE PAIN STARTED A MONTH AGO ALONG WITH HEARING A GRINDING AND SQUEAKING NOISE. PATIENT IS REPORTING THAT HE HAS SEEN THE SURGEON, HAS HAD THE BLOOD WORK, MRI AND X-RAYS DONE AND IS AWAITING RESULTS. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0DEG 30MM, CAT# NLS-300000B, LOT# 34095901; TRIDENT 0DEG X3 INSERT 36MM ID, CAT# 623-00-36E, LOT# MKA36R; TRIDENT HEMISPHERICAL CLUSTER HOLE SHELL, CAT# 502-11-54E, LOT# 35992104; DELTA V-40 CERAMIC HEAD 36/-5, CAT# 6570-0-036, LOT# 36162301; 6.5 CANCELLOUS BONE SCREW 25MM, CAT# 2030-6525-1, LOT# MKDMRM. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S PAIN. AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02249 | 2012-11-07 | 2012-04-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT HIP REJUVENATE/ABG II | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT HAS BEEN EXPERIENCING PAIN AROUND WAIST AREA ON THE LEFT SIDE AND ALSO A LITTLE BELOW THE WAIST AREA. THE PT STATES THAT HE IS FAVORING HIS LEFT LEG. THE PT IS REPORTING THAT THE PAIN HAS BEEN AROUND FOR THE LAST FOUR TO SIX MONTHS. PT IS SCHEDULED TO HAVE BLOOD WORK AND MRI (B)(6) 2012. | AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADD'L INFO HAS BEEN REQUESTED AND IF REC'D, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02227 | 2012-11-07 | 2012-10-15 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT, MECHANICALLY ASSISTED CREVICE CORROSION WITH ADVERSE LOCAL TISSUE REACTION, PSEUDO TUMOR WITH EFFUSION, WITH TOXICITY SECONDARY TO COBALT S/P MODULAR NECK/STEM IMPLANT. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR. INFO RECEIVED INDICATED THAT THE PT REQUESTED TO SEND THE EXPLANTED DEVICES TO INDEPENDENT FACILITY FOR TESTING. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS: MFR # 9616680-2012-00994. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02224 | 2012-11-07 | 2010-09-14 | I, n, j, u, r, y | Manufacturer report | UNKNOWN #7 REJUVENATE STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: PT HAS BEEN EXPERIENCING PAIN IN HIP AREA SINCE SURGERY. PT STATES THAT PAIN WAKES HER UP AT NIGHT. PT IS REPORTING THAT SHE SEES NO SWELLING BUT FEELS INFLAMMATION AND ALSO TENDERNESS IN HIP. PT ALSO SAYS THAT SHE HAS SORENESS WHEN WALKING IN THE JOINT AREA. PT STATES THAT SHE HAS HAD SIX WEEKS OF PHYSICAL THERAPY BUT IS STILL EXPERIENCING PAIN IN HER HIP. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: REJUVENATE MODULAR NECK 34MM V40, CAT # NLS-340000B, LOT # 29682601; RESTORATION ADM X3 INS 28/48, CAT # 1236-2-848, LOT # 32613101; RESTORATION ADM CUP W/HA, CAT # 1235-2481, LOT # G2820726; 8MM +4 LFIT V40 HEAD, CAT # 6260-9-228, LOT # 28440804. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02225 | 2012-11-07 | 2012-10-11 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT, PT COMPLAINED OF PAIN 7 WEEKS AGO WITH NUMBNESS AND TINGLING TO LEG. PT HAS HER HIP ASPIRATED BY RADIOLOGY DISCOVERING PUS AND DISCOLORED FLUID ON OCTOBER 11 AND SCHEDULED FOR RESECTION SURGERY ON 10/12. DR (B)(6) SENT INTRA-OPERATIVE CULTURES AND TISSUE SAMPLES, IMPLANTS WERE RESECTED VERY LITTLE FRETTING AND CORROSION NOTED ON BOTH TAPERS, WILL CLEAN AND SUBMIT FOR ANALYSIS. CEMENT SPACER WAS PLACED IN PT. | ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. WHEN COMPLETED, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS: MFR # 9616680-2012-00992. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02233 | 2012-11-07 | 2012-04-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT HIP REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: PATIENT IS EXPERIENCING SORENESS IN HIP AREA WHEN SEATED FOR LONG PERIODS. PATIENT STATES THAT THE PAIN STARTED 4-6 MONTHS AGO. PATIENT ALSO STATES THAT HE HAS HAD BLOOD WORK AND AN MRI DONE. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02232 | 2012-11-07 | 2012-04-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT HIP REJUVENATE/ABGII | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT IS COMPLAINING OF PAIN AND SWELLING IN HIS LEFT HIP. PATIENT HAS BEEN IN PAIN OFF AND ON FOR AT LEAST SIX MONTHS. WHILE PATIENT HAD IMPLANT SURGERY ON HIS RIGHT HIP, PATIENT STATES THAT HE HAD 30CC OF FLUID DRAINED FROM HIS LEFT HIP. PATIENT REPORTS THAT THE FLUID CONTAINED COBALT. PATIENT STATES THAT BLOOD WORK DONE IN (B)(6) SHOWS THE COBALT LEVEL IS 8.8. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02226 | 2012-11-07 | 2012-10-15 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT, POST OPERATIVE DEVELOPMENT MECHANICALLY ASSISTED CREVICE CORROSION WITH ADVERSE LOCAL SOFT TISSUE REACTION WITH TISSUE RECORDS (PSUEDOTUMAR). | AN EVAL OF THE REPORTED DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS RETAINED BY THE PT AND WAS NOT RETURNED TO THE MFR. THIS EVENT BECAME A LEGAL CLAIM. NO ADD'L INFO IS AVAILABLE AT THIS TIME DUE TO THE ONGOING LITIGATION. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THIS IS THE SAME PT/EVENT AS: MFR # 9616680-2012-00993. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02252 | 2012-11-07 | 2012-10-18 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORHTOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM, PLEASE BE AWARE THAT THIS OFFICE HAS BEEN RETAINED TO REPRESENT THE FOLLOWING INDIVIDUAL IN RELATION TO CLAIMS FOR PERSONAL INJURY RESULTING FROM THE IMPLANTATION OF STRYKER'S RECALLED REJUVENATE OR ABG II HIP DEVICE. UPDATE: THIS IS A FORMER CLAIM. PLAINTIFF ALLEGES LEFT REJUVENATE HIP IMPLANTED ON (B)(6) 2011 FAILED CAUSING INJURY AND ELEVATED METAL ION LEVELS. PLAINTIFF HAS NOT YET SCHEDULED REVISION SURGERY. | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02223 | 2012-11-07 | 2011-05-09 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT: PATIENT HAS BEEN EXPERIENCING PAIN SINCE REHABILITATION POST SURGERY. PT IS REPORTING BEING UNABLE TO PUT ON HER SOCK AND EXPERIENCING FATIGUE. SHE IS EXPERIENCING SWELLING AND PROBLEMS WALKING. PT IS ALSO REPORTING THAT SHE IS LIMPING. PT STATES THAT SHE HAS BEEN TESTED FOR CHROMIUM AND COBALT, HAD X-RAYS, AND HAD MRIS DONE. PT IS ALSO REPORTING THAT SHE IS SCHEDULED FOR REVISION SURGERY ON (B)(6) 2012, AT (B)(6). | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: REJUVENATE MODULAR NECK, CAT# NLV-3408007, LOT# 27305402. TRIDENT HEMISPHERICAL CLUSTER HOLE SHELL, CAT# 502-11-52E, LOT# 36555701. TRIDENT 0DEG X3 INSERT 36MM ID, CAT#623-00-36E, LOT# 35077702. DELTA V-40 CERAMIC HEAD 35/-2,5, CAT# 6570-0-436, LOT# 35077702, 6.5 CANCELLOUS BONE SCREW 25MM, CAT# 2030-6525-1, LOT# MJT4HN. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. THIS EVENT BECAME A LEGAL CLAIM. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME DUE TO THE ONGOING LITIGATION. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02250 | 2012-11-07 | 2012-10-18 | I, n, j, u, r, y | Manufacturer report | UNKNOWN ABGII OR REJUVENATE HIP SYSTEM | IMPLANT | STRYKER ORHTOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM, PLEASE BE AWARE THAT THIS OFFICE HAS BEEN RETAINED TO REPRESENT THE FOLLOWING INDIVIDUAL IN RELATION TO CLAIMS FOR PERSONAL INJURY RESULTING FORM THE IMPLANTATION OF STRYKER'S RECALLED REJUVENATE OR ABGG II HIP DEVICE. | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02235 | 2012-11-07 | 2012-01-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT THE PATIENT RECEIVED A LETTER FROM HER SURGEON EARLIER THIS YEAR. THE PATIENT STATES THAT SHE BEGAN HAVING A SHARP SHOOTING PAIN THAT RUNS BETWEEN THE RIGHT KNEE AND THE OUTSIDE OF HER HIP IN (B)(6) 2012. THE PAIN OCCURS WHEN SHE WALKS AND PUTS WEIGHT ON THE LEG. THE PAIN CAN LAST FROM A FEW MINUTES TO A FEW HOURS. SHE OCCASIONALLY HAS SWELLING IN THE SAME AREA. SHE STATES THAT ACCORDING TO HER DOCTOR IT APPEARS THAT THE STEM MAY HAVE MOVED INSIDE THE BONE. SHE TOOK BLOOD TEST AND IS SCHEDULED FOR AN MRI. SHE WILL HAVE A FOLLOW-UP VISIT ONCE ALL TEST RESULTS ARE RECEIVED. | CATALOG NUMBERS AND LOT CODES OF OTHER DEVICES LISTED IN THIS REPORT: DESCRIPTION: REJUVENATE MODULAR NECK 30MM V40, CAT # NLS-300000B, LOT # 35368002; DESCRIPTION: DELTA V-40 CERAMIC HEAD 28/0, CAT # 6570-0-128, LOT # 37385802; DESCRIPTION: RESTORATION (TM) ADM. CUP W/HA, CAT # 1235-2-501, LOT # G3035019; DESCRIPTION: RESTORATION ADM X3 INS 28/50, CAT # 1236-2-850, LOT # 35039601. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S EXPERIENCE. A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-11
2249697-2012-02231 | 2012-11-07 | 2010-01-13 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT HIP REJUVENATE/ABG II | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT HAS EXPERIENCED PAIN AND DISCOMFORT WHEN GOING UP AND DOWN THE STAIRS. PATIENT ALSO COMPLAINS THAT HE HAD TROUBLE GETTING INTO HIS CAR. PATIENT EXPLAINS THAT THE PAIN IS TOWARDS THE FRONT OF HIS HIP. PATIENT IS SCHEDULED TO SEE HIS SURGEON FOR TESTING ON (B)(6) 2012. | AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER, BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02245 | 2012-11-09 | 2012-10-15 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 10 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-100000S |  | IT WAS REPORTED THAT, SURGEON DID A LEFT HIP REVISION OF A REJUVENATE STEM BECAUSE PT WAS COMPLAINING OF PAIN IN LEFT HIP. REMOVED BIPOLAR COMPONENT, STEM AND MODULAR NECK. | THE FOLLOWING OTHER HIP DEVICES WERE ALSO LISTED IN THIS REPORT: LRG TAP PRI MOD NCK 0DEG 38MM, CAT# NLS-380000B, LOT# 32143201, 26MM STD LFIT V40 HEAD, CAT# 6260-9-126, LOT# 38574405, UHR BIPOLAR 26X58MM, CAT# UH1-58-26, LOT# MKRL77. IT CANNOT BE DETERMINED WHICH, IF ANY OF THESE DEVICES MAY HAVE CAUSED OR CONTRIBUTED TO THE PT'S PAIN. AN EVAL OF THE REPORTED DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR. PT HAS LEGAL REPRESENTATION AND OBTAINED IMPLANTS. ADD'L INFO INCLUDING X-RAYS AND MEDICAL RECORDS WAS NOT MADE AVAILABLE EITHER. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02305 | 2012-11-09 | 2012-10-18 | I, n, j, u, r, y | Manufacturer report | UNKNOWN ABGII OR REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS | NA | UNKNOWN |  | IT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM, PLEASE BE AWARE THAT THIS OFFICE HAS BEEN RETAINED TO REPRESENT THE FOLLOWING INDIVIDUAL IN RELATION TO CLAIM FOR PERSONAL INJURY RESULTING FROM THE IMPLANTATION OF STRYKER'S RECALLED REJUVENATE OR ABG II HIP DEVICE. | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION, NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED, IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02298 | 2012-11-09 | 2012-10-18 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT SURGEON REVISED PATIENT'S LEFT HIP DUE TO PAIN, INCREASED METAL IONS IN BLOOD AND FLUID BUILD UP. WHILE PERFORMING THE REVISION SURGEON ALSO NOTED SIGNIFICANT NECROTIC TISSUE AND SIGNIFICANT METAL WEAR AT THE STEM NECK JUNCTION WHICH INCLUDED SOME SORT OF BLACK MATERIAL. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER DUE TO HOSPITAL POLICY. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01024 | 2012-11-09 | 2012-10-18 | I, n, j, u, r, y | Manufacturer report | UNKNOWN PRODUCT - ABGII OR REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM, PLEASE BE AWARE THAT THIS OFFICE HAS BEEN RETAINED TO REPRESENT THE FOLLOWING INDIVIDUAL IN RELATION TO CLAIM FOR PERSONAL INJURY RESULTING FROM THE IMPLANTATION OF STRYKER'S RECALLED REJUVENATE OR ABG II HIP DEVICE. | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION, NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED, IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02304 | 2012-11-09 | 2012-10-18 | I, n, j, u, r, y | Manufacturer report | UNKNOWN ABGII OR REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNKNOWN |  | IT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM, PLEASE BE AWARE THAT THIS OFFICE HAS BEEN RETAINED TO REPRESENT THE FOLLOWING INDIVIDUAL IN RELATION TO CLAIM FOR PERSONAL INJURY RESULTING FROM THE IMPLANTATION OF STRYKER'S RECALLED REJUVENATE OR ABG II HIP DEVICE. | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION, NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED, IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02301 | 2012-11-09 | 2012-10-18 | I, n, j, u, r, y | Manufacturer report | UNKNOWN ABGII OR REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNKNOWN |  | IT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM, PLEASE BE AWARE THAT THIS OFFICE HAS BEEN RETAINED TO REPRESENT THE FOLLOWING INDIVIDUAL IN RELATION TO CLAIM FOR PERSONAL INJURY RESULTING FROM THE IMPLANTATION OF STRYKER'S RECALLED REJUVENATE OR ABG II HIP DEVICE. | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION, NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED, IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01020 | 2012-11-09 | 2012-10-18 | I, n, j, u, r, y | Manufacturer report | UNKNOWN ABGII OR REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM, PLEASE BE AWARE THAT THIS OFFICE HAS BEEN RETAINED TO REPRESENT THE FOLLOWING INDIVIDUAL IN RELATION TO CLAIM FOR PERSONAL INJURY RESULTING FROM THE IMPLANTATION OF STRYKER'S RECALLED REJUVENATE OR ABG II HIP DEVICE. | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION, NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. IT ADDITIONAL INFORMATION IS RECEIVED, IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02269 | 2012-11-09 | 2012-10-15 | I, n, j, u, r, y | Manufacturer report | DELTA V-40 CERAMIC HEAD 32/+4 | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | STRYKER ORTHOPAEDICS-MAHWAH | NA | 6570-0-232 |  | IT WAS REPORTED THAT DR (B)(6) BROUGHT PATIENT TO SURGERY FOR EXPLANTS OF THA AND PLACEMENT OF ANTIBIOTIC SPACER. HE DID DIRECT LATERAL APPROACH. WHEN JOINT WAS INCISED A PUS LIKE FLUID WAS DISCOVERED. CULTURES WERE TAKEN. AFTER CAREFUL DEBRIDEMENT, DR. POPPED HEAD OF TRUNNION AND DISSOCIATED THE LINER FROM THE CUP. ABNORMAL BURNISHING WAS DISCOVERED ON THE HEAD IN THE HEAD/TRUNNION JUNCTION. BURNISHING AND PITTING WERE ALSO NOTED ON THE TRUNNION BUT STEM WAS WELL FIXED. ABNORMAL WEAR WAS ALSO DISCOVERED ON THE ACETABULAR LINER IN THE LIPPED/ELEVATED AREA. PARTS WERE REPLACED WITH A 32+4 LFIT K-40 HEAD AND XFIRE LINER AS THE STEM AND CUP WERE FIXED AND CULTURE WAS NEGATIVE. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR DUE TO HOSPITAL POLICY. ADD¿L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD¿L INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02243 | 2012-11-09 | 2011-10-03 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: PT HAS BEEN IN CONSTANT PAIN SINCE HAVING SURGERY. PT IS REPORTING HAVING A THROBBING ACHE. PT IS ALSO COMPLAINING OF STIFFNESS, WALKING ON A BRUISE AND NUMBNESS WHEN SHE IS INACTIVE FOR PERIODS OF TIME. PT STATES THAT SHE HAS PAIN IN HER FEMUR AREA. PT ALSO STATES THAT SHE CAN'T STOOP, SQUAT, USE A LADDER, PULL, OR LIFT. BENDING AT THE WAIST FEELS LIKE IT IS TOUCHING HER HIP AND THAT THE IMPLANT IS LOOSE. PT STATES THAT SHE HAS HAD BLOOD WORK, X-RAYS AND MRI DONE. THE X-RAYS AND MRI SHOWS NO MASS. THE BLOOD WORK FROM (B)(6) 2012, SHOW THAT COBALT LEVELS ARE AT 14.3 AND THE CHROMIUM LEVEL IS AT 2.0. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01022 | 2012-11-09 | 2012-10-18 | I, n, j, u, r, y | Manufacturer report | UNKNOWN ABGII OR REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM, PLEASE BE AWARE THAT THIS OFFICE HAS BEEN RETAINED TO REPRESENT THE FOLLOWING INDIVIDUAL IN RELATION TO CLAIM FOR PERSONAL INJURY RESULTING FROM THE IMPLANTATION OF STRYKER'S RECALLED REJUVENATE OR ABG II HIP DEVICE. | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION, NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED, IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02260 | 2012-11-09 | 2012-10-15 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | HIP JOINT METAL/CERAMIC/POLYMER SEMI-CONSTRAINED CEMENTED OR NONPOROUS UNCEMENTE | STRYKER ORTHOPAEDICS-MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT, THROUGH ATTORNEY FOR THE PT, AS A RESULT OF A LEGAL CLAIM, THAT ALLEGEDLY THE PT SUSTAINED INJURIES AND COMPLICATIONS STEMMING FROM HER LEFT HIP IMPLANT. | THE INFO IN THIS REPORT WAS PROVIDED BY STRYKER ORTHOPAEDICS LEGAL AFFAIRS DEPARTMENT. NO ADD'L INFO IS AVAILABLE AT THIS TIME. THE DEVICE IS NOT AVAILABLE DUE TO THE ONGOING LITIGATION. SHOULD DEVICE OR ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02297 | 2012-11-09 | 2012-10-18 | I, n, j, u, r, y | Manufacturer report | UNKNOWN ABGII OR REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM, PLEASE BE AWARE THAT THIS OFFICE HAS BEEN RETAINED TO REPRESENT THE FOLLOWING INDIVIDUAL IN RELATION TO CLAIM FOR PERSONAL INJURY RESULTING FROM THE IMPLANTATION OF STRYKER'S RECALLED REJUVENATE OR ABG II HIP DEVICE. | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION, NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED, IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01023 | 2012-11-09 | 2011-07-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK 34MM V40 | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | NLS-340000B |  | IT WAS REPORTED THAT THE PT HAD AN MRI IN (B)(6) 2012 THAT SHOWED DETERIORATION IN BOTH JOINTS OF THE HIP. ALSO SHE SAID HIS BLOOD TEST LEVELS ARE ACCEPTABLE "FOR NOW." | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-11
2249697-2012-02276 | 2012-11-09 | 2011-02-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT; PATIENT COMPLAINS OF PAIN THAT BEGAN A COUPLE OF MONTHS AGO. PAIN IS CONSTANT BUT IS ALSO ASSOCIATED WITH MOVEMENT. HER COBALT LEVEL IS 7 AND MRI SHOWS DETERIORATION OF THE IMPLANT. PATIENT IS SCHEDULED FOR FURTHER BLOOD WORK AND FOLLOW UP. PER PATIENT'S SPOUSE, THE SURGEON IS CONSIDERING A REVISION SURGERY. THE SURGEON WILL DETERMINE IF THE PATIENT REQUIRES A REVISION WHEN ALL THE FURTHER BLOOD WORK IS COMPLETED. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01010 | 2012-11-09 | 2010-05-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT: PATIENT IS IN A GREAT DEAL OF PAIN IN HIP AREA. PATIENT HAS BEEN TO SEE DOCTOR AND HAD BLOOD WORK, MRI AND X-RAYS DONE. PATIENT IS SCHEDULED TO SEE SURGEON TO RECEIVE RESULTS OF THE TESTS THAT WERE PERFORMED. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02286 | 2012-11-09 | 2010-08-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT IS REPORTED THAT: PATIENT REPORTS THAT SHE NEEDS REVISION SURGERY DUE TO TWO TUMORS. PATIENT STATES THAT SHE CANNOT RAISE HER LEG FULLY AND HAS DIFFICULTY RUNNING. PATIENT HAS PAIN ON AND OFF. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01016 | 2012-11-09 | 2010-08-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT IS REPORTED THAT: PATIENT REPORTS THAT SHE NEEDS REVISION SURGERY DUE TO TWO TUMORS. PATIENT STATES THAT SHE CANNOT RAISE HER LEG FULLY AND HAS DIFFICULTY RUNNING. PATIENT HAS PAIN ON AND OFF. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01001 | 2012-11-09 | 2012-11-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | NLS-380000B |  | IT WAS REPORTED THAT THE PATIENT HAS NO PHYSICAL SYMPTOMS. PATIENT STATED THAT HIS BLOOD LEVELS ARE ELEVATED. SURGEON HAS COMMUNICATED TO PATIENT THAT HIS HIP NEEDS TO COME OUT. PATIENT WILL FOLLOW UP WITH SURGEON TO SCHEDULE REVISION DATE. ADDITIONAL INFORMATION INDICATED THAT THE DOCTOR REVISED PATIENT'S LEFT HIP DUE TO DRASTICALLY ELEVATED COBALT AND CHROMIUM LEVELS. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-11
9616680-2012-01006 | 2012-11-09 | 2011-02-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK 34MM V40 | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | NLS-340000B |  | IT WAS REPORTED THAT: PATIENT COMPLAINS OF PAIN THAT BEGAN A COUPLE OF MONTHS AGO. PAIN IS CONSTANT BUT IS ALSO ASSOCIATED WITH MOVEMENT. HER COBALT LEVEL IS 7 AND MRI SHOWS DETERIORATION OF THE IMPLANT. PATIENT IS SCHEDULED FOR FURTHER BLOOD WORK AND FOLLOW UP. PER PATIENT'S SPOUSE, THE SURGEON IS CONSIDERING A REVISION SURGERY. THE SURGEON WILL DETERMINE IF THE PATIENT REQUIRES A REVISION WHEN ALL THE FURTHER BLOOD WORK IS COMPLETED. THEY WANT STRYKER TO PROVIDE ASSISTANCE WITH ALL THE ASSOCIATED EXPENSES WITH THE RECALL. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02280 | 2012-11-09 | 2012-10-16 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 12 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-120000S |  | IT WAS REPORTED THAT IMPLANTS WERE EXPLANTED DUE TO ADVERSE LOCAL TISSUE REACTION. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER DUE TO HOSPITAL POLICY. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED, BUT NOT MADE AVAILABLE. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE AT LATER TIME, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02251 | 2012-11-09 | 2012-10-15 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 10 | IMPLANT | STRYKER ORTHOPEDICS MAHWAH | NA | SPT-100000S |  | IT WAS REPORTED THAT, PT CAME IN COMPLAINING AND HAD SWELLING OF THE LEFT HIP. SURGEON DRAINED ARE AND A MILKY LIQUID CAME OUT. THERE WAS NO SIGN OF INFECTION AND NO BACTERIAL INFECTION. SURGEON REVISED AND NOTED THAT THE HIP DID APPEAR TO HAVE AN A-CELLULAR/VASCULAR TISSUE. SURGEON NOTICED A SLIGHT CORROSION OF THE NECK AT JUNCTURE OF THE REJUVENATE STEM. | WHEN COMPLETED, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02281 | 2012-11-09 | 2010-05-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: PATIENT IS IN A GREAT DEAL OF PAIN IN HIP AREA. PATIENT HAS BEEN TO SEE DOCTOR AND HAD BLOOD WORK, MRI AND X-RAYS DONE. PATIENT IS SCHEDULED TO SEE SURGEON TO RECEIVE RESULTS OF THE TESTS THAT WERE PERFORMED. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01027 | 2012-11-09 | 2012-10-18 | I, n, j, u, r, y | Manufacturer report | UNKNOWN ABGII OR REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS | NA | UNKNOWN |  | IT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM, PLEASE BE AWARE THAT THIS OFFICE HAS BEEN RETAINED TO REPRESENT THE FOLLOWING INDIVIDUAL IN RELATION TO CLAIM FOR PERSONAL INJURY RESULTING FROM THE IMPLANTATION OF STRYKER'S RECALLED REJUVENATE OR ABG II HIP DEVICE. | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION, NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED, IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01014 | 2012-11-09 | 2012-01-25 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE MODULAR NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT IS REPORTED THAT: PATIENT IS EXPERIENCING PAIN SINCE JULY. MRI SHOWS FLUID IN JOINTS. PATIENT HAS BEEN IN PHYSICAL THERAPY. PATIENT REPORTS THAT SURGEON MAY NEED TO PERFORM REVISION SURGERY. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02294 | 2012-11-09 | 2012-10-17 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT COMPONENTS WERE EXPLANTED DUE TO ADVERSE LOCAL TISSUE REACTION. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR DUE TO HOSPITAL POLICY. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED, BUT NOT MADE AVAILABLE. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02283 | 2012-11-09 | 2012-01-25 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE MODULAR STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT IS REPORTED THAT: PATIENT IS EXPERIENCING PAIN SINCE (B)(6). MRI SHOWS FLUID IN JOINTS. PATIENT HAS BEEN IN PHYSICAL THERAPY. PATIENT REPORTS THAT SURGEON MAY NEED TO PERFORM REVISION SURGERY. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02242 | 2012-11-09 | 2012-10-15 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT: PT PRESENTED WITH COMPLAINT OF GROIN PAIN. ADD¿L INFO: PT¿S SON CALLED STATING THAT HIS MOTHER HAD PRIMARY SURGERY 14 MONTHS AGO. SHE HAD CORROSION OF THE RECALLED STEM AND WAS REVISED THIS PAST (B)(6) 2012. DURING THE SURGERY, THEY FOUND AN INFECTIOUS POCKET. ALL COMPONENTS WERE REMOVED AND SHE HAS DRAINS PUT IN. PRIOR TO THE REVISION, THE PT HAD SEVERE PAIN AND WOULD WAKE UP SCREAMING IN THE MIDDLE OF THE NIGHT AND WAS HARDLY ABLE TO WALK. SHE HAD BONE SCAN, MRI AND BLOOD WORK DONE AND IT SHOWED NOTHING GOOD. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR DUE TO HOSPITAL POLICY. ADD¿L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADD¿L INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01003 | 2012-11-09 | 2012-10-15 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER IRELAND, LTD. OSTEONICS (ORTHOPAEDICS DIVI | NA | UNK |  | IT WAS REPORTED THAT METALLOSIS AT JUNCTURE OF THE MODULAR NECK TO STEM ON REJUVENATE IMPLANT. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD¿L INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02266 | 2012-11-09 | 2012-10-15 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT METALLOSIS AT JUNCTURE OF THE MODULAR NECK TO STEM ON REJUVENATE IMPLANT. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD¿L INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02292 | 2012-11-09 | 2012-10-17 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT, PATIENT PRESENTED WITH RIGHT HIP PAIN. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER DUE TO HOSPITAL POLICY. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WILL BE REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01029 | 2012-11-12 | 2012-08-02 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | NLS-380000B |  | IT WAS REPORTED THAT THE PT COMPLAINS OF HIP PAIN THAT STARTED A FEW MONTHS AGO. PT IS SCHEDULED FOR AN MRI TOMORROW AND WILL BE FOLLOWING UP WITH HIS SURGEON. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-11
9616680-2012-01037 | 2012-11-12 | 2010-10-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK 34MM V40 | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | NLS-340000B |  | IT WAS REPORTED THAT PATIENT COMPLAINS OF SORENESS IN HIP AREA AND BACK. BLOOD TESTS SHOWED COBALT LEVELS WERE HIGH. CHROMIUM LEVELS ARE NORMAL. MRI IS SCHEDULED AND PATIENT WILL HAVE FLUID ASPIRATED FROM JOINT. SHE WILL FOLLOW UP WITH SURGEON. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01032 | 2012-11-12 | 2012-06-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK 30MM V40 | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | NLS-300000B |  | IT WAS REPORTED THAT: HUSBAND (B)(6) CALLED FOR HIS WIFE. HE CLAIMED THAT HIS WIFE HAS HAD PAIN IN HER RIGHT LEG FOR TWO TO THREE MONTHS. HER MRI SHOWED TUMOR AROUND THE IMPLANT. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01042 | 2012-11-12 | 2011-12-12 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK 30MM V40 | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | NLS-300000B |  | IT WAS REPORTED THAT PATIENT RECEIVED LETTER FROM DOCTOR AND THE DOCTOR MENTIONED THAT SHE WILL PROBABLY NEED A HIP REVISION SURGERY. SHE HAS AN MRI SCHEDULED FOR (B)(6) /2012. SHE HAS HAD BLOODWORK AND X-RAYS COMPLETED ON (B)(6) 2012. SHE HAS HAD TENDERNESS SINCE THE SURGERY AND HER BLOOD LEVELS ARE ELEVATED. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01040 | 2012-11-12 | 2011-10-25 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK 38MM V40 | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | NLV-380800G |  | IT WAS REPORTED THAT PATIENT SAID THAT HE HAD NO SYMPTOMS OTHER THAN LITTLE DISCOMFORT WALKING. HIS MRI SHOWED FLUID ACCUMULATION AT THE JOINT. HIS BLOOD TESTS RESULTS INDICATED HIGH LEVEL OF CHROMIUM AND COBALT. PATIENT SAID THAT HE HAS NO SWELLING. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01028 | 2012-11-12 | 2011-08-31 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK 34MM V40 | UNKNOWN | STRYKER ORTHOPAEDICS CORK | NA | NLS-340000B |  | PT STATES THAT SHE HAS DISCOMFORT IN THE LEFT HIP. SHE WENT TO THE DOCTOR ON (B)(6) 2012. SHE HAD BLOOD LABS DONE ON (B)(6) 2012. SHE RECEIVED THE RESULTS FROM HER DOCTOR'S OFFICE ON (B)(6) 2012 SAYING THAT HER LEVELS WERE ELEVATED. SHE IS BEING SCHEDULED FOR AN MRI THROUGH THE DOCTOR'S OFFICE. THE PT STATES THAT SHE IS HAVING TROUBLE MANEUVERING STAIRS. SHE WAS GIVEN A PRESCRIPTION FOR ANTI-INFLAMMATORY MEDICATION ON (B)(6) 2012. SHE BEGAN HAVING TENDERNESS AND PERIODIC NERVE TINGLING IN THE HIP (B)(6) 2012. SHE ALSO HAS OCCASIONAL DISCOMFORT IN THE GROIN AREA. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. CATALOG NUMBER AND LOT CODE OF OTHER DEVICE LISTED IN THIS REPORT: CAT # NLS-300000B, LOT # 36267401, DESCRIPTION: REJUVENATE MODULAR NECK 30MM V40. IT CANNOT BE DETERMINED WHICH MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02334 | 2012-11-13 | 2012-10-22 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT, ALTR. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01047 | 2012-11-13 | 2011-05-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT IS REPORTED THAT: PATIENT WAS EXPERIENCING PAIN IN HIP. PATIENT REPORTS THAT BLOOD TESTS SHOWED ELEVATED COBALT LEVELS AND A LOT OF METAL IN BLOOD. PATIENT STATES THAT MUSCLE WAS CORRODING, AND THIS LED TO REVISION SURGERY IN (B)(6) 2012. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02332 | 2012-11-13 | 2011-09-12 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IS REPORTED THAT: PATIENT COMPLAINS OF PAIN AT HIP JOINT SINCE THE SURGERY; ELEVATED METAL LEVELS AND FRETTING. REVISION SCHEDULED FOR (B)(6) 2013. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02333 | 2012-11-13 | 2011-05-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT IS REPORTED THAT: PATIENT WAS EXPERIENCING PAIN IN HIP. PATIENT REPORTS THAT BLOOD TESTS SHOWED ELEVATED COBALT LEVELS AND A LOT OF METAL IN BLOOD. PATIENT STATES THAT MUSCLE WAS CORRODING, AND THIS LED TO REVISION SURGERY IN (B)(6) 2012. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01046 | 2012-11-13 | 2011-09-12 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT IS REPORTED THAT: PATIENT COMPLAINS OF PAIN AT HIP JOINT SINCE THE SURGERY; ELEVATED METAL LEVELS AND FRETTING. REVISION SCHEDULED FOR (B)(6) 2013. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01049 | 2012-11-14 | 2009-12-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | NLS-340000B |  | IT WAS REPORTED THAT PT'S WIFE CALLED AND SAID THAT SHE LEARNED ABOUT THE RECALL FROM DR. ON (B)(6) 2012 DURING HER HUSBAND'S FOLLOW UP VISIT. SHE SAID THAT HER HUSBAND HAD AN MRI IN (B)(6) 2012 THAT SHOWED DETERIORATION IN BOTH JOINTS OF THE HIP. ALSO SHE SAID HIS BLOOD TEST LEVELS ARE ACCEPTABLE FOR NOW. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02340 | 2012-11-14 | 2012-01-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: PAIN FIRST STARTED IN (B)(6) 2012. THE PT STATES THAT THE PAIN IS USUALLY IN THE GROIN AREA DOWN TO THE TOP OF THIGH. PT IS ALSO EXPERIENCING NUMBNESS BELOW THE KNEE. PT STATES THAT THE KNEE ALWAYS FEELS NUMB. PT STATE THAT IN, (B)(6) 2012, PAIN WAS INCREASING AND BECOMING WORSE AND BY JUNE THE PAIN WAS SEVERE. PT IS REPORTING THAT THE PAIN ALSO WS IN THE MIDDLE OF HER BUTTOCK AREA. PT STATES THAT SHE IS NOW LIMPING AND HAS A FOOT DROP. PT HAS HAD BLOOD WORK AND THE RESULTS ARE PENDING. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01061 | 2012-11-14 | 2011-12-13 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT: PT HAD PRIMARY SURGERY AND STATED THAT BETWEEN TIME OF SURGERY AND REVISION SHE PULLED QUAD MUSCLE AND HAD TO REJUVENATE IMPLANT REMOVED. PT STATES THAT WHEN SHE GOT HOME FROM THE PRIMARY SURGERY, SHE WAS HAVING PROBLEMS WITH HIP. PT STATES THAT THE HIP FELT LIKE DEAD WEIGHT, PT WENT INTO THE EMERGENCY ROOM AND HER HIP WAS DRAINED TO RELIEVE FLUID AND SWELLING. PT ENDED UP HAVING REVISION SURGERY ON LEFT HIP ON (B)(6) 2011 AT (B)(6) MEDICAL CENTER. PT STATES THAT SHE HAD ANOTHER STRYKER IMPLANT BUT DOESN'T KNOW WHICH ONE. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01054 | 2012-11-14 | 2012-07-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICSS CORK | NA | UNK |  | IT WAS REPORTED THAT: PT IS EXPERIENCING DISCOMFORT IN HIP AREA. PT IS ALSO REPORTING PAIN IN UPPER BUTTOCK AREA. PT STATES THAT THE PAIN RADIATES FROM THE HIP TO THE FRONT OF THE KNEE. PAIN FIRST STARTED IN (B)(6). PT STATES THAT SHE HAS HAD BLOOD WORK AND X-RAYS ON (B)(6) 2012. PT IS SCHEDULED TO HAVE MRI DONE. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01058 | 2012-11-14 | 2012-01-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT: PAIN FIRST STARTED IN (B)(6) 2012. THE PT STATES THAT THE PAIN IS USUALLY IN THE GROIN AREA DOWN TO THE TOP OF THIGH. PT IS ALSO EXPERIENCING NUMBNESS BELOW THE KNEE. PT STATES THAT THE KNEE ALWAYS FEELS NUMB. PT STATE THAT IN, (B)(6) 2012, PAIN WAS INCREASING AND BECOMING WORSE AND BY (B)(6) THE PAIN WAS SEVERE. PT IS REPORTING THAT THE PAIN ALSO WAS IN THE MIDDLE OF HER BUTTOCK AREA. PT STATES THAT SHE IS NOW LIMPING AND HAS A FOOT DROP. PT HAS HAD BLOOD WORK AND THE RESULTS ARE PENDING. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02344 | 2012-11-14 | 2011-07-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: PT IS EXPERIENCING PAIN. PT IS REPORTING THAT THE PAIN HAS EXISTED SINCE SHE HAD SURGERY AND HAS NOT GOTTEN ANY BETTER. PT STATES THAT HIP IMPLANT IS AFFECTING HER WALK. PT ALSO STATES THAT WHEN SHE STANDS FOR A PERIOD OF TIME AND DOING CHORES SUCH AS COOKING, THE THIGH MUSCLE WOULD PARALYZE AND THE ONLY WAY IT WOULD MOVE IS IF IT IS MANUALLY MOVED WITH HER HANDS. PT STATES THAT SURGEON HAS TAKEN BLOOD TESTS AND X-RAYS AND WANTS TO SCHEDULE AN MRI. PT STATES THAT THE COBALT LEVEL IN HER BLOOD IS ELEVATED. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT REMAINED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01063 | 2012-11-14 | 2011-07-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT: PT IS EXPERIENCING PAIN. PT IS REPORTING THAT THE PAIN HAS EXISTED SINCE SHE HAD SURGERY AND HAS NOT GOTTEN ANY BETTER. PT STATES THAT HIP IMPLANT IS AFFECTING HER WALK. PT ALSO STATES THAT WHEN SHE STANDS FOR A PERIOD OF TIME AND DOING CHORES SUCH AS COOKING, THE THIGH MUSCLE WOULD PARALYZE AND THE ONLY WAY IT WOULD MOVE IS IF IT IS MANUALLY MOVED WITH HER HANDS. PT STATES THAT SURGEON HAS TAKEN BLOOD TESTS AND X-RAYS AND WANTS TO SCHEDULE AN MRI. PT STATES THAT THE COBALT LEVEL IN HER BLOOD IS ELEVATED. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT REMAINED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01065 | 2012-11-14 | 2010-10-10 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT: PT IS REPORTING DISCOMFORT IN HIP SINCE LAST (B)(6). PT IS REPORTING OFF AND ON PAIN IN HIP AREA. PT STATES THAT THERE IS NO SWELLING OR SHARP PAIN, IT IS AN ACHE. PT ALSO STATES THAT SHE IS CURRENTLY USING A CANE WITH A SEAT IF SHE HAS TO WALK FOR LONG PERIODS OF TIME. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT REMAINED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01068 | 2012-11-14 | 2012-10-23 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT, PT HAS HIGH COBALT LEVELS. | AN EVAL OF THE REPORTED DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS RETAINED BY THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02343 | 2012-11-14 | 2011-12-13 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP STEM | IMPLANT | SRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: PT HAD PRIMARY SURGERY AND STATED THAT BETWEEN TIME OF SURGERY AND REVISION SHE PULLED QUAD MUSCLE AND HAD TO REJUVENATE IMPLANT REMOVED. PT STATES THAT WHEN SHE GOT HOME FROM THE PRIMARY SURGERY, SHE WAS HAVING PROBLEMS WITH HIP. PT STATES THAT THE HIP FELT LIKE DEAD WEIGHT, PT WENT INTO THE EMERGENCY ROOM AND HER HIP WAS DRAINED TO RELIEVE FLUID AND SWELLING. PT ENDED UP HAVING REVISION SURGERY ON LEFT HIP ON (B)(6) 2011 AT (B)(6) MEDICAL CENTER. PT STATES THAT SHE HAD ANOTHER STRYKER IMPLANT BUT DOESN'T KNOW WHICH ONE. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02346 | 2012-11-14 | 2010-10-10 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: PT IS REPORTING DISCOMFORT IN HIP SINCE LAST (B)(6). PT IS REPORTING OFF AND ON PAIN IN HIP AREA. PT STATES THAT THERE IS NO SWELLING OR SHARP PAIN, IT IS AN ACHE. PT ALSO STATES THAT SHE IS CURRENTLY USING A CANE WITH A SEAT IF SHE HAS TO WALK FOR LONG PERIODS OF TIME. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT REMAINED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02350 | 2012-11-14 | 2012-10-23 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 10 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-100000S |  | IT WAS REPORTED THAT THERE WAS HIP PAIN. | AN EVAL OF THE REPORTED DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02354 | 2012-11-14 | 2012-10-25 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 11 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-110000S |  | IT WAS REPORTED THAT, MECHANICALLY ASSISTED CREVICE CORROSION AND SECONDARY LOCAL ADVERSE SOFT TISSUE REACTION WITH PSEUDOTUMOR FORMATION SECONDARY TO REJUVENATE MODULAR NECK-STEM FAILURE. | AN EVAL OF THE REPORTED DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS RETAINED BY THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02339 | 2012-11-14 | 2012-07-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: PT IS EXPERIENCING DISCOMFORT IN HIP AREA. PT IS ALSO REPORTING PAIN IN UPPER BUTTOCK AREA. PT STATES THAT THE PAIN RADIATES FROM THE HIP TO THE FRONT OF THE KNEE. PAIN FIRST STARTED IN (B)(6). PT STATES THAT SHE HAS HAD BLOOD WORK AND X-RAYS ON (B)(6) 2012. PT IS SCHEDULED TO HAVE MRI DONE. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02349 | 2012-11-14 | 2012-10-23 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, PT HAS HIGH COBALT LEVELS. | AN EVAL OF THE REPORTED DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS RETAINED BY THE PT AND WAS NOT RETURNED TO THE MFR. ADD'L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
MW5029227 | 2012-11-15 | 2012-10-10 | I, n, j, u, r, y | Voluntary report |  | ABG II MODULAR HIP | STRYKER |  |  |  | MY STRYKER ABG II HIP TOTAL IMPLANT IS CAUSING LOT OF PAIN AND LOSS OF INCOME. MY ORTHOPEDIC SURGEON SAYS IT NEEDS TO BE REPLACED. I AM LOSING WORK, NOW ABOUT HALF DAY PER DAY. THE PAIN BEGAN EARLIER THAN (B)(6), BUT IT WAS LOW PAIN, AND WHEN IT GOT ENOUGH TO WORRY ME I WENT TO MY ORTHOPEDIC SURGEON. HE SAID IT IS A RECALLED HIP IMPLANT. I HAVE BEEN READING ON FDA WEB SITE AND STRYKER WEB SITE. STRYKER SAYS THEY NOTIFIED USERS APRIL 23, 2012, BUT THEY DIDN'T NOTIFY ME. THAT COULD HAVE SAVE A LOT OF PAIN. THANK YOU FOR THIS PLACE TO LET YOU KNOW ABOUT IT. |  | Y | Y | 0 |  | 2012-11
2249697-2012-02369 | 2012-11-15 | 2012-10-24 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRIFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT, PT'S LEFT HIP WAS REVISED DUE TO PAIN AND HIGH COBALT CHROME SERUM LEVELS. | ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. WHEN COMPLETED, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02365 | 2012-11-15 | 2012-11-05 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT SURGEON REVISED PT'S LEFT HIP DUE TO PAIN. SURGEON STATED THAT HE NOTICED SOME FLUID COLLECTION IN THE PT'S HIP WHILE DOING THE REVISION. REPLACED WITH A RESTORATION MODULAR HIP. SURGEON WILL BE REVISING THE PT'S RIGHT HIP IN A FEW WEEKS AS THIS PT HAD A BILATERAL REJUVENATE HIP DONE 2011. | ONCE THE INVESTIGATION HAS BEEN COMPLETED ANY ADDITIONAL INFO WILL BE REPORTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02362 | 2012-11-15 | 2012-11-05 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNKNOWN |  | IT WAS REPORTED THAT SURGEON REVISED PATIENT'S LEFT REJUVENATE HIP AS PATIENT WAS HAVING PAIN. SURGEON REPLACED WITH RESTORATION MODULAR HIP. | DUE TO THE ONGOING LITIGATION NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01084 | 2012-11-15 | 2012-06-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK 30MM | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | NLS-300000B |  | IT WAS REPORTED THAT; PT STATES THAT SHE BEGAN TO EXPERIENCE PAIN AND DISCOMFORT IN HER RIGHT HIP ABOUT 2-3 MONTHS AGO. THE PAIN BEGAN SUBTLY BUT HAS INCREASED IN INTENSITY AND FREQUENCY OVER THE PAST FEW MONTHS. THE PT STATES THAT SHE CANNOT STAND FOR MORE THAN 20 MINS WITHOUT EXPERIENCING PAIN IN THE FRONT THIGH MUSCLE IN HER RIGHT LEG, ON THE SIDE OF HER RIGHT HIP AND IN THE BUTTOCKS. SHE STATES THAT SHE SEES SOME VISIBLE SWELLING AT THE TOP OF THE IMPLANT INCISION. SHE HAS TROUBLE GETTING UP FROM A CHAIR AND OUT OF A CAR. SHE CAN'T LIE ON HER RIGHT SIDE FOR LONG WITHOUT PAIN. WHEN THE PAIN GETS INTENSE SHE MUST LIE DOWN AND TAKE ASPIRIN. SHE TRIES TO CONTINUE LIGHT EXERCISE BY DAILY GOING FOR SLOW WALKS. THE PT IS CURRENTLY WAITING FOR THE DOCTOR'S STAFF TO SCHEDULE HER FOR AN APPOINTMENT. FURTHER REPORTED, THAT, MECHANICALLY ASSISTED CREVICE CORROSION AND SECONDARY ADVERSE LOCAL SOFT TISSUE REACTION AND PSEUDO TUMOR FORMATION AFTER REJUVENATE FEMORAL COMPONENT DUE TO MODULAR NECK - STEM MECHANISM OF FAILURE. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01073 | 2012-11-15 | 2010-12-09 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | NLS-340000B |  | IT WAS REPORTED THAT THE PT HAS DISCOMFORT PERIODICALLY, SHE IS UNSURE OF WHEN THE PROBLEMS STARTED EXACTLY. SHE REPORTS THAT SHE FEELS DISCOMFORT WHEN SHE SITS FOR TOO LONG. SHE STATES THAT HER BLOOD WORK SHOW ELEVATED LEVELS. HER MRI WAS DONE ONE WEEK AGO. SHE HAS MET WITH SURGEON AND SHE REPORTS TO HAVE FLUID BUILD UP. SHE WILL NEED A REVISION SURGERY TO REMOVE STEM ONLY. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01107 | 2012-11-16 | 2012-02-15 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT: PATIENT IS EXPERIENCING INCREASED PAIN IN HIP AREA. PATIENT STATES THAT, IF SEATED TOO LONG, HIP BECOMES STIFF AND PAINFUL. PATIENT IS REPORTING THAT SHE HAD BLOOD WORK AND MRI. PATIENT IS ALSO REPORTING THAT DOCTOR STATES THAT THERE IS NOTHING DEFINITIVE CURRENTLY. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02390 | 2012-11-16 | 2012-10-25 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT, THE SURGEON REVISED PT'S RIGHT REJUVENATE HIP DUE TO ELEVATED ION LEVELS, PAIN, AND COLLECTION OF FLUID IN THE JOINT. DURING THE REVISION, METALLOSIS AND FRETTING AT THE NECK/STEM JUNCTION WAS NOTICED ALONG WITH A SIGNIFICANT AMOUNT TISSUE AT THE JOINT. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR DUE TO HOSPITAL POLICY. ADD'L INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED, BUT NOT MADE AVAILABLE. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02378 | 2012-11-16 | 2012-02-15 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: PATIENT IS EXPERIENCING INCREASED PAIN IN HIP AREA. PATIENT STATES THAT, IF SEATED TOO LONG, HIP BECOMES STIFF AND PAINFUL. PATIENT IS REPORTING THAT SHE HAD BLOOD WORK AND MRI. PATIENT IS ALSO REPORTING THAT DOCTOR STATES THAT THERE IS NOTHING DEFINITIVE CURRENTLY. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02379 | 2012-11-16 | 2017-10-25 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | PROSTHESIS, HIP, SEMI-CONSTRAINED, UNCEMENTED, METAL/POLYMER, NON-POROUS, CALICU | STRYKER ORTHOPAEDICS-MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED, PATIENT IS EXPERIENCING OCCASIONAL PAIN DOWN THE FRONT OF THE THIGH. THE PAIN IS ALSO ON THE SIDE OF HIS LEG. PATIENT STATES THAT IF HE IS SEATED FOR A PERIOD OF TIME, THEY FEEL THE PAIN ON THE SIDE. PATIENT STATES THAT HE HAS BEEN EXPERIENCING THIS SINCE POST OPERATION. PATIENT STATES THAT HE HAS HAD AN MRI AND BLOOD WORK. RESULTS CURRENTLY ARE NOT SHOWING ANYTHING. PATIENT STATES THAT HE SHOULD RETURN TO THE DOCTOR WITHIN A YEAR FOR A CHECK UP. | AT THIS TIME, THE PATIENT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED BUT BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02382 | 2012-11-16 | 2012-10-23 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT PATIENT HAD COMPLAINED OF PAIN. SHE TESTED POSITIVE FOR AN INFECTION. HER CHROMIUM LEVEL WAS 4 AND COBALT LEVEL WAS 1.3. DR (B)(6) REMOVED ALL IMPLANTS. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02373 | 2012-11-16 | 2012-11-02 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE MODULAR NECK HIP IMPLANTS | IMPLANTS | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED, "NOTICE WAS REC'D FROM ATTORNEY ALLEGING THAT PT HAD SUSTAINED INJURIES AND DAMAGES RESULTING FROM THE RECALLED REJUVENATE MODULAR NECK HIP IMPLANT." | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION, NO ADD'L INFO IS AVAILABLE AT THIS TIME. IF ADD'L INFO IS REC'D, IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
MW5027746 | 2012-11-16 | 2005-02-07 | I, n, j, u, r, y | Voluntary report | TRIDENT HEMISPHERICAL ACETABULAR SHELL | STRYKER CERAMIC ON CERAMIC SYSTEM/LEFT HIP | STRYKER |  | 500-01-6OG |  | ACETABULAR SHELL, INSERT FEMORAL HEAD AND HIP STEM. I HAD PAIN PROBLEMS WITH THIS HIP ALL THE TIME SINCE THEN 6 MONTH CHECK UPS, 1 YEAR CHECK UPS. ALWAYS TOLD TO DO MORE THERAPY AND THE HIPS BEING RECALLED ONLINE, HIP PAIN, POPPING, LOOSENING, SWELLING, BONE LOSS, GROIN PAIN, ETC. |  | N | Y | LAY USER/PATIENT |  | 2012-11
9616680-2012-01116 | 2012-11-16 | 2010-06-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK | UNKNOWN | STRYKER ORTHOPAEDICS CORK | NA | NLS-380000B |  | IT WAS REPORTED THAT: PT COMPLAINS OF PAIN. PT WILL REQUIRE A REVISION SURGERY BECAUSE TESTS HAVE SHOWN FLUID BUILDUP IN THREE AREAS OF HIS LEFT HIP JOINT. THIS HAS DESTROYED THE MUSCLE IN HIS LEG AND HE HAS LOST MUSCLE STRENGTH. TESTS HAVE ALSO SHOWN TITANIUM IN HIS BLOOD. | ADD'L INFO HAS BEEN REQUESTED AND IF REC'D WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
MW5027905 | 2012-11-17 | 2011-10-25 | I, n, j, u, r, y | Voluntary report | STRYKER REJUVENATE | STRYKER REJUVENATE | STRYKER CORP | REJUVENATE #9 132 |  |  | AFTER DIAGNOSIS OF DETERIORATING RIGHT HIP, (ADVANCED RIGHT HIP ARTHRITIS) CONTRACTED WITH DR (B)(6) TO REPLACE HIP WITH A STRYKER REJUVENATE AND PREPARED FOR SURGICAL IMPLANT ON (B)(6) 2011. THREE MONTHS POST-RECOVERY BEGAN TO EXPERIENCE HIP PAIN SIMILAR TO ORIGINAL SYMPTOMS AND AFTER APPROX. EIGHT MONTHS POST-OP RETURNED TO DR (B)(6) FOR SOME PAIN REDUCTION SINCE I WAS EXPERIENCING LITTLE RELIEF FROM THE ORIGINAL PAIN REDUCTION. SHORTLY THEREAFTER I RECEIVED A NOTICE INDICATING THE STRYKER HIP WAS BEING RECALLED. HAVE SINCE BEEN MONITORED BY DR (B)(6) FOR FURTHER COMPLICATIONS. ON (B)(6) 2012¿BLOOD ANALYSIS, (B)(6) 2012 ¿ PHYSICAL THERAPY, (B)(6) 2012 ¿ PHYSICAL THERAPY, (B)(6) 2012 ¿ PHYSICAL THERAPY, (B)(6) 2012 ¿ PHYSICAL THERAPY. |  | N | Y | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02401 | 2012-11-19 | 2012-01-30 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 12 | IMPLANT | STRYKER ORHTOPAEDICS MAHWAH | NA | SPT-120000S |  | IT WAS REPORTED THAT THE PATIENT STARTED NOTICING AROUND (B)(6) 2012, TROUBLE WITH RANGE OF MOTION. PATIENT IS REPORTING POSTEROLATERAL PAIN AND GROIN PAIN. PATIENT ALSO REPORTS THAT HE HAD TROUBLE GETTING UP FROM A CHAIR AND THAT STANDING ON ONE LEG CAUSES SIGNIFICANT PAIN. PATIENT STATES THAT AT NIGHT HE EXPERIENCES SIGNIFICANT NIGHT PAIN AND THAT HE HAS TO SLEEP ON HIS BACK BECAUSE HE CANNOT SLEEP ON HIS SIDE. PATIENT IS ALSO REPORTING ELEVATED LEVELS OF COBALT (8.6) AND CHROMIUM (1.4). PATIENT IS SCHEDULED FOR BILATERAL REVISION. | ADDITIONAL INFORMATION REPORTED VIA SALES REP CONFIRMED THAT THE PATIENT HAD LEFT HIP REPLACED (B)(6) 2012 AND HAD RELIEF FOR 3 MONTHS AND PROGRESSED TO CHRONIC PAIN REQUIRING STEROID INJECTION AND CONTINUOUS ANT-INFLAMMATORY REGIMENT. A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-11
9616680-2012-01122 | 2012-11-19 | 2010-09-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT: PATIENT IS CURRENTLY EXPERIENCING PAIN IN HER RIGHT HIP. PATIENT REPORTS THAT SHE IS CURRENTLY HAVING DIFFICULTY BENDING HER KNEE. PATIENT IS ALSO EXPERIENCING PAIN IN THE KNEE AND CALF AREA. PATIENT IS ALSO EXPERIENCING PAIN GOING UP AND DOWN THE STAIRS AND HAS TO MANEUVER ONE STAIR AT A TIME. PATIENT STATES THAT THE PAIN IS ALSO IN HER BACK AREA AND THE SHE CAN'T BEND OVER TO PICK ANYTHING UP. PATIENT STATES THAT SHE HAS NOT BEEN TO SEE DR. (B)(6) MD BUT HAS BEEN TO SEE OTHER DOCTORS AND HAS RECEIVED CORTISONE INJECTIONS IN HER LEG FOR THE PAIN. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02395 | 2012-11-19 | 2010-09-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT IS CURRENTLY EXPERIENCING PAIN IN HER RIGHT HIP. PATIENT REPORTS THAT SHE IS CURRENTLY HAVING DIFFICULTY BENDING HER KNEE. PATIENT IS ALSO EXPERIENCING PAIN IN THE KNEE AND CALF AREA. PATIENT IS ALSO EXPERIENCING PAIN GOING UP AND DOWN THE STAIRS AND HAS TO MANEUVER ONE STAIR AT A TIME. PATIENT STATES THAT THE PAIN IS ALSO IN HER BACK AREA AND THAT SHE CAN'T BEND OVER TO PICK ANYTHING UP. PATIENT STATES THAT SHE HAS NOT BEEN TO SEE DR (B)(6), BUT HAS BEEN TO SEE OTHER DOCTORS AND HAS RECEIVED CORTISONE INJECTIONS IN HER LEG FOR THE PAIN. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01129 | 2012-11-19 | 2012-11-05 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT: "PATIENT HAS INJURIES SUSTAINED AS A RESULT OF THE IMPLANTATION OF A STRYKER REJUVENATE HIP REPLACEMENT." | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED IT WILL BE REPORTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02403 | 2012-11-19 | 2012-11-05 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED, "PATIENT HAS INJURIES SUSTAINED AS A RESULT OF THE IMPLANTATION OF A STRYKER REJUVENATE HIP REPLACEMENT." | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02399 | 2012-11-19 | 2012-04-17 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 12 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-120000S |  | IT WAS REPORTED THAT THE PATIENT STATES THAT SIX WEEKS AGO, HE COULD PICK UP ITEMS OFF THE FLOOR. NOW PATIENT STATES THAT HE CANNOT TOUCH ANKLES TO DRY IN THE SHOWER. PATIENT IS ALSO REPORTING THAT STANDING ON ONE LEG CAUSES SIGNIFICANT PAIN. PATIENT STATES THAT AT NIGHT HE HAS TO SLEEP ON HIS BACK BECAUSE HE CANNOT SLEEP ON HIS SIDE. PATIENT IS ALSO REPORTING ELEVATED LEVELS OF COBALT (8.6) AND CHROMIUM (1.4). PATIENT IS SCHEDULED FOR BILATERAL REVISION. ADDITIONAL INFORMATION RECEIVED FROM SALES REP: IT WAS REPORTED THAT HAD RT HIP REPLACED ON (B)(6) 2012. NEVER REALLY FELT GREAT AND HAS PROGRESSIVELY GOTTEN WORSE IN TERMS OF PAIN. HAS REQUIRED STEROID INJECTION AND CONTINUOUS USE OF ANT-INFLAMMATORY MEDS. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-11
9616680-2012-01141 | 2012-11-20 | 2011-11-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK 127/132 34MM V40 | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | NLS-340000B |  | IT WAS REPORTED THAT THE PT IS EXPERIENCING SORENESS IN GROIN AREA WHEN BENDING OVER. EXPERIENCES PAIN IN GROIN AREA AFTER DAILY 30 MINUTE WALKS. NOTICES SWELLING IN RIGHT LEG AND EXPERIENCING RIGHT KNEE PAIN. ADDITIONAL INFORMATION REPORTED BY THE SALES REP INDICATED THAT THERE WAS A REJUVENATE REVISION. BLOOD SERUM ION LEVEL WAS ELEVATED. PT'S COBALT LEVEL WAS 10 AND CHROMIUM LEVEL WAS LESS THAN 1. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-11
2249697-2012-02418 | 2012-11-20 | 2010-08-11 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT THE PT WAS EXPERIENCING EXTREME PAIN IN HIP AND TENDON. PT HAD MRI AND BLOOD WORK PERFORMED AND IT WAS FOUND THAT PT HAS ELEVATED BLOOD LEVELS. PT HAD A PARTIAL HIP REPLACEMENT ON (B)(6) 2012 ALONG WITH TENDON REPAIR. PT STATES THAT TENDON DAMAGE WAS DONE DUE TO ORIGINAL IMPLANT. PARTIAL IMPLANT REPLACEMENT WAS PERFORMED AT (B)(6) MEDICAL CTR. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | LAY USER/PATIENT |  | 2012-11
2249697-2012-02408 | 2012-11-20 | 2012-10-29 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED MECHANICALLY ASSISTED CREVICE CORROSION WITH ADVERSE LOCAL SOFT TISSUE REACTION AND PSEUDO TUMOR FORMATION SECONDARY TO REJUVENATE MODULAR FEMORAL COMPONENT. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-11
2249697-2012-02398 | 2012-11-20 | 2012-10-29 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT THE PT IS HURTING AND HAS HIGHER METAL ION LEVELS AND WANTS THE PROSTHESIS REMOVED AND A NEW ONE PUT IN. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-11
9616680-2012-01144 | 2012-11-20 | 2010-08-11 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK 34MM | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT THE PT WAS EXPERIENCING EXTREME PAIN IN HIP AND TENDON. PT HAD MRI AND BLOOD WORK PERFORMED AND IT WAS FOUND THAT PT HAS ELEVATED BLOOD LEVELS. PT HAD A PARTIAL HIP REPLACEMENT ON (B)(6) 2012, ALONG WITH TENDON REPAIR. PT STATES THAT TENDON DAMAGE WAS DONE DUE TO ORIGINAL IMPLANT. PARTIAL IMPLANT REPLACEMENT WAS PERFORMED AT (B)(6). | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | LAY USER/PATIENT |  | 2012-11
2249697-2012-02410 | 2012-11-20 | 2012-10-31 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT COMPONENTS WERE EXPLANTED DUE TO ADVERSE LOCAL TISSUE REACTION. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR DUE TO HOSPITAL POLICY. ADDITIONAL INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED, BUT NOT MADE AVAILABLE. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02414 | 2012-11-20 | 2010-06-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT IS EXPERIENCING PAIN IN HIS UPPER THIGH AND STATES THAT THE PAIN COMES AND GOES IN HIS RIGHT BUTT CHEEK. PT IS REPORTING THAT PAIN THIS YEAR STARTED TO BECOME NOTICEABLE. PATIENT ALSO STATES THAT LAST WEEKEND PAIN WAS UNBEARABLE. PATIENT REPORTS THAT THE PAIN IN THE THIGH AREA IS A BURNING PAIN AND ALSO FEELS LIKE THIGH DOESN¿T WANT TO WORK. PATIENT STATES THAT HE IS SCHEDULED FOR BLOOD WORK AND MRI ON SATURDAY, (B)(6) 2012. | ADD¿L INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | LAY USER/PATIENT |  | 2012-11
2249697-2012-02415 | 2012-11-20 | 2012-11-05 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT HAS INJURIES SUSTAINED AS A RESULT OF THE IMPLANTATION OF A STRYKER REJUVENATE HIP REPLACEMENT. | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01139 | 2012-11-20 | 2012-11-05 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP NECK | IMPLANT | STRYKER IRELAND, LTD. OSTEONICS (ORTHOPAEDICS DIVI | NA | UNK |  | IT WAS REPORTED THAT THE PT HAS INJURIES SUSTAINED AS A RESULT OF THE IMPLANTATION OF A STRYKER REJUVENATE HIP REPLACEMENT. | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION NO ADD¿L INFO IS AVAILABLE AT THIS TIME. IF ADD¿L INFO IS RECEIVED IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02416 | 2012-11-20 | 2011-01-18 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT IS EXPERIENCING PAIN IN HIS GROIN AND IN HIS HIP AREA. PATIENT IS REPORTING THAT THE PAIN FEELS LIKE A PULL OR TEAR IN THE BELLY BUTTON AREA. (B)(6) IS ALSO STATING THAT IT FEELS LIKE A STRAIN IN THE SAME GENERAL AREA WHEN WALKING. PT ALSO STATES THAT FOR THE LAST FEW DAYS HE HAS BEEN EXPERIENCING BLEEDING WHEN USING THE LAVATORY. PT STATES THAT HE HAS NOT HAD ANY TESTING YET IN REGARDS TO HIS HIP. | ADD¿L INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01140 | 2012-11-20 | 2012-01-18 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER IRELAND, LTD. OSTEONICS (ORTHOPAEDICS DIVI | NA | UNK |  | IT WAS REPORTED THAT THE PT IS EXPERIENCING PAIN IN HIS GROIN AND IN HIS HIP AREA. PATIENT IS REPORTING THAT THE PAIN FEELS LIKE A PULL OR TEAR IN THE BELLY BUTTON AREA. PAT IS ALSO STATING THAT IT FEELS LIKE A STRAIN IN THE SAME GENERAL AREA WHEN WALKING. PT ALSO STATES THAT FOR THE LAST FEW DAYS HE HAS BEEN EXPERIENCING BLEEDING WHEN USING THE LAVATORY. PT STATES THAT HE HAS NOT HAD ANY TESTING YET IN REGARDS TO HIS HIP. | ADD¿L INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02400 | 2012-11-20 | 2012-10-29 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT DOCTOR REVISED PATIENT DUE TO ALTR. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER DUE TO HOSPITAL POLICY. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED, BUT NOT MADE AVAILABLE. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02407 | 2012-11-20 | 2012-10-29 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT THE PATIENT COMPLAINED OF PAIN SO THE SURGEON REVISED LEFT HIP. | AN EVALUATION OF THE REPORTED DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS RETAINED BY THE HOSPITAL AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION INCLUDING X-RAYS AND MEDICAL RECORDS HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02420 | 2012-11-20 | 2012-11-06 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT DOCTOR (B)(6) REVISED THE PTS RIGHT HIP DUE TO ADVERSE LOCAL TISSUE REACTION ALTR. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02430 | 2012-11-20 | 2012-11-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP SYSTEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT IS ALLEGED THROUGH THE FILING OF A LAWSUIT, THAT AFTER THE PT'S RIGHT TOTAL HIP REPLACEMENT, THE PT BEGAN EXPERIENCING SEVERE HIP PAIN. NUMEROUS EFFORTS TO DIAGNOSE THE PAIN WERE MADE, WHICH LED TO A REVISION SURGERY AT (B)(6). | THE INFO IN THIS REPORT WAS PROVIDED BY STRYKER ORTHOPAEDICS' LEGAL AFFAIRS DEPARTMENT. NO ADDITIONAL INFO IS AVAILABLE AT THIS TIME. THE DEVICE IS NOT AVAILABLE DUE TO THE ONGOING LITIGATION. SHOULD DEVICE OR ADDITIONAL INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02438 | 2012-11-21 | 2012-11-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | REJUVENATE STEM EXTRACTOR | INSTRUMENT | STRYKER ORTHOPAEDICS MAHWAH | NA | 1601-1650 |  | IT WAS REPORTED THAT, SURGEON WAS USING REJUVENATE STEM EXTRACTOR TOOL WHILE DOING A REVISION OF PATIENT'S LEFT HIP ON (B)(6) 2012 AND THE TURNING KNOB CAME OFF WHILE IN USE. THE SURGEON USED A BACK UP REJUVENATE STEM EXTRACTOR TO COMPLETE THE CASE WITHOUT ANY DELAY. | METHOD - REVIEW OF DEVICE HISTORY, COMPLAINT HISTORY, SURGICAL PROTOCOL & IFU. VISUAL INSPECTION WA PERFORMED ALONGSIDE A DISCUSSION WITH MATERIAL ANALYSIS. THE RESULTS INDICATED THAT THE THUMB KNOB BROKE IN AN OVERLOAD CONDITION AND THA NO MANUFACTURING OR MATERIAL DEFECTS WERE OBSERVED. REVIEW OF THE DEVICE HISTORY RECORDS INDICATES THAT ALL DEVICES WERE MANUFACTURED AND ACCEPTED INTO FINAL STOCK WITH NO REPORTED DISCREPANCIES. THE PRODUCT EXPERIENCE REPORTING DATABASE SHOWS THERE HAVE BEEN NO OTHER EVENTS FOR THE REPORTED LOT ID. THE ROOT CAUSE WAS DETERMINED TO BE USER MISUSE AS THUMB KNOB BROKE OFF IN AN OVERLOAD CONDITION DUE TO EXCESSIVE TIGHTENING. | N | Y | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02440 | 2012-11-21 | 2012-10-31 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT THE PT HAD ELEVATED COBALT LEVEL OF 18 AND WANTED THE REVISION. THEREFORE, THE SURGEON REVISED LEFT HIP. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-11
MW5027814 | 2012-11-26 | 2009-03-05 | I, n, j, u, r, y | Voluntary report | STRYKER REJUVENATE SPT | TOTAL HIP REPLACEMENT | STRYKER | 690-00-22E | 690-22-22E |  | I HAD A TOTAL HIP REPLACEMENT ON (B)(6) 2009. I RECEIVED A STRYKER REJUVENATE SPT. PRIOR TO THIS SURGERY, I HAD PAIN AND CATCHING/CLICKING IN MY RIGHT HIP, IT WAS VERY PAINFUL TO RIDE IN ANY VEHICLE. THE PAIN WAS EFFECTING MY KNEES AS I TRIED TO FAVOR MY HIP. ALL THIS INCREASED SINCE THE SURGERY, PLUS THE WEAKNESS IN MY HIP. I HAD TO LEARN TO SLEEP ON MY BACK AS TURNING WAS A BIG PROBLEM AND SLEEPING ON MY SIDE IMPOSSIBLE. I HAVE HAD MANY ROUND OF PT OVER THE LAST 3+ YEARS AND MANY BOUTS OF BURSITIS, AT LEAST 4 TIMES SWELLING TO THE SIZE OF A TENNIS BALL. NONE OF THE PT OR THE EXERCISES SEEMED TO STRENGTHEN THE MUSCLES IN MY RIGHT HIP. MY RIGHT HIP DISLOCATED ON (B)(6) 2011, AS I WAS SITTING IN A CHAIR PUTTING ON A PAIR OF SHOES. I WAS TAKEN MY AMBULANCE TO THE HOSPITAL AND HAD TO STAY OVERNIGHT. IN (B)(6) 2012, I GOT THE LETTER OF RECALL AND WENT IN FOR MRI AND BLOOD TESTS, IT WAS CONFIRMED I HAD A VERY HIGH LEVEL OF HEAVY METAL POISONING. PT FROM (B)(6) 2009 THROUGH (B)(6) 2012. REASON FOR USE: ADVANCED RIGHT HIP ARTHRITIS. |  | Y | Y | LAY USER/PATIENT |  | 2012-11
MW5027816 | 2012-11-26 | 2009-10-01 | I, n, j, u, r, y | Voluntary report | STRYKER REJUVENATE/ABG II | MODULAR STEM | STRYKER |  |  |  | IN (B)(6) 2009, I UNDERWENT A LEFT TOTAL HIP ARTHROPLASTY. PRIOR TO THE SURGERY, I HAD SEVERE LEFT HIP PAIN THAT INTERFERED SIGNIFICANTLY WITH MY ABILITY TO WALK. FOLLOWING RECOVERY FROM THE SURGERY, MY LEFT HIP HAS BEEN PAIN FREE. DURING (B)(6) 2012, I WAS NOTIFIED BY THE ORTHOPEDIC SURGEON THAT THE HIP THAT HAD BEEN IMPLANTED IN MY BODY WAS BEING RECALLED DUE TO CONCERNS WITH TOXICITY IN MANY HIP REPLACEMENT PTS. UP TO THIS POINT, I CONTINUE TO BE PAIN FREE. HOWEVER, RECENT EVALUATIONS ORDERED BY THE ORTHOPEDIC SURGEON SHOWED SOME SIGNS OF LEFT HIP DETERIORATIONS. IN 6 MONTHS, MY HIP STATUS WILL BE RE-EVALUATED. ACCORDING TO MY PHYSICIAN, THERE IS A STRONG POSSIBILITY I WILL, NEED MY ARTIFICIAL HIP REPLACED UTILIZING A MORE INVASIVE PROCEDURE. THIS WILL RESULT IN A LONGER RECOVERY WITH RESULTANT WAGE LOSS! IT WILL REQUIRE HOSPITALIZATION AND INTERVENTIONS TO PREVENT DAMAGE. REASON FOR USE: LEFT HIP ARTHRITIS. |  | N | Y | 0 |  | 2012-11
9616680-2012-01176 | 2012-11-26 | 2012-04-02 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK 38MM V40 | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | NLS-380000B |  | IT WAS REPORTED THAT: PATIENT COMPLAINS OF PAIN AND SWELLING AT JOINT SITE. HE HAS BEEN HAVING HEADACHES, GI UPSET, LOSS OF BALANCE AND INTERMITTENT INABILITY TO BEAR WEIGHT ON HIS LEFT HIP FOR 2 - 3 MONTHS. HE HAS EXTREME FATIGUE AND HIS SURGEON HAS SCHEDULED HIM FOR AN UPCOMING REVISION SURGERY. MRI AND BLOOD TESTS HAVE SHOWN INCREASED COBALT LEVELS AND PATIENT HAS BEEN TOLD THAT HE HAS COBALT TOXICITY. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02462 | 2012-11-27 | 2012-11-05 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE RIGHT HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | PATIENT IN PAIN. SURGEON WILL BE REVISING THE PATIENT'S RIGHT HIP IN A FEW WEEKS AS THIS PATIENT HAD A BILATERAL REJUVENATE HIP DONE 2011. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02456 | 2012-11-27 | 2012-11-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | REJUVENATE STEM EXTRACTOR | INSTRUMENT | STRYKER ORTHOPAEDICS MAHWAH | NA | 1601-1650 |  | UPON RECEIPT OF DEVICES RELATED TO (B)(4), IT WAS NOTICED THAT THE REJUVENATE STEM EXTRACTOR WAS RETURNED WITH THE REJUVENATE STEM STILL ATTACHED. IT WAS ALSO NOTED THAT THE KNOB WAS NOT ATTACHED TO THE THREADED SHAFT AND NOT RETURNED WITH THE DEVICE. SPOKE WITH SALES REP ON (B)(4) 2012 AND SALES REP STATED THAT THE KNOB BROKE OFF OF THE EXTRACTOR AND THEREFORE THEY COULD NOT RELEASE EXPLANT FROM THE INSTRUMENT AND RETURNED THE PCS TOGETHER. | WHEN COMPLETED, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | N | Y | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01198 | 2012-11-27 | 2012-11-05 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | PATIENT IN PAIN. SURGEON WILL BE REVISING THE PATIENT'S RIGHT HIP IN A FEW WEEKS AS THIS PATIENT HAD A BILATERAL REJUVENATE HIP DONE 2011. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01171 | 2012-11-27 | 2011-12-12 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK 34MM V40 | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | NLS-340000B |  | IT WAS REPORTED THAT PATIENT HAS PAIN. PATIENT HAS FLUID IN JOINT THAT KEEPS COMING BACK EVERY TIME IT IS REMOVED. PATIENT WILL PROBABLY BE RECEIVING REVISION SURGERY. SUBSEQUENT INFORMATION REPORTED BY THE SALES REP INDICATED THAT THE PATIENT HAS INCREASING PAIN OVER PAST YEAR AND HAS HAD AN INFUSION AND WAS ASPIRATED AND WAS EXTENDED FOR SCREENING. SURGEON THINKS PATIENT MAY BE HAVING A METAL REACTION AND SHOULD BE REVISED. ADDITIONAL INFORMATION INDICATED THAT THE PATIENT HAD REVISION SURGERY ON (B)(6) 2012. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-11
2249697-2012-02459 | 2012-11-27 | 2012-11-05 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT THIS WAS A FAILED REJUVENATE ALTR. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY REPORT WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02461 | 2012-11-27 | 2012-05-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 10 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-100000S |  | IT IS REPORTED THAT: THE PATIENT HAS BILATERAL REJUVENATE HIP IMPLANTS. THE PATIENT BEGAN EXPERIENCING PAIN IN HIS LEFT HIP IN (B)(6) 2012. THE PATIENT SAW HIS SURGEON ON (B)(6) 2012 BECAUSE OF PAIN AND A SEVERE LACK OF MOBILITY IN THE LEFT HIP. THE PAIN IS VERY INTENSE AND AT TIMES THE HIP LOCKS UP. WHEN THIS OCCURS, THE PATIENT HAS TO WALK BACKWARDS OR LATERALLY TO REMAIN MOBILE. THE PATIENT STATES THAT HE CANNOT LIFT HIS LEG FORWARD MORE THAN SIX INCHES, PUT ON HIS SOCKS OR SLEEP LYING ON HIS LEFT SIDE. THE PATIENT DESCRIBES THE PAIN AS AN INTENSE TOOTHACHE THAT DOESN'T GO AWAY. THE SURGEON SENT HIM FOR BLOOD TEST, AN MRI AND AN ASPIRATION. THE BLOOD TEST INDICATED ELEVATED METAL LEVELS AND THE MRI SHOWED A PSEUDO TUMOR. THE PATIENT WAS REFERRED BY HIS SURGEON TO DR (B)(6). FOR A SECOND OPINION. THE PATIENT REPEATED THE BLOOD TESTS AND MRI. A REVISION SURGERY IS SCHEDULED TENTATIVELY FOR (B)(6) 2013. ADDITIONAL INFORMATION RECEIVED FROM SALES REP ON (B)(6) 2012: IT WAS REPORTED THAT, THE PATIENT HAD A LT THA ON (B)(6) 2011 AND HAS NEVER DONE WELL, IS EXPERIENCING CHRONIC PAIN. THE PATIENT HAS BILATERAL REJUVENATE HIPS. THE RIGHT IS DOING WELL. HE CURRENTLY WANTS HIS LEFT HIP REVISED. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01170 | 2012-11-27 | 2010-02-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | NLS-300000B |  | IT WAS REPORTED THAT THE PATIENT IS COMPLAINING THAT HE CAN'T SLEEP ON IT, PAIN AND CAN'T STAND OR SIT ON IT. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-11
9616680-2012-01195 | 2012-11-27 | 2009-10-22 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT: PATIENT IS EXPERIENCING PAIN IN HIS RIGHT HIP. PATIENT IS REPORTING THAT THE LONGER HE WALKS THE MORE PAIN HE IS IN. THE PAIN HAS INCREASED IN THE LAST TWO OR THREE MONTHS. THERE IS NO SWELLING IN THE HIP AREA. HE HAS HAD BLOOD WORK AND AN MRI DONE. PATIENT IS REPORTING THAT HE IS SCHEDULED FOR REVISION SURGERY ON (B)(6), 2012. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02457 | 2012-11-27 | 2009-10-22 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: PATIENT IS EXPERIENCING PAIN IN HIS RIGHT HIP. PATIENT IS REPORTING THAT THE LONGER HE WALKS THE MORE PAIN HE IS IN. THE PAIN HAS INCREASED IN THE LAST TWO OR THREE MONTHS. THERE IS NO SWELLING IN THE HIP AREA. HE HAS HAD BLOOD WORK AND AN MRI DONE. PATIENT IS REPORTING THAT HE IS SCHEDULED FOR REVISION SURGERY ON (B)(6) 2012. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED FOR THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01203 | 2012-11-28 | 2012-04-19 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT: PATIENT IS EXPERIENCING PAIN BUT IT IS MANAGEABLE. PATIENT ALSO STATE THAT SHE HAS BLOOD WORK SCHEDULED. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY REPORT WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01200 | 2012-11-28 | 2009-12-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT: PATIENT STARTED EXPERIENCING PAIN IN HER HIP AREA AT THE END OF 2010. PATIENT REPORTED THAT SHE HAD A CORTISONE INJECTION FOR THE FIRST TIME BECAUSE OF THE PAIN IN (B)(6) 2010. SHE RECEIVED HER SECOND CORTISONE INJECTION IN (B)(6) 2012. HER PAIN IS A DEEP, DULL ACHY PAIN IN HER HIP TO THE BACK OF HER BUTTOCK AREA. SHE IS STILL USING A CANE IF SHE IS DOING PROLONGED WALKING. PATIENT STATES THAT SHE HAS HAD X-RAYS, BLOOD WORK AND MRI. SHE IS CURRENTLY AWAITING RESULTS FROM THE MRI FROM THE SURGEON. DOCTOR IS CURRENTLY MONITORING HER. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED, AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY REPORT WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02465 | 2012-11-28 | 2012-11-06 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT PATIENT PRESENTED WITH HIP PAIN. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY REPORT WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02466 | 2012-11-28 | 2012-04-19 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: PATIENT IS EXPERIENCING PAIN BUT IT IS MANAGEABLE. PATIENT ALSO STATES THAT SHE HAS BLOOD WORK SCHEDULED. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY REPORT WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02468 | 2012-11-28 | 2010-12-21 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT PAINFUL HIP AFTER REPLACEMENT WITH REJUVENATE HIP STEM IMPLANT. EXPECTED DATE OF REVISION: (B)(6) 2012. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY REPORT WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02470 | 2012-11-28 | 2012-11-06 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT PATIENT'S RIGHT HIP WAS REVISED DIE TO PAIN AND A PSEUDO TUMOR. SURGEON REPLACED WITH A RESTORATION MODULAR HIP. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY REPORT WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01206 | 2012-11-28 | 2010-12-21 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE 34MM MODULAR NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT PAINFUL HIP AFTER REPLACEMENT WITH REJUVENATE HIP STEM IMPLANT. EXPECTED DATE OF REVISION: (B)(6) 2012. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY REPORT WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02474 | 2012-11-28 | 2012-11-06 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT SURGEON REVISED PT'S LEFT HIP BECAUSE OF PAIN AND A PSEUDO TUMOR. SURGEON REPLACED WITH A RESTORATION MODULAR SYSTEM. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR DUE TO HOSPITAL POLICY. ADDITIONAL INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02463 | 2012-11-28 | 2009-12-28 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: PATIENT STARTED EXPERIENCING PAIN IN HER HIP AREA AT THE END OF 2010. PATIENT REPORTED THAT SHE HAD A CORTISONE INJECTION FOR THE FIRST TIME BECAUSE OF THE PAIN IN (B)(6) 2011. SHE RECEIVED HER SECOND CORTISONE INJECTION IN (B)(6) 2012. HER PAIN IS A DEEP, DULL, ACHY PAIN IN HER HIP TO THE BACK OF HER BUTTOCK AREA. SHE IS STILL USING A CANE IF SHE IS DOING PROLONGED WALKING. PATIENT STATES THAT SHE HAS HAD X-RAYS, BLOOD WORK AND MRI. SHE IS CURRENTLY AWAITING RESULTS FROM THE MRI FROM HER SURGEON. DOCTOR IS CURRENTLY MONITORING HER. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02493 | 2012-11-29 | 2012-11-07 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT COMPONENTS WERE EXPLANTED DUE TO ADVERSE LOCAL TISSUE REACTION. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02496 | 2012-11-29 | 2012-11-07 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT, DR (B)(6) DID THE FIRST STEP OF A TWO STEP REVISION ON THIS PT. HE PUT IN A #5 127 DEGREE HFX STEM, 36 HEAD, AND 36 POLY WITH A CEMENT SPACER. | AN EVALUATION OF THE REPORTED DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS RETAINED BY THE HOSPITAL AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION INCLUDING X-RAYS AND MEDICAL RECORDS WAS NOT MADE AVAILABLE EITHER. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02500 | 2012-11-29 | 2012-11-07 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT THE PATIENT HAD PAIN, AN ELEVATED COBALT LEVEL AND PSEUDO TUMOR SO THE SURGEON REVISED THE LEFT HIP. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS RETAINED BY THE HOSPITAL AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02477 | 2012-11-29 | 2012-11-02 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT THE PATIENT EXPERIENCED ADVERSE LOCAL TISSUE REACTION. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-11
9616680-2012-01207 | 2012-11-29 | 2012-02-02 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK 38MM | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | NLS-380000B |  | IT IS REPORTED THAT THE PATIENT REPORTS PAIN AT THE HIP JOINT. PATIENT RECEIVES STEROID SHOTS AT THE JOINT TO RELIEVE PAIN. SURGEON TO SEE HER IN 4 WEEKS. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-11
2249697-2012-02501 | 2012-11-29 | 2012-11-07 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | REJUVENATE STEM EXTRACTOR | INSTRUMENT | STRYKER ORTHOPAEDICS MAHWAH | NA | 1601-1650 |  | IT WAS REPORTED THAT, THE WING NUT SNAPPED OFF DURING REMOVAL. A REPLACEMENT WAS ON HAND SO THERE WERE NO ADVERSE CONSEQUENCES TO THE PT. | WHEN COMPLETED, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | N | Y | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02469 | 2012-11-29 | 2012-02-02 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT IS REPORTED THAT THE PATIENT REPORTS PAIN AT THE HIP JOINT. PATIENT RECEIVES STEROID SHOTS AT THE JOINT TO RELIEVE PAIN. SURGEON TO SEE HER IN 4 WEEKS. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-11
2249697-2012-02494 | 2012-11-29 | 2012-11-07 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT PT PRESENTED WITH HIP PAIN. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER DUE TO HOSPITAL POLICY. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01258 | 2012-11-30 | 2010-01-11 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT: PT IS EXPERIENCING PAIN BUT IT IS NOT UNBEARABLE. PT IS REPORTING THAT THE PAIN IS MORE OF AN ACHINESS. PT STATES THAT HE HAS HAD BLOOD WORK AND EVERYTHING LOOKS OK. PT IS SUPPOSED TO HAVE AN MRI BUT HAS NOT SCHEDULED YET. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. DEVICE INFO HAS NOT BEEN PROVIDED AT THIS TIME. ADDITIONAL INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02508 | 2012-11-30 | 2011-01-25 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT THE PT IS EXPERIENCING ABRASIONS ON HIS ANKLE. PT ALSO STATES THAT PT IS EXPERIENCING SOME PAIN. PT IS REPORTING THAT HE CANNOT SLEEP IN HIS BED AND THAT CURRENTLY HE SLEEPS IN A (B)(6) RECLINER. PT IS ALSO REPORT THAT HE WALKS WITH A CANE. PT IS STATING THAT HIS WHOLE LEG IS DISCOLORED. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | OTHER |  | 2012-11
9616680-2012-01257 | 2012-11-30 | 2012-05-22 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT: PT HAS BEEN EXPERIENCING PAIN. PT ALSO STATES THAT SHE HAS NERVE DAMAGE. PT DESCRIBES THE PAIN AS A DULL ACHE IN THE HIP AREA. PT ALSO STATES THAT SHE HAS PAIN BELOW HER KNEE DUE TO THE NERVE DAMAGE. PT HAS BEEN TO SEE THE NEUROLOGIST ABOUT THE NERVE DAMAGE BUT HAS NOT HAD ANY TESTING DONE YET. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADDITIONAL INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WILL BE REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
9616680-2012-01268 | 2012-11-30 | 2011-05-09 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT: PATIENT STARTED EXPERIENCING PAIN ON (B)(6) 2012. PATIENT STATES THAT SHE CANNOT PUT ANY WEIGHT ON THAT LEG. PATIENT IS REPORTING THAT THE PAIN IS IN THE MIDDLE OF THE THIGH (THREE INCHES ABOVE THE KNEECAP). PATIENT ALSO IS STATING THAT IT FEELS LIKE DEEP TISSUE PAIN. PATIENT SAYS SHE CANNOT PUT ON A SOCK AND ALSO IS NOT AS FLEXIBLE. PATIENT HAS NOT BEEN TO SEE DOCTOR YET AS PAIN IS RECENT. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. DEVICE INFORMATION HAS NOT BEEN PROVIDED AT THIS TIME. INFORMATION RECEIVED FROM THE PATIENT INDICATED THAT REPORTED DEVICE MAY BE EITHER REJUVENATE OR ABGII. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WILL BE REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02512 | 2012-11-30 | 2010-01-11 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: PT IS EXPERIENCING PAIN BUT IT IS NOT UNBEARABLE. PT IS REPORTING THAT THE PAIN IS MORE OF AN ACHINESS. PT STATES THAT HE HAS HAD BLOOD WORK AND EVERYTHING LOOKS OK. PT IS SUPPOSED TO HAVE AN MRI BUT HAS NOT SCHEDULED YET. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. DEVICE INFO HAS NOT BEEN PROVIDED AT THIS TIME. ADDITIONAL INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02511 | 2012-11-30 | 2012-05-22 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: PT HAS BEEN EXPERIENCING PAIN. PT ALSO STATES THAT SHE HAS NERVE DAMAGE. PT DESCRIBES THE PAIN AS A DULL ACHE IN THE HIP AREA. PT ALSO STATES THAT SHE HAS PAIN BELOW HER KNEE DUE TO THE NERVE DAMAGE. PT HAS BEEN TO SEE THE NEUROLOGIST ABOUT THE NERVE DAMAGE BUT HAS NOT HAD ANY TESTING DONE YET. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADDITIONAL INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WILL BE REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-11
2249697-2012-02515 | 2012-12-04 | 2012-11-12 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDIC MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT MECHANICALLY ASSISTED CREVICE CORROSION AND SECONDARY ADVERSE LOCAL SOFT TISSUE REACTION AND PSEUDO TUMOR FORMATION AFTER REJUVENATE FEMORAL COMPONENT DUE TO MODULAR NECK-STEM MECHANISM OF FAILURE. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02520 | 2012-12-04 | 2011-05-09 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: PATIENT STARTED EXPERIENCING PAIN ON (B)(6), 2012. PATIENT STATES THAT SHE CANNOT PUT ANY WEIGHT ON THAT LEG. PATIENT IS REPORTING THAT THE PAIN IS IN THE MIDDLE OF THE THIGH (THREE INCHES ABOVE THE KNEECAP). PATIENT ALSO IS STATING THAT IT FEELS LIKE DEEP TISSUE PAIN. PATIENT SAYS SHE CANNOT PUT ON A SOCK AND ALSO IS NOT AS FLEXIBLE. PATIENT HAS NOT BEEN TO SE DOCTOR YET AS PAIN IS RECENT. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. DEVICE INFORMATION HAS NOT BEEN PROVIDED AT THIS TIME. INFORMATION RECEIVED FROM THE PATIENT INDICATOR THAT REPORTED DEVICE MAY BE EITHER REJUVENATE OR ABGII. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WILL BE REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02519 | 2012-12-04 | 2012-11-13 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | UNKNOWN | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT COMPONENTS WERE EXPLANTED DUE TO ADVERSE LOCAL TISSUE REACTION. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01278 | 2012-12-04 | 2011-06-07 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT: PT IS EXPERIENCING PAIN IN HIS HIP AREA. PT STATES THAT WHEN WALKING HE STARTS TO HAVE PAIN AND CANNOT WALK LONG DISTANCES. PT IS SCHEDULED TO HAVE DOCTOR'S APPOINTMENT WEEK OF (B)(6) 2012. PT IS SUPPOSE TO HAVE BLOOD WORK PRIOR TO APPOINTMENT WITH DOCTOR. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADDITIONAL INFO PERTAINING TO THE DEVICE REFERENCE IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01273 | 2012-12-04 | 2011-10-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT HIP REJUVENATE/ABGII NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT: PATIENT EXPERIENCED PAIN IN THE HIP AREA WHEN LIFTING ITEMS. PATIENTS LEG SWELLS FROM TIME TO TIME. PATIENT IS GOING TO DOCTOR SOON TO HAVE BLOOD WORK DONE TO HAVE COBALT LEVELS CHECKED. THE PATIENT HAS MOVED FROM (B)(6) (WHERE THE SURGERY WAS PERFORMED) TO (B)(6). THE PATIENT WILL CONTINUE TO KEEP IN CONTACT WITH THE SURGEON. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. DEVICE INFORMATION HAS NOT BEEN PROVIDED AT THIS TIME. INFORMATION RECEIVED FROM THE PATIENT INDICATED THAT REPORTED DEVICE MAY BE EITHER REJUVENATE OR ABGII. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENT REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02534 | 2012-12-04 | 2011-06-07 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: PT IS EXPERIENCING PAIN IN HIS HIP AREA. PT STATES THAT WHEN WALKING HE STARTS TO HAVE PAIN AND CANNOT WALK LONG DISTANCES. PT IS SCHEDULED TO HAVE DOCTOR'S APPOINTMENT WEEK OF (B)(6) 2012. PT IS SUPPOSED TO HAVE BLOOD WORK PRIOR TO APPOINTMENT WITH DOCTOR. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02535 | 2012-12-04 | 2011-12-12 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: PT IS EXPERIENCING PAIN ON THE SIDE OF HER RIGHT HIP. PT IS ALSO EXPERIENCING PAIN IN HER GROIN AREA AND ON HER THIGH AREA. PT IS REPORTING THAT SHE DOES NOT HAVE ANY SWELLING ON HER HIP THAT IS NOTICEABLE. PT ALSO STATES THAT THE PAIN HAS BEEN PRESENT SINCE SURGERY. PT STATES THAT SHE HAS A PRESCRIPTION FOR THE BLOOD WORK BUT HAS NOT MADE AN APPOINTMENT YET. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADDITIONAL INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | Y | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01275 | 2012-12-04 | 2010-09-03 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT: PATIENT STARTED HAVING PAIN TWO WEEKS AGO IN HER LEG. PATIENT STATES THAT THE PAIN IS IN HER GROIN AREA. ALSO THAT THE PAIN IS IN HER HIP AREA DOWN TO THE MIDDLE OF HER THIGH. PATIENT IS ALSO REPORTING THAT SHE CAN'T STAND UP STRAIGHT AND THAT THE PAIN IS RADIATING TO HER BACK. PATIENT STATES THAT SHE IS SCHEDULED FOR A CHECK UP ON (B)(6) 2012. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENT REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01279 | 2012-12-04 | 2011-12-12 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT: IS EXPERIENCING PAIN ON THE SIDE OF HER RIGHT HIP. PT IS ALSO EXPERIENCING PAIN IN HER GROIN AREA AND ON HER THIGH AREA. PT IS REPORTING THAT SHE DOES NOT SEE ANY SWELLING ON HER HIP THAT IS NOTICEABLE. PT ALSO STATES THAT THE PAIN HAS BEEN PRESENT SINCE SURGERY. PT STATES THAT SHE HAS A PRESCRIPTION FOR THE BLOOD WORK BUT HAS NOT MADE AN APPOINTMENT YET. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADDITIONAL INFO PERTAINING TO THE DEVICE REFERENCE IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02561 | 2012-12-04 | 2012-11-12 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT, PT COMPLAINED OF PAIN, HAS PSEUDO TUMOR AND ELEVATED COBALT SO THE SURGEON REVISED. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR DUE TO HOSPITAL POLICY. ADDITIONAL INFO (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02570 | 2012-12-04 | 2011-06-13 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: PT IS EXPERIENCING SORENESS AND PAIN IN THE HIP AREA. PT ALSO IS REPORTING THAT HE IS HAVING PAIN IN THE BACK OF HIS LEG. PT IS STATING THAT THE PAIN FIRST STARTED IN (B)(6) 2012. PT STATES THAT HE HAD AN APPOINTMENT ON (B)(6) 2012, IN WHICH HE HAD BLOOD WORK DONE AND HE IS SUPPOSED TO SCHEDULE AN MRI. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADDITIONAL INFO PERTAINING TO THE DEVICE REFERENCE IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL REPORTS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02548 | 2012-12-04 | 2012-11-13 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT THE PATIENT COMPLAINED OF PAIN WORSENING OVER TIME, MRI SHOWED LARGE FLUID COLLECTION AROUND HIP. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01300 | 2012-12-04 | 2011-09-22 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER IRELAND, LTD. OSTEONICS (ORTHOPAEDICS DIVI | NA | UNK |  | IT WAS REPORTED THAT: PT IS EXPERIENCING PAIN AND STATES THAT SHE HAS SWELLING IN HER HIP AREA. PATIENT IS REPORTING THAT ANY TYPE OF ACTIVITY IS PAINFUL. SHE IS UNABLE TO PUT ON SOCKS AND SHOES. SHE ALWAYS FEELING TIRED. SHE STARTED HAVING PAIN WITHIN THREE MONTHS AFTER HAVING SURGERY. PATIENT SAW HER DOCTOR IN (B)(6) 2012 AND HAD BLOOD WORK AND FOUND OUT THAT SHE HAS A LARGE CYST IN THE FRONT OF HER HIP. PATIENT IS SCHEDULED TO SEE DOCTOR ON (B)(6) 2012. SHE WILL GET AN IDEA ABOUT WHEN SHE IS TO HAVE REVISION SURGERY. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD¿L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD¿L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01301 | 2012-12-04 | 2012-11-13 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER IRELAND, LTD. OSTEONICS (ORTHOPAEDICS DIVI | NA | UNK |  | IT WAS REPORTED THAT: ¿PATIENT HAS INJURIES SUSTAINED AS A RESULT OF THE IMPLANTATION OF A STRYKER REJUVENATE HIP REPLACEMENT¿. | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION NO ADD¿L INFO IS AVAILABLE AT THIS TIME. IF ADD¿L INFO IS RECEIVED IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02569 | 2012-12-04 | 2011-12-14 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT: PT COMPLAINS OF PAIN THAT BEGAN A FEW WEEKS AGO. MRI AND BLOOD WORK RESULTS PENDING. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFO INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01299 | 2012-12-04 | 2011-06-13 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT: PT IS EXPERIENCING SORENESS AND PAIN IN THE HIP AREA. PT ALSO IS REPORTING THAT HE IS HAVING PAIN IN THE BACK OF HIS LEG. PT IS STATING THAT THE PAIN FIRST STARTED IN LATE (B)(6) 2012. PT STATES THAT HE HAD AN APPOINTMENT ON (B)(6) 2012 IN WHICH HE HAD BLOOD WORK DONE AND HE IS SUPPOSED TO SCHEDULE AN MRI. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADDITIONAL INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02571 | 2012-12-04 | 2011-09-22 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: PT IS EXPERIENCING PAIN AND STATES THAT SHE HAS SWELLING IN HER HIP AREA. PATIENT IS REPORTING THAT ANY TYPE OF ACTIVITY IS PAINFUL. SHE IS UNABLE TO PUT ON SOCKS AND SHOES. SHE ALWAYS FEELING TIRED. SHE STARTED HAVING PAIN WITHIN THREE MONTHS AFTER HAVING SURGERY. PATIENT SAW HER DOCTOR IN (B)(6) 2012 AND HAD BLOOD WORK AND FOUND OUT THAT SHE HAS A LARGE CYST IN THE FRONT OF HER HIP. PATIENT IS SCHEDULED TO SEE DOCTOR ON (B)(6) 2012. SHE WILL GET AN IDEA ABOUT WHEN SHE IS TO HAVE REVISION SURGERY. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MFR. ADD¿L INFO PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADD¿L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02574 | 2012-12-04 | 2012-11-13 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: ¿PATIENT HAS INJURIES SUSTAINED AS A RESULT OF THE IMPLANTATION OF A STRYKER REJUVENATE HIP REPLACEMENT¿. | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION NO ADD¿L INFO IS AVAILABLE AT THIS TIME. IF ADD¿L INFO IS RECEIVED IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01302 | 2012-12-04 | 2012-11-13 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER IRELAND, LTD. OSTEONICS (ORTHOPAEDICS DIVI | NA | UNK |  | IT WAS REPORTED THAT: ¿PATIENT HAS INJURIES SUSTAINED AS A RESULT OF THE IMPLANTATION OF A STRYKER REJUVENATE HIP REPLACEMENT¿. | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION NO ADD¿L INFO IS AVAILABLE AT THIS TIME. IF ADD¿L INFO IS RECEIVED IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02573 | 2012-12-04 | 2012-11-13 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: ¿PATIENT HAS INJURIES SUSTAINED AS A RESULT OF THE IMPLANTATION OF A STRYKER REJUVENATE HIP REPLACEMENT¿. | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION, NO ADD¿L INFO IS AVAILABLE AT THIS TIME. IF ADD¿L INFO IS RECEIVED, IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02527 | 2012-12-04 | 2010-09-03 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: PATIENT STARTED HAVING PAIN TWO WEEKS AGO IN HER LEG. PATIENT STATES THAT THE PAIN IS IN HER GROIN AREA. ALSO THAT THE PAIN IS IN HER HIP AREA DOWN TO THE MIDDLE OF HER THIGH. PATIENT IS ALSO REPORTING THAT SHE CAN'T STAND UP STRAIGHT AND THAT THE PAIN IS RADIATING TO HER BACK. PATIENT STATES THAT SHE IS SCHEDULED FOR A CHECK UP ON (B)(6) 2012. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENT REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02549 | 2012-12-06 | 2011-11-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT IS REPORTED THAT THE PATIENT WAS INFORMED OF THE RECALL DURING AN ANNUAL CHECK UP ON (B)(6) 2012. X-RAYS AND BLOOD WORK WERE PERFORMED. THE PATIENT'S COBALT LEVELS WERE ELEVATED. THE PATIENT HAS BEEN EXPERIENCING DISCOMFORT IN THE RIGHT BUTTOCKS FOR APPROXIMATELY TWO WEEKS. THE DISCOMFORT IS NOT CONSTANT OR DEBILITATING. THE PATIENT IS SCHEDULING A FOLLOW-UP APPOINTMENT WITH THE SURGEON FOR ADDITIONAL TESTING IN THE NEAR FUTURE. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02596 | 2012-12-06 | 2012-11-15 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE SURGEON REMOVED A REJUVENATE STEM FROM PT'S LEFT SIDE DUE TO EXTREME PAIN. SURGEON REPLACED WITH A ZIMMER STEM. | IT WAS NOTED THAT THE DEVICES ARE NOT AVAILABLE FOR EVAL. ADD'L IN HAS BEEN REQUESTED AND IF REC'D, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-12
9616680-2012-01335 | 2012-12-06 | 2010-10-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT IS REPORTED, THE PT LEARNED OF THE RECALL THROUGH A TELEVISION ADVERTISEMENT AND CONFIRMED THE RECALL STATUS OF HIS IMPLANT THROUGH HIS SURGEON. THE PT REPORTS HAVING PAIN THAT EXTENDS FROM THE TOP OF HIS HIP DOWN TO HIS KNEE, SWELLING IN THE HIP, AND NUMBNESS ON THE BOTTOM OF HIS RIGHT FOOT FOR APPROX ONE YR. THE PT FREQUENTLY HAS A POPPING SENSATION IN THE HIP. HE HAS A PROMINENT LIMP AND SOMETIMES NEED A CANE. THE PT SAW HIS SURGEON ON (B)(6) 2012. AN MRI AND BLOOD TEST WERE PRESCRIBED. THE PT STATES, HE IS UNABLE TO PAY FOR THE TESTING BECAUSE HE DOES NOT HAVE INSURANCE. | AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED BUT BELIEVES THEM TO BE REJUVENATE. ADD'L INFO HAS BEEN REQUESTED AND IF REC'D WILL BE SUBMITTED ON A SUPPLEMENTAL REPORT. DEVICE REMAINS IMPLANTED. | Y | Y | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02593 | 2012-12-06 | 2012-11-15 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE SURGEON REMOVED A REJUVENATE STEM FROM PT'S LEFT SIDE DUE TO EXTREME PAIN. SURGEON DISCOVERED INFECTION AND REPLACED COMPONENTS WITH A J AND J PROSTILAC STEM/ANTIBIOTIC SPACER. | IT WAS NOTED THAT THE DEVICES ARE NOT AVAILABLE FOR EVAL. ADD'L INFO HAS BEEN REQUESTED AND IF REC'D, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-12
2249697-2012-02537 | 2012-12-06 | 2012-11-14 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZ MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT REJUVENATE STEM AND NECK WERE EXTRACTED. THE CUP WAS LOOSE AND EXTRACTED. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01309 | 2012-12-06 | 2020-05-12 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT HIP REJUVENATE/ ABGII NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT THE PT HAS PAIN AT THE HIP JOINT. ANYTIME THE PT MOVES IT HURTS. PT HAD CHECK-UP VISIT (B)(6) 2012. PT IS GOING TO HAVE AN ULTRA SOUND DONE, BECAUSE, PT HAS A PACE MAKER AND CAN NOT HAVE A MRI. | AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02584 | 2012-12-06 | 2010-06-21 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-07.0000S |  | IT WAS REPORTED THAT THE PT IS EXPERIENCING PAIN ON THE FRONT AND BACK OF HER HIP AREA. PT STATES THAT THE PAIN IS WHERE STEM IS CONNECTED TO THE BALL JOINT. THE PAIN FIRST STARTED AT THE END OF (B)(6) 2012. PT IS REPORTING THAT IF SHE LIES ON HER LEFT SIDE, THE PAIN WAKES HER UP AND IS ACHY AND SHE HAS TO MOVE IN A DIFFERENT POSITION. SHE ALSO MENTIONS THAT WHEN USING THE STAIRS, SHE FEELS MORE STRESS. PT STATES THAT SHE HAD BLOOD WORK AND SHE IS AWAITING THE RESULTS. PT STATES THAT SHE HAD TO SCHEDULE AN MRI. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFO WAS REQUESTED. SHOULD ADDITIONAL INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02598 | 2012-12-06 | 2010-10-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN RIGHT REJUVENATE STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT IS REPORTED, THE PT LEARNED OF THE RECALL THROUGH A TELEVISION ADVERTISEMENT AND CONFIRMED THE RECALL STATUS OF HIS IMPLANT THROUGH HIS SURGEON. THE PT REPORTS HAVING PAIN THAT EXTENDS FROM THE TOP OF HIS HIP DOWN TO HIS KNEE, SWELLING IN THE HIP, AND NUMBNESS ON THE BOTTOM OF HIS RIGHT FOOT FOR APPROX ONE YR. THE PT FREQUENTLY HAS A POPPING SENSATION IN THE HIP. HE HAS A PROMINENT LIMP AND SOMETIMES NEED A CANE. THE PT SAW HIS SURGEON ON (B)(6) 2012. AN MRI AND BLOOD TEST WERE PRESCRIBED. THE PT STATES, HE IS UNABLE TO PAY FOR THE TESTING BECAUSE HE DOES NOT HAVE INSURANCE. | AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED BUT BELIEVES THEM TO BE REJUVENATE. ADD'L INFO HAS BEEN REQUESTED AND IF REC'D WILL BE SUBMITTED ON A SUPPLEMENTAL REPORT. DEVICE REMAINS IMPLANTED. | Y | Y | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02576 | 2012-12-06 | 2012-11-18 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT THE DR REVISED FEMORAL STEM AND NOTED DARK DISCOLORATION ON NECK/STEM TAPER. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02581 | 2012-12-06 | 2020-05-12 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT HIP REJUVENATE/ ABGII STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT HAS PAIN AT THE THIP JOINT. ANYTIME THE PT MOVES IT HURTS. PT HAD CHECK-UP VISIT (B)(6) 2012. PT IS GOING TO HAVE AN ULTRA SOUND DONE, BECAUSE PT HAS A PACE MAKER AND CAN NOT HAVE A MRI. | AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER, BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02591 | 2012-12-06 | 2012-11-14 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT, THE PATIENT COMPLAINED OF PAIN AND X-RAYS REVEALED PSEUDO TUMOR SO THE SURGEON REVISED. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED, BUT NOT MADE AVAILABLE. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02577 | 2012-12-06 | 2012-11-14 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT, THE PATIENT COMPLAINED OF EXTREME PAIN SO THE SURGEON REVISED. A LOT OF FLUID WAS NOTICED BY THE SURGEON BEFORE THE IMPLANTS WERE REMOVED. | DUE TO HOSPITAL POLICY DEVICES WILL NOT BE RETURNED. IF THE DEVICE OR ADDITIONAL INFORMATION BECOMES AVAILABLE IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. ADDITIONAL DEVICES: (B)(4), 26MM STD LFIT V40 HEAD, LOT# 39272603; (B)(4), UHR BIPOLAR 26X52MM, LOT#MJE7V9. | Y | Y | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02592 | 2012-12-06 | 2012-11-07 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | THE FOLLOWING WAS REPORTED: "TAKE OUT AND REIMPLANT." | ADD'L INFO HAS BEEN REQUESTED AND IF REC'D, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-12
MW5028128 | 2012-12-07 | 2010-10-05 | I, n, j, u, r, y | Voluntary report | STRYKER REJUVENATE | HIP | STRYKER CORP | SPT09000S |  |  | DEGENERATIVE JOINT DISEASE OF RIGHT HIP. HIP PAIN FOR PAST TWO YEARS. PAINFUL GETTING IN AND OUT OF THE CAR. PAINFUL TO LIE ON MY RIGHT SIDE. IF I SIT FOR A PERIOD OF TIME ¿ IT¿S VERY PAINFUL TO GET UP AND THE PAIN CAUSES ME TO LIMP. TWO HOURS IN THE CAR AND MY HIP STARTS TO ACHE AND I MUST STOP. EVENT: LOSS OF SIX PINTS OF BLOOD. CONCERNED THAT THEY COULD LOSE ME. DATES OF USE: (B)(6) 2010-(B)(6) 2012. |  | N | Y | LAY USER/PATIENT |  | 2012-12
MW5028127 | 2012-12-07 | 2012-08-14 | N, o, , a, n, s, w, e, r, , p, r, o, v, i, d, e, d | Voluntary report | STRYKER HIP | HIP | STRYKER |  |  |  | (B)(6) 2012, I FELL, EITHER THE HIP GAVE OUT OR I COULDN¿T LIFT MY FOOT HIGH ENOUGH TO GET OVER THE CURB. I FELL FACE FIRST, CAUGHT MYSELF WITH MY RIGHT HAND. BROKE MY HIP THAT HAD BEEN REPLACED (B)(6) 2010 AND FRACTURED MY RIGHT HAND. I HAVE HAD NOTHING BUT PROBLEMS SINCE IT WAS REPLACED IN 2010. I HAD BEEN GOING TO DOCTOR FREQUENTLY IN THOSE 2.5 YEARS COMPLAINING OF PAIN. MY HIP HAS ACTUALLY BEEN RECALLED SHORTLY BEFORE I FELL. SURGERY WAS DONE (B)(6) 2012. IT WAS A 7 HR SURGERY, I ENDED UP WITH INFECTION. BACK AND FORTH BETWEEN A REHAB FACILITY AND BACK IN THE HOSPITAL AGAIN. AT THE PRESENT TIME I AM STILL ON TWO ANTIBIOTICS, LOTS OF PAIN IN HIP AND TINGLING AND NUMBNESS IN MY RIGHT HAND, AFTER SURGERY INFECTION ACTUALLY COULD HAVE BEEN LIFE THREATENING. I HAD TO HAVE A PIC LINE PUT IN FOR ANTIBIOTICS AND BLOOD RETRIEVAL. |  | Y | Y | HEALTH PROFESSIONAL |  | 2012-12
MW5028188 | 2012-12-10 | 2012-12-07 | I, n, j, u, r, y | Voluntary report | STRYKER REJUVINATE HIP PROSTHESES | HIP PROSTHESIS, STEM FEMORAL REJUVENATE | STRYKER ORTHOPAEDICS | NLV-300800Y |  |  | PT READMITTED FOR REPLACEMENT HIP ARTHROPLASTY PROSTHESES. DIAGNOSIS OR REASON FOR USE: OSTEOARTHRITIS. |  | N | Y | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01340 | 2012-12-12 | 2019-05-02 | I, n, j, u, r, y | Manufacturer report | ABGII MODULAR LONG NECK | PROSTHESIS, HIP, SEMI-CONSTRAINED, UNCEMENTED, METAL / POLYMER, NON-POROUS, CALC | STRYKER ORTHOPAEDICS-MAHWAH | NA | 4845-4-415 |  | IT WAS REPORTED THAT THE PT HAS BEEN EXPERIENCING PAIN. PT HAS TENDERNESS IN HIP JOINT. HE DOES NOT FEEL MUCH PAIN WHEN WALKING EXCEPT WHEN GOING UP STEPS OR GOING UP AN INCLINE. | AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER, BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02606 | 2012-12-12 | 2019-05-02 | I, n, j, u, r, y | Manufacturer report | ABGII. MODULAR STEM | PROSTHESIS, HIP, SEMI-CONSTRAINED, UNCEMENTED, METAL / POLYMER, NON-POROUS, CALC | STRYKER ORTHOPAEDICS-MAHWAH | NA | 4845-4-107 |  | IT WAS REPORTED THAT THE PT HAS BEEN EXPERIENCING PAIN. PT HAS TENDERNESS IN HIP JOINT. HE DOES NOT FEEL MUCH PAIN WHEN WALKING EXCEPT WHEN GOING UP STEPS OR GOING UP AN INCLINE. | AT THIS TIME, THE PT WAS UNABLE TO IDENTIFY THE DEVICES IMPLANTED, HOWEVER, BELIEVES THEM TO BE EITHER REJUVENATE OR ABG II. ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01347 | 2012-12-12 | 2012-11-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | NLS-380000B |  | IT IS REPORTED THAT THE PT RECEIVED THE RECALL LETTER FROM HER SURGEON. THE PT REPORTS THAT SHE BEGAN HAVING A CONSTANT PAIN IN THE RIGHT HIP AREA, THIGH AND LOWER BACK APPROX ONE WEEK AGO. THE PT WILL CONTACT HER SURGEON TO SCHEDULE AN EVAL. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-12
2249697-2012-02640 | 2012-12-12 | 2012-11-20 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT PT WAS COMPLAINING OF PERSISTENT HIP PAIN. | AN EVALUATION OF THE REPORTED DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS RETAINED BY THE PT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCE IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02638 | 2012-12-12 | 2012-11-29 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT COMPONENTS WERE EXPLANTED DUE TO ADVERSE LOCAL TISSUE REACTION. | AN EVAL OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MFR DUE TO HOSPITAL POLICY. ADD'L INFO PERTAINING TO THE DEVICE(S) REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED BUT WAS NOT MADE AVAILABLE. SHOULD ADD'L INFO BECOME AVAILABLE, THE EVAL SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01351 | 2012-12-12 | 2010-12-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT: PT HAS HAD MRI AND BLOOD WORK DONE. PT STATES THAT COBALT AND CHROMIUM LEVELS ARE ELEVATED AND HE IS SCHEDULED FOR REVISION SURGERY IN (B)(6) 2013. | ADD'L INFO HAS BEEN REQUESTED AND IF REC'D WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02636 | 2012-12-12 | 2012-11-16 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT THERE WERE IMPINGEMENT ISSUES DISCOVERED AND CUP WAS TOO MEDIALIZED SO THE SURGEON REVISED. CULTURES WERE TAKEN OF FLUID FOUND IN THE HIP AND BLACKNESS WAS SEEN AROUND THE MODULAR NECK. | ADD'L INFO HAS BEEN REQUESTED AND IF REC'D, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-12
2249697-2012-02622 | 2012-12-12 | 2012-11-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT IS REPORTED THAT THE PT RECEIVED THE RECALL LETTER FROM HER SURGEON. THE PT REPORTS THAT SHE BEGAN HAVING A CONSTANT PAIN IN THE RIGHT HIP AREA, THIGH AND LOWER BACK APPROX ONE WEEK AGO. THE PT WILL CONTACT HER SURGEON TO SCHEDULE AN EVAL. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-12
9616680-2012-01346 | 2012-12-12 | 2010-06-30 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | NLS-340000B |  | IT WAS REPORTED THROUGH A LEGAL CORRESPONDENCE THAT: "PT HAS PERSONAL INJURIES SUSTAINED ON (B)(6) 2010 AS THE RESULT OF THE RECALLED STRYKER REJUVENATE MODULAR STEM." | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION NO ADD'L INFO IS AVAILABLE AT THIS TIME. IF ADD'L INFO IS RECEIVED IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-12
2249697-2012-02620 | 2012-12-12 | 2010-06-30 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THROUGH A LEGAL CORRESPONDENCE THAT: "PT HAS PERSONAL INJURIES SUSTAINED ON (B)(6) 2010 AS THE RESULT OF THE RECALLED STRYKER REJUVENATE MODULAR STEM." | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION NO ADD'L INFO IS AVAILABLE AT THIS TIME. IF ADD'L INFO IS RECEIVED IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-12
2249697-2012-02671 | 2012-12-14 | 2012-11-26 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT, MECHANICALLY ASSISTED CREVICE CORROSION AND SECONDARY ADVERSE LOCAL SOFT TISSUE REACTION AND PSEUDO TUMOR FORMATION AFTER REJUVENATE FEMORAL COMPONENT DUE TO MODULAR NECK-STEM MECHANISM FAILURE. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS RETAINED BY THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02652 | 2012-12-14 | 2012-06-20 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT DR. (B)(6) STATES THE PT PRESENTED INCREASED PAIN IN JOINT AREA, ELEVATED CHROME AND COBALT SERUM LEVELS AND A POSITIVE MARS SCAN. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02669 | 2012-12-14 | 2012-11-21 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED ADVERSE LOCAL TISSUE REACTION. | AN EVALUATION OF THE REPORTED DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION, INCLUDING X-RAYS AND MEDICAL RECORDS, WAS REQUESTED BUT NOT MADE AVAILABLE. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01364 | 2012-12-14 | 2012-11-09 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT IS REPORTED THAT: THE PATIENT HAS BEEN EXPERIENCING PAIN, SWELLING AND STIFFNESS IN THE RIGHT HIP SINCE ABOUT (B)(6) 2012. THE PAIN OCCURS WHEN STANDING FROM A SEATED POSITION AND WALKING. THE PAIN OCCURS A FEW TIMES DAILY BUT IS NOT CONSTANT. THE PATIENT COMPLETED BLOOD WORK AND SAW HIS SURGEON ON (B)(6) 2012. THE PATIENT'S COBALT LEVELS ARE ELEVATED. PATIENT IS BEING SCHEDULED FOR AN MRI IN THE NEAR FUTURE. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01399 | 2012-12-14 | 2012-12-03 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED, PT HAS INJURIES SUSTAINED AS A RESULT OF THE IMPLANTATION OF A STRYKER REJUVENATE HIP REPLACEMENT. PT RECENTLY COMPLETED MRI AND BLOOD TEST. ALSO REPORTS HAVING SOME DISCOMFORT IN RIGHT THIGH SINCE 2011. | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION NO ADDITIONAL INFO IS NOT AVAILABLE AT THIS TIME. IF ADDITIONAL INFO IS RECEIVED IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | Y | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02684 | 2012-12-14 | 2012-02-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT IS REPORTED, THE PT HAS HAD CONSTANT, MODERATE PAIN IN THE LATERAL UPPER THIGH, SWELLING AND CLICKING IN THE HIP FOR THE PAST FOUR MONTHS. BLOOD TESTS SHOW A COBALT LEVEL OF 7.0, SEDIMENTATION RATE OF 34 AND C REACTIVE PROTEIN OF 38. A MARS MRI IS SCHEDULED FOR (B)(6) 2012. THE SURGEON RECOMMENDS REVISION SURGERY. | DEVICE REMAINS IMPLANTED. ADDITIONAL INFO WAS REQUESTED AND IF RECEIVED WILL BE PROVIDED ON A SUPPLEMENTAL REPORT. | Y | Y | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02687 | 2012-12-14 | 2012-12-03 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | UNKNOWN | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED, PT HAS INJURIES SUSTAINED AS A RESULT OF THE IMPLANTATION OF A STRYKER REJUVENATE HIP REPLACEMENT. ADDITIONAL EVENT DESCRIPTION VIA SALES REP - (B)(6) 2012: IT WAS REPORTED THAT THE PT WAS REVISED FOR ALTR. THE CUP WAS LEFT INTACT, WITH A NEW LINER. NO FURTHER INFO IS AVAILABLE PER HOSPITAL PROTOCOL. | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION NO ADDITIONAL INFO IS NOT AVAILABLE AT THIS TIME. IF ADDITIONAL INFO IS RECEIVED IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | Y | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01381 | 2012-12-14 | 2012-01-10 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | NLS-340000B |  | IT WAS REPORTED THAT: PATIENT HAD TOTAL HIP REPLACEMENT USING STRYKER REJUVENATE IMPLANT. EIGHT MONTHS POST-OP AND PATIENT IS STILL HAVING ACHING PAIN IN THE POSTERIOR ANTEROLATERAL HIP. ADDITIONAL INFORMATION: PATIENT UNDERWENT REVISION SURGERY ON (B)(6) 2014. | A REVIEW OF THE DEVICE HISTORY RECORDS INDICATES THAT THE REPORTED DEVICES WERE MANUFACTURED AND ACCEPTED INTO FINAL STOCK MET SPECIFICATIONS. THE COMPLAINT HISTORY REVIEW INDICATED THAT THERE HAVE BEEN SIMILAR EVENTS FOR THE REPORTED FAMILY. THE EVENT WAS CONFIRMED. A VOLUNTARY RECALL RA 2012-067 WAS INITIATED FOR ABGII AND REJUVENATE MODULAR STEMS AND NECKS DUE TO THE POTENTIAL RISKS ASSOCIATED WITH THESE DEVICES. THE REPORTED EVENT IS CONSIDERED TO BE UNDER THE SCOPE OF THIS RECALL. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01385 | 2012-12-14 | 2012-11-21 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED BY THE PATIENT THROUGH AN ATTORNEY THAT PATIENT IS RECIPIENT OF REJUVENATE HIP IMPLANT AND WANTS TO OPEN A CLAIM. | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION, NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02672 | 2012-12-14 | 2012-11-21 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED BY THE PATIENT THROUGH AN ATTORNEY THAT PATIENT IS RECIPIENT OF REJUVENATE HIP IMPLANTS AND WANTS TO OPEN A CLAIM. | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION, NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01405 | 2012-12-14 | 2012-02-01 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT IS REPORTED, THE PT HAS HAD CONSTANT, MODERATE PAIN IN THE LATERAL UPPER THIGH, SWELLING AND CLICKING IN THE HIP FOR THE PAST FOUR MONTHS. BLOOD TESTS SHOW A COBALT LEVEL OF 7.0, SEDIMENTATION RATE OF 34 AND C REACTIVE PROTEIN OF 38. A (B)(6) MRI IS SCHEDULED FOR (B)(6) 2012. THE SURGEON RECOMMENDS REVISION SURGERY. | DEVICE REMAINS IMPLANTED. ADDITIONAL INFO WAS REQUESTED AND IF RECEIVED WILL BE PROVIDED ON A SUPPLEMENTAL REPORT. | Y | Y | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02641 | 2012-12-14 | 2012-11-20 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT, THE MODULAR STEM WAS REPLACED BY A ZIMMER MONOBLOCK STEM-PRESSFIT. THE INSERT WAS REPLACED WITH A STRYKER 36MM X3 INSERT. | AN EVALUATION OF THE REPORTED DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS RETAINED BY THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02670 | 2012-12-14 | 2012-01-10 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT: PATIENT HAD TOTAL HIP REPLACEMENT USING STRYKER REJUVENATE IMPLANT. EIGHT MONTHS POST-OP AND PATIENT IS STILL HAVING ACHING PAIN IN THE POSTERIOR ANTEROLATERAL HIP. ADDITIONAL INFORMATION RECEIVED: PATIENT UNDERWENT REVISION SURGERY ON (B)(6) 2014. | A REVIEW OF THE DEVICE HISTORY RECORDS INDICATES THAT THE REPORTED DEVICES WERE MANUFACTURED AND ACCEPTED INTO FINAL STOCK MET SPECIFICATIONS. THE COMPLAINT HISTORY REVIEW INDICATED THAT THERE HAVE BEEN SIMILAR EVENTS FOR THE REPORTED FAMILY. THE EVENT WAS CONFIRMED. A VOLUNTARY RECALL RA 2012-067 WAS INITIATED FOR ABGII AND REJUVENATE MODULAR STEMS AND NECKS DUE TO THE POTENTIAL RISKS ASSOCIATED WITH THESE DEVICES. THE REPORTED EVENT IS CONSIDERED TO BE UNDER THE SCOPE OF THIS RECALL. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02643 | 2012-12-14 | 2012-11-09 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP SYSTEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT IS REPORTED THAT: THE PATIENT HAS BEEN EXPERIENCING PAIN, SWELLING AND STIFFNESS IN THE RIGHT HIP SINCE ABOUT (B)(6) 2012. THE PAIN OCCURS WHEN STANDING FROM A SEATED POSITION AND WALKING. THE PAIN OCCURS A FEW TIMES DAILY BUT IS NOT CONSTANT. THE PATIENT COMPLETED BLOOD WORK AND SAW HIS SURGEON ON (B)(6) 2012. THE PATIENT'S COBALT LEVELS ARE ELEVATED. PATIENT IS BEING SCHEDULED FOR AN MRI IN THE NEAR FUTURE. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED IN THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02674 | 2012-12-14 | 2012-12-02 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT, DOCTOR STATES PT PRESENTED WITH INCREASED PAIN IN JOINT AREA, ELEVATED CHROME AND COBALT SERUM LEVELS, POSITIVE MARS SCAN. | IT IS NOT KNOWN AT THIS TIME IF THE DEVICE WILL BE RETURNED FOR EVALUATION. ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. ADDITIONAL DEVICES: TRIDENT 10 X 3 INSERT 32MM: 623-10-32D, ID LOT # MKEDPE. 6.5 CANCELLOUS BONE SCREW 20MM: 2030-6520, LOT # MJHMV3. TRIDENT PSL HA CLUSTER 48MM: 542-11-48D, LOT # MKE6A3. 6.5 CANCELLOUS BONE SCREW 16MM: 2030-6516-1, LOT #MKJ981. DELTA V-40 CERAMIC HEAD 32/0: 6570-0-132, LOT # 6438602. | Y | Y | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02677 | 2012-12-14 | 2012-12-03 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED, PT HAS INJURIES SUSTAINED AS A RESULT OF THE IMPLANTATION OF A STRYKER REJUVENATE HIP REPLACEMENT. PT RECENTLY COMPLETED MRI AND BLOOD TEST. ALSO REPORTS HAVING SOME DISCOMFORT IN RIGHT THIGH SINCE 2011. | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION, NO ADDITIONAL INFO IS AVAILABLE AT THIS TIME. IF ADDITIONAL INFO IS RECEIVED, IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | Y | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02686 | 2012-12-14 | 2012-11-23 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT, THE MODULAR STEM, HEAD, LINER WERE REMOVED. THE STEM WAS REPLACED WITH A ZIMMER 8' STEM + 36 HEAD. HE USED A (B)(4) FOR A NEW INSERT, HE LEFT A CUP IN. | AN EVALUATION OF THE REPORTED DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS RETAINED BY THE PATIENT AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02675 | 2012-12-14 | 2012-11-26 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP SYSTEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THROUGH FILING OF A LAWSUIT THAT THE PT ALLEGEDLY REC'D A REJUVENATE HIP SYSTEM. IT IS FURTHER ALLEGED THAT AFTER IMPLANTATION THE PT BEGAN TO EXPERIENCE SIGNIFICANT DISCOMFORT IN THE AREA OF THE DEVICE AND IS SCHEDULED FOR REVISION SURGERY. | ADD'L INFO HAS BEEN REQUESTED AND IF REC'D, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01392 | 2012-12-14 | 2012-11-26 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THROUGH FILING OF A LAWSUIT THAT THE PT ALLEGEDLY REC'D A REJUVENATE HIP SYSTEM. IT IS FURTHER ALLEGED THAT AFTER IMPLANTATION THE PT BEGAN TO EXPERIENCING SIGNIFICANT DISCOMFORT IN THE AREA OF THE DEVICE AND IS SCHEDULED FOR REVISION SURGERY. | ADD'L INFO HAS BEEN REQUESTED AND IF REC'D, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02673 | 2012-12-14 | 2012-12-02 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT, DOCTOR STATES PATIENT PRESENTED WITH INCREASED PAIN IN JOINT AREA, ELEVATED CHROME AND COBALT SERUM LEVELS, POSITIVE MARS SCAN. | IT IS NOT KNOWN AT THIS TIME IF THE DEVICE WILL BE RETURNED FOR EVALUATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. ADDITIONAL DEVICES: 623-10-36E TRIDENT 10 X 3 INSERT 36MM, LOT # MKPVJD; 2030-6520-1 6.5 CANCELLOUS BONE SCREW 20MM, LOT# MKN983; 542-11-50E TRIDENT PSL HA CLUSTER 50MM, LOT# MKMY94; 2030-6516-1 6.5 CANCELLOUS BONE SCREW 16MM, LOT# MKJ2XT; 6570-0-136 DELTA V-40 CERAMIC HEAD 36/0, LOT# 37896301. | Y | Y | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01394 | 2012-12-14 | 2012-11-26 | I, n, j, u, r, y | Manufacturer report | REJUVENATE HIP NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT: "PT HAS INJURIES SUSTAINED AS A RESULT OF THE IMPLANTATION OF A STRYKER REJUVENATE HIP REPLACEMENT." | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION NO ADD'L INFO IS AVAILABLE AT THIS TIME. IF ADD'L INFO IS REC'D IT WILL BE REPORTED ON A SUPPLEMENTAL. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01400 | 2012-12-14 | 2012-11-26 | I, n, j, u, r, y | Manufacturer report | REJUVENATE HIP NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT: "PT HAS INJURIES SUSTAINED AS A RESULT OF THE IMPLANTATION OF A STRYKER REJUVENATE HIP REPLACEMENT." | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION NO ADD'L INFO IS AVAILABLE AT THIS TIME. IF ADD'L INFO IS REC'D IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02679 | 2012-12-14 | 2012-11-26 | I, n, j, u, r, y | Manufacturer report | REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: "PT HAS INJURIES SUSTAINED AS A RESULT OF THE IMPLANTATION OF A STRYKER REJUVENATE HIP REPLACEMENT." | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION NO ADD'L INFO IS AVAILABLE AT THIS TIME. IF ADD'L INFO IS REC'D IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01403 | 2012-12-14 | 2012-11-26 | I, n, j, u, r, y | Manufacturer report | REJUVENATE HIP NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT: "PATIENT HAS INJURIES SUSTAINED AS A RESULT OF THE IMPLANTATION OF A STRYKER REJUVENATE HIP REPLACEMENT". | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02676 | 2012-12-14 | 2012-11-26 | I, n, j, u, r, y | Manufacturer report | REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: "PT HAS INJURIES SUSTAINED AS A RESULT OF THE IMPLANTATION OF A STRYKER REJUVENATE HIP REPLACEMENT. | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION NO ADD'L INFO IS AVAILABLE AT THIS TIME. IF ADD'L INFO IS REC'D IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02645 | 2012-12-14 | 2012-11-19 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT PATIENT HAD REVISED SURGERY FOR REMOVAL OF THE STRYKER REJUVENATE STEM. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | OTHER |  | 2012-12
2249697-2012-02681 | 2012-12-14 | 2012-11-26 | I, n, j, u, r, y | Manufacturer report | REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT: "PATIENT HAS INJURIES SUSTAINED AS A RESULT OF THE IMPLANTATION OF A STRYKER REJUVENATE HIP REPLACEMENT". | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02688 | 2012-12-14 | 2012-11-22 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP SYSTEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT, ALL THE SURGEON TOLD ME WAS THAT THE PATIENT FELL. I WAS NOT ABLE TO GET THE IMPLANT BECAUSE THE HOSPITAL SENT IT TO (B)(6) IN RISK MANAGEMENT. | AN EVALUATION OF THE REPORTED DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS RETAINED BY THE HOSPITAL AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02690 | 2012-12-14 | 2012-11-26 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT, THE SURGEON REVISED PATIENT'S LEFT HIP DUE TO PAIN AND ELEVATED COBALT LEVEL OF 3.9. REPLACED WITH A NEW CUP, LINER, HEAD AND EXETER STEM. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER DUE TO HOSPITAL POLICY. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01411 | 2012-12-14 | 2012-11-26 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT, THE SURGEON REVISED PATIENT'S LEFT HIP DUE TO PAIN AND ELEVATED COBALT LEVEL OF 3.9. REPLACED WITH A NEW CUP, LINER, HEAD AND EXETER STEM. | AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER DUE TO HOSPITAL POLICY. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01410 | 2012-12-14 | 2012-12-03 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK 127/132 | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | NLS-340000B |  | IT WAS REPORTED, PATIENT HAS INJURIES SUSTAINED AS A RESULT OF THE IMPLANTATION OF A STRYKER REJUVENATE HIP REPLACEMENT. IT WAS REPORTED THAT THE PATIENT WAS REVISED FOR ALTR. THE CUP WAS LEFT INTACT, WITH A NEW LINER. | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO ONGOING LITIGATION NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED IT WILL BE REPORTED IN A SUPPLEMENTAL REPORT. | Y | Y | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01409 | 2012-12-14 | 2012-11-22 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT, ALL THE SURGEON TOLD ME WAS THAT THE PATIENT FELL. I WAS NOT ABLE TO GET THE IMPLANT BECAUSE THE HOSPITAL SENT IT TO R.S. IN RISK MANAGEMENT. | AN EVALUATION OF THE REPORTED DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS RETAINED BY THE HOSPITAL AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION PERTAINING TO THE DEVICE REFERENCED IN THIS REPORT (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01425 | 2012-12-17 | 2011-10-10 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED, PT HAS PERSONAL INJURIES SUSTAINED ON (B)(6) 2011 AS A RESULT OF THE RECALLED STRYKER REJUVENATE MODULAR STEM. | THE EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO ONGOING LITIGATION NO ADDITIONAL INFO IS AVAILABLE AT THIS TIME. IF ADDITIONAL INFO IS RECEIVED, IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01415 | 2012-12-17 | 2012-07-18 | I, n, j, u, r, y | Manufacturer report | REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THROUGH THE ATTORNEY FOR THE PT, AS A RESULT OF A LEGAL CLAIM THAT, PLAINTIFF ALLEGES FAILURE OF THE REJUVENATE HIP THAT WAS IMPLANTED. ALLEGEDLY AFTER SYMPTOMS WORSENED, PLAINTIFF UNDERWENT A REVISION SURGERY ON (B)(6) 2012. DURING THAT SURGERY, IT WAS DISCOVERED THAT, IN FACT, THERE WAS SIGNIFICANT EVIDENCE OF HEAVY METAL TOXICITY INCLUDING A LARGE PSEUDOTUMOR FORMATION AND SOFT TISSUE NECROSIS AT THE PROXIMAL FEMUR. | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION NO ADDITIONAL INFO IS AVAILABLE AT THIS TIME. IF ADDITIONAL INFO IS RECEIVED IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02706 | 2012-12-17 | 2011-10-10 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE HIP STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED, PT HAS PERSONAL INJURIES SUSTAINED ON (B)(6) 2011 AS A RESULT OF THE RECALLED STRYKER REJUVENATE MODULAR STEM. | THE EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO ONGOING LITIGATION NO ADDITIONAL INFO IS AVAILABLE AT THIS TIME. IF ADDITIONAL INFO IS RECEIVED, IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | Y | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02705 | 2012-12-17 | 2012-08-21 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THROUGH THE ATTORNEY FOR THE PATIENT, AS A RESULT OF A LEGAL CLAIM THAT, PLAINTIFF ALLEGES FAILURE OF THE REJUVENATE HIP THAT WAS IMPLANTED. ALLEGEDLY DIAGNOSTIC WORKUP, INCLUDING A THREE-PHASE BONE SCAN CONDUCTED ON (B)(6) 2012, REVEALED THE ABSENCE OF DEVICE LESSENING OR INFECTION. FURTHER DIAGNOSTIC WORKUP, INCLUDING MAGNETIC RESONANCE IMAGING (MRI) CONDUCTED ON (B)(6) 2012, REVEALED THE EXISTENCE OF A SIGNIFICANT FLUID COLLECTION ABOUT THE HIP PROSTHESIS. PLAINTIFF UNDERWENT A REVISION SURGERY ON (B)(6) 2012. DURING THAT SURGERY, IT WAS DISCOVERED THAT, IN FACT, THERE WAS SIGNIFICANT EVIDENCE OF HEAVY METAL TOXICITY INCLUDING A LARGE PSEUDOTUMOR FORMATION AND SOFT TISSUE NECROSIS AT THE PROXIMAL FEMUR. | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01424 | 2012-12-17 | 2012-08-21 | I, n, j, u, r, y | Manufacturer report | REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THROUGH THE ATTORNEY FOR THE PATIENT, AS A RESULT OF A LEGAL CLAIM, THAT PLAINTIFF ALLEGES FAILURE OF THE REJUVENATE HIP THAT WAS IMPLANTED. ALLEGEDLY, DIAGNOSTIC WORKUP, INCLUDING A THREE-PHASE BONE SCAN CONDUCTED ON (B)(6) 2012, REVEALED THE ABSENCE OF DEVICE LOOSENING OR INFECTION. FURTHER DIAGNOSTIC WORKUP, INCLUDING MAGNETIC RESONANCE IMAGING (MRI) CONDUCTED ON (B)(6) 2012, REVEALED THE EXISTENCE OF A SIGNIFICANT FLUID COLLECTION ABOUT THE HIP PROSTHESIS. PLAINTIFF UNDERWENT A REVISION SURGERY ON (B)(6) 2012. DURING THAT SURGERY, IT WAS DISCOVERED THAT, IN FACT, THERE WAS SIGNIFICANT EVIDENCE OF HEAVY METAL TOXICITY INCLUDING A LARGE PSEUDOTUMOR FORMATION AND SOFT TISSUE NECROSIS AT THE PROXIMAL FEMUR. | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION, NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED, IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02702 | 2012-12-17 | 2012-05-24 | I, n, j, u, r, y | Manufacturer report | LEFT HIP REJUVENATE STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT BROKE HIP IN (B)(6) 2012. PATIENT WAS EXPERIENCING EXTREME PAIN. PATIENT HAD A REVISION WITH A NON-STRYKER HIP. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01421 | 2012-12-17 | 2012-05-24 | I, n, j, u, r, y | Manufacturer report | LEFT HIP REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT THE PT BROKE HIP IN (B)(6) 2012. PT WAS EXPERIENCING EXTREME PAIN. PT HAD A REVISION WITH A NON-STRYKER HIP. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01418 | 2012-12-17 | 2010-07-19 | I, n, j, u, r, y | Manufacturer report | LEFT HIP REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT THE PT IS EXPERIENCING PAIN AND FLUID RETENTION. PT SALES THAT THEY CAN HARDLY WALK AND THEY WILL BE HAVING A REVISION SURGERY SOON. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02696 | 2012-12-17 | 2012-07-18 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THROUGH THE ATTORNEY FOR THE PT, AS A RESULT OF A LEGAL CLAIM THAT, PLAINTIFF ALLEGES FAILURE OF THE REJUVENATE HIP THAT WAS IMPLANTED. ALLEGEDLY AFTER SYMPTOMS WORSENED, PLAINTIFF UNDERWENT A REVISION SURGERY ON (B)(6) 2012. DURING THAT SURGERY, IT WAS DISCOVERED THAT, IN FACT, THERE WAS SIGNIFICANT EVIDENCE OF HEAVY METAL TOXICITY INCLUDING A LARGE PSEUDOTUMOR FORMATION AND SOFT TISSUE NECROSIS AT THE PROXIMAL FEMUR. | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION NO ADDITIONAL INFO IS AVAILABLE AT THIS TIME. IF ADDITIONAL INFO IS RECEIVED IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02697 | 2012-12-17 | 2012-11-29 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT MECHANICALLY ASSISTED CREVICE CORROSION AND SECONDARY ADVERSE LOCAL SOFT TISSUE REACTION AND PSEUDOTUMOR FORMATION AFTER REJUVENATE FEMORAL COMPONENT DUE TO MODULAR NECK-STEM MECHANISM OF FAILURE. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02699 | 2012-12-17 | 2012-07-19 | I, n, j, u, r, y | Manufacturer report | LEFT HIP REJUVENATE STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT IS EXPERIENCING PAIN AND FLUID RETENTION. PT STATES THAT THEY CAN HARDLY WALK AND THEY WILL BE HAVING A REVISION SURGERY SOON. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02729 | 2012-12-18 | 2012-02-13 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT: PATIENT ORIGINALLY CALLED (B)(6) 2012 AND WAS ASYMPTOMATIC AND EXPERIENCING NO PAIN. PATIENT WAS REPORTING THAT SHE HAD BLOOD WORK AND X-RAYS PERFORMED ON (B)(6) 2012. PATIENT ALSO STATES THAT SHE HAD AN MRI ON (B)(6) 2012. PATIENT HAS SINCE RECEIVED RESULTS FROM THE SURGEON AND IS SCHEDULED FOR REVISION SURGERY ON (B)(6) 2013. PATIENT ALSO STATES THAT FLUID WAS ASPIRATED FROM HER HIP. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02722 | 2012-12-18 | 2012-07-16 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THROUGH THE ATTORNEY FOR THE PATIENT, AS A RESULT OF A LEGAL CLAIM, THAT PLAINTIFF ALLEGES FAILURE OF THE REJUVENATE HIP THAT WAS IMPLANTED. ALLEGEDLY PLAINTIFF BEGAN EXPERIENCING SIGNIFICANT PAIN AND DISCOMFORT IN THE AREA OF THE DEFECTIVE DEVICE. DIAGNOSTIC WORKUP REVEALED ELEVATED COBALT LEVELS INDICATING THE PRESENCE OF HEAVY METAL ION CONTAMINATION. PLAINTIFF UNDERWENT A REVISION SURGERY ON (B)(6) 2012. | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO ONGOING LITIGATION NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01434 | 2012-12-18 | 2012-07-16 | I, n, j, u, r, y | Manufacturer report | REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THROUGH THE ATTORNEY FOR THE PATIENT, AS A RESULT OF A LEGAL CLAIM, THAT PLAINTIFF ALLEGES FAILURE OF THE REJUVENATE HIP THAT WAS IMPLANTED. ALLEGEDLY PLAINTIFF BEGAN EXPERIENCING SIGNIFICANT PAIN AND DISCOMFORT IN THE AREA OF THE DEFECTIVE DEVICE. DIAGNOSTIC WORKUP REVEALED ELEVATED COBALT LEVELS INDICATING THE PRESENCE OF HEAVY METAL ION CONTAMINATION. PLAINTIFF UNDERWENT A REVISION SURGERY ON (B)(6) 2012. | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION, NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02733 | 2012-12-18 | 2011-07-21 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT BEGAN EXPERIENCING PAIN IN (B)(6) 2012. THE PATIENT REPORTS SLIGHT PAIN ON THE OUTSIDE OF THE LEFT HIP AT THE INCISION SITE. THE PATIENT ALSO REPORTS CONSTANT, INTENSE GROIN PAIN. THE PATIENT STATES THEY CANNOT STAND FOR SIGNIFICANT PERIODS OF TIME AND MUST USE A CANE TO MAINTAIN STABILITY. THE PATIENT HAD AN MRI AND TWO HIP ASPIRATIONS. BLOOD TESTS RECORD COBALT LEVELS OF 4.7 AND 5.1 RESPECTIVELY OVER A TWO MONTH PERIOD. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02714 | 2012-12-18 | 2012-11-28 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT PT WAS HAVING PAIN, ELEVATED COBALT LEVELS SO PT WAS REVISED. | AN EVALUATION OF THE REPORTED DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS RETAINED BY THE HOSPITAL AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION (INCLUDING X-RAYS AND MEDICAL RECORDS) HAS BEEN REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02720 | 2012-12-18 | 2012-12-02 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT, DR (B)(6) STATES PATIENT PRESENTED WITH INCREASED PAIN IN JOINT AREA, ELEVATED CHROME AND COBALT SERUM LEVELS, POSITIVE MARS SCAN. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02713 | 2012-12-18 | 2012-08-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE LEFT STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THROUGH THE ATTORNEY FOR THE PATIENT, AS A RESULT OF A LEGAL CLAIM THAT, PLAINTIFF ALLEGES EXPERIENCING SIGNIFICANT PAIN AND DISCOMFORT IN THE AREA OF THE DEFECTIVE DEVICE. THE LEFT REJUVENATE HIP WAS EXPLANTED ON (B)(6) HOSPITAL IN (B)(6). | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION, NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED, IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01430 | 2012-12-18 | 2012-08-01 | I, n, j, u, r, y | Manufacturer report | REJUVENATE LEFT NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THROUGH THE ATTORNEY FOR THE PATIENT, AS A RESULT OF A LEGAL CLAIM THAT, PLAINTIFF ALLEGES EXPERIENCING SIGNIFICANT PAIN AND DISCOMFORT IN THE AREA OF THE DEFECTIVE DEVICE. THE LEFT REJUVENATE HIP WAS EXPLANTED ON (B)(6) 2012 AT (B)(6) HOSPITAL. | THIS EVENT WAS REPORTED THROUGH AN ATTORNEY, AS A RESULT OF A LEGAL CLAIM. DUE TO THE ONGOING LITIGATION, NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01442 | 2012-12-18 | 2011-07-21 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT THE PATIENT BEGAN EXPERIENCING PAIN IN (B)(6) 2012. THE PATIENT REPORTS SLIGHT PAIN ON THE OUTSIDE OF THE LEFT HIP AT THE INCISION SIGHT. THE PATIENT ALSO REPORTS CONSTANT, INTENSE GROIN PAIN. THE PATIENT STATES THEY CANNOT STAND FOR SIGNIFICANT PERIODS OF TIME AND MUST REGULARLY USE A CANE TO MAINTAIN STABILITY. THE PATIENT HAD AN MRI AND TWO HIP ASPIRATIONS. BLOOD TEST RECORD COBALT LEVELS OF 4.7 AND 5.1 RESPECTIVELY OVER A TWO MONTH PERIOD. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01437 | 2012-12-18 | 2012-02-13 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | NLS-34000B |  | IT WAS REPORTED THAT: PT ORIGINALLY CALLED ON (B)(6)2012 AND WAS ASYMPTOMATIC AND EXPERIENCING NO PAIN. PT WAS REPORTING THAT SHE HAD BLOOD WORK AND X-RAYS PERFORMED ON (B)(6) 2012. PT ALSO STATES THAT SHE HAD AN MRI ON (B)(6) 2012. PT HAS SINCE RECEIVED RESULTS FROM THE SURGEON AND IS SCHEDULED FOR REVISION SURGERY IN (B)(6) 2012. PT ALSO STATES THAT FLUID WAS ASPIRATED FROM HER HIP. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02731 | 2012-12-18 | 2012-11-29 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT, MECHANICALLY ASSISTED CREVICE CORROSION AND SECONDARY ADVERSE LOCAL SOFT TISSUE REACTION AND PSEUDOTUMOR FORMATION AFTER REJUVENATE FEMORAL COMPONENT DUE TO MODULAR NECK-STEM MECHANISM OF FAILURE. | ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01446 | 2012-12-19 | 2011-08-03 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT HIP REJUVENATE NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | UNK |  | IT WAS REPORTED THAT THE PT IS HAVING PAIN IN THE HIP AREA. PT STATES THAT HE HAS EXPERIENCED THE PAIN FOR THE LAST SIX TO SEVEN MONTHS. THE PT VISITED THE SURGEON AND IS SCHEDULED FOR AN MRI. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02747 | 2012-12-19 | 2011-08-03 | I, n, j, u, r, y | Manufacturer report | UNKNOWN LEFT HIP REJUVENATE STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT THE PT IS HAVING PAIN IN THE HIP AREA. PT STATES THAT HE HAS EXPERIENCED THE PAIN FOR THE LAST SIX TO SEVEN MONTHS. THE PT VISITED THE SURGEON AND IS SCHEDULED FOR AN MRI. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02754 | 2012-12-19 | 2012-12-05 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-070000S |  | IT WAS REPORTED THAT, SURGEON REVISED PATIENT'S RIGHT HIP DUE TO A FAILED REJUVENATE STEM. PATIENT HAD HIP PAIN, SALES REP BELIEVES PATIENT HAD ELEVATED ION LEVELS AS WELL. | HOSPITAL RETAINED DEVICES. IF THE DEVICES OR ADDITIONAL INFORMATION BECOMES AVAILABLE IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02752 | 2012-12-19 | 2012-12-02 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-0700000S |  | IT WAS REPORTED THAT, DR (B)(6) STATES PT PRESENTED WITH INCREASED PAIN IN JOINT AREA, ELEVATED CHROME AND COBALT SERUM LEVELS, POSITIVE MARS SCAN. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02753 | 2012-12-19 | 2012-11-30 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 10 | UNK | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-100000S |  | IT WAS REPORTED THAT THE PT HAD A REJUVENATE REVISION. BLOOD SERUM LEVEL WAS ELEVATED. THE PT'S COBALT LEVEL WAS 6.9 AND THE CHROMIUM LEVEL WAS 1.12. | ADDITIONAL INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02746 | 2012-12-19 | 2012-11-29 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-080000S |  | IT WAS REPORTED THAT, INCISION WAS MADE IN THE DIRECT ANTERIOR APPROACH DUE TO THE LOCATION OF THE ABSCESS ON PT. INCISION WAS MADE AND ABSCESS WAS DRAINED. DR. (B)(6) CONTINUED HIS APPROACH INTO THE JOINT. AFTER ADEQUATE EXPOSURE WAS ACHIEVED HE PROCEEDED TO KNOCK THE BIPOLAR COMPONENT AND HEAD OFF WITH A BONE TAMP AND MALLET. NEXT, HE REMOVED THE FEMORAL NECK USING THE NECK REMOVAL TOOL. THERE WAS BLACK RESIDUE ON THE NECK AT THE BASE OF THE IMPLANT. THERE WERE NO SIGNS OF "BLACK" RESIDUE ON THE SOFT TISSUE. AFTER NECK WAS REMOVED, DR. (B)(4) PROCEEDED TO REMOVE THE STEM. STEM WAS EXPLANTED SUCCESSFULLY, AND THE... | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02751 | 2012-12-19 | 2012-11-30 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT, SURGEON REMOVED REJUVENATE STEM AND REPLACED IT WITH A RESTORATION MODULAR STEM. PSEUDO TUMOR. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02756 | 2012-12-19 | 2012-12-04 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT, PATIENT PRESENTED WITH HIP PAIN. SURGEON STATES THAT WORKUP CONSISTENT WITH ALTR. | DEVICE IS NOT AVAILABLE. IF THE DEVICE OR ADDITIONAL INFORMATION BECOMES AVAILABLE IT WILL BE REPORTED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01454 | 2012-12-19 | 2012-09-28 | I, n, j, u, r, y | Manufacturer report | REJUVENATE MODULAR NECK | IMPLANT | STRYKER ORTHOPAEDICS CORK | NA | NLS-3000000B |  | IT WAS REPORTED THAT PER DOCTOR (B)(6) OFFICE, THIS PATIENT IS EXPERIENCING DIFFICULTIES WITH THEIR HIP IMPLANT AND HAS HAD BLOOD TESTS AND MRI. UPDATE: PATIENT STATED THAT BOTH HIPS ARE EXPERIENCING "CLICKING AND POPPING" AND GROIN PAIN. PATIENT IS EXPERIENCING DIFFICULTY LIFTING LEGS GOING UPSTAIRS. PATIENT IS HAVING DIFFICULTY STANDING TO GET OUT OF THE CAR AND HAS TO POP HIS HIPS BACK INTO PLACE. PATIENT IS SCHEDULED FOR REVISION SURGERY ON (B)(6) 2012 FOR RIGHT HIP. LEFT HIP WILL BE REVISED APPROXIMATELY 8 WEEKS AFTER RIGHT HIP. ADDITIONALLY, REPORTED ON (B)(4) 2012 VIA SALES REP: PATIENT WAS REVISED FROM A REJUVENATE TO A CEMENT SPACER ON (B)(6) 2012 DUE TO A DISCOVERED INFECTION (GRAM STAIN). THE PATIENT HAD COMPLAINED OF PAIN AND WAS DIAGNOSED WITH A 22 METAL BLOOD SERUM LEVEL ON (B)(6) 2012. THE CEMENTED STEM ANTIBIOTIC SPACER WAS REMOVED AND A NEW SECURE-FIT STEM, HEAD, TRIDENT PSL AND MDM LINER AND POLY WAS IMPLANTED. | AN EVALUATION OF THE REPORTED DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS RETAINED BY THE HOSPITAL AND WAS NOT RETURNED TO THE MANUFACTURER. ADDITIONAL INFORMATION WAS REQUESTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE EVALUATION SUMMARY WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02757 | 2012-12-19 | 2012-12-02 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | KNA | SPT-080000S |  | IT WAS REPORTED THAT, DR. (B)(6) STATES PATIENT PRESENTED WITH INCREASED PAIN IN JOINT AREA, ELEVATED CHROME AND COBALT SERUM LEVELS, POSITIVE MARS SCAN. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
9616680-2012-01465 | 2012-12-19 | 2012-12-05 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE NECK | IMPLANT | STRYKER IRELAND, LTD. OSTEONICS (ORTHOPAEDICS DIVI | NA | UNK |  | IT WAS REPORTED THAT METALLOSIS AT JUNCTURE OF THE MODULAR NECK CONNECTED TO THE HIP STEM. | IT IS NOT KNOWN AT THIS TIME IF THE PATIENT WAS REVISED. ADD'L INFO WAS REQUESTED AND IF RECEIVED WILL BE PROVIDED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02760 | 2012-12-19 | 2012-12-05 | I, n, j, u, r, y | Manufacturer report | UNKNOWN REJUVENATE STEM | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | UNK |  | IT WAS REPORTED THAT METALLOSIS AT JUNCTURE OF THE MODULAR NECK CONNECTED TO THE HIP STEM. | IT IS NOT KNOWN AT THIS TIME IF THE PT WAS REVISED. ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED WILL BE PROVIDED ON A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02762 | 2012-12-19 | 2012-12-02 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT, DR. (B)(6) STATES PATIENT PRESENTED WITH INCREASED PAIN IN JOINT AREA, ELEVATED CHROME AND COBALT SERUM LEVELS, POSITIVE MARS SCAN. | ADD'L INFO HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. | Y | N | HEALTH PROFESSIONAL |  | 2012-12
2249697-2012-02749 | 2012-12-19 | 2012-12-02 | I, n, j, u, r, y | Manufacturer report | REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | IMPLANT | STRYKER ORTHOPAEDICS MAHWAH | NA | SPT-090000S |  | IT WAS REPORTED THAT DR (B)(6) STATES PRESENTED INCREASED PAIN IN JOINT AREA, ELEVATED CHROME AND COBALT SERUM LEVELS, POSITIVE MARS SCAN. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | Y | N | HEALTH PROFESSIONAL |  | 2012-12

## Monthly Summary
Month | Count | Date
2012-06 | 7 | 2012-06-01
2012-07 | 47 | 2012-07-01
2012-08 | 133 | 2012-08-01
2012-09 | 114 | 2012-09-01
2012-10 | 195 | 2012-10-01
2012-11 | 157 | 2012-11-01
2012-12 | 115 | 2012-12-01

## Event Types
Event Type | Count
I, n, j, u, r, y | 761
M, a, l, f, u, n, c, t, i, o, n | 4
N, o, , a, n, s, w, e, r, , p, r, o, v, i, d, e, d | 3

## Brand Names
Brand Name | Count
REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 7 | 85
REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 8 | 75
REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 9 | 59
UNKNOWN REJUVENATE NECK | 39
UNKNOWN RIGHT REJUVENATE HIP | 37
UNKNOWN REJUVENATE HIP DEVICE | 35
UNKNOWN LEFT REJUVENATE HIP | 34
UNKNOWN REJUVENATE HIP STEM | 27
REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 10 | 26
UNKNOWN RIGHT REJUVENATE/ABG II HIP | 12
REJUVENATE MODULAR NECK | 12
UNKNOWN REJUVENATE STEM | 10
UNKNOWN REJUVENATE HIP | 10
UNKNOWN RIGHT HIP REJUVENATE | 10
UNKNOWN REJUVENATE RIGHT HIP | 10
UNKNOWN RIGHT REJUVENATE/ABGII HIP | 8
UNK LEFT REJUVENATE HIP | 8
UNKNOWN REJUVENATE | 8
REJUVENATE MODULAR NECK 34MM V40 | 7
UNKNOWN LEFT REJUVENATE/ABG II HIP | 7
UNKNOWN RIGHT REJUVENATE | 6
UNKNOWN REJUVENATE LEFT HIP | 6
UNKNOWN LEFT HIP REJUVENATE | 6
UNKNOWN REJUVENATE HIP SYSTEM | 5
REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 11 | 4
UNKNOWN LEFT HIP REJUVENATE/ABG II | 4
REJUVENATE STEM EXTRACTOR | 4
REJUVENATE HIP STEM | 4
UNK RIGHT REJUVENATE HIP | 4
REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 12 | 3
REJUVENATE MODULAR NECK 38MM V40 | 3
REJUVENATE MODULAR NECK 30MM V40 | 3
ABGII. MODULAR STEM | 3
UNKNOWN LEFT REJUVENATE HIP SYSTEM | 3
UNKNOWN ABGII OR REJUVENATE HIP STEM | 3
UNKNOWN RIGHT REJUVENATE MODULAR HIP | 3
UNK LEFT REJUVENATE/ABG II HIP | 3
REJUVENATE HIP NECK | 3
UNKNOWN RIGHT HIP - REJUVENATE/ABGII | 3
UNKNOWN LEFT REJUVENATE | 3
UNKNOWN RIGHT REJUVENATE/AGBII HIP | 3
REJUVENATE STEM | 3
UNKNOWN ABGII OR REJUVENATE NECK | 3
STRYKER REJUVENATE | 3
REJUVENATE NECK | 3
UNK RIGHT REJUVENATE/ABG II HIP | 3
UNKNOWN LEFT REJUVENATE/ABGII HIP | 3
UNK REJUVENATE STEM | 2
UNKNOWN LEFT REJUVENATE MODULAR HIP | 2
UNKNOWN LEFT REJUVENATE MODULAR STEM | 2
REJUVENATE MODULAR HIP STEM | 2
REJUVENATE HIP COMPONENTS | 2
UNKNOWN LEFT HIP REJUVENATE/ABG II HIP | 2
UNKNOWN LEFT REJUVENATE/AGBII HIP | 2
UNK REJUVENATE HIP DEVICE | 2
UNKNOWN LEFT HIP REJUVENATE/ABGII | 2
REJUVENATE MODULAR NECK 34MM | 2
UNKNOWN LEFT REJUVENATE HIP STEM | 2
UNK REJUVENATE MODULAR NECK | 2
UNKNOWN LEFT REJUVENATE NECK | 2
LEFT HIP REJUVENATE NECK | 2
UNKNOWN RIGHT REJUVENATE STEM | 2
LEFT REJUVENATE HIP | 2
LEFT HIP REJUVENATE STEM | 2
UNKNOWN REJUVENATE HIP NECK | 2
UNKNOWN RIGHT REJUVENATE MODULAR NECK HIP STEM | 2
REJUVENATE LEFT STEM | 1
UNKNOWN LEFT REJUVENATE STEM | 1
UNK LEFT REJUVENATE/AGBII HIP | 1
REJUVENATE STRGHT PRFIT TMZ MOD STEM SIZE 9 | 1
REJUVENATE MODULAR NECK34MM V40 | 1
UNKNOWN LEFT HIP REJUVENATE/ ABGII NECK | 1
UNKNOWN LEFT HIP REJUVENATE/ ABGII STEM | 1
STRYKER HIP | 1
REJUVENATE STRGHT TMZF MOD STEM SIZE 8 | 1
UNKNOWN LEFT HIP REJUVENATE NECK | 1
UNKNOWN RIGHT REJUVENATE HIP DEVICE | 1
UNK RIGHT REJUVENATE/ABGII HIP | 1
STRYKER REJUVINATE HIP PROSTHESES | 1
UNKNOWN RIGHT REJUVENATE NECK | 1
UNKNOWN RIGHT REJUVENATE MODULAR HIP STEM | 1
ABGII MODULAR LONG NECK | 1
UNK LEFT REJUVENATE/ABGII HIP | 1
REJUVENATE MODULAR NECK 127/132 | 1
REJUVENATE LEFT NECK | 1
UNK LFT HIP REJUVENATE | 1
UNK - RIGHT HIP REJUVENATE/ABGII | 1
REJUVENATE MODULAR NECK 38MM | 1
UNKNOWN LEFT HIP REJUVENATE/ABGII NECK | 1
UNKNOWN ABGII OR REJUVENATE HIP SYSTEM | 1
UNKNOWN REJUVENATE STEM SIZE 8 | 1
UNKNOWN ABGII OR REJUVENATE HIP | 1
1
UNKNOWN PRODUCT - ABGII OR REJUVENATE NECK | 1
REJUVENATE STRGHT PRIFIT TMZF MOD STEM SIZE 7 | 1
REJUVENATE STRGHT TMZF MOD STEM SIZE 7 | 1
REJUVENATE MODULAR NECK 30MM | 1
UNKNOWN REJUVENATE MODULAR NECK HIP IMPLANTS | 1
TRIDENT HEMISPHERICAL ACETABULAR SHELL | 1
UNKNOWN RIGHT HIP REJUVENATE HIP | 1
DELTA V-40 CERAMIC HEAD 32/+4 | 1
REJUVENATE MODULAR NECK 127/132 34MM V40 | 1
UNKNOWN #7 REJUVENATE STEM | 1
UNKNOWN REJUVENATE MODULAR STEM | 1
REVOLUTION CEMENT MIXING SYSTEM WITH FEMORAL | 1
STRYKER REJUVENATE SPT | 1
STRYKER REJUVENATE/ABG II | 1
UNKNOWN REJUVENATE RIGHT HIP STEM | 1
UNKNOWN REJUVENATE 34MM MODULAR NECK | 1
REJUVENATE MODULAR NECK 42MM V40 | 1
UNKNOWN LEFT REJUVENATE MODULAR NECK | 1
UNKNOWN LEFT HIP REJUVENATE /ABGII | 1
UNKNOWN REJUVENATE/ABGII RIGHT HIP | 1
UNKNOWN REJUVENATE #9 STEM LEFT | 1
UNKNOWN LEFT HIP REJUVENATE SIZE 9 | 1
UNKNOWN SIZE 12 REJUVENATE STEM | 1
UNKNOWN 34 MM REJUVENATE NECK 132 DEG ANGLE NEUTRAL VERSIO | 1
REJUVENATE STRAIGHT PRFIT TMZF MOD STEM SIZE 10 | 1
REJUVENATE RIGHT HIP | 1
UNKNOWN REJUVENATE 34MM NECK 127 DEGREE X 8 ANTEVE | 1
RIGHT UNKNOWN REJUVENATE | 1
UNKNOWN - RIGHT REJUVENATE/ABG II HIP | 1
UNKNOWN PRODUCT LEFT HIP REJUVENATE | 1
UNKNOWN PRODUCT LEFT STRYKER REJUVENATE HIP | 1
UNKNOWN PRODUCT LEFT STRYKER REJUVENATE MODULAR HI | 1
UNKNOWN PRODUCT RIGHT REJUVENATE HIP | 1
UNKNOWN PRODUCT LEFT HIP REJUVENATE MODULAR HIP ST | 1
UNKNOWN PRODUCT RIGHT HIP REJUVENATE/ABG II STRYKE | 1
UNKNOWN PRODUCT REJUVENATE LEFT HIP | 1
UNKNOWN HIP REJUVENATE | 1
REJUVENATE STRIGHT PRFIT TMZF MOD STEM SIZE 8 | 1
REJUVENATE STRGHT PFIT TMZF MOD STEM SIZE 8 | 1
UNKNOWN #8 REJUVENATE STEM | 1
REJUVENATE MODULAR SIZE 8 STEM | 1
UNKNOWN PRODUCT REJUVENATE | 1
STRYKER REJUVENATE MODULAR FEMORAL STEM | 1
UNK REJUVENATE STEM - HIP RIGHT | 1
REJUVENATE STRIGTH PRFIT TMZF MOD STEM SIZE 8 | 1
UNK REJUVENATE 34 NECK | 1
UNK SIZE 7 REJUVENATE STEM | 1
REJUVENATE NECK 30MM | 1
REJUVENATE STRIGHT PRFIT TMZF MOD STEM SIZE 9 | 1
OMNIFIT EON 132 | 1
C-TAPER COCR LFIT HEAD 32MM/0 | 1
REJUVENATE STEM INSERTER | 1
UNKNOWN PRODUCT RIGHT HIP REJUVENATE/ABGII | 1
REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE 6 | 1
ACCOLADE PLUS TMZF HIP STEM #3 | 1
UNK REJUVENATE 127, 38 NECK | 1
UNKNOWN REJUVENATE HIP COMPONENTS | 1
UNKNOWN - LEFT REJUVENATE/ABG II HIP | 1
UNKNOWN RIGHT HIP-REJUVENATE/ABGII | 1
UNKNOWN LEFT HIP - REJUVENATE/ABGII | 1
V40 COCR LFIT HEAD 36MM/+5 | 1
TRIDENT 0° X3 INSERT 36MM ID | 1
UNK REJUVENATE MODULAR HIP STEM (LEFT) | 1
UNKNOWN RIGHT HIP/REJUVENATE IMPLANT | 1
UNKNOWN SIZE 7 REJUVENATE STEM | 1
UNKNOWN RIGHT REJUVENATE #9 HIP | 1
UNKNOWN LEFT REJUVENATE/ABG II MODULAR HIP | 1
STRYKER | 1
UNKNOWN LEFT HIP REJUVENATE/ ABGII | 1
UNKNOWN RIGHT HIP REJUVENATE/ABGII | 1
UNKNOWN REJUVENATE #10 STEM | 1
REJUVENATE LEFT HIP | 1
UNKNOWN TRIDENT HEMISPHERICAL HA CLUSTER SHELL | 1
UNKNOWN REJUVENATE MODULAR HIP | 1
ABG II MODULAR V 40 SHORT NECK | 1
TRIDENT ALUMINA INSERT | 1
UNKNOWN RIGHT HIP REJUVENATE MODULAR STEM | 1
REJUVENATE SPT | 1
UNKNOWN REJUVENATE FEMORAL IMPLANT | 1
REJUVENATE TOTAL HIP | 1
UNKNOWN_CORK_PRODUCT | 1
UNK LEFT HIP REJUVENATE | 1
REJUVENATE STEM #10 | 1
UNKNOWN PRODUCT RIGHT REJUVENATE/ABG II (ONE OF TH | 1
REJUVENATE STRGHT PRFIT TMZF MOD STEM SIZE | 1
UNKNOWN REJUVENATE/ABGII HIP | 1
UNKNOWN LEFT HIP REJUVENATE STEM | 1

## Case Metadata
Case Name | MDL Number | MDL Filing Date | Court | Settlement Amount | MAUDE Start Date | MAUDE End Date | Total Records | Date Range
Stryker Rejuvenate Hip Implant | 2441 | 2012-12-20 | D. Minnesota | $1.4+ billion (2014) | 2012-06-20 | 2012-12-20 | 768 | 2012-06-27 00:00:00 to 2012-12-19 00:00:00