## All Reports
report_number | date_received | date_of_event | event_type | report_source_code | brand_name | generic_name | manufacturer_d_name | model_number | catalog_number | device_class | event_description | manufacturer_narrative | adverse_event_flag | product_problem_flag | device_operator | manufacturer_g1_name | year_month
1822565-2009-01143 | 2009-11-12 | 2009-08-11 | I, n, j, u, r, y | Manufacturer report | HARRIS/GALANTE POROUS TOTAL HIP ACETABULAR COMPONENT | HIP PROSTHESIS | ZIMMER, INC. | NA | 00652204600 |  | IT IS REPORTED THAT THE DEVICE WAS IMPLANTED IN 1989, AND REVISED IN 2009, DUE TO THE ACETABULAR SHELL FRACTURING. | EVALUATION SUMMARY: THE DEVICES WERE IN VIVO FOR APPROXIMATELY 19 YEARS AND 10 MONTHS. SEM ANALYSIS OF THE SHELL FRACTURE SURFACES WAS CONDUCTED ON THE LARGE PIECE AND ONE OF THE SMALL PIECES. FRACTURE ON THE SMALL PIECE APPEARED MECHANICALLY DEFORMED IN THE POST FRACTURE CONDITION. TYPICAL FRACTURE SURFACE FEATURES ARE SHOW IN THE ATTACHED MICROGRAPHS. FATIGUE STRIATIONS AND CRYSTALLOGRAPHIC FRACTURE FEATURES WERE OBSERVED. THE SHELL FRACTURE APPEARED TO HAVE OCCURRED BY FATIGUE. EDS ANALYSIS OF THE SHELL MATERIAL SHOWED A TITANIUM PEAK IN SPECTRUM TYPICAL OF CPTI (COMMERCIALLY PURE TITANIUM). THE PRESENCE OF A BONE GRAFT SEEMS TO INDICATE THE PATIENT'S BONE QUALITY WAS POOR WHICH COULD HAVE LEAD TO INADEQUATE BONE SUPPORT FOR THE SHELL AND/OR THE POSSIBLE COLLAPSE OF THE GRAFT. THE WORN SCREW HEAD SEEMS TO INDICATE THAT THE SHELL AND MALLEOLAR SCREW SOMEHOW CAME INTO PROLONGED CONTACT, POSSIBLY AS A RESULT OF POOR BONE QUALITY ALLOWING COMPONENT MIGRATION. THE PROLONGED CONTACT AND INADEQUATE FIXATION COULD LEAD TO WEAR AND EVENTUAL FRACTURE OF THE SHELL, BUT THIS CAN NOT BE DETERMINED WITHOUT ALL FRAGMENTS OF THE SHELL OR CONCLUSIVE EVIDENCE OF SHELL-SCREW CONTACT. BASED ON THE AVAILABLE INFORMATION A DEFINITIVE ROOT CAUSE CAN NOT BE ASCERTAINED AT THIS TIME. EVALUATION: DEVICE HISTORY RECORDS INDICATE ALL COMPONENTS WERE MANUFACTURED AND INSPECTED TO SPECIFICATION. SCANNING ELECTRON MICROSCOPY (SEM) ANALYSIS / ENERGY DISPERSIVE SPECTROSCOPY (EDS) ANALYSIS WAS PERFORMED. NO MANUFACTURING ABNORMALITIES COULD BE DETECTED BY EITHER METHOD. | Y | N | LAY USER/PATIENT |  | 2009-11
9613350-2009-00643 | 2009-11-13 | 2009-10-27 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 58/52 CODE R | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.058 |  | IT WAS REPORTED THAT THE REVISION SURGERY OCCURRED DUE TO PAIN. | NO PRODUCT WAS RETURNED FOR EVAL. REVIEW OF THE DEVICE HISTORY RECORDS WAS ALSO NOT POSSIBLE AS THE ITEM NUMBERS AND/OR LOT NUMBERS WERE NOT PROVIDED. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF A DEVICE DEFICIENCY CAUSING OR CONTRIBUTING TO THE REPORTED PROBLEM. BASED ON THE AVAILABLE INFO, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL INFO BE RECEIVED THAT CHANGES THIS CONCLUSION, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-11
9613350-2009-00639 | 2009-11-13 | 2009-10-07 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 58/52 CODE R | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.058 |  | IT WAS REPORTED, THAT DUE TO A PT REQUEST A REVISION SURGERY WAS PERFORMED. NO APPARENT BONE INGROWTH/ONGROWTH ON THE CUP. | NO PRODUCT WAS RETURNED FOR EVAL. REVIEW OF THE DEVICE HISTORY RECORDS WAS ALSO NOT POSSIBLE AS THE ITEM NUMBERS AND/OR LOT NUMBERS WERE NOT PROVIDED. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF A DEVICE DEFICIENCY CAUSING OR CONTRIBUTING TO THE REPORTED PROBLEM. BASED ON THE AVAILABLE INFO, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL INFO BE RECEIVED THAT CHANGES THIS CONCLUSION, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-11
9613350-2009-00640 | 2009-11-13 | 2009-09-30 | I, n, j, u, r, y | Manufacturer report | DUROM FEMORAL COMPONENT 48 CODE N | DUROM HIP RESURFACING SYSTEM, FEMORAL COMPONE | ZIMMER GMBH |  | 01.00211.148 |  | IT WAS REPORTED THAT THE REVISION SURGERY OCCURRED DUE TO BREAKAGE OF THE PIN BECAUSE OF LOOSENING. IT IS ALSO REPORTED THAT THE PT HAS NECROSIS IN HIS FEMUR. | THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. | Y | N | LAY USER/PATIENT |  | 2009-11
9613350-2009-00633 | 2009-11-13 | 2009-10-06 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 58/52 R | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.158 |  | IT WAS REPORTED THAT PT COMPLAINED OF HIP WEAKNESS, PAIN AND SQUEAKING WITH AMBULATION. PT HAD 58MM OD/52MM ID DUROM ACETABULAR SHELL AND COMPETITOR FEMORAL STEM WITH A COMPETITOR 53MM FEMORAL HEAD. ONCE THE HIP JOINT WAS EXPOSED, THERE WAS MILD LEVELS OF METALOSIS. THE DUROM SHELL WAS THEN REMOVED WITH NO EVIDENCE OF BONE ONGROWTH. THE DUROM SHELL WAS REVISED TO A TM ACETABULAR SHELL WITH 40MM LONGEVITY LINER AND NEW COMPETITOR 40MM FEMORAL HEAD WAS IMPLANTED. | ZIMMER RECEIVED THE DUROM 58/52 CUP AND THE 53MM DIAMETER COMPETITOR FEMORAL HEAD. THE COMBINATION OF THE COMPETITOR 53MM HEAD WITH THE DUROM CUP IS NOT CONSISTENT WITH THE MFR LABELING FOR THE DEVICE. THE MISMATCH OF LARGER DIAMETER OF THE HEAD THAN THE INNER DIAMETER OF THE CUP MAY BE A CONTRIBUTOR TO THE METALOSIS AND OTHER SYMPTOMS OBSERVED. THE ZIMMER DEVICE WAS MFG TO SPECIFICATION AND THERE IS NO INDICATION THAT THE ZIMMER DEVICE FAILED TO PERFORM AS INTENDED OR CONTRIBUTED TO THE OUTCOME OBSERVED. THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED AT THIS TIME AND ZIMMER CONSIDERS THIS CASE CLOSED. THE COMPETITOR DEVICE WAS FORWARDED TO THIS FIRM. | Y | N | LAY USER/PATIENT |  | 2009-11
9613350-2009-00642 | 2009-11-13 | 2009-10-24 | I, n, j, u, r, y | Manufacturer report | DUROM FEMORAL COMPONENT 44 CODE J | DUROM HIP RESURFACING SYSTEM, FEMORAL COMPONENT | ZIMMER GMBH |  | 01.00211.144 |  | IT WAS REPORTED, THAT THE REVISION SURGERY OCCURRED DUE TO PAIN. | THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. | Y | N | LAY USER/PATIENT |  | 2009-11
9613350-2009-00647 | 2009-11-13 | 2009-06-08 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 52/46 L | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.152 |  | IT WAS REPORTED BY PT'S ATTORNEY THAT PT UNDERWENT A LEFT HIP ARTHROPLASTY IN 2008. POST-OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. A REVISION TOOK PLACE IN 2009. | NO PRODUCT WAS RETURNED FOR EVAL. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF A DEVICE DEFICIENCY CAUSING OR CONTRIBUTING TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL INFO BE RECEIVED THAT CHANGES THIS CONCLUSION, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-11
9613350-2009-00641 | 2009-11-13 | 2009-10-19 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 48/42 CODE H | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.048 |  | IT WAS REPORTED THAT A REVISION SURGERY OCCURRED DUE TO PAIN. DURING THE SURGERY, A SMALL FLUID FILLED CYST WAS FOUND AND THERE WAS NO BONE INGROWTH. | NO PRODUCT WAS RETURNED FOR EVAL. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF A DEVICE DEFICIENCY CAUSING OR CONTRIBUTING TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL INFO BE RECEIVED THAT CHANGES THIS CONCLUSION, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. . | Y | N | LAY USER/PATIENT |  | 2009-11
9613350-2009-00649 | 2009-11-17 | 2009-10-09 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 50/44 J | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.150 |  | IT WAS REPORTED BY PATIENT. THAT PATIENT UNDERWENT HIP ARTHROPLASTY IN 2007. POST-OP, SHE EXPERIENCED PAIN AND WAS REVISED IN 2009. | THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME, AND ZIMMER CONSIDERS THIS CASE CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-11
9613350-2009-00650 | 2009-11-17 | 2009-11-05 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 46/40 F | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.146 |  | IT WAS REPORTED BY PATIENT, THAT PATIENT UNDERWENT A RIGHT HIP ARTHROPLASTY ON IN 2008. POST-OP, SHE EXPERIENCES PAIN DUE TO POSSIBLE ACETABULAR LOOSENING, AND A REVISION IS SCHEDULED IN 2009. | INFORMATION WAS FORWARDED FROM ZIMMER WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME, AND ZIMMER CONSIDERS THIS CASE CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-11
9613350-2009-00648 | 2009-11-17 | 2009-11-05 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 60/54 T | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.160 |  | IT WAS REPORTED BY PATIENT'S ATTORNEY THAT PATIENT UNDERWENT A HIP ARTHROPLASTY IN 2007. POST-OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. A REVISION HAD BEEN SCHEDULED FOR 2010. | THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME, AND ZIMMER CONSIDERS THIS CASE CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-11
9613350-2009-00654 | 2009-11-17 | 2009-11-11 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 58/52 R | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.158 |  | IT WAS REPORTED THAT PATIENT CLAIMED ABOUT PERSISTENT PAIN AND STIFFNESS. REMOVED DUROM COMPONENTS. | THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME, AND ZIMMER CONSIDERS THIS CASE CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-11
9613350-2009-00655 | 2009-11-17 | 2009-08-11 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 52/46 L | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.152 |  | IT WAS REPORTED BY PATIENT, THAT PATIENT UNDERWENT A LEFT HIP ARTHROPLASTY IN 2008. POST-OP, SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING, AND WAS REVISED FOUR MONTHS LATER. | INFORMATION WAS FORWARDED FROM ZIMMER WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME, AND ZIMMER CONSIDERS THIS CASE CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-11
9613350-2009-00652 | 2009-11-17 | 2009-10-27 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 56/50 P | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.156 |  | IT WAS REPORTED THAT PATIENT UNDERWENT A RIGHT HIP ARTHROPLASTY IN 2007. POST-OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING, AND WAS REVISED IN 2009. | INFORMATION WAS FORWARDED FROM ZIMMER WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME, AND ZIMMER CONSIDERS THIS CASE CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-11
1822565-2009-01158 | 2009-11-24 | 2009-09-30 | O, t, h, e, r | Manufacturer report | TRABECULAR METAL MODULAR ACETABULAR SYSTEM SHELL WITH CLUSTER HOLES POROUS | HIP PROSTHESIS | ZIMMER, INC. | NA | 00620205222 |  | IT IS REPORTED THAT THE TM CUP WAS POSITIONED IN THE WOUND, BUT NOT IMPACTED. | EVALUATION SUMMARY: ACCORDING TO THE SUBMITTED INFO, THERE WAS SOME ISSUE WITH EITHER THE THREADS OF EITHER THE POLAR HOLE OF THE SHELL OR THE INSERTER. THE POSSIBILITY OF THREAD DAMAGE BETWEEN OPENING THE PACKAGE AND INSERTION CANNOT BE DETERMINED SINCE THE PRODUCT WAS NOT RETURNED. THERE IS ALSO NO MENTION OF THE USE OF AN IMPACTION ADAPTER AS PRESCRIBED IN THE ETCHING ALONG THE SIDE OF THE HYBRID OFFSET SHELL INSERTER AND IT ITS INSTRUCTIONS FOR USE. BECAUSE USE OF THE ADAPTER PREVENTS FULL LOADING OF THE THREADS DURING IMPACTION, IF USED WITHOUT AN ADAPTER, IT MAY BE POSSIBLE TO EFFECTIVELY LOCK THESE TWO PARTS TOGETHER THROUGH MUTUAL THREAD DAMAGE DURING IMPACTION. WITHOUT MORE INFO, THE PROBABLE CAUSE CANNOT BE DETERMINED. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADD'L SUBSTANTIVE INFO BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | N | Y | HEALTH PROFESSIONAL |  | 2009-11
1822565-2009-01169 | 2009-12-01 | 2009-11-02 | I, n, j, u, r, y | Manufacturer report | TRILOGY ACETABULAR SYSTEM LONGEVITY CROSSLINKED POLYETHYLENE LINER | HIP PROSTHESIS | ZIMMER, INC. | NA | 00630505636 |  | IT IS REPORTED THAT THE DEVICE WAS IMPLANTED IN 2009 AND REVISED NINE MONTHS LATER. INTRAOPERATIVELY, BEFORE THE LINER WAS REMOVED, THE PHYSICIAN NOTICED A CRACK IN THE SUPERIOR MEDIAL PORTION OF THE LINER. THE PATIENT HAD A HISTORY OF DISCOLORATION. | EVALUATION SUMMARY: NO PRODUCT WAS RETURNED FOR EVALUATION. ALSO, NO X-RAYS AND/OR OPERATIVE NOTES WERE RETURNED FOR REVIEW. LINER FRACTURE COULD BE DUE TO (BUT NOT LIMITED TO) ONE OR MORE OF THE FOLLOWING: PATIENT'S WEIGHT, PATIENT'S ACTIVITY, COMPONENT MALALIGNMENT, HEAD-NECK IMPINGEMENT, OR NON-RECOMMENDED HEAD/LINER/STEM/SHELL COMBINATIONS. THE PATIENT'S HEIGHT, WEIGHT, ACTIVITY LEVEL, AND AGE ARE UNKNOWN. THE PATIENT HAD A HISTORY OF DISLOCATION AND IT IS POSSIBLE LINER FRACTURE/DAMAGE OCCURRED DURING REDUCTION OF THE JOINT. BASED ON THE AVAILABLE INFORMATION, A DEFINITIVE ROOT CAUSE CAN NOT BE ASCERTAINED AT THIS TIME. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
9613350-2009-00664 | 2009-12-02 | 2009-11-24 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 58/52 R | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.158 |  | IT WAS REPORTED BY PATIENT THAT SHE UNDERWENT LEFT THA IN 2007. SINCE IMPLANT, SHE HAS EXPERIENCED GROIN PAIN AND LIMITATION OF LEG MOVEMENT. HER NOW TREATING PHYSICIAN HAS ADVISED THAT THE CUP AND STEM ARE LOOSE AND THAT SHE NEEDS A REVISION PROCEDURE, WHICH IS NOT YET SCHEDULED PENDING OUTCOME OF PELVIC MRI AND BIOPSY OF LEG TUMORS. | INFORMATION WAS FORWARDED FROM ZIMMER WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER CONSIDERS THIS CASE CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
9613350-2009-00665 | 2009-12-02 | 2009-07-01 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 58/52 R | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.158 |  | IT WAS REPORTED THAT ACETABULAR IMPLANT LOOSENING WHICH RESULTED IN COMPLETE REVISION OF THA. | INFORMATION WAS FORWARDED FROM ZIMMER, WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER CONSIDERS THIS CASE CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
9613350-2009-00668 | 2009-12-02 | 2009-11-23 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 48/42 H | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.148 |  | IT WAS REPORTED BY THE PATIENT THAT SHE UNDERWENT RIGHT TOTAL HIP REPLACEMENT IN 2006. PATIENT STATES THAT SHE HAS EXPERIENCED PAIN SINCE THE TIME OF IMPLANT UNTIL REVISION IN 2009 FOR LOOSENING OF THE ACETABULAR CUP. | THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER CONSIDERS THIS CASE CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
9613350-2009-00669 | 2009-12-02 | 2009-02-12 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 52/46 L | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.152 |  | IT WAS REPORTED THAT DUROM ACETABULAR IMPLANT LOOSENED WHICH RESULTED IN COMPLETE REVISION IN 2009. | INFORMATION WAS FORWARDED FROM ZIMMER WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER CONSIDERS THIS CASE CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
9613350-2009-00670 | 2009-12-02 | 2009-08-01 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 56/50 P | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.156 |  | IT WAS REPORTED THAT DUROM ACETABULAR IMPLANT LOOSENED WHICH RESULTED IN COMPLETE REVISION IN 2009. | INFORMATION WAS FORWARDED FROM ZIMMER WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER CONSIDERS THIS CASE CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
9613350-2009-00651 | 2009-12-02 | 2009-03-30 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.154 |  | IT WAS REPORTED BY PATIENT THAT PATIENT UNDERWENT A LEFT HIP ARTHROPLASTY IN 2008. POST-OP, SHE EXPERIENCED PAIN AND WAS REVISED IN 2009. | NO PRODUCT WAS RETURNED FOR EVALUATION. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF A DEVICE DEFICIENCY CAUSING OR CONTRIBUTING TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL INFORMATION BE RECEIVED THAT CHANGES THIS CONCLUSION, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
9613350-2009-00666 | 2009-12-02 | 2009-10-29 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 48/42 H | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.148 |  | IT WAS REPORTED BY THE PATIENT THAT HE UNDERWENT THA IN 2008. POST-OP, THE ACETABULAR AND FEMORAL COMPONENTS LOOSENED AND PATIENT WAS REVISED IN 2009. | THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER CONSIDERS THIS CASE CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
1822565-2009-01181 | 2009-12-03 | 2009-10-13 | I, n, j, u, r, y | Manufacturer report | VERSYS HIP SYSTEM FEMORAL STEM POROUS FULLCOAT BOWED NECK TAPER REVISION OFFSET | HIP PROSTHESIS | ZIMMER, INC. | NA | 00784301682 |  | IT IS REPORTED THAT THE DEVICE WAS IMPLANTED IN 2008, AND REVISED IN 2009, DUE TO HETEROTOPIC OSSIFICATION. | EVALUATION SUMMARY: NO PRODUCT FOR EVAL. ALSO, NO X-RAYS AND/OR OPERATIVE NOTES WERE RETURNED FOR REVIEW. THE DEVICE WAS IN VIVO FOR APPROXIMATELY 19 MONTHS. THE PT WAS REVISED DUE TO HETEROTOPIC OSSIFICATION, WHICH IS A CLINICAL CONDITION WHERE BONE MATERIAL GROWS IN SOFT TISSUES OF THE BODY. ACCORDING TO LITERATURE, THE CAUSE OF HETEROTOPIC OSSIFICATION MAY BE LINKED TO GENDER, TRAUMA, INJURY OR NEUROLOGICAL CONDITION. BASED ON THE AVAILABLE INFO, A DEFINITIVE ROOT CAUSE CANNOT BE ASCERTAINED AT THIS TIME. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADD'L SUBSTANTIVE INFO BE RECEIVED, THE COMPLAINT WILL BE REOPENED, ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
1822565-2009-01174 | 2009-12-03 | 2009-10-02 | I, n, j, u, r, y | Manufacturer report | ZMR HIP SYSTEM FEMORAL BODY REVISION NITRIDED POROUS CONE BODY | HIP PROSTHESIS | ZIMMER, INC. | NA | 00999301745 |  | IT IS REPORTED THAT THE DEVICE WAS IMPLANTED IN 2009 AND THAT THE PT WAS REVISED ONE MONTH LATER DUE TO THE STEM BREAKING BELOW THE JUNCTION. THE SALES ASSOCIATE INDICATED THAT THE PT WAS IMMENSELY OVERWEIGHT. | EVALUATION SUMMARY: SEM ANALYSIS INDICATES THAT THE FRACTURE COULD HAVE INITIATED WITH FRETTING FATIGUE AND FINISHING WITH FAST FINAL FRACTURE. RETURNED, THE FRACTURE STEM AND BODY SHOW SIGNIFICANT BONE ONGROWTH TO THE STEM WITH NONE PRESENT ON THE BODY. THE PACKAGE INSERT AND/OR SURGICAL TECHNIQUE CONTAINS STATEMENTS WARNING THAT DEVICE FRACTURES MAY OCCUR WHERE EXCESSIVE PT WEIGHT, EXCESSIVE PT ACTIVITY, AND/OR LACK OF PROXIMAL SUPPORT ARE INVOLVED. THE SALES ASSOCIATE'S OBSERVATION INDICATES A LIKELIHOOD OF EXCESSIVE PT WEIGHT THOUGH THIS CANNOT BE CONCLUDED WITHOUT SPECIFIC PT BODY WEIGHT MEASUREMENTS. THE ABSENCE OF BONE ONGROWTH TO THE ZMR BODY WITH SIGNIFICANT ONGROWTH SHOWN ON THE STEM LIKELY INDICATES A LACK PR PROXIMAL SUPPORT. IN THIS CASE, IT IS POSSIBLE THAT THE COMBINATION OF THE PT'S WEIGHT AND LACK OF PROXIMAL SUPPORT CONTRIBUTED TO THE STEM FRACTURE. WITHOUT FURTHER INFO, A DEFINITIVE CAUSE CANNOT BE DETERMINED. DEVICE HISTORY RECORDS INDICATE ALL COMPONENTS WERE MANUFACTURED AND INSPECTED TO SPECIFICATION. SCANNING ELETRON MICROSCOPY (SEM) ANALYSIS/ENERGY DISPERSIVE SPECTROSCOPY (EDS) ANALYSIS WAS PERFORMED. NO MFG ABNORMALITIES COULD BE DETECTED BY EITHER METHOD. | Y | N | LAY USER/PATIENT |  | 2009-12
1822565-2009-01173 | 2009-12-03 | 2009-11-06 | I, n, j, u, r, y | Manufacturer report | TRILOGY ACETABULAR LONGEVITY CROSSLINKED POLYETHYLENE LINER | HIP PROSTHESIS | ZIMMER, INC. | NA | 00631005032 |  | IT IS REPORTED THAT THE DEVICES WERE IMPLANTED IN 2007 AND THAT THE PT WAS REVISED IN 2009 BECAUSE OF PAIN DUE TO POSSIBLE JOINT LOOSENING. | THERE WERE NO PARTS RETURNED, NO PT INFO PROVIDED, NO X-RAYS, AND NO SURGICAL NOTES. DUE TO LACK OF INFO, IT IS NOT POSSIBLE TO DETERMINE THE PROBABLE CAUSE OF THE JOINT LOOSENING. : EVALUATION: NO PRODUCTS WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFO BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
1822565-2009-01180 | 2009-12-03 | 2009-11-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRABECULAR METAL MODULAR ACETABULAR SYSTEM SHELL WITH CLUSTER HOLES POROUS | HIP PROSTHESIS | ZIMMER, INC. | NA | 00620205422 |  | IT IS REPORTED THAT DURING A TOTAL HIP ARTHROPLASTY, THE SCREW WENT COMPLETELY THROUGH THE SHELL. POST-OP X-RAYS CONFIRMED THE HEAD OF THE SCREW TO BE ON THE BACK SIDE OF THE SHELL. NO UNUSUAL AMOUNT OF FORCE WAS USED TO IMPLANT THE SCREW. TWO MORE SCREWS WERE IMPLANTED WITHOUT A PROBLEM USING THE SAME TECHNIQUE. | (B)(4). EVAL SUMMARY: THE TM MODULAR ACETABULAR SHELL AND TRILOGY BONE SCREW WERE NOT AVAILABLE FOR ANALYSIS AND THEIR CONDITIONS ARE UNK. THE SURGICAL NOTES ARE UNAVAILABLE, AND IT IS UNK IF THE SCREW WAS INSERTED ACCORDING TO THE SURGICAL TECHNIQUE. THE INSTRUMENTS USED ARE ALSO UNK. BASED ON THE AVAILABLE INFO AND OBSERVATIONS, SCREW PULL-THROUGH THE ACETABULAR SHELL HOLE MAY HAVE BEEN DUE TO ONE OR A COMBINATION OF MECHANISMS. THE SCREW MAY HAVE BEEN INSERTED THROUGH THE SHELL AT AN EXTREME ANGLE. THIS MAY HAVE BEEN POSSIBLE IF THE PILOT HOLE WAS DRILLED WITHOUT THE USE OF A DRILL GUIDE, OR IF THE DRILL GUIDE WAS NOT INSERTED INTO THE SCREW HOLE, WHICH WOULD ALLOW FOR EXTREME ANGULATION. ALTERNATIVELY, EXCESSIVE TORQUE MAY HAVE BEEN APPLIED WHEN ATTEMPTING TO FIXATE THE SCREW. HOWEVER, NO DEFINITIVE CAUSE ANALYSIS CAN BE COMPLETED WITH CERTAINTY. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADD'L SUBSTANTIVE INFO BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | N | Y | HEALTH PROFESSIONAL |  | 2009-12
9613350-2009-00653 | 2009-12-04 | 2009-10-29 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 48/42 CODE H | DUROM ACETABULAR COMPONENT AND METASUL LDH | ZIMMER GMBH | N/A | 01.00214.048 |  | IT HAS NOW BEEN REPORTED THAT A PRODUCT LIABILITY CLAIM WAS RAISED. THE PATIENT WAS IMPLANTED A DUROM ACETABULAR COMPONENT 48/42 CODE H ON THE LEFT SIDE ON (B)(6) 2008. THE PATIENT WAS REVISED ON (B)(6) 2009 DUE TO PAIN AND FLUID COLLECTION. | THIS CASE WAS REOPENED ON (B)(6) 2017 TO ENTER ADDITIONAL INFORMATION WHICH WAS RECEIVED ON (B)(6) 2017. SINCE THIS CASE IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTIVE ACTION WHICH WAS REPORTED TO THE FDA IN JULY 2008 AS CORRECTION Z-2415/2426-2008, THERE WILL BE NO FURTHER INVESTIGATION AND ZIMMER (B)(4) WILL CLOSE THIS CASE ONCE AGAIN. ZIMMER¿S REFERENCE NUMBER OF THIS FILE IS (B)(4). | Y | N | LAY USER/PATIENT |  | 2009-12
1822565-2009-01187 | 2009-12-09 | 2009-11-04 | I, n, j, u, r, y | Manufacturer report | ZMR HIP SYSTEM FEMORAL STEM REVISION NITRIDED POROUS | HIP PROSTHESIS | ZIMMER, INC. | NA | 00998116532 |  | IT IS REPORTED THAT THE DEVICE WAS IMPLANTED IN 2007 AND THAT THE PATIENT WAS REVISED IN 2009 DUE TO FEMORAL PAIN. THE PATIENT DENIES FALLING, BUT STATED TO DOCTORS SHE STUMBLED A FEW MONTHS AGO. BEFORE REVISION SURGERY TO REMOVE IMPLANTS, THERE APPEARED TO BE AN IMPLANT THAT HAD BROKEN OR COME APART AT JUNCTION. AFTER REMOVAL, ONE COULD SEE WHERE IT HAD BROKEN. | EVALUATION SUMMARY: SEM ANALYSIS INDICATES THAT THE FRACTURE OCCURRED BY FATIGUE, WITH DEBRIS PRESENT IN THE REGION WHERE THE FRACTURE ORIGINATED. AS RETURNED, THE FRACTURED STEM AND BODY SHOW SIGNIFICANT BONE ONGROWTH TO THE STEM WITH VERY LITTLE PRESENT ON THE BODY. THE PACKAGE INSERT AND/OR SURGICAL TECHNIQUE CONTAINS STATEMENTS WARNING THAT DEVICE FRACTURES MAY OCCUR WHERE EXCESSIVE PATIENT'S WEIGHT, EXCESSIVE PATIENT'S ACTIVITY, AND/OR LACK OF PROXIMAL SUPPORT ARE INVOLVED. NO INFORMATION IS INCLUDED IN THE ALLEGED COMPLAINT RELATING TO THE PATIENT'S BUILD, HEIGHT, WEIGHT, OR LEVEL OF ACTIVITY. THE ABSENCE OF BONE ONGROWTH TO THE ZMR BODY WITH SIGNIFICANT ONGROWTH SHOWN ON THE STEM SUGGESTS THE POSSIBILITY OF A LACK OF PROXIMAL SUPPORT. INSUFFICIENT PROXIMAL SUPPORT OF THE STEM BODY MAY BE A CONTRIBUTING FACTOR. HOWEVER, IT IS NOT POSSIBLE TO DETERMINE THE ROOT CAUSE OF THE STEM FRACTURE WITH THE INFORMATION PRESENT. EVALUATION: DEVICE HISTORY RECORDS INDICATE ALL COMPONENTS WERE MANUFACTURED AND INSPECTED TO SPECIFICATION. SCANNING ELECTRON MICROSCOPY (SEM) ANALYSIS / ENERGY DISPERSIVE SPECTROSCOPY (EDS) ANALYSIS WAS PERFORMED. NO MANUFACTURING ABNORMALITIES COULD BE DETECTED BY EITHER METHOD. | Y | N | LAY USER/PATIENT |  | 2009-12
1822565-2009-01185 | 2009-12-09 | 2009-11-05 | I, n, j, u, r, y | Manufacturer report | ZIMMER M/L TAPER HIP PROSTHESIS WITH KINECTIV TECHNOLOGY MODULAR FEMORAL STEM PR | HIP PROSTHESIS | ZIMMER, INC. | NA | 00771301200 |  | IT IS REPORTED THAT THE DEVICES WERE IMPLANTED IN 2009 AND THAT THE PT WAS REVISED SIX MONTHS LATER DUE TO INFECTION. THE NECK WOULD NOT DISSOCIATE FROM THE STEM DURING A REVISION SURGERY. | EVAL SUMMARY: THE RECEIVED DEVICES WERE REVIEWED. THE SURGICAL NOTES ARE UNAVAILABLE, AND IT IS UNK WHICH INSTRUMENTS WERE USED WHILE ATTEMPTING TO SEPARATE THE NECK FROM THE STEM BODY. UPON REVIEW, THE KINECTIV NECK AND M/L TAPER WITH KINECTIV TECHNOLOGY STEM WERE DESIGNED TO REMAIN ASSEMBLED WHEN TENSILE FORCE UP TO 500N IS APPLIED TO SEPARATE THE COMPONENTS. THIS IS A NECESSARY DESIGN REQUIREMENT TO PREVENT DISASSEMBLY OF THE COMPONENTS IN-VIVO. HENCE, IT IS MOST LIKELY THAT THE IMPACT FORCE APPLIED VIA THE NECK EXTRACTION INSTRUMENTS WAS IN THE SAME DIRECTION AS THE AXIS OF THE BONE, AND RESULTED IN THE KINECTIV NECK AND M/L TAPER ASSEMBLY TO BE REMOVED RATHER THAN THE KINECTIV NECK DISASSEMBLING FROM THE IMPLANTED STEM. HOWEVER, THE EXACT CAUSE OF THE CURRENT SITUATION CANNOT BE CONCLUSIVELY DETERMINED. DEVICE HISTORY RECORDS INDICATE ALL COMPONENTS WERE MANUFACTURED AND INSPECTED TO SPECIFICATION. SCANNING ELECTRON MICROSCOPY (SEM) ANALYSIS/ENERGY DISPERSIVE SPECTROSCOPY (EDS) ANALYSIS WAS PERFORMED. NO MANUFACTURING ABNORMALITIES COULD BE DEFECTED BY EITHER METHOD. | Y | N | LAY USER/PATIENT |  | 2009-12
1822565-2009-01192 | 2009-12-10 | 2009-10-29 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRABECULAR METAL MODULAR ACETABULAR SYSTEM SHELL WITH CLUSTER HOLES POROUS | HIP PROSTHESIS | ZIMMER, INC. | NA | 00620205622 |  | IT IS REPORTED THAT DURING SURGERY ON (B) (6) 2009, THE SURGEON REAMED UP TO SIZE 50. ATTEMPTED TO IMPACT 56 TM SHELL WHICH SEATED PROUD. MADE SEVERAL ATTEMPTS TO ASSURE PREPARED SOCKET WAS CLEARED OF SOFT TISSUE, OSTEOPHYTES, AND OTHER DEBRIS. MADE SECOND ATTEMPT TO IMPACT 56 CUP. VERIFIED REAMER SIZE OF 56, OPENED ANOTHER 56 TM CLUSTER SHELL WITH REPEAT EXPERIENCE. REAMED WITH 57 REAMER AND IMPACTED 58 TRILOGY FM CUP. SURGERY WAS DELAYED 30 MINUTES. | (B) (4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. | N | Y | HEALTH PROFESSIONAL |  | 2009-12
9613350-2009-00656 | 2009-12-15 | 2009-06-24 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 52/46 L | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.152 |  | IT WAS REPORTED BY PATIENT'S ATTORNEY THAT PATIENT UNDERWENT A RIGHT HIP ARTHROPLASTY IN 2006. POST-OP, SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED IN 2009. | THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER CONSIDERS THIS CASE CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
9613350-2009-00661 | 2009-12-15 | 2009-11-02 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 56/50 P | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.156 |  | IT WAS REPORTED THAT PATIENT HAD PERSISTENT HIP PAIN. SURGEON DECIDED TO REVISE THE IMPLANTS. | INFORMATION WAS FORWARDED FROM ZIMMER WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER CONSIDERS THIS CASE CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
9613350-2009-00658 | 2009-12-15 | 2009-11-19 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 56/50 P | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.156 |  | IT WAS REPORTED BY PATIENT, THAT PATIENT UNDERWENT A HIP ARTHROPLASTY IN 2007. POST-OP, HE HAS EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. A REVISION IS SCHEDULED, BUT DATE IN UNKNOWN. | THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER CONSIDERS THIS CASE CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
9613350-2009-00657 | 2009-12-15 | 2009-09-09 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 46/40 F | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.146 |  | IT WAS REPORTED BY PATIENT THAT PATIENT UNDERWENT A RIGHT HIP ARTHROPLASTY IN 2008. POST-OP, SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION WAS PERFORMED IN 2009. | INFORMATION WAS FORWARDED FROM ZIMMER WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER CONSIDERS THIS CASE CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
2648920-2009-00065 | 2009-12-16 | 2009-11-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY HIP SYSTEM SELF TAPPING BONE SCREW | HIP PROSTHESIS | ZIMMER, MFG. B.V. | NA | 00625006515 |  | IT IS REPORTED THAT DURING SURGERY ON (B)(6) 2009, THE SCREW PASSED THROUGH THE SCREW HOLE TO STAY AT THE ABDOMINAL CAVITY. THE SURGEON DOES NOT INTEND TO REMOVE THE SCREW. | EVALUATION SUMMARY: THE RECEIVED X-RAYS WERE REVIEWED. IT WAS OBSERVED THAT A BONE SCREW IS LOCATED IN THE ABDOMINAL CAVITY. AS PER THE SURGICAL TECHNIQUE IT IS RECOMMENDED TO AVOID SCREW PLACEMENT THOUGH THE SHELL INTO THE ANTERIOR INFERIOR QUADRANT OF THE ACETABULUM TO PREVENT INJURY TO THE NEUROVASCULAR STRUCTURES. THE DEVICES WERE NOT AVAILABLE FOR ANALYSIS, AND THE DEVICE CONDITIONS ARE UNKNOWN. THE SURGERY NOTES ARE UNAVAILABLE, AND IT IS UNKNOWN IF THE SCREWS WERE IMPLANTED ACCORDING TO THE SURGICAL TECHNIQUE. THE INITIAL POSITION OF THE SCREW INTRA-OP IS UNKNOWN. THE INSTRUMENTS USED ARE ALSO UNKNOWN. BASED ON THE ABOVE INFORMATION AND OBSERVATIONS, SCREW PULL-THROUGH THE ACETABULAR SHELL HOLE MAY HAVE BEEN DUE TO ONE OR A COMBINATION OF THE FOLLOWING MECHANISMS. THE SCREW MAY HAVE BEEN ATTEMPTED TO BE INSERTED THROUGH THE SHELL AT AN EXTREME ANGLE. THIS MAY HAVE BEEN POSSIBLE IF THE PILOT HOLE WAS DRILLED WITHOUT THE USE OF A DRILL GUIDE, OR IF THE DRILL GUIDE WAS NOT INSERTED INTO THE SCREW HOLE, WHICH WOULD ALLOW FOR EXTREME ANGULATION. ALTERNATIVELY, EXCESSIVE TORQUE MAY HAVE BEEN APPLIED WHEN ATTEMPTING TO FIXATE THE SCREW. BASED ON THE AVAILABLE INFORMATION A DEFINITIVE ROOT CAUSE CANNOT BE ASCERTAINED AT THIS TIME. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS WAS ALSO NOT POSSIBLE AS THE PRODUCT AND/OR LOT NUMBERS REQUIRED FOR RETRIEVAL WERE UNAVAILABLE. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE AVAILABLE INFORMATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE RE-OPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | N | Y | HEALTH PROFESSIONAL |  | 2009-12
1822565-2009-01209 | 2009-12-16 | 2009-08-27 | I, n, j, u, r, y | Manufacturer report | ZIMMER M/L TAPER HIP PROSTHESIS FEMORAL STEM PRESS-FIT PLASMA SPRAYED | HIP PROSTHESIS | ZIMMER, INC. | NA | 00771101000 |  | IT IS REPORTED THAT THE DEVICES WERE IMPLANTED IN 2009 AND THAT THE PATIENT WAS REVISED THREE WEEKS LATER DUE TO INFECTION. | THE STERILIZATION PROCESS FOR THIS DEVICE WAS VALIDATED IN ACCORDANCE WITH FDA REGULATIONS AND ISO STANDARDS TO A STERILITY ASSURANCE LEVEL OF 1.0 X 10-6 OR BETTER. IN ADDITION, THE MANUFACTURING LOT SPECIFIED IN THIS COMPLAINT WAS PROCESSED ACCORDING TO THE VALIDATED STERILIZATION PROCESS PARAMETERS AND MET ALL THE ACCEPTANCE CRITERIA FOR STERILITY RELEASE. THEREFORE, IT IS HIGHLY UNLIKELY THAT THE SPECIFIED DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT INFECTION. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
1822565-2009-01211 | 2009-12-16 | 2009-11-19 | I, n, j, u, r, y | Manufacturer report | UNK ZIMMER CERAMIC ZIRCONIUM OXIDE FEMORAL HEAD | HIP PROSTHESIS | ZIMMER, INC. | NA | UNK |  | IT IS REPORTED BY PATIENT'S COUNSEL THAT PATIENT UNDERWENT RIGHT TOTAL HIP ARTHROPLASTY IN 2000. IN 2008, PATIENT EXPERIENCED A SUDDEN ONSET OF PAIN AND UNDERWENT REVISION IN WHICH IT WAS NOTED THAT THE FEMORAL HEAD HAD FRAGMENTED. | EVAL SUMMARY: NO PRODUCT WAS RETURNED FOR EVALUATION. NO PRE-OP OR POST-OP X-RAYS WERE RETURNED FOR REVIEW. SURGICAL NOTES WERE NOT AVAILABLE. PATIENT'S INFORMATION SUCH AS HEIGHT, WEIGHT, AND ACTIVITY LEVEL IS UNK. BASED ON THE AVAILABLE INFORMATION, A DEFINITIVE ROOT CAUSE CANNOT BE ASCERTAINED AT THIS TIME. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS WAS ALSO NOT POSSIBLE AS THE PRODUCT AND/OR LOT NUMBERS REQUIRED FOR RETRIEVAL WAS UNAVAILABLE. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON AVAILABLE INFORMATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
1822565-2009-01195 | 2009-12-16 |  | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MODULAR ACETABULAR SYSTEM LINER PROVISIONAL STANDARD FOR USE WITH SHELL PROVISIO | HIP INSTRUMENT | ZIMMER, INC. | NA | 00627005828 |  | IT IS REPORTED THAT WHEN TRYING TO INSERT THE LINER WITH SCREW, THE LINER FRACTURED. | EVAL: AS RETURNED, MATERIAL IS MISSING AROUND THE POLAR HOLE OF THE PROVISIONAL. THIS FAILURE HAS BEEN PREVIOUSLY CHARACTERIZED BY DEVELOPMENT. THIS FAILURE MAY BE A RESULT OF (BUT NOT LIMITED TO) THE FOLLOWING SCENARIOS EXCESSIVE TIGHTENING OF THE POLAR SCREW DURING ASSEMBLY, MATERIAL BECOMING BRITTLE DUE TO CLEANING/AUTOCLAVE PROCESSES, DEVICE FATIGUE DUE TO USAGE OVER TIME, ATTEMPTS TO REMOVE THE PROVISIONAL WITH THE SCREW INSTALLED, AND/OR IMPACTION DURING ASSEMBLY. THE INSTRUMENT WAS DIMENSIONALLY ANALYZED AND FOUND TO MEET PRINT SPECIFICATIONS. THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. THE PROVISIONAL HAD A POTENTIAL FIELD AGE OF APPROXIMATELY 6.5 YEARS AT THE TIME OF FAILURE. THE EXACT CAUSE OF THE FAILURE CAN NOT BE DETERMINED. | N | Y | HEALTH PROFESSIONAL |  | 2009-12
1822565-2009-01213 | 2009-12-16 | 2009-08-27 | I, n, j, u, r, y | Manufacturer report | TRABECULAR METAL MODULAR ACETABULAR SYSTEM SHELL WITH CLUSTER HOLES POROUS | HIP PROSTHESIS | ZIMMER, INC. | NA | 006202052022 |  | IT IS REPORTED THAT THE PATIENT WAS REVISED IN 2009 BECAUSE OF DISLOCATION OF CONSTRAINED LINER. IMPLANT DATE IS UNK. | THE EXACT TIME THE COMPONENTS WERE IN-VIVO IS UNK. THE MATING STEM COMPONENT AND CONDITION THEREOF IS UNK. THE SURGICAL NOTES FROM THE PRIMARY SURGERY ARE UNAVAILABLE. IT IS UNK IF THE COMPONENTS WERE IMPLANTED WITH THE CORRECT ORIENTATION AS PER THE SURGICAL TECHNIQUE. X-RAY IMAGES POST-PRIMARY SURGERY AND FROM SUCCESSIVE PATIENT VISITS ARE UNAVAILABLE. THE INSTRUMENTS USED IN THE PRIMARY SURGERY ARE ALSO UNK. PATIENT SEX, AGE, HEIGHT, WEIGHT, AND ACTIVITY LEVEL ARE UNK. THE REHABILITATION REGIME, BOTH PHYSICAL AND PHARMACOLOGICAL, AND COMPLIANCE THEREOF IS UNK. BASED ON THE ABOVE INFORMATION AND OBSERVATIONS, DISLOCATION MAY HAVE BEEN DUE TO ONE OR A COMBINATION OF THE FOLLOWING MECHANISMS: UNSATISFACTORY PLACEMENT OF THE COMPONENT PARTS (SHELL AND/OR STEM). THIS MAY HAVE LED TO IMPINGEMENT (LEVERING) OF THE STEM AGAINST THE LINER, WHICH COULD LEAD TO FEMORAL HEAD/LINER DISLOCATION; UNSATISFACTORY LINER IMPLANTATION. IF THE LINER WAS NOT SEATED OR ORIENTATED AS PER THE SURGICAL TECHNIQUE, THE MISALIGNMENT OF THE LINER MAY ALSO LEAD TO IMPINGEMENT OF THE STEM AGAINST THE LINER; EXTENT OF SOFT-TISSUE CONSTRAINT. INCORRECT SOFT TISSUE AND/OR POOR FUNCTIONAL MUSCLES AROUND THE HIP MAY NOT BE ABLE TO PROVIDE FUNCTIONAL SUPPORT TO THE JOINT; PATIENT BEHAVIOR. HOWEVER THE EXACT CAUSE OF THE CURRENT SITUATION CANNOT BE CONCLUSIVELY DETERMINED. THE DEVICE HISTORY RECORDS INDICATE THAT THE TM MODULAR CLUSTER-HOLED SHELL, TRILOGY LONGEVITY CONSTRAINED LINER, AND COCR FEMORAL HEAD WERE MANUFACTURED AND INSPECTED ACCORDING TO SPECIFICATIONS. NO MANUFACTURING LOT INFORMATION FOR THE BONE SCREWS WAS AVAILABLE FOR MANUFACTURING EVALUATION. | Y | N | LAY USER/PATIENT |  | 2009-12
1822565-2009-01198 | 2009-12-16 | 2009-11-20 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | REPLACEMENT OFFSET HEAD SEATER CUP | HIP INSTRUMENT | ZIMMER, INC. | NA | 00780401802 |  | IT IS REPORTED THAT THE PLASTIC IMPACTOR PAD BROKE WHEN IMPACTING. | EVAL RESULTS & CONCLUSIONS - AS RETURNED, TWO PEGS HAVE FRACTURED OFF THE HEADSEATER CUP. THE LOT NUMBER OF THE CUP IS UNK, SO A POTENTIAL FIELD AGE CAN NOT BE CALCULATED NOR CAN THE MANUFACTURING RECORDS BE REVIEWED. THE INSTRUMENT WAS DIMENSIONALLY ANALYZED AND FOUND TO MEET PRINT SPECIFICATIONS. THIS IS A PREVIOUSLY ADDRESSED DESIGN ISSUE WHERE A CAPA HAS BEEN GENERATED TO REDESIGN THE CAP AND MITIGATE THE RISK OF THE TABS BREAKING 00-7804-018-02 HAS BEEN OBSOLETED AND REPLACED WITH 00-7804-018-03 SINCE (B) (6) 2007. | N | Y | HEALTH PROFESSIONAL |  | 2009-12
2648920-2009-00066 | 2009-12-16 | 2009-11-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY ACETABULAR SYSTEM BONE SCREW SELF-TAPPING | HIP PROSTHESIS | ZIMMER MFG. B.V. | NA | 00625006540 |  | IT IS REPORTED THAT DURING A ROUTINE HIP REPLACEMENT, UPON INSERTING ONE OF THE SCREWS IN THE ACETABULUM IT BECAME STUCK. THE SURGEON CLAIMED THE PT HAD VERY HARD BONE. THE HEAD OF THE SCREW BROKE OFF, BUT NOT PROTRUDING THROUGH THE CUP. THE SURGEON TOOK THE X-RAY AND DECIDED TO LEAVE IT IN THE PT. THE HEAD AND PROXIMAL PORTION OF THE SCREW THAT BROKE OFF WAS REMOVED FROM THE PT. | (B) (4). EVAL SUMMARY: NO PRODUCT WAS RETURNED FOR EVAL. ALSO, NO X-RAYS AND/OR OPERATIVE NOTES WERE RETURNED FOR EVAL. BASED ON THE ABOVE INFO AND OBSERVATIONS, IT IS MOST LIKELY THAT THE SCREW FRACTURE OCCURRED DUE TO EXCESSIVE STRESS APPLIED VIA AN EXCESSIVE BENDING MOMENT, OVER-TORQUE, AND/OR POOR QUALITY OF PILOT HOLE. IF THE SCREW WAS NOT ORIENTED PERPENDICULAR TO THE SHELL HOLE, IT MAY HAVE CONTACTED THE SCREW HOLE EDGE AND MAY HAVE CAUGHT THE LOWER THREADS, AND WHEN EXCESSIVE TORQUE WAS APPLIED IT MAY HAVE CAUSED THE SCREW TO FRACTURE. BASED ON THE AVAILABLE INFO, A DEFINITIVE ROOT CAUSE CANNOT BE DETERMINED. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADD'L SUBSTANTIVE INFO BE RECEIVED, THE COMPLAINT WILL BE RE-OPENED. ZIMMER, INC CONSIDERS THE INVESTIGATION CLOSED. | N | Y | HEALTH PROFESSIONAL |  | 2009-12
1822565-2009-01212 | 2009-12-16 | 2009-11-14 | I, n, j, u, r, y | Manufacturer report | ZMR HIP SYSTEM FEMORAL STEM REVISION TAPER W/COMPRESSION NUT | HIP PROSTHESIS | ZIMMER, INC. | NA | 00998201718 |  | IT IS REPORTED THAT THE DEVICES WERE IMPLANTED IN 2007 AND REVISED IN 2009, DUE TO THE STEM BREAKING AT THE SHAFT INTERFACE AND SEPSIS. | THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. | Y | N | LAY USER/PATIENT |  | 2009-12
1822565-2009-01208 | 2009-12-16 | 2009-08-11 | I, n, j, u, r, y | Manufacturer report | TRILOGY ACETABULAR SYSTEM LONGEVITY CROSSLINKED POLYETHYLENE LINER | HIP PROSTHESIS | ZIMMER, INC. | NA | 00630505632 |  | IT IS REPORTED THAT THE DEVICE WAS IMPLANTED IN 2009 AND REVISED FIVE MONTHS LATER, DUE TO DISLOCATION. THE PATIENT HAS UNDER GONE THREE REDUCTIONS PREVIOUSLY TO THE REVISION. | NO PRODUCT WAS RETURNED FOR EVALUATION. ALSO, NO X-RAYS WERE RETURNED FOR REVIEW. THE DEVICES WERE IN VIVO FOR APPROXIMATELY 4 MONTHS. THE PATIENT'S HEIGHT, WEIGHT, AND ACTIVITY LEVEL ARE UNK. THE REHABILITATION REGIME, BOTH PHYSICAL AND PHARMACOLOGICAL AND COMPLIANCE THEREOF IS UNK. BASED ON THE AVAILABLE INFORMATION AND OBSERVATIONS, THE DISLOCATIONS MAY HAVE BEEN DUE TO ONE OR A COMBINATION OF THE FOLLOWING MECHANISMS: UNSATISFACTORY PLACEMENT OF THE COMPONENT PARTS (SHELL AND/OR STEM), EXTENT OF COMPONENT STABILITY, EXTENT OF SOFT TISSUE TENSION AND PATIENT BEHAVIOR. BASED ON THE AVAILABLE INFORMATION, A DEFINITIVE ROOT CAUSE CAN NOT BE ASCERTAINED AT THIS TIME. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
9613350-2009-00663 | 2009-12-16 | 2007-09-06 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 52/46 L | DUROM ACETABULAR COMPONENT | ZIMMER GMBH |  | 01.00214.152 |  | IT WAS REPORTED BY PATIENT'S COUNSEL THAT PATIENT UNDERWENT LEFT HIP RESURFACING IN 2007. AT TWO MONTHS POST-OP, THE PATIENT WAS REVISED DUE TO LOOSENING OF THE ACETABULAR COMPONENT. | INFORMATION WAS FORWARDED FROM ZIMMER INC WHICH MARKETS THE DEVICES IN THE U.S. NO PRODUCT WAS RETURNED FOR EVALUATION. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF A DEVICE DEFICIENCY CAUSING OR CONTRIBUTING TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL INFORMATION BE RECEIVED THAT CHANGES THIS CONCLUSION, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
1822565-2009-01196 | 2009-12-16 | 2009-10-27 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HGP II CUP FLEXIBLE DRILL FOR USE ONLY WITH HGP II CUP | HIP INSTRUMENT | ZIMMER, INC. | NA | 00661100301 |  | IT IS REPORTED THAT THE DRILL BIT FRACTURED DURING SURGERY AND REMAINS EMBEDDED IN THE PT'S BONE. | EVAL: IT IS ALLEGED THAT THE TIP OF A FLEXIBLE DRILL BROKE AND WAS LEFT IN THE PT. THE FLEXIBLE DRILL WAS NOT RETURNED FOR REVIEW. THIS FAILURE HAS BEEN PREVIOUSLY CHARACTERIZED BY DEVELOPMENT. THE FAILURE MAY BE ATTRIBUTED TO (BUT NOT LIMITED TO) ONE OF THE FOLLOWING SITUATIONS EXCESSIVE FORCE APPLIED DURING USAGE, CONTINUED OPERATION OF THE DRILL AFTER IT WAS STUCK AGAINST HARD MATERIAL, RAPID FORWARD AND REVERSAL OF DIRECTION OF ROTATION WHEN THE DEVICE IS STUCK, EXCESSIVE BENDING AROUND CORNERS, DEVICE WEAR DUE TO USAGE WITH TIME, AND/OR LOW SPEED/HIGH FOR QUE OPERATION. THE INSTRUMENT HAD A POTENTIAL FIELD AGE OF APPROXIMATELY 23 MONTHS AT THE TIME OF FAILURE. THE MANUFACTURING RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING INDICATING THAT THE INSTRUMENT WAS MANUFACTURED TO SPECIFICATION. DUE TO LACK OF SURGERY DETAILS, AN UNK USAGE HISTORY, AND TYPE OF DRIVER USED, THE EXACT CAUSE CAN NOT BE KNOWN WITH CERTAINTY. | N | Y | HEALTH PROFESSIONAL |  | 2009-12
9613350-2009-00673 | 2009-12-16 | 2009-11-20 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 60/54 T | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.160 |  | IT WAS REPORTED BY THE PATIENT THAT HE UNDERWENT TOTAL LEFT HIP ARTHROPLASTY IN 2007. POST-OP, HE EXPERIENCED PAIN AND LIMITATION IN LIFTING HIS LEFT LEG. IN 2009, IT WAS RECOMMENDED THAT HIS ACETABULAR COMPONENTS BE REVISED. THE REVISION PROCEDURE HAS NOT YET BEEN SCHEDULED. | THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAY, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER CONSIDERS THIS CASE CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
9613350-2009-00672 | 2009-12-16 | 2009-07-14 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 50/44 J | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.150 |  | IT WAS REPORTED BY PATIENT COUNSEL THAT PATIENT UNDERWENT RIGHT TOTAL HIP ARTHROPLASTY IN 2006. POST-OP, PATIENT EXPERIENCED PAIN AND WAS REVISED IN 2008, DUE TO LOOSENING OF THE ACETABULAR COMPONENT. | THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAY, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER CONSIDERS THIS CASE CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
9613350-2009-00692 | 2009-12-16 | 2009-08-25 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 50/44 J | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.150 |  | IT WAS REPORTED BY PATIENT'S COUNSEL THAT PATIENT UNDERWENT THA IN 2007 AND WAS REVISED IN 2009 FOR LOOSENING. | INFORMATION WAS FORWARDED FROM ZIMMER WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAY, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER CONSIDERS THIS CASE CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
9613350-2009-00674 | 2009-12-16 | 2009-10-09 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 58/52 R | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.158 |  | IT WAS REPORTED THAT PER CORRESPONDING CASE REPORT FORM, PATIENT'S COMPLICATION ACETABULAR IMPLANT LOOSENING WHICH RESULTED IN COMPLETE REVISION IN 2009. | NO PRODUCT WAS RETURNED FOR EVALUATION. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF A DEVICE DEFICIENCY CAUSING OR CONTRIBUTING TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL INFORMATION BE RECEIVED THAT CHANGES THIS CONCLUSION, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
9613350-2009-00676 | 2009-12-16 | 2009-11-16 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 48/42 H | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.148 |  | IT WAS REPORTED THAT THE CUP WAS LOOSE, 0% INGROWTH. EQUATORIAL FINS WERE THE ONLY PART OF CUP THAT HAD ANY FIXATION. | INFORMATION WAS FORWARDED FROM ZIMMER WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAY, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER CONSIDERS THIS CASE CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
9613350-2009-00695 | 2009-12-16 | 2009-11-30 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 48/42 H | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.148 |  | IT WAS REPORTED BY PT'S COUNSEL THAT PT UNDERWENT THA IN 2006 AND WAS REVISED IN 2008 FOR LOOSENING. | INFO WAS FORWARDED FROM ZIMMER WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN 07/2008. SHOULD ADD'L INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME, AND ZIMMER CONSIDERS THIS CASE CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
9613350-2009-00677 | 2009-12-16 | 2009-09-04 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 48/42 H | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.148 |  | IT WAS REPORTED THAT CUP WAS LOOSE, 0% INGROWTH. EQUATORIAL FINS WERE THE ONLY PART OF CUP THAT HAD ANY FIXATION. | INFORMATION WAS FORWARDED FROM ZIMMER WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAY, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER CONSIDERS THIS CASE CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
9613350-2009-00697 | 2009-12-16 | 2009-10-01 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 56/50 P | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.002143156 |  | IT WAS REPORTED BY PT'S COUNSEL THAT PT UNDERWENT LEFT THA IN 2007. PT WAS REVISED IN 2009 FOR LOOSENING OF THE ACETABULAR CUP. | INFO WAS FORWARDED FROM ZIMMER WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN 07/2008. SHOULD ADD'L INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME, AND ZIMMER CONSIDERS THIS CASE CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
9613350-2009-00698 | 2009-12-16 | 2009-03-24 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 56/50 P | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.156 |  | IT WAS REPORTED BY PT'S COUNSEL THAT PT UNDERWENT THA IN 2008 AND WAS REVISED IN 2009 FOR LOOSENING. | INFO WAS FORWARDED FROM ZIMMER WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN 07/2008. SHOULD ADD'L INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME, AND ZIMMER CONSIDERS THIS CASE CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
9613350-2009-00700 | 2009-12-16 | 2008-12-02 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 56/50 P | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.156 |  | IT WAS REPORTED BY PT'S COUNSEL THAT PT UNDERWENT LEFT HIP THA IN 2006. THE PT WAS REVISED IN 2008 FOR LOOSENING OF THE ACETABULAR CUP. | INFO WAS FORWARDED FROM ZIMMER WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN 07/2008. SHOULD ADD'L INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME, AND ZIMMER CONSIDERS THIS CASE CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
9613350-2009-00699 | 2009-12-16 | 2009-10-24 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.154 |  | IT IS REPORTED BY PT'S COUNSEL THAT PT UNDERWENT THA IN 2007 AND WAS REVISED IN 2009 FOR LOOSENING. | THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN 07/2008. SHOULD ADD'L INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME, AND ZIMMER CONSIDERS THIS CASE CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
9613350-2009-00696 | 2009-12-16 | 2009-08-01 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.154 |  | IT WAS REPORTED BY PT'S COUNSEL THAT PT UNDERWENT RIGHT THA IN 2007. PT WAS REVISED IN 2009 FOR LOOSENING OF THE ACETABULAR CUP. | THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN 07/2008. SHOULD ADD'L INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME, AND ZIMMER CONSIDERS THIS CASE CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
1822565-2009-01215 | 2009-12-17 | 2009-11-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | INTER-OP ACETABULAR SYSTEM THREADED STRAIGHT SHELL IMPACTOR ROD | HIP INSTRUMENT | ZIMMER, INC. | NA | 936600010 |  | IT IS REPORTED THAT THE SURGERY WAS DELAYED FOR 30 MINUTES DUE TO THE FACT THAT AFTER IMPACTING THE CUP INTO THE ACETABULUM, THE INSERTER WOULD NOT SCREW OFF AND DISASSOCIATE FROM THE IMPLANT. THE CUP WAS TAKEN OUT OF THE BODY AND STILL WOULD NOT SCREW OFF OF THE INSERTER. AFTER MUCH FORCE, THE CUP FINALLY CAME LOOSE. AGAIN THE CUP WAS ONLY HAND TIGHTENED ON THE INSERTER, IMPACTED INTO THE ACETABULUM AND AGAIN WOULD NOT DISASSOCIATE FROM THE IMPLANT. ANOTHER INSERTER OF THE SAME CATALOG NUMBER WAS TRIED AND THE SAME THING HAPPENED. | EVAL SUMMARY: THE MATING ACETABULAR SHELL COMPONENT IS UNK AND THE CONDITION THEREOF IS UNK. THE EXACT FIELD AGES OF THE INTER-OP SHELL ALIGNERS AND THE FREQUENCY OF THEIR USE IS UNK. THE SURGICAL NOTES ARE UNAVAILABLE. THE SURGICAL TECHNIQUE USED IS ALSO UNK. BASED ON THE ABOVE INFO AND OBSERVATIONS, THE SEIZING OF THE SHELL ALIGNERS ONTO THE SHELL MAY HAVE BEEN DUE TO ONE OR A COMBINATION OF THE FOLLOWING MECHANISMS. THE SHELL ALIGNERS MAY HAVE BEEN CROSS-THREADED. OVER-TIGHTENED, OR OVER-TORQUED ONTO THE SHELLS, WHICH MAY HAVE CAUSED THE ASSEMBLY TO SEIZE SUCCEEDING IMPACTION. ALTERNATIVELY, AS THE CONDITION OF THE MATING ACETABULAR SHELL COMPONENT IS UNK, THREAD DAMAGE OF THE POLAR HOSE OF THE MATING SHELL AFTER THE INITIAL ATTEMPT TO SEAT THE SHELL MAY HAVE INCREASED THE DIFFICULTY IN ASSEMBLY AND DISASSEMBLY OF THE COMPONENTS. HOWEVER, THE EXACT CAUSE OF THE CURRENT SITUATION CANNOT BE CONCLUSIVELY DETERMINED. BOTH INSTRUMENTS WERE RETURNED IN GOOD CONDITION WITH ONLY MINOR THREAD DEFORMATION. THE IMPACTOR RODS WERE FOUND TO BE CONFORMING TO THE APPROPRIATE THREAD AND RING FUNCTIONAL GAUGES. THE DEVICES WERE DIMENSIONALLY ANALYZED AND FOUND TO MEET PRINT SPECS WHERE COULD BE MEASURED AND THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. THE IMPACTORS HAVE POTENTIAL FIELD AGES OF LESS THAN 1.5 YEARS. | N | Y | HEALTH PROFESSIONAL |  | 2009-12
9613350-2009-00704 | 2009-12-21 | 2009-12-11 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 50/44 J | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.150 |  | IT WAS REPORTED BY THE PATIENT THAT SHE UNDERWENT RIGHT THA ON (B) (6) 2006. IN (B) (6) 2009, SHE BEGAN EXPERIENCING INTERMITTENT HIP PAIN AND IN OCTOBER OF 2009, A REVISION OF THE ACETABULAR COMPONENT DUE TO LOOSENING WAS RECOMMENDED AND SCHEDULED FOR (B) (6) 2010. | (B) (4). THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2009-12
9613350-2009-00703 | 2009-12-21 | 2009-12-11 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 52/46 L | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.152 |  | IT WAS REPORTED BY THE PATIENT THAT SHE UNDERWENT RIGHT HIP THA IN (B) (6) 2007. POST-OP, PATIENT CONTINUED TO EXPERIENCE PAIN AND LOOSENING OF THE CUP WAS CONFIRMED BY HER SURGEON. REVISION OF THE CUP WAS RECOMMENDED AND A REVISION PROCEDURE IS SCHEDULED FOR (B) (6) 2009. | (B) (4). THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2009-12
9613350-2009-00702 | 2009-12-21 | 2009-12-11 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.154 |  | IT WAS REPORTED BY PATIENT'S COUNSEL THAT PATIENT UNDERWENT LEFT THA ON (B) (6) 2007. IN (B) (6) OF 2009, HIS SURGEON RECOMMENDED A REVISION OF THE ACETABULAR COMPONENT DUE TO LOOSENING. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2009-12
9613350-2009-00632 | 2009-12-21 | 2009-10-21 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 52/46 L | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.152 |  | IT WAS REPORTED BY PATIENT THAT PATIENT UNDERWENT A HIP ARTHROPLASTY IN (B) (6) 2006. POST-OP, SHE EXPERIENCES PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. A REVISION IS SCHEDULED FOR (B) (6) 2009. | (B) (4). THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2009-12
1822565-2009-01223 | 2009-12-22 |  | I, n, j, u, r, y | Manufacturer report | ZIMMER APS NATURAL-HIP SYSTEM PROSTHESIS FEMORAL STEM POROUS | HIP PROSTHESIS | ZIMMER, INC. | NA | 00671000202 |  | IT IS REPORTED THAT AFTER INSERTION OF THE APS NATURAL STEM DURING SURGERY IN 2009, THE SURGEON RECOGNIZED A CRACK OF THE CALCAR AND USED A CABLE TO PROVIDE SUPPORT AND ALLOW FOR THE FEMUR TO HEAL. | THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. | Y | N | HEALTH PROFESSIONAL |  | 2009-12
1822565-2009-01228 | 2009-12-22 | 2009-10-29 | I, n, j, u, r, y | Manufacturer report | ZIMMER MIS HIP INSTRUMENTATION HYBRID OFFSET SHELL INSERTER | HIP INSTRUMENT | ZIMMER, INC. | NA | 00780402520 |  | IT IS REPORTED THAT THE SURGEON HAD TO EXTEND THE INCISION TO ACCOMMODATE A ZIMMER TRILOGY INTRODUCER DUE TO THE FACT THAT THE OFFSET SHELL INSERTER WOULD NOT FUNCTION PROPERLY. | EVALUATION SUMMARY: IT IS UNKNOWN WHICH MATING SHELL IMPACTION ADAPTER WAS USED DURING SURGERY. IT IS ALSO UNKNOWN WHETHER THE INSTRUCTIONS PER THE SURGICAL TECHNIQUE WERE FOLLOWED. IT IS POSSIBLE THAT THE THREADS ON THE INSERTER WERE NOT FULLY ENGAGED WITH THE CUP THREADS AT THE TIME OF ASSEMBLY. THIS FAILURE CAN BE CAUSED BY (BUT NOT LIMITED TO) ONE OR A COMBINATION OF THE FOLLOWING: THE SHELL IMPACTION ADAPTER SQUARE FEATURE NOT BEING ALIGNED PROPERLY AND ASSEMBLED FLUSH WITH THE SQUARE MATING FEATURE IN THE CONTINUUM SHELL DURING ASSEMBLY; DURING USE OF THE BALL HEX HEAD DRIVER, THE INSERTER BOLT MAY SIMPLY DURING ASSEMBLY; 2) DURING USE OF THE BALL HEX HEAD DRIVER, THE INSERTER BOLT MAY SIMPLY NOT HAVE BEEN TIGHTENED SECURELY ENOUGH DURING ASSEMBLY TO THE SHELL. HOWEVER, WITHOUT THE DEVICE, IT IS NOT POSSIBLE TO DETERMINE THE ROOT CAUSE WITH CERTAINTY AT THIS TIME. RESULTS: NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2009-12
1822565-2009-01235 | 2009-12-22 |  | I, n, j, u, r, y | Manufacturer report | UNKNOWN APS NATURAL HIP STEM | HIP PROSTHESIS | ZIMMER, INC. | NA | UNK |  | IT IS REPORTED THAT AFTER INSERTION OF THE APS NATURAL STEM, SURGEON RECOGNIZED A CRACK OF THE CALCAR AND USED A CABLE TO PROVIDE SUPPORT AND ALLOW FOR THE FEMUR TO HEAL. SURGERY DATE IS UNKNOWN. | EVALUATION SUMMARY: THE ITEM NUMBER AND LOT NUMBER WERE NOT PROVIDED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. NO X-RAYS WERE PROVIDED. IT IS POSSIBLE THAT THE INTRA-OPERATIVE FEMORAL FRACTURE MAY HAVE RESULTED FROM (BUT NOT LIMITED TO) ONE OR A COMBINATION OF THE FOLLOWING: IMPROPER RASPING TECHNIQUE; SELECTING/INSERTING A STEM TOO LARGE FOR THE FEMUR; VARIABILITY IN THE MECHANICAL PROPERTIES OF BONE; EXCESSIVE IMPACTION FORCE DURING IMPLANT INSERTION. HOWEVER, IT IS NOT POSSIBLE TO DEFINITIVELY DETERMINE THE CAUSE OF THE FEMORAL FRACTURE AT THIS TIME. EVALUATION CODES: NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS WERE ALSO NOT POSSIBLE AS THE PRODUCT AND/OR LOT NUMBERS REQUIRED FOR RETRIEVAL WERE UNAVAILABLE. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE AVAILABLE INFORMATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2009-12
1822565-2009-01234 | 2009-12-22 | 2009-11-30 | I, n, j, u, r, y | Manufacturer report | UNKNOWN APS NATURAL HIP STEM | HIP PROSTHESIS | ZIMMER, INC. | NA | UNK |  | IT IS REPORTED THAT DURING SURGERY IN 2009, AFTER INSERTION OF THE APS NATURAL STEM, THE SURGEON RECOGNIZED A SMALL CRACK OF THE CALCAR AND PLACED A CABLE AROUND THE PROXIMAL FEMUR TO PROVIDE SUPPORT AND ALLOW FOR BONE TO HEAL. | EVALUATION SUMMARY: THE ITEM NUMBER AND LOT NUMBER WERE NOT PROVIDED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. NO X-RAYS WERE PROVIDED. IT IS POSSIBLE THAT THE INTRA-OPERATIVE FEMORAL FRACTURE MAY HAVE RESULTED FROM (BUT NOT LIMITED TO) ONE OR A COMBINATION OF THE FOLLOWING: IMPROPER RASPING TECHNIQUE; SELECTING/INSERTING A STEM TOO LARGE FOR THE FEMUR; VARIABILITY IN THE MECHANICAL PROPERTIES OF BONE; EXCESSIVE IMPACTION FORCE DURING IMPLANT INSERTION. HOWEVER, IT IS NOT POSSIBLE TO DEFINITIVELY DETERMINE THE CAUSE OF THE FEMORAL FRACTURE AT THIS TIME. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS WERE ALSO NOT POSSIBLE AS THE PRODUCT AND/OR LOT NUMBERS REQUIRED FOR RETRIEVAL WERE UNAVAILABLE. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE AVAILABLE INFORMATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2009-12
1822565-2009-01233 | 2009-12-22 | 2009-11-30 | I, n, j, u, r, y | Manufacturer report | UNKNOWN APS NATURAL HIP STEM | HIP PROSTHESIS | ZIMMER, INC. | NA | UNK |  | IT IS REPORTED THAT DURING SURGERY IN 2009, AFTER INSERTION OF THE APS NATURAL STEM, THE SURGEON RECOGNIZED A SMALL CRACK OF THE CALCAR AND PLACED A CABLE AROUND THE PROXIMAL FEMUR TO PROVIDE SUPPORT AND ALLOW FOR BONE TO HEAL. | EVALUATION SUMMARY: THE ITEM NUMBER AND LOT NUMBER WERE NOT PROVIDED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. NO X-RAYS WERE PROVIDED. IT IS POSSIBLE THAT THE INTRA-OPERATIVE FEMORAL FRACTURE MAY HAVE RESULTED FROM (BUT NOT LIMITED TO) ONE OR A COMBINATION OF THE FOLLOWING: IMPROPER RASPING TECHNIQUE; SELECTING/INSERTING A STEM TOO LARGE FOR THE FEMUR; VARIABILITY IN THE MECHANICAL PROPERTIES OF BONE; EXCESSIVE IMPACTION FORCE DURING IMPLANT INSERTION. HOWEVER, IT IS NOT POSSIBLE TO DEFINITIVELY DETERMINE THE CAUSE OF THE FEMORAL FRACTURE AT THIS TIME. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS WERE ALSO NOT POSSIBLE AS THE PRODUCT AND/OR LOT NUMBERS REQUIRED FOR RETRIEVAL WERE UNAVAILABLE. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE AVAILABLE INFORMATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2009-12
1822565-2009-01232 | 2009-12-22 | 2009-11-12 | I, n, j, u, r, y | Manufacturer report | NATURAL-HIP SYSTEM 12/14 TAPER POROUS STEM COLLARLESS | HIP PROSTHESIS | ZIMMER, INC. | NA | 735402102 |  | IT IS REPORTED THAT THE DEVICES WERE IMPLANTED IN 2008 AND THAT THE PT WAS REVISED IN 2009 DUE TO LOOSENING AND DISLOCATION. | EVALUATION SUMMARY: NO PRODUCT WAS RETURNED FOR EVALUATION. ALSO, NO X-RAYS WERE RETURNED FOR REVIEW. THE REASON FOR REVISION IS UNK. SINCE THE REASON FOR THE REVISION SURGERY IS UNK, AND NO PRODUCT WAS RETURNED, THE PROBABLE CAUSE CANNOT BE DETERMINED. EVALUATION CODES: NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADD'L SUBSTANTIVE INFO BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
1822565-2009-01227 | 2009-12-22 | 2009-11-23 | I, n, j, u, r, y | Manufacturer report | ZIMMER M/L TAPER HIP PROSTHESIS WITH KINECTIV TECHNOLOGY MODULAR FEMORAL STEM PR | HIP PROSTHESIS | ZIMMER, INC. | NA | 00771300700 |  | IT IS REPORTED THAT THE DEVICE WAS IMPLANTED IN 2009 AND REVISED THE SAME DAY BECAUSE THE PT EXPERIENCED A DISLOCATED HIP. | EVALUATION SUMMARY: THE DEVICES WERE NOT RETURNED FOR EVALUATION AND PT X-RAYS WERE NOT INCLUDED. SUCH AN EARLY DISLOCATION COULD BE DUE TO (BUT NOT LIMITED TO) ONE OR MORE OF THE FOLLOWING: IMPROPER SHELL SIZE SELECTION. INCORRECT COMPONENT POSITIONING. PT ANATOMY PRONE TO DISLOCATION. IMPROPER ANATOMICAL POSITION ASSUMED BY PT. SOFT TISSUE IMBALANCE. PT FALL. ACCORDING TO THE ADVERSE EVENT REPORT SUBMITTED, THE SURGEON INDICATES THE ALLEGED EVENT WAS NOT RELATED TO THE DEVICE. HOWEVER, THE EXACT CAUSE OF DISLOCATION CANNOT BE DETERMINED WITH CERTAINTY AT THIS TIME. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADD'L SUBSTANTIVE INFO BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
1822565-2009-01224 | 2009-12-22 | 2009-11-20 | I, n, j, u, r, y | Manufacturer report | HARRIS PRECOAT PLUS WITH DISTAL PLUG FEMORAL STEM | HIP PROSTHESIS | ZIMMER, INC. | NA | 00902803000 |  | IT IS REPORTED THAT THE DEVICE WAS IMPLANTED IN 1997 AND THAT THE PT WAS REVISED IN 2009, DUE TO A FEMUR FRACTURE BELOW STEM TIP. | THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. | Y | N | LAY USER/PATIENT |  | 2009-12
1822565-2009-01222 | 2009-12-22 | 2007-09-03 | I, n, j, u, r, y | Manufacturer report | ZIMMER M/L TAPER HIP PROSTHESIS WITH KINECTIV TECHNOLOGY MODULAR FEMORAL STEM PR | HIP PROSTHESIS | ZIMMER, INC. | NA | 00771301000 |  | IT IS REPORTED THAT THE DEVICE WAS IMPLANTED IN 2007 AND THAT THE PT WAS GETTING OUT OF HER ROCKING CHAIR AND HEARD A POP. THE PT PRESENTED AT THE OFFICE THE FOLLOWING DAY AND X-RAYS CONFIRMED THAT PT HAD DISLOCATED HER HIP AND WAS REVISED THE FOLLOWING MONTH. | EVALUATION SUMMARY: THE DEVICES WERE NOT RETURNED FOR EVALUATION AND PT X-RAYS WERE NOT INCLUDED. DISLOCATION COULD BE DUE TO (BUT NOT LIMITED TO) ONE OR MORE OF THE FOLLOWING: IMPROPER SHELL SIZE SELECTION, INCORRECT COMPONENT POSITIONING, FEMORAL COMPONENT IMPINGEMENT, PT ANATOMY PRONE TO DISLOCATION, IMPROPER ANATOMICAL POSITION ASSUMED BY PT, SOFT TISSUE IMBALANCE, JOINT LAXITY; HOWEVER, THE EXACT CAUSE OF DISLOCATION CANNOT BE DETERMINED WITH CERTAINTY AT THIS TIME. ACCORDING TO THE ZIMMER HIP REGISTRY ADVERSE EVENT REPORT COMPLETED BY THE SURGEON, THE ALLEGED EVENT WAS NOT RELATED TO THE DEVICE. NO PRODUCT WAS RETURNED. DEVICE HISTORY RECORDS INDICATE ALL COMPONENTS WERE MANUFACTURED AND INSPECTED TO SPECIFICATION. NO MFG ABNORMALITIES COULD BE DETECTED. | Y | N | LAY USER/PATIENT |  | 2009-12
1822565-2009-01226 | 2009-12-22 | 2009-11-06 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EPSILON ACETABULAR SYSTEM DURASUL CONSTRAINED INSERT | HIP PROSTHESIS | ZIMMER, INC. | NA | 438032055 |  | IT IS REPORTED THAT THE SURGEON WAS UNABLE TO SEAT THE CONSTRAINING RING ON THE LINER. THE SURGEON FELT THAT THE LINER LOOKED DIFFERENT THAN PAST EPSILON DURASUL CONSTRAINED LINERS. THE REP ALSO SAID THAT THE "STICKERS" IN THE BOX DID NOT SEEM NORMAL. SURGERY WAS DELAYED BY 30 MINUTES. ZIMMER WAS MADE AWARE OF THE 30 MINUTE DELAY ON (B)(6) 2009, WHICH MADE THIS COMPLAINT REPORTABLE. | EVALUATION SUMMARY: THE DEVICE WAS UNAVAILABLE FOR ANALYSIS AND THE DEVICE CONDITION IS UNKNOWN. THE SURGICAL NOTES ARE UNAVAILABLE, AND IT IS UNKNOWN IF THE METAL LOCKING RING WAS ASSEMBLED AS PER THE SURGICAL TECHNIQUE (B)(4). THE INSTRUMENTS USED ARE ALSO UNKNOWN. BASED ON THE ABOVE INFORMATION AND OBSERVATIONS, THE DIFFICULTY ASSEMBLING THE METAL LOCKING RING TO THE LINER MAY HAVE BEEN DUE TO ONE OR A COMBINATION OF THE FOLLOWING MECHANISMS. SOFT TISSUE MAY HAVE IMPEDED THE ASSEMBLY, PREVENTING ENGAGEMENT OF THE METAL LOCKING RING WITH THE LINER. ALTERNATIVELY, THE METAL LOCKING RING MAY HAVE BEEN MISORIENTED, AS IT IS DESIGNED TO MATE IN ONE ORIENTATION, WITH DIFFERENT SIZED FINGER TABS TO CORRELATE WITH THE FINGER SIZES OF THE LINER. HOWEVER, THE EXACT CAUSE OF THE CURRENT SITUATION CANNOT BE CONCLUSIVELY DETERMINED. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | N | Y | HEALTH PROFESSIONAL |  | 2009-12
1822565-2009-01221 | 2009-12-22 | 2009-11-30 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRABECULAR METAL ACETABULAR REVISION SYSTEM - CAGE, LEFT SHORT FLANGE | HIP PROSTHESIS | ZIMMER, INC. | NA | 00712407668 |  | IT IS REPORTED THAT THE INNER, STERILE PACKAGING WAS WELDED TOGETHER WITH THE OUTER PACKAGE. THE SURGERY STAFF DID NOT EXPECT THAT AND THE PRODUCT NEARLY FELL TO THE FLOOR DURING SURGERY. AS THERE IS ALWAYS ONLY ONE IMPLANT PER SIZE DELIVERED, THERE WOULD HAVE BEEN NO REPLACEMENT PART FOR THE ONGOING SURGERY. | THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETED. | N | Y | HEALTH PROFESSIONAL |  | 2009-12
1822565-2009-01236 | 2009-12-23 | 2009-09-29 | I, n, j, u, r, y | Manufacturer report | ZMR HIP SYSTEM FEMORAL BODY REVISION TAPER SPOUT BODY | HIP PROSTHESIS | ZIMMER, INC. | NA | 00999301755 |  | IT IS REPORTED THAT THE DEVICES WERE IMPLANTED IN 2005 AND REVISED IN 2009 DUE TO THE STEM FRACTURING. | EVALUATION SUMMARY: NO PRODUCT WAS RETURNED FOR EVALUATION. ALSO, NO X-RAYS AND/OR OPERATIVE REPORTS WERE RETURNED FOR REVIEW. THE DEVICES WERE IN VIVO FOR APPROX 3 YRS AND 10 MONTHS. IMAGES PROVIDED SHOW STEM FRACTURED JUST INSIDE THE TAPER ON THE PROXIMAL BODY. THE PACKAGE INSERT CONTAINS STATEMENTS WARNING THAT DEVICE FRACTURES MAY OCCUR WHERE EXCESSIVE PT WEIGHT, EXCESSIVE PT ACTIVITY, OR LACK OR PROXIMAL SUPPORT OF THE BODY ARE INVOLVED. BASED ON THE AVAILABLE INFO, A DEFINITIVE ROOT CAUSE CANNOT BE ASCERTAINED AT THIS TIME, THOUGH PT WEIGHT MAY BE A CONTRIBUTING FACTOR. NO PRODUCT WAS RETURNED. REVIEWED OF THE DEVICE HISTORY RECORDS FOR LOT# 68141000 DID NOT FIND ANY DEVIATIONS OR ANOMALIES. REVIEW OF THE DEVICE HISTORY RECORDS WAS ALSO NOT POSSIBLE FOR THE UNK ZMR TAPER STEM AS THE PRODUCT AND/OR LOT NUMBERS REQUIRED FOR RETRIEVAL WERE UNAVAILABLE. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFO BE RECEIVED, THE COMPLAINT WILL BE RE-OPENED. ZIMMER, INC CONSIDERS THE INVESTIGATION CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
1822565-2009-01251 | 2009-12-23 | 2007-02-28 | I, n, j, u, r, y | Manufacturer report | VERSYS HIP SYSTEM EPOCH FULLCOAT PROSTHESIS FEMORAL STEM | HIP PROSTHESIS | ZIMMER, INC. | NA | 00408801406 |  | IT IS REPORTED THAT TWO HOURS AFTER INITIAL SURGERY IN 2007, THE PT WAS DIAGNOSED WITH A FEMORAL FRACTURE WHICH RESULTED IN A COMPLETE REVISION AND ADDING OF CABLES ON THE SAME DAY. | EVALUATION SUMMARY: THE EXPLANTED DEVICE WAS NOT RETURNED FOR EVALUATION. IT IS UNLIKELY THAT THE IMPLANT WAS DIRECTLY INVOLVED WITH A FRACTURE IN THE FEMUR. WITH THE INFO AVAILABLE, THE PROBABLE CAUSE OF THIS FRACTURE CANNOT BE DETERMINED. FACTORS THAT COULD BE INVOLVED INCLUDE EXCESSIVE IMPACTION OF THE STEM, IMPROPER FEMUR PREPARATION (I.E. REAMED CANAL), AND THE NATURAL VARIATION IN BONE MATERIAL PROPERTIES. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFO BE RECEIVED, THE COMPLAINT WILL BE RE-OPENED. ZIMMER, INC CONSIDERS THE INVESTIGATION CLOSED. | Y | N | LAY USER/PATIENT |  | 2009-12
9613350-2009-00753 | 2010-01-04 | 2009-12-03 | I, n, j, u, r, y | Manufacturer report | DUROM RESURFACING | NONE | ZIMMER GMBH |  |  |  | COMMENT FROM THE SURGEON: REGARDING SPECIFICALLY THE CUP BEING LOOSE, I HAVE ACTUALLY REVISED A DUROM RESURFACING WHICH WAS DONE IN CHORLEY ON A PT FROM OUR AREA. THERE WERE A VARIETY OF REASONS FOR REVISING THIS HIP WHICH INCLUDED A LOT OF PAIN, MASSIVE HETEROTOPIC OSSIFICATION AND STEEP INCLINATION OF THE CUP ANGLE. THE ONE THING I WAS NOT EXPECTING TO FIND WHEN I WENT IN WAS THAT THE HIP WAS LOOSE, AS THE X-RAY LOOKED ABSOLUTELY FINE IN TERMS OF CUP FIXATION. HOWEVER, WHEN I GOT IN, I FOUND THE ONLY THING HOLDING THE SOCKET IN PLACE WAS A LITTLE RIM OF BONE THAT HAD GROWN AROUND THE EDGE OF IT. ONCE THIS WAS REMOVED, THE CUP JUST FELL OUT AND THERE WAS NO BONDING OF THE CUP OF THE UNDERLYING BONE WHATSOEVER. | NO PRODUCT WAS RETURNED FOR EVALUATION. REVIEW OF THE DEVICE HISTORY RECORDS WAS ALSO NOT POSSIBLE AS THE ITEM NUMBERS AND/OR LOT NUMBERS WERE NOT PROVIDED. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF A DEVICE DEFICIENCY CAUSING OR CONTRIBUTING TO THE REPORTED PROBLEM. BASED ON THE AVAILABLE INFORMATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL INFORMATION BE RECEIVED THAT CHANGES THIS CONCLUSION, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. (B)(4). | Y | N | LAY USER/PATIENT |  | 2010-01
9613350-2009-00754 | 2010-01-04 | 2009-12-03 | I, n, j, u, r, y | Manufacturer report | DUROM CUP | NONE | ZIMMER GMBH |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS NECESSARY DUE TO PAIN. | NO PRODUCT WAS RETURNED FOR EVALUATION. REVIEW OF THE DEVICE HISTORY RECORDS WAS ALSO NOT POSSIBLE AS THE ITEM NUMBERS AND/OR LOT NUMBERS WERE NOT PROVIDED. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF A DEVICE DEFICIENCY CAUSING OR CONTRIBUTING TO THE REPORTED PROBLEM. BASED ON THE AVAILABLE INFORMATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL INFORMATION BE RECEIVED THAT CHANGES THIS CONCLUSION, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. (B)(4). | Y | N | LAY USER/PATIENT |  | 2010-01
1822565-2010-00004 | 2010-01-05 | 2009-04-30 | I, n, j, u, r, y | Manufacturer report | TRILOGY ACETABULAR SYSTEM LONGEVITY CROSSLINKED POLYETHYLENE LINER STANDARD | HIP PROSTHESIS | ZIMMER, INC. | NA | 00630506040 |  | IT IS REPORTED THAT THE DEVICES WERE IMPLANTED ON (B) (6) 2009 AND THAT THE PATIENT WAS REVISED ON (B) (6) 2009 DUE TO ETO - EXTENDED TROCHANTERIC OSTEOTOMY. | (B) (4). EVALUATION SUMMARY: AS IS NOTED ON THE HIP REGISTRY ADVERSE EVENTS REPORT DATED (B) (6) 2009, IT IS MOST LIKELY THAT THE COMPONENTS DID NOT CONTRIBUTE TO THE ALLEGED FAILURE OF NON-UNION OF THE ETO. THE SURGICAL NOTES FROM THE FIRST REVISION SURGERY IN (B) (6) 2009 AS WELL AS THE SECOND REVISION SURGERY IN (B) (6) 2009 ARE NOT AVAILABLE. X-RAY IMAGES POST-FIRST REVISION SURGERY AND FROM SUCCESSIVE PATIENT VISITS ARE UNAVAILABLE. IT IS UNKNOWN IF THE COMPONENTS WERE IMPLANTED WITH THE CORRECT ORIENTATION AND FIT AS PER THE SURGICAL TECHNIQUE. PATIENT ACTIVITY LEVEL IS ALSO UNKNOWN. BASED ON THE ABOVE INFORMATION AND OBSERVATIONS, IT IS MOST LIKELY THE NON-UNION OF THE ETO WAS DUE TO PATIENT FACTORS CONTRIBUTING TO BONE QUALITY, OF WHICH ARE NOT IN THE CONTROL OF THE DEVICE MANUFACTURER. HOWEVER, THE EXACT CAUSE OF THE CURRENT SITUATION CANNOT BE CONCLUSIVELY DETERMINED. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | LAY USER/PATIENT |  | 2010-01
1822565-2010-00003 | 2010-01-05 | 2009-12-02 | I, n, j, u, r, y | Manufacturer report | ZMR HIP SYSTEM FEMORAL STEM REVISION TAPER W/COMPRESSION NUT | HIP PROSTHESIS | ZIMMER, INC. | NA | 00998201918 |  | IT IS REPORTED THAT THE DEVICES WERE IMPLANTED IN 2005 AND REVISED IN 2009, DUE TO THE STEM FRACTURING. | THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. | Y | N | LAY USER/PATIENT |  | 2010-01
9613350-2010-00001 | 2010-01-12 | 2009-09-30 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 50/44 J | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.150 |  | IT IS REPORTED BY PT'S ATTORNEY THAT PT UNDERWENT A LEFT HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, SHE EXPERIENCED PAIN AND WAS REVISED ON (B) (6) 2008. | INFO WAS FORWARDED FROM (B) (4), ZIMMER, INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THIS INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-01
9613350-2010-00008 | 2010-01-12 | 2009-12-21 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 64/58 X | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.164 |  | IT IS REPORTED THAT CUP WAS LOOSE. NO INGROWTH DETECTED UPON REMOVAL. | INFO WAS FORWARDED FROM (B) (4), ZIMMER, INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THIS INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-01
9613350-2010-00009 | 2010-01-12 | 2010-01-04 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 48/42 H | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.148 |  | IT IS REPORTED BY PT THAT PT UNDERWENT A HIP ARTHROPLASTY IN (B) (6) OF 2007. POST-OP, SHE EXPERIENCES PAIN AND A REVISION IS SCHEDULED FOR (B) (6) 2010. | INFO WAS FORWARDED FROM (B) (4), ZIMMER, INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THIS INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-01
9613350-2010-00010 | 2010-01-12 | 2010-01-07 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 48/42 H | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.148 |  | IT IS REPORTED BY PT'S ATTORNEY THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2008. POST-OP, SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION IS SCHEDULED FOR AN UNK DATE. | (B) (4). THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IIN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-01
9613350-2010-00011 | 2010-01-12 | 2009-12-22 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 56/50 P | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.156 |  | IT IS REPORTED BY PT THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2009. | INFO WAS FORWARDED FROM (B) (4), ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-01
9613350-2010-00012 | 2010-01-12 | 2010-01-04 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 46/40 F | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.146 |  | IT IS REPORTED BY PT'S ATTORNEY THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2008. POST-OP, SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WILL BE REVISED ON (B) (6) 2010. | (B) (4). THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED, THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-01
9613350-2009-00755 | 2010-01-12 | 2009-12-23 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 50/44 J | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.150 |  | IT WAS REPORTED BY PT THAT PT UNDERWENT A RIGHT HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION IS SCHEDULED FOR (B) (6) 2010. | INFO WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THIS INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-01
9613350-2010-00002 | 2010-01-12 | 2010-01-04 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.154 |  | IT IS REPORTED BY PT THAT PT UNDERWENT A RIGHT HIP ARTHROPLASTY ON (B) (6) 2008. SHE EXPERIENCES PAIN AND A REVISION HAS BEEN RECOMMENDED BY HER PHYSICIAN. | INFO WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER (B) (4) WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THIS INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-01
9613350-2010-00007 | 2010-01-12 | 2009-10-10 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | UNK |  | IT IS REPORTED BY PT'S ATTORNEY THAT PT UNDERWENT A RIGHT HIP ARTHROPLASTY ON (B) (6) 2008. POST-OP, SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2008. | INFO WAS FORWARDED FROM (B) (4), ZIMMER, INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THIS INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-01
9613350-2009-00740 | 2010-01-12 | 2009-08-11 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 50/44 J | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.150 |  | IT IS REPORTED BY PT'S ATTORNEY THAT PT UNDERWENT A LEFT HIP ARTHROPLASTY ON (B) (6) 2008. POST-OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6), 2009. | INFO WAS FORWARDED FROM (B) (4), ZIMMER, INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THIS INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (64. | Y | N | LAY USER/PATIENT |  | 2010-01
9613350-2009-00736 | 2010-01-12 | 2009-11-25 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 62/56 V | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.162 |  | IT IS REPORTED THAT PT PRESENTED AND WAS DX WITH ACETABULAR IMPLANT LOOSENING WHICH RESULTED IN COMPLETE REVISION OF THA. | INFO WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER, INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THIS INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-01
1822565-2010-00012 | 2010-01-12 | 2009-12-14 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | HGP II SHELL DEPTH GAUGE | HIP INSTRUMENT | ZIMMER, INC. | NA | 00661109800 |  | IT IS REPORTED THAT UPON INSERTION OF THE DISTAL END OF THE GAUGE, THE TIP FRACTURED WHEN RETRACTING THE THUMB RING. | THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. | N | Y | HEALTH PROFESSIONAL |  | 2010-01
9613350-2009-00743 | 2010-01-12 | 2009-12-17 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 46/40 F | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.146 |  | IT IS REPORTED BY PT'S FIANCEE THAT PT UNDERWENT A HIP ARTHROPLASTY IN (B) (6) OF 2007. POST-OP, SHE EXPERIENCED PAIN AND WAS REVISED ON (B) (6) 2009. | INFO WAS FORWARDED FROM (B) (4), ZIMMER, INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THIS INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-01
1822565-2010-00025 | 2010-01-13 | 2009-12-10 | I, n, j, u, r, y | Manufacturer report | TRABECULAR METAL MODULAR ACETABULAR SYSTEM SHELL WITH MULTIHOLES POROUS | HIP PROSTHESIS | ZIMMER, INC. | NA | 00620206020 |  | IT IS REPORTED THAT THE DEVICES WERE IMPLANTED IN 2009 AND THAT THE PT WAS REVISED THREE MONTHS LATER BECAUSE THE SHELL RIPPED OUT OF THE PT'S PELVIS. | EVAL SUMMARY: THE SHELL WAS RETURNED WITH A TM AUGMENT CEMENTED TO ITS SURFACE AND FOUR SCREWS IN PLACE THROUGH THE CEMENT. ONE OF THE FOUR SCREWS WAS BROKEN DISTAL TO THE CEMENT WHERE IT INTERFACED WITH BONE. TWO ADD'L SCREWS WERE RETURNED UNASSEMBLED TO THE SHELL. THE SHELL SHOWED MINOR BONE INGROWTH. THE SHELL AND AUGMENT SHOW NO VISIBLE SIGNS OF DAMAGE. THE LOCKING RING IS IN PLACED AND UNDAMAGED. THE CONSTRAINED LINER WAS NOT RETURNED AND ITS CONDITION IS UNK. IT IS UNK IF THE BROKEN SCREW WAS PRESENT AT THE TIME OF REVISION OR IT BROKE DURING REVISION. PT'S WEIGHT AND ACTIVITY LEVEL ARE UNK. THE PT'S QUALITY OF ACETABULAR BONE IS UNK. THE MINOR BONY-IN-GROWTH SUGGESTS THE IMPLANT WAS HELD IN PLACE PRIMARILY BY THE SCREWS. IT IS POSSIBLE THE IMPLANT EXPERIENCED FORCES THAT EXCEEDED THE PULL OUT STRENGTH OF THE SCREWS FROM THE BONE. HOWEVER, BASED ON THE INFO ABOVE, A DEFINITIVE CAUSE OF THE EXPERIENCE OBSERVED CANNOT BE DETERMINED WITH CERTAINTY. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADD'L SUBSTANTIVE INFO BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | LAY USER/PATIENT |  | 2010-01
1822565-2010-00024 | 2010-01-13 | 2009-11-10 | I, n, j, u, r, y | Manufacturer report | ZIMMER M/L TAPER HIP PROSTHESIS WITH KINECTIV TECHNOLOGY | HIP PROSTHESIS | ZIMMER, INC. | NA | 00771300900 |  | IT WAS REPORTED THAT THE DEVICE IMPLANTED IN 2009 AND THAT THE PT WAS REVISED THE FOLLOWING MONTH DUE TO A FRACTURE OF THE FEMORAL SHAFT. | EVALUATION SUMMARY: THE DEVICES WERE NOT RETURNED, AND X-RAYS WERE NOT PROVIDED. IT IS UNK IF THE COMPONENTS WERE IMPLANTED WITH THE CORRECT ORIENTATION AND FIT AS PER THE SURGICAL TECHNIQUE. PT ACTIVITY LEVEL POST-OP IS ALSO UNK. IT IS POSSIBLE THE POST-OPERATIVE FRACTURE WAS A RESULT OF THE PT FALLING DOWN STAIRS WITHIN 4 WEEKS OF THE PRIMARY SURGERY AS INDICATED PER THE INCLUDED PRODUCT EXPERIENCE REPORT. ACCORDING TO THE SUBMITTED ADVERSE EVENTS REPORT THE SURGEON INDICATED THE EVENT WAS NOT RELATED TO THE DEVICE AND MOST LIKELY WAS CAUSED BY THE FALL. A DEFINITIVE ROOT CAUSE ANALYSIS CANNOT BE PERFORMED WITH CERTAINTY AT THIS TIME. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADD'L SUBSTANTIVE INFO BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. | Y | N | LAY USER/PATIENT |  | 2010-01
1822565-2010-00023 | 2010-01-13 | 2009-12-16 | I, n, j, u, r, y | Manufacturer report | M/L TAPER HIP PROSTHESIS FEMORAL STEM PRESS-FIT PLASMA SPRAYED | HIP PROSTHESIS | ZIMMER, INC. | NA | 00771101010 |  | IT IS REPORTED THAT THE DEVICE WAS IMPLANTED IN 2006 AND THAT THE PT WAS REVISED IN 2008, DUE TO INFECTION. | EVALUATION: THE STERILIZATION PROCESS FOR THIS DEVICE WAS VALIDATED IN ACCORDANCE WITH FDA REGULATIONS AND ISO STANDARDS TO A STERILITY ASSURANCE LEVEL OF 1.0 X 10-6 OR BETTER. IN ADDITION, THE MANUFACTURING LOT SPECIFIED IN THIS COMPLAINT WAS PROCESSED ACCORDING TO THE VALIDATED STERILIZATION PROCESS PARAMETERS AND MET ALL THE ACCEPTANCE CRITERIA FOR STERILITY RELEASE. THEREFORE, IT IS HIGHLY UNLIKELY THAT THE SPECIFIED DEVICE CAUSED OR CONTRIBUTED TO THE PT INFECTION. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION NOT INDICATED. SHOULD ADD'L SUBSTANTIVE INFO BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC CONSIDERS THE INVESTIGATION CLOSED. | Y | N | LAY USER/PATIENT |  | 2010-01
1822565-2010-00022 | 2010-01-13 | 2009-12-11 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | EPSILON ACETABULAR SYSTEM DURASUL CONSTRAINED INSERT | HIP PROSTHESIS | ZIMMER, INC. | NA | 438032055 |  | IT IS REPORTED THAT DURING SURGERY ON (B) (6) 2009, THE 55 EPSILON CONSTRAINED LINER WAS IMPLANTED INTO A WELL FIXED 55MM CONVERGE LINER. LOCKING RING WAS IMPACTED USING THE APPROPRIATE TECHNIQUE AND INSTRUMENTATION. LOCKING RING FAILED TO SEAT INTO THE LINER. A SECOND LINER WAS NOT AVAILABLE. SURGEON EXPLANTED LINER AND CONVERGE CUP, DUE TO FAILURE. SURGERY WAS DELAYED 45 MINUTES. | EVALUATION SUMMARY: THE RECEIVED DEVICE WAS REVIEWED. IT WAS OBSERVED THAT IN ITS RETURNED CONDITION THE LINER HAD TWO SCRATCHES ON THE OUTER SURFACE APPROX 2 CM IN LENGTH AND LOCATED BELOW THE FLANGE FRACTURE; FLANGE WAS FRACTURED AT THE SHELL/FLANGE GROOVE, WITH THREE HOLES WHICH RESULTED IN DEFORMED FLANGE GEOMETRY; ONE OF THESE HOLES RESULTED IN A SHEARED POLY EDGE, WHICH PROTRUDED INTO THE INNER SURFACE; IT ALSO HAD A DIAGONAL CUT/FRACTURE BEGINNING IN THE CENTER OF THE FRACTURED FINGER LOCATION; ONE FINGER SHEARED OFF IN A CONCAVE "V" FRACTURE PATTERN AND HAD SCRATCHES ON THE INNER SURFACE, FLANGE SURFACE AND TIP OF FINGER. THE METAL LOCKING RING HAD NICKS IN TWO LOCATIONS APPROX 1 CM ON EITHER SIDE OF THE LOT NUMBER ETCH, SCUFF MARKS BELOW "THIS FACE" ETCH OF THE "THIS FACE TOWARD ACETABULUM" ETCH, NICKS/GOUGES ON EDGES OF "THIS FACE TOWARD FEMUR" SURFACE AS WELL AS BOTH FINGER EDGES, EXTENDING TO THE TOP OF THE FINGER EDGES, AND GOUGES ON TOP FINGER SURFACES AROUND PEG HOLES. THE GOUGES WERE MORE SIGNIFICANT ON THE SINGLE HOLE AND PEG HOLE CLOSEST TO THE "THIS FACE TOWARD FEMUR ETCH". THE OPERATIVE NOTES ARE UNAVAILABLE, AND IT IS UNK IF THE METAL LOCKING RING WAS ATTEMPTED TO BE ASSEMBLED AS PER THE SURGICAL TECHNIQUE. BASED ON THE ABOVE INFO AND OBSERVATIONS, THE DIFFICULTY ASSEMBLING THE METAL LOCKING RING TO THE LINER MAY HAVE BEEN DUE TO ONE OR MORE OF THE FOLLOWING MECHANISMS: SOFT TISSUE IMPINGEMENT, PREVENTING ENGAGEMENT OF THE METAL LOCKING LINER WITH THE LINER OR THE METAL LOCKING RING OR IMPACTOR MAY HAVE BEEN MISORIENTED. BASED ON THE AVAILABLE INFO, A DEFINITIVE ROOT CAUSE CANNOT BE ASCERTAINED AT THIS TIME. EVALUATION: DEVICE HISTORY RECORDS INDICATE ALL COMPONENTS WERE MANUFACTURED AND INSPECTED TO SPECIFICATION. | N | Y | HEALTH PROFESSIONAL |  | 2010-01
1822565-2010-00021 | 2010-01-13 | 2009-12-16 | I, n, j, u, r, y | Manufacturer report | TRILOGY ACETABULAR SYSTEM LONGEVITY CROSSLINKED POLYETHYLENE LINER | HIP PROSTHESIS | ZIMMER, INC. | NA | 00630505036 |  | IT IS REPORTED THAT THE DEVICE WAS IMPLANTED IN 2009, AND REVISED ELEVEN MONTHS LATER, DUE TO REPEAT DISLOCATIONS. | EVAL SUMMARY: THE FOLLOWING EVENT COULD BE DUE TO ONE OR MORE OF THE FOLLOWING MECHANISMS: UNSATISFACTORY PLACEMENT OF THE COMPONENT PARTS (SHELL, STEM, OR LINER), EXTENT OF COMPONENT STABILITY, EXTENT OF SOFT TISSUE TENSION OR PT BEHAVIOR. BASED ON THE AVAILABLE INFO, A DEFINITIVE ROOT CAUSE CAN NOT BE ASCERTAINED AT THIS TIME. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFO BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-01
1591413 | 2010-01-14 | 2009-11-23 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | DUROM | IMPLANT, HIP, ACETABULAR CUP | ZIMMER GMBH | * | * |  | NOTES FROM DISCHARGE SUMMARY: "THE PATIENT WAS KNOWN TO HAVE A FAILED LEFT TOTAL HIP ARTHROPLASTY, ACETABULAR COMPONENT. ON BONE SCAN, THE COMPONENT WAS LOOSE AND SHE WAS HAVING SIGNIFICANT AMOUNT OF PAIN. SHE WAS ADMITTED TO THE HOSPITAL IN FALL 2009, WHEN SHE UNDERWENT A LEFT HIP REVISION, ACETABULAR COMPONENT. SHE TOLERATED THE PROCEDURE WELL AND WAS RETURNED TO RECOVERY ROOM IN STABLE CONDITION." ON POSTOPERATIVE DAY #3, THE PATIENT WAS DISCHARGED TO AN EXTENDED CARE FACILITY FOR REHABILITATION. |  | N | Y | INVALID DATA |  | 2010-01
9613350-2009-00737 | 2010-01-15 | 2009-02-01 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.154 |  | IT IS REPORTED THAT PT WAS REVISED DUE TO AN INFECTION IN HIS RIGHT HIP. | NO PRODUCT WAS RETURNED FOR EVALUATION. THE INFO REPORTED BY THE OPERATING SURGEON INDICATED THAT THE REVISION SURGERY WAS NOT RELATED TO OR CAUSED BY THE ZIMMER DEVICES. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF A DEVICE DEFICIENCY CAUSING OR CONTRIBUTING TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADD'L INFO BE RECEIVED THAT CHANGES THIS CONCLUSION, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-01
9613350-2009-00707 | 2010-01-18 | 2009-11-18 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 48/42 CODE H | METASUL LARGE DIAMETER HEAD 42/H | ZIMMER GMBH |  | 01.00181.420 |  | IT WAS REPORTED THAT REVISION WAS NECESSARY DUE TO A COBALT ALLERGY. HIP PAIN STARTED APPROXIMATELY 4 MONTHS POST-OP. | THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETED. | Y | N | LAY USER/PATIENT |  | 2010-01
MW5014426 | 2010-01-18 | 2008-06-26 | I, n, j, u, r, y | Voluntary report | VERSYS 12/14 FEMORAL HEAD | ARTIFICIAL HIP | ZIMMER | PART NO: | 00-8018-028-02 |  | PT UNDERWENT BIPOLAR HEMIARTHROPLASTY/ENDOPROSTHESIS FOR RIGHT HIP FRACTURE, SUING ZIMMER BRAND VERSYS 12/14 FEMORAL HEAD 28MM. SIX MONTHS POST SURGERY, PT DEVELOPED CONSTANT BURNING IN THE RIGHT BUTTOCK AND SOME NUMBNESS. ON 11/13/2009 THE FDA ISSUED A RECALL CLASS 2, (B) (4). THE LOT LISTED IS 60378581, AND THE LOT NUMBER REFLECTED IN PT'S (B) (6) HEALTH SERVICES RECORDS IS (B) (6). PT IS CONCERNED AS TO WHETHER THE RECALL NOTICE OUGHT TO INCLUDE HER DEVICE, AND WHETHER A DEFECT OF THE DEVICE IS CAUSING CURRENT SYMPTOMS. PT WOULD LIKE ADVICE ON WHAT EVALUATION/TESTING IS NECESSARY TO DETERMINE WHETHER THE DEVICE IS DEFECTIVE AND FAILING. DATES OF USE: (B) (6) 2008 - (B) (6) 2010. DIAGNOSIS OR REASON FOR USE: FRACTURED HIP. |  | N | Y | HEALTH PROFESSIONAL |  | 2010-01
9613350-2009-00733 | 2010-01-19 | 2009-11-19 | I, n, j, u, r, y | Manufacturer report | METASUL DUROM, COMPONENT FOR ACETABULUM, UNCEMENTED, 52/46, CODE L | DUROM ACETABULAR COMPONENT | ZIMMER GMBH | N/A | 01.00214.052 |  | A PRODUCT LIABILITY CLAIM WAS RAISED. IT WAS REPORTED THAT THE PATIENT WAS IMPLANTED A METASUL DUROM, COMPONENT FOR ACETABULUM, UNCEMENTED, 52/46, CODE L AND WAS REVISED AFTER FIVE YEARS AND FOUR MONTHS IN VIVO DUE TO LOOSENING AND PAIN. | THIS CASE HAS BEEN REOPENED TO ENTER ADDITIONAL INFORMATION (IMPLANTATION DATE AND CORRECTED EXPLANTATION DATE) MEDICAL PRODUCTS: METASUL LDH, HEAD ADAPTER, L, +4, TAPER 12/14-18/20, REF# 01.00185.147 LOT#: 2190922; METASUL LDH, HEAD, 46, CODE L, TAPER 18/20, REF# 01.00181.460, LOT# 2189275; CLS SPOTORNO, STEM, 135, UNCEMENTED, 11.25, TAPER 12/14, REF# 290039112, LOT#: 2213445. DHR REVIEW RESULTS: THE QUALITY RECORDS INDICATE THAT THESE COMPONENTS MET ALL REQUIREMENTS TO PERFORM AS INTENDED. TREND ANALYSIS: TREND HAS BEEN IDENTIFIED AND CAPA WAS INITIATED AND PERFORMED. REVIEW OF EVENT DESCRIPTION: PATIENT BEING REVISED DUE TO LOOSENING AND PAIN. OTHER INFORMATION & SOURCES: REVIEW OF THE COMPLAINT RELEVANT DOCUMENTS DID NOT LEAD TO NEW INFORMATION REGARDING THE REPORTED EVENT. SINCE THIS CASE IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTIVE ACTION WHICH WAS REPORTED TO THE FDA IN JULY 2008 AS CORRECTION Z-2415/2426-2008, THERE WILL BE NO FURTHER INVESTIGATION AND ZIMMER GMBH WILL CLOSE THIS CASE ONCE AGAIN. ZIMMER¿S REFERENCE NUMBER OF THIS FILE IS (B)(4). | Y | Y | LAY USER/PATIENT |  | 2010-01
1822565-2010-00036 | 2010-01-20 | 2009-12-09 | I, n, j, u, r, y | Manufacturer report | CENTRALIGN PRECOAT HIP PROSTHESIS FEMORAL STEM | HIP PROSTHESIS | ZIMMER, INC. | NA | 00980000200 |  | THE STEM IS BROKEN. | THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. | Y | N | HEALTH PROFESSIONAL |  | 2010-01
2648920-2010-00002 | 2010-01-26 | 2009-12-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | VERSA-FX TIV SELF TAPPING BONE SCREW | HIP PROSTHESIS | ZIMMER MFG. B.V., | NA | 48231904400 |  | IT IS REPORTED THAT WHEN THE SURGEON WAS INSERTING THE SCREW, THE SCREW BROKE UNDER THE SCREW HEAD. THE SCREW PART WAS EXTRACTED FROM THE PATIENT'S FEMUR. ITEM NOT SOLD IN THE USA. | THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. | N | Y | HEALTH PROFESSIONAL |  | 2010-01
9613350-2009-00691 | 2010-02-02 | 2009-10-13 | I, n, j, u, r, y | Manufacturer report | METASUL LARGE DIAMETER HEAD 52/R | PROSTHESIS, HIP | ZIMMER GMBH |  | 01.00181.520 |  | IT WAS REPORTED THAT A REVISION SURGERY OCCURRED DUE TO A POSSIBLE ALLERGIC REACTION AND PAIN. | (B) (4). NO PRODUCT WAS RETURNED FOR EVALUATION. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. BASED ON AN EXTENSIVE INVESTIGATION OF EVENTS REPORTED FROM SEVERAL USER FACILITIES OUTSIDE THE U.S., ZIMMER IDENTIFIED THAT THE MOST PROBABLE CAUSE FOR THE OUTCOME OBSERVED WAS A LOOSE OR UNSTABLE CUP THAT RESULTED FROM USE OF SURGICAL TECHNIQUES NOT CONSISTENT WITH THE MFR'S RECOMMENDATIONS. AS A CORRECTIVE ACTION, A RETRAINING PROGRAM FOR USERS OUTSIDE THE U.S. WAS INITIATED IN NOV 2009 AND REPORTED TO THE NATIONAL COMPETENT AUTHORITIES AS REQUIRED. THE DUROM CUP REPORTED IN THIS CASE IS NOT MARKETED IN THE U.S. A SIMILAR CUP, COMPATIBLE WITH THE METASUL LDH FEMORAL HEAD, IS CLEARED IN THE U.S. AND A CORRECTIVE ACTION FOR THIS PRODUCT WAS REPORTED TO THE FDA IN JULY 2008 AS CORRECTION Z-2415/2426/2008. NO FURTHER ACTION IS INDICATED AND ZIMMER CONSIDERS THE INVESTIGATION CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00032 | 2010-02-03 | 2010-01-13 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 50/44 CODE J | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.050 |  | IT WAS REPORTED THAT THE REVISION SURGERY OCCURRED DUE TO PAIN. | NO PRODUCT WAS RETURNED FOR EVALUATION. REVIEW OF THE DEVICE HISTORY RECORDS WAS ALSO NOT POSSIBLE AS THE ITEM NUMBERS AND/OR LOT NUMBERS WERE NOT PROVIDED. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF A DEVICE DEFICIENCY CAUSING OR CONTRIBUTING TO THE REPORTED PROBLEM. BASED ON THE AVAILABLE INFORMATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL INFORMATION BE RECEIVED THAT CHANGES THIS CONCLUSION, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00033 | 2010-02-03 | 2010-01-11 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 54/48 CODE N | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.054 |  | IT WAS REPORTED THAT THE REVISION SURGERY OCCURRED DUE TO PAIN. | NO PRODUCT WAS RETURNED FOR EVALUATION. REVIEW OF THE DEVICE HISTORY RECORDS WAS ALSO NOT POSSIBLE AS THE ITEM NUMBERS AND/OR LOT NUMBERS WERE NOT PROVIDED. IT IS NO SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF A DEVICE DEFICIENCY CAUSING OR CONTRIBUTING TO THE REPORTED PROBLEM. BASED ON THE AVAILABLE INFORMATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL INFORMATION BE RECEIVED THAT CHANGES THIS CONCLUSION, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-02
1822565-2010-00047 | 2010-02-03 | 2007-08-03 | I, n, j, u, r, y | Manufacturer report | TRABECULAR METAL PRIMARY HIP PROSTHESIS FEMORAL STEM PRESS-FIT COLLARLESS | HIP PROSTHESIS | ZIMMER, INC. | NA | 00786401320 |  | IT IS REPORTED BY PT'S ATTORNEY THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2007. POST OP, HE EXPERIENCED PAIN DUE TO AN E-COLI INFECTION AND WAS REVISED ON (B) (6) 2007. | EVALUATION SUMMARY: THE STERILIZATION PROCESS FOR THIS DEVICE WAS VALIDATED IN ACCORDANCE WITH FDA REGULATIONS AND ISO STANDARDS TO A STERILITY ASSURANCE LEVEL OF 1.0 X 10-6 OR BETTER. IN ADDITION, THE MFG LOT SPECIFIED IN THIS COMPLAINT WAS PROCESSED ACCORDING TO THE VALIDATED STERILIZATION PROCESS PARAMETERS AND MET ALL THE ACCEPTANCE CRITERIA FOR STERILITY RELEASE. THEREFORE, IT IS HIGHLY UNLIKELY THAT THE SPECIFIED DEVICE CAUSED OR CONTRIBUTED TO THE PT INFECTION. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADD'L SUBSTANTIVE INFO BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-02
9613350-2010-00003 | 2010-02-04 | 2009-12-03 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | RESURFACING IMPACTOR HEAD FOR HANDLE | DUROM HIP RESURFACING SYSTEM, FEMORAL COMPONENT | ZIMMER GMBH |  | 01.00219.817 |  | IT IS REPORTED THAT THE IMPACTOR HEAD COULD NOT BE REMOVED FROM IMPACTOR HANDLE. | THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. (B) (4) | N | Y | HEALTH PROFESSIONAL |  | 2010-02
9613350-2010-00040 | 2010-02-04 | 2010-01-05 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | UNK |  | IT IS REPORTED BY PT THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2008. POST-OP, SHE EXPERIENCES PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION IS SCHEDULED FOR (B) (6) 2010. | THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00005 | 2010-02-04 | 2009-06-01 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 52/46 L | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.152 |  | IT IS REPORTED THAT PT PRESENTED TO THE OFFICE AND WAS DX WITH ACETABULAR IMPLANT LOOSENING. THE PT IS SCHEDULED FOR REVISION ON (B) (6) 2010. | NO PRODUCT WAS RETURNED FOR EVALUATION. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF A DEVICE DEFICIENCY CAUSING OR CONTRIBUTING TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL INFORMATION BE RECEIVED THAT CHANGES THIS CONCLUSION, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00038 | 2010-02-04 | 2009-08-21 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 58/52 R | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.158 |  | IT IS REPORTED BY THE PT'S ATTORNEY THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2006, POST-OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2009. | THE MANUFACTURER DID NOT RECEIVE THE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBER WERE RECEIVED FOR EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADD'L INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00030 | 2010-02-04 | 2010-01-07 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 52/46 L | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.152 |  | IT IS REPORTED BY PT'S ATTORNEY THAT PT UNDERWENT A RIGHT HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION HAS EITHER TAKEN PLACE OR IS SCHEDULED. | THE MANUFACTURER DID NOT RECEIVE THE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBER WERE RECEIVED FOR EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADD'L INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00039 | 2010-02-04 | 2009-02-16 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 58/52 R | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.158 |  | IT IS REPORTED BY PT'S ATTORNEY THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2009. | THE MANUFACTURER DID NOT RECEIVE THE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBER WERE RECEIVED FOR EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADD'L INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00031 | 2010-02-04 | 2007-09-10 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.154 |  | IT IS REPORTED BY PT'S ATTORNEY THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2006. POST-OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. A REVISION WAS PERFORMED ON (B) (6) 2007. | THE MANUFACTURER DID NOT RECEIVE THE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBER WERE RECEIVED FOR EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADD'L INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00034 | 2010-02-04 | 2010-01-07 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.154 |  | IT IS REPORTED BY PATIENT'S ATTORNEY THAT PATIENT UNDERWENT A LEFT HIP ARTHROPLASTY (B) (6) 2007. POST-OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION HAS EITHER TAKEN PLACE OR IS SCHEDULED. | THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAY, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00035 | 2010-02-04 | 2007-08-03 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 58/52 R | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.158 |  | IT IS REPORTED BY PATIENT'S ATTORNEY THAT PATIENT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, HE EXPERIENCED PAIN DUE TO AN E-COLI INFECTION AND WAS REVISED ON (B) (6) 2007. SURGEON DID INDICATE THERE WAS NO BONE INGROWTH ON THE CUP. | THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00036 | 2010-02-04 | 2010-01-07 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 50/44 J | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.150 |  | IT IS REPORTED BY PATIENT'S ATTORNEY THAT PATIENT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2008. POST-UP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION IS SCHEDULED FOR AN UNKNOWN DATE. | THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00037 | 2010-02-04 | 2010-01-07 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 48/42 H | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.148 |  | IT IS REPORTED BY PATIENT'S ATTORNEY THAT PATIENT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND REVISION IS SCHEDULED. | THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00014 | 2010-02-04 | 2009-03-27 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.154 |  | IT IS REPORTED BY THE PT'S ATTORNEY THAT PT UNDERWENT A HIS ARTHROPLASTY ON (B) (6) 2008. POST-OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2009. | THE MANUFACTURER DID NOT RECEIVE THE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBER WERE RECEIVED FOR EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADD'L INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00004 | 2010-02-04 | 2009-06-01 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 48/42 H | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.148 |  | IT IS REPORTED THAT PATIENT PRESENTED TO THE OFFICE AND WAS DX WITH ACETABULAR IMPLANT LOOSENING. THE PATIENT IS SCHEDULED FOR REVISION ON (B) (6) 2010. | NO PRODUCT WAS RETURNED FOR EVALUATION. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF A DEVICE DEFICIENCY CAUSING OR CONTRIBUTING TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL INFORMATION BE RECEIVED THAT CHANGES THIS CONCLUSION, AN AMENDED MEDICAL DEVICE REPORT WILL FILED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00013 | 2010-02-04 | 2009-11-03 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.154 |  | IT IS REPORTED BY PT'S ATTORNEY THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION TOOK PLACE ON (B) (6) 2008. | THE MANUFACTURER DID NOT RECEIVE THE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBER WERE RECEIVED FOR EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADD'L INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-02
1621683 | 2010-02-08 | 2009-12-14 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | LONGEVITY HIGHLY CROSSLINKED POLYETHYLENE STANDARD LINER | PROSTHESIS, HIP | ZIMMER, INC. | NA | 00-6305-050-32 |  | PATIENT HAD LEFT TOTAL HIP ARTHROPLASTY (THA). REVISION ONE MONTH AFTER IMPLANT, DUE TO PAIN AND BY IMAGING INDICATED THAT THE POLYETHYLENE LINER HAD UNLOCKED AND SLID OUT OF POSITION, LEAVING THE JOINT IN A SUBLUXED POSITION. IN THE OPERATING ROOM FOUND LINER TO HAVE SPLIT. CUP AND LINER REPLACED. |  | N | Y | INVALID DATA |  | 2010-02
9613350-2010-00049 | 2010-02-08 | 2009-05-27 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 52/46 L | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.152 |  | IT IS REPORTED THAT PATIENT STARTED TO EXPERIENCE PAIN. THERE WAS NO SPECIFIC EVENT THAT CAUSED THE PAIN. AFTER CHECKING THE PATIENT'S CHARTS, SURGEON DID FIND THAT SHE FELL IN (B) (6) 2009. | THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETED. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00048 | 2010-02-08 | 2010-01-07 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 50/44 J | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.150 |  | IT IS REPORTED THAT PT UNDERWENT A HIP ARTHROPLASTY IN (B) (6) 2008. POST-OP, SHE EXPERIENCES PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION IS SCHEDULED FOR (B) (6) 2010. | THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADD'L INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-02
1822565-2010-00055 | 2010-02-09 | 2008-10-31 | I, n, j, u, r, y | Manufacturer report | UNKNOWN TRILOGY ACETABULAR SHELL | HIP PROSTHESIS | ZIMMER, INC. | NA | UNK |  | IT IS REPORTED BY PT'S ATTORNEY THAT PT UNDERWENT A RIGHT HIP REVISION ARTHROPLASTY ON (B) (6) 2004. POST OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2008. | EVAL SUMMARY: THE ALLEGATION IS THAT THE ACETABULAR SHELL BECAME LOOSE APPROXIMATELY 3.75 YEARS POST-OPERATIVELY. IT IS STATED THAT THE COMPONENTS WERE USED IN A REVISION CASE. NO PRE-OPERATIVE OR POST-OPERATIVE X-RAYS WERE PROVIDED, SO THE EXTENT OF PRE-EXISTING DAMAGE TO THE ACETABULUM IS NOT KNOWN. NO INFO BEYOND THE NAME AND SEX OF THE PT HAS BEEN PROVIDED, SO THE AGE, WEIGHT, HEIGHT, ACTIVITY LEVEL, AND ANY ADDITIONAL DETAILS ARE NOT KNOWN. ADDITIONALLY, NO SURGICAL NOTES FROM THE SURGICAL OR REVISION SURGERY WERE PROVIDED AND THE DEVICES WERE NOT RETURNED FOR EVAL. THE PROBABLE CAUSE OF THIS EVENT CANNOT BE DETERMINED AT THIS TIME DUE TO LACK OF AVAILABLE INFO. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS WAS ALSO NOT POSSIBLE AS THE PRODUCT AND/OR LOT NUMBERS REQUIRED FOR RETRIEVAL WERE UNAVAILABLE. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE AVAILABLE INFO, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFO BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER INC CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-02
1822565-2010-00058 | 2010-02-09 | 2009-12-25 | I, n, j, u, r, y | Manufacturer report | ZIMMER M/L TAPER HIP PROSTHESIS WITH KINECTIV TECHNOLOGY | HIP PROSTHESIS | ZIMMER, INC. | NA | 00771301000 |  | SITE REPORTS THAT PT WAS REVISED DUE TO FRACTURE OF THE FEMORAL SHAFT. | EVAL SUMMARY: DEVICES WERE NOT RETURNED AND X-RAYS WERE NOT PROVIDED. IT IS UNK IF THE COMPONENTS WERE IMPLANTED WITH THE CORRECT ORIENTATION AND FIT WAS PER SURGICAL TECHNIQUE. PT'S ACTIVITY LEVEL POST-OP IS ALSO UNK. ACCORDING TO THE ATTACHED DEMOGRAPHIC EVAL, THE PT'S CONCURRENT MEDICAL CONDITIONS INCLUDED OSTEOPOROSIS. THE PROVIDED PT'S OPERATIVE INFO ALSO CLASSIFIED THE PT'S FEMUR AS A DORR TYPE C WITH STOVE PIPE SHAPE, DRAMATIC THINNING OF THE A/P CORTICES, AND A VERY WIDE IM CANAL. IN ADDITION TO THE ABOVE CONTRIBUTING FACTORS, SUCH AN EARLY POST-OP FEMORAL SHAFT FRACTURE MAY HAVE BEEN CAUSED BY SELECTING A STEM SIZE THAT WAS TOO LARGE FOR THE FEMUR OR THE PT MAY HAVE FALLEN. HOWEVER, THE ROOT CAUSE OF THE ALLEGED EVENT CANNOT BE DETERMINED WITH CERTAINTY AT THIS TIME BASED ON THE PROVIDED INFO. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFO BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER INC CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-02
1822565-2010-00059 | 2010-02-09 | 2010-01-06 | I, n, j, u, r, y | Manufacturer report | VERSYS HIP SYSTEM ZIRCONIUM OXIDE CERAMIC FEMORAL HEAD | HIP PROSTHESIS | ZIMMER, INC. | NA | 00811802801 |  | PT HEARD A POP SOUND AND NOW ON X-RAY YOU CAN SEE PIECES OF WHAT APPEARS TO BE THE CERAMIC HEAD AROUND THE JOINT. REVISION SURGERY IS SCHEDULED FOR (B) (6) 2010. | EVAL SUMMARY: NO PRODUCT WAS RETURNED. THE DEVICE WAS IN-VIVO APPROXIMATELY 9 YEARS BEFORE THE ALLEGED EVENT OCCURRED. THIS ZIRCONIA HEAD IS WITHIN THE SCOPE OF A RECALL IN 2001 DUE TO POTENTIAL FRACTURING OF THE FEMORAL HEAD. THE RECALL IS COMPLETE AND NO FURTHER MANUFACTURE OR MARKETING OF THE ZIRCONIA FEMORAL HEAD 8118 FAMILY OCCURS. THE DEVICE WAS IMPLANTED BEFORE THE RECALL TOOK EFFECT. | Y | N | HEALTH PROFESSIONAL |  | 2010-02
1822565-2010-00057 | 2010-02-09 | 2010-01-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRABECULAR METAL MODULAR ACETABULAR SYSTEM SHELL WITH CLUSTER HOLES POROUS | HIP PROSTHESIS | ZIMMER, INC. | NA | 00620205422 |  | SURGEON IMPLANTED MODULAR TRILOGY CUP. ONLY AFTER IMPLANTATION DID HE REALIZE THAT THE CUP WAS MISSING THE LOCKING RING. IT WAS NOT PRESENT AROUND THE CIRCUMFERENCE OF THE CUP AS IT SHOULD BE, NOR WAS IT EVIDENT IN THE PACKAGING. RATHER THAN REMOVE THE CUP AND CAUSE FURTHER TRAUMA TO THE PATIENT, THE SURGEON WAS FORCED TO OPEN A TM MODULAR MULTIHOLDED 54 SHELL, AND REMOVE THE LOCKING MECHANISM FROM THIS CUP, AND PUT IT INTO THE ONE INSIDE THE PATIENT. THE SURGEON WAS THEN ABLE TO IMPLANT A LINER AND COMPLETE THE OPERATION. SURGERY WAS DELAYED BY 40 MINUTES. | EVALUATION SUMMARY: ALL OF THE SHELLS IN THE LOT WERE SHOWN TO BE ASSEMBLED WITH THE APPROPRIATE SUBCOMPONENTS, INCLUDING THE LOCKING RINGS. PRODUCT SURVEILLANCE ATTEMPTED TO PERFORM A STOCK INVESTIGATION OF THE COMPLAINT LOT, HOWEVER, NO STOCK WAS AVAILABLE. THE STOCK INVESTIGATION WAS EXPANDED TO ITEMS PACKAGED BY THE SAME OPERATOR A MONTH PRIOR TO AND A MONTH FOLLOWING THE MANUFACTURE DATE OF THE COMPLAINT LOT. AGAIN, NO STOCK WAS AVAILABLE. LOCKING RINGS ARE SOLD INDIVIDUALLY IN CASE ONE IS BENT DURING THE SURGERY AND COMMON OPERATING ROOM PRACTICE IS TO HAVE ADDITIONAL IMPLANTS AVAILABLE FOR UNKNOWN PATIENT CONDITIONS. THIRTEEN OTHER IMPLANTS WERE IDENTIFIED AS POTENTIAL DONOR IMPLANTS (PACKAGED WITH THE SAME LOCKING RING). THEREFORE, A DONOR LOCKING RING WILL MOST LIKELY BE AVAILABLE DURING SURGERY. ZIMMER PRODUCT SURVEILLANCE HAS NOT RECEIVED ANY REPORTS OF MISSING LOCKING RINGS OF ACETABULAR SHELLS MANUFACTURED AROUND THE SAME TIME AS THE COMPLAINT LOT. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | N | Y | HEALTH PROFESSIONAL |  | 2010-02
9613350-2010-00017 | 2010-02-10 | 2009-07-28 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 60/54 T | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.160 |  | IT IS REPORTED BY PT'S ATTORNEY THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) /2009. | THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN 07/2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. OUR CASE (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00055 | 2010-02-10 | 2009-10-08 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | UNKNOWN |  | IT IS REPORTED BY PT THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2009. | THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN 07/2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. OUR CASE (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00053 | 2010-02-10 | 2010-01-11 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 56/50 P | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.156 |  | IT IS REPORTED BY PT'S ATTORNEY THAT PT UNDERWENT A LEFT HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION IS SCHEDULED. | THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN 07/2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. OUR CASE (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00025 | 2010-02-10 | 2010-01-15 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.154 |  | IT IS REPORTED THAT THE DUROM CUP WAS LOOSE. METAL DEBRIS PRESENT. CUP REMOVED AND REPLACED WITH TM CUP AND 36MM LINER/HEAD. | NO PRODUCT WAS RETURNED FOR EVAL. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF A DEVICE DEFICIENCY CAUSING OR CONTRIBUTING TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADD'L INFO BE RECEIVED THAT CHANGES THIS CONCLUSION, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. OUR CASE (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00023 | 2010-02-10 | 2010-01-15 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.154 |  | IT IS REPORTED BY PT THAT PT UNDERWENT A LEFT HIP ARTHROPLASTY IN (B) (6) 2008. POST-OP, SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION IS SCHEDULED FOR (B) (6) 2010. | THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN 07/2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. OUR CASE (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00019 | 2010-02-10 | 2009-12-14 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 62/56 V | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.162 |  | IT IS REPORTED BY PT'S ATTORNEY THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2006. POST-OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2009. | THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN 07/2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. OUR CASE (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00021 | 2010-02-10 | 2009-09-30 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 52/46 L | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.002143152 |  | IT IS REPORTED BY PT'S ATTORNEY THAT PT UNDERWENT A RIGHT HIP ARTHROPLASTY ON (B) (6) 2008. POST-OP, SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2008. | THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN 07/2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. OUR CASE (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00051 | 2010-02-10 | 2009-09-30 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 52/46 L | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.152 |  | IT IS REPORTED BY PT THAT PT UNDERWENT A RIGHT HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2009. | THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN 07/2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. OUR CASE (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00054 | 2010-02-10 | 2010-01-01 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 50/44 J | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.150 |  | IT WAS REPORTED THAT PT WAS DX WITH ACETABULAR IMPLANT LOOSENING WHICH RESULTED IN A COMPLETE REVISION. | THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN 07/2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. OUR CASE (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00056 | 2010-02-10 | 2010-01-19 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.154 |  | IT IS REPORTED BY PT'S ATTORNEY THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, SHE EXPERIENCES PAIN AND WILL BE REVISED AT AN UNK DATE. | THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN 07/2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. OUR CASE (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-02
1822565-2010-00062 | 2010-02-10 | 2010-01-12 | I, n, j, u, r, y | Manufacturer report | ZMR HIP SYSTEM FEMORAL STEM REVISION TAPER W/COMPRESSION NUT | HIP PROSTHESIS | ZIMMER, INC. | NA | 00998201718 |  | THE PATIENT WAS REVISED DUE TO THE TAPER STEM AND CONE BODY FRACTURING. | EVALUATION SUMMARY: PATIENT'S WEIGHT AND ACTIVITY LEVEL WERE NOT PROVIDED, PRE-REVISION X-RAYS WERE NOT AVAILABLE FOR EVALUATION, AND THE STEM WAS NOT RETURNED, MAKING ASSESSEMENT OF BONE ON-GROWTH AND PROXIMAL SUPPORT IMPOSSIBLE. THE PACKAGE INSERT AND/OR SURGICAL TECHNIQUE CONTAIN STATEMENTS WARNING THAT DEVICE FRACTURES MAY OCCUR WHERE EXCESSIVE PATIENT'S WEIGHT, EXCESSIVE PATIENT'S ACTIVITY, AND/OR LACK OF PROXIMAL SUPPORT ARE INVOLVED. THE ROOT CAUSE OF THIS COMPLAINT CANNOT BE DETERMINED FROM THE INFORMATION PROVIDED. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE RE-OPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-02
9613350-2010-00018 | 2010-02-10 | 2009-09-02 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 60/54 T | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.160 |  | IT IS REPORTED BY PT'S ATTORNEY THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION WAS PERFORMED ON (B) (6) 2009. | THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN 07/2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. OUR CASE (B) () | Y | N | LAY USER/PATIENT |  | 2010-02
1822565-2010-00063 | 2010-02-11 | 2007-07-31 | I, n, j, u, r, y | Manufacturer report | UNKNOWN VERSYS HIP SYSTEM STEM | HIP PROSTHESIS | ZIMMER, INC. | NA | UNK |  | IT IS REPORTED BY THE PATIENT'S ATTORNEY THAT THE PATIENT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2006. POST-OP, HE EXPERIENCED PAIN DUE TO AN INFECTION AND WAS REVISED. | EVALUATION SUMMARY: THE STERILIZATION PROCESS FOR THIS DEVICE WAS VALIDATED IN ACCORDANCE WITH FDA REGULATIONS AND ISO STANDARDS TO A STERILITY ASSURANCE LEVEL OF 1.0 X 10-6 OR BETTER. IN ADDITION, BEFORE EACH MANUFACTURING LOT IS RELEASED, THE "CERTIFICATE OF PROCESSING" FROM (B) (4) (STERILIZATION SUPPLIER) IS REVIEWED FOR CONFORMANCE. THIS CERTIFICATE LISTS THE MAXIMUM, MINIMUM, AND ACTUAL GAMMA DOSE IN KGY. THEREFORE, IT IS HIGHLY UNLIKELY THAT THE SPECIFIED DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT INFECTION. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS WAS ALSO NOT POSSIBLE AS THE PRODUCT AND/OR LOT NUMBERS REQUIRED FOR RETRIEVAL WERE UNAVAILABLE. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE AVAILABLE INFORMATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-02
9613350-2010-00052 | 2010-02-11 | 2007-07-31 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | UNK |  | IT IS REPORTED BY PATIENT'S ATTORNEY THAT PATIENT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2006. POST-OP, HE EXPERIENCED PAIN DUE TO AN INFECTION AND THE ACETABULAR COMPONENT WAS REMOVED ON (B) (6) 2007. | NO PRODUCT WA RETURNED FOR EVALUATION. THE INFORMATION REPORTED BY THE OPERATING SURGEON INDICATED THAT THE REVISION SURGERY WAS NOT RELATED TO OR CAUSED BY THE ZIMMER DEVICES. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF A DEVICE DEFICIENCY CAUSING OR CONTRIBUTING TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL INFORMATION BE RECEIVED THAT CHANGES THIS CONCLUSION, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-02
1822565-2010-00065 | 2010-02-15 | 2010-01-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY ACETABULAR SYSTEM CUP POSITIONER | HIP INSTRUMENT | ZIMMER, INC. | NA | 00626001800 |  | IT IS REPORTED THAT THE SURGEON WAS SCREWING THE ACETABULAR SHELL ONTO THE HANDLE WHEN THE END OF THE HANDLE BROKE OFF IN HIS HAND. | EVAL SUMMARY: VISUAL INSPECTION OF THE RETURNED PARTS DETERMINED THAT THERE WAS DEFORMATION OF THE LEADING THREADS AT THE TIP WHERE THE DEVICE MATES TO THE ACETABULAR SHELLS, BUT A FUNCTIONAL CHECK DETERMINED THAT THE THREADS STILL MATED PROPERLY WITH TRILOGY/TM MODULAR SHELLS AND COULD STILL BE USED. IT WAS REPORTED THAT THE IMPACTION CAP DETACHED WHILE SCREWING THE DEVICE INTO A SHELL. CONSIDERING THE RELATIVELY LOW TORQUE FORCES APPLIED TO THREAD THE SHELL ON, IT IS OKAY THAT THERE WAS/WERE FATIGUE CRACK(S) ALREADY PRESENT ON/IN THE DEVICE BEFORE THIS OPERATION. THE DEVICE MAY HAVE BEEN DAMAGED WHILE NOT IN USE (I.E. DROPPED DURING CLEANING), WHILE BEING USED IN A NON-RECOMMENDED FASHION, OR MULTIPLE OTHER REASONS, BUT THE CAUSE CANNOT BE DETERMINED BECAUSE THE USE HISTORY OF THE DEVICE IS NOT AVAILABLE. EVAL: AS RETURNED, THE WELD HOLDING THE CAP (IMPACTION SURFACE) ON THE INSTRUMENT HAS FAILED. THE IMPACTION SURFACE HAS EVIDENCE OF BEING WELL USED OVER THE LIFETIME OF THE DEVICE (5.5 YEARS). ADDITIONALLY, THE LEADING THREADS THAT ENGAGE WITH THE ACETABULAR SHELL HAVE BEEN DAMAGED. THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. | N | Y | HEALTH PROFESSIONAL |  | 2010-02
1822565-2010-00078 | 2010-02-16 | 2009-04-15 | I, n, j, u, r, y | Manufacturer report | VERSYS HIP SYSTEM EPOCH FULLCOAT HIP PROSTHESIS FEMORAL STEM FIBER METAL | HIP PROSTHESIS | ZIMMER, INC. | NA | 00408801506 |  | FAT EMBOLISM SYNDROME OCCURRED 12 HOURS AFTER SURGERY LEADING TO ALTERED MENTAL STATUS AND MILD HYPOXIA FOR THREE DAYS. | EVALUATION SUMMARY: IT IS UNKNOWN IF THE FEMORAL CANAL WAS PREPARED AS ACCORDING TO THE SURGICAL TECHNIQUE. THE INSTRUMENTS USED ARE UNKNOWN, AND IT IS UNKNOWN IF THE STEM WAS INSERTED WITH THE CORRECT ORIENTATION AND FIT AS PER THE SURGICAL TECHNIQUE. PATIENT FACTORS SUCH AS BONE QUALITY ARE ALSO UNKNOWN. FAT EMBOLISM SYNDROME (FES) IS A KNOWN COMPLICATION THAT MAY ARISE WHEN THERE IS A CHANCE FOR FAT TO ENTER THE CIRCULATORY SYSTEM, SUCH AS A LONG-BONE FRACTURE OR TOTAL HIP ARTHROPLASTY. RASPING TECHNIQUE, LAVAGE, ANESTHESIA, AND PATIENT FACTORS ARE ALL FACTORS THAT MAY CONTRIBUTE TO FES. TO DECREASE THE POTENTIAL FOR FES, PHARMACEUTICAL HEPARIN AND PULSE LAVAGE MAY BE EMPLOYED. DUE TO LACK OF SUFFICIENT INFORMATION, THE PROBABLE CAUSE OF FES IS UNKNOWN. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-02
1822565-2010-00069 | 2010-02-16 | 2010-01-19 | O, t, h, e, r | Manufacturer report | KINECTIV TECHNOLOGY MODULAR NECK | HIP PROSTHESIS | ZIMMER, INC. | NA | 00784801200 |  | WHEN THE KINECTIV MODULAR NECK E WAS TAKEN OUT OF ITS PACKAGING, THE SURGEON NOTICED A MARK ON THE STEM TAPER PORTION OF THE MODULAR NECK. HE THEN TURNED THE MODULAR NECK OVER AND THE SAME TYPE OF MARK WAS ON THE OTHER SIDE AS WELL. SINCE NO REPLACEMENT WAS AVAILABLE, HE DECIDED TO IMPLANT IT. | EVAL SUMMARY: DURING SURGERY AND PRIOR TO IMPLANTATION, THE SURGEON NOTICED MARKS ON THE TAPER PORTION OF A KINECTIV MODULAR NECK E. DESPITE THE MARK, THE NECK WAS IMPLANTED EVEN THOUGH IT WAS ALLEGED THE IMPLANT MAY HAVE BEEN DAMAGED. THE DEVICE HAS A CONFORMING DEVICE HISTORY RECORD. THE DEVICE WAS NOT RETURNED AND PICTURES WERE PROVIDED. FROM THE PHOTOS ALONE, IT CANNOT BE DETERMINED IF THE MARK IS A SCRATCH OR IF IT IS A SUBSTANCE STUCK TO THE PART. THE PACKAGE INSERT PROVIDED WITH THE NECK IMPLANT PROVIDES THE FOLLOWING WARNING "DO NOT USE ANY COMPONENT IF DAMAGE IS FOUND OR CAUSED DURING SETUP OR INSERTION." HOWEVER, THE ROOT CAUSE OF THE MARK CANNOT BE DETERMINED WITH CERTAINTY AT THIS TIME BASED ON THE PROVIDED INFO. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFO BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | N | Y | HEALTH PROFESSIONAL |  | 2010-02
9613350-2009-00705 | 2010-02-16 | 2009-11-16 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 56/50 CODE P | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.056 |  | IT WAS REPORTED THAT PT WAS COMPLAINING ABOUT HIP PAIN. ACETABULAR CUP WAS LOOSE. A REVISION SURGERY WAS PERFORMED AT 12 MONTHS POST-OP. | THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETED. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
1822565-2010-00079 | 2010-02-16 | 2009-07-27 | I, n, j, u, r, y | Manufacturer report | VERSYS HIP SYSTEM EPOCH FULLCOAT HIP PROSTHESIS FEMORAL STEM FIBER METAL | HIP PROSTHESIS | ZIMMER, INC. | NA | 00408801426 |  | FAT EMBOLISM SYNDROME OCCURRED 12 HOURS AFTER SURGERY LEADING TO SEMI-COMATOSE STATE FOR 4 DAYS. | EVALUATION SUMMARY: IT IS UNKNOWN IF THE FEMORAL CANAL WAS PREPARED AS ACCORDING TO THE SURGICAL TECHNIQUE. THE INSTRUMENTS USED ARE UNKNOWN, AND IT IS UNKNOWN IF THE STEM WAS INSERTED WITH THE CORRECT ORIENTATION AND FIT AS PER THE SURGICAL TECHNIQUE. PATIENT FACTORS SUCH AS BONE QUALITY ARE ALSO UNKNOWN. FAT EMBOLISM SYNDROME (FES) IS A KNOWN COMPLICATION THAT MAY ARISE WHEN THERE IS A CHANCE FOR FAT TO ENTER THE CIRCULATORY SYSTEM, SUCH AS A LONG-BONE FRACTURE OR TOTAL HIP ARTHROPLASTY. RASPING TECHNIQUE, LAVAGE, ANESTHESIA, AND PATIENT FACTORS ARE ALL FACTORS THAT MAY CONTRIBUTE TO FES. TO DECREASE THE POTENTIAL FOR FES, PHARMACEUTICAL HEPARIN AND PULSE LAVAGE MAY BE EMPLOYED. DUE TO LACK OF SUFFICIENT INFORMATION, THE PROBABLE CAUSE OF FES IS UNKNOWN. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-02
2648920-2010-00005 | 2010-02-16 | 2010-01-15 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY ACETABULAR SYSTEM SHELL WITH CLUSTER HOLES POROUS | HIP PROSTHESIS | ZIMMER MANUFACTURING B.V. | NA | 00620006022 |  | THE SURGEON PREPARED THE ACETABULUM FOR A 60MM TRILOGY CLUSTER SHELL #6200-60-22. HE IMPLANTED A 60 SHELL, LOT #61112456. HE INSERTED THE SCREWS AND THEN TOOK AN X-RAY TO VERIFY THE CUP POSITION. THE LOCKING RING VISIBLE IN THE X-RAY WAS MALALIGNED AND AS A RESULT, THE LINER WOULD NOT ENGAGE APPROPRIATELY. THE LOCKING RING BECAME DEFORMED AND A SECOND CUP HAD TO BE OPENED (LOT 61340489) AS REPLACEMENT RINGS WERE NOT AVAILABLE THROUGH THE DISTRIBUTOR OFFICE. SURGERY WAS PROLONGED BY 45 MINUTES. | EVALUATION SUMMARY: THE DEFORMED RING WAS NOT RETURNED. THE INTRA-OPERATIVE X-RAY SHOWS THE RING NOT SEATED IN THE MATING GROOVE IN THE SHELL. THE RETURNED SHELL FROM WHICH THE RING WAS SWAPPED SHOWS NO SIGNS OF DAMAGE. IT IS UNKNOWN IF THE INSTRUCTIONS ON IMPLANTING THE SHELL; ASSEMBLING AND IMPACTING THE LINER INTO THE SHELL DESCRIBED IN THE SURGICAL TECHNIQUE WERE FOLLOWED. THE RING MAY HAVE BECOME DISLODGED DURING IMPLANTATION OF THE SHELL AND SUBSEQUENTLY DAMAGED UPON ATTEMPTING TO ASSEMBLE THE LINER INTO THE SHELL. ZIMMER DOES NOT OFFER REPLACEMENT LOCKING RINGS AS A STOCKED PART NUMBER. BASED ON THE INFORMATION AVAILABLE, THE DEFINITIVE CAUSE CANNOT BE DETERMINED. . EVALUATION CODES: REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | N | Y | HEALTH PROFESSIONAL |  | 2010-02
1822565-2010-00080 | 2010-02-16 | 2009-02-23 | I, n, j, u, r, y | Manufacturer report | VERSYS HIP SYSTEM EPOCH FULLOCAT HIP PROSTHESIS FEMORAL STEM FIBER METAL | HIP PROSTHESIS | ZIMMER, INC. | NA | 00408801606 |  | FAT EMBOLISM SYNDROME OCCURRED 12 HOURS AFTER SURGERY LEADING TO A COMATOSE STATE FOR 5 DAYS AND SLOW NEUROLOGIC IMPROVEMENTS FOR AN ADDITIONAL SEVERAL WEEKS. | EVALUATION SUMMARY: IT IS UNKNOWN IF THE FEMORAL CANAL WAS PREPARED AS ACCORDING TO THE SURGICAL TECHNIQUE. THE INSTRUMENTS USED ARE UNKNOWN, AND IT IS UNKNOWN IF THE STEM WAS INSERTED WITH THE CORRECT ORIENTATION AND FIT AS PER THE SURGICAL TECHNIQUE. PATIENT FACTORS SUCH AS BONE QUALITY ARE ALSO UNKNOWN. FAT EMBOLISM SYNDROME (FES) IS A KNOWN COMPLICATION THAT MAY ARISE WHEN THERE IS A CHANCE FOR FAT TO ENTER THE CIRCULATORY SYSTEM, SUCH AS A LONG-BONE FRACTURE OR TOTAL HIP ARTHROPLASTY. RASPING TECHNIQUE, LAVAGE, ANESTHESIA, AND PATIENT FACTORS ARE ALL FACTORS THAT MAY CONTRIBUTE TO FES. TO DECREASE THE POTENTIAL FOR FES, PHARMACEUTICAL HEPARIN AND PULSE LAVAGE MAY BE EMPLOYED. DUE TO LACK OF SUFFICIENT INFORMATION, THE PROBABLE CAUSE OF FES IS UNKNOWN. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-02
1822565-2010-00081 | 2010-02-16 | 2007-08-15 | I, n, j, u, r, y | Manufacturer report | VERSYS HIP SYSTEM EPOCH FULLCOAT HIP PROSTHESIS FEMORAL STEM FIBER METAL | HIP PROSTHESIS | ZIMMER, INC. | NA | 00408801726 |  | FAT EMBOLISM SYNDROME OCCURRED IMMEDIATELY POST OPERATIVELY, LEADING TO COMATOSE STATE FOR ONE WEEK. | EVALUATION SUMMARY: IT IS UNKNOWN IF THE FEMORAL CANAL WAS PREPARED AS ACCORDING TO THE SURGICAL TECHNIQUE. THE INSTRUMENTS USED ARE UNKNOWN, AND IT IS UNKNOWN IF THE STEM WAS INSERTED WITH THE CORRECT ORIENTATION AND FIT AS PER THE SURGICAL TECHNIQUE. PATIENT FACTORS SUCH AS BONE QUALITY ARE ALSO UNKNOWN. FAT EMBOLISM SYNDROME (FES) IS A KNOWN COMPLICATION THAT MAY ARISE WHEN THERE IS A CHANCE FOR FAT TO ENTER THE CIRCULATORY SYSTEM, SUCH AS A LONG-BONE FRACTURE OR TOTAL HIP ARTHROPLASTY. RASPING TECHNIQUE, LAVAGE, ANESTHESIA, AND PATIENT FACTORS ARE ALL FACTORS THAT MAY CONTRIBUTE TO FES. TO DECREASE THE POTENTIAL FOR FES, PHARMACEUTICAL HEPARIN AND PULSE LAVAGE MAY BE EMPLOYED. DUE TO LACK OF SUFFICIENT INFORMATION, THE PROBABLE CAUSE OF FES IS UNKNOWN. : EVALUATION: NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-02
1526350-2010-00011 | 2010-02-18 | 2010-01-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | PULSAVAC PLUS HIP KIT | PULSAVAC PLUS HIP KIT | ZIMMER ORTHOPAEDIC SURGICAL PRODUCTS | NA | 00-5150-482-00 |  | IT WAS REPORTED THAT THE PULSAVAC PLUS UNIT GOT HOT DURING NORMAL USE. THERE WAS NO REPORT OF INJURY OR ADVERSE PT CONSEQUENCES ASSOCIATED WITH THIS INCIDENT. | THE DEVICE WAS NOT RETURNED TO THE MFR FOR EVAL. IF THE DEVICE SHOULD BE RETURNED, A F/U MEDWATCH REPORT WILL BE SUBMITTED. | N | Y | HEALTH PROFESSIONAL |  | 2010-02
9613350-2010-00061 | 2010-02-19 | 2009-08-11 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 56/50 P | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.156 |  | IT IS REPORTED BY THE PATIENT'S ATTORNEY THAT THE PATIENT UNDERWENT A LEFT HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2009. | THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00068 | 2010-02-19 | 2010-01-22 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 50/44 J | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.150 |  | IT IS REPORTED BY THE PATIENT'S ATTORNEY THAT THE PATIENT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. A REVISION DATE IS UNKNOWN. | THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00067 | 2010-02-19 | 2009-05-01 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 56/50 P | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.156 |  | IT IS REPORTED BY THE PATIENT'S ATTORNEY THAT THE PATIENT UNDERWENT A LEFT HIP ARTHROPLASTY ON (B) (6) 2008. POST-OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION TOOK PLACE IN (B) (6) 2009. | THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00066 | 2010-02-19 | 2010-01-22 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 56/50 P | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.156 |  | IT IS REPORTED BY THE PATIENT'S ATTORNEY THAT THE PATIENT UNDERWENT A RIGHT HIP ARTHROPLASTY ON (B) (6) 2008. POST-OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION TOOK PLACE ON (B) (6) 2009. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00064 | 2010-02-19 | 2010-01-22 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 52/46 L | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.152 |  | IT IS REPORTED BY THE PATIENT'S ATTORNEY THAT THE PATIENT UNDERWENT A LEFT HIP ARTHROPLASTY ON (B) (6) 2006. POST-OP, SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. A REVISION DATE IS UNKNOWN. | THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00063 | 2010-02-19 | 2009-09-01 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 56/50 P | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.156 |  | IT IS REPORTED THAT THE PATIENT PRESENTED AND WAS DX WITH ACETABULAR IMPLANT LOOSENING WHICH RESULTED IN REVISION. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00062 | 2010-02-19 | 2010-01-20 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 56/50 P | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.156 |  | IT IS REPORTED BY PATIENT THAT HE UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, HE EXPERIENCES PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION IS SCHEDULED FOR (B) (6) 2010. | THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00065 | 2010-02-19 | 2010-01-22 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 52/46 L | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.152 |  | IT IS REPORTED BY THE PATIENT'S ATTORNEY THAT THE PATIENT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2008. POST-OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION HAS NOT BEEN SCHEDULED. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00756 | 2010-02-23 | 2008-02-12 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 60/54 T | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.160 |  | IT WAS REPORTED THAT PT PRESENTED WITH ACETABULAR IMPLANT LOOSENING, WHICH RESULTED IN AN ACETABULAR REVISION ON (B)(6) 2008. | INFO WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC, WHICH MARKETS THE DEVICES IN THE US. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADD¿L INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B)(4) CONSIDERS THIS CASE CLOSED. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. (B)(4). | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00278 | 2010-02-23 | 2006-12-01 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 50/44 CODE J | DUROM HIP RESURFACING SYSTEM, FEMORAL C | ZIMMER GMBH |  | 01.00214.050 |  | IT WAS REPORTED THAT THERE WAS A POOR OSTEOINTEGRATION ON THE ACETABULAR SIDE. | (B) (4). CONCLUSION: BASED ON THE INVESTIGATION PERFORMED, THE REASON FOR THE POOR OSTEOINTEGRATION COULD NOT BE EVALUATED. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00273 | 2010-02-23 | 2006-12-07 | I, n, j, u, r, y | Manufacturer report | DUROM FEMORAL COMPONENT 52 CODE R | DUROM HIP RESURFACING SYSTEM, FEMORAL C | ZIMMER GMBH |  | 01.00211.152 |  | IT WAS REPORTED THAT THE DUROM FEMUR COMPONENT WAS LOOSENING AND PT WAS IN PAIN. | (B) (4). CONCLUSION: THE INVESTIGATION SHOWS THAT THERE IS NO SIGN OF A MATERIAL OR PRODUCTION FAILURE. FEMUR COMPONENT WAS IMPLANTED WITH TOO MUCH CEMENT. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00288 | 2010-02-23 | 2007-06-22 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 48/42 CODE H | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH |  | 01.00214.048 |  | IT WAS REPORTED THAT PATIENT WAS IN SEVERE PAIN, ACTIVE AND IN REST. | (B) (4). CONCLUSION: THE INVESTIGATION SHOWS THAT THERE IS NO SIGN OF A PRODUCTION OR MATERIAL FAILURE. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00289 | 2010-02-23 | 2007-06-22 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 52/46 CODE L | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH |  | 01.00214.052 |  | IT WAS REPORTED THAT SURGEON REAMED TO CUP SIZE, BUT WHEN HE TRIED THE 52 AND 54 SIZE BOTH WERE TOO LOOSE. | (B) (4). CONCLUSION: FOR BOTH PRODUCTS NO DEVIATION FROM THE PRODUCTION PROCESS WAS FOUND. A RE-MEASUREMENT SHOWS THAT BOTH CUPS HAVE THE CORRECT SIZE ACCORDING TO THE LASER MARKING. IT IS NOT POSSIBLE TO FINALLY DETERMINE THE CAUSE FOR THE MISSING PRIMARY FIXATION OF THE CUPS. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00290 | 2010-02-23 | 2007-08-13 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 50/44 CODE J | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH |  | 01.00214.050 |  | IT WAS REPORTED THAT THE DUROM COMPONENTS WERE REVISED DUE TO TISSUE INFLAMMATORY. | CONCLUSION: THE INVESTIGATION SHOWS THAT THERE IS NO SIGN OF A MATERIAL OR PRODUCTION FAILURE. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00487 | 2010-02-23 | 2009-04-16 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 48/42 CODE H | DUROM ACETABULAR COMPONENT | ZIMMER GMBH |  | 01.00214.048 |  | IT WAS REPORTED THAT PT WAS EXPERIENCING PAIN. SURGEON'S OBSERVATION ALSO REVEALED CYST/PSEUDOTUMOR. | CONCLUSION: THE QUALITY RECORDS INDICATE THAT ALL THE SPECIFIED CHARACTERISTICS (MATERIAL, MEASUREMENTS, SURFACE, AND SO ON) WERE IN CONFORMITY WITH THE REQUIREMENTS IN FORCE AT THE TIME OF MFG. THE INVESTIGATION INCLUDING THE EVAL OF THE DEVICE SHOWS THAT THERE IS NO SIGN OF A MATERIAL OR PRODUCTION FAILURE. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00284 | 2010-02-23 | 2007-03-01 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 48/42 CODE H | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH |  | 01.00214.048 |  | IT WAS REPORTED THAT PT WAS EXPERIENCING PAIN. | (B) (4). CONCLUSION: OUR INVESTIGATION SHOWS THAT THE WEAR RATES CAN BE CONSIDERED LOW. THERE IS NO SIGN OF A PRODUCTION OR MATERIAL FAILURE. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00423 | 2010-02-23 | 2009-06-11 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | UNK |  | IT WAS REPORTED BY PT'S COUNSEL THAT PT UNDERWENT TOTAL HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, PT EXPERIENCED PAIN AND LOOSENING OF THE CUP. HER SURGEON HAS RECOMMENDED A REVISION, WHICH IS NOT SCHEDULED AT THIS TIME. | (B) (4). THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00491 | 2010-02-23 | 2009-06-04 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 54/48 CODE N | DUROM ACETABULAR COMPONENT | ZIMMER GMBH |  | 01/00214.054 |  | IT WAS REPORTED THAT PT WAS REVISED DUE TO PAIN. AT TIME OF REVISION SURGERY, DUROM CUP WAS REMOVED WITH USE OF PUNCH ON RIM AND IT CAME OUT. ZERO BONE INGROWTH ON CUP. THERE WAS SIGNIFICANT BLACK CORROSION/RESIDUE AT THE HEAD/ADAPTER JUNCTION. | CONCLUSION: NO PRODUCT WAS RETURNED FOR EVAL. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF A DEVICE DEFICIENCY CAUSING OR CONTRIBUTING TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL INFO BE RECEIVED THAT CHANGES THIS CONCLUSION, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00247 | 2010-02-23 | 2006-09-22 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 60/54 CODE T | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH |  | 01.00214.060 |  | IT WAS REPORTED THAT PT WAS REVISED DUE TO A FEMORAL NECK FRACTURE IN THE FOURTH POSTOPERATIVE WEEK. | (B) (4). CONCLUSION: THE INVESTIGATION SHOWS THAT THERE IS NO SIGN OF A PRODUCTION OR MATERIAL FAILURE. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00244 | 2010-02-23 | 2006-06-05 | I, n, j, u, r, y | Manufacturer report | DUROM FEMORAL COMPONENT 52 CODE R | DUROM HIP RESURFACING SYSTEM, FEMORAL C | ZIMMER GMBH |  | 01.00211.152 |  | IT WAS REPORTED THAT THE HEAD COMPONENT AND THE SHAFT BROKE OFF. | (B) (4). CONCLUSION: BASED ON THE APPEARANCE OF THE EXPLANTS, IT HAS TO BE ASSUMED THAT THE FRACTURE OF THE SLIM STEM OF THE FEMORAL COMPONENT CAN BE ASCRIBED TO OVERLOAD. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00485 | 2010-02-23 | 2009-04-16 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 50/44 CODE J | DUROM ACETABULAR COMPONENT | ZIMMER GMBH |  | 01.00214.050 |  | IT WAS REPORTED THAT PT WAS EXPERIENCING PAIN. SURGEON'S OBSERVATION ALSO REVEALED CYST/PSEUDOTUMOR. | CONCLUSION: THE QUALITY RECORDS INDICATE THAT ALL THE SPECIFIED CHARACTERISTICS (MATERIAL, MEASUREMENTS, SURFACE, AND SO ON) WERE IN CONFORMITY WITH THE REQUIREMENTS IN FORCE AT THE TIME OF MFG. THE INVESTIGATION INCLUDING THE EVALUATION OF THE DEVICE SHOWS THAT THERE IS NO SIGN OF A MATERIAL OR PRODUCTION FAILURE. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00490 | 2010-02-23 | 2009-04-13 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 48/42 CODE H | DUROM ACETABULAR COMPONENT | ZIMMER GMBH |  | 01.00214.048 |  | IT WAS REPORTED THAT AFTER A BLOOD TEST IT WAS SHOWED AN INCREMENT OF CHROMIUM AND COBALT IONS LEVELS IN WHOLE BLOOD. SO, IT WAS NECESSARY TO EXPLANT THE CUP AND THE HEAD. | THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETED. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPONENT FILES. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00474 | 2010-02-23 | 2009-02-04 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 60/54 CODE T | DUROM ACETABULAR COMPONENT | ZIMMER GMBH |  | 01.00214.060 |  | IT WAS REPORTED THAT THE DEVICE WAS IMPLANTED ON (B) (6) 2007. THERE WAS AN INITIAL GOOD RESULT. THERE WAS INCREASED PAIN 6 MONTHS POST-OP. THEREFORE, IT HAD TO BE REVISED ON (B) (6) 2009. THERE WAS NO BONY INGROWTH AND THE CUP WAS LOOSE. | THIS FOLLOW-UP REPORT IS BEING FILLED TO RELAY ADDITIONAL INFORMATION. CONCOMITANT MEDICAL PRODUCTS: ITEM: 290039100, CATALOG #: CLS SPOTORNO, STEM, 135°, UNCEMENTED, 10.0, TAPER 12/14 LOT #: 2366889; ITEM: 0100181540, CATALOG #: METASULÂ® LDHÂ®, HEAD, 54, CODE T, TAPER 18/20, LOT #: 2336402. NEW INFORMATION DOES NOT CHANGE THE INVESTIGATION-RESULTS OF THIS INCIDENT, A FOLLOW UP REPORT WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. THE NEED FOR CORRECTIVE MEASURES IS NOT INDICATED AND ZIMMER GMBH CONSIDERS THIS CASE AS CLOSED. ZIMMER BIOMET'S REFERENCE NUMBER OF THIS FILE IS (B)(4). | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00483 | 2010-02-23 | 2009-03-19 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 52/46 CODE L | DUROM ACETABULAR COMPONENT | ZIMMER GMBH |  | 01.00214.052 |  | IT WAS REPORTED THAT DUROM CUP HAD MERGED TO A VERTICAL POSITION. PATIENT WAS FEMALE HIGHLY MEDICATED. PER SURGEON, SHE WAS A PROBLEMATIC PATIENT WITH LONG MEDICAL HISTORY. | CONCLUSION: THE QUALITY RECORDS INDICATE THAT ALL THE SPECIFIED CHARACTERISTICS (MATERIAL, MEASUREMENTS, SURFACE, AND SO ON) WERE IN CONFORMITY WITH THE REQUIREMENTS IN FORCE AT THE TIME OF MANUFACTURING. THE INVESTIGATION INCLUDING THE EVALUATION OF THE DEVICE SHOWS THAT THERE IS NO SIGN OF A MATERIAL OR PRODUCTION FAILURE. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00484 | 2010-02-23 | 2009-03-16 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 56/50 CODE P | DUROM ACETABULAR COMPONENT | ZIMMER GMBH |  | 01.00214.056 |  | IT WAS REPORTED THAT THE PATIENT UNDERWENT REVISION FOR PAIN AND DISABILITY FROM PAIN. AT REVISION SURGERY, NO BONE INGROWTH WAS FOUND IN THE DUROM CUP (NO EXPLANT REQUIRED), AND CORROSION ON STEM TAPER MAY BE POSSIBLE. | CONCLUSION: NO PRODUCT WAS RETURNED FOR EVALUATION. REVIEW OF THE DEVICE HISTORY RECORDS WAS ALSO NOT POSSIBLE AS THE ITEM NUMBERS AND/OR LOT NUMBERS WERE NOT PROVIDED. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF A DEVICE DEFICIENCY CAUSING OR CONTRIBUTING TO THE REPORTED PROBLEM. BASED ON THE AVAILABLE INFORMATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL INFORMATION BE RECEIVED THAT CHANGES THIS CONCLUSION, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00486 | 2010-02-23 | 2009-04-16 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 50/44 CODE J | DUROM ACETABULAR COMPONENT | ZIMMER GMBH |  | 01.00214.050 |  | IT WAS REPORTED THAT PATIENT WAS EXPERIENCING PAIN. SURGEON'S OBSERVATION ALSO REVEALED CYST/PSEUDOTUMOR. | CONCLUSION: THE QUALITY RECORDS INDICATE THAT ALL THE SPECIFIED CHARACTERISTICS (MATERIAL, MEASUREMENTS, SURFACE AND SO ON) WERE IN CONFORMITY WITH THE REQUIREMENTS IN FORCE AT THE TIME OF MANUFACTURING. THE INVESTIGATION INCLUDING THE EVALUATION OF THE DEVICE SHOWS THAT THERE IS NO SIGN OF A MATERIAL OR PRODUCTION FAILURE. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00467 | 2010-02-23 | 2008-11-15 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 50/44 CODE J | DUROM ACETABULAR COMPONENT | ZIMMER GMBH |  | 01.00214.050 |  | IT WAS REPORTED THAT PATIENT WAS PRESENTED WITH PAIN AND SWELLING. REVISION PROCEDURE WAS REQUIRED TO REMOVE THE IMPLANTS. | CONCLUSION: NO PRODUCT WAS RETURNED FOR EVALUATION. REVIEW OF THE DEVICE HISTORY RECORDS WAS ALSO NOT POSSIBLE AS THE ITEM NUMBERS AND/OR LOT NUMBERS WERE NOT PROVIDED. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF A DEVICE DEFICIENCY CAUSING OR CONTRIBUTING TO THE REPORTED PROBLEM. BASED ON THE AVAILABLE INFORMATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL INFORMATION BE RECEIVED THAT CHANGES THIS CONCLUSION, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00468 | 2010-02-23 | 2008-10-28 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 50/44 CODE J | DUROM ACETABULAR COMPONENT | ZIMMER GMBH |  | 01.00214.050 |  | IT WAS REPORTED THAT REVISION SURGERY WAS REQUIRED. REASONS FOR REVISION SURGERY WERE DISLOCATION AND LOOSENING. | CONCLUSION: NO PRODUCT WAS RETURNED FOR EVALUATION. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF A DEVICE DEFICIENCY CAUSING OR CONTRIBUTING TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL INFORMATION BE RECEIVED THAT CHANGES THIS CONCLUSION, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00470 | 2010-02-23 | 2009-01-29 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 58/52 CODE R | DUROM ACETABULAR COMPONENT | ZIMMER GMBH |  | 01.00214.058 |  | IT WAS REPORTED THAT REVISION SURGERY TOOK PLACE DUE TO PAIN. | CONCLUSION: THE INVESTIGATION SHOWS THAT THERE IS NO SIGN OF A MATERIAL OR PRODUCTION FAILURE. THE QUALITY RECORDS INDICATE THAT ALL THE SPECIFIED CHARACTERISTICS (MATERIAL, MEASUREMENTS, SURFACE, AND SO ON) WERE IN CONFORMITY WITH THE REQUIREMENTS IN FORCE AT THE TIME OF MANUFACTURING. THE INVESTIGATION SHOWS THAT THERE IS NO SIGN OF A MATERIAL OR PRODUCTION FAILURE. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00471 | 2010-02-23 | 2009-01-20 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 60/54 CODE T | DUROM ACETABULAR COMPONENT | ZIMMER GMBH |  | 01.00214.060 |  | IT WAS REPORTED THAT REVISION SURGERY TOOK PLACE DUE TO PAIN. | CONCLUSION: NO PRODUCT WAS RETURNED FOR EVALUATION. REVIEW OF THE DEVICE HISTORY RECORDS WAS ALSO NOT POSSIBLE AS THE ITEM NUMBERS AND/OR LOT NUMBERS WERE NOT PROVIDED. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF A DEVICE DEFICIENCY CAUSING OR CONTRIBUTING TO THE REPORTED PROBLEM. BASED ON THE AVAILABLE INFORMATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL INFORMATION BE RECEIVED THAT CHANGES THIS CONCLUSION, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00463 | 2010-02-23 | 2008-09-29 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 52/46 CODE L | DUROM ACETABULAR COMPONENT | ZIMMER GMBH |  | 01.00214.052 |  | IT WAS REPORTED THAT REVISION HAS TAKEN PLACE DUE TO PAIN AND NO BONE INGROWTH. | CONCLUSION: NO PRODUCT WAS RETURNED FOR EVALUATION. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF A DEVICE DEFICIENCY CAUSING OR CONTRIBUTING TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL INFORMATION BE RECEIVED THAT CHANGES THIS CONCLUSION, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00464 | 2010-02-23 | 2008-07-29 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 52/46 CODE L | DUROM ACETABULAR COMPONENT | ZIMMER GMBH |  | 01.00214.052 |  | IT WAS REPORTED THAT PATIENT EXPERIENCED PERSISTING PAIN IN THE GROIN WITH CYST FORMATION AND POSSIBLE ALLERGIC REACTION BASED ON METAL DEBRIS. | CONCLUSION: THE QUALITY RECORDS INDICATE THAT ALL THE SPECIFIED CHARACTERISTICS (MATERIAL, MEASUREMENTS, SURFACE, AND SO ON) WERE IN CONFORMITY WITH THE REQUIREMENTS IN FORCE AT THE TIME OF MANUFACTURING. THE WEAR RATES MEASURED ON THE ARTICULATIONS SURFACES CAN BE CONSIDERED LOW. DURING THE HISTOLOGICAL INVESTIGATION, ONLY A SMALL AMOUNT OF METALLIC PARTICLES IN THE TISSUE COULD BE FOUND. THE HISTOLOGY POINTED TO METAL SENSITIVITY. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00245 | 2010-02-23 | 2006-09-05 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | RESURFACING STYLUS CONNECTOR | DUROM HIP RESURFACING SYSTEM, FEMORAL C | ZIMMER GMBH |  | 01.00219.811 |  | IT WAS REPORTED THAT THE ADAPTER WOULD NOT FIT OVER THE CENTRING JIG. | (B) (4). CONCLUSION: OUR INTERNAL INVESTIGATIONS AND FUNCTION CHECKS HAVE SHOWN THAT THE STYLUS CONNECTOR COULD BE EASILY ATTACHED ONTO THE CENTERING JIG. WE CANNOT RECONSTRUCT THE COMPLAINT REASON. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. | N | Y | HEALTH PROFESSIONAL |  | 2010-02
9613350-2009-00465 | 2010-02-23 | 2008-10-27 | I, n, j, u, r, y | Manufacturer report | DUROM ACETAB. COMP. 58/52 CODE | DUROM ACETAB. COMP. 58/52 CODE | ZIMMER GMBH |  | 01.00214.058 |  | IT WAS REPORTED THAT THE LEFT DUROM SOCKET WAS LOOSENING. THE ML TAPER STEM WAS REMOVED BECAUSE OF A PROBABLE INFECTION. | (B) (4). CONCLUSION: NO PRODUCT WAS RETURNED FOR EVALUATION. REVIEW OF THE DEVICE HISTORY RECORDS WAS ALSO NOT POSSIBLE AS THE LOT NUMBERS WERE NOT PROVIDED. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF A DEVICE DEFICIENCY CAUSING OR CONTRIBUTING TO THE REPORTED PROBLEM. BASED ON THE AVAILABLE INFORMATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL INFORMATION BE RECEIVED THAT CHANGES THIS CONCLUSION, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. THIS MDR IS BEING SUBMITTED LATE, AS THE ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00466 | 2010-02-23 | 2008-11-05 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 60/54 CODE T | DUROM ACETABULAR COMPONENT | ZIMMER GMBH |  | 01.00214.060 |  | IT WAS REPORTED THAT REVISION SURGERY WAS REQUIRED DUE TO PAIN. | (B) (4). (B) (4). CONCLUSION: THE QUALITY RECORDS INDICATE THAT ALL THE SPECIFIED CHARACTERISTICS (MATERIAL, MEASUREMENTS, SURFACE, AND SO ON) WERE IN CONFORMITY WITH THE REQUIREMENTS IN FORCE AT THE TIME OF MANUFACTURING. SINCE NO PARTS WERE SENT TO US FOR INVESTIGATION, WE ARE NOT IN A POSITION TO TAKE A STAND CONCERNING THE OCCURRENCE REFERRED TO IN THE COMPLAINT. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00461 | 2010-02-23 | 2008-04-14 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 62/56 CODE V | DUROM ACETABULAR COMPONENT | ZIMMER GMBH |  | 01.00214.062 |  | IT WAS REPORTED THAT IN (B) (6) 2004 A DUROM CUP WITH A LARGE DIAMETER HEAD AND A CEMENTED EMERAUDE STEM WERE IMPLANTED. THE POSTOPERATIVE CHECK-UP SHOWED THAT THE PROSTHESIS WAS IN THE CORRECT POSITION. AT THE 2-MONTH CHECK-UP, THE X-RAYS SHOWED THAT THE CUP HAD TILTED INTO A VALGUS POSITION. THEREFORE, THE SURGEON RECOMMENDED A REVISION SURGERY BUT THE PATIENT WHO WAS VERY SATISFIED WITH THE RESULT, REPEATEDLY REJECTED THE REVISION SURGERY. THE PATIENT REPORTED A GOOD RESULT APART FROM AN OCCASIONAL NOISE AT CERTAIN MOVEMENTS. ON (B) (6) 2008, THE PATIENT CAME TO THE HOSPITAL WITHIN AN EMERGENCY DUE TO A FALL FROM THE ROOF AND A FEMORAL FRACTURE AROUND AND UNDER THE STEM. THE HIP PROSTHESIS WAS REVISED. | CONCLUSION: OUR INVESTIGATION SHOWED THAT DUE TO AN EARLY LOOSENING THE CUP HAD TILTED INTO A VALGUS POSITION. SINCE THE PATIENT REJECTED THE REVISION SURGERY THE PAIRING ARTICULATED IN A SUBLUXATED POSITION OVER A CERTAIN TIME PERIOD WHICH LED TO HIGH WEAR. THERE ARE NO SIGNS OF MATERIAL OR PRODUCTION FAILURE. THE QUALITY RECORDS INDICATE THAT ALL THE SPECIFIED CHARACTERISTICS (MATERIAL, MEASUREMENTS, SURFACE, AND SO ON) WERE IN CONFORMITY WITH THE REQUIREMENTS IN FORCE AT THE TIME OF MANUFACTURING. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00497 | 2010-02-23 | 2008-06-13 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 54/48 CODE N | DUROM ACETABULAR COMPONENT 54/48 CODE N | ZIMMER GMBH |  | 01.00214.054 |  | IT WAS REPORTED THAT THE PATIENT WAS REVISED. REASON FOR SURGERY IS UNKNOWN. | (B) (4). CONCLUSION: THE QUALITY RECORDS INDICATE THAT ALL SPECIFIED CHARACTERISTICS (MATERIAL, MEASUREMENTS, SURFACE AND SO ON) WERE IN CONFORMITY WITH THE REQUIREMENTS IN FORCE AT THE TIME OF MANUFACTURING. THE INVESTIGATION INCLUDING THE EVALUATION OF THE DEVICE SHOWS THAT THERE IS NO SIGN OF A MATERIAL OR PRODUCTION FAILURE. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. (B) (4). (WITHIN THE RECEIVED PRODUCT EXPERIENCE REPORT, WE RECEIVED TOW CASES OF TWO DIFFERENT PATIENTS. THIS IS CASE 2 OF 2 WITH THIS REC-NUMBER). | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00291 | 2010-02-23 | 2007-08-29 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 54/48 CODE N | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH |  | 01.00214.054 |  | IT WAS REPORTED THAT THE IMPLANT WAS A 56MM AND NOT A 54MM AS REQUESTED BY THE SURGEON. ANOTHER 54MM WAS OPENED AND IMPLANTED. | (B) (4). CONCLUSION: BOTH THESE LOTS WERE BLISTER-PACKAGED AND SEALED ON THE SAME DAY IN THE CLEAN ROOM. AFTER THE CLEAN ROOM PACKAGING PROCESS, BOTH LOTS WERE LABELED WITH BAR-CODE LABELS. A MIX-UP OCCURRED ON THE PACKAGING LINE WHILE LABELING THE TWO LOTS OF BLISTER-PACKS WHICH WERE EXCHANGED FOR EACH OTHER. PROCESS CORRECTIVE ACTIONS HAVE BEEN IDENTIFIED AND IMPLEMENTED TO PREVENT FURTHER ISSUES. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. | N | Y | OTHER |  | 2010-02
9613350-2009-00236 | 2010-02-23 | 2006-06-12 | I, n, j, u, r, y | Manufacturer report | DUROM FEMORAL COMPONENT 50 CODE P | DUROM HIP RESURFACING SYSTEM, FEMORAL C | ZIMMER GMBH |  | 01.00211.150 |  | IT WAS REPORTED THAT PATIENT EXPERIENCED PAIN SINCE HE HAD A DUROM FEMORAL CUP IMPLANTED. | (B) (4). CONCLUSION: THE INVESTIGATION SHOWS NO SIGNS OF A MATERIAL OR PRODUCTION FAILURE. IT IS NOT POSSIBLE TO EVALUATE IF THERE WAS BONE RESORPTION OR NOT, SINCE A STEAM STERILIZATION WAS CARRIED OUT BY THE HOSPITAL. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00280 | 2010-02-23 | 2007-03-06 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 54/48 CODE N | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH |  | 01.00214.054 |  | IT WAS REPORTED THAT 24 HRS AFTER OPERATION, ACETABULAR CUP MIGRATED SUPERIORLY, LOWER LIMB EXTENSIVE SWELLING COLOURED BLACK AND YELLOW. | (B) (4). CONCLUSION: NO PRODUCT WAS RETURNED FOR EVAL. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF A DEVICE DEFICIENCY CAUSING OR CONTRIBUTING TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL INFO BE RECEIVED THAT CHANGES THIS CONCLUSION, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00296 | 2010-02-23 | 2007-02-15 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 52/46 L | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH |  | 01.00214.152 |  | IT WAS REPORTED THAT THE DURCOM ACETABULAR CUP WAS LOOSE. | (B) (4). CONCLUSION: THE INVESTIGATION SHOWS THAT THERE IS NO SIGN OF A MATERIAL OR PRODUCTION FAILURE. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2009-00356 | 2010-02-23 | 2009-04-27 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 48/42 H | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.148 |  | IT WAS REPORTED BY PT THAT SHE UNDERWENT TOTAL HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, SHE EXPERIENCED PAIN AND HER SURGEON RECOMMENDED A REVISION OF THE DUROM ACETABULAR CUP. THE REVISION PROCEDURE IS PLANNED FOR (B) (6) 2009. | THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-02
1822565-2010-00102 | 2010-02-25 | 2010-01-29 | I, n, j, u, r, y | Manufacturer report | ZMR HIP SYSTEM FEMORAL STEM REVISION TAPER | HIP PROSTHESIS | ZIMMER, INC. | NA | 00998201818 |  | THE ZMR STEM WAS IMPLANTED APPROXIMATELY 5 YEARS AGO. THE STEM WAS FOUND FRACTURED AT THE JUNCTION AND REVISION WAS PERFORMED. THERE WAS NO PROXIMAL BONE AND THE PROXIMAL PART REMOVED EASILY, BUT THE DISTAL STEM WAS WELL-FIXED. | THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETED. | Y | N | HEALTH PROFESSIONAL |  | 2010-02
9613350-2010-00081 | 2010-02-25 | 2010-01-16 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 56/50 P | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.156 |  | IT IS REPORTED THAT A DUROM CUP WAS IMPLANTED ON (B) (6) 2009. THE CUP LOOSENED AND REVISION DONE ON (B) (6) 2010. | THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETED. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00029 | 2010-02-25 | 2010-01-29 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 52/46 L | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.152 |  | IT IS REPORTED BY PATIENT THAT PATIENT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2008. POST-OP, SHE EXPERIENCES PAIN DUET TO POSSIBLE ACETABULAR LOOSENING. A REVISION WAS SCHEDULED FOR (B) (6) 2010 AND POSTPONED AGAIN. | (B) (4). THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00028 | 2010-02-25 | 2010-01-27 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 50/44 J | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.150 |  | IT IS REPORTED BY PATIENT'S ATTORNEY THAT HE UNDERWENT A RIGHT HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. A REVISION IS SCHEDULED FOR (B) (6) 2010. | (B) (4). THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00027 | 2010-02-25 | 2010-01-27 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 50/44 J | DUROM ACETABULAR COMONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.150 |  | IT IS REPORTED BY PATIENT'S ATTORNEY THAT PATIENT UNDERWENT A LEFT HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. A REVISION IS SCHEDULED FOR (B) (6) OF 2010. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00080 | 2010-02-25 | 2009-11-23 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 48/42 H | DUROM ACETABULAR COMONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.148 |  | IT IS REPORTED BY PATIENT THAT PATIENT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2006. POST-OP, SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2009. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00079 | 2010-02-25 | 2009-11-30 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 58/52 R | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.158 |  | IT IS REPORTED BY PATIENT THAT PATIENT UNDERWENT A LEFT HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2009. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00078 | 2010-02-25 | 2009-09-02 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 50/44 J | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.150 |  | IT IS REPORTED BY PATIENT'S ATTORNEY THAT PATIENT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2008. POST-OP, SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2009. | (B) (4). THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4)CONSIDERS THIS CASE CLOSED. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00077 | 2010-02-25 | 2010-01-25 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.154 |  | IT IS REPORTED BY PATIENT THAT PATIENT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2006. POST-OP, SHE EXPERIENCES PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION IS SCHEDULED FOR (B) (6) 2010. | (B) (4). THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-02
1822565-2010-00105 | 2010-02-25 | 2010-01-26 | I, n, j, u, r, y | Manufacturer report | VERSYS HIP SYSTEM FEMORAL HEAD 12/14 TAPER CERAMIC | HIP PROSTHESIS | ZIMMER, INC. | NA | 00811802801 |  | PT PRESENTED WITH FRACTURED FEMORAL HEAD. | EVAL SUMMARY: THE DEVICE WAS IMPLANTED IN 2000, AND IN-VIVO APPROX 10 YEARS BEFORE THE ALLEGED EVENT OF FEMORAL HEAD FRACTURE. THE LOT OF THIS DEVICE IS WITHIN THE SCOPE OF A RECALL IN 2001 DUE TO THE FRACTURING OF THE FEMORAL HEAD. THE RECALL IS COMPLETE AND NO FURTHER MANUFACTURE OR MARKETING OF THE ZIRCONIA FEMORAL HEAD 8118 FAMILY OCCURS. THE DEVICE WAS IMPLANTED BEFORE THE RECALL TOOK EFFECT. NO CAPA IS REQUIRED. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADD'L SUBSTANTIVE INFO BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-02
9613350-2009-00724 | 2010-02-26 | 2009-11-16 | I, n, j, u, r, y | Manufacturer report | DUROM FEMORAL COMPONENT 44 CODE J | DUROM HIP RESURFACING SYS, FEMORAL COMPONE | ZIMMER GMBH |  | 01.00211.144 |  | IT IS REPORTED THAT PT WAS TAKEN TO HOSPITAL DUE TO FRACTURE OF THE FEMORAL NECK. | (B) (4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-02
9613350-2010-00100 | 2010-03-02 | 2009-05-06 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 56/50 P | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.156 |  | IT WAS REPORTED BY THE PT'S ATTORNEY THAT THE PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2008. POST OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2009 BY (B) (6). | INFO WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WHERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADD'L INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-03
9613350-2010-00069 | 2010-03-02 | 2008-12-30 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 52/46 L | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.152 |  | IT IS REPORTED BY THE PATIENT'S ATTORNEY THAT THE PATIENT UNDERWENT A RIGHT HIP ARTHROPLASTY ON (B) (6) 2007. POST OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2008. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN 07/08. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-03
9613350-2010-00103 | 2010-03-02 | 2010-02-15 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 48/42 H | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 0100214148 |  | IT WAS REPORTED BY THE PATIENT THAT SHE UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2007. POST OP, SHE EXPERIENCED PAIN DUE TO THE POSSIBLE ACETABULAR LOOSENING. A REVISION SURGERY IS SCHEDULED FOR (B) (6) 2010. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. (B) (4), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN 07/08. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-03
9613350-2010-00102 | 2010-03-02 | 2010-02-08 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 58/52 R | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.158 |  | IT IS REPORTED THAT X-RAYS SHOWED A LOOSE CUP. THE UP FACE WAS FACING FLAT TO THE FLOOR. AT THE REVISION, THE CUP CAME OUT VERY EASILY WITH NO BONE INGROWTH. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER (B) (4) ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN 07/08. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-03
9613350-2010-00101 | 2010-03-02 | 2007-05-22 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 50/44 J | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.150 |  | IT IS REPORTED BY PATIENT THAT PATIENT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2007. POST OP, SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2010 BY DR (B) (6) OF (B) (6). | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN 07/08. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-03
9613350-2010-00059 | 2010-03-02 | 2007-10-27 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 46/40 F | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.146 |  | IT IS REPORTED BY PT THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2007. | INFO WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WHERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADD'L INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-03
9613350-2010-00058 | 2010-03-02 | 2010-02-02 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 50/44 J | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.150 |  | IT IS REPORTED BY PT THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2006. POST-OP, SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION IS SCHEDULED FOR (B) (6) 2010. | INFO WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WHERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADD'L INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-03
9613350-2010-00057 | 2010-03-02 | 2007-03-06 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 62/56 V | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.162 |  | IT IS REPORTED BY PT THAT PT UNDERWENT A RIGHT HIP ARTHROPLASTY ON (B) (4) 2006. POST-OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (4) 2007. | INFO WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WHERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADD'L INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-03
9613350-2010-00070 | 2010-03-02 | 2010-02-08 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 58/52 R | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.158 |  | IT IS REPORTED THE PATIENT UNDERWENT A RIGHT HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION SURGERY IS SCHEDULED FOR (B) (6) 2010. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN 07/08. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-03
9613350-2010-00050 | 2010-03-02 | 2010-02-03 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.154 |  | IT IS REPORTED BY PT THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (4) 2008. POST-OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION IS BEING SCHEDULED. | INFO WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WHERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008 . SHOULD ADD'L INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-03
9613350-2010-00060 | 2010-03-02 | 2010-01-18 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | UNK |  | IT IS REPORTED BY PT THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, SHE EXPERIENCED PAIN DUE TO ACETABULAR LOOSENING AND A REVISION WILL BE SCHEDULED. | THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-03
9613350-2010-00099 | 2010-03-03 | 2009-12-08 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 48/42 H | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.148 |  | IT IS REPORTED BY PT'S ATTORNEY THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, SHE EXPERIENCES PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2009. | INFO WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-03
9613350-2010-00072 | 2010-03-03 | 2010-02-12 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 54/48N | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.154 |  | IT IS REPORTED BY THE PT THAT SHE UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2008. POST OP, SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. A REVISION IS SCHEDULED FOR (B) (6) 2010. | INFO WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-03
9613350-2010-00076 | 2010-03-03 | 2010-02-17 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 56/50 P | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.156 |  | IT WAS REPORTED BY THE PT'S ATTORNEY THAT THE PT UNDERWENT A LEFT HIP ARTHROPLASTY ON (B) (6) 2007. POST OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. A REVISION SURGERY HAS BEEN SCHEDULED. | INFO WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-03
9613350-2010-00083 | 2010-03-03 | 2009-12-03 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | UNK |  | IT WAS REPORTED BY THE PT'S ATTORNEY THAT SHE UNDERWENT A LEFT HIP ARTHROPLASTY ON (B) (6) 2008. POST OP, SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2009. | INFO WAS FORWARDED FORM THE OWNER ESTABLISHMENT, ZIMMER ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. NO PRODUCT WAS RETURNED FOR EVAL. REVIEW OF THE DEVICE HISTORY RECORDS WAS ALSO NOT POSSIBLE AS THE ITEM NUMBERS AND/OR LOT NUMBERS WERE NOT PROVIDED. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF A DEVICE DEFICIENCY CAUSING OR CONTRIBUTING TO THE REPORTED PROBLEM. BASED ON THE AVAILABLE INFO, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL INFO BE RECEIVED THAT CHANGES THIS CONCLUSION, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-03
1822565-2010-00125 | 2010-03-03 | 2010-02-01 | I, n, j, u, r, y | Manufacturer report | TRILOGY LONGEVITY HIGHLY CROSSLINKED POLYETHYLENE CONSTRAINED LINER | HIP PROSTHESIS | ZIMMER, INC. | NA | 00633405028 |  | DISLOCATED TWICE, FIRST TIME CLOSED REDUCTION, SECOND TIME REMOVED AND REPLACED WITH LARGER HEAD AND CONSTRAINED LINER. | EVALUATION SUMMARY: THE DEVICE WAS IN-VIVO APPROXIMATELY 6 WEEKS. THE TRILOGY LONGEVITY CONSTRAINED LINER WAS NOT AVAILABLE FOR ANALYSIS, AND ITS CONDITION IS UNKNOWN. THE MATING ACETABULAR SHELL, FEMORAL HEAD, AND FEMORAL STEM COMPONENTS ARE UNKNOWN, AND THEIR CONDITIONS ARE ALSO UNKNOWN. THE SURGICAL NOTES FROM THE PRIMARY SURGERY ARE UNAVAILABLE. IT IS UNKNOWN IF THE COMPONENTS WERE IMPLANTED WITH THE CORRECT ORIENTATION AS PER THE SURGICAL TECHNIQUE. X-RAY IMAGES POST-PRIMARY SURGERY AND FROM SUCCESSIVE PATIENT VISITS ARE UNAVAILABLE. THE INSTRUMENTS USED IN PRIMARY SURGERY ARE ALSO UNKNOWN. PATIENT ACTIVITY LEVEL IS UNKNOWN. THE REHABILITATION REGIMEN, BOTH PHYSICAL AND PHARMACOLOGICAL, AND COMPLIANCE THEREOF ARE UNKNOWN. BASED ON THE ABOVE INFORMATION AND OBSERVATIONS, DISLOCATION MAY HAVE BEEN DUE TO ONE OR A COMBINATION OF THE FOLLOWING MECHANISMS: UNSATISFACTORY PLACEMENT OF THE COMPONENT PARTS (SHELL AND/OR STEM). THIS MAY HAVE LED TO IMPINGEMENT (LEVERING) OF THE STEM AGAINST THE LINER, WHICH COULD LEAD TO FEMORAL HEAD/LINER DISLOCATION. UNSATISFACTORY LINER IMPLANTATION. IF THE LINER WAS NOT SEATED OR ORIENTATED AS PER THE SURGICAL TECHNIQUE, THE MISALIGNMENT OF THE LINER MAY ALSO LEAD TO IMPINGEMENT OF THE STEM AGAINST THE LINER. EXTENT OF SOFT-TISSUE CONSTRAINT. INCORRECT SOFT TISSUE TENSION AND/OR POOR FUNCTIONAL MUSCLES AROUND THE HIP MAY NOT BE ABLE TO PROVIDE FUNCTIONAL SUPPORT TO THE JOINT. PATIENT BEHAVIOR. HOWEVER, THE EXACT CAUSE OF THE CURRENT SITUATION CANNOT BE CONCLUSIVELY DETERMINED. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-03
9613350-2010-00084 | 2010-03-03 | 2010-02-17 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 56/50 P | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.156 |  | IT WAS REPORTED BY THE PT'S ATTORNEY THAT THE PT UNDERWENT A RIGHT HIP ARTHROPLASTY ON (B) (6) 2008. POST OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. A REVISION SURGERY HAS BEEN SCHEDULED. | INFO WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-03
9613350-2010-00085 | 2010-03-03 | 2010-02-17 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 52/46 L | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.152 |  | IT WAS REPORTED BY THE PT'S ATTORNEY THAT THE PT UNDERWENT A LEFT HIP ARTHROPLASTY ON (B) (6) 2008. POST OP, SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION SURGERY HAS BEEN SCHEDULED. | INFO WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-03
9613350-2010-00086 | 2010-03-03 | 2010-02-16 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 58/52 R | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.158 |  | IT WAS REPORTED BY THE PT THAT HE UNDERWENT A RIGHT HIP ARTHROPLASTY ON (B) (6) 2007. POST OP, HE EXPERIENCED PAIN DUE TO ACETABULAR LOOSENING AND A REVISION SURGERY HAS BEEN SCHEDULED. | INFO WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-03
9613350-2010-00087 | 2010-03-03 | 2010-02-16 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.154 |  | IT WAS REPORTED BY THE PT THAT HE UNDERWENT A LEFT HIP ARTHROPLASTY IN (B) (6) 2007. POST OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION SURGERY IS SCHEDULED. | INFO WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLANTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-03
9613350-2010-00088 | 2010-03-03 | 2015-04-06 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL LDH | ZIMMER GMBH |  | 01.00214.154 |  | IT HAS NOW BEEN REPORTED THAT THE PATIENT WAS IMPLANTED A DUROM US ACETABULAR COMPONENT 54/48 N ON THE RIGHT SIDE. THE PATIENT WAS REVISED ON (B)(6) 2015 DUE TO PAIN AND LOOSENING. | THIS CASE WAS REOPENED ON AUGUST 18, 2015 TO ENTER ADDITIONAL INFORMATION WHICH HAD BEEN RECEIVED ON AUGUST 14, 2015. SINCE THIS CASE IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A NOTIFICATION IN JULY 2008, ZIMMER (B)(4) WILL CLOSE THIS CASE ONCE AGAIN. ZIMMER'S REFERENCE NUMBER OF THIS FILE IS (B)(4). | Y | N | LAY USER/PATIENT |  | 2010-03
9613350-2010-00071 | 2010-03-03 | 2010-01-07 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 56/50 P | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.156 |  | IT IS REPORTED BY THE PT THAT HE UNDERWENT A LEFT HIP ARTHROPLASTY ON (B) (6) 2007. POST OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2010. | INFO WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-03
1822565-2010-00126 | 2010-03-03 | 2010-02-01 | I, n, j, u, r, y | Manufacturer report | TRILOGY LONGEVITY HIGHLY CROSSLINKED POLYETHYLENE CONSTRAINED LINER | HIP PROSTHESIS | ZIMMER, INC. | NA | 00633405028 |  | PATIENT DISLOCATED, CLOSED REDUCTION SURGERY. | EVALUATION SUMMARY: THE DEVICE WAS IN-VIVO APPROXIMATELY 4 MONTHS. THE TRILOGY LONGEVITY CONSTRAINED LINER WAS NOT AVAILABLE FOR ANALYSIS, AND ITS CONDITION IS UNKNOWN. THE MATING ACETABULAR SHELL, FEMORAL HEAD, AND FEMORAL STEM COMPONENTS ARE UNKNOWN, AND THEIR CONDITIONS ARE ALSO UNKNOWN. THE SURGICAL NOTES FROM THE PRIMARY SURGERY ARE UNAVAILABLE. THE COMPONENTS THAT WERE REVISED IN THE (B) (6) 2010 SURGERY ARE UNKNOWN. IT IS UNKNOWN IF THE COMPONENTS WERE IMPLANTED WITH THE CORRECT ORIENTATION AS PER THE SURGICAL TECHNIQUE. X-RAY IMAGES POST PRIMARY SURGERY AND FROM SUCCESSIVE PATIENT VISITS ARE UNAVAILABLE. THE INSTRUMENTS USED IN THE PRIMARY SURGERY ARE ALSO UNKNOWN. THE REHABILITATION REGIMEN, BOTH PHYSICAL AND PHARMACOLOGICAL, AND COMPLIANCE THEREOF ARE UNKNOWN. BASED ON THE ABOVE INFORMATION AND OBSERVATIONS, DISLOCATION MAY HAVE BEEN DUE TO ONE OR A COMBINATION OF THE FOLLOWING MECHANISMS: UNSATISFACTORY PLACEMENT OF THE COMPONENT PARTS (SHELL AND/OR STEM). THIS MAY HAVE LED TO IMPINGEMENT (LEVERING) OF THE STEM AGAINST THE LINER, WHICH COULD LEAD TO FEMORAL HEAD/LINER DISLOCATION. UNSATISFACTORY LINER IMPLANTATION. IF LINER WAS NOT SEATED OR ORIENTATED AS PER THE SURGICAL TECHNIQUE, THE MISALIGNMENT OF THE LINER MAY ALSO LEAD TO IMPINGEMENT OF THE STEM AGAINST THE LINER. EXTENT OF SOFT-TISSUE CONSTRAINT. INCORRECT SOFT TISSUE TENSION AND/OR POOR FUNCTIONAL MUSCLES AROUND THE HIP MAY PRECLUDE FUNCTIONAL SUPPORT TO THE JOINT. PATIENT BEHAVIOR. HOWEVER, THE EXACT CAUSE OF THE CURRENT SITUATION CANNOT BE CONCLUSIVELY DETERMINED. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-03
9613350-2010-00089 | 2010-03-03 | 2010-02-17 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 56/50 P | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.156 |  | IT IS REPORTED BY PT'S ATTORNEY THAT PT UNDERWENT A RIGHT HIP ARTHROPLASTY ON (B) (6) 2008. POST-OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. A REVISION IS SCHEDULED. | INFO WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-03
1822565-2010-00127 | 2010-03-03 | 2010-02-02 | I, n, j, u, r, y | Manufacturer report | VERSYS HIP SYSTEM REVISION FEMORAL STEM BEADED FULLCOAT PLUS | HIP PROSTHESIS | ZIMMER, INC. | NA | 00784301108 |  | VERSYS FULLCOAT FRACTURE WITH FEMORAL FRACTURE. | EVALUATION SUMMARY: THE RECEIVED X-RAY WAS REVIEWED. IT WAS OBSERVED THAT THE VERSYS BEADED FULLCOAT STEM WAS FRACTURED AT APPROXIMATELY MID-STEM. THE DEVICE WAS IN-VIVO APPROXIMATELY 6 YEARS. NO MANUFACTURING LOT INFORMATION WAS AVAILABLE FOR MANUFACTURING EVALUATION. THE DEVICE WAS NOT AVAILABLE FOR ANALYSIS. THE MATING ACETABULAR SHELL, LINER, AND FEMORAL HEAD COMPONENTS ARE UNKNOWN AND THEIR CONDITIONS ARE UNKNOWN. X-RAY IMAGES PRE-2002 SURGERY AND FROM SUCCESSIVE PATIENT VISITS ARE UNAVAILABLE. IT IS UNKNOWN IF THE FEMORAL STEM WAS IMPLANTED AS PER THE SURGICAL TECHNIQUE. PATIENT ACTIVITY IS ALSO UNKNOWN. BASED ON THE ABOVE INFORMATION AND OBSERVATIONS, STEM FRACTURE MAY HAVE BEEN DUE TO ONE OR A COMBINATION OF THE FOLLOWING MECHANISMS: EXTREME LOADING DUE TO PATIENT WEIGHT AND/OR PATIENT ACTIVITY; DISTAL IN-GROWTH WITH LACK OF PROXIMAL IN-GROWTH; STEM LOOSENING; COMPONENT MALALIGNMENT. HOWEVER, NO DEFINITIVE CAUSE ANALYSIS CAN BE COMPLETED WITH CERTAINTY. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS WAS ALSO NOT POSSIBLE AS THE PRODUCT AND/OR LOT NUMBERS REQUIRED FOR RETRIEVAL WERE UNAVAILABLE. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE AVAILABLE INFORMATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-03
1822565-2010-00130 | 2010-03-03 |  | I, n, j, u, r, y | Manufacturer report | TRABECULAR METAL MODULAR ACETABULAR SYSTEM SHELL WITH MULTI-HOLES | HIP PROSTHESIS | ZIMMER, INC. | NA | 00620206020 |  | TRABECULAR METAL SHELL WAS IMPLANTED IN A REVISION CASE, WITH ONE SCREW INTO THE SUPERIOR ILIUM. SHELL WAS LOOSE, AND THE SCREW HAD WORKED OUT THROUGH THE SHELL AND WAS STILL SECURE IN THE PATIENT'S BONE. REVISION WAS PERFORMED. | THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETED. | Y | N | HEALTH PROFESSIONAL |  | 2010-03
1822565-2010-00140 | 2010-03-09 | 2010-01-11 | I, n, j, u, r, y | Manufacturer report | KINECTIV TECHNOLOGY MODULAR NECK | HIP PROSTHESIS | ZIMMER, INC. | NA | 00784801300 |  | PT REPORTED AT 6 MONTHS POST-OP THAT HE HAD BEEN HAVING INTERMITTENT PAIN X3 MONTHS - THIS HAD BECOME SIGNIFICANT (AFTER HE HAD INCREASED ACTIVITY), NO EVIDENCE OF INFECTION OR LOOSENING AT THAT TIME. PT RETURNED 1 MONTH LATER, PAIN WAS NOT SUBSIDING AND ASPIRATION ARTHROGRAM WAS PERFORMED, APPROXIMATELY ASPIRATE GREW STAPH AUREUS. PT WAS SCHEDULED FOR GIRDLESTONE PROCEDURE WITH RE-IMPLANTATION TO FOLLOW IN 2-3 MONTHS. | EVALUATION SUMMARY: IT IS UNLIKELY THE DEVICE CONTRIBUTED TO THE REPORTED INFECTION. THE EXACT CAUSE OF THE INFECTION CAN NOT BE DETERMINED WITH CERTAINTY AT THIS TIME. ACCORDING TO THE PER THE PT'S PAIN INCREASED AS PT ACTIVITY INCREASED BUT THE ACTUAL ACTIVITIES WERE NOT DESCRIBED. PT ACTIVITY AND/OR POSSIBLE MALPOSITIONED COMPONENTS MAY HAVE CONTRIBUTED TO THE REPORTED PAIN DUE TO THE FACT THE PAIN WAS REPORTED WITHIN 3 MONTHS OF THE INITIAL SURGERY. HOWEVER, THE EXACT CAUSE OF THE PAIN CAN NOT BE DETERMINED WITH CERTAINTY AT THIS TIME. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADD'L SUBSTANTIVE INFO BE RECEIVED, THE COMPLAINT WILL BE REOPENED, ZIMMER CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-03
1822565-2010-00141 | 2010-03-09 | 2009-11-11 | I, n, j, u, r, y | Manufacturer report | TRILOGY ACETABULAR SYSTEM LONGEVITY CROSSLINKED POLYETHYLENE LINER | HIP PROSTHESIS | ZIMMER, INC. | NA | 00631005032 |  | THE LINER WAS USED ON (B) (6) 2007 AS PART OF A ROUTINE TOTAL HIP REPLACEMENT. THE PT WAS RETURNED TO THEATRE ON (B) (6) 2009 AND THE LINER HAD CRACKED. THE PT CAME IN WITH A DISLOCATED HIP AND WAS SUSPECTED OF HAVING A FALL. | THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. | Y | N | HEALTH PROFESSIONAL |  | 2010-03
1822565-2010-00133 | 2010-03-09 | 2009-12-30 | I, n, j, u, r, y | Manufacturer report | TRABECULAR METAL MODULAR ACETABULAR SYSTEM SHELL WITH CLUSTER HOLES | HIP PROSTHESIS | ZIMMER, INC. | NA | 00620205422 |  | PT HAD THA ON (B) (6) 2007 AND HAD CONTINUOUS PAIN IN RIGHT HIP. HER RIGHT LEG WAS GIVING OUT AND HAD GREAT PAIN AND STIFFNESS WHEN RISING FROM A CHAIR. SHE SAW SEVERAL SURGEONS AND HAD X-RAYS AND BONE SCAN DONE WHICH SHOWED NO PROBLEMS. TREATING ORTHOPAEDIC SURGEON REFERRED HER TO SPINAL CORD SPECIALIST WHO GAVE SHOTS AND MADE THE PAIN GO AWAY FOR A FEW DAYS BUT, WOULD RETURN. SHE WAS THEN REFERRED TO A NEW SURGEON. IT WAS DISCOVERED THAT THE ZIMMER SHELL WAS IRRITATING THE HIP MUSCLE THAT WAS CAUSING THE PAIN. | EVALUATION SUMMARY: PT'S ACETABULAR CUP, LINER, FEMORAL HEAD AND MODULAR NECK WERE REVISED DUE TO PERSISTENT, SEVERE HIP PAIN. MEDICAL INTERVENTIONS TO DIAGNOSE AND TREAT THE PAIN INCLUDED X-RAYS, BONE SCANS, SPINAL EXAM AND INJECTIONS, AND HIP INJECTIONS; ALL TREATMENTS RESULTED IN TEMPORARY ALLEVIATION OF PAIN. THE SIZE OF THE ACETABULAR CUP USED IN THE REVISION SURGERY IS LARGER THAN THE PRIMARY, INDICATING THAT CUP SIZE IS NOT THE LIKELY CAUSE OF THE PAIN. DURING THE REVISION SURGERY, THE PRIMARY CUP WAS OBSERVED TO BE WELL FIXED AND NOT LOOSE, INDICATING THAT LOOSENESS OF THE CUP IS NOT A LIKELY CAUSE OF THE PAIN. THE PRIMARY SURGEON DID NOT INDICATE ANY INFO ABOUT THE PLACEMENT OF THE CUP OTHER THAN THAT "EXCELLENT FIXATION FOUND" AFTER IMPACTION, AND THAT HE UTILIZED A IMAGE INTENSIFIER TO "CONFIRM APPROPRIATE POSITIONING." NO X-RAYS FROM THE ORIGINAL SURGERY ARE AVAILABLE FOR REVIEW OR COMPARISON WITH THE REVISING SURGEON'S OBSERVATIONS. THE REVISING SURGEON STATED THAT THE CUP APPEARED "QUITE LARGE" AND NOTED "CONSIDERABLE AMOUNT OF EXPOSED POSTERIOR IMPLANT METAL PRESENT BEFORE REACHING BONE." THE MOST LIKELY CAUSE OF THE PAIN THAT LED TO REVISION SURGERY WAS IRRITATION OF SOFT TISSUES BY THE EXPOSED RIM OF THE POSTERIOR REGION OF THE SHELL OR BY SOFT TISSUE IMPINGEMENT, CAUSED BY EXCESSIVE ANTEVERSION OF THE SHELL. BASED ON THE OBSERVED INSTABILITY WHEN THE NEW CUP WAS PLACED IN A MORE ANATOMIC POSITION, IT IS POSSIBLE THAT THIS PLACEMENT WAS CHOSEN TO MITIGATE RISK OF DISLOCATION. THE REVISING SURGEON'S USE OF A NECK WITH GREATER ANTEVERSION AND OFFSET ALLOWED FOR JOINT STABILITY WITH A MORE OPTIMAL PLACEMENT. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADD'L SUBSTANTIVE INFO BE RECEIVED, THE COMPLAINT WILL BE REOPENED, ZIMMER INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-03
1822565-2010-00132 | 2010-03-09 | 2010-01-26 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CONTINUUM ACETABULAR SYSTEM TRABECULAR METAL SHELL | HIP PROSTHESIS | ZIMMER, INC. | NA | 00875706202 |  | THE SURGEON WAS USING THE CONTINUUM CUP FOR HIS 2ND TIME ON THIS OCCURRENCE. HE WAS USING THE OFFSET/CURVED CUP IMPACTOR WITH THE ROTATIONAL CONTROLLING CAP. THE CUP WAS ATTACHED TO THE IMPACTOR ON THE BACK TABLE, WITH CARE TO MAKE SURE THAT THE SQUARE DIMPLE WAS ALIGNED AND THEN THE THREADS WERE ADVANCED LOCKING ONTO THE CUP. UPON INSERTION OF THE CUP, THE IMPACTOR SNAPPED/SEPARATED COMPLETELY FROM THE CUP WHILE ENGAGING THE PREPARED ACETABULUM AND TM AUGMENT. THE CUP IMPACTOR WAS ABLE TO BE RE-ATTACHED TO THE CUP WHILE THE CUP WAS IN PATIENT AFTER SEVERAL MINUTES OF TRYING. I CANNOT RECALL IF THE TIP HAD TO BE REMOVED IN ORDER TO GET THE THREADS TO ENGAGE. TITANIUM THREAD WAS FOUND STRIPPED OUT OF CUP AND IN THE WOUND. CUP WAS THEN FINISHED SEATING AND SCREWS WERE PLACED. | THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. | N | Y | HEALTH PROFESSIONAL |  | 2010-03
1822565-2010-00135 | 2010-03-09 | 2010-02-01 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CONTINUUM ACETABULAR SYSTEM TRABECULAR METAL SHELL WITH CLUSTER HOLES POROUS | HIP PROSTHESIS | ZIMMER, INC. | NA | 00875706001 |  | THE SURGEON WAS USING THE CONTINUUM CUP FOR HIS 4TH OR 5TH TIME ON THIS OCCURRENCE. HE WAS USING THE OFFSET/CURVED CUP IMPACTOR WITH THE ROTATIONAL CONTROLLING CAP. THE CUP WAS ATTACHED TO THE IMPACTOR ON THE BACK TABLE, WITH CARE TO MAKE SURE THAT THE SQUARE DIMPLE WAS ALIGNED AND THEN THE THREADS WERE ADVANCED LOCKING ONTO THE CUP. UPON INSERTION OF THE CUP THE IMPACTOR SNAPPED/SEPARATED COMPLETELY FROM THE CUP WHILE ENGAGING THE PREPARED ACETABULUM. THE CUP IMPACTOR OF CHOICE COULD NOT BE RE-ATTACHED TO THE CUP WHILE CUP WAS IN PATIENT. TITANIUM THREAD WAS FOUND STRIPPED OUT OF CUP AND IN THE WOUND. CUP WAS FULLY SEATED, HOWEVER, THERE WASN'T A WAY TO VERIFY THE CORRECT PLACEMENT OF CUP, AND REMOVAL OF CUP AT THIS TIME WAS NOT EASY. DECISION WAS MADE TO ACCEPT POSITIONING FOR THE TIME BEING AND CONTINUE ON TO THE FEMORAL COMPONENT. UPON COMPLETION OF THE FEMORAL COMPONENT, IT WAS REALIZED THAT THE VERSION OF THE CUP WAS NOT CORRECT AND IT HAD TO BE EXPLANTED. THE ONLY CUP INSERTER/IMPACTOR THAT WE COULD GET TO SCREW INTO THE CUP SUFFICIENTLY FOR EXPLANTING WAS THE OLD STRAIGHT TRILOGY IMPACTOR, USED WITHOUT A CAP AND ENGAGING INTO THE VERY DEEP THREADS OF THE HOLE. WE THEN INSERTED A TM MODULAR CUP WITH THE SAME CURVED IMPACTOR USING BOWTIE RIM IMPACTOR TIP AND HAD ZERO ISSUES. SURGERY WAS PROLONGED BY 35 MINUTES. | THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. | N | Y | HEALTH PROFESSIONAL |  | 2010-03
2648920-2010-00009 | 2010-03-10 | 2010-02-09 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY ACETABULAR SYSTEM SHELL WITH CLUSTER HOLES POROUS | HIP PROSTHESIS | ZIMMER MANUFACTURING B.V. | NA | 00620005622 |  | WE WERE UNABLE TO GET 2 DIFFERENT LINERS TO LOCK INTO THIS TRILOGY CLUSTER SHELL. THE SURGEON TRIED FOR 45 MINUTES TO MAKE THIS SHELL WORK. HE CLEANED THE RIM FOR OSTEOPHYTES, CLEANED AND CLEARED ALL SOFT TISSUE, IRRIGATED, AND CHECKED THE SCREW HEADS. THE LINER STILL WOULD NOT LOCK DOWN. AT THAT POINT HE CHOSE TO TAKE THE CUP OUT AND IMPLANT ANOTHER 56 MM TRILOGY CUP, WHICH WAS A MULTI-HOLE BECAUSE THAT IS ALL WE HAD FOR BACK UP. THE LINER LOCKED INTO THIS SHELL JUST FINE. SURGERY WAS PROLONGED BY 45-60 MINUTES. | EVAL SUMMARY: THE RECEIVED DEVICES WERE REVIEWED. IN THEIR RETURNED CONDITIONS, IT WAS OBSERVED THAT THE TRILOGY CLUSTER-HOLED SHELL HAD SCRATCHES IN SEVERAL PLACES, TWO PATCHES OF SURFACE "DELAMINATION", A CIRCULAR BUFFED MARK CENTERED ABOUT THE POLAR HOLD INDICATING WHERE MATED WITH SHELL IMPACTOR, AND THIRD BODY PARTICLES IN VARIOUS LOCATIONS ON ONE HALF OF THE POROUS SUFACE. THE LOCKING RING HAD THIRD BODY SUBSTANCE AT APEX OF RING, THE SURFACE WAS BUFFED IN VARIOUS LOCATIONS, AND IT WAS BENT IN A "WAVE" PATTERN. WHEN PLACED ON A FLAT SURFACE, THE LOCKING RING TABS AND APEX OF THE RING WERE ON THE SURFACE, WITH THE REMAINING RING BENT UPWARDS TO APPROX 1 MM. THE TRILOGY UHMWPE STANDARD LINER HAD SCRATCHES ON OUTER SURFACE IN VARIOUS LOCATIONS, HAD ONE BUFF MARK IN AN ARCH SHAPE ON OUTER SURFACE LOCATED APPROX 3 CM FROM THE LINER EDGE, THE POLAR BOSS EDGE WAS STRIPPED, THE BOTTOM EDGE OF FLANGE WAS FRACTURED IN A WEDGE SHAPE APPROX 2 CM LONG STARTING AT ONE SCALLOP, AND HAD SCRATCHES ON CHAMFERED EDGE AND FLAT SURFACE OF THE FLANGE LOCATED UNDER "56 X 36" ENGRAVING. IT IS UNK IF THE LINER WAS ATTEMPTED TO BE INSERTED AS PER SURGICAL TECHNIQUE. AS THE SURGICAL NOTES ARE UNAVAILABLE, IT IS MOST LIKELY THAT THE LOCKING RING WAS MISALIGNED IN THE SHELL AND WHEN THE POLY LINER WAS ATTEMPTED TO BE INSERTED, THE LOCKING RING MAY HAVE DEFORMED. THIS MAY HAVE CAUSED THE LOCKING RING TO NOT MATER PROPERLY WITH THE POLY LINER. HOWEVER, THE EXACT CAUSE OF THE CURRENT SITUATION CANNOT BE CONCLUSIVELY DETERMINED. EVAL: REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADD'L SUBSTANTIVE INFO BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | N | Y | HEALTH PROFESSIONAL |  | 2010-03
2648920-2010-00010 | 2010-03-11 | 2009-11-09 | I, n, j, u, r, y | Manufacturer report | VERSYS HIP SYSTEM FEMORAL HEAD 1 | HIP PROSTHESIS | ZIMMER MANUFACTURING B.V. | NA | 00801803202 |  | THE PRODUCT EXPERIENCE REPORT STATES: "IT WAS REPORTED THAT THOUGH ABOVE MENTIONED FEMORAL HEAD ONCE DISSOCIATED FROM THE STEM IN A WEEK AFTER IMPLANTATION , THE CORRECTION SURGERY HAS BEEN CONDUCTED TO REASSEMBLE THEM IMMEDIATELY FOR CONTINUOUS USE." | EVAL SUMMARY - NO PRODUCT WAS RETURNED AS IT REMAINS IN THE PT. IT IS UNK WHETHER PROPER SURGICAL TECHNIQUE WAS FOLLOWED FOR IMPACTING THE FEMORAL HEAD TO THE FEMORAL STEM. BASED ON THE ABOVE INFO AND OBSERVATIONS, THE DISSOCIATION OF THE FEMORAL HEAD FROM THE FEMORAL STEM MAY HAVE BEEN DUE TO ONE OR A COMBINATION OF THE FOLLOWING OCCURRENCES. A PROBABLE CONTRIBUTOR IS THAT THE FEMORAL HEAD DID NOT MATE PROPERLY WITH THE TAPER DURING SURGERY. THIS COULD BE DUE TO SOFT TISSUE OR THIRD BODY PARTICLES (BONE CEMENT) ON THE TAPER DURING IMPACTION OR INSUFFICIENT IMPACTION TO FULLY ENGAGE THE HEAD ONTO THE TAPER. ALTERNATIVELY, PT ACTIVITY LEVEL MAY HAVE AGGRAVATED AN IMPROPERLY SEATED FEMORAL HEAD AND CAUSED IT TO DISSOCIATE. BASED ON THE AVAILABLE INFO, A DEFINITIVE ROOT CAUSE CAN NOT BE ASCERTAINED AT THIS TIME. EVAL - NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFO BE RECEIVED, THE COMPLAINT WILL RE-OPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-03
1822565-2010-00152 | 2010-03-11 | 2010-02-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CONTINUUM TRILOGY IT ALLOFIT IT ACETABULAR SYSTEMS DOME IMPACTOR | HIP INSTRUMENT | ZIMMER, INC. | NA | 00879001540 |  | IT IS REPORTED THAT AS THE SURGEON WAS IMPACTING A CONTINUUM LINER, THE YELLOW DOME IMPACTOR FRACTURED INTO PIECES. | INFORMATION WAS RECEIVED FROM A FOREIGN SOURCE WHO IS NOT REQUIRED TO COMPLETE FORM 3500A. EVALUATION SUMMARY: IT IS UNKNOWN WHETHER THE DOME IMPACTOR WAS ASSEMBLED TO THE UNIVERSAL HANDLE AND IMPACTED PROPERLY PER INSTRUCTIONS IN THE SURGICAL TECHNIQUE. IMPROPER ASSEMBLY AND USE COULD HAVE LEAD TO THE REPORTED FRACTURE BUT THE ROOT CAUSE CANNOT BE DETERMINED WITH CERTAINTY AT THIS TIME. EVALUATION: NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS WAS ALSO NOT POSSIBLE AS THE PRODUCT AND/OR LOT NUMBERS REQUIRED FOR RETRIEVAL WERE UNAVAILABLE. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE AVAILABLE INFORMATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE RE-OPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | N | Y | HEALTH PROFESSIONAL |  | 2010-03
9613350-2010-00104 | 2010-03-15 | 2010-02-10 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT | DUROM ACETABULAR COMPONENT AND METASUL LDH | ZIMMER GMBH |  | UNK |  | IT IS REPORTED BY PT THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2009. POST-OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION IS SCHEDULED FOR THE SUMMER OF 2010. | (B) (4). THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X- RAYS, OR OTHER SOURCE DOCUMENTS TO REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-03
1822565-2010-00153 | 2010-03-15 | 2010-02-16 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ZIMMER MIS HIP INSTRUMENTATION HYBRID OFFSET SHELL INSERTER | HIP INSTRUMENT | ZIMMER, INC. | NA | 00780402520 |  | IT IS REPORTED THAT THE INSERTER STRIPPED, CAUSING THE SURGEON TO INSERT A PLASTIC LINER INSTEAD OF THE ORIGINALLY PLANNED METAL ONE. | THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. | N | Y | HEALTH PROFESSIONAL |  | 2010-03
1822565-2010-00162 | 2010-03-16 | 2010-02-17 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CPT HIP SYSTEM 12/14 TAPER INSTRUMENTATION STEM INSERTER | HIP INSTRUMENT | ZIMMER, INC. | NA | 00833401300 |  | THE INSERTER BROKE WHEN THE SCREW-THREAD BROKE OFF IN THE STEM. | THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETED. | N | Y | LAY USER/PATIENT |  | 2010-03
1822565-2010-00159 | 2010-03-16 | 2007-07-13 | I, n, j, u, r, y | Manufacturer report | ZIMMER M/L TAPER HIP PROSTHESIS FEMORAL STEM PRESS-FIT PLASMA-SPRAYED | HIP PROSTHESIS | ZIMMER, INC. | NA | 00771101300 |  | REVISION DUE TO INFECTION. | EVALUATION SUMMARY: THE STERILIZATION PROCESS FOR THIS DEVICE WAS VALIDATED IN ACCORDANCE WITH FDA REGULATIONS AND ISO STANDARDS TO A STERILITY ASSURANCE LEVEL OF 1.0 X 10-6 OR BETTER. IN ADDITION, THE MFG LOT SPECIFIED IN THIS COMPLAINT WAS PROCESSED ACCORDING TO THE VALIDATED STERILIZATION PROCESS PARAMETERS AND MET ALL THE ACCEPTANCE CRITERIA FOR STERILITY RELEASE. THEREFORE, IT IS HIGHLY UNLIKELY THAT THE SPECIFIED DEVICE CAUSED OR CONTRIBUTED TO THE PT INFECTION. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECT THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFO BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-03
1822565-2010-00164 | 2010-03-16 | 2010-02-08 | I, n, j, u, r, y | Manufacturer report | TRILOGY ACETABULAR SYSTEM LONGEVITY / CROSSLINKED POLYTHYLENE LINER STANDARD | HIP PROSTHESIS | ZIMMER, INC. | NA | 00630505036 |  | ON (B) (6) 2010, SURGERY WAS PERFORMED ON THIS PT TO REPAIR A TENDON IN THE HIP. THE PT HAD UNDERGONE A TOTAL HIP REVISION (B) (6) 2007. UPON ENTRY INTO THE JOINT, THE SURGEON NOTICED 2 PIECES OF PLASTIC. FURTHER EXAMINATION IN-VIVO INDICATED VISUALLY A FRACTURE OF THE POLY LINER AT THE RIM. THE LINER WAS STILL INTACT WITHIN THE ACETABULAR CUP WITH THE HIP PROPERLY LOCATED. LINER AND HEAD WERE REMOVED AND REPLACED. SURGEON NOTED NO IMPINGEMENT THROUGHOUT THE RANGE OF MOTION WITH THE 32MM LINER. | EVALUATION SUMMERY: THE LINER DIMENSIONS AS RETURNED CANNOT BE EVALUATED DUE TO THE FRACTURE. SEM ANALYSIS OF THE FRACTURED LINER SHOWS EVIDENCE THAT THE FRACTURE OCCURRED BY FATIGUE, ORIGINATING ON THE OUTER DIAMETER, PROPAGATING, ALONG THE OUTER AND INNER DIAMETERS. CAUSES OF LINER FRACTURE LISTED IN THE LITERATURE INCLUDE TRAUMA, PT ACTIVITY, PT WEIGHT IMPINGEMENT (CONTRIBUTING FACTORS: CUP POSITION OR GEOMETRY OF THE CUP/HEAD/STEM SYSTEM, PT MOVEMENTS), DISLOCATION (CONTRIBUTING FACTORS: PT NEUROMUSCULAR DEFICIENCY, INADEQUATE TISSUE TENSION, CUP POSITION, PRIOR SURGERY ON THE HIP), REDUCED STRENGTH RELATIVE TO NON-CROSSLINKED LINERS, LINER OXIDATION, AND RELATIVELY THIN POLYETHYLENE AT THE CUP RIM. NOTES OF THE REVISION SURGERY IN 2007 WHEN THE TRILOGY LONGEVITY LINER WAS IMPLANTED INDICATE THE CAUSE OF THAT REVISION WAS RECURRENT DISLOCATION. THE SURGEON EXPLICITLY STATES IN THE NOTES THAT A 36MM LINER/HEAD WAS SELECTED TO MITIGATE THE POTENTIAL FOR FUTURE DISLOCATIONS. THE SURGEON OBSERVED STABILITY OF THE JOINT THROUGHOUT THE RANGE OF MOTION AT THE TIME OF THE SURGERY. NO X-RAYS ARE INCLUDED; EVALUATION OF THE CUP POSITION IS NOT POSSIBLE. PREVIOUSLY REFERENCED FOLLOW-UP ALSO NOTED NO PAIN WITH RANGE OF MOTION IN THE RIGHT HIP. AT THE FINAL REVISION SURGERY, THE 36 MM LINER WAS REPLACED WITH A 32MM LINER, AND NO IMPINGEMENT WAS OBSERVED BY THE SURGEON THROUGHOUT THE RANGE OF MOTION. THE OBSERVATIONS MAKE IMPINGEMENT LESS LIKELY, THOUGH IT CANNOT BE RULED OUT. IT IS UNK WHETHER THE TORN TENDON WAS CAUSED BY TRAUMA, DISLOCATION, OR BY THE FRACTURED LINER. THE ROOT CAUSE OF THE LINER FRACTURE CAN NOT BE DETERMINED FROM THE INFO PROVIDED. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFO BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-03
1822565-2010-00169 | 2010-03-17 | 2010-02-12 | I, n, j, u, r, y | Manufacturer report | HARRIS PRECOAT PLUS WITH DISTAL PLUG FEMORAL STEM | HIP PROSTHESIS | ZIMMER, INC. | NA | 00902801000 |  | PAIN IN PT, X-RAY CONFORMATION OF LOOSENING, STEM REMOVED RELATIVELY EASILY. TRILOGY LINER CHANGED AT THE SAME TIME. | (B) (4). EVAL SUMMARY: THE RECEIVED X-RAY IMAGE (PRE-2010 REVISION SURGERY) WAS REVIEWED, AND IT WAS OBSERVED THAT THE CEMENT MANTLE APPEARS TO HAVE DISASSOCIATED FROM THE SURROUNDING BONE, AND THE DISTAL MANTLE APPEARS TO BE BROKEN. THE DEVICES WERE IN-VIVO APPROX 12.5 YEARS. THE DEVICE HISTORY RECORDS OF THE STEM INDICATE THAT THE COMPONENT WAS MANUFACTURED AND INSPECTED ACCORDING TO SPECIFICATIONS THE STEM AND HEAD WERE UNAVAILABLE FOR ANALYSIS, AND THEIR CONDITIONS ARE UNK. THE MATING ACETABULAR SHELL COMPONENT IS UNK, AND ITS CONDITION IS ALSO UNK. SURGICAL NOTES FROM THE PRIMARY SURGERY AS WELL AS THE REVISION SURGERY ARE UNAVAILABLE, AND IT IS UNK IF THE COMPONENTS WERE IMPLANTED WITH THE CORRECT ORIENTATION AND FIT AS PER THE SURGICAL TECHNIQUE. X-RAY IMAGES POST-PRIMARY SURGERY AND FROM SUCCESSIVE PT VISITS PRIOR TO THE 2010 REVISION SURGERY ARE UNAVAILABLE. PT AGE, HEIGHT, WEIGHT, AND ACTIVITY LEVEL ARE ALSO UNKNOWN. BASED ON THE ABOVE INFO AND OBSERVATIONS, STEM LOOSENING MAY HAVE BEEN DUE TO ONE OR A COMBINATION OF THE FOLLOWING MECHANISMS: MISALIGNMENT OF THE FEMORAL STEM WITHIN THE CEMENT MANTLE. IF THE STEM WAS INSERTED IN A VARUS POSITION, THIS MAY HAVE INCREASED THE CEMENT MANTLE STRESSES WHICH MAY HAVE RESULTED IN DISTAL CEMENT MANTLE FRACTURE AND SUBSEQUENT STEM LOOSENING. INADEQUATE CEMENT MANTLE FIXATION. CEMENT MANTLE FRACTURE MAY HAVE LED TO STEM LOOSENING. PT ACTIVITY LEVEL AND WEIGHT. HOWEVER, NO DEFINITIVE CAUSE ANALYSIS CAN BE COMPLETED WITH CERTAINTY. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFO BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER INC CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-03
1822565-2010-00170 | 2010-03-17 | 2010-02-17 | I, n, j, u, r, y | Manufacturer report | VERSYS HIP SYSTEM PROVISIONAL FEMORAL HEAD NECK TAPER STANDARD NECK LENGTH | HIP INSTRUMENT | ZIMMER, INC. | NA | 00789503202 |  | IT IS REPORTED THAT THE HEAD TRIAL WAS DROPPED INTO THE PT AND COULD NOT BE RETRIEVED. THE DOCTOR LOOKED AROUND FOR AN HOUR AND A HALF, BUT COULD NOT FIND IT ANYWHERE. THEY CALLED IN X-RAY, BUT X-RAY COULDN'T FIND IT EITHER. THE DOCTOR THEN DECIDED TO CALL IT AND SEWED THE PT UP, LEAVING IT IN. | (B) (4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. | Y | N | HEALTH PROFESSIONAL |  | 2010-03
MW5015201 | 2010-03-17 | 2010-02-23 | I, n, j, u, r, y | Voluntary report | ZIMMER | FEMORAL TRABECULAR METAL SIZE 12; 128 MM STEM LENGTH | ZIMMER DEPTULA |  |  |  | THIS IS A (B) (6) FEMALE WHO FELL ON (B) (6) 2010. THE FALL RESULTED IN A RIGHT HIP FEMORAL NECK FRACTURE. THE PT HAD A RIGHT HIP HEMI-ARTHROPLASTY PERFORMED ON (B) (6) 2010. THE PT'S HOSPITALIZATION WAS UNEVENTFUL AND SHE WAS TRANSFERRED TO A (B) (6) NURSING FACILITY ON (B) (6) 2010 TO CONTINUE HER RECOVERY AND REHABILITATION. THE PT WAS DOING WELL UNTIL (B) (6) 2010, WHEN SHE COMPLAINED OF SEVERE PAIN WHILE PERFORMING PHYSICAL THERAPY AT THE (B) (6) NURSING FACILITY. THE PT WAS TRANSFERRED BACK TO THE HOSPITAL WHERE X-RAYS SHOWED A DE-COAPTATION OF THE RIGHT HIP PROSTHESIS FEMORAL STEM FROM THE FEMORAL HEAD. THE FEMORAL HEAD AND STEM WERE STILL IN THE ACETABULAR SOCKET. THE TRUNNION OF THE STEM WAS IMPACTED ON THE LATERAL EDGE OF THE ACETABULUM AND EXTRAARTICULAR. THE PT RETURNED TO SURGERY FOR REPAIR OF THE RIGHT HIP PROSTHESIS. INTRAOPERATIVE FINDINGS IDENTIFIED THAT THE BIPOLAR HEAD AND SHELL WERE INTACT. THE ACETABULUM WAS IN THE SOCKET, BUT THE FEMORAL STEM WAS DISENGAGED. THE PT WAS CONVERTED TO A TOTAL HIP REPLACEMENT. DATES OF USE: (B) (6) 2010 - (B) (6) 2010. DIAGNOSIS OR REASON FOR USE: RIGHT HIP FRACTURE. |  | Y | Y | HEALTH PROFESSIONAL |  | 2010-03
9613350-2010-00124 | 2010-03-18 | 2009-09-21 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL LDH | ZIMMER GMBH |  | 01.00214.154 |  | IT IS REPORTED BY THE PT'S ATTORNEY THAT THE PT UNDERWENT A RIGHT HIP ARTHROPLASTY ON (B) (6) 2008. POST-OP, SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2009. | (B) (4). THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN 07/2008. SHOULD ADD'L INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-03
9613350-2010-00121 | 2010-03-18 | 2010-02-19 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 56/50 P | DUROM ACETABULAR COMPONENT AND METASUL LDH | ZIMMER GMBH |  | 01.00214.156 |  | IT IS REPORTED BY THE PT THAT PT UNDERWENT A HIP ARTHROPLASTY IN (B) (6) 2007. POST-OP, SHE EXPERIENCES PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION IS SCHEDULED FOR (B) (6) 2010. | (B) (4). THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN 07/2008. SHOULD ADD'L INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-03
9613350-2010-00120 | 2010-03-18 | 2009-04-15 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 48/42 H | DUROM ACETABULAR COMPONENT AND METASUL LDH | ZIMMER GMBH |  | 01.00214.148 |  | IT IS REPORTED BY PATIENT'S ATTORNEY THAT PATIENT UNDERWENT A RIGHT HIP ARTHROPLASTY ON (B) (6) 2006. POST-OP, SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2009. | (B) (4). THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-03
1822565-2010-00171 | 2010-03-18 | 2010-02-19 | I, n, j, u, r, y | Manufacturer report | TRILOGY ACETABULAR SYSTEM LINER STANDARD LONGEVITY CROSSLINKED POLYETHYLENE | HIP PROSTHESIS | ZIMMER, INC. | NA | 00630505036 |  | PATIENT HAD SYMPTOMS OF SUBLUXATION AND CLUNKING IN HER LEFT HIP. INTRAOPERATIVELY, PHYSICIAN NOTICED A CRACK IN POSTERIOR ASPECT OF LINER. PROCEEDED TO DISLOCATE HIP AND WHILE DOING SO, PIECE OF POLY COMPLETELY BROKE OFF. | EVALUATION SUMMARY: THE LINER WAS NOT RETURNED. CAUSES OF LINER CRACKING AND/OR FRACTURE LISTED IN THE LITERATURE INCLUDED TRAUMA, PATIENT ACTIVITY, PATIENT WEIGHT, IMPINGEMENT (CONTRIBUTING FACTORS: CUP POSITION OR GEOMETRY OF THE CUP/HEAD/STEM SYSTEM, PATIENT MOVEMENTS), DISLOCATION (CONTRIBUTING FACTORS: PATIENT NEUROMUSCULAR DEFICIENCY, INADEQUATE TISSUE TENSION, CUP POSITION, PRIOR SURGERY ON THE HIP), REDUCED STRENGTH RELATIVE TO NON-CROSSLINKED LINERS, LINER OXIDATION, AND RELATIVELY THIN POLYETHYLENE AT THE CUP RIM. NO X-RAYS ARE INCLUDED; THEREFORE, EVALUATION OF THE CUP IS NOT POSSIBLE. THE LINER WAS NOT RETURNED; THEREFORE, ANALYSIS OF THE FRACTURE IS NOT POSSIBLE. PATIENT BUILD NOTED AS SMALL MAKES EXCESSIVE WEIGHT AN UNLIKELY FACTOR. THE ROOT CAUSE OF THE LINER CRACKING AND SUBSEQUENT FRACTURE CANNOT BE DETERMINED FROM THE INFORMATION PROVIDED. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-03
1822565-2010-00172 | 2010-03-18 | 2010-02-16 | I, n, j, u, r, y | Manufacturer report | VERSYS HIP SYSTEM FEMORAL STEM FIBER METAL TAPER COLLARLESS | HIP PROSTHESIS | ZIMMER, INC. | NA | 00786201300 |  | POST-OPERATIVELY, PATIENT COMPLAINED OF NOT BEING ABLE TO POINT FOOT FORWARD. X-RAY DETERMINED SUBSIDENCE OF ABOUT 1 CM WITH ROTATION OF STEM. MALPOSITIONING WAS CONFIRMED DUE TO SMALL FRACTURE OF PROXIMAL FEMUR. | EVALUATION SUMMARY: THE DEVICES WERE IN-VIVO APPROXIMATELY 2.5 MONTHS. THE DEVICES ARE UNAVAILABLE FOR ANALYSIS AND THE DEVICE CONDITIONS ARE UNKNOWN. IT IS UNKNOWN IF THE OBSERVED FRACTURE OF THE PROXIMAL FEMUR OCCURRED DURING STEM INSERTION DURING THE PRIMARY SURGERY, OR SUBSEQUENT TO STEM SUBSIDENCE. X-RAY IMAGES POST-PRIMARY SURGERY AND FROM SUCCESSIVE PATIENT VISITS ARE UNAVAILABLE. IT IS UNKNOWN IF THE STERN WAS IMPLANTED WITH THE CORRECT ORIENTATION AND FIT AS PER THE SURGICAL TECHNIQUE. THE INSTRUMENTS USED IN THE PRIMARY SURGERY ARE ALSO UNKNOWN. PATIENT'S HEIGHT AND ACTIVITY LEVEL ARE UNKNOWN. THE REHABILITATION PROGRAM, BOTH PHYSICAL AND PHARMACOLOGICAL, AND COMPLIANCE THEREOF IS UNKNOWN. BASED ON THE ABOVE INFORMATION AND OBSERVATIONS, STEM SUBSIDENCE MAY HAVE BEEN DUE TO ONE OR A COMBINATION OF THE FOLLOWING MECHANISMS: INADEQUATE PRESS FIT BETWEEN THE STEM AND PREPARED FEMUR. THIS MAY HAVE NOT PROVIDED THE NECESSARY INITIAL STABILIZATION OF THE STEM, WHICH UPON LOAD-BEARING MAY HAVE RESULTED IN STEM SUBSIDENCE. THIS MAY HAVE NOT PROVIDED THE NECESSARY INITIAL STABILIZATION OF THE STEM, WHICH, UPON LOAD-BEARING, MAY HAVE RESULTED IN STEM SUBSIDENCE. ALTERNATIVELY, RASPING TECHNIQUE AND/OR PATIENT BEHAVIOUR MAY HAVE ALSO CONTRIBUTED TO THE ALLEGED EVENT OF STEM SUBSIDENCE. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-03
1822565-2010-00173 | 2010-03-18 | 2010-02-10 | I, n, j, u, r, y | Manufacturer report | ZIMMER ANATOMIC HIP PROSTHESIS FEMORAL STEM | HIP PROSTHESIS | ZIMMER, INC. | NA | 00660001321 |  | ALL COMPONENTS WERE REMOVED DUE TO LIKELY INFECTION. | EVALUATION SUMMARY: THE STERILIZATION PROCESS FOR THIS DEVICE WAS VALIDATED IN ACCORDANCE WITH FDA REGULATIONS AND ISO STANDARDS TO A STERILITY ASSURANCE LEVEL OF 1.0 X 10-6 OR BETTER. IN ADDITION, THE MANUFACTURING LOT SPECIFIED IN THIS COMPLAINT WAS PROCESSED ACCORDING TO THE VALIDATED STERILIZATION PROCESS PARAMETERS AND MET ALL THE ACCEPTANCE CRITERIA FOR STERILITY RELEASE. THEREFORE, IT IS HIGHLY UNLIKELY THAT THE SPECIFIED DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT INFECTION. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-03
2648920-2010-00011 | 2010-03-18 | 2010-02-12 | I, n, j, u, r, y | Manufacturer report | TRILOGY ACETABULAR SYSTEM SHELL WITH HOLES POROUS | HIP PROSTHESIS | ZIMMER MANUFACTURING B.V. | NA | 00620005420 |  | PRIMARY HIP ARTHROPLASTY WAS PERFORMED IN (B) (6) 2003. REVISION WAS PERFORMED ON (B) (6) 2010, DUE TO PAIN AND RADIOGRAPHIC LOOSENING. INFECTION WAS RULED OUT. CUP AND LINER WERE REMOVED. | EVALUATION SUMMARY: THE DEVICES WERE IN-VIVO APPROXIMATELY 6.5 YEARS. THE DEVICES WERE UNAVAILABLE FOR ANALYSIS, AND THEIR CONDITIONS ARE UNKNOWN. THE SURGICAL NOTES FROM THE PRIMARY SURGERY AS WELL AS REVISION SURGERY ARE UNAVAILABLE, AND IT IS UNKNOWN IF THE TRILOGY ACETABULAR SYSTEM WAS IMPLANTED WITH THE CORRECT ORIENTATION AND FIT AS PER THE SURGICAL TECHNIQUE. INSTRUMENTS USED IN THE PRIMARY SURGERY ARE UNKNOWN. X-RAY IMAGES POST-PRIMARY SURGERY AND FROM SUCCESSIVE PATIENT VISITS ARE UNAVAILABLE. PATIENT'S HEIGHT AND WEIGHT ARE UNKNOWN. BASED ON THE ABOVE INFORMATION AND OBSERVATIONS, SHELL LOOSENING AND PAIN MAY HAVE BEEN DUE TO ONE OR A COMBINATION OF THE FOLLOWING MECHANISMS: INITIAL SHELL INSTABILITY AND/OR INITIAL PRESS-FIT CONDITION OF THE SHELL, DEBRIS BETWEEN THE SHELL AND REAMED ACETABULUM DURING IMPLANTATION, IMPROPER SHELL SIZE FOR PATIENT, MISALIGNMENT OF SHELL, IMPINGEMENT, POOR BONE QUALITY, PATIENT ACTIVITY POST-OP. HOWEVER, THE EXACT CAUSE OF THE CURRENT SITUATION CANNOT BE CONCLUSIVELY DETERMINED. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-03
9613350-2010-00125 | 2010-03-19 | 2010-01-29 | I, n, j, u, r, y | Manufacturer report | EU DUROM ACETABULAR COMPONENT DUROM CUP | DUROM ACETABULAR COMPONENT AND METASUL LDH | ZIMMER GMBH |  | UNK |  | IT WAS REPORTED THAT A REVISION SURGERY OCCURRED DUE TO PAIN. | THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-03
9613350-2010-00015 | 2010-03-22 | 2008-02-18 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 60/54 T | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.160 |  | IT IS REPORTED BY THE PATIENT'S ATTORNEY THAT THE PATIENT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2008. POST-OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2008. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-03
9613350-2010-00131 | 2010-03-23 | 2009-04-23 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMNPONENT 48/42 H | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.148 |  | IT IS REPORTED BY PATIENT'S ATTORNEY THAT PATIENT UNDERWENT A LEFT HIP ARTHROPLASTY ON (B) (6) 2008. POST-OP, SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2009 BY DR (B) (6). | (B) (4). THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLANTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-03
9613350-2010-00127 | 2010-03-23 | 2010-02-24 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.154 |  | IT IS REPORTED BY PATIENT'S ATTORNEY THAT PATIENT UNDERWENT A RIGHT HIP ARTHROPLASTY IN (B) (6) OF 2007. POST-OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION IS SCHEDULED. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-03
9613350-2010-00129 | 2010-03-23 | 2010-02-26 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 56/50 P | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.156 |  | IT IS REPORTED BY PATIENT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2008. POST-OP, HE EXPERIENCES PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. A REVISION HAS BEEN SCHEDULED FOR (B) (6) 2010. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-03
9613350-2010-00130 | 2010-03-23 | 2010-02-25 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 52/46 L | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.152 |  | IT IS REPORTED BY PATIENT'S ATTORNEY THAT PATIENT UNDERWENT A LEFT HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, SHE EXPERIENCES PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION IS SCHEDULED (B) (6) 2010 BY DR (B) (6). | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-03
9613350-2010-00126 | 2010-03-23 | 2010-02-23 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.154 |  | IT IS REPORTED BY PATIENT THAT PATIENT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2006. POST-OP, HE EXPERIENCES PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. A REVISION IS SCHEDULED FOR (B) (6) 2010. | (B) (4). THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBER WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4) | Y | N | LAY USER/PATIENT |  | 2010-03
1822565-2010-00176 | 2010-03-23 | 2010-02-22 | O, t, h, e, r | Manufacturer report | CONTINUUM TRILOGY IT ALLOFIT IT ACETABULAR SYSTEMS LONGEVITY HIGHLY CROSSLINKED | HIP PROSTHESIS | ZIMMER, INC. | NA | 00875101336 |  | REPORTED DIRECTLY FROM THE PRODUCT EXPERIENCE REPORT: "ON (B)(6) 2010, THE SURGEON USED A 58MM TM REVISION CUP FOR AN ACETABULAR REVISION. WHEN THE CUP HAD BEEN IMPLANTED, HE OPTED TO USE A 36MM INTERNAL DIAMETER LINER. ONLY 36MM LINERS THAT HAD BEEN SENT IN THE JOINTLINE LOAN KIT WERE CONTINUUM LINERS, RATHER THAN REVISION LINERS. AS A RESULT, A 58MM (LL)/36MM CONTINUUM LINER WAS OPENED AND GIVEN TO THE SURGEON AND IMPLANTED INTO THE PATIENT. AT THIS TIME THE SURGEON COMMENTED THAT THE LINER WAS VERY SMOOTH AND OPTED TO USE AN OSCILLATING SAW TO SCORE THE BACK OF THE LINER BEFORE CEMENTING IT INTO THE CUP. BEFORE DOING SO, I INFORMED THE SURGEON THAT BY SCORING THE BACK OF THE LINER, HE WOULD BECOME LIABLE FOR ANY SUBSEQUENT LINER FAILURE. THE SURGEON ACKNOWLEDGED WHAT I HAD SAID, AND CONTINUED TO SCORE THE CONTINUUM LINER BEFORE CEMENTING IT INTO THE TM REVISION CUP, WHERE HE REPORTED IT SAT PROUD BY APPROXIMATELY 1 TO 2MM. THE SURGEON WAS HAPPY THAT THE OPERATION HE HAD PERFORMED WAS SUCCESSFUL IN SPITE OF THE INCORRECT LINER BEING USED. THE ARTICULATING SURFACE IS CORRECT AND THE CEMENTING OF THE LINER INTO THE SHELL WAS EXPECTED. BY "ROUGHING UP" THE OUTER SURFACE OF THE LINER, THE SURGEON BELIEVES THAT THE IMPLANT WILL WORK EXACTLY AS REQUIRED FOR THIS PATIENT." | EVALUATION SUMMARY: THE PACKAGE INSERT FOR THIS DEVICE STATES, "ZIMMER IS NOT LIABLE FOR COMPLICATIONS ARISING FROM THE USE OF THE DEVICE OUTSIDE OF ITS INDICATED USES, SURGICAL TECHNIQUE OR JUDGMENT, PRODUCT SELECTION, AND SIMILAR MATTERS OUTSIDE THE CONTROL OF ZIMMER." NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. THERE IS NO ALLEGED FAILURE OR ADVERSE EVENT REPORTED, AND THIS IS A CASE OF USING THE DEVICE OUTSIDE OF IT'S INDICATED USE. | N | Y | HEALTH PROFESSIONAL |  | 2010-03
1822565-2010-00187 | 2010-03-24 | 2009-12-02 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ZIMMER M/L TAPER HIP PROSTHESIS SYSTEM RASP | HIP INSTRUMENT | ZIMMER, INC. | NA | 00771200460 |  | IT IS REPORTED THAT THE BROACH BROKE OFF IN THE PT'S CANAL AND WAS NOT NOTICED UNTIL THE POST-OP X-RAYS. | EVAL SUMMARY: IT IS UNK WHETHER PROPER SURGICAL TECHNIQUE WAS FOLLOWED FOR THE RASPING OF THE FEMORAL CANAL. ADDITIONALLY, EXACT FIELD AGE OF THE INSTRUMENT, FREQUENCY OF USE, AND PT FACTORS WHICH MAY CONTRIBUTE TO INSTRUMENT PERFORMANCE SUCH AS BONE QUALITY AND AGE ARE NOT KNOWN. BASED ON THE KNOWN INFO AND OBSERVATIONS, THE BROACH FRACTURE WAS MOST LIKELY CAUSED BY REPEATED OFF-AXIS LOADING. IT IS ALSO POSSIBLE THAT THE PT BONE QUALITY, DEBRIS IN THE INTRAMEDULLARY CANAL, OR DULLNESS OF THE INSTRUMENT MAY HAVE CONTRIBUTED TO THE FRACTURE. HOWEVER, THE EXACT CAUSE OF THIS SITUATION CANNOT BE DETERMINED WITH CERTAINTY AT THIS TIME. AS RETURNED, THE TIP OF THE TAPERED RASP HAS FRACTURED OFF THE DEVICE AND WAS LEFT IN THE PT AS INDICATED ON THE PER DEVICE RECORDS INDICATE ALL COMPONENTS WERE MANUFACTURED AND INSPECTED TO SPECIFICATION. SEM ANALYSIS WAS PERFORMED ON THE FRACTURED SURFACE WHICH INDICATED THAT THE INSTRUMENT FAILED DUE TO REVERSE BENDING FATIGUE LOADING. THE RASP WAS DIMENSIONALLY ANALYZED AND FOUND TO BE CONFORMING. THE TEETH ON THE RASP WERE FOUND TO EXHIBIT SOME DAMAGE. THE RASP HAD A POTENTIAL FIELD AGE OF APPROXIMATELY 17 MONTHS AT THE TIME OF FAILURE. | Y | N | HEALTH PROFESSIONAL |  | 2010-03
1822565-2010-00186 | 2010-03-24 | 2010-02-17 | I, n, j, u, r, y | Manufacturer report | TRILOGY ACETABULAR SYSTEM LONGEVITY CROSSLINKED POLYETHYLENE LINER | HIP PROSTHESIS | ZIMMER, INC. | NA | 00632005032 |  | REVISION DUE TO DISLOCATION. | (B) (4). EVAL SUMMARY: THE ORIGINAL TRILOGY LONGEVITY LINER WAS REMOVED AND RETURNED. AS RETURNED, THE LINER HAS A RIM FRACTURE NEAR THE ELEVATION TRANSITION POINT. THE LINER ALSO HAS NUMEROUS MARKS ON THE BACKSIDE OF THE ELEVATED PORTION WHICH MAY HAVE BEEN CAUSED DURING REMOVAL. PT DETAILS OTHER THAN NAME WERE NOT PROVIDED. SEM ANALYSIS INDICATES THE FRACTURE OCCURRED UNDER REPEATED LOADING. LINER FRACTURE CAN BE DUE TO (BUT NOT LIMITED TO) ONE OR MORE OF THE FOLLOWING: PT WEIGHT; PT ACTIVITY; INCORRECT COMPONENT POSITIONING; HEAD-NECK IMPINGEMENT; NON-RECOMMENDED HEAD/LINER/SHELL/STEM COMBINATIONS. DISLOCATION CAN BE DUE TO (BUT NOT LIMITED TO) ONE OR MORE OF THE FOLLOWING: INCORRECT COMPONENT POSITIONING; FEMORAL COMPONENT IMPINGEMENT; PT ANATOMY PRONE TO DISLOCATION; IMPROPER ANATOMICAL POSITION ASSUMED BY PT. BASED ON THE PROVIDED INFO, THE EXACT CAUSE OF LINER FRACTURE CANNOT BE DETERMINED WITH CERTAINTY AT THIS TIME. DEVICE HISTORY RECORDS INDICATE ALL COMPONENTS WERE MFG AND INSPECTED TO SPECIFICATION. SCANNING ELECTRON MICROSCOPY (SEM) ANALYSIS / ENERGY DISPERSIVE SPECTROSCOPY (EDS) ANALYSIS WAS PERFORMED. NO MFG ABNORMALITIES COULD BE DETECTED BY EITHER METHOD. | Y | N | HEALTH PROFESSIONAL |  | 2010-03
1822565-2010-00185 | 2010-03-24 | 2010-02-08 | I, n, j, u, r, y | Manufacturer report | UNK ZMR HIP SYSTEM FEMORAL | HIP PROSTHESIS | ZIMMER, INC. | NA | UNK |  | FRACTURE OF BROKEN ZMR STEM. REVISING HOSP WILL NOT RELEASE STEM FOR REVIEW. | (B) (4). EVAL SUMMARY: NO INFO IS INCLUDED REGARDING PART NUMBERS OR LOT NUMBERS. NO X-RAYS ARE INCLUDED WHICH WOULD ALLOW ANALYSIS OF THE LEVEL OF PROXIMAL SUPPORT. PARTS ARE NOT INCLUDED WHICH WOULD ALLOW SEM ANALYSIS TO DETERMINE MECHANICAL FAILURE MODE. NO INFO IS INCLUDED OUTLINING PERTINENT PAST MEDICAL HISTORY OR SEQUENCE OF EVENTS LEADING TO FRACTURE OF THE STEM. THERE IS INSUFFICIENT INFO TO DETERMINE ROOT CAUSE. EVALUATION: NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS WAS ALSO NOT POSSIBLE AS THE PRODUCT AND/OR LOT NUMBERS REQUIRED FOR RETRIEVAL WERE UNAVAILABLE. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE AVAILABLE INFO, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFO BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-03
1822565-2010-00183 | 2010-03-24 | 2009-12-01 | I, n, j, u, r, y | Manufacturer report | TRILOGY ACETABULAR SYSTEM LONGEVITY CROSSLINKED POLYETHYLENE LINER | HIP PROSTHESIS | ZIMMER, INC. | NA | 00630505032 |  | PT HAD LEFT TOTAL HIP ARTHROPLASTY. REVISION WAS PERFORMED DUE TO PAIN; IMAGING INDICATED THAT THE POLYETHYLENE LINER HAD UNLOCKED AND SLID OUT OF POSITION, LEAVING THE JOINT IN A SUBLUXED POSITION. INTRA-OPERATIVELY, THE LINER WAS FOUND TO HAVE SPLIT. THE CUP AND LINER WERE REPLACED. | (B) (4). EVAL SUMMARY: IT IS UNK WHETHER THE LINER DISLOCATED FROM THE CUP BEFORE OR AFTER THE LINER FRACTURED. NO X-RAYS ARE INCLUDED; THEREFORE, EVAL OF THE CUP AND LINER POSITION IS NOT POSSIBLE. THE LINER WAS NOT RETURNED; THEREFORE, ANALYSIS OF THE FRACTURE IS NOT POSSIBLE. ADDITIONALLY, NO INFO IS PROVIDED ON THE ASSOCIATED COMPONENTS, SO PRODUCT COMPATIBILITY CANNOT BE CONFIRMED. THE ROOT CAUSE OF THE LINER CRACKING AND SUBSEQUENT FRACTURE CANNOT BE DETERMINED FROM THE INFO PROVIDED. EVALUATION: NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE AVAILABLE INFO, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFO BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER INC CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-03
1822565-2010-00182 | 2010-03-24 | 2009-07-31 | I, n, j, u, r, y | Manufacturer report | VERSYS HIP SYSTEM EPOCH FULLCOAT HIP PROSTHESIS FEMORAL STEM FIBER METAL | HIP PROSTHESIS | ZIMMER, INC. | NA | 00408802006 |  | PER ADVERSE EVENT FORM, PT EXPERIENCED A FRACTURE OF THE FEMORAL SHAFT WHICH RESULTED IN REVISION OF THE FEMORAL STEM AND HEAD. | (B) (4). EVAL SUMMARY: PT INFO INDICATES PAST MEDICAL HISTORY OF OSTEOPOROSIS, AND PRIOR FAILED FRACTURE TREATMENT USING OPEN REDUCTION WITH INTERNAL FIXATION OF THE FEMORAL NECK ON THE SAME SIDE. THIS INDICATES THAT THE BONE HAD PREVIOUSLY DEMONSTRATED WEAKNESS LEADING TO FRACTURE OF THE FEMUR. ALSO, THE SURGICAL DEVICE INFO PAGE OF THE PT INFO INDICATES THAT THERE WAS MEDIAL, POSTERIOR, AND ANTERIOR DEFICIENCY OF THE FEMORAL BONE IN THE PROXIMAL REGION. THE OPERATIVE INFO PAGE OF THE PT INFO INDICATES A DORR TYPE C CLASSIFICATION (INCLUDES DRAMATIC A/P CORTICAL THINNING, AND WIDE INTRAMEDULLARY CANAL). STEM UTILIZED IN THIS SURGERY WAS A STANDARD VERSYS EPOCH FULLCOAT. SURGICAL TECHNIQUE (B) (4) STATES THAT THESE STEMS ARE BEST FOR DORR TYPE A AND B FEMURS. THE SURGICAL TECHNIQUE STATES THAT VERSYS EPOCH FULLCOAT RN I, STEMS ARE THE BEST CONFIGURATION FOR DORR TYPE C. IT IS UNK WHETHER THIS INFO WAS CONSIDERED IN PRE-OPERATIVE PLANNING. THE INDICATION FROM THE ADVERSE EVENT REPORT IS THAT THE EVENT WAS NOT RELATED TO THE DEVICE. NO X-RAYS ARE INCLUDED WHICH WOULD ALLOW EVAL OF THE FEMORAL SHAPE. WHILE THE PT INFO INDICATES A DORR TYPE C FEMUR, AND THE SURGICAL TECHNIQUE INDICATES THAT THE STEM SELECTED IS A BETTER FIT FOR DORR TYPE A AND U FEMURS, THE MOST LIKELY CAUSE OF THE FEMORAL FRACTURE IS THE PT'S OWN OSTEOPOROSIS. THIS IS SUPPORTED BY LENGTH OF TIME FROM THE SURGERY TO THE EVENT, AS WELL AS THE INFO CONTAINED IN THE ADVERSE EVENT REPORT AND THE PT INFO. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS WAS ALSO NOT POSSIBLE AS THE LOT NUMBER REQUIRED FOR RETRIEVAL WERE UNAVAILABLE. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE AVAILABLE INFO, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFO BE RECEIVED, THE COMPLAINT WILL BE REOPENED. | Y | N | HEALTH PROFESSIONAL |  | 2010-03
1822565-2010-00192 | 2010-03-25 | 2010-01-08 | I, n, j, u, r, y | Manufacturer report | TRILOGY ACETABULAR SYSTEM LINER | HIP PROSTHESIS | ZIMMER, INC. | NA | 00612004628 |  | REVISION DUE TO LUXATION. WHEN IMPLANT WAS REMOVED, THERE WERE SCRATCHES AND DELAMINATION OF THE INNER SURFACE OF THE LINER DESPITE THE FACT THAT THE LINER HAD ONLY BEEN IN THE PATIENT 1.5 MONTHS. | EVALUATION SUMMARY: WHEN IMPLANT WAS REMOVED, THERE WERE SCRATCHES AND DELAMINATION OF THE INNER SURFACE OF THE INSERT, DESPITE THE LINER HAVING ONLY BEEN IN PATIENT 1-1/2 MONTHS. THE RETURNED LINER SHOWS CONSIDERABLE SCORING ON THE ARTICULAR SURFACE AND A CIRCULAR AREA OF DELAMINATION OPPOSITE THE ELEVATION. THERE IS ALSO CONSIDERABLE DAMAGE ON THE FACE OF THE ELEVATION. IN ADDITION, THERE IS DAMAGE UNDER THE LIP OF THE ELEVATION AND BELOW THE LIP AT THE LOCKING GROOVE (WHICH MAY HAVE OCCURRED DURING LINER REMOVAL). THE X-RAYS SHOW THE FEMORAL HEAD DISLOCATED FROM THE LINER. IN ADDITION, THE SHELL APPEARS TO BE ORIENTED AT LESS THAN 45 DEGREE (MORE HORIZONTAL). THE CONDITIONS OF THE STEM, FEMORAL HEAD, AND ACETABULAR SHELL AT THE TIME OF REVISION ARE UNKNOWN. ONE OR A COMBINATION OF THE FOLLOWING MECHANISMS MAY HAVE CONTRIBUTED TO THE DISLOCATION OBSERVED: UNSATISFACTORY PLACEMENT OF THE COMPONENT PARTS. EXTENT OF COMPONENT STABILITY. EXTENT OF SOFT TISSUE TENSION. PATIENT BEHAVIOR. REPEATED LUXATION OF THE FEMORAL HEAD INTO AND OUT OF THE LINER MAY HAVE GENERATED THIRD PARTY DEBRIS, WHICH IN COMBINATION WITH THE BIOMECHANICAL FORCES EXPERIENCED BY THE LINER AS A RESULT, LUXATION MAY HAVE CONTRIBUTED TO THE DAMAGE OBSERVED. DEFINITIVE CAUSE CANNOT BE DETERMINED. DEVICE HISTORY RECORDS INDICATE ALL COMPONENTS WERE MANUFACTURED AND INSPECTED TO SPECIFICATION. SCANNING ELECTRON MICROSCOPY (SEM) ANALYSIS / ENERGY DISPERSIVE SPECTROSCOPY (EDS) ANALYSIS WAS PERFORMED. NO MANUFACTURING ABNORMALITIES COULD BE DETECTED BY EITHER METHOD. | Y | N | HEALTH PROFESSIONAL |  | 2010-03
1822565-2010-00195 | 2010-03-25 | 2009-05-07 | I, n, j, u, r, y | Manufacturer report | TRILOGY ACETABULAR SYSTEM LONGEVITY CROSSLINKED POLYETHYLENE LINER STANDARD | HIP PROSTHESIS | ZIMMER, INC. | NA | 00630506040 |  | SITE REPORTS THAT PATIENT HAD AN ADVERSE EVENT OF "WOUND DRAINAGE" WHICH RESULTED IN TREATMENT OF I&D AND ACETABULAR LINER EXCHANGE. | EVALUATION SUMMARY: THE DEVICE WAS IN-VIVO APPROXIMATELY 7 DAYS. IT IS UNKNOWN WHETHER THE CORRECT SURGICAL TECHNIQUE WAS FOLLOWED FOR THE SEATING OF THE LINER IN THE SHELL. DETAILS ON THE INSTRUMENTS USED FOR IMPLANTATION ARE UNAVAILABLE. ADDITIONALLY, IT IS UNKNOWN WHICH DEVICES WERE USED IN COMBINATION WITH THE LINER AND WHAT CONDITION THESE COMPONENTS WERE IN. THE STERILIZATION PROCESS FOR THIS DEVICE WAS VALIDATED IN ACCORDANCE WITH FDA REGULATIONS AND ISO STANDARDS TO A STERILITY ASSURANCE LEVEL OF 1.0 X 10-6 OR BETTER. IN ADDITION, THE MANUFACTURING LOT SPECIFIED IN THIS COMPLAINT WAS PROCESSED ACCORDING TO THE VALIDATED STERILIZATION PROCESS PARAMETERS AND MET ALL THE ACCEPTANCE CRITERIA FOR STERILITY RELEASE. THEREFORE, IT IS HIGHLY UNLIKELY THAT THE SPECIFIED DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT INFECTION. BASED ON THE INFORMATION PROVIDED, THE CAUSE OF THIS EVENT WAS NOT LIKELY RELATED TO THE DEVICE. HOWEVER, THE EXACT CAUSE CANNOT BE DETERMINED. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-03
9613350-2010-00174 | 2010-04-01 | 2010-03-08 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 52/46 H | DUROM ACETABULAR COMPONENT AND METASUL LDH | ZIMMER GMBH |  | 01.00214.152 |  | IT IS REPORTED BY THE PATIENT THAT THE PATIENT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2008. POST-OP PATIENT EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION IS SCHEDULED. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00179 | 2010-04-01 | 2010-03-09 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 52/46 L | DUROM ACETABULAR COMPONENT AND METASUL LDH | ZIMMER GMBH |  | 0100214152 |  | IT IS REPORTED BY PT'S ATTORNEY THAT PT UNDERWENT A HIP ARTHROPLASTY IN (B) (6) 2007. POST-OP, PT EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. REVISION IS SCHEDULED. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER ((B) (4)), WHICH MARKETS THE DEVICE IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00178 | 2010-04-01 | 2010-02-08 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 60/54 T | DUROM ACETABULAR COMPONENT AND METASUL LDH | ZIMMER GMBH |  | 0100214160 |  | IT IS REPORTED BY PT THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, PT EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2010 | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00177 | 2010-04-01 | 2010-03-02 | I, n, j, u, r, y | Manufacturer report | DUROM US ACTETABULAR COMPONENT | DUROM ACETABULAR COMPONENT AND METASUL LDH | ZIMMER GMBH |  | UNK |  | IT IS REPORTED BY PT THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2006. POST-OP, PT EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. A REVISION IS SCHEDULED. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00176 | 2010-04-01 | 2010-03-02 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 54/48 H | DUROM ACETABULAR COMPONENT AND METASUL LDH | ZIMMER GMBH |  | 01.00214.154 |  | IT WAS REPORTED THAT PT WAS DX WITH ACETABULAR IMPLANT LOOSENING WHICH RESULTED IN A REVISION. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00175 | 2010-04-01 | 2008-07-08 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 52/46 H | DUROM ACETABULAR COMPONENT AND METASUL LDH | ZIMMER GMBH |  | 01.00214.152 |  | IT IS REPORTED BY PT'S ATTORNEY THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2008. POST-OP, PT EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND PT WAS REVISED ON (B) (6) 2008. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00172 | 2010-04-01 | 2010-03-05 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 54/48 H | DUROM ACETABULAR COMPONENT AND METASUL LDH | ZIMMER GMBH |  | 01.00214.154 |  | IT IS REPORTED BY THE PATIENT ATTORNEY THAT THE PATIENT UNDERWENT A RIGHT HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION IS SCHEDULED. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00171 | 2010-04-01 | 2010-03-04 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 54/48 H | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH |  | 01.00214.154 |  | IT IS REPORTED BY THE PATIENT ATTORNEY THAT THE PATIENT UNDERWENT A LEFT HIP ARTHROPLASTY ON (B) (6) 2008. POST-OP HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION IS SCHEDULED. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4)CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00173 | 2010-04-01 | 2013-05-23 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 54/48 H | DUROM ACETABULAR COMPONENT AND METASUL LDH | ZIMMER GMBH |  | 01.00214.154 |  | IT IS REPORTED BY THE PATIENT ATTORNEY THAT THE PATIENT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION IS SCHEDULED. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00169 | 2010-04-01 | 2010-03-05 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 56/50 H | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH |  | 01.00214.156 |  | IT IS REPORTED BY THE PATIENT ATTORNEY THAT THE PATIENT UNDERWENT A LEFT HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION IS SCHEDULED. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00168 | 2010-04-01 | 2010-03-02 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 48/42 H | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH |  | 01.00214.148 |  | IT IS REPORTED BY THE PATIENT ATTORNEY THAT THE PATIENT UNDERWENT A LEFT ARTHROPLASTY ON (B) (6) 2007. POST-OP SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2009. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00167 | 2010-04-01 | 2010-03-02 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 44/38 H | DUROM ACETABULAR COMPONENT AND METASUL LDH | ZIMMER GMBH |  | 01.00214.144 |  | IT IS REPORTED BY PATIENT THAT PATIENT UNDERWENT A LEFT HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND REVISION IS SCHEDULED FOR (B) (6) 2010. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00134 | 2010-04-01 | 2010-02-26 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 58/52 R | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH |  | 01.00214.158 |  | IT WAS REPORTED THAT PATIENT PRESENTED AND WAS DX WITH ACETABULAR IMPLANT LOOSENING WHICH RESULTED IN DEVICE REMOVAL (FEMORAL HEAD, ACETABULAR COMPONENT AND HEAD ADAPTER). | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00170 | 2010-04-01 | 2010-03-05 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 48/42 H | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH |  | 01.00214.148 |  | IT IS REPORTED BY THE PATIENT ATTORNEY THAT THE PATIENT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2006. POST-OP SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION IS SCHEDULED. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00164 | 2010-04-06 | 2008-03-19 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 56/50 P | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.156 |  | IT IS REPORTED BY PT'S ATTORNEY THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2006. POST OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. A REVISION TOOK PLACE ON (B) (6) 2008. | INFO WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN (B) (4). THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00147 | 2010-04-06 | 2010-03-15 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 56/50 P | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.156 |  | IT IS REPORTED BY PATIENT THAT PATIENT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2007. POST OP, HE EXPERIENCES PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION IS SCHEDULED. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00148 | 2010-04-06 | 2010-03-09 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | NA |  | IT IS REPORTED BY PATIENT'S ATTORNEY THAT PATIENT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2006. POST OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. A REVISION IS SCHEDULED. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC (B) (4), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00146 | 2010-04-06 | 2010-03-15 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.154 |  | IT IS REPORTED BY PATIENT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2007. POST OP, PATIENT EXPERIENCES PAIN DE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION IS SCHEDULED. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00145 | 2010-04-06 | 2010-03-09 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 58/52 R | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.158 |  | IT IS REPORTED BY PATIENT'S ATTORNEY THAT PATIENT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2006. POST OP, PATIENT EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. A REVISION IS SCHEDULED. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICE, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00166 | 2010-04-06 | 2010-03-24 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 54/48 N | DRUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.154 |  | IT IS REPORTED BY PATIENT'S ATTORNEY THAT PATIENT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, SHE EXPERIENCES PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. A REVISION HAS NOT TAKEN PLACE. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN (B) (4). THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00165 | 2010-04-06 | 2010-03-24 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 52/46 L | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.152 |  | IT IS REPORTED BY PATIENT'S ATTORNEY THAT PATIENT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2008. POST OP, HE EXPERIENCES PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. A REVISION HAS NOT BEEN SCHEDULED. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN (B) (4). THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00162 | 2010-04-06 | 2009-12-17 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.154 |  | IT IS REPORTED BY PT'S ATTORNEY, THE PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2007. POST OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. A REVISION TOOK PLACE ON (B) (6) 2009. | INFO WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICE IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00163 | 2010-04-06 | 2009-11-05 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 52/46 L | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.152 |  | IT IS REPORTED BY PT'S ATTORNEY THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2009. POST OP, HE EXPERIENCES PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. A REVISION TOOK PLACE ON (B) (6) 2009. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN (B) (4). NO PRODUCT WAS RETURNED FOR EVAL. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF A DEVICE DEFICIENCY CAUSING OR CONTRIBUTING TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL INFO BE RECEIVED THAT CHANGES THIS CONCLUSION, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00160 | 2010-04-06 | 2010-03-17 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 52/46 L | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.152 |  | IT IS REPORTED BY PT THAT PT UNDERWENT A HIP ARTHROPLASTY IN (B) (6) 2007. POST OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION IS SCHEDULED FOR (B) (6) 2010 WITH DR (B) (6). | INFO WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00159 | 2010-04-06 | 2010-03-18 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 58/52 R | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.158 |  | IT IS REPORTED BY PT'S ATTORNEY THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2007. POST OP, HE EXPERIENCES PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. IT IS UNK WHETHER A REVISION IS PLANNED OR HAS TAKEN PLACE. | INFO WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00158 | 2010-04-06 | 2008-10-01 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 56/50 P | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.156 |  | IT IS REPORTED BY PT THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) OF 2007. POST-OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION TOOK PLACE IN (B) (6) OF 2008. | INFO WS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICE IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00157 | 2010-04-06 | 2008-08-25 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 50/44 J | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.150 |  | IT IS REPORTED BY PT'S ATTORNEY THAT PT UNDERWENT A RIGHT HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2008. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER ((B) (4)5), WHICH MARKETS THE DEVICE IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00156 | 2010-04-06 | 2010-01-29 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 52/46 L | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.152 |  | IT IS REPORTED BY PT'S ATTORNEY THAT PT UNDERWENT A LEFT HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2010. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER ((B) (4)), WHICH MARKETS THE DEVICE IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00152 | 2010-04-06 | 2009-09-24 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.154 |  | IT IS REPORTED BY PATIENT'S ATTORNEY THE PATIENT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2008. POST OP, SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND REVISION TOOK PLACE ON (B) (6). | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
2648920-2010-00014 | 2010-04-06 | 2010-02-25 | I, n, j, u, r, y | Manufacturer report | TRILOGY ACETABULAR SYSTEM SHELL WITH CLUSTER | HIP PROSTHESIS | ZIMMER | NA | 00620005622 |  | CUP SPUN OUT OF PT'S ACETABULUM LEAVING THE SCREW IN PLACE IN THE PT. THE SCREW HEAD PULLED THROUGH THE HOLE AND REMAINED IN THE PT. | (B) (6). EVAL SUMMARY: WITH THE INFO PROVIDED, IT IS NOT POSSIBLE TO DEFINITIVELY CONCLUDE THE CAUSE OF DEVICE FAILURE. PT INFO SUCH AS HEIGHT, WEIGHT, PT'S BONE QUALITY, AND ACTIVITY LEVEL WERE NOT REPORTED. ADDITIONALLY, SURGICAL NOTES WERE NOT PROVIDED, SO THE AMOUNT OF PRESS FIT, IF ANY, USED DURING SHELL INSERTION IS UNK. CONSIDERING THE DAMAGE TO THE SCREW HEAD AND THE FIBER METAL SHELL COATING NEAR THE SCREW HOLE, IT SEEMS LIKELY THAT EXCESSIVE TORQUE WAS APPLIED TO THE SCREW DURING INSERTION. ANOTHER POSSIBILITY IS THAT, POST-OPERATIVELY, AN EXCESSIVE LOADING OF THE SHELL CAUSED THE SHELL TO ROTATE WITHIN THE ACETABULUM WHICH, BECAUSE THE SCREW WAS WELL-FIXED IN BONE, FORCED THE SCREW TO PULL-THROUGH THE SCREW HOLE. WITHOUT MORE INFO, THE PROBABLE CAUSE CANNOT BE DETERMINED. EVAL: DEVICE HISTORY RECORDS INDICATE ALL COMPONENTS WERE MFG AND INSPECTED TO SPEC. SCANNING ELECTRON MICROSCOPY (SEM) ANALYSIS/ENERGY DISPERSIVE SPECTROSCOPY (EDS) ANALYSIS WAS PERFORMED. NO MFG ABNORMALITIES COULD BE DETECTED BY EITHER METHOD. | Y | N | HEALTH PROFESSIONAL |  | 2010-04
1822565-2010-00209 | 2010-04-06 | 2006-08-21 | I, n, j, u, r, y | Manufacturer report | VERSYS HIP SYSTEM FEMORAL STEM FIBER METAL | HIP PROSTHESIS | ZIMMER, INC. | NA | 00784101500 |  | IT IS REPORTED BY PT'S ATTORNEY THAT PT UNDERWENT A RIGHT HIP ARTHROPLASTY ON (B) (6) 1999. POST OP, HE EXPERIENCED PAIN DUE TO THE SUBSIDING OF THE STEM INTO THE FEMUR. A REVISION TOOK PLACE ON (B) (6) 2006. | (B) (4). EVAL SUMMARY: NO PRODUCT HAS BEEN RETURNED FOR EVAL. NO X-RAY AFTER PRIMARY SURGERY SHOWING THE ORIENTATION OF THE CONSTRUCT HAS BEEN MADE AVAILABLE. HOWEVER, SUBSIDENCE OF FEMORAL STEM MAY BE CAUSED DUE TO THE LACK OF BIOLOGICAL FIXATION BETWEEN THE FEMORAL STEM AND THE BONE. THE PAIN MIGHT HAVE BEEN CAUSED DUE TO THE INSTABILITY OF THE STEM. NO INFO HAS BEEN PROVIDED REGARDING THE CONDITION OF THE BIPOLAR SHELL/LINER AFTER RETRIEVAL. BASED ON THE LIMITED INFO AVAILABLE, NO DEFINITE CAUSE CAN BE DETERMINED WITH CERTAINTY. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADD'L SUBSTANTIVE INFO BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-04
9613350-2010-00161 | 2010-04-06 | 2008-03-23 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.154 |  | IT IS REPORTED BY PT THAT PT UNDERWENT A RIGHT HIP ARTHROPLASTY IN (B) (6) OF 2009. POST OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. A REVISION HAS BEEN SCHEDULED. | INFO WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN (B) (4). NO PRODUCT WAS RETURNED FOR EVAL. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF A DEVICE DEFICIENCY CAUSING OR CONTRIBUTING TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL INFO BE RECEIVED THAT CHANGES THIS CONCLUSION, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00153 | 2010-04-06 | 2010-03-18 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.154 |  | IT IS REPORTED BY PATIENT'S ATTORNEY THAT PATIENT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2008. POST OP, HE EXPERIENCES PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. IT IS UNKNOWN WHETHER A REVISION IS PLANNED OR HAS TAKEN PLACE. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00149 | 2010-04-06 | 2010-03-08 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 48/42 H | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.148 |  | DR (B) (6) REVISED THE DUROM COMPONENT FOR ASEPTIC LOOSENING ON (B) (6) 2010. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC (B) (4), WHICH MARKETS THE DEVICES IN THE US. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00155 | 2010-04-06 | 2008-05-22 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.154 |  | IT IS REPORTED BY PATIENT'S ATTORNEY THAT PATIENT UNDERWENT A RIGHT HIP ARTHROPLASTY ON (B) (6) 2007. POST-OP, SHE EXPERIENCED PAIN AND WAS REVISED ON (B) (6) 2008. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00154 | 2010-04-06 | 2010-02-05 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.154 |  | IT IS REPORTED BY PATIENT THAT PATIENT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2006. POST OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2010. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. ((B) (4)), WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
1822565-2010-00202 | 2010-04-06 | 2010-02-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | ZMR HIP SYSTEM OTJ METAL PROVISIONAL BODY | HIP INSTRUMENT | ZIMMER INC. | NA | 00999604645 |  | IT IS REPORTED THAT AFTER INSERTING THE DISTAL PORTION OF THE STEM, THE SURGEON WAS TRIALING WITH THE METAL (B)(4) TRIALS. HE TAPPED THE TRIAL ON THE STEM. AFTER HE TRIALED THE BODY, HE WANTED TO CHANGE THE VERSION. HOWEVER, HE COULD NOT LOOSEN THE JACK SCREW. THE METAL BODY TRIAL WOULD NOT RELEASE FROM THE TAPER. IT WAS UNABLE TO BE REMOVED ON THE BACK TABLE EITHER. | EVAL SUMMARY: THE PER STATES THAT THE SURGEON "TAPPED" THE TRIAL ONTO THE STEM. THIS AGREES WITH THE PROCEDURE DESCRIBED IN THE SURGICAL TECHNIQUE. IT IS MOST LIKELY THAT THE SURGEON OVERESTIMATED THE STRENGTH OF THE "TAP" NEEDED TO TEMPORARILY FIX THE PROVISIONAL BODY ONTO THE STEM. THE SURGEON APPEARS TO HAVE FOLLOWED THE METHOD SUGGESTED BY THE SURGICAL TECHNIQUE TO FREE THE STUCK ASSEMBLY AND RESOLVE ISSUE. EVALUATION: AS RETURNED THE STEM IS SEIZED WITH IN THE PROXIMAL TRIAL. CORROSION IS NOTED AT THE JUNCTION BETWEEN THE TRIAL AND THE STEM. IT IS NOT KNOWN, HOWEVER, WHEN THIS CORROSION WAS FIRST EVIDENT. DEVICE HISTORY RECORDS FOR BOTH THE STEM AND THE PROXIMAL TRIAL WERE REVIEWED AND NO ANOMALIES WERE NOTED. THE TRIAL HAD A POTENTIAL FIELD AGE OF APPROX 15 MONTHS AT THE TIME OF THE INCIDENT. NO FURTHER INFO IS AVAILABLE AT THIS TIME. | N | Y | HEALTH PROFESSIONAL |  | 2010-04
1822565-2010-00211 | 2010-04-07 | 2010-03-09 | I, n, j, u, r, y | Manufacturer report | CONTINUUM ACETABULAR SYSTEM TRABECULAR METAL | HIP PROSTHESIS | ZIMMER, INC. | NA | 00875705001 |  | SURGEON FIT AND SEATED CUP AND USED ONE SCREW TO AFFIX CUP, USED THE CENTERPULSE LINER IMPACTOR TO SEAT THE LINER. IT WOULD NOT SEAT. THE DOCTOR THEN USED THE TRILOGY METAL IMPACTOR AND THE LINER WOULD NOT SEAT. THE DOCTOR THEN USED THE CONTINUUM IMPACTOR WITH THE 32 HEAD. THE LINER PULLED OUT OF THE CUP AND STUCK TO THE IMPACTOR WHEN THE DOCTOR REMOVED THE IMPACTOR. THE DOCTOR REFUSED TO OPEN AND TRY TO INSTALL THE SECOND LINER. THE CUP WAS REMOVED AND A DIFFERENT LINER WAS INSTALLED. | EVALUATION SUMMARY: THE ACETABULAR LINER WAS VISUALLY EXAMINED AND NO DAMAGE WAS DETECTED THAT WOULD BE EXPECTED IF THE LINER WAS ATTEMPTED TO BE LOCKED INTO THE ACETABULAR SHELL AND IMPACTED. ADDITIONALLY, THE ACETABULAR SHELL EXHIBITED MINIMAL AMOUNTS OF BONE ADHERED TO THE TRABECULAR METAL. THE RETURNED IMPLANTS WERE FOUND TO BE CONFORMING AND WERE ASSEMBLED PROPERLY. AT THIS TIME, A NONCONFORMANCE COULD NOT BE FOUND WITH THE RETURNED IMPLANTS. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | N | Y | HEALTH PROFESSIONAL |  | 2010-04
9613350-2010-00144 | 2010-04-07 | 2001-08-29 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 62/56 V | DUROM ACETABULAR COMPONENT AND METASUL LDH | ZIMMER GMBH |  | 01.00214.162 |  | IT IS REPORTED THAT REVISION SURGERY OCCURRED DUE DISLOCATION. THE DISLOCATION OCCURRED WHEN THE PATIENT ATTEMPTED TO CROSS HIS LEGS AT ONE DAY POST-OP. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER, WHICH MARKETS THE DEVICES IN U.S. NO PRODUCT WAS RETURNED FOR EVALUATION. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF A DEVICE DEFICIENCY CAUSING OR CONTRIBUTING TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL INFORMATION BE RECEIVED THAT CHANGES THIS CONCLUSION, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. ZIMMER CONSIDERS THE INVESTIGATION CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
1822565-2010-00210 | 2010-04-07 | 2010-03-01 | I, n, j, u, r, y | Manufacturer report | ZIMMER M/L TAPER HIP PROSTHESIS WITH KINECTIV | HIP PROSTHESIS | ZIMMER, INC. | NA | 00771301000 |  | PATIENT PRESENTED AND WAS DIAGNOSED WITH ACUTE INFECTION OF THE RIGHT HIP S/P RIGHT THA WHICH RESULTED IN REVISION OF ALL COMPONENTS. | EVALUATION SUMMARY: THE STERILIZATION PROCESS FOR THIS DEVICE WAS VALIDATED IN ACCORDANCE WITH FDA REGULATIONS AND ISO STANDARDS TO A STERILITY ASSURANCE LEVEL OF 1.0 10-6 OR BETTER. IN ADDITION, THE MANUFACTURING LOT SPECIFIED IN THIS COMPLAINT WAS PROCESSED ACCORDING TO THE VALIDATED STERILIZATION PROCESS PARAMETERS AND MET ALL THE ACCEPTANCE CRITERIA FOR STERILITY RELEASE. THEREFORE, IT IS HIGHLY UNLIKELY THAT THE SPECIFIED DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT INFECTION. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-04
MW5015524 | 2010-04-10 | 2009-12-31 | I, n, j, u, r, y | Voluntary report | ZIMMER DUROM ACETABULUM | HIP JOINT | ZIMMER | DUROM |  |  | SEVERE PAIN IN HIP PROSTHESIS, ZIMMER DUROM ACETABULUM, CONSTANT ACHING. DATES OF USE: (B) (6) 2007 - (B) (6) 2010. DIAGNOSIS: TOTAL HIP ARTHROPLASTY. |  | Y | Y | 0 |  | 2010-04
1822565-2010-00219 | 2010-04-13 | 2010-02-01 | I, n, j, u, r, y | Manufacturer report | VERSYS HIP SYSTEM FEMORAL STEM BEADED FULLCOAT | HIP PROSTHESIS | ZIMMER, INC. | NA | 00784301106 |  | THE PT HAS A HISTORY OF FALLING AND DISLOCATING. THE TOTAL WAS DONE SOMEWHERE IN TEXAS. IT LOOKED TO BE VERY PROUD. | EVAL SUMMARY: THE RECEIVED DEVICES WERE REVIEWED VIA BOTH VISUAL INSPECTION AND SEM ANALYSIS. SEM ON THE FRACTURED SURFACE ON THE PROXIMAL PORTION OF THE STEM SHOWED CRYSTALLOGRAPHIC FRACTURE AND FATIGUE. THE EXACT TIME IN-VIVO IS UNK. THE MATING ACETABULAR SHELL AND LINER ARE UNK, AND THEIR CONDITIONS ARE ALSO UNK. THE SURGICAL NOTES FROM THE PRIMARY SURGERY AND FROM SUCCESSIVE PT VISITS ARE UNAVAILABLE, AND IT IS UNK IF THE STEM WAS IMPLANTED WITH THE CORRECT ORIENTATION AND FIT AS PER THE SURGICAL TECHNIQUE. X-RAY IMAGES POST-PRIMARY SURGERY AND FROM SUCCESSIVE PT VISITS ARE UNAVAILABLE. PT HEIGHT AND ACTIVITY LEVEL ARE ALSO UNK. BASED ON THE ABOVE INFO AND OBSERVATIONS, THE FEMORAL STEM FRACTURE MAY HAVE BEEN DUE TO ONE OR A COMBINATION OF THE FOLLOWING MECHANISMS: EXTREME LOADING DUE TO PT WEIGHT AND/OR PT ACTIVITY; DISTAL IN-GROWTH WITH LACK OF PROXIMAL IN-GROWTH AND/OR SUPPORT; STEMS LOOSENING; COMPONENT MISALIGNMENT. HOWEVER, THE EXACT CAUSE OF THE CURRENT SITUATION CANNOT BE CONCLUSIVELY DETERMINED. EVAL: DEVICE HISTORY RECORDS INDICATE ALL COMPONENTS WERE MANUFACTURED AND INSPECTED TO SPECIFICATION. SCANNING ELECTRON MICROSCOPY (SEM) ANALYSIS / ENERGY DISPERSIVE SPECTROSCOPY (EDS) ANALYSIS WAS PERFORMED. NO MFG ABNORMALITIES COULD BE DETECTED BY EITHER METHOD. | Y | N | HEALTH PROFESSIONAL |  | 2010-04
9613350-2010-00180 | 2010-04-14 |  | I, n, j, u, r, y | Manufacturer report | DUROM HIP | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH | NA | UNK |  | IT IS REPORTED BY THE PATIENT THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6), 2007. POST OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION IS SCHEDULED FOR (B) (6), 2010. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
1822565-2010-00227 | 2010-04-14 | 2008-09-05 | I, n, j, u, r, y | Manufacturer report | TRILOGY ACETABULAR SYSTEM CONSTRAINED LINER | HIP PROSTHESIS | ZIMMER, INC. | NA | 00615405032 |  | IT IS REPORTED BY PATIENT'S ATTORNEY THAT THE PATIENT UNDERWENT A RIGHT HIP ARTHROPLASTY ON (B) (6) 2004. POST OP, SHE EXPERIENCED PAIN DUE TO A DISLOCATION AND FRACTURE OF REINFORCING RING. A REVISION TOOK PLACE ON (B) (6) 2008. SUBSEQUENTLY, SEVERAL ADDITIONAL DISLOCATIONS HAVE TAKEN PLACE. HOWEVER, NO REVISION HAS BEEN SCHEDULED. | (B) (4). THE CONSTRAINED LINER AND FEMORAL HEAD WERE IN VIVO APPROXIMATELY 4 YEARS. THE PATIENT SUFFERED MULTIPLE DISLOCATIONS OF THE RIGHT HIP DURING THIS PERIOD. IT IS UNKNOWN WHETHER PROPER SURGICAL TECHNIQUE WAS FOLLOWED. ADDITIONALLY, INFORMATION ON THE FEMORAL STEM AND ACETABULAR SHELL USED IN COMBINATION WITH THE ALLEGED PARTS AND WHETHER THEY WERE REPLACED IS NOT AVAILABLE. INSTRUMENTS USED, X-RAY IMAGES AND SURGICAL NOTES WERE NOT PROVIDED. BASED ON THE INFORMATION GIVEN AND THE ABOVE OBSERVATIONS, THE REVISION MAY HAVE BEEN CAUSED BY ONE OR A COMBINATION OF THE FOLLOWING OCCURRENCES: REDUCTION OF THE FEMORAL HEAD AFTER THE RESTRAINING RING WAS ASSEMBLED; IMPROPER ASSEMBLY OF THE CONSTRAINING RING AND LINER UPON IMPLANTATION; MULTIPLE REDUCTIONS OF THE FEMORAL HEAD INTO THE LINER PRIOR TO THE RESTRAINING RING ASSEMBLY, DAMAGING THE COMPONENTS; INCORRECT POSITIONING OF THE SHELL UPON IMPLANTATION; MUSCLE AND FIBROUS TISSUE LAXITY AROUND THE HIP; PATIENT FACTORS SUCH AS BEHAVIOR AND/OR HIGH PATIENT BMI. PLEASE NOTE PATIENT IS A (B) (6) FEMALE, (B) (6), WITH A PAST MEDICAL HISTORY OF ATRIAL FIBRILLATION WITH RVR, PRIMARY PULMONARY HYPERTENSION, HTN, ARTHRITIS, DVT, PE, ORAL THRUSH, ASTHMA, HYPOVENTILATION ASSOCIATED WITH OBESITY SYNDROME, C. DIFF DIARRHEA, OSTEOARTHRITIS, RHEUMATOID ARTHRITIS, ACUTE PANCREATITIS, AND SPINAL STENOSIS. HOWEVER, THE EXACT CAUSE OF THIS SITUATION CANNOT BE CONCLUSIVELY DETERMINED. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. | Y | N | HEALTH PROFESSIONAL |  | 2010-04
1822565-2010-00226 | 2010-04-14 | 2009-05-30 | I, n, j, u, r, y | Manufacturer report | VERSYS HIP SYSTEM FEMORAL STEM BEADED FULLCOAT | HIP PROSTHESIS | ZIMMER, INC. | NA | 00784301316 |  | TWO MONTHS AFTER SURGERY, I STARTED TO EXPERIENCE PAIN IN MY QUAD FEMUR AREA, I BELIEVE THAT IT IS WHAT IS REFERRED TO AS START UP PAIN. TWO DOCTORS HAVE NOW CONFIRMED THAT IT IS A LOOSE STEM. I HAVE HAD SEVERAL BLOOD DRAWS AND FLUID EXTRACTION WHICH WAS CONFIRMED THAT NO INFECTION WAS PRESENT. I AM SCHEDULED FOR SURGERY ON (B) (6) 2010. | (B) (4). EVALUATION SUMMARY: IMPLANT WAS NOT RETURNED. ORIGINAL PROCEDURE REPORT DOCUMENTS SURGEON. | Y | N | HEALTH PROFESSIONAL |  | 2010-04
9613350-2010-00075 | 2010-04-14 | 2010-03-09 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULUM COMPONENT 56/50 CODE P | DUROM ACETABULAR COMPONENT AND METASUL LDH LA | ZIMMER GMBH |  | 01.00214.056 |  | IT IS REPORTED THAT A REVISION SURGERY OCCURRED DUE TO A POSSIBLE ALLERGIC REACTION AND PAIN. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER (B) (4), WHICH MARKETS THE DEVICES IN (B) (4). AN AMENDED FOLLOW-UP REPORT WIL BE SUBMITTED AS SOON AS THE DEVICES HAVE BEEN RECEIVED AND INVESTIGATION IS COMPLETE. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
1822565-2010-00231 | 2010-04-15 | 2010-03-18 | I, n, j, u, r, y | Manufacturer report | TRILOGY ACETABULAR SYSTEM LONGEVITY CROSSLINK | HIP PROSTHESIS | ZIMMER, INC. | NA | 00631005632 |  | REVISED DUE TO GROIN PAIN. | (B) (4). EVALUATION SUMMARY: THE DEVICES WERE IN-VIVO FOR APPROXIMATELY 2.5 YEARS. NO MANUFACTURING LOT INFORMATION WAS AVAILABLE FOR MANUFACTURING EVALUATION FOR THE VERSYS HA/TCP POROUS FEMORAL STEM. THE DEVICES WERE UNAVAILABLE FOR EVALUATION AND THE DEVICE CONDITIONS ARE UNKNOWN. THE MATING ACETABULAR SHELL IS UNKNOWN, AND ITS CONDITION IS ALSO UNKNOWN. THE SURGICAL NOTES FROM THE PRIMARY SURGERY AND FROM THE REVISION SURGERY ARE UNAVAILABLE, AND IT IS UNKNOWN IF THE COMPONENTS WERE IMPLANTED AS ACCORDING TO THE SURGICAL TECHNIQUE. X-RAY IMAGES POST-PRIMARY SURGERY AND FROM SUCCESSIVE PATIENT VISITS ARE UNAVAILABLE. PATIENT HEIGHT, WEIGHT, AND ACTIVITY LEVEL ARE UNKNOWN. DUE TO LACK OF SUFFICIENT INFORMATION, THE PROBABLE CAUSE OF PAIN IS UNKNOWN. EVALUATION: NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED ZIMMER, INC CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-04
1822565-2010-00234 | 2010-04-15 | 2010-02-19 | I, n, j, u, r, y | Manufacturer report | TRILOGY ACETABULAR SYSTEM CONSTRAINED LINER | HIP PROSTHESIS | ZIMMER, INC. | NA | 00615405832 |  | IT WAS REPORTED THAT ABOVE MENTIONED FEMORAL HEAD WAS DISLOCATED FROM THE CONSTRAINED LINER WITHOUT DISSOCIATION OF THE CONSTRAINED RING. | THE DEVICE WAS IN VIVO APPROXIMATELY 2 MONTHS. NO DEVICES WERE RETURNED FOR EVALUATION, THEREFORE, THEIR CONDITIONS ARE UNKNOWN. IT IS UNKNOWN WHETHER PROPER TECHNIQUE WAS FOLLOWED FOR THE CORRECT ORIENTATION OF THE ACETABULAR SHELL AND THE ASSEMBLY OF THE RING, FEMORAL HEAD AND LINER. NO INFORMATION ON INSTRUMENTS USED DURING THE IMPLANTATION WAS PROVIDED. PATIENT FACTORS SUCH AS HEIGHT/WEIGHT, BEHAVIOR, REHABILITATION PROTOCOL AND WHETHER THE PROTOCOL WAS FOLLOWED WERE NOT PROVIDED. SURGICAL NOTES AND X-RAYS WERE UNAVAILABLE. BASED ON THE ABOVE INFORMATION AND OBSERVATIONS, THE REVISION DUE TO PAIN AND CHRONIC DISLOCATION MAY HAVE BEEN CAUSED BY ONE OR A COMBINATION OF THE FOLLOWING OCCURRENCES: INCORRECT POSITIONING OF THE SHELL UPON IMPLANTATION; IMPROPER ASSEMBLY OF THE RESTRAINING RING AND LINER UPON IMPLANTATION; MULTIPLE REDUCTIONS OF THE FEMORAL HEAD INTO THE LINER PRIOR TO THE RESTRAINING RING ASSEMBLY, POTENTIALLY DAMAGING THE COMPONENTS; REDUCTION OF THE FEMORAL HEAD AFTER THE RESTRAINING RING WAS ASSEMBLED; MUSCLE AND FIBROUS TISSUE LAXITY AROUND THE HIP COULD POTENTIALLY LEAD TO INSUFFICIENT FUNCTIONAL SUPPORT TO THE JOINT; PATIENT FACTORS SUCH AS PATIENT BEHAVIOR. HOWEVER, THE EXACT CAUSE OF THIS SITUATION CANNOT BE CONCLUSIVELY DETERMINED. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. | Y | N | HEALTH PROFESSIONAL |  | 2010-04
9613350-2010-00181 | 2010-04-15 | 2010-03-15 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 60/54 T | NONE | ZIMMER GMBH | NA | 01.00214.160 |  | IT IS REPORTED THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2007. POST OP, HE EXPERIENCED PAIN DUE TO ACETABULAR LOOSENING. A REVISION TOOK PLACE ON (B) (6) 2010. | INFO WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC., WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-04
9613350-2010-00182 | 2010-04-15 | 2010-02-28 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABULAR COMPONENT 56/50 P | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH | NA | 01.00214.156 |  | IT WAS REPORTED THAT THE PT WAS REVISED DUE TO A LOOSE ACETABULAR COMPONENT. | INFO WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC., WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | HEALTH PROFESSIONAL |  | 2010-04
2648920-2010-00016 | 2010-04-20 | 2010-03-24 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | MULTIPOLAR BIPOLAR CUP LINER | HIP PROSTHESIS | ZIMMER | NA | 00500104222 |  | IT IS REPORTED THAT CRACK LINES DEVELOPED IN THE LINER WHILE IMPACTING WHICH PROLONGED SURGERY BY 30 MINUTES. | THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. | N | Y | HEALTH PROFESSIONAL |  | 2010-04
1822565-2010-00240 | 2010-04-21 | 2010-03-18 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | CONTINUUM | HIP PROSTHESIS | ZIMMER, INC. | NA | 00875101440 |  | CONTINUUM CUP PERFECTLY IMPLANTED, INTRAOPERATIVELY, IT WAS IMPOSSIBLE TO STABILIZE PE INLAY INTO THE CUP: THE FEELING WAS THAT THE CUP WAS SLIGHTLY BIGGER THAN PE. OF COURSE SIZES OF BOTH COMPONENT WERE DOUBLE CHECKED. SURGERY ENDED CHANGING INLAY WITH A METASUL ONE. | THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETED. | N | Y | HEALTH PROFESSIONAL |  | 2010-04
1822565-2010-00243 | 2010-04-21 | 2010-03-23 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY ACETABULAR SYSTEM CUP POSITIONER | HIP INSTRUMENT | ZIMMER, INC. | NA | 00626001800 |  | IT IS REPORTED THAT DURING THE SURGERY, THE POSITIONER DISASSEMBLED, THEN ITS SET SCREW AND PIN DROPPED INTO THE WOUND. THE PIECES WERE RETRIEVED. | EVALUATION SUMMARY: HAD THE SET SCREW BEEN INADEQUATELY THREADED INTO THE SPRING PLUNGER BODY DURING MANUFACTURE, IT WOULD LIKELY HAVE FALLEN APART MUCH EARLIER IN THE LIFE OF THE INSTRUMENT. THE DURATION OF USE EFFECTIVE USE OF THE INSTRUMENT MAKES ASSEMBLY ERROR UNLIKELY. THE SET SCREW AND SPRING PLUNGER BODY WERE REASSEMBLED BY DEVELOPMENT AS A PART OF THIS INVESTIGATION, AND THE SET SCREW THREADS INTO BODY TO FULL DEPTH AND WITH PROPER FIT. THIS MAKES DEFECTIVE SUBCOMPONENTS UNLIKELY. A REVIEW OF THE COMPLAINT HISTORY FOR THIS INSTRUMENT SHOWS NO OTHER EVENTS OF THE SPRING PLUNGER DISASSEMBLING DURING USE, MAKING INADEQUATE DESIGN AN UNLIKELY CONTRIBUTING FACTOR FOR THIS PARTICULAR EVENT. THE MOST LIKELY CAUSE OF THE SPRING PLUNGER OF THE CUP POSITIONER DISASSEMBLING DURING USE CANNOT BE DETERMINED FROM THE INFORMATION PROVIDED. HOWEVER, VIBRATION FROM NORMAL USE AS WELL AS POSSIBLE LACK OF ANY MEANS PRESENT TO LOCK THE SET SCREW IN PLACE IN THE SPRING PLUNGER MAY BE CONTRIBUTING FACTORS. AS RETURNED, THE SPRING FROM INSIDE THE SPRING PLUNGER IS MISSING, HOWEVER, THE PHOTO FROM THE COMPLAINT FILE IN (B)(4) SHOWS THE SPRING FROM THE SPRING PLUNGER WAS CAPTURED INSIDE THE COMPRESSION SPRING IN THE END OF THE CUP POSITIONER. MANUFACTURING RECORDS FOR THIS LOT WERE REVIEWED BY AND FOUND TO BE CONFIRMING. THE FIELD AGE OF THE PART IS UP TO 19 MONTHS. | N | Y | HEALTH PROFESSIONAL |  | 2010-04
9613350-2010-00184 | 2010-04-22 |  | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 48/42 H | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH | NA | 01.00214.148 |  | EVENT TIMING: AFTER SURGERY. IT WAS REPORTED THAT A REVISION SURGERY IS BEING SCHEDULED DUE TO PAIN AND SUSPECTED LOOSENING OF THE DUROM CUP. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT IT IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | HEALTH PROFESSIONAL |  | 2010-04
9613350-2010-00186 | 2010-04-22 | 2010-03-18 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 52/46 CODE L | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH | NA | 0100214052 |  | IT WAS REPORTED THAT PATIENT HAD TO UNDERGO A REVISION SURGERY DUE TO A PSEUDO TUMOR. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT YET RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, OR WHEN THE PRODUCTS HAVE BEEN RETURNED AND EVALUATED IN OUR LABS AND INVESTIGATION RESULTS ARE AVAILABLE, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | HEALTH PROFESSIONAL |  | 2010-04
9613350-2010-00185 | 2010-04-22 |  | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 58/52 R | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH | NA | 0100214158 |  | IT IS REPORTED THAT PATIENT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2008. POST OP, HE EXPERIENCES PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. THE STATUS OF REVISION IS UNKNOWN. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC. WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THIS IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | HEALTH PROFESSIONAL |  | 2010-04
9613350-2010-00183 | 2010-04-22 | 2015-06-02 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH | NA | 01.00214.154 |  | A PRODUCT LIABILITY CLAIM WAS RAISED. IT HAS NOW BEEN REPORTED THAT THE PATIENT WAS IMPLANTED A DUROM ACETABULAR COMPONENT 54/48 N ON THE RIGHT SIDE. THE PATIENT WAS REVISED ON (B)(6) 2015 DUE TO PAIN, METALLOSIS, ELEVATED METAL IONS, PSEUDOTUMOR, CORROSION. | ADDITIONAL INFORMATION WAS RECEIVED ON (B)(4) 2017. THE DEVICES WERE RETURNED AND THE RESULT OF THE VISUAL EXAMINATION HAS BEEN MADE AVAILABLE. DHR REVIEW: THE QUALITY RECORDS INDICATE THAT THESE COMPONENTS MET ALL REQUIREMENTS TO PERFORM AS INTENDED. REVIEW OF EVENT DESCRIPTION: PATIENT UNDERWENT REVISION DUE TO METALLOSIS, PSEUDOTUMOR, CORROSION, ELEVATED METAL IONS. REVIEW OF RECEIVED DATA SURGICAL REPORT : REVIEW OF SURGICAL REPORT DID NOT LEAD TO NEW INFORMATION REGARDING THE REPORTED EVENT. OTHER INFORMATION & SOURCES: REVIEW OF THE COMPLAINT RELEVANT DOCUMENTS DID NOT LEAD TO NEW INFORMATION REGARDING THE REPORTED EVENT. COMPLEX PHYSIOLOGICAL REACTIONS LIKE PSEUDOTUMOR OR METAL ALLERGY ARE KNOWN RISKS FOR THIS KIND OF METAL-ON-METAL IMPLANTS AS STATED IN ZIMMER¿S ""INSTRUCTION LEAFLET FOR ENDOPROSTHESIS"" DEVICES ANALYSIS VISUAL EXAMINATION: FOLLOWING COMPONENTS HAVE BEEN RETURNED FOR INVESTIGATION: DUROM CUP AND LDH HEAD AND HEAD ADAPTER ALL RECEIVED COMPONENTS SHOW DAMAGE FROM REVISION SURGERY SUCH AS NICKS AND SCRATCHES. ADDITIONALLY, FOLLOWING OBSERVATIONS ARE DONE: THE DUROM CUP SHOWS NONE BONE ONGROWTH ON THE ANCHORING SIDE. THE INNER SURFACE OF THE HEAD ADAPTER SHOWS SOME SURFACE CHANGES IN FORM OF FRETTING CORROSION THE HEAD TAPER SURFACE IS INCONSPICUOUS. CONCLUSION: NO FURTHER INVESTIGATION REQUIRED AS THIS ISSUE IS KNOWN (ERROR PATTERN: POTENTIAL EARLY REVISION OF THE ACETABULAR COMPONENT DUE TO LOOSENING, IMPLANT MIGRATION OR UNRESOLVED PAIN, HIGHER ION RELEASE). AT LEAST ONE OF THESE ERROR PATTERNS IS OBSERVED IN THIS EVENT. THE RESULT OF THE EXAMINATION DID NOT CHANGE THE RESULTS OF THE PREVIOUS ASSESSMENT IN WHICH ZIMMER IMPLEMENTED A NOTIFICATION IN JULY 2008. THE DISTRIBUTION OF THIS PRODUCT WAS CEASED IN THE MEANWHILE. THEREFORE, ZIMMER GMBH CONSIDERS THIS CASE AS CLOSED AGAIN. ZIMMER¿S REFERENCE NUMBER OF THIS FILE IS (B)(4). | Y | Y | HEALTH PROFESSIONAL |  | 2010-04
9613350-2010-00187 | 2010-04-23 | 2009-11-17 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 56/50 P | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH | NA | 0100214156 |  | IT IS REPORTED THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6), 2008. POST OP, HE EXPERIENCES PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. A REVISION TOOK PLACE ON (B) (6), 2009. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC., WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AND AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | HEALTH PROFESSIONAL |  | 2010-04
9613350-2010-00188 | 2010-04-23 | 2008-07-08 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 50/44 J | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH | NA | 0100214150 |  | IT WAS REPORTED THAT A REVISION SURGERY OCCURRED DUE TO IMPLANT LOOSENING. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC., WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FORM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | HEALTH PROFESSIONAL |  | 2010-04
9613350-2010-00189 | 2010-04-23 | 2010-03-10 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 46/40 CODE F | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH | NA | 0100214046 |  | IT IS REPORTED THAT PT WAS REVISED DUE TO METALLOSIS. | INFO WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED OR OTHER SOURCE DOCUMENTS FOR REVIEW WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADD'L INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B)(4) CONSIDERS THIS CASE AS CLOSED. (B)(4). | Y | N | HEALTH PROFESSIONAL |  | 2010-04
1822565-2010-00257 | 2010-04-27 | 2010-03-12 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TOTAL HIP INSTRUMENTATION STRAIGHT SHELL INSERT | HIP INSTRUMENT | ZIMMER, INC. | NA | 00780401520 |  | IT IS REPORTED THAT UPON IMPACTION OF THE CONTINUUM ACETABULAR COMPONENT, THE DISTAL THREADS OF THE CUP INSERTER BROKE OFF AND THE CUP WAS DISENGAGED. | EVALUATION SUMMARY: IT IS UNKNOWN WHETHER PROPER SURGICAL TECHNIQUE WAS FOLLOWED FOR INSERTING THE CUP INTO THE ACETABULUM. ADDITIONALLY, INFORMATION ON FREQUENCY OF USE IS UNKNOWN. BASED ON AVAILABLE INFORMATION AND OBSERVATIONS, THE DISTAL THREAD BREAKAGE MAY HAVE BEEN DUE TO ONE OR A COMBINATION OF THE FOLLOWING OCCURRENCES: THE DEVICE WAS USED WITHOUT AN ADAPTER, WHICH COULD HAVE LED TO HIGH STRESS ON THE THREADS CAUSING THEM TO BREAK. INSUFFICIENT THREAD ENGAGEMENT DURING IMPACTION MAY HAVE LED TO HIGH STRESS ON THE DISTAL THREADS CAUSING THEM TO BREAK. OFF-AXIS IMPACTION OF THE COMPONENT. THE DEVICE WAS INADVERTENTLY DROPPED OR DAMAGED DURING CLEANING. HOWEVER, NO DEFINITIVE CAUSE ANALYSIS CAN BE COMPLETED WITH CERTAINTY. AS RETURNED, THE HEX BOLT APPEARS TO BE JAMMED BACK INTO THE INSTRUMENT CAUSING LESS EXPOSURE OF THREADS. THE HEX BOLT MAY BE UP AGAINST THE WELD PIN CAUSING THE HEX BOLT TO BE JAMMED. THE THREADS ON THE INSTRUMENT DO NOT APPEAR TO BE DAMAGED AS ALLEGED IN THE COMPLAINT. THE INSTRUMENT HAD A FIELD AGE OF APPROXIMATELY 16 MONTHS AT THE TIME OF RETURN. THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. AN ADDITIONAL CONTRIBUTING FACTOR MAY BE ASSOCIATED WITH DIFFICULTY ASSEMBLING THE CONTINUUM ACETABULAR SYSTEMS SHELL INSERTER ADAPTER ONTO THE CUP (THE ADAPTER GOES BETWEEN THE INSERTER AND CUP). ON 03/19/2010, ZIMMER INITIATED A DEVICE REMOVAL OF THE ADAPTER VIA ZIMMER RECALL NUMBER 1822565-03/02-2010-003-R. | N | Y | HEALTH PROFESSIONAL |  | 2010-04
1822565-2010-00248 | 2010-04-27 | 2009-12-14 | I, n, j, u, r, y | Manufacturer report | ZIMMER LONGEVITY LINER STANDARD | HIP PROSTHESIS | ZIMMER, INC. | NA | 00630505032 |  | PATIENT HAD LEFT TOTAL HIP ARTHOPLASTY. REVISION IN 2009 DUE TO PAIN. IMAGING INDICATED THAT THE POLYETHYLENE LINER HAD UNLOCKED AND SLID OUT OF POSITION, LEAVING THE JOINT IN A SUBLUXED POSITION. IN THE OPERATING ROOM, FOUND LINER TO HAVE SPLIT. CUP AND LINER REPLACED. | (B) (4). EVALUATION SUMMARY: NO DEVICES WERE RETURNED FOR EVALUATION, THEREFORE THE DEVICE CONDITION IS UNKNOWN. IT IS UNKNOWN WHETHER PROPER SURGICAL TECHNIQUE WAS FOLLOWED. TIME IN VIVO WAS APPROXIMATELY 1 MONTH. WHETHER ACETABULAR SCREWS WERE USED IS UNKNOWN. NO PHOTOS OR X-RAYS WERE RETURNED FOR EVALUATION. INFORMATION ON INSTRUMENTATION USED, MATING ACETABULAR SHELL, FEMORAL HEAD AND FEMORAL STEM, AND PATIENT FACTORS SUCH AS REHABILITATION PROTOCOL, LEVEL OF COMPLIANCE WITH REHABILITATION PROTOCOL, AND PATIENT ACTIVITY LEVEL IS UNAVAILABLE. WITH THE INFORMATION GIVEN, THE EXACT CAUSE OF THIS SITUATION CANNOT BE DETERMINED WITH CERTAINTY AT THIS TIME. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES, IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE RE-OPENED ZIMMER INC CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-04
1822565-2010-00251 | 2010-04-27 | 2010-03-26 | I, n, j, u, r, y | Manufacturer report | TRILOGY ACETABULAR SYSTEMS LINER 10 DEG ELEV | HIP PROTHESIS | ZIMMER, INC. | NA | 00611004826 |  | IT IS REPORTED THAT THE SURGEON WHO REVISED A POLY LINER IS CLAIMING THAT IT HAS WEAR ON ITS BACK SIDE. | (B) (4). THIS REPORT WILL BE AMENDED WHEN OUR INVESTIGATION IS COMPLETE. | Y | N | HEALTH PROFESSIONAL |  | 2010-04
9613350-2010-00196 | 2010-04-29 | 2010-04-06 | I, n, j, u, r, y | Manufacturer report | DUROM ACETABULAR COMPONENT 62/56 CODE V | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH | NA | 0100214062 |  | IT IS REPORTED THAT A REVISION SURGERY HAS OCCURRED DUE TO MOBILIZATION OF THE CUP. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC., WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT YET RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AND/OR THE SAMPLES BE RETURNED AND AN EVALUATION RESULT BE AVAILABLE, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | HEALTH PROFESSIONAL |  | 2010-04
9613350-2010-00197 | 2010-04-29 | 2009-10-01 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 52/46 L | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH | NA | 0100214162 |  | IT IS REPORTED THAT PATIENT UNDERWENT A REVISION SURGERY ON (B) (6) 2009, DUE TO PAIN AND POSSIBLE ACETABULAR LOOSENING. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC., WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | HEALTH PROFESSIONAL |  | 2010-04
9613350-2010-00195 | 2010-04-29 |  | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 56/50 P | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH | NA | 0100214156 |  | IT IS REPORTED THAT PATIENT UNDERWENT A RIGHT HIP ARTHROPLASTY ON (B) (6) 2008. POST OP, HE EXPERIENCES PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. IT IS UNK WHETHER A REVISION IS PLANNED OR HAS TAKEN PLACE. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC., WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THIS IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | HEALTH PROFESSIONAL |  | 2010-04
9613350-2010-00193 | 2010-04-29 |  | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 58/52 R | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH | NA | 0100214158 |  | IT IS REPORTED THAT PATIENT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2007. POST OP, SHE EXPERIENCES PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION IS SCHEDULED FOR (B) (6) 2010. | THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THIS IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AND/OR THE COMPLAINT SAMPLES BE RETURNED AND AN EVAL REPORT BE AVAILABLE, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | HEALTH PROFESSIONAL |  | 2010-04
9613350-2010-00192 | 2010-04-29 | 2009-04-24 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH | NA | 0100214154 |  | IT WAS REPORTED THAT A REVISION SURGERY OCCURRED DUE TO PAIN AND POSSIBLE ACETABULAR LOOSENING. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC., WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT YET RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AND/OR THE COMPLAINT SAMPLES BE RETURNED AND EVALUATION RESULT BE AVAILABLE, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | HEALTH PROFESSIONAL |  | 2010-04
9613350-2010-00198 | 2010-04-29 | 2009-12-01 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 52/46 L | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH | NA | 0100214152 |  | IT WAS REPORTED THAT PATIENT UNDERWENT A REVISION SURGERY DUE TO SUSPECTED ACETABULAR IMPLANT LOOSENING. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC., WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | HEALTH PROFESSIONAL |  | 2010-04
9613350-2010-00194 | 2010-04-29 |  | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 56/50 P | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH | NA | 0100214156 |  | AFTER SURGERY. IT IS REPORTED THAT PATIENT UNDERWENT A LEFT HIP ARTHROPLASTY ON (B) (6) 2006. POST OP, HE EXPERIENCES PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. IT IS UNK WHETHER A REVISION IS PLANNED OR HAS TAKEN PLACE. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC., WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THIS EVENT IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | HEALTH PROFESSIONAL |  | 2010-04
2648920-2010-00017 | 2010-05-04 | 2010-03-22 | M, a, l, f, u, n, c, t, i, o, n | Manufacturer report | TRILOGY ACETABULAR SYSTEM SHELL WITH HOLES | HIP PROSTHESIS | ZIMMER | NA | 00620004820 |  | IT IS REPORTED THAT UPON IMPACTING THE LINER INTO THE CUP, THE SURGEON WAS UNABLE TO SIT THE LINER PROPERLY. IT WAS NOT POSSIBLE TO LOCK THE LINER INTO PLACE. SURGEON PROCEEDED TO REAM TO 50MM TO USE ONE SIZE UP AS THE CONSIGNMENT CARRIED ONLY ONE CUP. | OTHER DEVICE USED: CATALOG #00630504832, TRILOGY ACETABULAR SYSTEM LONGEVITY CROSSLINKED POLYETHYLENE LINER, LOT #61320591. EVAL SUMMARY: IT IS UNCLEAR IF THE SURGICAL TECHNIQUE WAS FOLLOWED SINCE NO SURGICAL NOTES WERE SUBMITTED - ONLY THE DESCRIPTION IN THE PRODUCT EXPERIENCE REPORT. ASSEMBLY OF THE RETURNED PRODUCTS WAS ATTEMPTED AND THE LOCKING RING WOULD NOT EXPAND ENOUGH TO ALLOW THE LINER TO SNAP INTO PLACE. THE TABS AT THE OPEN END OF THE RING WOULD NOT SNAP BACK INTO PLACE AND MOVE FREELY INDICATING THE LINER IS NOT SEATED. IT IS POSSIBLE THAT THE LOCKING RING WAS IN THE WAY OF THE LINER DURING LINER INSERTION AND MAY HAVE DEFORMATION OF THE RING WHICH COULD HAVE LEAD TO THE INABILITY TO SEAT THE LINER. HOWEVER, THE ROOT CAUSE OF THE INABILITY TO SEAT THE LINER CANNOT BE DETERMINED WITH CERTAINTY. THE LINER MAY HAVE BEEN AUTOCLAVED. THE POLAR BOSS WAS CUT OFF FROM THE LINER DURING THE SURGERY. THE LOCKING RING ON THE SHELL FLOATS FREELY AND DOES NOT APPEAR TO BE DAMAGED. THE INNER SPHERICAL RADIUS SHOWS SOME DAMAGE NEAR ONE OF THE SCREW HOLES. DEVICE HISTORY RECORDS INDICATE ALL COMPONENTS WERE MANUFACTURED, PACKAGED, AND INSPECTED TO SPECIFICATION. | N | Y | HEALTH PROFESSIONAL |  | 2010-05
1822565-2010-00270 | 2010-05-05 | 2010-03-15 | I, n, j, u, r, y | Manufacturer report | EPSILON METASUL STANDARD INSERT | HIP PROSTHESIS | ZIMMER, INC. | NA | 437228065 |  | A 65MM METASUL LINER WAS REMOVED AFTER 10+ YEARS DUE TO PAIN. SURGEON DESCRIBED AS METALLOSIS. A NEW LINER WAS IMPLANTED INTO THE INTEROP SHELL WITH NO PROBLEMS. SURGEON WOULD ALSO LIKE MENTIONED THAT THE EXTENSIVE BONE LOSS AND SOFT TISSUE NECROSIS OF THE ADJACENT TISSUES OF THE HIP JOINT IS VERY UNUSUAL FOR A 10 YR OLD IMPLANT WITH NORMAL WEAR. ALSO THERE WERE NO OBVIOUS SIGNS OF POLY OR METAL WEAR ON THE ARTICULAR SIDE OR BACK SIDE OF THE POLY/METAL INSERT. | (B) (4). EVALUATION SUMMARY: ACTIVITY LEVEL OF THE PATIENT IS NOT KNOWN. X-RAYS WERE NOT PROVIDED. THE CAUSE OF THE PATIENT'S PAIN REQUIRING REVISION SURGERY CANNOT BE DETERMINED FROM THE INFORMATION PROVIDED. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-05
1822565-2010-00269 | 2010-05-05 | 2010-04-08 | I, n, j, u, r, y | Manufacturer report | TRILOGY ACETABULAR SYSTEM LONGEVITY LINER | HIP PROSTHESIS | ZIMMER, INC. | NA | 00620004620 |  | PATIENT REPORTS SHE WAS REVISED DUE TO DISLOCATION. | (B) (4). EVALUATION SUMMARY: DEVICES WERE IN VIVO APPROXIMATELY 3 MONTHS. NO PRODUCT WAS RETURNED. THE OPERATIVE NOTES FROM BOTH THE PRIMARY THR AND THE REVISION WERE NOT PROVIDED. PATIENT WEIGHT AND ACTIVITY LEVEL ARE UNKNOWN. POSSIBLE CAUSES FOR THE LINER DISASSOCIATING FROM THE SHELL INCLUDE IMPINGEMENT AND/OR IMPROPER SURGICAL TECHNIQUE RESULTING IN THE IMPROPER ASSEMBLY OF COMPONENTS DURING INITIAL SURGERY (ANTIROTATION TABS MIS-ALIGNED, LINER NOT FULLY SEATED INTO CUP, LOCKING RING UNABLE TO MOVE FREELY FROM SIDE-TO-SIDE AFTER LINER INSERTION). GIVEN THE INFORMATION PROVIDED, A DEFINITIVE CAUSE FOR THE REPORTED INCIDENT CANNOT BE DETERMINED WITH CERTAINTY. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY AND EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE INFORMATION BE RECEIVED, THE COMPLAINT WILL BE REOPENED. ZIMMER, INC. CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-05
9613350-2010-00199 | 2010-05-06 |  | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 52/46 L | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH | NA | 0100214152 |  | EVENT TIMING: AFTER SURGERY. IT IS REPORTED THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2007. POST OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. INFORMATION ABOUT A REVISION IS UNKNOWN. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC., WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THIS IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADD'L INFO BECOME AVAILABLE THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | HEALTH PROFESSIONAL |  | 2010-05
9613350-2010-00201 | 2010-05-06 | 2010-03-09 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 58/52 R | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH | N/A | 0100214158 |  | NO CHANGES IN DESCRIPTION. | ADDITIONAL INFORMATION WHICH WAS RECEIVED ON JUL 26, 2019. MEDICAL DEVICES: PART#: 0100181520 LOT#: 2359044; PART#: 0100185146 LOT#: 2363718. SINCE THIS CASE IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTIVE ACTION WHICH WAS REPORTED TO THE FDA IN JULY 2008 AS CORRECTION Z-2415/2426-2008, THERE WILL BE NO FURTHER INVESTIGATION AND ZIMMER GMBH WILL CLOSE THIS CASE ONCE AGAIN. ZIMMER¿S REFERENCE NUMBER OF THIS FILE IS (B)(4). | Y | N | HEALTH PROFESSIONAL |  | 2010-05
9613350-2010-00202 | 2010-05-06 |  | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 60/54 T | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH | NA | 0100214160 |  | IT IS REPORTED THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2007. POST OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. A REVISION IS SCHEDULED FOR (B) (6) 2010. SURGERY DETAILS OF THE REVISION SURGERY ARE YET UNKNOWN. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC., WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE IMPENDING REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADD'L INFO BECOME AVAILABLE AND/OR THE DEVICE(S) BE RETURNED FOR EVAL AND AN INVESTIGATION RESULT BE AVAILABLE, THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | HEALTH PROFESSIONAL |  | 2010-05
9613350-2010-00203 | 2010-05-06 |  | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH | NA | 0100214154 |  | IT IS REPORTED THAT PT UNDERWENT A HIP ARTHROPLASTY IN (B) (6) 2007. POST OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. A REVISION IS SCHEDULED FOR APRI(B) (6) 2010. DETAILS OF THE REVISION SURGERY YET UNKNOWN. | INFO WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC., WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE IMPENDING REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADD'L INFO BECOME AVAILABLE AND/OR THE DEVICE(S) BE RETURNED FOR EVAL AND AN INVESTIGATION RESULT BE AVAILABLE, THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | HEALTH PROFESSIONAL |  | 2010-05
9613350-2010-00204 | 2010-05-06 | 2008-11-26 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 52/46 L | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH | NA | 0100214152 |  | IT IS REPORTED THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2007. POST OP, SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2008. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC., WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADD'L INFO BECOME AVAILABLE AND/OR THE DEVICE(S) BE RETURNED FOR EVAL AND AN INVESTIGATION RESULT BE AVAILABLE, THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | HEALTH PROFESSIONAL |  | 2010-05
9613350-2010-00205 | 2010-05-06 | 2008-12-03 | I, n, j, u, r, y | Manufacturer report | DUROM HIP | NONE | ZIMMER GMBH | NA | DUROM HIP GENERIC |  | IT IS REPORTED THAT PT UNDERWENT A HIP ARTHROPLASTY ON (B) (6) 2007. POST OP, SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND WAS REVISED ON (B) (6) 2008. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC., WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THE REVISION SURGERY IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADD'L INFO BECOME AVAILABLE AND/OR THE DEVICE(S) BE RETURNED FOR EVAL AND AN INVESTIGATION RESULT BE AVAILABLE, THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | HEALTH PROFESSIONAL |  | 2010-05
9613350-2010-00206 | 2010-05-06 |  | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 52/46 L | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH | NA | 0100214152 |  | IT IS REPORTED BY PT'S ATTORNEY THAT PT UNDERWENT A LEFT HIP ARTHROPLASTY ON (B) (6) 2007. POST OP, HE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. THE STATUS OF A REVISION IS UNKNOWN. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC., WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THIS EVENT IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE AND/OR THE DEVICE(S) BE RETURNED FOR EVALUATION AND AN INVESTIGATION RESULT BE AVAILABLE, THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSE. (B) (4). | Y | N | HEALTH PROFESSIONAL |  | 2010-05
9613350-2010-00207 | 2010-05-06 |  | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH | NA | 0100214154 |  | IT IS REPORTED THAT PT UNDERWENT A RIGHT HIP ARTHROPLASTY ON (B) (6) 2008. POST OP, HE EXPERIENCES PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. THE STATUS OF A REVISION IS UNKNOWN | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT ZIMMER INC., WHICH MARKETS THE DEVICES IN THE U.S. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFO PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THIS EVENT IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADD'L INFO BECOME AVAILABLE AND/OR THE DEVICE(S) BE RETURNED FOR EVALUATION AND AN INVESTIGATION RESULT BE AVAILABLE, THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | HEALTH PROFESSIONAL |  | 2010-05
9613350-2010-00208 | 2010-05-06 |  | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 56/50 P | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH | NA | 0100214156 |  | IT IS REPORTED THAT PATIENT UNDERWENT A HIP ARTHROPLASTY IN (B) (4) OF 2007. POST OP, HE EXPERIENCES PAIN DUE TO POSSIBLE ACETABULAR LOOSENING. STATUS OF A REVISION IS UNKNOWN. | INFORMATION WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC., WHICH MARKETS THE DEVICES IN THE U.S. THE MANUFACTURER DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE RECEIVED FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THIS EVENT IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE AND/OR THE DEVICE(S) BE RETURNED FOR EVALUATION AND AN INVESTIGATION RESULT BE AVAILABLE, THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURES IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | HEALTH PROFESSIONAL |  | 2010-05
1822565-2010-00297 | 2010-05-11 | 2010-02-04 | I, n, j, u, r, y | Manufacturer report | TRILOGY LONGEVITY CONSTRAINED LINER | HIP PROSTHESIS | ZIMMER, INC. | NA | 00633406236 |  | THREE MONTHS POST-OP, PT PRESENTED IN THE ER WITH A DISLOCATION. FEMORAL HEAD HAD DISASSOCIATED FROM LINER. INTRAOPERATIVELY, IT WAS DISCOVERED THAT THE RING WAS STILL SECURED TO THE LINER, AND LINER TO THE SHELL. HEAD HAD PULLED OUT. | (B) (4). EVAL SUMMARY: PARTS WERE IN VIVO APPROXIMATELY 3 MONTHS PRIOR TO DISLOCATION. THAT, IN CONJUNCTION WITH THE DISLOCATION AND SUBSEQUENT EXPLANTATION PREVENTED ACCURATE DIMENSIONAL ANALYSIS. INFO ON THE MATING DEVICES AND INSTRUMENTATION USED WAS NOT PROVIDED. IT IS UNK IF THE CONSTRAINED LINER WAS IMPLANTED ACCORDING TO THE PROPER SURGICAL TECHNIQUE. PT FACTORS SUCH AS COMPLIANCE WITH REHABILITATION PROTOCOL ARE UNAVAILABLE. NO X-RAYS, PHOTOS, OR SURGICAL NOTES WERE RETURNED. BASED ON THE ABOVE INFO AND OBSERVATIONS, THE DISSOCIATION OF THE HEAD FROM THE LINER MAY HAVE BEEN DUE TO ONE OR A COMBINATION OF THE FOLLOWING EVENTS: INCORRECT POSITIONING OF THE SHELL UPON IMPLANTATION; IMPROPER ASSEMBLY OF THE RESTRAINING RING AND LINER; MUSCLE AND FIBROUS TISSUE LAXITY AROUND THE HIP; PT FACTORS SUCH AS ACTIVITY LEVEL OR COMPLIANCE WITH REHABILITATION PROTOCOL. HOWEVER, THE EXACT CAUSE OF THIS SITUATION CANNOT BE DETERMINED WITH CERTAINTY AT THIS TIME. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. ZIMMER, INC CONSIDERS THE INVESTIGATION CLOSED. | Y | N | HEALTH PROFESSIONAL |  | 2010-05
9613350-2010-00218 | 2010-05-12 | 2008-08-22 | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 48/42 H | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH | NA | 0100214148 |  | IT IS REPORTED BY PT THAT PT UNDERWENT A RIGHT HIP ARTHROPLASTY ON (B) (6), 2008. POST OP, SHE EXPERIENCED PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION TOOK PLACE ON (B) (6), 2008. | INFO WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC, WHICH MARKETS THE DEVICES IN THE UNITED STATES. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE REC'D FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THIS EVENT IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADD'L INFO BECOME AVAILABLE AND / OR THE DEVICE(S) BE RETURNED FOR EVAL AND AN INVESTIGATION RESULT BE AVAILABLE, THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURE IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | HEALTH PROFESSIONAL |  | 2010-05
1822565-2010-00322 | 2010-05-12 | 2010-03-11 | I, n, j, u, r, y | Manufacturer report | VERSYS REVISION FEMORAL STEM BEADED | HIP PROSTHESIS | ZIMMER, INC. | NA | 00784301408 |  | IT IS REPORTED THAT THE DEVICE WAS NOT THE CORRECT SIZE RESULTING IN THE SURGERY BEING PROLONGED BY 45 MINUTES. | (B) (4). POSSIBLE CAUSES OF IMPROPER FIT OF THE STEM INSIDE THE RASPED FEMORAL CANAL ARE USER ERROR WITH RESPECT TO THE RASP SELECTION, USER ERROR WITH RESPECT TO RASP USE, AND DULL OR BENT RASPS. THE CAUSE OF THE SURGEON'S DIFFICULTY CAN NOT BE DETERMINED FROM THE INFORMATION PROVIDED. ALL MANUFACTURING RECORDS FOR ALL THE RETURNED STEMS WERE REVIEWED AND FOUND TO BE CONFORMING. ALL RETURNED STEMS WERE DIMENSIONALLY CHECKED AND FOUND TO BE WITH SPECIFICATION. DEVICE HISTORY RECORDS INDICATE ALL COMPONENTS WERE MANUFACTURED INSPECTED, AND PACKAGED TO SPECIFICATION. | N | Y | HEALTH PROFESSIONAL |  | 2010-05
1822565-2010-00324 | 2010-05-12 | 2010-03-18 | I, n, j, u, r, y | Manufacturer report | VERSYS HIP SYSTEM FEMORAL STEM | HIP PROSTHESIS | ZIMMER, INC. | NA | 65786201600 |  | PATIENT WAS REVISED DUE TO INFECTION. | (B) (4). THE STERILIZATION PROCESS FOR THIS DEVICE WAS VALIDATED IN ACCORDANCE WITH FDA REGULATIONS AND ISO STANDARDS TO A STERILITY ASSURANCE LEVEL OF 1.0 X 10-6 OR BETTER. IN ADDITION, THE MFG LOT SPECIFIED IN THIS COMPLAINT WAS PROCESSED ACCORDING TO THE VALIDATED STERILIZATION PROCESS PARAMETERS AND MET ALL THE ACCEPTANCE CRITERIA FOR STERILITY RELEASE. THEREFORE, IT IS HIGHLY UNLIKELY THAT THE SPECIFIED DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT INFECTION. IT WAS NOTED THAT FOR THREE OF THE FOUR DEVICES (ALL EXCEPT FOR THE METAL ACETABULAR SHELL) WAS IMPLANTED ON A DATE THAT WAS PAST THE EXPIRATION DATE CLEARLY MARKED ON THE BOXES. ZIMMER EXPECTS THAT ALL PRODUCTS THAT ARE LABELED WITH AN EXPIRATION DATE WILL NOT BE USED IF THAT DATE IS EXCEEDED. NO PRODUCT WAS RETURNED. REVIEW OF THE DEVICE HISTORY RECORDS DID NOT FIND ANY DEVIATIONS OR ANOMALIES. IT IS NOT SUSPECTED THAT THE PRODUCT FAILED TO MEET SPECIFICATIONS. THE INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. BASED ON THE INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. SHOULD ADDITIONAL SUBSTANTIVE... | Y | N | HEALTH PROFESSIONAL |  | 2010-05
9613350-2010-00219 | 2010-05-12 |  | I, n, j, u, r, y | Manufacturer report | DUROM US ACETABAR COMPONENT 54/48 N | DUROM ACETABULAR COMPONENT AND METASUL | ZIMMER GMBH | NA | 0100214154 |  | IT IS REPORTED THAT PT UNDERWENT A HIP ARTHROPLASTY IN (B) (6) OF 2008. POST OP, HE EXPERIENCES PAIN DUE TO POSSIBLE ACETABULAR LOOSENING AND A REVISION IS SCHEDULED. | INFO WAS FORWARDED FROM THE OWNER ESTABLISHMENT, ZIMMER INC, WHICH MARKETS THE DEVICES IN THE UNITED STATES. THE MFR DID NOT RECEIVE EXPLANTED DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW. WHERE LOT NUMBERS WERE REC'D FOR THE EXPLANTED DEVICES, THE DEVICE HISTORY RECORDS WERE REVIEWED AND FOUND TO BE CONFORMING. WHILE A CAUSE FOR THIS SPECIFIC CLINICAL EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED, THE COMMON CLINICAL PRESENTATION AND THE DATE OF THE ORIGINAL IMPLANTATION SUGGEST THAT THIS EVENT IS RELATED TO THE ISSUES FOR WHICH ZIMMER IMPLEMENTED A CORRECTION IN JULY 2008. SHOULD ADD'L INFO BECOME AVAILABLE AND / OR THE DEVICE(S) BE RETURNED FOR EVAL AND AN INVESTIGATION RESULT BE AVAILABLE, THAT CHANGES THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE FILED. THE NEED FOR FURTHER CORRECTIVE MEASURE IS NOT INDICATED AT THIS TIME AND ZIMMER (B) (4) CONSIDERS THIS CASE CLOSED. (B) (4). | Y | N | LAY USER/PATIENT |  | 2010-05

## Monthly Summary
Month | Count | Date
2009-11 | 15 | 2009-11-01
2009-12 | 61 | 2009-12-01
2010-01 | 29 | 2010-01-01
2010-02 | 101 | 2010-02-01
2010-03 | 64 | 2010-03-01
2010-04 | 70 | 2010-04-01
2010-05 | 17 | 2010-05-01

## Event Types
Event Type | Count
I, n, j, u, r, y | 320
M, a, l, f, u, n, c, t, i, o, n | 34
O, t, h, e, r | 3

## Brand Names
Brand Name | Count
DUROM US ACETABULAR COMPONENT 56/50 P | 25
DUROM US ACETABULAR COMPONENT 54/48 N | 23
DUROM US ACETABULAR COMPONENT 52/46 L | 23
DUROM US ACETABULAR COMPONENT 50/44 J | 17
DUROM US ACETABULAR COMPONENT 48/42 H | 15
DUROM US ACETABULAR COMPONENT 58/52 R | 12
DUROM US ACETABAR COMPONENT 54/48 N | 10
DUROM ACETABULAR COMPONENT | 8
DUROM ACETABULAR COMPONENT 50/44 CODE J | 7
DUROM ACETABULAR COMPONENT 48/42 CODE H | 7
DUROM US ACETABULAR COMPONENT 60/54 T | 7
DUROM US ACETABAR COMPONENT 52/46 L | 6
DUROM US ACETABAR COMPONENT 58/52 R | 6
TRILOGY ACETABULAR SYSTEM LONGEVITY CROSSLINKED POLYETHYLENE LINER | 6
DUROM US ACETABAR COMPONENT 56/50 P | 5
DUROM ACETABULAR COMPONENT 52/46 CODE L | 5
TRABECULAR METAL MODULAR ACETABULAR SYSTEM SHELL WITH CLUSTER HOLES POROUS | 5
DUROM US ACETABULAR COMPONENT 46/40 F | 5
DUROM ACETABULAR COMPONENT 54/48 CODE N | 5
DUROM US ACETABULAR COMPONENT 54/48 H | 4
VERSYS HIP SYSTEM EPOCH FULLCOAT HIP PROSTHESIS FEMORAL STEM FIBER METAL | 4
DUROM US ACETABAR COMPONENT 48/42 H | 4
DUROM US ACETABULAR COMPONENT 62/56 V | 4
DUROM ACETABULAR COMPONENT 60/54 CODE T | 4
DUROM ACETABULAR COMPONENT 58/52 CODE R | 3
ZIMMER M/L TAPER HIP PROSTHESIS WITH KINECTIV TECHNOLOGY MODULAR FEMORAL STEM PR | 3
ZMR HIP SYSTEM FEMORAL STEM REVISION TAPER W/COMPRESSION NUT | 3
UNKNOWN APS NATURAL HIP STEM | 3
VERSYS HIP SYSTEM FEMORAL STEM BEADED FULLCOAT | 2
KINECTIV TECHNOLOGY MODULAR NECK | 2
ZIMMER M/L TAPER HIP PROSTHESIS WITH KINECTIV TECHNOLOGY | 2
DUROM ACETABULAR COMPONENT 56/50 CODE P | 2
TRILOGY ACETABULAR SYSTEM SHELL WITH CLUSTER HOLES POROUS | 2
TRILOGY ACETABULAR SYSTEM CUP POSITIONER | 2
TRILOGY ACETABULAR SYSTEM LONGEVITY CROSSLINKED POLYETHYLENE LINER STANDARD | 2
DUROM US ACETABAR COMPONENT 60/54 T | 2
DUROM FEMORAL COMPONENT 52 CODE R | 2
DUROM HIP | 2
TRILOGY ACETABULAR SYSTEM CONSTRAINED LINER | 2
EPSILON ACETABULAR SYSTEM DURASUL CONSTRAINED INSERT | 2
HARRIS PRECOAT PLUS WITH DISTAL PLUG FEMORAL STEM | 2
DUROM FEMORAL COMPONENT 44 CODE J | 2
ZIMMER MIS HIP INSTRUMENTATION HYBRID OFFSET SHELL INSERTER | 2
DUROM US ACETABULAR COMPONENT 52/46 H | 2
DUROM ACETABULAR COMPONENT 62/56 CODE V | 2
TRILOGY LONGEVITY HIGHLY CROSSLINKED POLYETHYLENE CONSTRAINED LINER | 2
VERSYS HIP SYSTEM PROVISIONAL FEMORAL HEAD NECK TAPER STANDARD NECK LENGTH | 1
ZIMMER M/L TAPER HIP PROSTHESIS FEMORAL STEM PRESS-FIT PLASMA-SPRAYED | 1
TRILOGY ACETABULAR SYSTEM LONGEVITY / CROSSLINKED POLYTHYLENE LINER STANDARD | 1
CPT HIP SYSTEM 12/14 TAPER INSTRUMENTATION STEM INSERTER | 1
CONTINUUM TRILOGY IT ALLOFIT IT ACETABULAR SYSTEMS LONGEVITY HIGHLY CROSSLINKED | 1
VERSYS HIP SYSTEM FEMORAL STEM FIBER METAL TAPER COLLARLESS | 1
DUROM US ACETABULAR COMNPONENT 48/42 H | 1
ZIMMER | 1
ZIMMER ANATOMIC HIP PROSTHESIS FEMORAL STEM | 1
TRILOGY ACETABULAR SYSTEM LINER STANDARD LONGEVITY CROSSLINKED POLYETHYLENE | 1
UNK ZMR HIP SYSTEM FEMORAL | 1
EU DUROM ACETABULAR COMPONENT DUROM CUP | 1
TRILOGY ACETABULAR SYSTEM SHELL WITH HOLES POROUS | 1
ZIMMER M/L TAPER HIP PROSTHESIS SYSTEM RASP | 1
HARRIS/GALANTE POROUS TOTAL HIP ACETABULAR COMPONENT | 1
TRILOGY ACETABULAR SYSTEM LINER | 1
CONTINUUM | 1
VERSYS REVISION FEMORAL STEM BEADED | 1
TRILOGY LONGEVITY CONSTRAINED LINER | 1
TRILOGY ACETABULAR SYSTEM LONGEVITY LINER | 1
EPSILON METASUL STANDARD INSERT | 1
TRILOGY ACETABULAR SYSTEM SHELL WITH HOLES | 1
TRILOGY ACETABULAR SYSTEMS LINER 10 DEG ELEV | 1
ZIMMER LONGEVITY LINER STANDARD | 1
TOTAL HIP INSTRUMENTATION STRAIGHT SHELL INSERT | 1
DUROM ACETABULAR COMPONENT 46/40 CODE F | 1
DUROM US ACETABAR COMPONENT 50/44 J | 1
DUROM ACETABULAR COMPONENT 54/48 N | 1
MULTIPOLAR BIPOLAR CUP LINER | 1
DUROM US ACTETABULAR COMPONENT | 1
TRILOGY ACETABULAR SYSTEM LONGEVITY CROSSLINK | 1
DUROM ACETABULUM COMPONENT 56/50 CODE P | 1
ZIMMER DUROM ACETABULUM | 1
ZIMMER M/L TAPER HIP PROSTHESIS WITH KINECTIV | 1
CONTINUUM ACETABULAR SYSTEM TRABECULAR METAL | 1
VERSYS HIP SYSTEM FEMORAL HEAD 1 | 1
ZMR HIP SYSTEM OTJ METAL PROVISIONAL BODY | 1
VERSYS HIP SYSTEM FEMORAL STEM FIBER METAL | 1
TRILOGY ACETABULAR SYSTEM SHELL WITH CLUSTER | 1
DUROM US ACETABULAR COMPONENT 44/38 H | 1
DUROM US ACETABULAR COMPONENT 56/50 H | 1
CONTINUUM TRILOGY IT ALLOFIT IT ACETABULAR SYSTEMS DOME IMPACTOR | 1
PULSAVAC PLUS HIP KIT | 1
CONTINUUM ACETABULAR SYSTEM TRABECULAR METAL SHELL WITH CLUSTER HOLES POROUS | 1
TRABECULAR METAL MODULAR ACETABULAR SYSTEM SHELL WITH MULTIHOLES POROUS | 1
DUROM US ACETABULAR COMPONENT 64/58 X | 1
DUROM CUP | 1
DUROM RESURFACING | 1
VERSYS HIP SYSTEM EPOCH FULLCOAT PROSTHESIS FEMORAL STEM | 1
ZMR HIP SYSTEM FEMORAL BODY REVISION TAPER SPOUT BODY | 1
TRABECULAR METAL ACETABULAR REVISION SYSTEM - CAGE, LEFT SHORT FLANGE | 1
NATURAL-HIP SYSTEM 12/14 TAPER POROUS STEM COLLARLESS | 1
ZIMMER APS NATURAL-HIP SYSTEM PROSTHESIS FEMORAL STEM POROUS | 1
INTER-OP ACETABULAR SYSTEM THREADED STRAIGHT SHELL IMPACTOR ROD | 1
HGP II CUP FLEXIBLE DRILL FOR USE ONLY WITH HGP II CUP | 1
TRILOGY ACETABULAR SYSTEM BONE SCREW SELF-TAPPING | 1
REPLACEMENT OFFSET HEAD SEATER CUP | 1
MODULAR ACETABULAR SYSTEM LINER PROVISIONAL STANDARD FOR USE WITH SHELL PROVISIO | 1
UNK ZIMMER CERAMIC ZIRCONIUM OXIDE FEMORAL HEAD | 1
ZIMMER M/L TAPER HIP PROSTHESIS FEMORAL STEM PRESS-FIT PLASMA SPRAYED | 1
TRILOGY HIP SYSTEM SELF TAPPING BONE SCREW | 1
ZMR HIP SYSTEM FEMORAL STEM REVISION NITRIDED POROUS | 1
TRILOGY ACETABULAR LONGEVITY CROSSLINKED POLYETHYLENE LINER | 1
ZMR HIP SYSTEM FEMORAL BODY REVISION NITRIDED POROUS CONE BODY | 1
VERSYS HIP SYSTEM FEMORAL STEM POROUS FULLCOAT BOWED NECK TAPER REVISION OFFSET | 1
DUROM FEMORAL COMPONENT 48 CODE N | 1
HGP II SHELL DEPTH GAUGE | 1
M/L TAPER HIP PROSTHESIS FEMORAL STEM PRESS-FIT PLASMA SPRAYED | 1
CONTINUUM ACETABULAR SYSTEM TRABECULAR METAL SHELL | 1
DUROM | 1
TRABECULAR METAL MODULAR ACETABULAR SYSTEM SHELL WITH CLUSTER HOLES | 1
TRABECULAR METAL MODULAR ACETABULAR SYSTEM SHELL WITH MULTI-HOLES | 1
VERSYS HIP SYSTEM REVISION FEMORAL STEM BEADED FULLCOAT PLUS | 1
DUROM US ACETABULAR COMPONENT 54/48N | 1
VERSYS HIP SYSTEM FEMORAL HEAD 12/14 TAPER CERAMIC | 1
ZMR HIP SYSTEM FEMORAL STEM REVISION TAPER | 1
DUROM FEMORAL COMPONENT 50 CODE P | 1
DUROM ACETAB. COMP. 58/52 CODE | 1
RESURFACING STYLUS CONNECTOR | 1
VERSYS HIP SYSTEM EPOCH FULLOCAT HIP PROSTHESIS FEMORAL STEM FIBER METAL | 1
UNKNOWN VERSYS HIP SYSTEM STEM | 1
VERSYS HIP SYSTEM ZIRCONIUM OXIDE CERAMIC FEMORAL HEAD | 1
UNKNOWN TRILOGY ACETABULAR SHELL | 1
LONGEVITY HIGHLY CROSSLINKED POLYETHYLENE STANDARD LINER | 1
RESURFACING IMPACTOR HEAD FOR HANDLE | 1
TRABECULAR METAL PRIMARY HIP PROSTHESIS FEMORAL STEM PRESS-FIT COLLARLESS | 1
METASUL LARGE DIAMETER HEAD 52/R | 1
VERSA-FX TIV SELF TAPPING BONE SCREW | 1
CENTRALIGN PRECOAT HIP PROSTHESIS FEMORAL STEM | 1
METASUL DUROM, COMPONENT FOR ACETABULUM, UNCEMENTED, 52/46, CODE L | 1
VERSYS 12/14 FEMORAL HEAD | 1
VERSYS HIP SYSTEM FEMORAL STEM | 1

## Case Metadata
Case Name | MDL Number | MDL Filing Date | Court | Settlement Amount | MAUDE Start Date | MAUDE End Date | Total Records | Date Range
Zimmer Durom Cup Hip Implant | 2158 | 2010-05-12 | D. New Jersey | $228 million (2013) | 2009-11-12 | 2010-05-12 | 357 | 2009-11-12 00:00:00 to 2010-05-12 00:00:00