## All Reports
report_number | date_received | date_of_event | event_type | report_source_code | manufacturer_name | brand_name | generic_name | device_problem | patient_age | patient_sex | event_description | manufacturer_narrative | year | month | year_month
MW5092020 | 2020-01-02 | 2019-10-24 | I, n, j, u, r, y | Voluntary report | SMITH & NEPHEW, INC. | BIRMINGHAM HIP RESURFACING SYSTEM | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL RESURFACING |  | 67 YR |  | REPORTER ALLEGES SHE EXPERIENCED EXCRUCIATING PAIN AND WAS ADVISED BY THE PAIN INSTITUTE TO HAVE HER HIPS CHECKED. BLOOD TESTS FOR METALS WAS DONE ON (B)(6) 2019 AND THE TEST SHOWED HIGH LEVELS OF CARBIDE IN HER BLOOD (ABOUT 3.9). SHE SAW HER PHYSICIAN ON (B)(6) 2019 WHO ADVISED SHE NEEDS A REVISION TO TAKE OUT THE IMPLANTS AND HAVE A TOTAL HIP REPLACEMENT. SHE HAS BEEN GETTING INJECTIONS AND IS SCHEDULED TO EXPLANT RIGHT HIP ON (B)(6) 2020 AND LEFT HIP ON (B)(6) 2020. SHE WAS ALSO INFORMED THE IMPLANTS HAVE DETERIORATED AND IF A REVISION IS NOT DONE, SHE MIGHT NEED TO USE A WHEEL CHAIR. REPORTER IS UPSET SHE HAS TO GO THROUGH THIS AGAIN AND SAID SHE IS AWARE THE DEVICE HAS BEEN RECALLED BUT THE MFR HAS NOT NOTIFIED HER. REPORTER HAS HAD A NAME CHANGE FROM STOKES TO KAMP. |  | 2020 | 1 | 2020-01
MW5092051 | 2020-01-06 | 2017-01-12 | I, n, j, u, r, y | Voluntary report | SMITH AND NEPHEW, INC. | SMITH AND NEPHEW METAL-ON-METAL BIRMINGHAM HIP IMPLANT | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/ METAL, RESURFACING-BHR |  | 63 YR |  | (B)(6). ON (B)(6) 2010, THE PATIENT RECEIVED A LEFT THA WITH A SMITH AND NEPHEW METAL-ON-METAL BIRMINGHAM HIP IMPLANT WITH A SIZE 52 CUP, 46 MM HEAD WITH A +0 NECK LENGTH AND A SIZE 4 STEM. IN (B)(6) OF 2014, PATIENT DEVELOPED HIGH FREQUENCY HEARING LOSS ACCORDING TO AN AUDIOGRAM. IN (B)(6) OF 2014, THE PATIENT¿S ECHOCARDIOGRAM WAS NOTABLE FOR GRADE 1 DIASTOLIC DYSFUNCTION. "ELECTROMYELOGRAM" AND NERVE CONDUCTION VELOCITY TESTS WERE ONLY POSITIVE FOR CARPAL TUNNEL SYNDROME AND EVIDENT POLYNEUROPATHY. ON (B)(6) 2015, HER PLASMA COBALT LEVEL WAS 97 MCG/L AND HER WHOLE BLOOD COBALT LEVEL WAS 97 MCG/L, AND HER SERUM CHROMIUM LEVEL WAS 47 MCG/L. THE PATIENT DEVELOPED AUDIBLE SQUEAKING AND MECHANICAL GRINDING OF THE LEFT HIP WITH REGULAR WALKING. THE GRINDING WAS ACCOMPANIED BY LATERAL THIGH PAIN AND SOME ILIOTIBIAL BAND PAIN ON THE LEFT SIDE, METAL SUPPRESSION MRI OF THE LEFT HIP SHOWED LATERAL CAPSULAR PSEUDOTUMOR AND POSSIBLE EARLY INVOLVEMENT OF THE HIP ABDUCTOR TENDONS AND POSTERIOR CAPSULAR THICKENING. RECHECK OF HER COBALT AND CHROMIUM LEVELS ON (B)(6) 2015 WAS NOTABLE FOR A URINE COBALT LEVEL 1,528.5 MCG/L, WHOLE BLOOD COBALT OF 111.0 MCG/L, URINE CHROMIUM OF 242.8MCG/L AND A WHOLE BLOOD CHROMIUM OF 37.4 MCG/L. THE LEFT HIP WAS REVISED ON (B)(6) 2017. FROZEN SECTION OF LEFT HIP TISSUE NOTABLE FOR FIBROUS STROMA WITH AREAS OF NECROSIS AND FOCAL HISTIOCYTIC INFILTRATES, BUT NEGATIVE FOR INFECTION. THE COBALT LEVEL OF THE LEFT HIP WAS 12,000 MCG/L AND THE CHROMIUM LEVEL OF THE LEFT HIP JOINT FLUID WAS 6,800 MCG/L. THE REVISION SURGEON NOTED MARKED INTRACAPSULAR METALLOSIS. THERE WAS EVIDENCE OF SEVERE CORROSION AT THE TRUNNION OF THE STEM AND HEAD BORE. THERE WAS SEVERE LYSIS ABOUT THE ACETABULAR COMPONENT AND THE COMPONENT WAS FIXED AND NOT LOOSE. ABDUCTION TENDONS AND CAPSULE WERE INTACT. THE REVISION IMPLANT WAS A ZIMMER CONTINUUM SOCKET DIAMETER 56 FIXED WITH 4 SCREWS, VIVACITY E POLYETHYLENE LINER, THE SMITH AND NEPHEW STEM WAS RETAINED AND A NEW SMITH AND NEPHEW OXONIUM HEAD SIZE +MM +12. THE PATIENT EXPERIENCED A POST-SURGICAL COMPLICATION OF DVT AT 13 DAYS POST REVISION SURGERY. ON (B)(6) 2017, CHROMIUM LEFT HIP FLUID, 6,800 MCG/L. FDA SAFETY REPORT ID # (B)(4). |  | 2020 | 1 | 2020-01
MW5092374 | 2020-01-21 | 2020-01-09 | I, n, j, u, r, y | Voluntary report | SMITH & NEPHEW ORTHOPAEDICS LTD. | SMITH & NEPHEW BIRMINGHAM HIP RESURFACING | PROSTHESIS, HIP, SEMI-CONSTRAINED/ METAL/METEL, RESURFACING |  | 189 YR |  | ON (B)(6) 2009, HE RECEIVED A LEFT TOTAL HIP RESURFACING ARTHROPLASTY FOR OSTEOARTHRITIS. THE IMPLANT USED A SMITH & NEPHEW BIRMINGHAM RESURFACING WITH A SIZE 54 CUP AND A SIZE 48 HEAD. ON (B)(6) 2018, HIS WHOLE BLOOD COBALT LEVEL WAS 5.1 MCG/L AND SERUM / PLASMA CHROMIUM LEVEL WAS 5.7 MCG/L. HE COMPLAINED OF REST TREMOR IN THE RIGHT HAND, FATIGUE, SOME MEMORY PROBLEMS, AND SOME BALANCE ISSUES. HIS FDG PET BRAIN SCAN WITH NEURO Q ANALYSIS NOTABLE FOR FOCAL AND GENERALIZED HYPOMETABOLISM CONSISTENT WITH CHRONIC TOXIC ENCEPHALOPATHY. CONFOUNDING FACTOR WAS PRESENCE OF ZIMMER METAL ON METAL DUROM THA ON THE RIGHT SIDE. THIS WAS REVISED ON (B)(6) 2018. FLUID WAS ASPIRATED FROM THE RIGHT HIP DURING REVISION, COBALT LEVEL OF THE RIGHT HIP FLUID WAS 440 MCG/L. AND CHROMIUM LEVEL OF RIGHT HIP FLUID WAS 190 MCG/L. FOLLOWING REVISION OF THE RIGHT HIP, THE ONLY POTENTIAL SOURCE OF COBALT REMAINING IN HIS BODY WAS THE RETAINED LEFT BHR, AND ON (B)(6) 2019, HIS BLOOD COBALT WAS 3.2 MCG/L, AND THE URINE COBALT WAS 17.6 MCG/L. ON (B)(6) 2019 BLOOD COBALT WAS 1.5 MCG/L AND URINE COBALT WAS 13.6 MCG/L. A METAL SUPPRESSION MRI OF THE LEFT BIRMINGHAM HIP, FROM (B)(6) 2018 SHOWED, AT THAT TIME, AN INTRACAPSULAR CYSTIC PSEUDOTUMOR, WHICH WAS ABOUT 6 CM IN DIAMETER. REPEAT METAL SUPPRESSION MRI OF THE LEFT HIP FROM SEPTEMBER OF 2019 SHOWED AGAIN THE CYSTIC PSEUDOTUMOR WITH PROGRESS OF MUSCLE EDEMA / ATROPHY OF THE LEFT ILIACUS MUSCLE. HIS COBALT LEVELS REMAIN PERSISTENTLY ELEVATED SUGGESTED POSSIBLE WEAR OF HIS BHR. ON (B)(6) 2020, HIS WHOLE BLOOD COBALT LEVEL WAS 1.4 MCG/L, AND HIS URINE COBALT LEVEL WAS 14.6 MCG/L. HE IS TAKING A SUPPLEMENTAL DOSE OF 600MG OF N-ACETYLCYSTEINE PER DAY FOR COBALT CHELATION. SINCE THE REVISION OF HIS RIGHT DUROM METAL ON METAL HIP, HE BELIEVES THAT HIS BALANCE ISSUES HAVE RESOLVED. HE ALSO BELIEVES HIS MEMORY HAS IMPROVED. HE STILL HAS A VISIBLE REST TREMOR ALTHOUGH IT IS NOT AS BAD AS IT WAS; 3425_MOM_BHR. FDA SAFETY REPORT ID# (B)(4). |  | 2020 | 1 | 2020-01
MW5092544 | 2020-01-27 | 2019-12-18 | I, n, j, u, r, y | Voluntary report | SMITH & NEPHEW ORTHOPAEDICS LTD. | SMITH & NEPHEW BIRMINGHAM HIP RESURFACING | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 58 YR |  | PT ID: (B)(6); ON (B)(6) 2009, SHE RECEIVED A RIGHT TOTAL HP ARTHROPLASTY FOR OSTEOARTHRITIS AND HIP DYSPLASIA. THE IMPLANT WAS A SMITH & NEPHEW BIRMINGHAM RESURFACING WITH A SIZE 50MM OD UP, 42MM BIRMINGHAM RESURFACING FEMORAL HEAD COMPONENT, 1 PACKAGE OF SIMPLEX P CEMENT WITH 1.2G TOBRAMYCIN. THE RIGHT THA WORKED WELL UNTIL ABOUT (B)(6) OF 2019, WHEN SHE BEGAN EXPERIENCING PROGRESSIVE GROIN AND LATERAL HIP PAIN WHEN SHE WORKS OUT ON A TREADMILL. A METAL SUPPRESSION MRI OF THE RIGHT HIP WAS OBTAINED ON (B)(6) 2019 SHOWED A DEFINITE ADVERSE REACTION TO METAL DEBRIS WITH INCREASED FLUID AT THE HIP AND INSUFFICIENCY OF THE POSTERIOR HIP CAPSULE WITH THE FLUID POOLING POSTERIORLY. HIP ABDUCTOR TENDONS MAY HAVE EARLY INVOLVEMENT AS WELL. PLAIN X-RAYS OF THE HIP AND PELVIS SHOW A WELL-POSITIONED BIRMINGHAM HIP RESURFACING WITH A LITTLE BIT OF FOCAL LYSIS AT THE INGROWTH ACETABULAR BONE INTERFACE AND A LITTLE BIT OF BONE ABSORPTION AROUND THE SUPERIOR FEMORAL NECK. ON (B)(6) 2019, HER URINE COBALT LEVEL WAS 38.2 MCG/L AND HER WHOLE BLOOD COBALT LEVEL WAS 15.7 MCG/L. ON (B)(6) 2019, WHILE STEPPING OUT OF HER CAR, SHE SLIPPED ON THE ICY RUNNING BOARDS OF HER TRUCK, AND LANDED ON HER LEFT SHOULDER AND HER LEFT TEMPLE. CT SCAN MADE OF THE HEAD ON THE SAME DAY OF THE FALL SHOWED NO EVIDENCE OF INTRACRANIAL PROCESS. SHE DID NOT LOSE CONSCIOUSNESS BUT FELT "FUNNY". EVER SINCE THEN SHE HAS HAD AN INTERMITTENT HEADACHE THAT IS HOLOCEPHALIC BUT WORSE IN FOCAL AREAS. SHE NOTICES WHEN SHE IS STRESSED OUT OR IS TRYING TO FOCUS AND IT FEELS LIKE VICE-GRIP TIGHTENING. SHE ALSO DEVELOPED SEVERE PHONOPHOBIA, WITH CERTAIN SOUNDS AGGRAVATING HER HEADACHE. SHE HAS ALSO DEVELOPED POOR MEMORY, IRRITABILITY, AND TROUBLE CONCENTRATING. NO PHOTOPHOBIA AND NO NAUSEA. SHE HAS TRIED BOTOX INJECTIONS OF THE NECK AND HEAD, WHICH WERE NOT HELPFUL. SHE UNDERWENT AN FDG PET BRAIN SCAN DUE TO CONCERN OF POSSIBLE ONGOING SYMPTOMS DUE TO COBALT TOXICITY. FDA SAFETY REPORT ID# (B)(4). |  | 2020 | 1 | 2020-01
3005975929-2020-00080 | 2020-03-24 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A BHR REVISION SURGERY WAS PERFORMED. |  | 2020 | 3 | 2020-03
3005975929-2020-00093 | 2020-03-25 | 2016-05-13 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | BIRMINGHAM HIP RESURFACING (BHR) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Male | IT WAS REPORTED THAT A BHR RIGHT HIP REVISION SURGERY WAS PERFORMED ON (B)(6) 2023. DUE TO PAIN, LIMITED MOBILITY, PSEUDOTUMOR AND ELEVATED METAL ION LEVELS. THE REVISION WAS COMPLETED WITH AN ACTIS SIZE 8 STANDARD OFFSET STEM, 28+1. 5 BIOLOX DELTA ENDOPOSTHETIC HEAD, AND A SMITH AND NEPHEW POLAR HEAD, 28.X 50 INSIDE TO OUTSIDE DIAMETERS. THE PRIMARY RIGHT HIP SURGERY WAS PERFORMED ON (B)(6) 2008, AND ON 13/MAY/2016 THE PATIENT UNDERWENT A RIGHT HIP ASPIRATION THAT INDICATED PINK, CLOUDY SYNOVIAL FLUID AS WELL AS ELEVATED METAL ION LEVELS. THE CURRENT STATE OF HEALTH OF THE PATIENT IS UNKNOWN. |  | 2020 | 3 | 2020-03
3005975929-2020-00130 | 2020-05-11 | 2019-09-09 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED DUE TO METALLOSIS WITH ELEVATED COBALT AND CHROMIUM LEVELS AND PAIN IN HIS RIGHT HIP. |  | 2020 | 5 | 2020-05
3005975929-2020-00131 | 2020-05-11 | 2019-09-09 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) CUP | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED DUE TO METALLOSIS WITH ELEVATED COBALT AND CHROMIUM LEVELS AND PAIN IN HIS RIGHT HIP. |  | 2020 | 5 | 2020-05
3005975929-2020-00135 | 2020-05-18 | 2020-03-31 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED DUE TO SEVERE PAIN, DISCOMFORT, INFLAMMATION AROUND THE HIP IMPLANT, OSTELOYSIS, REACTIVE CAPSULAR AND SYNOVIAL TISSUES. |  | 2020 | 5 | 2020-05
3005975929-2020-00134 | 2020-05-18 | 2019-11-05 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED FOR FAILURE OF THE LEFT HIP REPLACEMENT DUE TO METALLOSIS, ADVERSE TISSUE REACTION. ELEVATED METAL LEVELS AND FLUID. |  | 2020 | 5 | 2020-05
3005975929-2020-00136 | 2020-05-19 | 2019-11-05 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP MODULAR HEAD (BHMH) SLEEVE | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED FOR FAILURE OF THE LEFT HIP REPLACEMENT DUE TO METALLOSIS, ADVERSE TISSUE REACTION. ELEVATED METAL LEVELS AND FLUID. |  | 2020 | 5 | 2020-05
3005975929-2020-00137 | 2020-05-19 | 2019-10-14 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED DUE TO CHRONIC AND SEVERE PERSISTENT LEFT HIP PAIN, TOXIC ELEVATED CHRONIUM AND COBALT LEVELS, METALLOSIS AND ADVERSE LOCAL TISSUE REACTION. |  | 2020 | 5 | 2020-05
3005975929-2020-00140 | 2020-05-20 | 2019-10-14 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, AFTER A HIP REPLACEMENT SURGERY HAD BEEN PERFORMED, THE PATIENT STARTED EXPERIENCING CHRONIC AND SEVERE PERSISTENT PAIN, TOXIC ELEVATED CHRONIUM AND COBALT LEVELS, METALLOSIS AND ADVERSE LOCAL TISSUE REACTION. THE EVENT WAS TREATED BY PERFORMING A REVISION SURGERY. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 5 | 2020-05
3005975929-2020-00158 | 2020-05-27 | 2013-12-31 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD IMP | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT THE PATIENT UNDERWENT MEDICALLY-INDICATED REVISION OF THE BHR RIGHT HIP IMPLANT. THE REVISION SURGERY WAS PERFORMED DUE TO PAIN, LIMITED MOBILITY, METALLOSIS, AND ELEVATED METAL ION LEVELS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 5 | 2020-05
3005975929-2020-00149 | 2020-05-27 | 2017-10-11 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, AFTER A BHR SYSTEM HAD BEEN IMPLANTED ON THE RIGHT HIP, THE PATIENT SUFFERED, PAIN, LIMITED MOBILITY, ELEVATED METAL ION LEVELS, AND INFECTION, WHICH WAS LIKELY LINKED TO ALVAL. THE ADVERSE EVENT WAS TREATED BY PERFORMING A REVISION SURGERY. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 5 | 2020-05
3005975929-2020-00156 | 2020-05-27 | 2020-04-27 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, AFTER A BHR SURGERY HAD BEEN PERFORMED, THE PATIENT STARTED EXPERIENCING PAIN. A REVISION SURGERY WAS PERFORMED TO ADDRESS THIS ADVERSE EVENT. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 5 | 2020-05
3005975929-2020-00154 | 2020-05-27 | 2018-10-26 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, AFTER A LEFT-SIDED BHR SURGERY HAD BEEN PERFORMED, THE PATIENT STARTED EXPERIENCING PAIN, LIMITED MOBILITY, METALLOSIS, AND ELEVATED METAL ION LEVELS. THE ADVERSE EVENT WAS ADDRESSED BY PERFORMING A MEDICALLY-INDICATED REVISION SURGERY TO EXCHANGE THE IMPLANTS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 5 | 2020-05
3005975929-2020-00155 | 2020-05-27 | 2020-04-27 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, AFTER A BHR SURGERY HAD BEEN PERFORMED, THE PATIENT STARTED EXPERIENCING PAIN. A REVISION SURGERY WAS PERFORMED TO ADDRESS THIS ADVERSE EVENT. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 5 | 2020-05
3005975929-2020-00153 | 2020-05-27 | 2018-10-26 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, AFTER A LEFT-SIDED BHR SURGERY HAD BEEN PERFORMED, THE PATIENT STARTED EXPERIENCING PAIN, LIMITED MOBILITY, METALLOSIS, AND ELEVATED METAL ION LEVELS. THE ADVERSE EVENT WAS ADDRESSED BY PERFORMING A MEDICALLY-INDICATED REVISION SURGERY TO EXCHANGE THE IMPLANTS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 5 | 2020-05
3005975929-2020-00157 | 2020-05-27 | 2013-12-31 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT THE PATIENT UNDERWENT MEDICALLY-INDICATED REVISION OF THE BHR RIGHT HIP IMPLANT. THE REVISION SURGERY WAS PERFORMED DUE TO PAIN, LIMITED MOBILITY, METALLOSIS, AND ELEVATED METAL ION LEVELS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 5 | 2020-05
3005975929-2020-00150 | 2020-05-27 | 2018-08-29 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, AFTER A R3-THA SURGERY HAD BEEN PERFORMED, THE PATIENT SUFFERED, PAIN, EXTREMELY ELEVATED BLOOD METAL IONS, METALLOSIS, AND TRUNNIONS. THE EVENT WAS TREATED BY PERFORMING A MEDICALLY-INDICATED REVISION SURGERY. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 5 | 2020-05
3005975929-2020-00148 | 2020-05-27 | 2017-10-11 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, AFTER A BHR SYSTEM HAD BEEN IMPLANTED ON THE RIGHT HIP, THE PATIENT SUFFERED, PAIN, LIMITED MOBILITY, ELEVATED METAL ION LEVELS, AND INFECTION, WHICH WAS LIKELY LINKED TO ALVAL. THE ADVERSE EVENT WAS TREATED BY PERFORMING A REVISION SURGERY. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 5 | 2020-05
3005975929-2020-00165 | 2020-05-29 | 2020-02-23 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT REVISION SURGERY WAS PERFORMED DUE TO METALLOSIS. METAL ON METAL ACETABULAR CUP WAS EXCHANGED FOR A REDAPT MODULAR CUP. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 5 | 2020-05
1020279-2020-01922 | 2020-05-30 | 2015-05-18 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | BIRMINGHAM HIP MODULAR HEAD (BHMH) | PROSTHESIS, HIP, HEMI-, FEMORAL, METAL |  | 70 YR | Female | IT WAS REPORTED THAT, AFTER A RIGHT THA SURGERY HAD BEEN PERFORMED, THE CLINICAL SUBJECT EXPERIENCED AN UNSPECIFIED ADVERSE EVENT THAT WAS RESOLVED BY PERFORMING A REVISION SURGERY. AN UNSPECIFIED ANTHOLOGY SIZE 5 STEM, AN UNSPECIFIED SIZE 46 METAL FEMORAL HEAD, AN UNSPECIFIED SIZE 46/54 CHROME LINER AND AN UNSPECIFIED MIDLAND MEDICAL UNCEMENTED SIZE 54 ACETABULAR SHELL WERE EXPLANTED. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 5 | 2020-05
3005975929-2020-00182 | 2020-06-08 | 2018-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MOD HEAD IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | EU LEGAL CASE: IT WAS REPORTED THAT A BHR REVISION SURGERY WAS PERFORMED DUE TO METALLOSYS FOR (B)(6),REVISION AT (B)(6). 2008. NO FURTHER INFORMATION IS AVAILABLE AT THE MOMENT. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 6 | 2020-06
3005975929-2020-00192 | 2020-06-18 | 2020-05-20 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A BHR 42 FEMORAL HEAD WAS PUT IN 2007, WAS REMOVED DUE TO NECK FRACTURE AND METALLOSIS. NO FURTHER COMPLICATIONS DURING THE PROCEDURE HAVE BEEN REPORTED. |  | 2020 | 6 | 2020-06
3005975929-2020-00201 | 2020-06-19 | 2019-01-08 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | EU LEGAL CASE. IT WAS REPORTED THAT A BHR REVISION SURGERY WAS PERFORMED ON THE RIGHT HIP OF THE PATIENT, DUE TO UNKNOWN SYMPTOMS. NO FURTHER INFORMATION IS AVAILABLE AT THE MOMENT, ADDITIONAL INFORMATION WILL BE REQUESTED. |  | 2020 | 6 | 2020-06
3005975929-2020-00200 | 2020-06-19 | 2017-05-18 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT A BHR REVISION SURGERY WAS PERFORMED ON THE LEFT HIP ON (B)(6) 2017. THE REVISION WAS PERFORMED ON THE PATIENT BECAUSE OF PAIN AND HIGH METAL ION LEVELS. NO FURTHER INFORMATION IS AVAILABLE AT THE MOMENT, ADDITIONAL INFORMATION WILL BE REQUESTED. |  | 2020 | 6 | 2020-06
3005975929-2020-00203 | 2020-06-19 | 2019-05-23 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | BIRMINGHAM HIP RESURFACING (BHR) CUP | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | US LEGAL MDL - IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENT'S RIGHT HIP DUE TO PAIN, LIMITED MOBILITY, AND ELEVATED METAL ION LEVELS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 6 | 2020-06
3005975929-2020-00210 | 2020-06-19 | 2019-01-08 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | EU LEGAL CASE. IT WAS REPORTED THAT A BHR REVISION SURGERY WAS PERFORMED ON THE RIGHT HIP OF THE PATIENT, DUE TO UNKNOWN SYMPTOMS. NO FURTHER INFORMATION IS AVAILABLE AT THE MOMENT, ADDITIONAL INFORMATION WILL BE REQUESTED. |  | 2020 | 6 | 2020-06
3005975929-2020-00209 | 2020-06-19 | 2017-05-18 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | BIRMINGHAM HIP RESURFACING (BHR) CUP | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT A BHR REVISION SURGERY WAS PERFORMED ON THE LEFT HIP ON (B)(6) 2017. THE REVISION WAS PERFORMED ON THE PATIENT BECAUSE OF PAIN AND HIGH METAL ION LEVELS. NO FURTHER INFORMATION IS AVAILABLE AT THE MOMENT, ADDITIONAL INFORMATION WILL BE REQUESTED. |  | 2020 | 6 | 2020-06
3005975929-2020-00206 | 2020-06-19 | 2019-05-23 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | US LEGAL MDL - IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENT'S RIGHT HIP DUE TO PAIN, LIMITED MOBILITY, AND ELEVATED METAL ION LEVELS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 6 | 2020-06
3005975929-2020-00204 | 2020-06-19 | 2019-06-06 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, HEMI-, TRUNNION-BEARING, FEMORAL, METAL/POLYACETAL |  |  |  | US LEGAL. IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENTS RIGHT HIP DUE TO PAIN, ELEVATED METAL ION LEVELS, AND PSEUDOTUMORS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 6 | 2020-06
3005975929-2020-00211 | 2020-06-22 | 2019-06-06 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | BIRMINGHAM HIP MODULAR HEAD (BHMH) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | US LEGAL IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENTS RIGHT HIP DUE TO PAIN, ELEVATED METAL ION LEVELS, AND PSEUDOTUMORS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 6 | 2020-06
3005975929-2020-00217 | 2020-06-30 | 2020-06-03 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 77 YR |  | IT WAS REPORTED THAT, AFTER A RIGHT KNEE TKR SURGERY, THE PATIENT DEVELOPED SEPSIS. A REVISION SURGERY WAS PERFORMED TO TREAT THE EVENT IN WHICH A BHR HEAD WAS REMOVED. NO FURTHER INFORMATION AVAILABLE AT THE MOMENT. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 6 | 2020-06
3005975929-2020-00228 | 2020-07-03 | 2019-05-23 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, AFTER A BHR CONSTRUCT HAD BEEN IMPLANTED IN A RIGHT HIP ARTHROPLASTY, THE PATIENT PRESENTED CHRONIC AND SEVERE PERSISTENT PAIN, TOXIC ELEVATED CHROMIUM AND COBALT LEVELS, METALLOSIS, AND ADVERSE LOCAL TISSUE REACTION. A REVISION SURGERY WAS PERFORMED TO TREAT THE ADVERSE EVENT. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 7 | 2020-07
3005975929-2020-00230 | 2020-07-03 | 1901-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED A RIGHT HIP REVISION SURGERY WAS PERFORMED DUE TO PERIPROSTHETIC INFECTION. IT IS UNKNOWN THE OCCURRENCE DATE OF THE EVENT. NO FURTHER INFORMATION IS AVAILABLE AT THE MOMENT. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 7 | 2020-07
3005975929-2020-00231 | 2020-07-03 | 1901-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED A RIGHT HIP REVISION SURGERY WAS PERFORMED DUE TO PERIPROSTHETIC INFECTION. IT IS UNKNOWN THE OCCURRENCE DATE OF THE EVENT. NO FURTHER INFORMATION IS AVAILABLE AT THE MOMENT. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 7 | 2020-07
3005975929-2020-00235 | 2020-07-10 | 2017-11-16 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | KNEE, PTLLFMRTBIAL, SM-CNSTRND, CMNTD, POLYMER/METAL/POLYMER |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENTS LEFT HIP DUE TO PAIN, LIMITED MOBILITY, ELEVATED METAL ION LEVELS, SCIATIC NERVE DAMAGE, AND INFECTION WHICH WAS LIKELY LINKED TO ALVAL. |  | 2020 | 7 | 2020-07
3005975929-2020-00236 | 2020-07-10 | 2017-11-16 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) HEAD | KNEE, PTLLFMRTBIAL, SM-CNSTRND, CMNTD, POLYMER/METAL/POLYMER |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENTS LEFT HIP DUE TO PAIN, LIMITED MOBILITY, ELEVATED METAL ION LEVELS, SCIATIC NERVE DAMAGE, AND INFECTION WHICH WAS LIKELY LINKED TO ALVAL. |  | 2020 | 7 | 2020-07
3005975929-2020-00249 | 2020-07-14 | 2020-05-15 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, POROUS UNCEMENTED |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENTS RIGHT HIP DUE TO PAIN AND ELEVATED CHROMIUM AND COBALT LEVELS FROM METAL IONS FROM A FAILED RIGHT HIP RESURFACING ARTHROPLASTY CONSISTENT WITH METAL-ON-METAL WEAR. AMONG THE SURGICAL FINDINGS INCLUDED SOME STAINING TISSUE, AND THAT THE FEMORAL PORTION OF THE RESURFACING WAS GROSSLY LOOSE WITH THE UNDERLYING BONE BEING NECROTIC IN APPEARANCE, CONFIRMING HIS PRE-OPERATIVE DIAGNOSIS OF FAILED RIGHT BIRMINGHAM RESURFACING SECONDARY TO ASEPTIC LOOSENING AND INCREASED METAL IONS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 7 | 2020-07
3005975929-2020-00250 | 2020-07-14 | 2020-05-15 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, POROUS UNCEMENTED |  |  |  | US LEGAL MDL. IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENTS RIGHT HIP DUE TO PAIN AND ELEVATED CHROMIUM AND COBALT LEVELS FROM METAL IONS FROM A FAILED RIGHT HIP RESURFACING ARTHROPLASTY CONSISTENT WITH METAL-ON-METAL WEAR. AMONG THE SURGICAL FINDINGS INCLUDED SOME STAINING TISSUE, AND THAT THE FEMORAL PORTION OF THE RESURFACING WAS GROSSLY LOOSE WITH THE UNDERLYING BONE BEING NECROTIC IN APPEARANCE, CONFIRMING HIS PRE-OPERATIVE DIAGNOSIS OF FAILED RIGHT BIRMINGHAM RESURFACING SECONDARY TO ASEPTIC LOOSENING AND INCREASED METAL IONS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 7 | 2020-07
3005975929-2020-00251 | 2020-07-15 | 2018-07-16 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS,HIP,SEMI-CONSTRAINED(METAL CEMENTED ACETABULAR COMPONENT) |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENTS RIGHT HIP DUE TO PAIN, LIMITED MOBILITY AND ELEVATED METAL ION LEVELS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 7 | 2020-07
3005975929-2020-00254 | 2020-07-15 | 2019-10-21 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP MODULAR HEAD (BHMH) IMPL | PROSTHESIS, HIP, HEMI-, TRUNNION-BEARING, FEMORAL, METAL/POLYACETAL |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENTS LEFT HIP DUE TO SEVERE PAIN, LIMITED MOBILITY, ELEVATED COBALT AND CHROMIUM LEVELS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 7 | 2020-07
3005975929-2020-00253 | 2020-07-15 | 2019-10-21 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED (METAL CEMENTED ACETABULAR COMPONENT) |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENTS LEFT HIP DUE TO SEVERE PAIN, LIMITED MOBILITY, ELEVATED COBALT AND CHROMIUM LEVELS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 7 | 2020-07
3005975929-2020-00252 | 2020-07-15 | 2018-07-16 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, HEMI-, TRUNNION-BEARING, FEMORAL, METAL/POLYACETAL |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENTS RIGHT HIP DUE TO PAIN, LIMITED MOBILITY AND ELEVATED METAL ION LEVELS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 7 | 2020-07
3005975929-2020-00279 | 2020-08-03 | 2020-06-30 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | BIRMINGHAM HIP RESURFACING (BHR) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 73 YR | Male | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENTS LEFT HIP DUE TO PAIN, DISABILITY, AND MENTAL SUFFERING. |  | 2020 | 8 | 2020-08
3005975929-2020-00281 | 2020-08-07 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A PATIENT UNDERWENT A BHR REVISION SURGERY. IT IS UNKNOWN WHEN DID THE EVENT HAPPENED NOR HOW MANY DEVICES WERE EXPLANTED. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 8 | 2020-08
3005975929-2020-00280 | 2020-08-07 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A PATIENT UNDERWENT A BHR REVISION SURGERY. IT IS UNKNOWN WHEN DID THE EVENT HAPPENED NOR HOW MANY DEVICES WERE EXPLANTED. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 8 | 2020-08
3005975929-2020-00287 | 2020-08-09 | 1901-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENTS HIP DUE TO UNKNOWN REASONS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 8 | 2020-08
3005975929-2020-00288 | 2020-08-09 | 1901-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENTS HIP DUE TO UNKNOWN REASONS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 8 | 2020-08
3005975929-2020-00300 | 2020-08-18 | 1901-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT THE PATIENT REQUESTED A REVISION SURGERY DUE TO HIGH LEVEL OF METALLOSIS. IT IS UNKNOWN IF THE REVISION SURGERY HAS TAKEN PLACE YET OR NOT. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 8 | 2020-08
3005975929-2020-00301 | 2020-08-18 | 1901-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT THE PATIENT REQUESTED A REVISION SURGERY DUE TO HIGH LEVEL OF METALLOSIS. IT IS UNKNOWN IF THE REVISION SURGERY HAS TAKEN PLACE YET OR NOT. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 8 | 2020-08
3005975929-2020-00319 | 2020-08-21 | 2019-08-20 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, AFTER A BHR CONSTRUCT HAD BEEN IMPLANTED IN A LEFT HIP ON (B)(6) 2008, THE PATIENT EXPERIENCED, AS CONSEQUENCE OF A FAILED LEFT TOTAL HIP REPLACEMENT, PAIN AND POOR FUNCTION; COBALT TOXICITY; PSEUDOTUMOR WITH INVOLVEMENT OF DEEP CAPSULAR STRUCTURES AND EXTERNAL ROTATORS; ATROPHY AND NECROSIS OF TISSUE AND FLUID IN HIP JOINT. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2019 TO TREAT THE ADVERSE EVENT. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 8 | 2020-08
3005975929-2020-00306 | 2020-08-21 | 2020-07-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, 11 YEARS AFTER A BHR THA CONSTRUCT HAD BEEN IMPLANTED, THE PATIENT DEVELOPED UNBEARABLE PAIN. UPON GETTING AN MRI IT WAS NOTICED THAT ABOUT 1.5 LITERS OF PUS ACCUMULATED NEAR THE HIP IMPLANT. BY THE TIME THE PATIENT WAS ADMITTED, IT WAS A CONFIRMED CASE OF SEPTICEMIA, A LIFE-THREATENING DISORDER, AND THE PATIENT WAS OPERATED. UPON FURTHER DIAGNOSIS, IT WAS FOUND THAT DUE TO METAL IONS FORMED NEAR THE IMPLANT, THERE WAS A BIG CLOT THAT CAUSED THE INFECTION, WHICH LED TO SEPTICEMIA. IT IS UNKNOWN IF THE DEVICES WERE REVISED. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 8 | 2020-08
3005975929-2020-00302 | 2020-08-21 | 2019-10-16 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) CUP | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, AFTER A BHR THA CONSTRUCT HAD BEEN IMPLANTED ON THE RIGHT HIP, THE PATIENT EXPERIENCED PERSISTENT PAIN, SYNOVITIS AND METALLOSIS. A REVISION SURGERY WAS PERFORMED TO TREAT THE ADVERSE EVENT. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 8 | 2020-08
3005975929-2020-00303 | 2020-08-21 | 2019-10-16 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, AFTER A BHR THA CONSTRUCT HAD BEEN IMPLANTED ON THE RIGHT HIP, THE PATIENT EXPERIENCED PERSISTENT PAIN, SYNOVITIS AND METALLOSIS. A REVISION SURGERY WAS PERFORMED TO TREAT THE ADVERSE EVENT. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 8 | 2020-08
3005975929-2020-00305 | 2020-08-21 | 2019-11-20 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, AFTER A BHR-THA CONSTRUCT HAD BEEN IMPLANTED IN A THA SURGERY, THE PATIENT EXPERIENCED SEVERE PAIN, LIMITED MOBILITY, EXTENSIVE METALLOSIS, AND PSEUDOTUMOR. A REVISION SURGERY WAS PERFORMED ON THE PATIENT AS A RESULT OF THE PREMATURE FAILURE OF THE DEVICES. THE FEMORAL HEAD AND CUP WERE REPLACED. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 8 | 2020-08
3005975929-2020-00304 | 2020-08-21 | 2019-11-20 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, AFTER A BHR-THA CONSTRUCT HAD BEEN IMPLANTED IN A THA SURGERY, THE PATIENT EXPERIENCED SEVERE PAIN, LIMITED MOBILITY, EXTENSIVE METALLOSIS, AND PSEUDOTUMOR. A REVISION SURGERY WAS PERFORMED ON THE PATIENT AS A RESULT OF THE PREMATURE FAILURE OF THE DEVICES. THE FEMORAL HEAD AND CUP WERE REPLACED. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 8 | 2020-08
3005975929-2020-00307 | 2020-08-21 | 2020-07-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP MODULAR HEAD (BHMH) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, 11 YEARS AFTER A BHR THA CONSTRUCT HAD BEEN IMPLANTED, THE PATIENT DEVELOPED UNBEARABLE PAIN. UPON GETTING AN MRI IT WAS NOTICED THAT ABOUT 1.5 LITERS OF PUS ACCUMULATED NEAR THE HIP IMPLANT. BY THE TIME THE PATIENT WAS ADMITTED, IT WAS A CONFIRMED CASE OF SEPTICEMIA, A LIFE-THREATENING DISORDER, AND THE PATIENT WAS OPERATED. UPON FURTHER DIAGNOSIS, IT WAS FOUND THAT DUE TO METAL IONS FORMED NEAR THE IMPLANT, THERE WAS A BIG CLOT THAT CAUSED THE INFECTION, WHICH LED TO SEPTICEMIA. IT IS UNKNOWN IF THE DEVICES WERE REVISED. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 8 | 2020-08
3005975929-2020-00318 | 2020-08-21 | 2019-08-20 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, AFTER A BHR CONSTRUCT HAD BEEN IMPLANTED IN A LEFT HIP ON (B)(6) 2008, THE PATIENT EXPERIENCED, AS CONSEQUENCE OF A FAILED LEFT TOTAL HIP REPLACEMENT, PAIN AND POOR FUNCTION; COBALT TOXICITY; PSEUDOTUMOR WITH INVOLVEMENT OF DEEP CAPSULAR STRUCTURES AND EXTERNAL ROTATORS; ATROPHY AND NECROSIS OF TISSUE AND FLUID IN HIP JOINT. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2019 TO TREAT THE ADVERSE EVENT. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 8 | 2020-08
3005975929-2020-00326 | 2020-08-24 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, AFTER A THA IN WHICH A BHR CONSTRUCT HAD BEEN IMPLANTED ON THE LEFT HIP, THE IMPLANTS FAILED IN AN UNKNOWN WAY. THE PATIENT UNDERWENT A UNKNOWN REVISION SURGERY 10 YEARS AGO TO TREAT THE ADVERSE EVENT; THE PATIENT RECEIVED A METAL-ON-METAL SYSTEM INSTEAD. |  | 2020 | 8 | 2020-08
3005975929-2020-00322 | 2020-08-24 | 2017-01-12 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, SIX YEARS AFTER A REVISION SURGERY IN WHICH A BHR SYSTEM HAD BEEN REPLACED WITH A METAL-ON-METAL SYSTEM ON BOTH HIPS, THE PATIENT BECAME SEVERLY DISABLED AND WAS DIAGNOSED WITH METALLOSIS AND ARTHROPLASTIC COBALT ENCEPHALOPATHY. IT IS UNKNOWN IF/HOW THE PATIENT HAS BEEN TREATED. |  | 2020 | 8 | 2020-08
3005975929-2020-00321 | 2020-08-24 | 2017-01-12 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, SIX YEARS AFTER A REVISION SURGERY IN WHICH A BHR SYSTEM HAD BEEN REPLACED WITH A METAL-ON-METAL SYSTEM ON BOTH HIPS, THE PATIENT BECAME SEVERLY DISABLED AND WAS DIAGNOSED WITH METALLOSIS AND ARTHROPLASTIC COBALT ENCEPHALOPATHY. IT IS UNKNOWN IF/HOW THE PATIENT HAS BEEN TREATED. |  | 2020 | 8 | 2020-08
3005975929-2020-00320 | 2020-08-24 | 2017-01-12 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, SIX YEARS AFTER A REVISION SURGERY IN WHICH A BHR SYSTEM HAD BEEN REPLACED WITH A METAL-ON-METAL SYSTEM ON BOTH HIPS, THE PATIENT BECAME SEVERLY DISABLED AND WAS DIAGNOSED WITH METALLOSIS AND ARTHROPLASTIC COBALT ENCEPHALOPATHY. IT IS UNKNOWN IF/HOW THE PATIENT HAS BEEN TREATED. |  | 2020 | 8 | 2020-08
3005975929-2020-00323 | 2020-08-24 | 2017-01-12 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, SIX YEARS AFTER A REVISION SURGERY IN WHICH A BHR SYSTEM HAD BEEN REPLACED WITH A METAL-ON-METAL SYSTEM ON BOTH HIPS, THE PATIENT BECAME SEVERLY DISABLED AND WAS DIAGNOSED WITH METALLOSIS AND ARTHROPLASTIC COBALT ENCEPHALOPATHY. IT IS UNKNOWN IF/HOW THE PATIENT HAS BEEN TREATED. |  | 2020 | 8 | 2020-08
3005975929-2020-00324 | 2020-08-24 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, AFTER A THA IN WHICH A BHR CONSTRUCT HAD BEEN IMPLANTED ON THE RIGHT HIP, THE IMPLANTS FAILED IN AN UNKNOWN WAY. THE PATIENT UNDERWENT A UNKNOWN REVISION SURGERY 10 YEARS AGO TO TREAT THE ADVERSE EVENT; THE PATIENT RECEIVED A METAL-ON-METAL SYSTEM INSTEAD. |  | 2020 | 8 | 2020-08
3005975929-2020-00327 | 2020-08-24 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, AFTER A THA IN WHICH A BHR CONSTRUCT HAD BEEN IMPLANTED ON THE LEFT HIP, THE IMPLANTS FAILED IN AN UNKNOWN WAY. THE PATIENT UNDERWENT A UNKNOWN REVISION SURGERY 10 YEARS AGO TO TREAT THE ADVERSE EVENT; THE PATIENT RECEIVED A METAL-ON-METAL SYSTEM INSTEAD. |  | 2020 | 8 | 2020-08
3005975929-2020-00325 | 2020-08-24 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, AFTER A THA IN WHICH A BHR CONSTRUCT HAD BEEN IMPLANTED ON THE RIGHT HIP, THE IMPLANTS FAILED IN AN UNKNOWN WAY. THE PATIENT UNDERWENT A UNKNOWN REVISION SURGERY 10 YEARS AGO TO TREAT THE ADVERSE EVENT; THE PATIENT RECEIVED A METAL-ON-METAL SYSTEM INSTEAD. |  | 2020 | 8 | 2020-08
3005975929-2020-00334 | 2020-09-12 | 2019-09-05 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, AFTER A BHR CONSTRUCT HAD BEEN IMPLANTED IN THE RIGHT HIP ON (B)(6)2009, THE PATIENT EXPERIENCED CHRONIC, SEVERE, AND PERSISTENT PAIN, TOXIC ELEVATED CHROMIUM AND COBALT LEVELS, METALLOSIS, AND ADVERSE LOCAL TISSUE REACTION. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2019 TO TREAT THE ADVERSE EVENT. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 9 | 2020-09
3005975929-2020-00335 | 2020-09-12 | 2019-09-05 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, AFTER A BHR CONSTRUCT HAD BEEN IMPLANTED IN THE RIGHT HIP ON (B)(6)2009, THE PATIENT EXPERIENCED CHRONIC, SEVERE, AND PERSISTENT PAIN, TOXIC ELEVATED CHROMIUM AND COBALT LEVELS, METALLOSIS, AND ADVERSE LOCAL TISSUE REACTION. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2019 TO TREAT THE ADVERSE EVENT. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 9 | 2020-09
3005975929-2020-00338 | 2020-09-13 | 2011-10-13 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, FOUR YEARS AFTER A BHR CONSTRUCT HAD BEEN IMPLANTED, THE PATIENT EXPERIENCED AN UNSPECIFIED ADVERSE EVENT THAT LED TO REVISION SURGERY: THE BHR CONSTRUCT WAS REPLACED FOR A BHR-THA SYSTEM. |  | 2020 | 9 | 2020-09
3005975929-2020-00339 | 2020-09-13 | 2011-10-13 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, FOUR YEARS AFTER A BHR CONSTRUCT HAD BEEN IMPLANTED, THE PATIENT EXPERIENCED AN UNSPECIFIED ADVERSE EVENT THAT LED TO REVISION SURGERY: THE BHR CONSTRUCT WAS REPLACED FOR A BHR-THA SYSTEM. |  | 2020 | 9 | 2020-09
3005975929-2020-00353 | 2020-09-24 | 2020-08-28 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON (B)(6) 2020. THE REVISION SURGERY WAS PERFORMED DUE TO UNKNOWN REASONS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 9 | 2020-09
3005975929-2020-00361 | 2020-09-28 | 2019-05-13 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT THE PATIENT UNDERWENT MEDICALLY-INDICATED REVISION OF THE BHR RIGHT HIP IMPLANTS ON (B)(6) 2019. THE PATIENT REVISION SURGERY WAS PERFORMED DUE PROGRESSIVE HIP PAIN AND ELEVATED COBALT AND CHROMIUM BLOOD LEVELS IN HER BODY. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 9 | 2020-09
3005975929-2020-00365 | 2020-09-28 | 2019-11-13 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON (B)(6) 2019 ON THE PATIENT RIGHT HIP DUE TO METAL BEARING WEAR-CAUSING PARTICULATE REACTIVE SYNOVITIS RIGHT HIP; STATUS POST UNCEMENTED RIGHT TOTAL HIP ARTHROPLASTY WITH METAL-ON METAL BEARING; AND CHRONIC RIGHT HIP PAIN. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 9 | 2020-09
3005975929-2020-00364 | 2020-09-28 | 2019-05-13 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT THE PATIENT UNDERWENT MEDICALLY-INDICATED REVISION OF THE BHR RIGHT HIP IMPLANTS ON (B)(6) 2019. THE PATIENT REVISION SURGERY WAS PERFORMED DUE PROGRESSIVE HIP PAIN AND ELEVATED COBALT AND CHROMIUM BLOOD LEVELS IN HER BODY. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 9 | 2020-09
3005975929-2020-00363 | 2020-09-28 | 2019-11-13 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON (B)(6) 2019 ON THE PATIENT RIGHT HIP DUE TO METAL BEARING WEAR-CAUSING PARTICULATE REACTIVE SYNOVITIS RIGHT HIP; STATUS POST UNCEMENTED RIGHT TOTAL HIP ARTHROPLASTY WITH METAL-ON METAL BEARING; AND CHRONIC RIGHT HIP PAIN. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 9 | 2020-09
3005975929-2020-00370 | 2020-10-04 | 2020-07-22 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Male | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENT RIGHT HIP ON (B)(6) 2020 DUE TO UNEXPLAINED PAIN, AND ELEVATED COBALT AND CHROMIUM LEVELS. AMONG THE INTRA-OPERATIVE FINDINGS THERE WERE METALLOSIS, REACTIVE SYNOVIAL TISSUE WITH FIBRINOUS DEGENERATIVE AND EXTENSIVE MACROPHAGE REACTION TO PIGMENTED METALLIC DEBRIS, CONSISTENT WITH A FAILED BIRMINGHAM RIGHT HIP RESURFACING ARTHROPLASTY RESULTING FROM METAL-ON-METAL WEAR AND ADVERSE METAL REACTION. ONLY THE FEMORAL HEAD WAS EXPLANTED. THE CUP REMAINS WITHIN THE PATIENT. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 10 | 2020-10
3005975929-2020-00369 | 2020-10-04 | 2020-07-22 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Male | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENT RIGHT HIP ON (B)(6) 2020 DUE TO UNEXPLAINED PAIN, AND ELEVATED COBALT AND CHROMIUM LEVELS. AMONG THE INTRA-OPERATIVE FINDINGS THERE WERE METALLOSIS, REACTIVE SYNOVIAL TISSUE WITH FIBRINOUS DEGENERATIVE AND EXTENSIVE MACROPHAGE REACTION TO PIGMENTED METALLIC DEBRIS, CONSISTENT WITH A FAILED BIRMINGHAM RIGHT HIP RESURFACING ARTHROPLASTY RESULTING FROM METAL-ON-METAL WEAR AND ADVERSE METAL REACTION. ONLY THE FEMORAL HEAD WAS EXPLANTED. THE CUP REMAINS WITHIN THE PATIENT. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 10 | 2020-10
3005975929-2020-00381 | 2020-10-13 | 2018-10-21 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON (B)(6) 2018 ON THE PATIENT RIGHT HIP DUE TO FAILURE OF RIGHT THA WITH SYNOVITIS, PAIN IN THE HIP AND GROIN REGION, TISSUE AND NERVE DAMAGE, AND METALLOSIS. ONLY THE R3 METAL LINER, HEAD AND SLEEVE WERE EXPLANTED. A S&N 50X32MM +0 POLYETHELENE R3 XPLE LINER AND A S&N OXINIUM 32M +8 FEMORAL HEAD WERE IMPLANTED. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 10 | 2020-10
3005975929-2020-00395 | 2020-10-22 | 2020-10-05 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT THE PATIENT UNDERWENT A REVISION OF THE BHR THA HIP IMPLANT. PATIENT WAS REVISED TO BHR DUAL MOBILITY HEAD. THE PATIENT HAD ELEVATED LAB MARKERS. NO FURTHER INFORMATION IS AVAILABLE AT THE MOMENT. |  | 2020 | 10 | 2020-10
3005975929-2020-00394 | 2020-10-22 | 2020-10-05 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT THE PATIENT UNDERWENT A REVISION OF THE BHR THA HIP IMPLANT. PATIENT WAS REVISED TO BHR DUAL MOBILITY HEAD. THE PATIENT HAD ELEVATED LAB MARKERS. NO FURTHER INFORMATION IS AVAILABLE AT THE MOMENT. |  | 2020 | 10 | 2020-10
3005975929-2020-00404 | 2020-10-26 | 2020-09-17 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP BRIDGING | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT THE PATIENT UNDERWENT TWO REVISIONS DUE TO PAIN AND ELEVATED COBALT AND CHROMIUM LEVELS. PATIENT HAD A BILATERAL PRIMARY SURGERY ON DIFFERENT DATES AND STARTED TO PRESENT FEELINGS OF PAIN. UPON CHECKING IT WAS DISCOVERED THAT THE PATIENT HAD ABOVE NORMAL LEVELS OF COBALT AND CHROMIUM LEVELS ON HIS BLOOD, THEREFORE IT WAS DECIDED TO REVISE THE RIGHT HIP AND THEN THE LEFT HIP. |  | 2020 | 10 | 2020-10
3005975929-2020-00396 | 2020-10-26 | 2020-10-08 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, AFTER A BHR CONSTRUCT FROM SMITH & NEPHEW HAD BEEN IMPLANTED ON THE RIGHT HIP, THE PATIENT HAD HIGH METAL ION LEVELS. A REVISION SURGERY WAS SCHEDULED TO REMOVE THE IMPLANTS, BUT IT WAS CANCELLED FOR THE PATIENT DID NOT WANT FURTHER SMITH & NEPHEW PRODUCTS IMPLANTED. IT IS STILL UNKNOWN IF THE PATIENT PROBLEM HAS BEEN ADDRESSED YET. |  | 2020 | 10 | 2020-10
3005975929-2020-00399 | 2020-10-26 | 2020-10-08 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, AFTER A BHR CONSTRUCT FROM SMITH & NEPHEW HAD BEEN IMPLANTED ON THE LEFT HIP, THE PATIENT HAD HIGH METAL ION LEVELS. A REVISION SURGERY WAS SCHEDULED TO REMOVE THE IMPLANTS, BUT IT WAS CANCELLED FOR THE PATIENT DID NOT WANT FURTHER SMITH & NEPHEW PRODUCTS IMPLANTED. IT IS STILL UNKNOWN IF THE PATIENT PROBLEM HAS BEEN ADDRESSED YET. |  | 2020 | 10 | 2020-10
3005975929-2020-00401 | 2020-10-26 | 2020-06-18 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT THE PATIENT UNDERWENT TWO REVISIONS DUE TO PAIN AND ELEVATED COBALT AND CHROMIUM LEVELS. PATIENT HAD A BILATERAL PRIMARY SURGERY ON DIFFERENT DATES AND STARTED TO PRESENT FEELINGS OF PAIN. UPON CHECKING IT WAS DISCOVERED THAT THE PATIENT HAD ABOVE NORMAL LEVELS OF COBALT AND CHROMIUM LEVELS ON HIS BLOOD, THEREFORE IT WAS DECIDED TO REVISE THE RIGHT HIP AND THEN THE LEFT HIP. |  | 2020 | 10 | 2020-10
3005975929-2020-00403 | 2020-10-26 | 2020-09-17 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT THE PATIENT UNDERWENT TWO REVISIONS DUE TO PAIN AND ELEVATED COBALT AND CHROMIUM LEVELS. PATIENT HAD A BILATERAL PRIMARY SURGERY ON DIFFERENT DATES AND STARTED TO PRESENT FEELINGS OF PAIN. UPON CHECKING IT WAS DISCOVERED THAT THE PATIENT HAD ABOVE NORMAL LEVELS OF COBALT AND CHROMIUM LEVELS ON HIS BLOOD, THEREFORE IT WAS DECIDED TO REVISE THE RIGHT HIP AND THEN THE LEFT HIP. |  | 2020 | 10 | 2020-10
3005975929-2020-00402 | 2020-10-26 | 2020-06-18 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP BRIDGING | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT THE PATIENT UNDERWENT TWO REVISIONS DUE TO PAIN AND ELEVATED COBALT AND CHROMIUM LEVELS. PATIENT HAD A BILATERAL PRIMARY SURGERY ON DIFFERENT DATES AND STARTED TO PRESENT FEELINGS OF PAIN. UPON CHECKING IT WAS DISCOVERED THAT THE PATIENT HAD ABOVE NORMAL LEVELS OF COBALT AND CHROMIUM LEVELS ON HIS BLOOD, THEREFORE IT WAS DECIDED TO REVISE THE RIGHT HIP AND THEN THE LEFT HIP. |  | 2020 | 10 | 2020-10
3005975929-2020-00398 | 2020-10-26 | 2020-10-08 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, AFTER A BHR CONSTRUCT FROM SMITH & NEPHEW HAD BEEN IMPLANTED ON THE LEFT HIP, THE PATIENT HAD HIGH METAL ION LEVELS. A REVISION SURGERY WAS SCHEDULED TO REMOVE THE IMPLANTS, BUT IT WAS CANCELLED FOR THE PATIENT DID NOT WANT FURTHER SMITH & NEPHEW PRODUCTS IMPLANTED. IT IS STILL UNKNOWN IF THE PATIENT PROBLEM HAS BEEN ADDRESSED YET. |  | 2020 | 10 | 2020-10
3005975929-2020-00397 | 2020-10-26 | 2020-10-08 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, AFTER A BHR CONSTRUCT FROM SMITH & NEPHEW HAD BEEN IMPLANTED ON THE RIGHT HIP, THE PATIENT HAD HIGH METAL ION LEVELS. A REVISION SURGERY WAS SCHEDULED TO REMOVE THE IMPLANTS, BUT IT WAS CANCELLED FOR THE PATIENT DID NOT WANT FURTHER SMITH & NEPHEW PRODUCTS IMPLANTED. IT IS STILL UNKNOWN IF THE PATIENT PROBLEM HAS BEEN ADDRESSED YET. |  | 2020 | 10 | 2020-10
3005975929-2020-00406 | 2020-10-30 | 2020-06-15 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT PATIENT SUFFERED FROM A ELEVATED ION LEVELS. THE PATIENT IS A BILATERAL, THIS CASE IS FOR THE RIGHT HIP IMPLANTED ON (B)(6) 2006. THE PATIENT HAS NOT UNDERGONE A MEDICALLY-INDICATED REVISION SO FAR, BUT THE CONTRIBUTION OF THE RIGHT HIP IMPLANTS TO THE ELEVATED ION LEVELS COULD NOT BE DISCARDED. NO FURTHER INFORMATION IS AVAILABLE AT THE MOMENT. |  | 2020 | 10 | 2020-10
3005975929-2020-00405 | 2020-10-30 | 2020-06-15 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | ASSAY, NUCLEIC ACID AMPLIFICATION, BACILLUS ANTHRACIS |  |  |  | IT WAS REPORTED THAT PATIENT SUFFERED FROM A ELEVATED ION LEVELS. THE PATIENT IS A BILATERAL , THIS CASE IS FOR THE RIGHT HIP IMPLANTED ON (B)(6) 2006. THE PATIENT HASN'T UNDERGONE A MEDICALLY-INDICATED REVISION SO FAR, BUT THE CONTRIBUTION OF THE RIGHT HIP IMPLANTS TO THE ELEVATED ION LEVELS COULD NOT BE DISCARDED. NO FURTHER INFORMATION IS AVAILABLE AT THE MOMENT. |  | 2020 | 10 | 2020-10
3005975929-2020-00408 | 2020-11-01 | 2019-07-11 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENTS LEFT HIP ON (B)(6) 2019 DUE TO PAIN, METALLOSIS, ELEVATED METAL LEVELS, PSEUDOTUMOR, AND FAILURE OF LEFT TOTAL HIP REPLACEMENT. AMONG THE INTRA-OPERATIVE FINDINGS AND/OR DIAGNOSES WERE FOUND THE FAILURE OF LEFT HIP REPLACEMENT DUE TO METALLOSIS AND LOCAL ADVERSE TISSUE REACTION, ELEVATED METAL LEVELS, AND FLUID. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 11 | 2020-11
3005975929-2020-00407 | 2020-11-01 | 2019-07-11 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENTS LEFT HIP ON (B)(6) 2019 DUE TO PAIN, METALLOSIS, ELEVATED METAL LEVELS, PSEUDOTUMOR, AND FAILURE OF LEFT TOTAL HIP REPLACEMENT. AMONG THE INTRA-OPERATIVE FINDINGS AND/OR DIAGNOSES WERE FOUND THE FAILURE OF LEFT HIP REPLACEMENT DUE TO METALLOSIS AND LOCAL ADVERSE TISSUE REACTION, ELEVATED METAL LEVELS, AND FLUID. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 11 | 2020-11
3005975929-2020-00414 | 2020-11-03 | 2020-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A BHR-THA REVISION SURGERY WAS PERFORMED. THE REASON FOR WHY THE REVISION SURGERY WAS PERFORMED IS UNKNOWN. THE INFORMATION ABOUT THE SURGEON, HOSPITAL AND DATES ARE UNKNOWN. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 11 | 2020-11
3005975929-2020-00420 | 2020-11-10 | 2019-11-04 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, HEMI-, TRUNNION-BEARING, FEMORAL |  |  | Male | US LEGAL MDL. IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENT RIGHT HIP DUE TO PERSISTENT HIP PAIN AND LIMITATIONS OF DAY TO DAY LIVING, ELEVATED CHROMIUM AND COBALT LEVELS AND CONCERN FOR METALLOSIS. AMONG THE INTRA-OPERATIVE FINDINGS WAS PRESENCE OF SYNOVITIC-TYPE DARKISH STAINED TISSUE CONSISTENT WITH METAL ON METAL WEAR FROM A FAILED LEFT HIP RESURFACING ARTHROPLASTY. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 11 | 2020-11
3005975929-2020-00421 | 2020-11-10 | 2019-08-07 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, HEMI-, FEMORAL, METAL |  |  |  | LOCAL QARA. IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON PATIENT'S HIP ON (B)(6) 2019 DUE TO METALLOSYS. THE PATIENT OUTCOME IT IS UNKNOWN. |  | 2020 | 11 | 2020-11
3005975929-2020-00423 | 2020-11-10 | 2007-01-30 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT THE PATIENT PRESENT ELEVATED COBALT LEVELS IN HIS/HER BLOOD (ABOVE 7PPM). THE EVENT IS NOT CURRENTLY BEING TREATED AND IT IS ONGOING. |  | 2020 | 11 | 2020-11
3005975929-2020-00424 | 2020-11-10 | 2007-01-30 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT THE PATIENT PRESENT ELEVATED COBALT LEVELS IN HIS/HER BLOOD (ABOVE 7PPM). THE EVENT IS NOT CURRENTLY BEING TREATED AND IT IS ONGOING. |  | 2020 | 11 | 2020-11
3005975929-2020-00429 | 2020-11-10 | 2016-02-11 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT THE PATIENT SUFFERED FROM MIGRATION OF THE FEMORAL COMPONENT. THE EVENT WAS TREATED VIA REVISION SURGERY, WITH REMOVAL OF THE HEAD AND CUP. |  | 2020 | 11 | 2020-11
3005975929-2020-00419 | 2020-11-10 | 2019-11-04 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, HEMI-, TRUNNION-BEARING, FEMORAL |  |  | Male | *US LEGAL MDL* IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENT RIGHT HIP DUE TO PERSISTENT HIP PAIN AND LIMITATIONS OF DAY TO DAY LIVING, ELEVATED CHROMIUM AND COBALT LEVELS AND CONCERN FOR METALLOSIS. AMONG THE INTRA-OPERATIVE FINDINGS WAS PRESENCE OF SYNOVITIC-TYPE DARKISH STAINED TISSUE CONSISTENT WITH METAL ON METAL WEAR FROM A FAILED LEFT HIP RESURFACING ARTHROPLASTY. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 11 | 2020-11
3005975929-2020-00425 | 2020-11-10 | 2020-10-18 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, AFTER A BHR HAD BEEN PERFORMED, THE PATIENT PRESENTED WITH AN INFECTION IN HIS HIP, RESULTING FROM A TOOTH INFECTION THAT WENT UNTREATED FOR OVER 6 MONTHS. THE PATIENT'S HIP WAS WASHED OUT AND THE DEVICES HAVE BEEN RETAINED THUS FAR. THE INFORMATION ABOUT THE SURGEON, HOSPITAL AND DATE OF PRIMARY SURGERY IS UNKNOWN. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 11 | 2020-11
3005975929-2020-00431 | 2020-11-13 | 2020-10-21 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT THE PATIENT UNDERWENT A REVISION SURGERY OF THE BHR THA RIGHT HIP IMPLANT. THE DATE OF THE PRIMARY SURGERY IS UNKNOWN. ALL IMPLANTS WERE REMOVED. THE ACETABULAR SIDE AND FEMORAL HEAD WERE REPLACED WITH ZIMMER BIOMET TRIFLANGE CUSTOM IMPLANT. STEM WAS REVISED TO SIZE 16SO 240MM REDAPT SLEEVELESS STEM. THE REASON FOR WHY THE REVISION SURGERY WAS PERFORMED IS UNKNOWN. SURGEON DOES NOT BLAME IMPLANTS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 11 | 2020-11
3005975929-2020-00432 | 2020-11-13 | 2020-10-21 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT THE PATIENT UNDERWENT A REVISION SURGERY OF THE BHR THA RIGHT HIP IMPLANT. THE DATE OF THE PRIMARY SURGERY IS UNKNOWN. ALL IMPLANTS WERE REMOVED. THE ACETABULAR SIDE AND FEMORAL HEAD WERE REPLACED WITH ZIMMER BIOMET TRIFLANGE CUSTOM IMPLANT. STEM WAS REVISED TO SIZE 16SO 240MM REDAPT SLEEVELESS STEM. THE REASON FOR WHY THE REVISION SURGERY WAS PERFORMED IS UNKNOWN. SURGEON DOES NOT BLAME IMPLANTS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 11 | 2020-11
3005975929-2020-00435 | 2020-11-13 | 2013-02-04 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT THE PATIENT UNDERWENT A RIGHT HIP DEBRIDEMENT DUE TO AN INFECTION. NO IMPLANTS WERE REMOVED DURING THIS PROCEDURE. |  | 2020 | 11 | 2020-11
3005975929-2020-00436 | 2020-11-13 | 2013-02-04 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT THE PATIENT UNDERWENT A RIGHT HIP DEBRIDEMENT DUE TO AN INFECTION. NO IMPLANTS WERE REMOVED DURING THIS PROCEDURE. |  | 2020 | 11 | 2020-11
3005975929-2020-00447 | 2020-11-17 | 2020-08-12 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 59 YR |  | IT WAS REPORTED THAT THE, AFTER A BHR THA CONSTRUCT HAD BEEN IMPLANTED, THE PATIENT EXPERIENCED PAIN ON HIP, TOXIC LEVELS OF COBALT AND CHROMIUM, SECONDARY TO TRUNNIONS; JOINT EFFUSIONS ARE ALSO SHOWN ON MRI. A MEDICALLY-INDICATED REVISION SURGERY WAS PERFORMED TO ADDRESS THE ADVERSE EVENT. THE PATIENT IS DOING FINE. |  | 2020 | 11 | 2020-11
3005975929-2020-00449 | 2020-11-17 | 2019-04-09 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 56 YR |  | IT WAS REPORTED THAT, AFTER A BHR THA SURGERY HAD BEEN PERFORMED ON THE LEFT HIP, THE PATIENT EXPERIENCED SEVERE PAIN, LIMITED MOBILITY, AND ELEVATED COBALT AND CHROMIUM LEVELS. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2019 TO ADDRESS THIS EVENT. AMONG THE INTRA-OPERATIVE FINDINGS AND DIAGNOSES WERE METALLOSIS, AND SPECIFICALLY TRUNNIONOSIS AND CORROSION AT THE NECK-STEM JUNCTION, AS A RESULT OF THE PREMATURE FAILURE OF THE DEVICE. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 11 | 2020-11
3005975929-2020-00448 | 2020-11-17 | 2020-08-12 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 59 YR |  | IT WAS REPORTED THAT THE, AFTER A BHR THA CONSTRUCT HAD BEEN IMPLANTED, THE PATIENT EXPERIENCED PAIN ON HIP, TOXIC LEVELS OF COBALT AND CHROMIUM, SECONDARY TO TRUNIONOSIS; JOINT EFFUSIONS ARE ALSO SHOWN ON MRI. A MEDICALLY-INDICATED REVISION SURGERY WAS PERFORMED TO ADDRESS THE ADVERSE EVENT. THE PATIENT IS DOING FINE. |  | 2020 | 11 | 2020-11
3005975929-2020-00466 | 2020-11-26 | 2018-11-20 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | PLAINTIFF UNDERWENT MEDICALLY-INDICATED REVISION OF THE BHR RIGHT HIP IMPLANTS ON (B)(6) 2018. THE PATIENT REVISION SURGERY WAS PERFORMED DUE TO SEVERE PAIN, FAILED RIGHT HIP RESURFACING, STIFFNESS, LOSS OF MOBILITY, METALLOSIS, PSEUDOTUMOR, NECROSIS, AND ADVERSE TISSUE REACTION. THE CUP AND HEAD WERE EXPLANTED. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 11 | 2020-11
3005975929-2020-00477 | 2020-12-03 | 2009-10-16 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | PATIENT UNDERWENT MEDICALLY-INDICATED REVISION OF THE R3-THA RIGHT HIP IMPLANT. THE PATIENT REVISION SURGERY WAS PERFORMED ON (B)(6) 2009, DUE TO UNKNOWN REASONS. ONLY THE R3 METAL LINER, HEAD AND SLEEVE WERE EXPLANTED. |  | 2020 | 12 | 2020-12
3005975929-2020-00476 | 2020-12-03 | 2020-11-10 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | BIRMINGHAM HIP MODULAR HEAD (BHMH) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT THE PATIENT UNDERWENT REVISION SURGERY DUE TO METALLOSIS. |  | 2020 | 12 | 2020-12
3005975929-2020-00479 | 2020-12-03 | 2019-02-19 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | PATIENT UNDERWENT MEDICALLY-INDICATED REVISION OF THE R3-THA RIGHT HIP IMPLANT. THE PATIENT REVISION SURGERY WAS PERFORMED ON (B)(6) 2019, DUE TO SEVERE PAIN, ELEVATED COBALT AND CHROMIUM IONS, AND METALLOSIS. AMONG THE INTRA-OPERATIVE FINDINGS AND/OR DIAGNOSES THERE WAS A FAILED TOTAL HIP ARTHROPLASTY SECONDARY TO METAL ON METAL ADVERSE WEAR AND THE METAL STAINED TISSUE IN PLAINTIFF¿S HIP JOINT AS THE RESULT OF PREMATURE FAILURE OF THE DEVICE. ONLY THE R3 METAL LINER, HEAD AND SLEEVE WERE EXPLANTED. |  | 2020 | 12 | 2020-12
3005975929-2020-00485 | 2020-12-09 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENT LEFT HIP ON UNKNOWN DATE. THE REVISION SURGERY WAS PERFORMED DUE TO AN ADVERSE LOCAL TISSUE REACTION TO METAL DEBRIS AND GRAY-STAINED TISSUE AND JOINT FLUID; PSEUDOTUMOR; AND ELEVATED COBALT AND CHROMIUM LEVELS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 12 | 2020-12
3005975929-2020-00486 | 2020-12-09 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENT LEFT HIP ON UNKNOWN DATE. THE REVISION SURGERY WAS PERFORMED DUE TO AN ADVERSE LOCAL TISSUE REACTION TO METAL DEBRIS AND GRAY-STAINED TISSUE AND JOINT FLUID; PSEUDOTUMOR; AND ELEVATED COBALT AND CHROMIUM LEVELS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 12 | 2020-12
3005975929-2020-00487 | 2020-12-09 | 2019-06-26 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENT RIGHT HIP ON (B)(6) 2019. THE REVISION SURGERY WAS PERFORMED DUE TO AN ADVERSE LOCAL TISSUE REACTION TO METAL DEBRIS AND GRAY-STAINED TISSUE AND JOINT FLUID; PSEUDOTUMOR; AND ELEVATED COBALT AND CHROMIUM LEVELS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 12 | 2020-12
3005975929-2020-00488 | 2020-12-09 | 2019-06-26 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENT RIGHT HIP ON (B)(6) 2019. THE REVISION SURGERY WAS PERFORMED DUE TO AN ADVERSE LOCAL TISSUE REACTION TO METAL DEBRIS AND GRAY-STAINED TISSUE AND JOINT FLUID; PSEUDOTUMOR; AND ELEVATED COBALT AND CHROMIUM LEVELS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 12 | 2020-12
3005975929-2020-00511 | 2020-12-16 | 2020-08-18 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 60 YR | Male | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON (B)(6) 2020 ON THE PATIENT LEFT HIP. THE REVISION SURGERY WAS PERFORMED DUE TO PAIN, DIFFICULTY WALKING, AND SIGNIFICANTLY ELEVATED CHROMIUM AND COBALT LEVELS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 12 | 2020-12
3005975929-2020-00509 | 2020-12-16 | 2020-08-18 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 60 YR | Male | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON (B)(6) 2020 ON THE PATIENT RIGHT HIP. THE REVISION SURGERY WAS PERFORMED DUE TO PAIN, DIFFICULTY WALKING, AND SIGNIFICANTLY ELEVATED CHROMIUM AND COBALT LEVELS. AMONG THE INTRAOPERATIVE FINDINGS THERE WAS: PRESENCE OF METALLOSIS AND SYNOVITIS IN THE LINING OF THE JOINT CONSISTENT WITH METAL ON METAL WEAR FROM A FAILED LEFT HIP RESURFACING ARTHROPLASTY. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 12 | 2020-12
3005975929-2020-00507 | 2020-12-16 | 2020-09-03 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 64 YR | Female | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENT RIGHT HIP ON (B)(6) 2020. THE REVISION SURGERY WAS PERFORMED DUE TO UNEXPLAINED PAIN, SIGNIFICANTLY ELEVATED CHROMIUM AND COBALT LEVELS AND PSEUDOTUMOR FROM FAILED RIGHT HIP RESURFACING. AMONG THE INTRAOPERATIVE FINDINGS WERE: LARGE AMOUNT OF NECROTIC MATERIAL IN THE HIP FROM THE PSEUDOTUMOR, AS WELL AS ADDITIONAL NECROTIC DEBRIS AND METALLIC DEBRIS, CONSISTENT WITH METAL ON METAL WEAR FROM A FAILED RIGHT HIP RESURFACING ARTHROPLASTY. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 12 | 2020-12
3005975929-2020-00502 | 2020-12-16 | 2019-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | THE PATIENT HASN¿T UNDERGONE A MEDICALLY-INDICATED REVISION OF THE BHR RIGHT HIP IMPLANTS YET, AT THE BEGINNING OF 2019 THE PATIENT START HAVING PAIN ALSO THE PATIENT SUFFERED FROM ELEVATED LEVELS OF COBALT (20.4 MCG/L) AND CHROMIUM (2.9 MCG/L). ALTHOUGH THE RIGHT HIP HAS NOT BEEN REVISED, THE RIGHT HIP CONTRIBUTION TO THE ELEVATED ION LEVELS ON BLOOD CANNOT BY DISCARDED. |  | 2020 | 12 | 2020-12
3005975929-2020-00496 | 2020-12-16 | 2019-11-11 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENT LEFT HIP ON (B)(6) 2019. THE REVISION SURGERY WAS PERFORMED DUE TO HIP PAIN, AND OTHER COMPLICATIONS CAUSED BY THE FAILURE OF THE DEVICES. AMONG THE INTRAOPERATIVE FINDINGS THERE WAS A FLUID COLLECTION IN THE HIP JOINT. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 12 | 2020-12
3005975929-2020-00497 | 2020-12-16 | 2019-11-11 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENT LEFT HIP ON (B)(6) 2019. THE REVISION SURGERY WAS PERFORMED DUE TO HIP PAIN, AND OTHER COMPLICATIONS CAUSED BY THE FAILURE OF THE DEVICES. AMONG THE INTRAOPERATIVE FINDINGS THERE WAS A FLUID COLLECTION IN THE HIP JOINT. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 12 | 2020-12
3005975929-2020-00500 | 2020-12-16 | 2019-12-12 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENTS LEFT HIP ON (B)(6) 2019. THE REVISION SURGERY WAS PERFORMED DUE TO SHARP, DEBILITATING NERVE PAIN BEGINNING 2019; PAIN AND STIFFNESS; METALLOSIS; DIFFICULTY WALKING; INSOMNIA AND SLEEPLESSNESS; AND OTHER INJURIES. AMONG THE INTRA-OPERATIVE FINDINGS AND DIAGNOSES WERE: FAILED LEFT METAL ON METAL TOTAL HIP REPLACEMENT, SIGNIFICANT METALLOSIS, PSEUDOTUMOR, ELEVATED LEVELS OF COBALT (20.4 MCG/L) AND CHROMIUM (2.9 MCG/L), OSTEOLYSIS WITH METAL DEBRIS, AND OTHER DAMAGE AROUND THE LEFT-SIDE HIP JOINT. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 12 | 2020-12
3005975929-2020-00501 | 2020-12-16 | 2019-12-12 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENTS LEFT HIP ON (B)(6) 2019. THE REVISION SURGERY WAS PERFORMED DUE TO SHARP, DEBILITATING NERVE PAIN BEGINNING 2019; PAIN AND STIFFNESS; METALLOSIS; DIFFICULTY WALKING; INSOMNIA AND SLEEPLESSNESS; AND OTHER INJURIES. AMONG THE INTRA-OPERATIVE FINDINGS AND DIAGNOSES WERE: FAILED LEFT METAL ON METAL TOTAL HIP REPLACEMENT, SIGNIFICANT METALLOSIS, PSEUDOTUMOR, ELEVATED LEVELS OF COBALT (20.4 MCG/L) AND CHROMIUM (2.9 MCG/L), OSTEOLYSIS WITH METAL DEBRIS, AND OTHER DAMAGE AROUND THE LEFT-SIDE HIP JOINT. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 12 | 2020-12
3005975929-2020-00510 | 2020-12-16 | 2020-08-18 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 60 YR | Male | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON (B)(6) 2020 ON THE PATIENT RIGHT HIP. THE REVISION SURGERY WAS PERFORMED DUE TO PAIN, DIFFICULTY WALKING, AND SIGNIFICANTLY ELEVATED CHROMIUM AND COBALT LEVELS. AMONG THE INTRAOPERATIVE FINDINGS THERE WAS: PRESENCE OF METALLOSIS AND SYNOVITIS IN THE LINING OF THE JOINT CONSISTENT WITH METAL ON METAL WEAR FROM A FAILED LEFT HIP RESURFACING ARTHROPLASTY. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 12 | 2020-12
3005975929-2020-00503 | 2020-12-16 | 2019-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | THE PATIENT HASN¿T UNDERGONE A MEDICALLY-INDICATED REVISION OF THE BHR RIGHT HIP IMPLANTS YET, AT THE BEGINNING OF 2019 THE PATIENT START HAVING PAIN ALSO THE PATIENT SUFFERED FROM ELEVATED LEVELS OF COBALT (20.4 MCG/L) AND CHROMIUM (2.9 MCG/L). ALTHOUGH THE RIGHT HIP HAS NOT BEEN REVISED, THE RIGHT HIP CONTRIBUTION TO THE ELEVATED ION LEVELS ON BLOOD CANNOT BY DISCARDED. |  | 2020 | 12 | 2020-12
3005975929-2020-00515 | 2020-12-19 | 2020-11-24 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A PATIENT UNDERWENT REVISION SURGERY DUE TO PAIN AND METALLOSIS. THE PATIENT PRESENTED RISING COCR BLOOD LEVELS. THE DR. CONVERTED TO BHR DUAL MOBILITY IMPLANT. |  | 2020 | 12 | 2020-12
3005975929-2020-00516 | 2020-12-19 | 2020-11-24 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A PATIENT UNDERWENT REVISION SURGERY DUE TO PAIN AND METALLOSIS. THE PATIENT PRESENTED RISING COCR BLOOD LEVELS. THE DR. CONVERTED TO BHR DUAL MOBILITY IMPLANT. |  | 2020 | 12 | 2020-12
3005975929-2020-00526 | 2020-12-29 | 2018-07-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENTS LEFT HIP IMPLANT ON (B)(6) 2018. THE REVISION SURGERY WAS PERFORMED DUE TO BURNING SENSATIONS, PAIN, DISCOMFORT, DETERIORATING HEALTH, MENTAL FUZZINESS, SIGNIFICANT MALAISE, INSTABILITY, LACK OF BALANCE, METALLOSIS, COBALT, AND CHROMIUM POISONING AND OTHER SERIOUS NEUROLOGIC DEFICITS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 12 | 2020-12
3005975929-2020-00528 | 2020-12-29 | 2016-11-29 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT THE CLINICAL SUBJECT SUFFERED FROM ELEVATED ION LEVELS. THE EVENT WAS TREATED VIA REVISION SURGERY OF THE ACETABULAR CUP. |  | 2020 | 12 | 2020-12
3005975929-2020-00527 | 2020-12-29 | 2018-07-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENTS LEFT HIP IMPLANT ON (B)(6) 2018. THE REVISION SURGERY WAS PERFORMED DUE TO BURNING SENSATIONS, PAIN, DISCOMFORT, DETERIORATING HEALTH, MENTAL FUZZINESS, SIGNIFICANT MALAISE, INSTABILITY, LACK OF BALANCE, METALLOSIS, COBALT, AND CHROMIUM POISONING AND OTHER SERIOUS NEUROLOGIC DEFICITS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 12 | 2020-12
3005975929-2020-00531 | 2020-12-31 | 1901-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 59 YR | Male | IT WAS REPORTED THAT THE PATIENT UNDERWENT A REVISION SURGERY ON NOVEMBER 11TH ASSOCIATED TO PAIN AND ELEVATED ION LEVELS IN THE BLOOD. |  | 2020 | 12 | 2020-12
3005975929-2020-00534 | 2020-12-31 | 2018-04-13 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | US LEGAL MDL. IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENTS RIGHT HIP ON (B)(6) 2018. THE REVISION SURGERY WAS PERFORMED DUE TO METALLOSIS ELEVATED LEVELS OF COBALT AND CHROMIUM. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 12 | 2020-12
3005975929-2020-00535 | 2020-12-31 | 2018-04-13 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | US LEGAL MDL. IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENTS RIGHT HIP ON (B)(6) 2018. THE REVISION SURGERY WAS PERFORMED DUE TO METALLOSIS ELEVATED LEVELS OF COBALT AND CHROMIUM. THE PATIENT OUTCOME IS UNKNOWN. |  | 2020 | 12 | 2020-12
3005975929-2020-00539 | 2020-12-31 | 2019-11-09 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 82 YR |  | IT WAS REPORTED THAT THE PATIENT PRESENTED ELEVATED ION LEVELS. THE EVENT WAS TREATED VIA REVISION SURGERY OF THE ACETABULUM. |  | 2020 | 12 | 2020-12
3005975929-2020-00530 | 2020-12-31 | 2020-11-11 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, HEMI-, FEMORAL, METAL |  | 59 YR | Male | IT WAS REPORTED THAT THE PATIENT UNDERWENT A REVISION SURGERY ON NOVEMBER 11TH ASSOCIATED TO PAIN AND ELEVATED ION LEVELS IN THE BLOOD. |  | 2020 | 12 | 2020-12
3005975929-2021-00011 | 2021-01-15 | 2014-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿HIP INTERNATIONAL 2014 - CLINICAL SIGNIFICANCE OF HETEROTOPIC OSSIFICATION AFTER METAL-ON-METAL TOTAL HIP RESURFACING¿, 1 CASE OF DEBRIDEMENT, IRRIGATION, AND IV ANTIBIOTICS WERE REPORTED DUE TO WOUND INFECTION WHILE USING BHR-THA RESURFACING SYSTEM. |  | 2021 | 1 | 2021-01
3005975929-2021-00010 | 2021-01-15 | 2020-12-18 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿HIP INTERNATIONAL 2014 - CLINICAL SIGNIFICANCE OF HETEROTOPIC OSSIFICATION AFTER METAL-ON-METAL TOTAL HIP RESURFACING¿, A CASE OF FEMORAL NECK FRACTURE WAS REPORTED, WITH NO TREATMENT MENTIONED, WHILE USING BHR-THA RESURFACING SYSTEM. |  | 2021 | 1 | 2021-01
3005975929-2021-00009 | 2021-01-15 | 2014-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿HIP INTERNATIONAL 2014 - CLINICAL SIGNIFICANCE OF HETEROTOPIC OSSIFICATION AFTER METAL-ON-METAL TOTAL HIP RESURFACING¿, 2 DEEP VEIN THROMBOSIS CASES WERE REPORTED, WITH NO TREATMENT MENTIONED, WHILE USING BHR-THA RESURFACING SYSTEM. |  | 2021 | 1 | 2021-01
3005975929-2021-00008 | 2021-01-15 | 2014-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿HIP INTERNATIONAL 2014 - CLINICAL SIGNIFICANCE OF HETEROTOPIC OSSIFICATION AFTER METAL-ON-METAL TOTAL HIP RESURFACING¿, 3 REVISION SURGERIES, WHERE THE HEAD WAS REMOVED, WERE REPORTED DUE TO FEMORAL NECK FRACTURE WHILE USING BHR-THA RESURFACING SYSTEM. |  | 2021 | 1 | 2021-01
3005975929-2021-00012 | 2021-01-17 | 2019-12-05 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT THE PATIENT UNDERWENT REVISION SURGERY ON THE RIGHT HIP, AFTER A PRIMARY BHR RESURFACING SYSTEM SURGERY, DUE TO ELEVATED METAL ION LEVELS, INFLAMMATION, AND SWELLING. |  | 2021 | 1 | 2021-01
3005975929-2021-00018 | 2021-01-18 | 2019-12-05 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT THE PATIENT UNDERWENT REVISION SURGERY ON THE RIGHT HIP, AFTER A PRIMARY BHR RESURFACING SYSTEM SURGERY, DUE TO ELEVATED METAL ION LEVELS, INFLAMMATION, AND SWELLING. |  | 2021 | 1 | 2021-01
3005975929-2021-00024 | 2021-01-22 | 2019-12-23 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Female | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENT LEFT HIP ON A UNKNOWN DATE. THE REVISION SURGERY WAS PERFORMED DUE TO PAIN, LIMITED MOBILITY, OSTEOLYSIS, BEARING SURFACE WEAR, AND ELEVATED METAL ION LEVELS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 1 | 2021-01
3005975929-2021-00028 | 2021-01-22 | 2019-07-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) DUAL MOBILITY INSERT | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENT LEFT HIP ON (B)(6) 2019. THE PATIENT REVISION SURGERY WAS PERFORMED DUE TO SEVEN DISLOCATIONS OF HER HIP JOINT. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 1 | 2021-01
3005975929-2021-00023 | 2021-01-22 | 2019-12-23 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Female | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENT LEFT HIP. THE DATE IS UNKNOWN. THE REVISION SURGERY WAS PERFORMED DUE TO PAIN, LIMITED MOBILITY, OSTEOLYSIS, BEARING SURFACE WEAR, AND ELEVATED METAL ION LEVELS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 1 | 2021-01
3005975929-2021-00033 | 2021-01-25 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT THE PATIENT UNDERWENT REVISION SURGERY DUE TO ELEVATED METAL ION LEVELS ON HIS BLOOD. THE IMPLANT WAS IN VIVO FOR OVER 11 YEARS BUT AFTER 9 YEARS THE CHROMIUM LEVELS RAISED TO 193 NMOLS AND COBALT TO 453. THE PATIENT CONVERTED TO A THR CERAMIC ON CERAMIC SYSTEM. |  | 2021 | 1 | 2021-01
3005975929-2021-00034 | 2021-01-25 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT THE PATIENT UNDERWENT REVISION SURGERY DUE TO ELEVATED METAL ION LEVELS ON HIS BLOOD. THE IMPLANT WAS IN VIVO FOR OVER 11 YEARS BUT AFTER 9 YEARS THE CHROMIUM LEVELS RAISED TO 193 NMOLS AND COBALT TO 453. THE PATIENT CONVERTED TO A THR CERAMIC ON CERAMIC SYSTEM. |  | 2021 | 1 | 2021-01
3005975929-2021-00051 | 2021-01-26 | 2020-09-03 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿CLINICAL ORTHOPAEDICS AND RELATED RESEARCH 2020 - IS THE SURVIVORSHIP OF BIRMINGHAM HIP RESURFACING BETTER THAN SELECTED CONVENTIONAL HIP ARTHROPLASTIES IN MEN YOUNGER THAN 65 YEARS OF AGE? A STUDY FROM AUSTRALIA ORTHOPAEDIC ASSOCIATION NATIONAL JOINT REPLACEMENT REGISTRY.¿, 5 REVISION SURGERIES WERE REPORTED DUE TO MALPOSITION OF THE IMPLANTS WHILE HAVING BHR COMPONENTS. |  | 2021 | 1 | 2021-01
3005975929-2021-00050 | 2021-01-26 | 2020-09-03 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿CLINICAL ORTHOPAEDICS AND RELATED RESEARCH 2020 - IS THE SURVIVORSHIP OF BIRMINGHAM HIP RESURFACING BETTER THAN SELECTED CONVENTIONAL HIP ARTHROPLASTIES IN MEN YOUNGER THAN 65 YEARS OF AGE? A STUDY FROM AUSTRALIA ORTHOPAEDIC ASSOCIATION NATIONAL JOINT REPLACEMENT REGISTRY.¿, 5 REVISION SURGERIES WERE REPORTED DUE TO MALPOSITION OF THE IMPLANTS WHILE HAVING BHR COMPONENTS. |  | 2021 | 1 | 2021-01
3005975929-2021-00053 | 2021-01-26 | 2020-09-03 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿CLINICAL ORTHOPAEDICS AND RELATED RESEARCH 2020 - IS THE SURVIVORSHIP OF BIRMINGHAM HIP RESURFACING BETTER THAN SELECTED CONVENTIONAL HIP ARTHROPLASTIES IN MEN YOUNGER THAN 65 YEARS OF AGE? A STUDY FROM AUSTRALIA ORTHOPAEDIC ASSOCIATION NATIONAL JOINT REPLACEMENT REGISTRY.¿, 13 REVISION SURGERIES WERE REPORTED DUE TO UNSPECIFIED REASONS. THE REVISIONS WERE PERFORMED WHILE HAVING BHR COMPONENTS IMPLANTED. |  | 2021 | 1 | 2021-01
3005975929-2021-00052 | 2021-01-26 | 2020-09-03 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿CLINICAL ORTHOPAEDICS AND RELATED RESEARCH 2020 - IS THE SURVIVORSHIP OF BIRMINGHAM HIP RESURFACING BETTER THAN SELECTED CONVENTIONAL HIP ARTHROPLASTIES IN MEN YOUNGER THAN 65 YEARS OF AGE? A STUDY FROM AUSTRALIA ORTHOPAEDIC ASSOCIATION NATIONAL JOINT REPLACEMENT REGISTRY.¿, 13 REVISION SURGERIES WERE REPORTED DUE TO UNSPECIFIED REASONS. THE REVISIONS WERE PERFORMED WHILE HAVING BHR COMPONENTS IMPLANTED. |  | 2021 | 1 | 2021-01
3005975929-2021-00049 | 2021-01-26 | 2020-09-03 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿CLINICAL ORTHOPAEDICS AND RELATED RESEARCH 2020 - IS THE SURVIVORSHIP OF BIRMINGHAM HIP RESURFACING BETTER THAN SELECTED CONVENTIONAL HIP ARTHROPLASTIES IN MEN YOUNGER THAN 65 YEARS OF AGE? A STUDY FROM AUSTRALIA ORTHOPAEDIC ASSOCIATION NATIONAL JOINT REPLACEMENT REGISTRY.¿, 10 REVISION SURGERIES WERE REPORTED DUE TO OSTEONECROSIS WHILE HAVING BHR COMPONENTS IMPLANTED. |  | 2021 | 1 | 2021-01
3005975929-2021-00044 | 2021-01-26 | 2020-09-03 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿CLINICAL ORTHOPAEDICS AND RELATED RESEARCH 2020 - IS THE SURVIVORSHIP OF BIRMINGHAM HIP RESURFACING BETTER THAN SELECTED CONVENTIONAL HIP ARTHROPLASTIES IN MEN YOUNGER THAN 65 YEARS OF AGE? A STUDY FROM AUSTRALIA ORTHOPAEDIC ASSOCIATION NATIONAL JOINT REPLACEMENT REGISTRY.¿, 21 REVISION SURGERIES WERE REPORTED DUE TO LYSIS, WHILE USING BHR COMPONENTS. |  | 2021 | 1 | 2021-01
3005975929-2021-00047 | 2021-01-26 | 2020-09-03 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿CLINICAL ORTHOPAEDICS AND RELATED RESEARCH 2020 - IS THE SURVIVORSHIP OF BIRMINGHAM HIP RESURFACING BETTER THAN SELECTED CONVENTIONAL HIP ARTHROPLASTIES IN MEN YOUNGER THAN 65 YEARS OF AGE? A STUDY FROM (B)(6) ORTHOPAEDIC ASSOCIATION NATIONAL JOINT REPLACEMENT REGISTRY.¿, 12 REVISION SURGERIES WERE REPORTED DUE TO PAIN OF THE IMPLANT WHILE USING BHR COMPONENTS. |  | 2021 | 1 | 2021-01
3005975929-2021-00037 | 2021-01-26 | 2020-09-03 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿CLINICAL ORTHOPAEDICS AND RELATED RESEARCH 2020 - IS THE SURVIVORSHIP OF BIRMINGHAM HIP RESURFACING BETTER THAN SELECTED CONVENTIONAL HIP ARTHROPLASTIES IN MEN YOUNGER THAN 65 YEARS OF AGE? A STUDY FROM AUSTRALIA ORTHOPAEDIC ASSOCIATION NATIONAL JOINT REPLACEMENT REGISTRY.¿, 60 REVISION SURGERIES WERE REPORTED DUE TO LOOSENING OF THE IMPLANT WHILE USING BHR COMPONENTS. |  | 2021 | 1 | 2021-01
3005975929-2021-00038 | 2021-01-26 | 2020-09-03 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿CLINICAL ORTHOPAEDICS AND RELATED RESEARCH 2020 - IS THE SURVIVORSHIP OF BIRMINGHAM HIP RESURFACING BETTER THAN SELECTED CONVENTIONAL HIP ARTHROPLASTIES IN MEN YOUNGER THAN 65 YEARS OF AGE? A STUDY FROM AUSTRALIA ORTHOPAEDIC ASSOCIATION NATIONAL JOINT REPLACEMENT REGISTRY.¿, 45 REVISION SURGERIES WERE REPORTED DUE TO FRACTURE OF THE IMPLANT WHILE USING BHR COMPONENTS. |  | 2021 | 1 | 2021-01
3005975929-2021-00039 | 2021-01-26 | 2020-09-03 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿CLINICAL ORTHOPAEDICS AND RELATED RESEARCH 2020 - IS THE SURVIVORSHIP OF BIRMINGHAM HIP RESURFACING BETTER THAN SELECTED CONVENTIONAL HIP ARTHROPLASTIES IN MEN YOUNGER THAN 65 YEARS OF AGE? A STUDY FROM AUSTRALIA ORTHOPAEDIC ASSOCIATION NATIONAL JOINT REPLACEMENT REGISTRY.¿, 45 REVISION SURGERIES WERE REPORTED DUE TO FRACTURE OF THE IMPLANT WHILE USING BHR COMPONENTS. |  | 2021 | 1 | 2021-01
3005975929-2021-00041 | 2021-01-26 | 2020-09-03 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿CLINICAL ORTHOPAEDICS AND RELATED RESEARCH 2020 - IS THE SURVIVORSHIP OF BIRMINGHAM HIP RESURFACING BETTER THAN SELECTED CONVENTIONAL HIP ARTHROPLASTIES IN MEN YOUNGER THAN 65 YEARS OF AGE? A STUDY FROM AUSTRALIA ORTHOPAEDIC ASSOCIATION NATIONAL JOINT REPLACEMENT REGISTRY.¿, 32 REVISION SURGERIES WERE REPORTED DUE TO METAL RELATED PATHOLOGY ASSOCIATED TO THE IMPLANT, WHILE USING BHR COMPONENTS. |  | 2021 | 1 | 2021-01
3005975929-2021-00048 | 2021-01-26 | 2020-09-03 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿CLINICAL ORTHOPAEDICS AND RELATED RESEARCH 2020 - IS THE SURVIVORSHIP OF BIRMINGHAM HIP RESURFACING BETTER THAN SELECTED CONVENTIONAL HIP ARTHROPLASTIES IN MEN YOUNGER THAN 65 YEARS OF AGE? A STUDY FROM AUSTRALIA ORTHOPAEDIC ASSOCIATION NATIONAL JOINT REPLACEMENT REGISTRY.¿, 10 REVISION SURGERIES WERE REPORTED DUE TO OSTEONECROSIS WHILE HAVING BHR COMPONENTS IMPLANTED. |  | 2021 | 1 | 2021-01
3005975929-2021-00040 | 2021-01-26 | 2020-09-03 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿CLINICAL ORTHOPAEDICS AND RELATED RESEARCH 2020 - IS THE SURVIVORSHIP OF BIRMINGHAM HIP RESURFACING BETTER THAN SELECTED CONVENTIONAL HIP ARTHROPLASTIES IN MEN YOUNGER THAN 65 YEARS OF AGE? A STUDY FROM AUSTRALIA ORTHOPAEDIC ASSOCIATION NATIONAL JOINT REPLACEMENT REGISTRY.¿, 32 REVISION SURGERIES WERE REPORTED DUE TO METAL RELATED PATHOLOGY ASSOCIATED TO THE IMPLANT, WHILE USING BHR COMPONENTS. |  | 2021 | 1 | 2021-01
3005975929-2021-00046 | 2021-01-26 | 2020-09-03 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿CLINICAL ORTHOPAEDICS AND RELATED RESEARCH 2020 - IS THE SURVIVORSHIP OF BIRMINGHAM HIP RESURFACING BETTER THAN SELECTED CONVENTIONAL HIP ARTHROPLASTIES IN MEN YOUNGER THAN 65 YEARS OF AGE? A STUDY FROM AUSTRALIA ORTHOPAEDIC ASSOCIATION NATIONAL JOINT REPLACEMENT REGISTRY.¿, 12 REVISION SURGERIES WERE REPORTED DUE TO PAIN OF THE IMPLANT WHILE USING BHR COMPONENTS. |  | 2021 | 1 | 2021-01
3005975929-2021-00043 | 2021-01-26 | 2020-09-03 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿CLINICAL ORTHOPAEDICS AND RELATED RESEARCH 2020 - IS THE SURVIVORSHIP OF BIRMINGHAM HIP RESURFACING BETTER THAN SELECTED CONVENTIONAL HIP ARTHROPLASTIES IN MEN YOUNGER THAN 65 YEARS OF AGE? A STUDY FROM AUSTRALIA ORTHOPAEDIC ASSOCIATION NATIONAL JOINT REPLACEMENT REGISTRY.¿, 17 REVISION SURGERIES WERE REPORTED DUE TO INFECTION, WHILE USING BHR COMPONENTS. |  | 2021 | 1 | 2021-01
3005975929-2021-00042 | 2021-01-26 | 2020-09-03 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿CLINICAL ORTHOPAEDICS AND RELATED RESEARCH 2020 - IS THE SURVIVORSHIP OF BIRMINGHAM HIP RESURFACING BETTER THAN SELECTED CONVENTIONAL HIP ARTHROPLASTIES IN MEN YOUNGER THAN 65 YEARS OF AGE? A STUDY FROM AUSTRALIA ORTHOPAEDIC ASSOCIATION NATIONAL JOINT REPLACEMENT REGISTRY.¿ 21 REVISION SURGERIES WERE REPORTED DUE TO LYSIS, WHILE USING BHR COMPONENTS. |  | 2021 | 1 | 2021-01
3005975929-2021-00045 | 2021-01-26 | 2020-09-03 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿CLINICAL ORTHOPAEDICS AND RELATED RESEARCH 2020 - IS THE SURVIVORSHIP OF BIRMINGHAM HIP RESURFACING BETTER THAN SELECTED CONVENTIONAL HIP ARTHROPLASTIES IN MEN YOUNGER THAN 65 YEARS OF AGE? A STUDY FROM AUSTRALIA ORTHOPAEDIC ASSOCIATION NATIONAL JOINT REPLACEMENT REGISTRY.¿, 17 REVISION SURGERIES WERE REPORTED DUE TO INFECTION WHILE USING BHR COMPONENTS. |  | 2021 | 1 | 2021-01
3005975929-2021-00054 | 2021-01-27 | 2020-09-03 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿CLINICAL ORTHOPAEDICS AND RELATED RESEARCH 2020 - IS THE SURVIVORSHIP OF BIRMINGHAM HIP RESURFACING BETTER THAN SELECTED CONVENTIONAL HIP ARTHROPLASTIES IN MEN YOUNGER THAN 65 YEARS OF AGE? A STUDY FROM AUSTRALIA ORTHOPAEDIC ASSOCIATION NATIONAL JOINT REPLACEMENT REGISTRY.¿, 60 REVISION SURGERIES WERE REPORTED DUE TO LOOSENING OF THE IMPLANT WHILE USING BHR COMPONENTS. |  | 2021 | 1 | 2021-01
3005975929-2021-00057 | 2021-01-28 | 2009-04-07 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 64 YR |  | IT WAS REPORTED THAT, AFTER A RESURFACING HIP CONSTRUCT HAD BEEN IMPLANTED, THE PATIENT EXPERIENCED LOOSENING OF FEMORAL COMPONENT. A CONVERSION FROM RESURFACING HIP TO R3-THA WAS PERFORMED: THE RESURFACING HEAD WAS EXPLANTED, WHILE THE ORIGINAL ACETABULAR COMPONENT WAS RETAINED; AN ANTHOLOGY STEM WAS IMPLANTED, AND IT IS VERY LIKELY THAT A BHR MODULAR HEAD AND SLEEVE WERE ADDED AT THIS POINT. THE ISSUE WAS ONGOING. |  | 2021 | 1 | 2021-01
3005975929-2021-00060 | 2021-01-28 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT THE PATIENT UNDERWENT REVISION SURGERY DUE TO PERSISTENT HIP PAIN, ELEVATED CHROMIUM AND COBALT LEVELS (METALLOSIS DEBRIS FOUND) AND MRI CHANGES CONSISTENT WITH A TRUNNIONOSIS, ALL FROM A FAILED RIGHT TOTAL HIP ARTHROPLASTY. |  | 2021 | 1 | 2021-01
3005975929-2021-00061 | 2021-01-28 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 69 YR |  | IT WAS REPORTED THAT THE PATIENT UNDERWENT REVISION SURGERY DUE TO PERSISTENT HIP PAIN, ELEVATED CHROMIUM AND COBALT LEVELS (METALLOSIS DEBRIS FOUND) AND MRI CHANGES CONSISTENT WITH A TRUNNIONOSIS, ALL FROM A FAILED RIGHT TOTAL HIP ARTHROPLASTY. |  | 2021 | 1 | 2021-01
3005975929-2021-00069 | 2021-02-04 | 2020-09-30 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿STUDY TO ASSESS THE RATE OF ADVERSE REACTION TO METAL DEBRIS IN HIP RESURFACING AT A MINIMUM 13-YEAR FOLLOW-UP.¿, 1 REVISION SURGERY WAS REPORTED DUE TO PAINFUL PSEUDOTUMOR WHILE USING BHR COMPONENTS. |  | 2021 | 2 | 2021-02
3005975929-2021-00064 | 2021-02-04 | 2020-09-30 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿STUDY TO ASSESS THE RATE OF ADVERSE REACTION TO METAL DEBRIS IN HIP RESURFACING AT A MINIMUM 13-YEAR FOLLOW-UP.¿, 1 REVISION SURGERY WAS REPORTED DUE TO PAIN AND ELEVATED COBALT/CHROMIUM LEVEL WHILE USING BHR COMPONENTS. |  | 2021 | 2 | 2021-02
3005975929-2021-00065 | 2021-02-04 | 2020-09-30 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿STUDY TO ASSESS THE RATE OF ADVERSE REACTION TO METAL DEBRIS IN HIP RESURFACING AT A MINIMUM 13-YEAR FOLLOW-UP.¿, 1 REVISION SURGERY WAS REPORTED DUE TO PAIN AND ELEVATED COBALT/CHROMIUM LEVEL WHILE USING BHR COMPONENTS. |  | 2021 | 2 | 2021-02
3005975929-2021-00066 | 2021-02-04 | 2020-09-30 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿STUDY TO ASSESS THE RATE OF ADVERSE REACTION TO METAL DEBRIS IN HIP RESURFACING AT A MINIMUM 13-YEAR FOLLOW-UP.¿, 1 REVISION SURGERY WAS REPORTED DUE TO INEXPLICABLE PAIN WHILE USING BHR COMPONENTS. |  | 2021 | 2 | 2021-02
3005975929-2021-00068 | 2021-02-04 | 2020-09-30 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿STUDY TO ASSESS THE RATE OF ADVERSE REACTION TO METAL DEBRIS IN HIP RESURFACING AT A MINIMUM 13-YEAR FOLLOW-UP.¿, 1 REVISION SURGERY WAS REPORTED DUE TO PAINFUL PSEUDOTUMOR WHILE USING BHR COMPONENTS. |  | 2021 | 2 | 2021-02
3005975929-2021-00079 | 2021-02-04 | 2020-09-30 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿STUDY TO ASSESS THE RATE OF ADVERSE REACTION TO METAL DEBRIS IN HIP RESURFACING AT A MINIMUM 13-YEAR FOLLOW-UP.¿, 31 CASES OF ADVERSE REACTION TO METAL DEBRIS WITH EITHER ASYMPTOMATIC OR MINIMAL SYMPTOMS, PAIN AND STIFFNESS WERE REPORTED. |  | 2021 | 2 | 2021-02
3005975929-2021-00067 | 2021-02-04 | 2020-09-30 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿STUDY TO ASSESS THE RATE OF ADVERSE REACTION TO METAL DEBRIS IN HIP RESURFACING AT A MINIMUM 13-YEAR FOLLOW-UP¿, 1 REVISION SURGERY WAS REPORTED DUE TO VASCULAR NECROSIS AND HEAD LOOSENING WHILE USING BHR COMPONENTS. |  | 2021 | 2 | 2021-02
3005975929-2021-00077 | 2021-02-04 | 2020-09-30 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿STUDY TO ASSESS THE RATE OF ADVERSE REACTION TO METAL DEBRIS IN HIP RESURFACING AT A MINIMUM 13-YEAR FOLLOW-UP.¿, 1 REVISION SURGERY WAS REPORTED DUE TO NECK FEMUR FRACTURE WHILE USING BHR COMPONENTS. |  | 2021 | 2 | 2021-02
3005975929-2021-00076 | 2021-02-04 | 2020-09-30 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿STUDY TO ASSESS THE RATE OF ADVERSE REACTION TO METAL DEBRIS IN HIP RESURFACING AT A MINIMUM 13-YEAR FOLLOW-UP.¿, 1 REVISION SURGERY WAS REPORTED DUE TO PAIN AND ELEVATED COBALT/CHROMIUM LEVEL WHILE USING BHR COMPONENTS. |  | 2021 | 2 | 2021-02
3005975929-2021-00075 | 2021-02-04 | 2020-09-30 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿STUDY TO ASSESS THE RATE OF ADVERSE REACTION TO METAL DEBRIS IN HIP RESURFACING AT A MINIMUM 13-YEAR FOLLOW-UP.¿, 1 REVISION SURGERY WAS REPORTED DUE TO PAIN AND ELEVATED COBALT/CHROMIUM LEVEL WHILE USING BHR COMPONENTS. |  | 2021 | 2 | 2021-02
3005975929-2021-00074 | 2021-02-04 | 2020-07-31 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿RETURN TO WORK AFTER HIP RESURFACING¿, 1 HEMATOMA WAS SURGICALLY EVACUATED AT WEEK 3, WITH AN UNCOMPLICATED POSTOPERATIVE COURSE WAS REPORTED WHILE USING BHR COMPONENTS. |  | 2021 | 2 | 2021-02
3005975929-2021-00073 | 2021-02-04 | 2020-07-31 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿RETURN TO WORK AFTER HIP RESURFACING¿, 1 HIP JOINT LAVAGE TREATED WITH ANTIBIOTIC THERAPY FOR CASE OF INFECTION WAS REPORTED WHILE USING BHR COMPONENTS. |  | 2021 | 2 | 2021-02
3005975929-2021-00071 | 2021-02-04 | 2020-09-30 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿STUDY TO ASSESS THE RATE OF ADVERSE REACTION TO METAL DEBRIS IN HIP RESURFACING AT A MINIMUM 13-YEAR FOLLOW-UP.¿, 1 REVISION SURGERY WAS REPORTED DUE TO A LARGE PSEUDOTUMOR WHILE USING BHR COMPONENTS. |  | 2021 | 2 | 2021-02
3005975929-2021-00070 | 2021-02-04 | 2020-09-30 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿STUDY TO ASSESS THE RATE OF ADVERSE REACTION TO METAL DEBRIS IN HIP RESURFACING AT A MINIMUM 13-YEAR FOLLOW-UP.¿, 1 REVISION SURGERY WAS REPORTED DUE TO A LARGE PSEUDOTUMOR WHILE USING BHR COMPONENTS. |  | 2021 | 2 | 2021-02
3005975929-2021-00078 | 2021-02-04 | 2020-09-30 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT ON LITERATURE REVIEW ¿STUDY TO ASSESS THE RATE OF ADVERSE REACTION TO METAL DEBRIS IN HIP RESURFACING AT A MINIMUM 13-YEAR FOLLOW-UP.¿, 31 CASES OF ADVERSE REACTION TO METAL DEBRIS WITH EITHER ASYMPTOMATIC OR MINIMAL SYMPTONS, PAIN AND STIFNESS WERE REPORTED. |  | 2021 | 2 | 2021-02
3005975929-2021-00082 | 2021-02-14 | 2020-10-22 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | *US LEGAL MDL* IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENT LEFT HIP ON (B)(6) 2020. THE REVISION SURGERY WAS PERFORMED DUE TO LEFT HIP PAIN, LEFT HIP LIMITED MOTION, FAILURE OF LEFT HIP BHR DEVICE, METALLOSIS (ELEVATED COBALT AND CHROMIUM METAL IONS). THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 2 | 2021-02
3005975929-2021-00083 | 2021-02-14 | 2020-10-22 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | US LEGAL MDL. IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENT LEFT HIP ON (B)(6) 2020. THE REVISION SURGERY WAS PERFORMED DUE TO LEFT HIP PAIN, LEFT HIP LIMITED MOTION, FAILURE OF LEFT HIP BHR DEVICE, METALLOSIS (ELEVATED COBALT AND CHROMIUM METAL IONS). THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 2 | 2021-02
3005975929-2021-00084 | 2021-02-15 | 2020-09-17 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENT LEFT HIP ON (B)(6) 2020. THE REVISION SURGERY WAS PERFORMED DUE TO EXTREME LEFT HIP PAIN, LEFT HIP LIMITED MOTION, FAILURE OF CONSERVATIVE NON-SURGICAL TREATMENT MEASURES, FAILURE OF LEFT HIP BHR DEVICE, LEFT HIP PSEUDOTUMOR, ELEVATED COBALT AND CHROMIUM SERUM LEVELS, METALLOSIS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 2 | 2021-02
3005975929-2021-00085 | 2021-02-15 | 2020-09-17 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENT LEFT HIP ON (B)(6) 2020. THE REVISION SURGERY WAS PERFORMED DUE TO EXTREME LEFT HIP PAIN, LEFT HIP LIMITED MOTION, FAILURE OF CONSERVATIVE NON-SURGICAL TREATMENT MEASURES, FAILURE OF LEFT HIP BHR DEVICE, LEFT HIP PSEUDOTUMOR, ELEVATED COBALT AND CHROMIUM SERUM LEVELS, METALLOSIS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 2 | 2021-02
3005975929-2021-00132 | 2021-03-09 | 2020-12-14 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 57 YR | Male | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENT LEFT HIP ON (B)(6) 2020. THE REVISION SURGERY WAS PERFORMED DUE TO PAIN, STIFFNESS, MUSCLE PAIN, LOSS OF STRENGTH, LIMPING, INABILITY TO WALK PROPER, INABILITY TO RUN, INTERFERED WITH WORK. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 3 | 2021-03
3005975929-2021-00135 | 2021-03-11 | 2020-10-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENT´S RIGHT HIP ON (B)(6) 2020. THE REVISION SURGERY WAS PERFORMED DUE TO CHRONIC PAIN, HEARING LOSS, AND INCREASINGLY HIGHLY ELEVATED COBALT AND CHROMIUM LEVELS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 3 | 2021-03
3005975929-2021-00136 | 2021-03-11 | 2020-10-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENT´S RIGHT HIP ON (B)(6) 2020. THE REVISION SURGERY WAS PERFORMED DUE TO CHRONIC PAIN, HEARING LOSS, AND INCREASINGLY HIGHLY ELEVATED COBALT AND CHROMIUM LEVELS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 3 | 2021-03
3005975929-2021-00140 | 2021-03-11 | 2020-09-09 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENT'S LEFT HIP ON (B)(6) 2020. THE REVISION SURGERY WAS PERFORMED DUE TO PROGRESSIVE PAIN, AND EXTREMELY ELEVATED CHROMIUM AND COBALT LEVELS. AMONG THE INTRA OPERATIVE FINDINGS THERE WAS A CONFIRMED SEVERE METALLOSIS, TISSUE STAINING, METAL DEBRIS, ALONG WITH A SMALL PSEUDOTUMOR. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 3 | 2021-03
3005975929-2021-00154 | 2021-03-23 | 2021-01-19 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | US LEGAL MDL - IT WAS REPORTED THAT A REVISION SURGERY OF THE BHR IMPLANTS OF THE LEFT HIP ON (B)(6) 2021 WAS PERFORMED DUE TO CHRONIC SEVERE PAIN, DIFFICULTY WALKING, AND ELEVATED COBALT AND CHROMIUM LEVELS. AMONG THE INTRAOPERATIVE FINDINGS THERE WAS LYTIC DEBRIS, EXTENSIVE AND DESTRUCTIVE PELVIC OSTEOLYSIS CAUSING A MASSIVE AMOUNT OF PELVIC BONE LOSS, ALL CONSISTENT WITH TRUNNIONOSIS CAUSING AN ADVERSE LOCAL TISSUE REACTION, RESULTING IN A FAILED LEFT HIP RESURFACING ARTHROPLASTY. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 3 | 2021-03
3005975929-2021-00152 | 2021-03-23 | 2020-05-18 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | ''US LEGAL MDL'' IT WAS REPORTED THAT A REVISION SURGERY OF THE BHR IMPLANTS OF THE RIGHT HIP ON (B)(6) 2020 DUE TO PAIN IN HIP AND GROIN DUE TO A PSEUDOTUMOR AND METALLOSIS. PATIENT ALSO PRESENTED DIFFICULTY WALKING; STRENGTH LOSS, METAL DEBRIS, LARGE PSEUDOTUMOR, NECROTIC TISSUE, LARGE AMOUNTS OF CLOUDY FLUID PRESENT, CORROSION, ELEVATED LEVELS OF COBALT AND CHROMIUM, AND OTHER DAMAGE AROUND THE RIGHT-SIDE HIP JOINT. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 3 | 2021-03
3005975929-2021-00153 | 2021-03-23 | 2020-05-18 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | US LEGAL MDL - IT WAS REPORTED THAT A REVISION SURGERY OF THE BHR IMPLANTS OF THE RIGHT HIP ON (B)(6) 2020 DUE TO PAIN IN HIP AND GROIN DUE TO A PSEUDOTUMOR AND METALLOSIS. PATIENT ALSO PRESENTED DIFFICULTY WALKING; STRENGTH LOSS, METAL DEBRIS, LARGE PSEUDOTUMOR, NECROTIC TISSUE, LARGE AMOUNTS OF CLOUDY FLUID PRESENT, CORROSION, ELEVATED LEVELS OF COBALT AND CHROMIUM, AND OTHER DAMAGE AROUND THE RIGHT-SIDE HIP JOINT. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 3 | 2021-03
3005975929-2021-00155 | 2021-03-23 | 2021-01-19 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | *US LEGAL MDL* IT WAS REPORTED THAT A REVISION SURGERY OF THE BHR IMPLANTS OF THE LEFT HIP ON (B)(6) 2021 WAS PERFORMED DUE TO CHRONIC SEVERE PAIN, DIFFICULTY WALKING, AND ELEVATED COBALT AND CHROMIUM LEVELS. AMONG THE INTRAOPERATIVE FINDINGS THERE WAS LYTIC DEBRIS, EXTENSIVE AND DESTRUCTIVE PELVIC OSTEOLYSIS CAUSING A MASSIVE AMOUNT OF PELVIC BONE LOSS, ALL CONSISTENT WITH TRUNNIONOSIS CAUSING AN ADVERSE LOCAL TISSUE REACTION, RESULTING IN A FAILED LEFT HIP RESURFACING ARTHROPLASTY. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 3 | 2021-03
3005975929-2021-00169 | 2021-03-28 | 2021-03-05 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT THERE WAS A REVISION BHR HIP SURGERY DUE TO PAIN. NO OTHER COMPLICATIONS WERE REPORTED |  | 2021 | 3 | 2021-03
3005975929-2021-00166 | 2021-03-28 | 2019-10-11 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | *US LEGAL MDL* IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED ON THE PATIENT'S RIGHT HIP ON (B)(6) 2019 DUE TO TO PAIN, PSEUDOTUMOR, ELEVATED CHROMIUM AND COBALT LEVELS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 3 | 2021-03
3005975929-2021-00174 | 2021-04-06 | 2021-03-11 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, THE PATIENT UNDERWENT A REVISION SURGERY ON (B)(6) 2021 DUE TO HIGH METAL IONS ON BLOOD. THE BHR WAS ORIGINALLY IMPLANTED ON (B)(6) 2013. THE METAL BLOOD LEVELS WERE CO 17.2 AND CR 10.4. CONVERTED TO AN ANTHOLOGY WITH BDMH. PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 4 | 2021-04
3005975929-2021-00180 | 2021-04-12 | 2020-03-09 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT THE PLAINTIFF UNDERWENT A MEDICALLY INDICATED REVISION OF THE BHR-THA RIGHT HIP ON (B)(6) 2020 DUE TO PAIN, IMPLANT FAILURE AND ELEVATED ION LEVELS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 4 | 2021-04
3005975929-2021-00181 | 2021-04-12 | 2020-03-09 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT THE PLAINTIFF UNDERWENT A MEDICALLY INDICATED REVISION OF THE BHR-THA RIGHT HIP ON (B)(6) 2020 DUE TO PAIN, IMPLANT FAILURE AND ELEVATED ION LEVELS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 4 | 2021-04
3005975929-2021-00195 | 2021-04-26 | 2019-09-24 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 63 YR | Male | US LEGAL MDL. IT WAS REPORTED THAT THE PATIENT UNDERWENT MEDICALLY INDICATED REVISION OF THE BHR IMPLANTS OF THE LEFT HIP ON (B)(6) 2019 DUE TO PAIN, ELEVATED COBALT AND CHROMIUM LEVELS AND METALLOSIS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 4 | 2021-04
3005975929-2021-00196 | 2021-04-26 | 2019-09-24 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 63 YR | Male | US LEGAL MDL. IT WAS REPORTED THAT THE PATIENT UNDERWENT MEDICALLY INDICATED REVISION OF THE BHR IMPLANTS OF THE LEFT HIP ON (B)(6) 2019 DUE TO PAIN, ELEVATED COBALT AND CHROMIUM LEVELS AND METALLOSIS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 4 | 2021-04
3005975929-2021-00198 | 2021-04-28 | 2014-12-30 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 46 YR |  | US LEGAL MDL. IT WAS REPORTED THAT THE PATIENT UNDERWENT MEDICALLY INDICATED REVISION OF THE BHR IMPLANTS OF THE RIGHT HIP ON (B)(6) 2014 DUE TO IMPLANT FAILURE, ADVERSE LOCAL TISSUE REACTION, TISSUE DAMAGE, METALLOSIS, ELEVATED METAL LEVELS, ABNORMAL FLUID COLLECTION, PAIN, INFLAMMATION, AND DIFFICULTY WALKING. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 4 | 2021-04
3005975929-2021-00199 | 2021-04-28 | 2014-12-30 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 46 YR |  | US LEGAL MDL. IT WAS REPORTED THAT THE PATIENT UNDERWENT MEDICALLY INDICATED REVISION OF THE BHR IMPLANTS OF THE RIGHT HIP ON (B)(6) 2014 DUE TO IMPLANT FAILURE, ADVERSE LOCAL TISSUE REACTION, TISSUE DAMAGE, METALLOSIS, ELEVATED METAL LEVELS, ABNORMAL FLUID COLLECTION, PAIN, INFLAMMATION, AND DIFFICULTY WALKING. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 4 | 2021-04
3005975929-2021-00201 | 2021-05-04 | 2019-06-13 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | US LEGAL MDL. IT WAS REPORTED THAT THE PATIENT UNDERWENT MEDICALLY INDICATED REVISION OF THE BHR IMPLANTS OF THE RIGHT HIP ON (B)(6) 2019 DUE TO PAIN, INFLAMMATION, LOSS OF MOBILITY AND ELEVATED METAL LEVEL, ASEPTIC LOOSENING OF THE CUP. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00202 | 2021-05-04 | 2019-06-13 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | US LEGAL MDL. IT WAS REPORTED THAT THE PATIENT UNDERWENT MEDICALLY INDICATED REVISION OF THE BHR IMPLANTS OF THE RIGHT HIP ON (B)(6) 2019 DUE TO PAIN, INFLAMMATION, LOSS OF MOBILITY AND ELEVATED METAL LEVEL, ASEPTIC LOOSENING OF THE CUP. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00214 | 2021-05-10 | 2019-07-16 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD |  |  |  | *US LEGAL MDL* IT WAS REPORTED THAT THE PATIENT UNDERWENT MEDICALLY INDICATED REVISION OF THE BHR IMPLANTS OF THE LEFT HIP ON (B)(6) 2019. THE PATIENT REVISION SURGERY WAS PERFORMED DUE TO PAIN, ELEVATED METAL ION LEVELS, AND LARGE FLUID COLLECTION. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00215 | 2021-05-11 | 2019-07-16 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | US LEGAL MDL. IT WAS REPORTED THAT THE PATIENT UNDERWENT MEDICALLY INDICATED REVISION OF THE BHR IMPLANTS OF THE LEFT HIP ON (B)(6) 2019. THE PATIENT REVISION SURGERY WAS PERFORMED DUE TO PAIN, ELEVATED METAL ION LEVELS, AND LARGE FLUID COLLECTION. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00223 | 2021-05-12 | 2020-10-17 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED DUE TO THE PATIENT SUFFERED A BHR METAL ON METAL REACTION. ADDITIONAL DETAILS HAVE NOT BEEN PROVIDED AT THIS TIME. THE OUTCOME OF THE PATIENT IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00224 | 2021-05-12 | 2018-10-31 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, HEMI-, TRUNNION-BEARING, FEMORAL, METAL/POLYACETAL |  |  |  | ON THE LITERATURE ARTICLE NAMED "MODULAR JUNCTION MAY BE MORE PROBLEMATIC THAN BEARING WEAR IN METAL-ON-METAL TOTAL HIP ARTHROPLASTY", IT WAS REPORTED THAT, AFTER AN ANTHOLOGY STEM HAD BEEN IMPLANTED ON 1 PATIENT, AN UNSPECIFIED FRACTURE WAS REPORTED. IT WAS NOT REPORTED WHERE THE FRACTURE OCCURRED. IT IS ALSO UNKNOWN IF/HOW THE ADVERSE EVENT WAS TREATED. THE OUTCOME OF THE PATIENT IS ALSO UNKNOWN. ALL AVAILABLE INFORMATION HAS BEEN DISCLOSED. IF ADDITIONAL INFORMATION SHOULD BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. |  | 2021 | 5 | 2021-05
3005975929-2021-00222 | 2021-05-12 | 2020-10-17 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT A REVISION SURGERY WAS PERFORMED DUE TO THE PATIENT SUFFERED A BHR METAL ON METAL REACTION. ADDITIONAL DETAILS HAVE NOT BEEN PROVIDED AT THIS TIME. THE OUTCOME OF THE PATIENT IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00225 | 2021-05-12 | 2018-10-31 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, HEMI-, TRUNNION-BEARING, FEMORAL, METAL/POLYACETAL |  |  |  | ON THE LITERATURE ARTICLE NAMED "MODULAR JUNCTION MAY BE MORE PROBLEMATIC THAN BEARING WEAR IN METAL-ON-METAL TOTAL HIP ARTHROPLASTY", IT WAS REPORTED THAT, AFTER AN ANTHOLOGY STEM HAD BEEN IMPLANTED ON 1 PATIENT, 1 REVISION SURGERY WAS PERFORMED DUE TO UNSPECIFIED REASONS. IT IS UNKNOWN WHICH DEVICES WERE EXPLANTED. THE OUTCOME OF THE PATIENT IS ALSO UNKNOWN. ALL AVAILABLE INFORMATION HAS BEEN DISCLOSED. IF ADDITIONAL INFORMATION SHOULD BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. |  | 2021 | 5 | 2021-05
3005975929-2021-00241 | 2021-05-13 | 2018-10-31 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) |  |  |  | ON THE LITERATURE ARTICLE NAMED "MODULAR JUNCTION MAY BE MORE PROBLEMATIC THAN BEARING WEAR IN METAL-ON-METAL TOTAL HIP ARTHROPLASTY", IT WAS REPORTED THAT, AFTER A SMF STEM HAD BEEN IMPLANTED, AN UNSPECIFIED NUMBER OF PATIENTS EXHIBITED HIGH COBALT AND CHROMIUM IONS LEVEL IN BLOOD. THE MEAN CONCENTRATION WAS 7.2 G/L IN CO AND 11.4 G/L IN CR. IT IS UNKNOWN IF/HOW THE ISSUE WAS TREATED. |  | 2021 | 5 | 2021-05
3005975929-2021-00242 | 2021-05-13 | 2018-10-31 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI |  |  |  | ON THE LITERATURE ARTICLE NAMED "MODULAR JUNCTION MAY BE MORE PROBLEMATIC THAN BEARING WEAR IN METAL-ON-METAL TOTAL HIP ARTHROPLASTY", IT WAS REPORTED THAT, AFTER AN ANTHOLOGY STEM HAD BEEN IMPLANTED ON 1 PATIENT, 1 PATIENT HAD A VERY HIGH COBALT IONS MEASURE (APPROX. 12 G/L) AND IT WAS NOT ASSOCIATED WITH ARMD. IT IS UNKNOWN IF/HOW THE ISSUE WAS TREATED. |  | 2021 | 5 | 2021-05
3005975929-2021-00243 | 2021-05-13 | 2018-10-31 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI |  |  |  | ON THE LITERATURE ARTICLE NAMED "MODULAR JUNCTION MAY BE MORE PROBLEMATIC THAN BEARING WEAR IN METAL-ON-METAL TOTAL HIP ARTHROPLASTY", IT WAS REPORTED THAT, AFTER A SMF STEM HAD BEEN IMPLANTED, AN UNSPECIFIED NUMBER OF PATIENTS EXHIBITED HIGH COBALT AND CHROMIUM IONS LEVEL IN BLOOD. THE MEAN CONCENTRATION WAS 7.2 G/L IN CO AND 11.4 G/L IN CR. IT IS UNKNOWN IF/HOW THE ISSUE WAS TREATED. |  | 2021 | 5 | 2021-05
3005975929-2021-00237 | 2021-05-13 | 2021-03-03 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, AFTER A BHR CONSTRUCT HAD BEEN IMPLANTED ON (B)(6) 2010, THE PLAINTIFF EXPERIENCED SEVERE AND INCREASING HIP PAIN, SWELLING, ADVERSE LOCAL TISSUE REACTION, TISSUE DAMAGE, METALLOSIS, ELEVATED IONS LEVELS AND DISLOCATION OF THE HIP. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2021 TO TREAT THESE ADVERSE EVENTS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00240 | 2021-05-13 | 2018-10-31 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) |  |  |  | ON THE LITERATURE ARTICLE NAMED "MODULAR JUNCTION MAY BE MORE PROBLEMATIC THAN BEARING WEAR IN METAL-ON-METAL TOTAL HIP ARTHROPLASTY", IT WAS REPORTED THAT, AFTER AN ANTHOLOGY STEM HAD BEEN IMPLANTED ON 1 PATIENT, 1 PATIENT HAD A VERY HIGH COBALT IONS MEASURE (APPROX. 12 G/L) AND IT WAS NOT ASSOCIATED WITH ARMD. IT IS UNKNOWN IF/HOW THE ISSUE WAS TREATED. |  | 2021 | 5 | 2021-05
3005975929-2021-00235 | 2021-05-13 | 2018-10-31 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, HEMI-, TRUNNION-BEARING, FEMORAL, METAL/POLYACETAL |  |  |  | ON THE LITERATURE ARTICLE NAMED "MODULAR JUNCTION MAY BE MORE PROBLEMATIC THAN BEARING WEAR IN METAL-ON-METAL TOTAL HIP ARTHROPLASTY", IT WAS REPORTED THAT, AFTER A SMF STEM HAD BEEN IMPLANTED ON 4 PATIENTS, 4 REVISION SURGERIES WERE PERFORMED DUE TO ADVERSE REACTION TO METAL DEBRIS. DURING THESE SURGERIES, BLACKENED CORRODED DEBRIS WAS OBSERVED AT THE JUNCTION BETWEEN THE STEM AND THE MODULAR NECK. ALL STEMS WERE WELL FIXED AND REQUIRED AN EXTENDED FEMORAL OSTEOTOMY FOR THEIR REMOVAL. THE BEARINGS WERE CHANGED TO CERAMIC ON CERAMIC. |  | 2021 | 5 | 2021-05
3005975929-2021-00234 | 2021-05-13 | 2018-10-31 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | ON THE LITERATURE ARTICLE NAMED "MODULAR JUNCTION MAY BE MORE PROBLEMATIC THAN BEARING WEAR IN METAL-ON-METAL TOTAL HIP ARTHROPLASTY", IT WAS REPORTED THAT, AFTER A SMF STEM HAD BEEN IMPLANTED ON 4 PATIENTS, 4 REVISION SURGERIES WERE PERFORMED DUE TO ADVERSE REACTION TO METAL DEBRIS. DURING THESE SURGERIES, BLACKENED CORRODED DEBRIS WAS OBSERVED AT THE JUNCTION BETWEEN THE STEM AND THE MODULAR NECK. ALL STEMS WERE WELL FIXED AND REQUIRED AN EXTENDED FEMORAL OSTEOTOMY FOR THEIR REMOVAL. THE BEARINGS WERE CHANGED TO CERAMIC ON CERAMIC. |  | 2021 | 5 | 2021-05
3005975929-2021-00233 | 2021-05-13 | 2018-10-31 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | ON THE LITERATURE ARTICLE NAMED "MODULAR JUNCTION MAY BE MORE PROBLEMATIC THAN BEARING WEAR IN METAL-ON-METAL TOTAL HIP ARTHROPLASTY", IT WAS REPORTED THAT, AFTER A SMF STEM HAD BEEN IMPLANTED ON 1 PATIENT, 1 REVISION SURGERY WAS PERFORMED DUE TO ADVERSE REACTION TO METAL DEBRIS, SEVERE ACETABULAR OSTEOLYSIS AND ACETABULAR COMPONENT LOOSENING. DURING THIS SURGERY, BLACKENED CORRODED DEBRIS WAS OBSERVED AT THE JUNCTION BETWEEN THE STEM AND THE MODULAR NECK. THE STEM WAS WELL FIXED AND REQUIRED AN EXTENDED FEMORAL OSTEOTOMY TO REMOVE IT. THE BEARINGS WERE CHANGED TO CERAMIC ON CERAMIC. |  | 2021 | 5 | 2021-05
3005975929-2021-00232 | 2021-05-13 | 2018-10-31 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | ON THE LITERATURE ARTICLE NAMED "MODULAR JUNCTION MAY BE MORE PROBLEMATIC THAN BEARING WEAR IN METAL-ON-METAL TOTAL HIP ARTHROPLASTY", IT WAS REPORTED THAT, AFTER A SMF STEM HAD BEEN IMPLANTED ON 1 PATIENT, 1 REVISION SURGERY WAS PERFORMED DUE TO ADVERSE REACTION TO METAL DEBRIS, SEVERE ACETABULAR OSTEOLYSIS AND ACETABULAR COMPONENT LOOSENING. DURING THIS SURGERY, BLACKENED CORRODED DEBRIS WAS OBSERVED AT THE JUNCTION BETWEEN THE STEM AND THE MODULAR NECK. THE STEM WAS WELL FIXED AND REQUIRED AN EXTENDED FEMORAL OSTEOTOMY TO REMOVE IT. THE BEARINGS WERE CHANGED TO CERAMIC ON CERAMIC. |  | 2021 | 5 | 2021-05
3005975929-2021-00231 | 2021-05-13 | 2018-10-31 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | ON THE LITERATURE ARTICLE NAMED "MODULAR JUNCTION MAY BE MORE PROBLEMATIC THAN BEARING WEAR IN METAL-ON-METAL TOTAL HIP ARTHROPLASTY", IT WAS REPORTED THAT, AFTER AN ANTHOLOGY STEM HAD BEEN IMPLANTED ON 5 PATIENTS, 5 REVISION SURGERIES WERE PERFORMED DUE TO ADVERSE REACTION TO METAL DEBRIS. DURING THESE SURGERIES, BLACKENED CORRODED DEBRIS WAS OBSERVED AT THE JUNCTION BETWEEN THE STEM AND THE FEMORAL HEAD ADAPTER SLEEVE. WEAR AT THE SURFACE OF THE TRUNNION WAS ALSO OBSERVED. STEMS WERE WELL FIXED AND WERE NOT EXPLANTED. THE BEARINGS WERE CHANGED TO CERAMIC ON CERAMIC. |  | 2021 | 5 | 2021-05
3005975929-2021-00229 | 2021-05-13 | 2018-10-31 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, HEMI-, TRUNNION-BEARING, FEMORAL, METAL/POLYACETAL |  |  |  | ON THE LITERATURE ARTICLE NAMED "MODULAR JUNCTION MAY BE MORE PROBLEMATIC THAN BEARING WEAR IN METAL-ON-METAL TOTAL HIP ARTHROPLASTY", IT WAS REPORTED THAT, AFTER A SMF STEM HAD BEEN IMPLANTED ON 1 PATIENT, 1 REVISION SURGERY WAS PERFORMED DUE TO ADVERSE REACTION TO METAL DEBRIS. DURING THIS SURGERY, BLACKENED CORRODED DEBRIS WAS OBSERVED AT THE JUNCTION BETWEEN THE STEM AND THE MODULAR NECK. THE STEM WAS NOT WELL FIXED AND WAS EXPLANTED WITHOUT THE NEED OF EXTENDED FEMORAL OSTEOTOMY. THE BEARINGS WERE CHANGED TO CERAMIC ON CERAMIC. |  | 2021 | 5 | 2021-05
3005975929-2021-00226 | 2021-05-13 | 2018-10-31 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, HEMI-, TRUNNION-BEARING, FEMORAL, METAL/POLYACETAL |  |  |  | ON THE LITERATURE ARTICLE NAMED "MODULAR JUNCTION MAY BE MORE PROBLEMATIC THAN BEARING WEAR IN METAL-ON-METAL TOTAL HIP ARTHROPLASTY", IT WAS REPORTED THAT, AFTER AN ANTHOLOGY STEM HAD BEEN IMPLANTED ON 1 PATIENT, 1 REVISION SURGERY WAS PERFORMED DUE TO AN UNSPECIFIED INFECTION. IT IS UNKNOWN WHICH DEVICES WERE EXPLANTED. THE OUTCOME OF THE PATIENT IS ALSO UNKNOWN. ALL AVAILABLE INFORMATION HAS BEEN DISCLOSED. IF ADDITIONAL INFORMATION SHOULD BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. |  | 2021 | 5 | 2021-05
3005975929-2021-00227 | 2021-05-13 | 2018-10-31 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, HEMI-, TRUNNION-BEARING, FEMORAL, METAL/POLYACETAL |  |  |  | ON THE LITERATURE ARTICLE NAMED "MODULAR JUNCTION MAY BE MORE PROBLEMATIC THAN BEARING WEAR IN METAL-ON-METAL TOTAL HIP ARTHROPLASTY", IT WAS REPORTED THAT, AFTER AN ANTHOLOGY STEM HAD BEEN IMPLANTED ON 5 PATIENTS, 5 REVISION SURGERIES WERE PERFORMED DUE TO ADVERSE REACTION TO METAL DEBRIS. DURING THESE SURGERIES, BLACKENED CORRODED DEBRIS WAS OBSERVED AT THE JUNCTION BETWEEN THE STEM AND THE FEMORAL HEAD ADAPTER SLEEVE. WEAR AT THE SURFACE OF THE TRUNNION WAS ALSO OBSERVED. STEMS WERE WELL FIXED AND WERE NOT EXPLANTED. THE BEARINGS WERE CHANGED TO CERAMIC ON CERAMIC. |  | 2021 | 5 | 2021-05
3005975929-2021-00228 | 2021-05-13 | 2018-10-31 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, HEMI-, TRUNNION-BEARING, FEMORAL, METAL/POLYACETAL |  |  |  | ON THE LITERATURE ARTICLE NAMED "MODULAR JUNCTION MAY BE MORE PROBLEMATIC THAN BEARING WEAR IN METAL-ON-METAL TOTAL HIP ARTHROPLASTY", IT WAS REPORTED THAT, AFTER AN ANTHOLOGY STEM HAD BEEN IMPLANTED ON 1 PATIENT, 1 REVISION SURGERY WAS PERFORMED DUE TO AN UNSPECIFIED IMPINGEMENT. IT IS UNKNOWN WHICH DEVICES WERE EXPLANTED. THE OUTCOME OF THE PATIENT IS ALSO UNKNOWN. ALL AVAILABLE INFORMATION HAS BEEN DISCLOSED. IF ADDITIONAL INFORMATION SHOULD BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. |  | 2021 | 5 | 2021-05
3005975929-2021-00230 | 2021-05-13 | 2018-10-31 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | ON THE LITERATURE ARTICLE NAMED "MODULAR JUNCTION MAY BE MORE PROBLEMATIC THAN BEARING WEAR IN METAL-ON-METAL TOTAL HIP ARTHROPLASTY", IT WAS REPORTED THAT, AFTER A SMF STEM HAD BEEN IMPLANTED ON 1 PATIENT, 1 REVISION SURGERY WAS PERFORMED DUE TO ADVERSE REACTION TO METAL DEBRIS. DURING THIS SURGERY, BLACKENED CORRODED DEBRIS WAS OBSERVED AT THE JUNCTION BETWEEN THE STEM AND THE MODULAR NECK. THE STEM WAS NOT WELL FIXED AND WAS EXPLANTED WITHOUT THE NEED OF EXTENDED FEMORAL OSTEOTOMY. THE BEARINGS WERE CHANGED TO CERAMIC ON CERAMIC. |  | 2021 | 5 | 2021-05
3005975929-2021-00236 | 2021-05-13 | 2021-03-03 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, AFTER A BHR CONSTRUCT HAD BEEN IMPLANTED ON (B)(6) 2010, THE PLAINTIFF EXPERIENCED SEVERE AND INCREASING HIP PAIN, SWELLING, ADVERSE LOCAL TISSUE REACTION, TISSUE DAMAGE, METALLOSIS, ELEVATED IONS LEVELS AND DISLOCATION OF THE HIP. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2021 TO TREAT THESE ADVERSE EVENTS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00245 | 2021-05-14 | 2021-04-29 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI |  |  |  | *US LEGAL MDL* IT WAS REPORTED THAT, AFTER A THA CONSTRUCT HAD BEEN IMPLANTED ON (B)(6) 2009, THE PLAINTIFF EXPERIENCED PAIN AND METALLOSIS. A REVISION SURGERY HAS NOT BEEN PERFORMED YET DUE TO THE PLAINTIFF UNDERGOING HEALTH ISSUES, BUT MIGHT BE PERFORMED IN THE FUTURE, NO DATES HAD BEEN GIVEN YET. THE PLAINTIFF OUTCOME IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00248 | 2021-05-14 | 2019-01-02 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) |  | 52 YR | Male | US LEGAL MDL. IT WAS REPORTED THAT, AFTER A BHR CONSTRUCT HAD BEEN IMPLANTED ON (B)(6) 2011, THE PLAINTIFF EXPERIENCED PAIN, LIMITED MOBILITY, ELEVATED METAL ION LEVELS AND METALLOSIS. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2019 TO TREAT THIS ADVERSE EVENT. THE PLAINTIFF OUTCOME IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00247 | 2021-05-14 | 2018-08-08 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI |  |  |  | *US LEGAL MDL* IT WAS REPORTED THAT, AFTER A R3 THA CONSTRUCT HAD BEEN IMPLANTED ON (B)(6) 2009, THE PLAINTIFF EXPERIENCED PAIN AND METALLOSIS. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2018 TO TREAT THESE ADVERSE EVENTS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00257 | 2021-05-18 | 2019-05-14 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP ACET (STANDARD) | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP ACET (STANDARD) |  |  |  | ON THE LITERATURE ARTICLE NAMED "CONVERSION OF HIP RESURFACING WITH RETENTION OF MONOBLOCK ACETABULAR SHELL USING DUAL-MOBILITY COMPONENTS", IT WAS REPORTED THAT, FROM THE BC GROUP, 1 PATIENT UNDERWENT REVISION SURGERY DUE TO AN UNSPECIFIED FRACTURE. THE FEMORAL HEAD AND THE CUP WERE EXPLANTED. THE OUTCOME OF THE PATIENTS IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00253 | 2021-05-18 | 2019-05-14 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP ACET (STANDARD) | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP ACET (STANDARD) |  |  |  | ON THE LITERATURE ARTICLE NAMED "CONVERSION OF HIP RESURFACING WITH RETENTION OF MONOBLOCK ACETABULAR SHELL USING DUAL-MOBILITY COMPONENTS", IT WAS REPORTED THAT, FROM THE BC GROUP, 1 PATIENT UNDERWENT REVISION SURGERY DUE TO AN UNEXPLAINED PAIN. THE FEMORAL HEAD AND THE CUP WERE EXPLANTED. THE OUTCOME OF THE PATIENTS IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00255 | 2021-05-18 | 2019-05-14 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) |  |  |  | ON THE LITERATURE ARTICLE NAMED "CONVERSION OF HIP RESURFACING WITH RETENTION OF MONOBLOCK ACETABULAR SHELL USING DUAL-MOBILITY COMPONENTS", IT WAS REPORTED THAT, FROM THE BC GROUP, 1 PATIENT UNDERWENT REVISION SURGERY DUE TO AN UNSPECIFIED FRACTURE. THE FEMORAL HEAD AND THE CUP WERE EXPLANTED. THE OUTCOME OF THE PATIENTS IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00254 | 2021-05-18 | 2019-05-14 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) |  |  |  | ON THE LITERATURE ARTICLE NAMED "CONVERSION OF HIP RESURFACING WITH RETENTION OF MONOBLOCK ACETABULAR SHELL USING DUAL-MOBILITY COMPONENTS", IT WAS REPORTED THAT, FROM THE BC GROUP, 1 PATIENT UNDERWENT REVISION SURGERY DUE TO AN UNEXPLAINED PAIN. THE FEMORAL HEAD AND THE CUP WERE EXPLANTED. THE OUTCOME OF THE PATIENT IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00271 | 2021-05-18 | 2020-12-15 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW, INC. | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Male | *EU LEGAL* IT WAS REPORTED THAT, AFTER A LEFT BHR CONSTRUCT HAD BEEN IMPLANTED ON (B)(6) 2014, THE PLAINTIFF EXPERIENCED HIGH FEVER AND HIGH LEVELS OF COCR IN BLOOD. A TWO-STAGE REVISION SURGERY WAS PERFORMED TO TREAT THESE ADVERSE EVENTS; ON SEPTEMBER 2020 THE DEVICES WERE EXPLANTED, AND A CEMENT SPACER WAS APPLIED, LEAVING THE PLAINTIFF WITH NO MOBILITY. THEN ON (B)(6) 2020 THE REVISION SURGERY WAS COMPLETED WITH THE IMPLANTATION OF NEW PROSTHESIS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00269 | 2021-05-18 | 2019-05-14 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) |  |  |  | ON THE LITERATURE ARTICLE NAMED "CONVERSION OF HIP RESURFACING WITH RETENTION OF MONOBLOCK ACETABULAR SHELL USING DUAL-MOBILITY COMPONENTS", IT WAS REPORTED THAT, FROM THE BC GROUP, 20 PATIENTS UNDERWENT REVISION SURGERIES DUE TO COMPLICATIONS RELATED TO ELEVATED ION LEVELS AND ALVAL. THE FEMORAL HEAD AND THE CUP WERE EXPLANTED. THE OUTCOME OF THE PATIENTS IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00268 | 2021-05-18 | 2019-05-14 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) |  |  |  | ON THE LITERATURE ARTICLE NAMED "CONVERSION OF HIP RESURFACING WITH RETENTION OF MONOBLOCK ACETABULAR SHELL USING DUAL-MOBILITY COMPONENTS", IT WAS REPORTED THAT, FROM THE BC GROUP, 3 PATIENTS UNDERWENT REVISION SURGERIES DUE TO COMPLICATIONS RELATED TO IMPINGEMENT AND MALPOSITION. THE FEMORAL HEAD AND THE CUP WERE EXPLANTED. THE OUTCOME OF THE PATIENTS IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00256 | 2021-05-18 | 2019-05-14 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) |  |  |  | ON THE LITERATURE ARTICLE NAMED "CONVERSION OF HIP RESURFACING WITH RETENTION OF MONOBLOCK ACETABULAR SHELL USING DUAL-MOBILITY COMPONENTS", IT WAS REPORTED THAT, FROM THE RAC GROUP, 1 PATIENT UNDERWENT REVISION SURGERY DUE TO AN UNSPECIFIED FRACTURE. THE FEMORAL HEAD WAS EXPLANTED AND THE CUP WAS RETAINED. THE OUTCOME OF THE PATIENT IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00266 | 2021-05-18 | 2019-05-14 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) |  |  |  | ON THE LITERATURE ARTICLE NAMED "CONVERSION OF HIP RESURFACING WITH RETENTION OF MONOBLOCK ACETABULAR SHELL USING DUAL-MOBILITY COMPONENTS", IT WAS REPORTED THAT, FROM THE RAC GROUP, 14 PATIENTS UNDERWENT REVISION SURGERIES DUE TO COMPLICATIONS RELATED TO ELEVATED ION LEVELS AND ALVAL. THE FEMORAL HEAD WAS EXPLANTED AND THE CUP WAS RETAINED. THE OUTCOME OF THE PATIENTS IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00264 | 2021-05-18 | 2019-05-14 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) IMPL | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) IMPL |  |  |  | ON THE LITERATURE ARTICLE NAMED "CONVERSION OF HIP RESURFACING WITH RETENTION OF MONOBLOCK ACETABULAR SHELL USING DUAL-MOBILITY COMPONENTS", IT WAS REPORTED THAT, FROM THE RAC GROUP, 1 PATIENT UNDERWENT REVISION SURGERY DUE TO AN UNEXPLAINED PAIN. THE FEMORAL HEAD WAS EXPLANTED AND THE CUP WAS RETAINED. THE OUTCOME OF THE PATIENT IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00263 | 2021-05-18 | 2019-05-14 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) |  |  |  | ON THE LITERATURE ARTICLE NAMED "CONVERSION OF HIP RESURFACING WITH RETENTION OF MONOBLOCK ACETABULAR SHELL USING DUAL-MOBILITY COMPONENTS", IT WAS REPORTED THAT, FROM THE RAC GROUP, 4 YEARS AFTER THE FIRST REVISION WAS PERFORMED, ONE PATIENT HAD AN INTRA-PROSTHETIC DISLOCATION OF THE DUAL MOBILITY COMPONENT ASSOCIATED WITH EXTENSIVE POLY WEAR. THE EVENT RESULTED IN EXTENSIVE METAL DEBRIS REQUIRING REVISION OF THE ACETABULAR COMPONENT WITH BONE GRAFTING AND CONVERSION TO A CERAMIC ON POLYETHYLENE CONVENTIONAL ARTICULATION. |  | 2021 | 5 | 2021-05
3005975929-2021-00262 | 2021-05-18 | 2019-05-14 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP ACET (STANDARD) | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP ACET (STANDARD) |  |  |  | ON THE LITERATURE ARTICLE NAMED "CONVERSION OF HIP RESURFACING WITH RETENTION OF MONOBLOCK ACETABULAR SHELL USING DUAL-MOBILITY COMPONENTS", IT WAS REPORTED THAT, FROM THE BC GROUP, 19 PATIENTS UNDERWENT REVISION SURGERIES DUE TO COMPLICATIONS RELATED TO OSTEOLYSIS AND ASEPTIC LOOSENING. THE FEMORAL HEAD AND THE CUP WERE EXPLANTED. THE OUTCOME OF THE PATIENTS IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00258 | 2021-05-18 | 2019-05-14 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) |  |  |  | ON THE LITERATURE ARTICLE NAMED "CONVERSION OF HIP RESURFACING WITH RETENTION OF MONOBLOCK ACETABULAR SHELL USING DUAL-MOBILITY COMPONENTS", IT WAS REPORTED THAT, FROM THE BC GROUP, 19 PATIENTS UNDERWENT REVISION SURGERIES DUE TO COMPLICATIONS RELATED TO OSTEOLYSIS AND ASEPTIC LOOSENING. THE FEMORAL HEAD AND THE CUP WERE EXPLANTED. THE OUTCOME OF THE PATIENTS IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00259 | 2021-05-18 | 2019-05-14 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD |  |  |  | ON THE LITERATURE ARTICLE NAMED "CONVERSION OF HIP RESURFACING WITH RETENTION OF MONOBLOCK ACETABULAR SHELL USING DUAL-MOBILITY COMPONENTS", IT WAS REPORTED THAT, FROM THE RAC GROUP, 14 PATIENTS UNDERWENT REVISION SURGERIES DUE TO COMPLICATIONS RELATED TO ELEVATED ION LEVELS AND ALVAL. THE FEMORAL HEAD WAS EXPLANTED AND THE CUP WAS RETAINED. THE OUTCOME OF THE PATIENTS IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00260 | 2021-05-18 | 2019-05-14 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) DYSPLASIA SCREW | UNKN BIRMINGHAM HIP RESURFACING (BHR) DYSPLASIA SCREW |  |  |  | ON THE LITERATURE ARTICLE NAMED "CONVERSION OF HIP RESURFACING WITH RETENTION OF MONOBLOCK ACETABULAR SHELL USING DUAL-MOBILITY COMPONENTS", IT WAS REPORTED THAT, FROM THE RAC GROUP, 11 PATIENTS UNDERWENT REVISION SURGERIES DUE TO COMPLICATIONS RELATED TO OSTEOLYSIS AND ASEPTIC LOOSENING. THE FEMORAL HEAD WAS EXPLANTED AND THE CUP WAS RETAINED. THE OUTCOME OF THE PATIENTS IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00261 | 2021-05-18 | 2019-05-14 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP ACET (STANDARD) | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP ACET (STANDARD) |  |  |  | ON THE LITERATURE ARTICLE NAMED "CONVERSION OF HIP RESURFACING WITH RETENTION OF MONOBLOCK ACETABULAR SHELL USING DUAL-MOBILITY COMPONENTS", IT WAS REPORTED THAT, FROM THE BC GROUP, 20 PATIENTS UNDERWENT REVISION SURGERIES DUE TO COMPLICATIONS RELATED TO ELEVATED ION LEVELS AND ALVAL. THE FEMORAL HEAD AND THE CUP WERE EXPLANTED. THE OUTCOME OF THE PATIENTS IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00267 | 2021-05-18 | 2019-05-14 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP ACET (STANDARD) | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP ACET (STANDARD) |  |  |  | ON THE LITERATURE ARTICLE NAMED "CONVERSION OF HIP RESURFACING WITH RETENTION OF MONOBLOCK ACETABULAR SHELL USING DUAL-MOBILITY COMPONENTS", IT WAS REPORTED THAT, FROM THE BC GROUP, 3 PATIENTS UNDERWENT REVISION SURGERIES DUE TO COMPLICATIONS RELATED TO IMPINGEMENT AND MALPOSITION. THE FEMORAL HEAD AND THE CUP WERE EXPLANTED. THE OUTCOME OF THE PATIENTS IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00274 | 2021-05-19 | 2021-05-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT, AFTER BHR SURGERY HAD BEEN PERFORMED ON AN UNKNOWN DATE, THE PATIENT HAS EXPERIENCED PAIN AND HAS A FOREIGN BODY GRANULOMA. IT ALSO PRESENTS AN INCREASE IN THE CONCENTRATION OF METALLIC IONS. REVISION SURGERY OF THE BHR IMPLANTS WAS PERFORMED ON MAY 1, 2021 TO TREAT THIS ADVERSE EVENT. THE PATIENT'S OUTCOME IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00296 | 2021-05-31 | 2020-06-18 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI |  |  |  | *US LEGAL MDL* IT WAS REPORTED THAT, AFTER A R3 THA CONSTRUCT HAD BEEN IMPLANTED ON (B)(6) 2010, THE PLAINTIFF EXPERIENCED FAILED METAL ON METAL ARTICULATION, ELEVATED ION LEVELS, PAIN, AND DISABILITY. THE PLAINTIFF HAS FAILED THE CONSERVATIVE NONSURGICAL MANAGEMENT, WHICH INCLUDED AMBULATORY ASSISTIVE DEVICES, LOW IMPACT EXERCISE, AND ANTI INFLAMMATORY MEDICATION. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2020 TO TREAT THESE ADVERSE EVENTS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00295 | 2021-05-31 | 2020-02-06 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI |  | 72 YR |  | US LEGAL MDL. IT WAS REPORTED THAT, AFTER A R3 THA CONSTRUCT HAD BEEN IMPLANTED ON AN UNKNOWN DATE, THE PLAINTIFF EXPERIENCED FAILED METAL ON METAL ARTICULATION, ELEVATED ION LEVELS, PAIN, AND DISABILITY. THE PLAINTIFF HAS FAILED THE CONSERVATIVE NONSURGICAL MANAGEMENT, WHICH INCLUDED AMBULATORY ASSISTIVE DEVICES, LOW IMPACT EXERCISE, AND ANTI INFLAMMATORY MEDICATION. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2020 TO TREAT THESE ADVERSE EVENTS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00291 | 2021-05-31 | 2013-08-22 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | ON THE LITERATURE ARTICLE NAMED "HIGH FREQUENCY OF ADVERSE LOCAL TISSUE REACTIONS IN ASYMPTOMATIC PATIENTS WITH METAL-ON-METAL THA", IT WAS REPORTED THAT, AFTER A R3 SYSTEM HAD BEEN IMPLANTED ON 1 PATIENT, THE PATIENT SUFFERED FROM AN OSTEOLYTIC LESION OF 1.9X5.1 WHILE USING THE R3 MOM SYSTEM. IT IS UNKNOWN IF HOW THE ADVERSE LOCAL TISSUE REACTION WAS TREATED. THE OUTCOME OF THE PATIENT IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00290 | 2021-05-31 | 2021-05-18 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, AFTER A BHR CONSTRUCT HAD BEEN IMPLANTED ON (B)(6) 2010, THE PATIENT EXPERIENCED METALLOSIS OR ELEVATED METAL IONS. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2021 TO TREAT THIS ADVERSE EVENT. THE FEMORAL HEAD 52MM WAS EXPLANTED DURING THIS SURGERY AND IT WAS REPLACED WITH A BHR DUAL MOBILITY INSERT FROM SMITH AND NEPHEW. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 5 | 2021-05
3005975929-2021-00304 | 2021-06-01 | 2014-08-04 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI |  | 38 YR |  | *US LEGAL MDL* IT WAS REPORTED THAT, AFTER A BHR-THA CONSTRUCT HAD BEEN IMPLANTED ON THE PLAINTIFF'S RIGHT HIP ON 2012, THE PLAINTIFF EXPERIENCE LOOSENING OF THE FEMORAL HEAD AND PAIN. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2014 TO TREAT THIS ADVERSE EVENT. THE PLAINTIFF OUTCOME IS UNKNOWN. |  | 2021 | 6 | 2021-06
3005975929-2021-00311 | 2021-06-07 | 2021-05-17 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | *BILATERAL PATIENT* IT WAS REPORTED THAT, AFTER A BMHR CONSTRUCT HAD BEEN IMPLANTED ON THE PATIENT'S RIGHT AND LEFT HIP ON AN UNKNOWN DATE AROUND 2010, THE PATIENT EXPERIENCED PAIN IN BOTH HIPS. X-RAYS WERE TAKEN AND COBALT AND CHROME BLOOD ION LEVELS WERE ASSESSED, SHOWING SUGGESTIONS OF METAL ON METAL PARTICLE GENERATIONS AS A CAUSE OF ONGOING PAIN IN THE PATIENT'S HIPS. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2021 TO TREAT THIS ADVERSE EVENT, BUT IT IS UNKNOWN IN WHICH HIP WAS THIS PROCEDURE CONDUCTED. A BIRMINGHAM HIP RESURFACING (BHR) CUP, BIRMINGHAM MID-HEAD RESECTION (BMHR) FEMORAL HEAD AND UNKN BIRMINGHAM MID-HEAD RESECTION (BMHR) STEM WERE EXPLANTED. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 6 | 2021-06
3005975929-2021-00308 | 2021-06-07 | 1901-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | *BILATERAL PATIENT* IT WAS REPORTED THAT, AFTER A BMHR CONSTRUCT HAD BEEN IMPLANTED ON THE PATIENT'S RIGHT AND LEFT HIP ON AN UNKNOWN DATE AROUND 2010, THE PATIENT EXPERIENCED PAIN IN BOTH HIPS. X-RAYS WERE TAKEN AND COBALT AND CHROME BLOOD ION LEVELS WERE ASSESSED, SHOWING SUGGESTIONS OF METAL ON METAL PARTICLE GENERATIONS AS A CAUSE OF ONGOING PAIN IN THE PATIENT'S HIPS. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2021 TO TREAT THIS ADVERSE EVENT IN ONE OF THE PATIENT'S HIPS ((B)(4)), BUT IT IS UNKNOWN IF THE SECOND HIP BMHR IMPLANTS WERE REVISED OR IF A REVISION SURGERY HAS BEEN ALREADY SCHEDULED. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 6 | 2021-06
3005975929-2021-00330 | 2021-06-25 | 2020-02-26 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | US LEGAL MDL IT WAS REPORTED THAT, AFTER A BHR THA CONSTRUCT HAD BEEN IMPLANTED ON (B)(6) 2008, THE PLAINTIFF EXPERIENCED ELEVATED METAL ION LEVELS, HIP PAIN, PSEUDOTUMOR, AND OSTEOLYSIS. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2020 TO TREAT THESE ADVERSE EVENTS. DURING THE SURGERY, THE CUP WAS FOUND TO BE WRONGLY POSITIONED AND A MODERATE SIZED FLUID WAS COLLECTED. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 6 | 2021-06
3005975929-2021-00331 | 2021-06-25 | 2020-02-26 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | *US LEGAL MDL* IT WAS REPORTED THAT, AFTER A BHR THA CONSTRUCT HAD BEEN IMPLANTED ON (B)(6) 2008, THE PLAINTIFF EXPERIENCED ELEVATED METAL ION LEVELS, HIP PAIN, PSEUDOTUMOR, AND OSTEOLYSIS. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2020 TO TREAT THESE ADVERSE EVENTS. DURING THE SURGERY, THE CUP WAS FOUND TO BE WRONGLY POSITIONED AND A MODERATE SIZED FLUID WAS COLLECTED. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 6 | 2021-06
3005975929-2021-00343 | 2021-07-02 | 2018-05-22 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | *US LEGAL MDL* IT WAS REPORTED THAT, AFTER A BHR RESURFACING CONSTRUCT HAD BEEN IMPLANTED ON THE PLAINTIFF'S LEFT HIP ON (B)(6) 2008, THE PLAINTIFF EXPERIENCED SEVERE PAIN, DISCOMFORT, INFLAMMATION AROUND THE HIP IMPLANT, ELEVATED BLOOD METAL ION LEVELS AND METALLOSIS. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2018 TO TREAT THIS ADVERSE EVENT. THE PLAINTIFF'S OUTCOME IS UNKNOWN. |  | 2021 | 7 | 2021-07
3005975929-2021-00342 | 2021-07-02 | 2018-05-22 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | *US LEGAL MDL* IT WAS REPORTED THAT, AFTER A BHR RESURFACING CONSTRUCT HAD BEEN IMPLANTED ON THE PLAINTIFF'S LEFT HIP ON (B)(6) 2008, THE PLAINTIFF EXPERIENCED SEVERE PAIN, DISCOMFORT, INFLAMMATION AROUND THE HIP IMPLANT, ELEVATED BLOOD METAL ION LEVELS AND METALLOSIS. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2018 TO TREAT THIS ADVERSE EVENT. THE PLAINTIFF'S OUTCOME IS UNKNOWN. |  | 2021 | 7 | 2021-07
3005975929-2021-00344 | 2021-07-08 | 2021-05-27 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | *US LEGAL MDL* IT WAS REPORTED THAT, AFTER A BHR THA CONSTRUCT HAD BEEN IMPLANTED ON (B)(6) 2008, THE PLAINTIFF EXPERIENCED JOINT PAIN, METALLOSIS, ELEVATED ION LEVELS, AND DIFFICULTY TO AMBULATE. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2021 TO TREAT THESE ADVERSE EVENTS. DURING REVISION SURGERY, A CLOUDY MURKY FLUID WAS FOUND IN THE JOINT, AS WELL AS METAL DEBRIS, A PSEUDOTUMOR REQUIRING EXCISION, ADVERSE LOCAL TISSUE REACTION AND NEUROLYSIS OF THE SCIATIC NERVE. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 7 | 2021-07
3005975929-2021-00345 | 2021-07-08 | 2021-05-27 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | US LEGAL MDL - IT WAS REPORTED THAT, AFTER A BHR THA CONSTRUCT HAD BEEN IMPLANTED ON (B)(6) 2008, THE PLAINTIFF EXPERIENCED JOINT PAIN, METALLOSIS, ELEVATED ION LEVELS, AND DIFFICULTY TO AMBULATE. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2021 TO TREAT THESE ADVERSE EVENTS. DURING REVISION SURGERY, A CLOUDY MURKY FLUID WAS FOUND IN THE JOINT, AS WELL AS METAL DEBRIS, A PSEUDOTUMOR REQUIRING EXCISION, ADVERSE LOCAL TISSUE REACTION AND NEUROLYSIS OF THE SCIATIC NERVE. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 7 | 2021-07
3005975929-2021-00355 | 2021-07-12 | 2019-02-28 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | US LEGAL MDL. IT WAS REPORTED THAT, AFTER A SECOND REVISION SURGERY HAD BEEN PERFORMED ON PLAINTIFF'S LEFT HIP ON (B)(6) 2015 (COVERED UNDER (B)(4)), THE PLAINTIFF EXPERIENCED METALLOSIS, LOCAL ADVERSE TISSUE REACTION, ELEVATED METAL ION LEVELS, PSEUDOTUMOR FORMATION AND FLUID ACCUMULATION. A THIRD REVISION SURGERY WAS PERFORMED ON THE PLAINTIFF'S LEFT HIP ON (B)(6) 2019 TO TREAT THIS ADVERSE EVENT; THE BHR CUP WAS EXPLANTED AS WELL AS THE INSERT (DEPUY-COMPETITORS) AND THE FEMORAL HEAD (SMITH AND NEPHEW). THE PLAINTIFF'S OUTCOME IS UNKNOWN. PLAINTIFF'S FIRST REVISION SURGERY IS COVERED UNDER (B)(4). |  | 2021 | 7 | 2021-07
3005975929-2021-00359 | 2021-07-28 | 2021-06-21 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | *US LEGAL MDL* IT WAS REPORTED THAT, AFTER A BHR CONSTRUCT HAD BEEN IMPLANTED ON (B)(6) 2008, THE PLAINTIFF EXPERIENCED EXTREME HIP PAIN, ADVERSE LOCAL TISSUE REACTION, RIGHT HIP LIMITED MOTION, FAILURE OF RIGHT HIP AND METALLOSIS. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2021 TO TREAT THESE ADVERSE EVENTS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 7 | 2021-07
3005975929-2021-00360 | 2021-07-28 | 2021-06-21 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | *US LEGAL MDL* IT WAS REPORTED THAT, AFTER A BHR CONSTRUCT HAD BEEN IMPLANTED ON (B)(6) 2008, THE PLAINTIFF EXPERIENCED EXTREME HIP PAIN, ADVERSE LOCAL TISSUE REACTION, RIGHT HIP LIMITED MOTION, FAILURE OF RIGHT HIP AND METALLOSIS. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2021 TO TREAT THESE ADVERSE EVENTS. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 7 | 2021-07
3005975929-2021-00361 | 2021-07-28 | 2021-07-13 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT AFTER A R3 THA WAS PERFORMED ON AN UNKNOWN DATE, THE PATIENT EXPERIENCED ELEVATED COCR LEVELS ON BLOOD. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2021 TO TREAT THIS ADVERSE EVENT. DURING THE SURGERY THE MODULAR BHR HEAD, THE MODULAR BHR SLEEVE AND THE COCR LINER WERE EXPLANTED. A OXINIUM HEAD AND A XLPE LINER WERE IMPLANTED INSTEAD. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 7 | 2021-07
3005975929-2021-00374 | 2021-08-11 | 2019-08-28 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 59 YR |  | EU LEGAL. IT WAS REPORTED THAT THE PLAINTIFF UNDERWENT A SECOND REVISION SURGERY OF THE LEFT HIP ON (B)(6) 2019 DUE TO UNSPECIFIED REASONS. THE PATIENT UNDERWENT THE PRIMARY LEFT BHR SURGERY ON (B)(6) 2007 AND THE FIRST REVISION SURGERY ON (B)(6) 2009. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 8 | 2021-08
3005975929-2021-00372 | 2021-08-11 | 2009-04-08 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 49 YR |  | IT WAS REPORTED THAT, AFTER A BHR CONSTRUCT HAD BEEN IMPLANTED ON (B)(6) 2007, THE PLAINTIFF EXPERIENCED UNSPECIFIED SYMPTOMS. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2009 TO TREAT THIS ADVERSE EVENT. THE PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 8 | 2021-08
3005975929-2021-00373 | 2021-08-11 | 2007-12-05 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 47 YR |  | IT WAS REPORTED THAT THE PLAINTIFF AWAITS REVISION SURGERY OF HIS RIGHT HIP DUE TO ONGOING PAIN AND UNSPECIFIED INJURIES. THE REVISION SURGERY HAS NOT BEEN PERFORMED YET. THE PATIENT UNDERWENT THE PRIMARY RIGHT BHR SURGERY ON (B)(6) 2007. |  | 2021 | 8 | 2021-08
3005975929-2021-00382 | 2021-08-19 | 2010-05-27 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 53 YR |  | ''US LEGAL MDL'' IT WAS REPORTED THAT, AFTER A LEFT BHR CONSTRUCT HAD BEEN IMPLANTED ON (B)(6) 2006, THE PLAINTIFF EXPERIENCED A MECHANICAL FAILURE OF THE BIRMINGHAM HIP RESURFACING SYSTEM AND A NONUNION OF LEFT SUBCAPITAL FEMORAL NECK FRACTURE. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2010 TO TREAT THESE ADVERSE EVENTS. THE FEMORAL HEAD WAS EXPLANTED AND THE HIP WAS REVISED TO A BHR-THA SYSTEM. THE PATIENT WAS TRANSFERRED TO RECOVERY IN FAIR CONDITION. |  | 2021 | 8 | 2021-08
3005975929-2021-00378 | 2021-08-19 | 2021-01-04 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 64 YR | Female | US LEGAL MDL. IT WAS REPORTED THAT, AFTER A LEFT BHR THA CONSTRUCTS HAD BEEN IMPLANTED ON (B)(6) 2010 AS A PART OF A FIRST REVISION SURGERY, THE PLAINTIFF EXPERIENCED MECHANICAL COMPLICATIONS AND COBALT AND CHROMIUM TOXICITY. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2021 TO TREAT THESE ADVERSE EVENTS. THERE WAS SOME METALLOSIS FOUND AT THE TRUNNION. THERE WAS SOME OSTEOLYSIS FOUND IN THE ACETABULUM. THE METAL FEMORAL HEAD WAS REVISED TO A DUAL MOBILITY POLY INSERT WITH AN OXINIUM FEMORAL HEAD. THE CUP AND THE STEM WERE NOT EXPLANTED. |  | 2021 | 8 | 2021-08
3005975929-2021-00383 | 2021-08-23 | 2021-01-25 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 64 YR |  | *US LEGAL MDL* IT WAS REPORTED THAT, AFTER A RIGHT BHR CONSTRUCT HAD BEEN IMPLANTED ON (B)(6) 2006, THE PLAINTIFF EXPERIENCED A MECHANICAL COMPLICATION AND COBALT AND CHROMIUM TOXICITY. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2021 TO TREAT THESE ADVERSE EVENTS. SOME BROWN DISCOLORATION OF THE SYNOVIUM WAS FOUND, AS WELL AS OSTEOLYSIS AROUND THE ACETABULUM. THE FEMORAL HEAD WAS EXPLANTED AND THE CUP WAS RETAINED. THE HIP WAS REVISED TO A THA WITH A DUAL MOBILITY SYSTEM. THERE WERE NO INTRAOPERATIVE COMPLICATIONS AND THE PATIENT WAS TAKEN TO THE RECUPERATION ROOM IN STABLE CONDITION. |  | 2021 | 8 | 2021-08
3005975929-2021-00384 | 2021-08-23 | 2021-01-25 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 64 YR |  | US LEGAL MDL: IT WAS REPORTED THAT, AFTER A RIGHT BHR CONSTRUCT HAD BEEN IMPLANTED ON (B)(6) 2006, THE PLAINTIFF EXPERIENCED A MECHANICAL COMPLICATION AND COBALT AND CHROMIUM TOXICITY. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2021 TO TREAT THESE ADVERSE EVENTS. SOME BROWN DISCOLORATION OF THE SYNOVIUM WAS FOUND, AS WELL AS OSTEOLYSIS AROUND THE ACETABULUM. THE FEMORAL HEAD WAS EXPLANTED AND THE CUP WAS RETAINED. THE HIP WAS REVISED TO A THA WITH A DUAL MOBILITY SYSTEM. THERE WERE NO INTRAOPERATIVE COMPLICATIONS AND THE PATIENT WAS TAKEN TO THE RECUPERATION ROOM IN STABLE CONDITION. |  | 2021 | 8 | 2021-08
3005975929-2021-00388 | 2021-08-26 | 2019-02-26 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 62 YR |  | US LEGAL: IT WAS REPORTED THAT THE PLAINTIFF UNDERWENT A SECOND REVISION OF THE LEFT HIP ON (B)(6) 2019 DUE TO METAL-ON-METAL WEAR, OSTEOLYSIS, ADVERSE LOCAL TISSUE REACTION, TRUNIONOSIS, PSEUDOTUMOR AND DESTRUCTION/DETACHMENT OF ABDUCTOR MUSCLES. THE CUP AND THE STEM WERE NOTICED TO BE WELL FIXED, SO THEY WERE LEFT IN THE PATIENT. THE METAL FEMORAL HEAD WAS EXPLANTED AND REPLACED WITH A BHR DUAL MOBILITY LINER AND AN OXINIUM FEMORAL HEAD. THE PATIENT WAS TAKEN TO THE RECOVERY ROOM IN STABLE CONDITION. |  | 2021 | 8 | 2021-08
3005975929-2021-00391 | 2021-08-31 | 2020-08-18 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 59 YR | Male | IT WAS REPORTED THAT A REVISION SURGERY ON THE LEFT HIP OF THE PATIENT WAS PERFORMED ON (B)(6) 2020 DUE TO CHRONICALLY ELEVATED COBALT AND CHROMIUM LEVELS. DURING REVISION SURGERY, CORROSION WAS FOUND AROUND THE FEMORAL HEAD AND NECK JUNCTION WITH NO OBVIOUS EVIDENCE OF LOCAL TISSUE REACTION. THE PATIENT WAS STABLE AFTER THE REVISION. PRIMARY SURGERY WAS PERFORMED ON (B)(6) 2010. |  | 2021 | 8 | 2021-08
3005975929-2021-00393 | 2021-09-01 | 2020-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, AFTER A PRIMARY BHR RESURFACING CONSTRUCT HAD BEEN IMPLANTED ON THE PATIENT'S RIGHT HIP ON (B)(6) 2000, THE PATIENT EXPERIENCED THE FOLLOWING SYMPTOMS: SORE KNEE, PAIN, SWELLING, DISCOMFORT, MOVEMENT DIFFICULTIES AND TEETH AND NAIL BRITTLE. A REVISION SURGERY WAS PERFORM ON AN UNSPECIFIED DATE TO TREAT THIS ADVERSE EVENT; THE BHR CONSTRUCT WAS EXPLANTED AND REPLACED WITH AN UNSPECIFIED (B)(4) (COMPETITORS). SOON AFTER THE REVISION SURGERY OF THE RIGHT HIP, PATIENT UNDERWENT MEDICALLY INDICATED PRIMARY SURGERY OF THE LEFT HIP, ALTHOUGH IT IS NOT SPECIFIED THE TYPE OF CONSTRUCT THAT WAS IMPLANTED. OVER THE NEXT YEARS, PATIENT HAS BEEN EXPERIENCING WORSENING PAIN, INABILITY TO BEND OVER, NAILS AND TEETH BREAKAGE, FORMATION OF TUMORS, SWELLING AND LOW FOLATE IN BLOOD. THE CURRENT OUTCOME OF THE PATIENT IS NOT RECOVERED. |  | 2021 | 9 | 2021-09
3005975929-2021-00392 | 2021-09-01 | 2020-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  |  | IT WAS REPORTED THAT, AFTER A PRIMARY BHR RESURFACING CONSTRUCT HAD BEEN IMPLANTED ON THE PATIENT'S RIGHT HIP ON (B)(6) 2000, THE PATIENT EXPERIENCED THE FOLLOWING SYMPTOMS: SORE KNEE, PAIN, SWELLING, DISCOMFORT, MOVEMENT DIFFICULTIES AND TEETH AND NAIL BRITTLE. A REVISION SURGERY WAS PERFORM ON AN UNSPECIFIED DATE TO TREAT THIS ADVERSE EVENT; THE BHR CONSTRUCT WAS EXPLANTED AND REPLACED WITH AN UNSPECIFIED MATHYS CONSTRUCT (COMPETITORS). SOON AFTER THE REVISION SURGERY OF THE RIGHT HIP, PATIENT UNDERWENT MEDICALLY INDICATED PRIMARY SURGERY OF THE LEFT HIP, ALTHOUGH IT IS NOT SPECIFIED THE TYPE OF CONSTRUCT THAT WAS IMPLANTED. OVER THE NEXT YEARS, PATIENT HAS BEEN EXPERIENCING WORSENING PAIN, INABILITY TO BEND OVER, NAILS AND TEETH BREAKAGE, FORMATION OF TUMORS, SWELLING AND LOW FOLATE IN BLOOD. THE CURRENT OUTCOME OF THE PATIENT IS NOT RECOVERED. |  | 2021 | 9 | 2021-09
3005975929-2021-00400 | 2021-09-09 | 2017-06-28 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 58 YR | Male | (B)(6). IT WAS REPORTED THAT THE PLAINTIFF UNDERWENT MEDICALLY INDICATED RIGHT HIP THR PRIMARY SURGERY ON (B)(6) 2017, DUE TO GENU VALGUM CAUSED BY LEFT BHR PROTHESIS MALPOSITION. IT IS UNKNOWN WHAT SYSTEM WAS IMPLANTED AND THE SYSTEM MANUFACTURER. THE LEFT BHR PRIMARY SURGERY WAS PERFORMED ON (B)(6) 2007. ADDITIONALLY, PLAINTIFF HAD PRIMARY LEFT BHR REVISION SURGERY PERFORMED ON (B)(6) 2017 AND A BACK SURGERY PERFORMED ON (B)(6) 2016. THE PLAINTIFF'S OUTCOME IS UNKNOWN. |  | 2021 | 9 | 2021-09
1020279-2021-06920 | 2021-09-16 | 2021-07-20 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 58 YR | Male | IT WAS REPORTED THAT, AFTER A HIP REPLACEMENT SURGERY WAS PERFORMED ON 2009, THE PATIENT EXPERIENCED METALLOSIS, DISCOMFORT IN HIP JOINT, PAIN, HEADACHES, AND ANXIETY. THIS ADVERSE EVENT WILL BE SOLVED WITH A REVISION SURGERY ON (B)(6) 2021. THE PATIENT IS REQUESTING COMPENSATION. |  | 2021 | 9 | 2021-09
3005975929-2021-00414 | 2021-09-20 | 2021-08-30 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP ACET (STANDARD) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 55 YR |  | IT WAS REPORTED THAT, AFTER A RIGHT BHR CONSTRUCT HAD BEEN IMPLANTED IN 2010, THE PATIENT EXPERIENCED METALLOSIS. THIS ADVERSE EVENT WAS TREATED WITH A REVISION SURGERY PERFORMED ON (B)(6) 2021, IN WHICH THE BHR HEAD AND CUP WERE EXPLANTED. PROPHYLACTIC ANTIBIOTICS WERE GIVEN PRE AND POST SURGERY. NO RECORDS OF EXPLANTED DEVICES ARE AVAILABLE. NO FURTHER CONTACT WITH THE PATIENT, FOR WHICH PATIENT'S CURRENT HEALTH STATUS IS UNKNOWN. |  | 2021 | 9 | 2021-09
3005975929-2021-00425 | 2021-09-23 | 2010-08-23 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 63 YR | Male | (B)(6). IT WAS REPORTED THAT, AFTER A PRIMARY R3-THA CONSTRUCT HAD BEEN IMPLANTED ON THE PLAINTIFF¿S LEFT HIP ON (B)(6) 2010, THE PLAINTIFF HAS BEEN EXPERIENCING HIGH LEVELS OF COBALT AND CHROMIUM IN BLOOD, INTENSE PAIN AND AMBULATORY DIFFICULTIES. A REVISION SURGERY HAS BEEN MEDICALLY INDICATED, BUT STILL NOT CONDUCTED DUE TO A CURRENT WAITING LIST OF 12-18 MONTHS. |  | 2021 | 9 | 2021-09
3005975929-2021-00439 | 2021-09-27 | 2021-02-05 | D, e, a, t, h | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 53 YR | Male | ON THE LITERATURE ARTICLE NAMED "COBALT TOXICITY: A PREVENTABLE AND TREATABLE CAUSE FOR POSSIBLY LIFE THREATENING CARDIOMYOPATHY", IT WAS REPORTED THAT A PATIENT PRESENTED TO THE HOSPITAL WITH DECOMPENSATED HEART FAILURE, NEW-ONSET PAROXYSMAL ATRIAL TACHYCARDIA AND INCREASING LEFT HIP PAIN ON 2019. THE PATIENT HAD CONCURRENT DECOMPENSATED CARDIOMYOPATHY REQUIRING DOPAMINE AND FUROSEMIDE INFUSIONS IN THE PAST. DESPITE SUCCESSFUL CHELATING THERAPY AND HEART FAILURE TREATMENT, THE PATIENT PASSED AWAY SECONDARY TO COBALT TOXICITY INDUCED CARDIOMYOPATHY (CTCM). EVEN THOUGH, THE PATIENT´S SYMPTOMATIC HIP WAS THE LEFT ONE, HE ALSO UNDERWENT A RIGHT PRIMARY BHR SURGERY ON 2006. IT IS NOT POSSIBLE TO DISCARD THE INFLUENCE OF THE RIGHT SYSTEM IN THE COBALT TOXICITY RELATED EVENT. |  | 2021 | 9 | 2021-09
3005975929-2021-00440 | 2021-09-27 | 2021-02-05 | D, e, a, t, h | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 53 YR | Male | ON THE LITERATURE ARTICLE NAMED "COBALT TOXICITY: A PREVENTABLE AND TREATABLE CAUSE FOR POSSIBLY LIFE THREATENING CARDIOMYOPATHY", IT WAS REPORTED THAT A PATIENT PRESENTED TO THE HOSPITAL WITH DECOMPENSATED HEART FAILURE, NEW-ONSET PAROXYSMAL ATRIAL TACHYCARDIA AND INCREASING LEFT HIP PAIN ON 2019. THE PATIENT HAD CONCURRENT DECOMPENSATED CARDIOMYOPATHY REQUIRING DOPAMINE AND FUROSEMIDE INFUSIONS IN THE PAST. DESPITE SUCCESSFUL CHELATING THERAPY AND HEART FAILURE TREATMENT, THE PATIENT PASSED AWAY SECONDARY TO COBALT TOXICITY INDUCED CARDIOMYOPATHY (CTCM). EVEN THOUGH, THE PATIENT´S SYMPTOMATIC HIP WAS THE LEFT ONE, HE ALSO UNDERWENT A RIGHT PRIMARY BHR SURGERY ON 2006. IT IS NOT POSSIBLE TO DISCARD THE INFLUENCE OF THE RIGHT SYSTEM IN THE COBALT TOXICITY RELATED EVENT. |  | 2021 | 9 | 2021-09
3005975929-2021-00438 | 2021-09-27 | 2017-08-24 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP MODULAR HEAD (BHMH) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL |  |  |  | (B)(6) BILATERAL PATIENT. IT WAS REPORTED THAT, AFTER A PRIMARY R3-THA CONSTRUCT HAD BEEN IMPLANTED ON THE PLAINTIFF¿S LEFT HIP ON (B)(6) 2009, THE PLAINTIFF EXPERIENCED AN ALVAL REACTION CONSISTENT WITH METALLOSIS, ELEVATED METAL ION LEVELS IN BLOOD AND THE PRESENCE OF A SOLID MASS WITH BROWNISH COLOR, GREATER IN SIZE THAN THE ONE FOUND DURING THE REVISION SURGERY PERFORMED ON THE RIGHT HIP COVERED UNDER (B)(4). A REVISION SURGERY WAS PERFORMED ON (B)(6) 2017 TO TREAT THIS ADVERSE EVENT; THE UNKNOWN 44 MM DIAMETER BHR MODULAR HEAD AND UNKNOWN R3 ACETABULAR LINER CO-CR WERE EXPLANTED. SURGERY WAS COMPLETED WITHOUT ANY COMPLICATIONS. |  | 2021 | 9 | 2021-09
1020279-2021-07838 | 2021-11-04 | 2021-10-13 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD. | UNKNOWN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT A REVISION SURGERY HAD TO BE CONDUCTED TO EXCHANGE AN R3 42MM ID INTL COCR LINER 54MM AND AN UNKNOWN FEMORAL HEAD IN A BHR TO A POLY LINER AND OXINIUM HEAD DUE TO UNSPECIFIED REASONS. CURRENT PATIENT HEALTH STATUS IS UNKNOWN. |  | 2021 | 11 | 2021-11
3005975929-2021-00500 | 2021-11-08 | 2017-01-04 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 65 YR | Female | ¿US LEGAL - BILATERAL PATIENT. IT WAS REPORTED THAT AFTER A BILATERAL BHR HIP SURGERY, THE PATIENT EXPERIMENTED HIGH METAL ION LEVELS AND KIDNEY DISEASE ASSOCIATED WITH THOSE METAL IONS. THE PATIENT UNDERWENT A REVISION SURGERY OF THE RIGHT HIP TO ADDRESS THE CONDITION. EVEN THOUGH, NO REVISION SURGERY OR MEDICAL INTERVENTION HAS BEEN SCHEDULED TO INTERVENE THE LEFT HIP SYSTEM, ITS INFLUENCE CANNOT BE DISCARDED AS A POSSIBLE CAUSE OF THE ADVERSE EVENT REPORTED.¿ |  | 2021 | 11 | 2021-11
3005975929-2021-00498 | 2021-11-08 | 2017-01-04 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 65 YR | Female | (B)(6). BILATERAL PATIENT. IT WAS REPORTED THAT AFTER A BILATERAL BHR HIP SURGERY, THE PATIENT EXPERIMENTED HIGH METAL ION LEVELS AND KIDNEY DISEASE ASSOCIATED WITH THOSE METAL IONS. THE PATIENT UNDERWENT A REVISION SURGERY OF THE RIGHT HIP TO ADDRESS THE CONDITION. EVEN THOUGH, NO REVISION SURGERY OR MEDICAL INTERVENTION HAS BEEN SCHEDULED TO INTERVENE THE LEFT HIP SYSTEM, ITS INFLUENCE CANNOT BE DISCARDED AS A POSSIBLE CAUSE OF THE ADVERSE EVENT REPORTED.¿ |  | 2021 | 11 | 2021-11
3005975929-2021-00516 | 2021-11-09 | 2019-05-29 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT AFTER A BHR THA CONSTRUCT WAS IMPLANTED ON (B)(6) 2014, PLAINTIFF EXPERIENCED SEVERE PAIN, FAILED TOTAL LEFT HIP ARTHROPLASTY, STIFFNESS, LOSS OF MOBILITY, ADVERSE TISSUE REACTION, PSEUDOTUMOR, AND METALLOSIS. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2019, TO TREAT THESE ADVERSE EVENTS. PLAINTIFF OUTCOME IS UNKNOWN. |  | 2021 | 11 | 2021-11
3005975929-2021-00508 | 2021-11-09 | 2021-08-03 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 71 YR | Female | IT WAS REPORTED THAT, AFTER A RIGHT BHR ON (B)(6) 2011 DUE TO OSTEOARTHRITIS. THE PLAINTIFF HAS EXPERIENCED PAIN AND METALLOSIS. PLAINTIFF UNDERWENT REVISION SURGERY ON (B)(6), 2021 WHERE A FRACTURED OF THE PROXIMAL FEMUR BECAUSE OF PREVIOUS DAMAGE OF HIS FEMORAL NECK FROM METALLOSIS OCCURRED. DURING SURGERY, METALLOSIS WAS CONFIRMED WITH STAINING OF THE TISSUES AROUND THE JOINT COMPATIBLE WITH A PSEUDOTUMOR. THE PATIENT'S CONDITION IS UNKNOWN. |  | 2021 | 11 | 2021-11
3005975929-2021-00509 | 2021-11-09 | 2021-08-03 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 71 YR | Female | IT WAS REPORTED THAT, AFTER A RIGHT BHR ON (B)(6), 2011 DUE TO OSTEOARTHRITIS. THE PLAINTIFF HAS EXPERIENCED PAIN AND METALLOSIS. PLAINTIFF UNDERWENT REVISION SURGERY ON (B)(6), 2021 WHERE A FRACTURED OF THE PROXIMAL FEMUR BECAUSE OF PREVIOUS DAMAGE OF HIS FEMORAL NECK FROM METALLOSIS OCCURRED. DURING SURGERY, METALLOSIS WAS CONFIRMED WITH STAINING OF THE TISSUES AROUND THE JOINT COMPATIBLE WITH A PSEUDOTUMOR. THE PATIENT'S CONDITION IS UNKNOWN. |  | 2021 | 11 | 2021-11
3005975929-2021-00514 | 2021-11-09 | 2021-01-06 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Male | IT WAS REPORTED THAT THE PLAINTIFF UNDERWENT A MEDICALLY INDICATED REVISION SURGERY OF THE LEFT HIP IMPLANTS ON (B)(6) 2021 DUE TO METALLOSIS, ELEVATED METAL ION LEVELS, PAIN, REDUCED MOBILITY, OSTEOLYSIS AND CORTICAL DISRUPTION INVOLVING THE FEMORAL NECK. PRIMARY BHR SURGERY WAS PERFORMED ON (B)(6) 2015. IT WAS NOT REPORTED WHICH DEVICE OR DEVICES WERE EXPLANTED. THE HEALTH STATUS OF THE PATIENT IS UNKNOWN. |  | 2021 | 11 | 2021-11
3005975929-2021-00515 | 2021-11-09 | 2019-05-29 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT AFTER A BHR THA CONSTRUCT WAS IMPLANTED ON (B)(6) 2014, PLAINTIFF EXPERIENCED SEVERE PAIN, FAILED TOTAL LEFT HIP ARTHROPLASTY, STIFFNESS, LOSS OF MOBILITY, ADVERSE TISSUE REACTION, PSEUDOTUMOR, AND METALLOSIS. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2019, TO TREAT THESE ADVERSE EVENTS. PLAINTIFF OUTCOME IS UNKNOWN. |  | 2021 | 11 | 2021-11
3005975929-2021-00513 | 2021-11-09 | 2021-01-06 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Male | IT WAS REPORTED THAT THE PLAINTIFF UNDERWENT A MEDICALLY INDICATED REVISION SURGERY OF THE LEFT HIP IMPLANTS ON (B)(6) 2021 DUE TO METALLOSIS, ELEVATED METAL ION LEVELS, PAIN, REDUCED MOBILITY, OSTEOLYSIS AND CORTICAL DISRUPTION INVOLVING THE FEMORAL NECK. PRIMARY BHR SURGERY WAS PERFORMED ON (B)(6) 2015. IT WAS NOT REPORTED WHICH DEVICE OR DEVICES WERE EXPLANTED. THE HEALTH STATUS OF THE PATIENT IS UNKNOWN. |  | 2021 | 11 | 2021-11
3005975929-2021-00522 | 2021-11-10 | 2021-09-02 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Male | US LEGAL. IT WAS REPORTED THAT, AFTER A BHR-THA ON (B)(6) 2010. PLAINTIFF HAS EXPERIENCED RIGHT HIP METALLOSIS. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2021 TO TREAT THE ADVERSE EVENTS. THE OUTCOME OF THE PATIENT IS UNKNOWN. |  | 2021 | 11 | 2021-11
3005975929-2021-00519 | 2021-11-10 | 2021-06-16 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 52 YR | Male | US LEGAL. IT WAS REPORTED THAT AFTER A LEFT R3-THA CONSTRUCT WAS IMPLANTED ON (B)(6) 2011, PLAINTIFF EXPERIENCED METALLOSIS, AND ELEVATED COBALT AND CHROMIUM LEVELS. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2021 TO TREAT THESE ADVERSE EVENTS. PLAINTIFF OUTCOME IS UNKNOWN. |  | 2021 | 11 | 2021-11
3005975929-2021-00518 | 2021-11-10 | 2021-09-02 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Male | IT WAS REPORTED THAT, AFTER A BHR-THA ON (B)(6) 2010. PLAINTIFF HAS EXPERIENCED RIGHT HIP METALLOSIS. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2021 TO TREAT THE ADVERSE EVENTS. THE OUTCOME OF THE PATIENT IS UNKNOWN. |  | 2021 | 11 | 2021-11
3005975929-2021-00532 | 2021-11-12 | 1901-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 65 YR | Female | IT WAS REPORTED THAT, AFTER A PRIMARY UNSPECIFIED BHR CONSTRUCT HAD BEEN IMPLANTED ON THE PATIENT¿S RIGHT HIP ON (B)(6) 2010, THE PATIENT EXPERIENCED ELEVATED CHROMIUM LEVELS IN BLOOD (21.58 UG/L) AS REVEALED BY A BLOOD TEST RESULT PERFORMED ON (B)(6) 2019. PER PATIENT¿S CORRESPONDENCE, SHE HAS EXPERIENCED WEIGHT LOSS, AN UNSPECIFIED INFECTION IN THE BODY, CLOGGED BLOOD VESSELS, ADVANCED BONE DEGRADATION AND THE APPEARANCE OF A SLIGHT NOISE IN THE IMPLANT AREA. IT IS UNKNOWN IF A REVISION SURGERY HAS ALREADY BEEN CONDUCTED, INDICATED AND/OR SCHEDULED TO TREAT THIS ADVERSE EVENT. THE PATIENT DOES NOT COMPLAINT OF PAIN IN THE AREA OF THE PROSTHESIS. |  | 2021 | 11 | 2021-11
3005975929-2021-00531 | 2021-11-12 | 2019-12-20 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 65 YR | Female | IT WAS REPORTED THAT, AFTER A PRIMARY UNSPECIFIED BHR CONSTRUCT HAD BEEN IMPLANTED ON THE PATIENT¿S LEFT HIP ON (B)(6) 2010, THE PATIENT EXPERIENCED ELEVATED CHROMIUM LEVELS IN BLOOD (21.58 UG/L) AS REVEALED BY A BLOOD TEST RESULT PERFORMED ON (B)(6) 2019. PER PATIENT¿S CORRESPONDENCE, SHE HAS EXPERIENCED WEIGHT LOSS, AN UNSPECIFIED INFECTION IN THE BODY, CLOGGED BLOOD VESSELS, ADVANCED BONE DEGRADATION AND THE APPEARANCE OF A SLIGHT NOISE IN THE IMPLANT AREA. IT IS UNKNOWN IF A REVISION SURGERY HAS ALREADY BEEN CONDUCTED, INDICATED AND/OR SCHEDULED TO TREAT THIS ADVERSE EVENT. THE PATIENT DOES NOT COMPLAINT OF PAIN IN THE AREA OF THE PROSTHESIS. |  | 2021 | 11 | 2021-11
3005975929-2021-00530 | 2021-11-12 | 1901-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 65 YR | Female | IT WAS REPORTED THAT, AFTER A PRIMARY UNSPECIFIED BHR CONSTRUCT HAD BEEN IMPLANTED ON THE PATIENT¿S RIGHT HIP ON (B)(6) 2010, THE PATIENT EXPERIENCED ELEVATED CHROMIUM LEVELS IN BLOOD (21.58 UG/L) AS REVEALED BY A BLOOD TEST RESULT PERFORMED ON (B)(6) 2019. PER PATIENT¿S CORRESPONDENCE, SHE HAS EXPERIENCED WEIGHT LOSS, AN UNSPECIFIED INFECTION IN THE BODY, CLOGGED BLOOD VESSELS, ADVANCED BONE DEGRADATION AND THE APPEARANCE OF A SLIGHT NOISE IN THE IMPLANT AREA. IT IS UNKNOWN IF A REVISION SURGERY HAS ALREADY BEEN CONDUCTED, INDICATED AND/OR SCHEDULED TO TREAT THIS ADVERSE EVENT. THE PATIENT DOES NOT COMPLAINT OF PAIN IN THE AREA OF THE PROSTHESIS. |  | 2021 | 11 | 2021-11
3005975929-2021-00533 | 2021-11-12 | 2019-12-20 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 65 YR | Female | IT WAS REPORTED THAT, AFTER A PRIMARY UNSPECIFIED BHR CONSTRUCT HAD BEEN IMPLANTED ON THE PATIENT¿S LEFT HIP ON (B)(6) 2010, THE PATIENT EXPERIENCED ELEVATED CHROMIUM LEVELS IN BLOOD (21.58 UG/L) AS REVEALED BY A BLOOD TEST RESULT PERFORMED ON (B)(6) 2019. PER PATIENT¿S CORRESPONDENCE, SHE HAS EXPERIENCED WEIGHT LOSS, AN UNSPECIFIED INFECTION IN THE BODY, CLOGGED BLOOD VESSELS, ADVANCED BONE DEGRADATION AND THE APPEARANCE OF A SLIGHT NOISE IN THE IMPLANT AREA. IT IS UNKNOWN IF A REVISION SURGERY HAS ALREADY BEEN CONDUCTED, INDICATED AND/OR SCHEDULED TO TREAT THIS ADVERSE EVENT. THE PATIENT DOES NOT COMPLAINT OF PAIN IN THE AREA OF THE PROSTHESIS. |  | 2021 | 11 | 2021-11
3005975929-2021-00537 | 2021-11-22 | 2016-02-03 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Male | IT WAS REPORTED THAT AFTER A LEFT BHR CONSTRUCT WAS IMPLANTED ON (B)(6) 2012, PLAINTIFF EXPERIENCED DISABLING PAIN. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2016 TO TREAT THIS ADVERSE EVENT. DURING SURGERY, THE FEMORAL BHR FEMORAL HEAD WAS EXPLANTED, AND THE CUP WAS FOUND TO BE STABLE AND REMAINED IMPLANTED. THE HIP WAS REVISED TO A THA USING STRYKER COMPONENTS. PLAINTIFF OUTCOME IS UNKNOWN. |  | 2021 | 11 | 2021-11
3005975929-2021-00538 | 2021-11-22 | 2015-07-15 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Male | IT WAS REPORTED THAT AFTER A RIGHT BHR CONSTRUCT WAS IMPLANTED ON (B)(6) 2011, PLAINTIFF EXPERIENCED DISABLING PAIN, INABILITY TO AMBULATE, AND FRACTURE OF THE FEMORAL COMPONENT WITH MIGRATION. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2015 TO TREAT THESE ADVERSE EVENTS. DURING SURGERY, THE FEMORAL NECK WAS FOUND TO HAVE A SUBCAPITAL FRACTURE. THE BHR FEMORAL HEAD WAS EXPLANTED, AND THE CUP WAS FOUND TO BE STABLE AND REMAINED IMPLANTED. THE HIP WAS REVISED TO A THA USING STRYKER COMPONENTS. PLAINTIFF OUTCOME IS UNKNOWN. |  | 2021 | 11 | 2021-11
3005975929-2021-00541 | 2021-11-25 | 2021-07-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 55 YR | Male | IT WAS REPORTED THAT 11 YEARS AFTER A HIP REPLACEMENT WAS PERFORMED, THE PATIENT UNDERWENT A REVISION SURGERY BECAUSE OF BONE DESTRUCTION AND OSTEOLYSIS, RELATED TO METALLOSIS. DURING SURGERY, LARGE AMOUNTS OF DARK FLUID CAME OUT OF THE JOINT. THE CURRENT STATE OF HEALTH OF THE PATIENT IS UNKNOWN. |  | 2021 | 11 | 2021-11
3005975929-2021-00547 | 2021-11-30 | 2018-05-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 60 YR | Unknown | IT WAS REPORTED THAT AFTER A PRIMARY LEFT BHR HIP SURGERY WAS PERFORMED IN (B)(6) 2007, THE PLAINTIFF EXPERIENCED PAIN AND DISCOMFORT IN THE REGION NEARBY THE PROSTHESIS. THE PATIENT ALSO PRESENTED SUSTAINED ABNORMALLY ELEVATED COBALT AND CHROMIUM LEVELS IN BLOOD. THE PATIENT UNDERWENT A REVISION SURGERY IN (B)(6) 2018 TO ADDRESS THE ISSUE. THE CURRENT STATE OF HEALTH OF THE PATIENT IS UNKNOWN. |  | 2021 | 11 | 2021-11
3005975929-2021-00546 | 2021-11-30 | 2018-05-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 60 YR | Unknown | IT WAS REPORTED THAT AFTER A PRIMARY LEFT BHR HIP SURGERY WAS PERFORMED IN (B)(6) 2007, THE PLAINTIFF EXPERIENCED PAIN AND DISCOMFORT IN THE REGION NEARBY THE PROSTHESIS. THE PATIENT ALSO PRESENTED SUSTAINED ABNORMALLY ELEVATED COBALT AND CHROMIUM LEVELS IN BLOOD. THE PATIENT UNDERWENT A REVISION SURGERY IN (B)(6) 2018 TO ADDRESS THE ISSUE. THE CURRENT STATE OF HEALTH OF THE PATIENT IS UNKNOWN. |  | 2021 | 11 | 2021-11
3005975929-2021-00556 | 2021-12-07 | 2020-03-04 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 70 YR | Female | US LEGAL, IT WAS REPORTED THAT THE PLAINTIFF UNDERWENT A RIGHT HIP REVISION SURGERY IN (B)(6) 2020 DUE TO ADVANCED METALLOSIS (PREDOMINANTLY IN THE BURSA), METAL DEBRIS AND OSTEOLYSIS OF THE GREATER TROCHANTER. DURING THE REVISION, A FAIR AMOUNT OF BURSA WAS REMOVED. THE PRIMARY RIGHT BHR SURGERY WAS PERFORMED IN (B)(6) 2007. THE CURRENT STATE OF HEALTH OF THE PATIENT IN UNKNOWN. |  | 2021 | 12 | 2021-12
3005975929-2021-00557 | 2021-12-07 | 2010-01-05 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 46 YR | Male | IT WAS REPORTED THAT, AFTER A LEFT BIRMINGHAM RESURFACING ON (B)(6) 2006 DUE TO DEGENERATIVE JOINT DISEASE, PLAINTIFF HAS EXPERIENCE LOOSENING OF FEMORAL COMPONENT AND PAIN THAT WAS TREATED WITH MARCAINE AND LIDOCAINE INJECTIONS. AS THE PROBLEM WAS NOT SOLVED, A REVISION SURGERY WAS PERFORMED ON (B)(6) 2010 WHERE IT WAS CONVERTED TO BHR THA. PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 12 | 2021-12
3005975929-2021-00558 | 2021-12-07 | 2021-04-05 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 57 YR | Male | US LEGAL. IT WAS REPORTED THAT, AFTER A REVISION SURGERY PERFORMED ON (B)(6) 2010 WHERE AN INITIAL BIRMINGHAM RESURFACING WHERE CONVERTED TO BHR THA PLAINTIFF HAS EXPERIENCE A FEMORAL COMPONENT FAILURE. A SECOND REVISION SURGERY WAS PERFORMED ON (B)(6) 2021 WHERE THE CUP, SLEEVE AND HEAD WERE REMOVED AND A OXINIUM FEMORAL HEAD WAS IMPLANTED INSTEAD. PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 12 | 2021-12
3005975929-2021-00559 | 2021-12-07 | 2021-04-05 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 57 YR | Male | IT WAS REPORTED THAT, AFTER A REVISION SURGERY PERFORMED ON (B)(6) 2010 WHERE AN INITIAL LEFT BIRMINGHAM RESURFACING WAS CONVERTED TO A LEFT BHR THA, THE PLAINTIFF HAS EXPERIENCED PAIN, ASEPTIC LOOSENING, PSEUDOTUMOR AND ELEVATED CHROMIUM AND COBALT LEVELS. A SECOND REVISION SURGERY WAS PERFORMED ON (B)(6) 2021 WHERE THE CUP, SLEEVE AND HEAD WERE REMOVED AND A OXINIUM FEMORAL HEAD WAS IMPLANTED INSTEAD. PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 12 | 2021-12
3005975929-2021-00560 | 2021-12-07 | 2021-04-05 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) SLEEVE | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 57 YR | Male | *US LEGAL * IT WAS REPORTED THAT, AFTER A REVISION SURGERY PERFORMED ON (B)(6) 2010 WHERE AN INITIAL BIRMINGHAM RESURFACING WHERE CONVERTED TO BHR THA PLAINTIFF HAS EXPERIENCE A FEMORAL COMPONENT FAILURE. A SECOND REVISION SURGERY WAS PERFORMED ON (B)(6) 2021 WHERE THE CUP, SLEEVE AND HEAD WERE REMOVED AND A OXINIUM FEMORAL HEAD WAS IMPLANTED INSTEAD. PATIENT OUTCOME IS UNKNOWN. |  | 2021 | 12 | 2021-12
3005975929-2021-00567 | 2021-12-13 | 1901-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 55 YR | Male | IT WAS REPORTED THAT, AFTER A RIGHT HIP REPLACEMENT PERFORMED ON AN UNSPECIFIED DATE, THE PATIENT UNDERWENT A REVISION SURGERY OF THE RIGHT HIP DUE TO COBALT POISONING. THIS SURGERY WAS PERFORMED 8 WEEKS AFTER THE LEFT HIP WAS REVISED (COVERED UNDER (B)(4)). THE PATIENT STATED THAT NOW HE FEELS BETTER. |  | 2021 | 12 | 2021-12
3005975929-2021-00571 | 2021-12-15 | 2020-09-20 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 75 YR | Male | ''US LEGAL'': IT WAS REPORTED THAT, AFTER LEFT BHR THA ON (B)(6) 2008, THE PLAINTIFF HAS SUFFERED OF PAIN, LIMITED MOBILITY, ELEVATED COBALT CHROMIUM LEVELS, AND DEBRIS-FILLED PSEUDOTUMOR MEASURING 6 CENTIMETERS. PLAINTIFF UNDERWENT A REVISION SURGERY ON (B)(6) 2020. THE CURRENT HEALTH STATUS OF THE PATIENT IS UNKNOWN. |  | 2021 | 12 | 2021-12
3005975929-2021-00570 | 2021-12-15 | 2020-09-20 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 75 YR | Male | IT WAS REPORTED THAT, AFTER LEFT BHR THA ON (B)(6) 2008, THE PLAINTIFF HAS SUFFERED OF PAIN, LIMITED MOBILITY, ELEVATED COBALT CHROMIUM LEVELS, AND DEBRIS-FILLED PSEUDOTUMOR MEASURING 6 CENTIMETERS. PLAINTIFF UNDERWENT A REVISION SURGERY ON (B)(6) 2020. THE CURRENT HEALTH STATUS OF THE PATIENT IS UNKNOWN. |  | 2021 | 12 | 2021-12
3005975929-2021-00579 | 2021-12-27 | 2019-12-23 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Male | IT WAS REPORTED THAT AFTER A RIGHT BHR CONSTRUCT WAS IMPLANTED ON (B)(6) 2011, THE PLAINTIFF EXPERIENCED PAIN, ELEVATED METALLOSIS, PSEUDOTUMOR AND LOSS OF FUNCTION. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2019 TO TREAT THE ADVERSE EVENTS. PLAINTIFF OUTCOME IS UNKNOWN. |  | 2021 | 12 | 2021-12
3005975929-2022-00015 | 2022-01-12 | 2020-04-21 | D, e, a, t, h | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 71 YR | Male | US LEGAL - BILATERAL PATIENT. IT WAS REPORTED THAT AFTER THE PATIENT UNDERWENT A REVISION SURGERY OF THE LEFT HIP ON (B)(6) 2020. THE PATIENT UNDERWENT ANOTHER REVISION SURGERY OF THE RIGHT HIP ON (B)(6) 2020 DUE TO ELEVATED COBALT AND CHROMIUM LEVELS, PAIN, AND METAL DEBRIS. THE PRIMARY RIGHT BHR SURGERY WAS PERFORMED ON (B)(6) 2008. THE PATIENT DIED ON (B)(6) 2021, DUE TO A PULMONARY THROMBOEMBOLISM. HE ALSO SUFFERED CARDIOMYOPATHY AND OTHER SYSTEMIC CONDITIONS. |  | 2022 | 1 | 2022-01
3005975929-2022-00016 | 2022-01-12 | 2020-04-21 | D, e, a, t, h | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 71 YR | Male | US LEGAL - BILATERAL PATIENT. IT WAS REPORTED THAT AFTER THE PATIENT UNDERWENT A REVISION SURGERY OF THE LEFT HIP ON (B)(6) 2020. THE PATIENT UNDERWENT ANOTHER REVISION SURGERY OF THE RIGHT HIP ON (B)(6) 2020 DUE TO ELEVATED COBALT AND CHROMIUM LEVELS, PAIN, AND METAL DEBRIS. THE PRIMARY RIGHT BHR SURGERY WAS PERFORMED ON (B)(6) 2008. THE PATIENT DIED ON (B)(6) 2021, DUE TO A PULMONARY THROMBOEMBOLISM. HE ALSO SUFFERED CARDIOMYOPATHY AND OTHER SYSTEMIC CONDITIONS. |  | 2022 | 1 | 2022-01
3005975929-2022-00022 | 2022-01-18 | 2021-11-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 64 YR | Male | ''US LEGAL'': IT WAS REPORTED THAT THE PLAINTIFF SUFFERED FROM AN UNSPECIFIED ADVERSE EVENT AND WAS SCHEDULED FOR A RIGHT HIP REVISION SURGERY IN (B)(6) 2021 TO ADDRESS THE ISSUE. TO THE DATE, THERE IS NO CONFIRMATION THAT THIS PROCEDURE TOOK PLACE. THE CURRENT STATE OF HEALTH OF THE PATIENT IS UNKNOWN. THE RIGHT HIP BHR PRIMARY SURGERY WAS PERFORMED IN (B)(6) 2011 DUE TO DEGENERATIVE JOINT DECEASE. |  | 2022 | 1 | 2022-01
3005975929-2022-00024 | 2022-01-25 | 2021-07-20 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 58 YR | Male | IT WAS REPORTED THAT, AFTER A RIGHT HIP REPLACEMENT SURGERY WAS PERFORMED ON 2009, THE PATIENT EXPERIENCED METALLOSIS, OSTEOLYSIS, DISCOMFORT IN HIP JOINT, PAIN, HEADACHES, AND ANXIETY. THIS ADVERSE EVENT WAS SOLVED WITH A REVISION SURGERY ON (B)(6) 2021. DURING THE PROCEDURE, THE FEMORAL COMPONENT WAS NOTICED TO BE TOO SMALL FOR THE SIZE OF THE CUP AND A SIGNIFICANT AMOUNT OF OSTEOLYSIS WAS PRESENT BEHIND THE ACETABULAR COMPONENT. THE STEM WAS STABLE AND WITH NO SIGNS OF OSTEOLYSIS. THE PATIENT TOLERATED THE PROCEDURE WELL AND WAS TRANSFERRED TO RECOVERY. |  | 2022 | 1 | 2022-01
3005975929-2022-00032 | 2022-02-01 | 2021-06-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 59 YR | Male | US LEGAL. IT WAS REPORTED THAT THE PATIENT UNDERWENT A REVISION SURGERY OF THE RIGHT HIP ON (B)(6) 2021 DUE TO HIP PAIN, ELEVATED METAL LEVELS IN BLOOD, METALLOSIS, TISSUE INFLAMMATION AND NECROTIC TISSUE. THE PRIMARY BHR THA SURGERY WAS PERFORMED ON (B)(6) 2010. THE PATIENT CURRENT STATUS IS UNKNOWN. |  | 2022 | 2 | 2022-02
3005975929-2022-00046 | 2022-02-21 | 2020-10-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP MODULAR HEAD (BHMH) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT THE PLAINTIFF UNDERWENT A MEDICALLY INDICATED REVISION SURGERY OF THE RIGHT HIP ON (B)(6) 2020 DUE TO SEVERE PAIN, LIMITED MOBILITY, AND ELEVATED COBALT AND CHROMIUM LEVELS. INTRAOPERATIVE FINDINGS INCLUDED NOTABLE METALLOSIS, ABUNDANT BLACK FLUID, NECROTIC AND BLACKENED CAPSULAR TISSUES, AND MULTIPLE FLUID POCKETS WITH LIQUID BLACK FLUID AND PSEUDOTUMOR. THE OUTCOME OF THE PATIENT IS UNKNOWN. |  | 2022 | 2 | 2022-02
3005975929-2022-00047 | 2022-02-21 | 2020-02-24 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT, ON (B)(6) 2020 THE PLAINTIFF UNDERWENT A MEDICALLY INDICATED REVISION SURGERY OF THE RIGHT HIP DUE TO SEVERE PAIN, LIMITED MOBILITY AND ELEVATED COBALT AND CHROMIUM LEVELS. INTRA-OPERATIVE FINDINGS INCLUDED FLUID, CORROSION, AND NECROTIC / STAINED TISSUE. DURING THE REVISION SURGERY, THE METALLIC CUP AND FEMORAL HEAD WERE EXPLANTED. THE PRIMARY BHR THA RIGHT HIP SURGERY WAS PERFORMED ON (B)(6) 2008. THE OUTCOME OF THE PATIENT IS UNKNOWN. |  | 2022 | 2 | 2022-02
3005975929-2022-00048 | 2022-02-21 | 2020-02-24 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT, ON (B)(6) 2020 THE PLAINTIFF UNDERWENT A MEDICALLY INDICATED REVISION SURGERY OF THE RIGHT HIP DUE TO SEVERE PAIN, LIMITED MOBILITY AND ELEVATED COBALT AND CHROMIUM LEVELS. INTRA-OPERATIVE FINDINGS INCLUDED FLUID, CORROSION, AND NECROTIC / STAINED TISSUE. DURING THE REVISION SURGERY, THE METALLIC CUP AND FEMORAL HEAD WERE EXPLANTED. THE PRIMARY BHR THA RIGHT HIP SURGERY WAS PERFORMED ON (B)(6) 2008. THE OUTCOME OF THE PATIENT IS UNKNOWN. |  | 2022 | 2 | 2022-02
3005975929-2022-00049 | 2022-02-24 | 2021-12-13 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | *US LEGAL* IT WAS REPORTED THAT, AFTER AN R3-THA IMPLANTATION IN THE LEFT HIP IN 2011, THE PLAINTIFF EXPERIENCED PAIN AND ELEVATED LEVELS OF COBALT AND CHROMIUM. PLAINTIFF UNDERWENT LEFT REVISION SURGERY ON (B)(6) 2021 WHERE A PSEUDOTUMOR WAS FOUND. THE PATIENT'S OUTCOME IS UNKNOWN. |  | 2022 | 2 | 2022-02
3005975929-2022-00062 | 2022-03-14 | 2019-09-17 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT, ON LITERATURE REVIEW "MIDTERM OUTCOMES OF TOTAL HIP ARTHROPLASTY WITH A MODULAR BIRMINGHAM HEAD", FROM 9 PATIENTS THAT INITIALLY HAD EITHER A BHR-THA OR AN R3-THA PROSTHESIS CONSTRUCT IMPLANTED, AN UNSPECIFIED QUANTITY OF THEM WERE REVISED FOR SUSPECTED SOFT TISSUE REACTION AFTER A MEDIAN OF 68 MONTHS POSTOPERATIVELY. OF THESE, 5 HAD SOFT TISSUE ABNORMALITIES NOTED ON ULTRASOUND OR MRI, 3 HAD UNREMARKABLE IMAGING AND ONE WAS REVISED ON CLINICAL GROUNDS WITHOUT MRI. MEAN SERUM CHROMIUM LEVEL WAS 5.7 G/L AND SERUM COBALT WAS 8.1 G/L, WHILE MEAN SERUM TITANIUM WAS 2.0 G/L. ADDITIONAL DETAILS ON THE CLINICAL OUTCOME OF EACH PATIENT ARE UNKNOWN AT THIS TIME. NO FURTHER INFORMATION IS AVAILABLE IN REGARDS OF THE IDENTITY OF THE EXPLANTED DEVICES DURING THE REVISION SURGERY. |  | 2022 | 3 | 2022-03
3005975929-2022-00064 | 2022-03-14 | 2019-09-17 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT, ON LITERATURE REVIEW "MIDTERM OUTCOMES OF TOTAL HIP ARTHROPLASTY WITH A MODULAR BIRMINGHAM HEAD", ONE (1) PATIENT WHO UNDERWENT PRIMARY SURGERY WITH A R3-THA CONSTRUCT REQUIRED A REVISION SURGERY DUE TO DUE TO R3 LINER DISSOCIATION WITH INTRAOPERATIVELY NOTED ABNORMAL SOFT TISSUE METAL REACTION. ADDITIONAL DETAILS ON THE CLINICAL OUTCOME OF THIS PATIENT ARE UNKNOWN AT THIS TIME. NO FURTHER INFORMATION IS AVAILABLE IN REGARDS OF THE IDENTITY OF THE EXPLANTED DEVICES DURING THE REVISION SURGERY. |  | 2022 | 3 | 2022-03
3005975929-2022-00061 | 2022-03-14 | 2019-09-17 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT, ON LITERATURE REVIEW "MIDTERM OUTCOMES OF TOTAL HIP ARTHROPLASTY WITH A MODULAR BIRMINGHAM HEAD", FROM 9 PATIENTS THAT INITIALLY HAD EITHER A BHR-THA OR AN R3-THA PROSTHESIS CONSTRUCT IMPLANTED, AN UNSPECIFIED QUANTITY OF THEM WERE REVISED FOR SUSPECTED SOFT TISSUE REACTION AFTER A MEDIAN OF 68 MONTHS POSTOPERATIVELY. OF THESE, 5 HAD SOFT TISSUE ABNORMALITIES NOTED ON ULTRASOUND OR MRI, 3 HAD UNREMARKABLE IMAGING AND ONE WAS REVISED ON CLINICAL GROUNDS WITHOUT MRI. MEAN SERUM CHROMIUM LEVEL WAS 5.7 G/L AND SERUM COBALT WAS 8.1 G/L, WHILE MEAN SERUM TITANIUM WAS 2.0 G/L. ADDITIONAL DETAILS ON THE CLINICAL OUTCOME OF EACH PATIENT ARE UNKNOWN AT THIS TIME. NO FURTHER INFORMATION IS AVAILABLE IN REGARDS OF THE IDENTITY OF THE EXPLANTED DEVICES DURING THE REVISION SURGERY. |  | 2022 | 3 | 2022-03
3005975929-2022-00060 | 2022-03-14 | 2020-10-17 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) SLEEVED MODULAR | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "THE EARLY TO MEDIUM-TERM RESULTS OF A HEMISPHERICAL, POROUS COATED ACETABULAR SHELL WITH MULTIPLE DIFFERENT BEARING COMBINATIONS ARE EXCELLENT WITH THE EXCEPTION OF METAL-ON-METAL", (1) PATIENT HAD PAIN AND METALLOSIS AFTER A THA WHILE USING R3 ACETABULAR SHELL AND HIGH OFFSET POROUS-COATED SYNERGY STEM. THE EVENTS WERE TREATED BY A REVISION SURGERY 6 YEARS POST-OPERATIVELY. PATIENTS¿ OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 3 | 2022-03
3005975929-2022-00058 | 2022-03-14 | 2020-10-17 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) SLEEVED MODULAR | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "THE EARLY TO MEDIUM-TERM RESULTS OF A HEMISPHERICAL, POROUS COATED ACETABULAR SHELL WITH MULTIPLE DIFFERENT BEARING COMBINATIONS ARE EXCELLENT WITH THE EXCEPTION OF METAL-ON-METAL," TWO (2) PATIENTS HAD A TWO-STAGE REVISION SURGERY DUE TO A DEEP INFECTION AFTER A THA WHILE USING R3 ACETABULAR SHELL AND HIGH OFFSET POROUS-COATED SYNERGY STEM. THE EVENTS WERE TREATED BY A REVISION SURGERY. PATIENTS¿ OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 3 | 2022-03
3005975929-2022-00056 | 2022-03-14 | 2020-10-17 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) SLEEVED MODULAR | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "THE EARLY TO MEDIUM-TERM RESULTS OF A HEMISPHERICAL, POROUS COATED ACETABULAR SHELL WITH MULTIPLE DIFFERENT BEARING COMBINATIONS ARE EXCELLENT WITH THE EXCEPTION OF METAL-ON-METAL", TWO (2) PATIENTS HAD AN ADVERSE REACTION TO METAL DEBRIS (ARMD) WITH MARKED METALOSIS, CAUSING INCREASE IN COBALT-CHROMIUM ABOVE NORMAL LIMIT AFTER A THA WHILE USING R3 ACETABULAR SHELL AND HIGH OFFSET POROUS-COATED SYNERGY STEM. THE EVENTS WERE TREATED BY A REVISION SURGERY WHERE HEAD AND LINER WERE EXCHANGE. PATIENTS' OUTCOME ARE UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 3 | 2022-03
3005975929-2022-00066 | 2022-03-15 | 2019-09-17 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT, ON LITERATURE REVIEW "MIDTERM OUTCOMES OF TOTAL HIP ARTHROPLASTY WITH A MODULAR BIRMINGHAM HEAD", FROM 9 PATIENTS THAT INITIALLY HAD EITHER A BHR-THA (COVERED UNDER (B)(4)) OR AN R3-THA PROSTHESIS CONSTRUCT IMPLANTED, AN UNSPECIFIED QUANTITY OF THEM WERE REVISED FOR SUSPECTED SOFT TISSUE REACTION AFTER A MEDIAN OF 68 MONTHS POSTOPERATIVELY. OF THESE, 5 HAD SOFT TISSUE ABNORMALITIES NOTED ON ULTRASOUND OR MRI, 3 HAD UNREMARKABLE IMAGING AND ONE WAS REVISED ON CLINICAL GROUNDS WITHOUT MRI. MEAN SERUM CHROMIUM LEVEL WAS 5.7 G/L AND SERUM COBALT WAS 8.1 G/L, WHILE MEAN SERUM TITANIUM WAS 2.0 G/L. ADDITIONAL DETAILS ON THE CLINICAL OUTCOME OF EACH PATIENT ARE UNKNOWN AT THIS TIME. NO FURTHER INFORMATION IS AVAILABLE IN REGARDS OF THE IDENTITY OF THE EXPLANTED DEVICES DURING THE REVISION SURGERY. |  | 2022 | 3 | 2022-03
3005975929-2022-00069 | 2022-03-17 | 2022-02-25 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 73 YR | Male | IT WAS REPORTED THAT, AFTER A BHR RESURFACING FEMORAL HEAD 58MM WAS IMPLANTED IN (B)(6) 2010, A FLUID COLLECTION DURING A COMPUTED TOMOGRAPHY (CT) AROUND THE JOINT PRESENTED THE FOLLOWING METAL CONCENTRATIONS: COBALT 3.4UG/L AND CHROMIUM 0.9UG/L. THIS ADVERSE EVENT WAS TREATED WITH A REVISION SURGERY ON (B)(6) 2022. PATIENT'S CURRENT HEATH STATUS IS UNKNOWN. |  | 2022 | 3 | 2022-03
3005975929-2022-00090 | 2022-03-31 | 2022-03-09 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 63 YR | Male | IT WAS REPORTED THAT AFTER A BHR SURGERY ON (B)(6) 2007, THE PATIENT PRESENTED HIGH CO AND CR LEVELS. THE ADVERSE EVENT WAS TREATED VIA REVISION SURGERY ON (B)(6) 2022. PATIENT CURRENT STATUS IS UNKNOWN. |  | 2022 | 3 | 2022-03
3005975929-2022-00092 | 2022-04-01 | 2021-07-26 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Male | US LEGAL MDL: IT WAS REPORTED THAT, AFTER BHR TOTAL HIP ARTHROPLASTY WAS PERFORMED ON (B)(6) 2010, THE PATIENT EXPERIENCED PAIN, OSTEOLYSIS AND ELEVATED CHROMIUM AND COBALT LEVELS IN BLOOD. THIS ADVERSE EVENT WAS TREATED WITH A REVISION SURGERY ON (B)(6) 2021. DURING THE REVISION SURGERY, METALLIC DISCOLORATION OF THE SYNOVIUM, RETROACETABULAR OSTEOLYSIS AND METALLOSIS WERE FOUND. PATIENT'S CURRENT HEALTH STATUS IS UNKNOWN. |  | 2022 | 4 | 2022-04
3005975929-2022-00093 | 2022-04-01 | 2021-07-26 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Male | US LEGAL MDL, IT WAS REPORTED THAT, AFTER BHR TOTAL HIP ARTHROPLASTY WAS PERFORMED ON (B)(6) 2010, THE PATIENT EXPERIENCED PAIN, OSTEOLYSIS AND ELEVATED CHROMIUM AND COBALT LEVELS IN BLOOD. THIS ADVERSE EVENT WAS TREATED WITH A REVISION SURGERY ON (B)(6) 2021. DURING THE REVISION SURGERY, METALLIC DISCOLORATION OF THE SYNOVIUM, RETROACETABULAR OSTEOLYSIS AND METALLOSIS WERE FOUND. PATIENT'S CURRENT HEALTH STATUS IS UNKNOWN. |  | 2022 | 4 | 2022-04
3005975929-2022-00097 | 2022-04-01 | 2020-09-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 69 YR | Female | US LEGAL MDL. IT WAS REPORTED THAT, AFTER RIGHT TOTAL HIP SURFACE REPLACEMENT WAS PERFORMED ON AN UNSPECIFIED DATE, THE PATIENT EXPERIENCED PAIN, METALLOSIS AND A LOOSE FEMORAL HEAD. THIS ADVERSE EVENT WAS TREATED WITH A REVISION SURGERY ON (B)(6) 2020. DURING THE REVISION, SIGNIFICANT BONE LOSS WAS FOUND IN THE FEMORAL HEAD, AS WELL AS EXTENSIVE METALLOSIS IN THE SYNOVIUM. THE FEMORAL HEAD WAS EXPLANTED, AND THE HIP WAS REVISED TO A THA. PATIENT'S CURRENT HEALTH STATUS IS UNKNOWN. |  | 2022 | 4 | 2022-04
3005975929-2022-00096 | 2022-04-01 | 2020-09-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 69 YR | Female | US LEGAL MDL. IT WAS REPORTED THAT, AFTER RIGHT TOTAL HIP SURFACE REPLACEMENT WAS PERFORMED ON AN UNSPECIFIED DATE, THE PATIENT EXPERIENCED PAIN, METALLOSIS AND A LOOSE FEMORAL HEAD. THIS ADVERSE EVENT WAS TREATED WITH A REVISION SURGERY ON (B)(6) 2020. DURING THE REVISION, SIGNIFICANT BONE LOSS WAS FOUND IN THE FEMORAL HEAD, AS WELL AS EXTENSIVE METALLOSIS IN THE SYNOVIUM. THE FEMORAL HEAD WAS EXPLANTED, AND THE HIP WAS REVISED TO A THA. PATIENT'S CURRENT HEALTH STATUS IS UNKNOWN. |  | 2022 | 4 | 2022-04
3005975929-2022-00120 | 2022-04-05 | 2020-07-14 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 68 YR | Male | IT WAS REPORTED THAT, AFTER A BHR RESURFACING CONSTRUCT HAD BEEN IMPLANTED ON THE PATIENT¿S RIGHT HIP ON AN UNSPECIFIED DATE, THE PATIENT EXPERIENCED PERSISTENT PAIN AND LOOSENING OF THE FEMORAL COMPONENT. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2020 TO TREAT THIS ADVERSE EVENT. DURING THIS PROCEDURE, THE BHR FEMORAL HEAD WAS EXPLANTED AND REPLACED WITH A TOTAL HIP REPLACEMENT PROSTHESIS CONSTRUCT. INTRAOPERATIVELY, THE CUP WAS NOTICED TO BE WELL FIXED AND IN GOOD POSITION; THEREFORE, IT WAS NOT EXPLANTED. THE PATIENT¿S OUTCOME IS UNKNOWN. |  | 2022 | 4 | 2022-04
3005975929-2022-00122 | 2022-04-06 | 2021-03-03 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS,HIP,SEMICONSTRAINED,METALMETAL,RESURFACING |  |  | Unknown | IT WAS REPORTED THAT, AFTER A BHR THA WAS PERFORMED ON (B)(6) 2007 DUE TO SEVERE OSTEOARTHRITIS, PAIN, RESTRICTION OF MOTION AND INTERFERENCE WITH DAILY ACTIVITIES, THE PLAINTIFF EXPERIENCED FAILURE OF THA FOR METAL-ON-METAL COMPLICATION. THIS ADVERSE EVENT WAS TREATED WITH A REVISION SURGERY ON (B)(6) 2021, IN WHICH METAL DEBRIS, TRUNNION, AND CORROSION WERE NOTICED. PATIENT'S CURRENT HEALTH STATUS IS UNKNOWN. |  | 2022 | 4 | 2022-04
3005975929-2022-00128 | 2022-04-07 | 2021-03-15 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 47 YR | Male | US LEGAL. IT WAS REPORTED THAT THE PLAINTIFF UNDERWENT A LEFT HIP REVISION SURGERY DUE TO PAIN, STIFFNESS, AND A MECHANICAL FAILURE (LOOSENING) OF THE COMPONENTS. THE SURGEON ALSO NOTED THAT THE PATIENT HAD A LEG LENGTH DISCREPANCY. DURING THE REVISION SURGERY, BOTH BHR COMPONENTS WERE EXPLANTED AND REPLACED WITH A THA SYSTEM FROM A COMPETITOR. THE PATIENT TOLERATED THE PROCEDURE WELL. |  | 2022 | 4 | 2022-04
3005975929-2022-00131 | 2022-04-07 | 2021-07-17 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 51 YR | Female | US LEGAL. BILATERAL PATIENT. IT WAS REPORTED THAT, AFTER A LEFT HIP REPLACEMENT IN 2011, PLAINTIFF HAS EXPERIENCED A MECHANICAL COMPLICATION OF THE LEFT HIP REPLACEMENT, LEFT HIP PAIN, AND ONGOING DISCOMFORT WITH WEIGHT BEARING AND RANGE OF MOTION. PLAINTIFF UNDERWENT REVISION SURGERY ON THE LEFT HIP ON (B)(6) 2021. THE PATIENT'S OUTCOME IS UNKNOWN. |  | 2022 | 4 | 2022-04
3005975929-2022-00129 | 2022-04-07 | 2021-03-15 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 47 YR | Male | US LEGAL. IT WAS REPORTED THAT THE PLAINTIFF UNDERWENT A LEFT HIP REVISION SURGERY DUE TO PAIN, STIFFNESS, AND A MECHANICAL FAILURE (LOOSENING) OF THE COMPONENTS. THE SURGEON ALSO NOTED THAT THE PATIENT HAD A LEG LENGTH DISCREPANCY. DURING THE REVISION SURGERY, BOTH BHR COMPONENTS WERE EXPLANTED AND REPLACED WITH A THA SYSTEM FROM A COMPETITOR. THE PATIENT TOLERATED THE PROCEDURE WELL. |  | 2022 | 4 | 2022-04
3005975929-2022-00140 | 2022-04-08 | 2021-03-03 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | US LEGAL. IT WAS REPORTED THAT, AFTER A BHR THA WAS PERFORMED ON (B)(6) 2007 DUE TO SEVERE OSTEOARTHRITIS, PAIN, RESTRICTION OF MOTION AND INTERFERENCE WITH DAILY ACTIVITIES, THE PLAINTIFF EXPERIENCED FAILURE OF THA FOR METAL-ON-METAL COMPLICATION. THIS ADVERSE EVENT WAS TREATED WITH A REVISION SURGERY ON (B)(6) 2021, IN WHICH METAL DEBRIS, TRUNNION, AND CORROSION WERE NOTICED. PATIENT'S CURRENT HEALTH STATUS IS UNKNOWN. |  | 2022 | 4 | 2022-04
3005975929-2022-00144 | 2022-04-09 | 2022-03-18 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT, AFTER A THR PERFORMED ON (B)(6) 2011, THE PATIENT EXPERIENCED PAIN AND ELEVATED COBALT AND CHROMIUM LEVELS. THIS ADVERSE EVENT WAS SOLVED BY CONVERSION OF BHR TO BDMH ON (B)(6) 2022. CURRENT HEALTH STATUS OF PATIENT IS OK. |  | 2022 | 4 | 2022-04
3005975929-2022-00150 | 2022-04-13 | 2021-03-24 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 56 YR | Male | US LEGAL. IT WAS REPORTED THAT, AFTER A RIGHT BHR PERFORMED IN 2016. PLAINTIFF DEVELOPED PAIN AND EXTREMELY HIGH COBALT AND CHROMIUM ION LEVELS IN THE BLOOD. ON (B)(6) 2021, PLAINTIFF UNDERWENT A REVISION SURGERY ON THE RIGHT HIP. THE JOINT WAS NOTICED TO BE FULL OF METALLOSIS, METAL-STAINED INFLAMMATORY TISSUE AND GRANULATION, ALL THIS TISSUE WAS REMOVED. THE FEMORAL AND ACETABULAR COMPONENTS WERE EXPLANTED. THE CURRENT HEALTH STATE OF THE PLAINTIFF IS UNKNOWN. |  | 2022 | 4 | 2022-04
3005975929-2022-00160 | 2022-04-13 | 2017-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Female | US LEGAL: IT WAS REPORTED THAT, AFTER RIGHT BIRMINGHAM HIP RESURFACING SURGERY WAS PERFORMED IN 2013 DUE TO PAIN, STIFFNESS, CATCHING AND DECREASED ACTIVITY CAPABILITIES, THE PATIENT EXPERIENCED PAIN, SIGNIFICANTLY ELEVATED CHROMIUM AND COBALT IONS AND GLUTEUS TENDINOSIS. IT IS UNKNOWN IF ANY INTERVENTION WAS PERFORMED. |  | 2022 | 4 | 2022-04
3005975929-2022-00159 | 2022-04-13 | 2017-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Female | US LEGAL: IT WAS REPORTED THAT, AFTER RIGHT BIRMINGHAM HIP RESURFACING SURGERY WAS PERFORMED IN 2013 DUE TO PAIN, STIFFNESS, CATCHING AND DECREASED ACTIVITY CAPABILITIES, THE PATIENT EXPERIENCED PAIN, SIGNIFICANTLY ELEVATED CHROMIUM AND COBALT IONS AND GLUTEUS TENDINOSIS. IT IS UNKNOWN IF ANY INTERVENTION WAS PERFORMED. |  | 2022 | 4 | 2022-04
3005975929-2022-00155 | 2022-04-13 | 2021-05-13 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 61 YR | Male | US LEGAL: IT WAS REPORTED THAT, AFTER A THR SURGERY ON THE LEFT HIP, PERFORMED ON (B)(6) 2009 DUE TO ARTHRITIS AND DEFORMITY. PLAINTIFF UNDERWENT REVISION SURGERY ON LEFT HIP ON (B)(6) 2021, DUE TO ELEVATED COBALT AND CHROMIUM LEVELS. THERE WAS MINIMAL EVIDENCE OF METALLOSIS ON THE SURROUNDING TISSUE, HOWEVER THERE WERE BLACK RINGS OF METAL DEBRIS ON THE TRUNNION. LINER AND MODULAR HEAD WERE EXPLANTED. CURRENT HEALTH OF THE PLAINTIFF IS UNKNOWN. |  | 2022 | 4 | 2022-04
3005975929-2022-00151 | 2022-04-13 | 2021-03-24 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 56 YR | Male | US LEGAL: IT WAS REPORTED THAT, AFTER A RIGHT BHR PERFORMED IN 2016. PLAINTIFF DEVELOPED PAIN AND EXTREMELY HIGH COBALT AND CHROMIUM ION LEVELS IN THE BLOOD. ON (B)(6) 2021, PLAINTIFF UNDERWENT A REVISION SURGERY ON THE RIGHT HIP. THE JOINT WAS NOTICED TO BE FULL OF METALLOSIS, METAL-STAINED INFLAMMATORY TISSUE AND GRANULATION, ALL THIS TISSUE WAS REMOVED. THE FEMORAL AND ACETABULAR COMPONENTS WERE EXPLANTED. THE CURRENT HEALTH STATE OF THE PLAINTIFF IS UNKNOWN. |  | 2022 | 4 | 2022-04
3005975929-2022-00166 | 2022-04-18 | 2022-03-22 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT, AFTER A THR PERFORMED ON (B)(6) 2015, THE PATIENT EXPERIENCED METALLOSIS. THIS ADVERSE EVENT WAS SOLVED BY CONVERSION OF BHR TO BDMH ON (B)(6) 2022. CURRENT HEALTH STATUS OF PATIENT IS UNKNOWN. |  | 2022 | 4 | 2022-04
3005975929-2022-00173 | 2022-04-19 | 1901-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 40 YR | Male | US LEGAL. BILATERAL PATIENT. IT IS REPORTED THAT, AFTER A RIGHT BHR SURGERY WAS PERFORMED ON UNKNOWN DATE, PLAINTIFF HAS EXPERIENCING EXTREME HIP PAIN, LOCAL ADVERSE TISSUE REACTION, LIMITED MOTION OF THE RIGHT HIP, PREMATURE FAILURE OF THE RIGHT HIP BHR DEVICE, AND METALLOSIS. IT IS UNKNOWN IF OR HOW THE EVENT HAS BEEN TREATED. THE PATIENT'S CONDITION IS UNKNOWN. |  | 2022 | 4 | 2022-04
3005975929-2022-00172 | 2022-04-19 | 1901-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 40 YR | Male | US LEGAL. BILATERAL PATIENT. IT IS REPORTED THAT, AFTER A RIGHT BHR SURGERY WAS PERFORMED ON UNKNOWN DATE, PLAINTIFF HAS EXPERIENCING EXTREME HIP PAIN, LOCAL ADVERSE TISSUE REACTION, LIMITED MOTION OF THE RIGHT HIP, PREMATURE FAILURE OF THE RIGHT HIP BHR DEVICE, AND METALLOSIS. IT IS UNKNOWN IF OR HOW THE EVENT HAS BEEN TREATED. THE PATIENT'S CONDITION IS UNKNOWN. |  | 2022 | 4 | 2022-04
3005975929-2022-00176 | 2022-04-22 | 2021-08-17 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | BIRMINGHAM HIP RESURFACING (BHR) CUP | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Male | IT WAS REPORTED THAT THE PLAINTIFF UNDERWENT A RIGHT HIP REVISION SURGERY ON (B)(6) 2021 DUE TO HIP PAIN, METALLOSIS, ELEVATED METAL ION LEVELS AND PSEUDOTUMOR. THE CURRENT STATE OF HEALTH OF THE PATIENT IS UNKNOWN. THE PRIMARY RIGHT HIP BHR THA SURGERY WAS PERFORMED ON (B)(6) 2008. |  | 2022 | 4 | 2022-04
3005975929-2022-00178 | 2022-04-25 | 2021-10-05 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | US LEGAL,BILATERAL PATIENT. IT WAS REPORTED THAT THE PLAINTIFF UNDERWENT A LEFT HIP REVISION SURGERY ON (B)(6) 2021 DUE TO HIP PAIN, METALLOSIS, ELEVATED METAL ION LEVELS AND PSEUDOTUMOR. THIS PATIENT UNDERWENT A RIGHT HIP REVISION SURGERY DUE TO THE SAME REASONS ON (B)(6)2021 (COVERED UNDER CASE- (B)(4)). THE CURRENT STATE OF HEALTH OF THE PATIENT IS UNKNOWN. THE PRIMARY LEFT HIP BHR THA SURGERY WAS PERFORMED ON (B)(6)2008. |  | 2022 | 4 | 2022-04
3005975929-2022-00179 | 2022-04-25 | 2021-10-05 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | BILATERAL PATIENT. IT WAS REPORTED THAT THE PLAINTIFF UNDERWENT A LEFT HIP REVISION SURGERY ON (B)(6) 2021 DUE TO HIP PAIN, METALLOSIS, ELEVATED METAL ION LEVELS AND PSEUDOTUMOR. THIS PATIENT UNDERWENT A RIGHT HIP REVISION SURGERY DUE TO THE SAME REASONS ON (B)(6) 2021 (COVERED UNDER CASE-(B)(4)). THE CURRENT STATE OF HEALTH OF THE PATIENT IS UNKNOWN. THE PRIMARY LEFT HIP BHR THA SURGERY WAS PERFORMED ON (B)(6) 2008. |  | 2022 | 4 | 2022-04
3005975929-2022-00180 | 2022-04-25 | 2021-08-17 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | *US LEGAL* BILATERAL PATIENT. IT WAS REPORTED THAT THE PLAINTIFF UNDERWENT A RIGHT HIP REVISION SURGERY ON (B)(6) 2021 DUE TO HIP PAIN, METALLOSIS, ELEVATED METAL ION LEVELS AND PSEUDOTUMOR. THIS PATIENT UNDERWENT A LEFT HIP REVISION SURGERY DUE TO THE SAME REASONS ON (B)(6) 2021 (COVERED UNDER (B)(4)). THE CURRENT STATE OF HEALTH OF THE PATIENT IS UNKNOWN. THE PRIMARY RIGHT HIP BHR THA SURGERY WAS PERFORMED ON (B)(6) 2008. |  | 2022 | 4 | 2022-04
3005975929-2022-00184 | 2022-04-29 | 2022-02-16 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP INSTRUMENT | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT DURING A THR SURGERY, THE SURGEON WAS USING THE HEAD NECK SIZER TO GAUGE THE SIZE OF THE FEMORAL NECK. AS HE WAS ROTATING IT AROUND THE NECK, THE SIZER BROKE IN THE MIDDLE SEPARATING IT INTO TWO PIECES. THE PROCEDURE WAS FINISHED WITH A SMITH AND NEPHEW BACK UP DEVICE WITHOUT DELAY. THE PATIENT WAS NOT HARMED BEYOND THE REPORTED PROBLEM. |  | 2022 | 4 | 2022-04
3005975929-2022-00186 | 2022-05-04 | 2022-04-07 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS,,HIP,HEMI-,TRUNNIONBEARING,FEMORAL,METALPOLYACETAL |  |  | Male | IT WAS REPORTED THAT AFTER A THR PERFORMED ON AN UNKNOWN DATE THE PATIENT EXPERIENCED A FEMUR NECK FRACTURE OF TRAUMATIC NATURE. THIS WAS ADDRESSED VIA REVISION SURGERY ON (B)(6) 2022 AND THE BHR FEMORAL COMPONENT WAS EXPLANTED. THE PATIENT CURRENT STATUS IS UNKNOWN. |  | 2022 | 5 | 2022-05
3005975929-2022-00204 | 2022-05-06 | 2022-02-16 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "LONG-TERM BLOOD METAL ION LEVELS AND CLINICAL OUTCOME AFTER BIRMINGHAM HIP ARTHROPLASTY," 5 PATIENTS WITH BHR SURGERY HAD A. REVISION SURGERY DUE TO LOOSENING OF THE FEMORAL HEAD. FURTHER INFORMATION IS UNKNOWN |  | 2022 | 5 | 2022-05
3005975929-2022-00205 | 2022-05-06 | 2022-02-16 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "LONG-TERM BLOOD METAL ION LEVELS AND CLINICAL OUTCOME AFTER BIRMINGHAM HIP ARTHROPLASTY," 44 PATIENTS HAD PSEUDOTUMOR WITH RADIOLOGICALLY DIAGNOSIS. IT IS UNKNOWN HOW THIS WERE TREATED. FURTHER INFORMATION IS UNKNOWN |  | 2022 | 5 | 2022-05
3005975929-2022-00207 | 2022-05-06 | 2022-02-16 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, HEMI-, FEMORAL, METAL |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "LONG-TERM BLOOD METAL ION LEVELS AND CLINICAL OUTCOME AFTER BIRMINGHAM HIP ARTHROPLASTY," 2 PATIENTS WITH UNILATERAL BHR THA SURGERY AND 1 PATIENT WITH BILATERAL BHR THA HAD A REVISION SURGERY DUE TO METALLOSIS. FURTHER INFORMATION IS UNKNOWN. |  | 2022 | 5 | 2022-05
3005975929-2022-00208 | 2022-05-06 | 2022-02-16 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "LONG-TERM BLOOD METAL ION LEVELS AND CLINICAL OUTCOME AFTER BIRMINGHAM HIP ARTHROPLASTY," 1 PATIENT WITH BHR SURGERY HAD A REVISION SURGERY DUE TO LEG LENGTH DISCREPANCY. FURTHER INFORMATION IS UNKNOWN. |  | 2022 | 5 | 2022-05
3005975929-2022-00203 | 2022-05-06 | 2022-02-16 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "LONG-TERM BLOOD METAL ION LEVELS AND CLINICAL OUTCOME AFTER BIRMINGHAM HIP ARTHROPLASTY," 14 PATIENTS HAD METAL ION LEVELS ABOVE 10 PPB WITH NO REVISION SURGERY. IT IS UNKNOWN HOW THEY WERE TREATED. FURTHER INFORMATION IS UNKNOWN. |  | 2022 | 5 | 2022-05
3005975929-2022-00206 | 2022-05-06 | 2022-02-16 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP MODULAR HEAD (BHMH) | PROSTHESIS, HIP, HEMI-, FEMORAL, METAL |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "LONG-TERM BLOOD METAL ION LEVELS AND CLINICAL OUTCOME AFTER BIRMINGHAM HIP ARTHROPLASTY," 2 PATIENTS WITH UNILATERAL BHR THA SURGERY AND 1 PATIENT WITH BILATERAL BHR THA HAD A REVISION SURGERY DUE TO METALLOSIS. FURTHER INFORMATION IS UNKNOWN. |  | 2022 | 5 | 2022-05
3005975929-2022-00199 | 2022-05-06 | 2022-02-16 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS,HIP,HEMI-,TRUNNION-BEARING,FEMORAL,METALPOLYACETAL |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "LONG-TERM BLOOD METAL ION LEVELS AND CLINICAL OUTCOME AFTER BIRMINGHAM HIP ARTHROPLASTY ", 11 PATIENTS WITH BHR SURGERY HAD METAL ION LEVELS ABOVE 10 PPB. A REVISION SURGERY WAS PERFORMED TO TREAT THIS ADVERSE EVENT. FURTHER INFORMATION IS NOT KNOWN. |  | 2022 | 5 | 2022-05
3005975929-2022-00198 | 2022-05-06 | 2022-02-16 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS,HIP,HEMI-,TRUNNION-BEARING,FEMORAL,METALPOLYACETAL |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "LONG-TERM BLOOD METAL ION LEVELS AND CLINICAL OUTCOME AFTER (B)(6) HIP ARTHROPLASTY," 11 PATIENTS WITH BHR SURGERY HAD METAL ION LEVELS ABOVE 10 PPB. A REVISION SURGERY WAS PERFORMED TO TREAT THIS ADVERSE EVENT. FURTHER INFORMATION IS NOT KNOWN. |  | 2022 | 5 | 2022-05
3005975929-2022-00197 | 2022-05-06 | 2022-02-16 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS,HIP,HEMI-,TRUNNION-BEARING,FEMORAL, METAL/POLYACETAL |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "LONG-TERM BLOOD METAL ION LEVELS AND CLINICAL OUTCOME AFTER BIRMINGHAM HIP ARTHROPLASTY ", 4 PATIENTS WITH BHR SURGERY HAD A REVISION SURGERY DUE TO MECHANICAL IMPINGEMENT. FURTHER INFORMATION IS UNKNOWN. |  | 2022 | 5 | 2022-05
3005975929-2022-00196 | 2022-05-06 | 2022-04-04 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS,HIP,HEMI-,TRUNNION-BEARING,FEMORAL,METALPOLYACETAL |  |  | Female | IT WAS REPORTED THAT ON LITERATURE REVIEW "THE 20-YEAR RESULTS OF BIRMINGHAM HIP RESURFACING CONSECUTIVE COHORT SERIES, 1998-2001", 3 PATIENTS SUFFERED FROM AN METALLOSIS AT 7-8 YEARS AND REQUIRED A REVISION SURGERY. THERE'S NO MORE INFORMATION AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00192 | 2022-05-06 | 2022-04-04 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS,HIP,HEMI-,TRUNNION-BEARING,FEMORAL,METALPOLYACETAL |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "THE 20-YEAR RESULTS OF BIRMINGHAM HIP RESURFACING CONSECUTIVE COHORT SERIES, 1998-2001", 1 PATIENT SUFFERED FROM METALLOSIS AT 18 YEARS AND REQUIRED A REVISION SURGERY. THERE'S NO MORE INFORMATION AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00191 | 2022-05-06 | 2022-04-04 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS,HIP,HEMI-,TRUNNION-BEARING,FEMORAL,METALPOLYACETAL |  |  | Female | IT WAS REPORTED THAT ON LITERATURE REVIEW "THE 20-YEAR RESULTS OF BIRMINGHAM HIP RESURFACING CONSECUTIVE COHORT SERIES, 1998-2001", 1 PATIENT SUFFERED FROM METALLOSIS AT 18 YEARS AND REQUIRED A REVISION SURGERY. THERE'S NO MORE INFORMATION AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00190 | 2022-05-06 | 2022-04-04 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS,HIP,HEMI-,TRUNNION-BEARING,FEMORAL,METALPOLYACETAL |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "THE 20-YEAR RESULTS OF BIRMINGHAM HIP RESURFACING CONSECUTIVE COHORT SERIES, 1998-2001", 1 PATIENT SUFFERED FROM METALLOSIS AT 15 YEARS AND REQUIRED A REVISION SURGERY. THERE'S NO MORE INFORMATION AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00193 | 2022-05-06 | 2022-04-04 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS,HIP,HEMI-,TRUNNION-BEARING,FEMORAL,METALPOLYACETAL |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "THE 20-YEAR RESULTS OF BIRMINGHAM HIP RESURFACING CONSECUTIVE COHORT SERIES, 1998-2001", 3 PATIENTS SUFFERED FROM AN METALLOSIS AT 7-8 YEARS AND REQUIRED A REVISION SURGERY. THERE'S NO MORE INFORMATION AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00194 | 2022-05-06 | 2022-04-04 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS,HIP,HEMI-,TRUNNIONBEARING,FEMORAL,METALPOLYACETAL |  |  | Male | IT WAS REPORTED THAT ON LITERATURE REVIEW "THE 20-YEAR RESULTS OF BIRMINGHAM HIP RESURFACING CONSECUTIVE COHORT SERIES, 1998-2001", 1 PATIENTS SUFFERED FROM AN INFECTION AT 5 YEARS AND REQUIRED A REVISION SURGERY. THERE'S NO MORE INFORMATION AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00195 | 2022-05-06 | 2022-04-04 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS,HIP,HEMI-,TRUNNIONBEARING,FEMORAL,METALPOLYACETAL |  |  | Female | IT WAS REPORTED THAT ON LITERATURE REVIEW "THE 20-YEAR RESULTS OF BIRMINGHAM HIP RESURFACING CONSECUTIVE COHORT SERIES, 1998-2001", 1 PATIENT SUFFERED FROM METALLOSIS AT 15 YEARS AND REQUIRED A REVISION SURGERY. THERE'S NO MORE INFORMATION AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00200 | 2022-05-06 | 2022-02-16 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP IMPL | PROSTHESIS,HIP,HEMI-,TRUNNION-BEARING,FEMORAL,METALPOLYACETAL |  |  | Unknown |  |  | 2022 | 5 | 2022-05
3005975929-2022-00229 | 2022-05-07 | 2019-08-09 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "DOES MODULARITY OF METAL-ON-METAL HIP IMPLANTS INCREASE COBALT: CHROMIUM RATIO?", FROM 414 PATIENTS THAT HAD BHR, 41 PATIENTS HAD HIGH METAL IONS IN BLOOD THAT WAS ADDRESSED VIA REVISION SURGERY. THE PATIENTS OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00231 | 2022-05-07 | 2019-08-09 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "DOES MODULARITY OF METAL-ON-METAL HIP IMPLANTS INCREASE COBALT: CHROMIUM RATIO?", FROM 414 PATIENTS THAT HAD BHR, 49 PATIENTS HAD A PSEUDOTUMOR THAT WAS ADDRESSED VIA REVISION SURGERY. THE PATIENTS OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00236 | 2022-05-07 | 2019-08-09 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS,HIP,HEMI-,TRUNNION-BEARING,FEMORAL,METALPOLYACETAL |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "DOES MODULARITY OF METAL-ON-METAL HIP IMPLANTS INCREASE COBALT: CHROMIUM RATIO? BLOOD METAL IONS LEVELS WERE TAKE TO ALL 35 PATIENTS THAT HAD BMHR. THE COBALT MEAN WAS OF 2.8 PPB (MIN 0.18 PPB, MAX 10 PPB) AND CHROMIUM MEAN WAS OF 3 PPB (MIN 0.36 PPB, MAX 12 PPB), IT IS UNKNOWN THE EXACT QUANTITY OF PATIENTS THAT HAD MORE LEVELS THAN ACCEPTABLE AND HOW THIS WAS ADDRESSED. FURTHER INFORMATION IS NOT KNOWN FOR THE MOMENT. |  | 2022 | 5 | 2022-05
3005975929-2022-00232 | 2022-05-07 | 2019-08-09 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "DOES MODULARITY OF METAL-ON-METAL HIP IMPLANTS INCREASE COBALT: CHROMIUM RATIO," FROM 414 PATIENTS THAT HAD BHR, 10 PATIENTS PRESENTED INFECTION THAT WAS ADDRESSED VIA REVISION SURGERY. THE PATIENTS OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00237 | 2022-05-07 | 2019-08-09 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "DOES MODULARITY OF METAL-ON-METAL HIP IMPLANTS INCREASE COBALT: CHROMIUM RATIO?", BLOOD METAL IONS LEVELS WERE TAKE TO ALL 54 PATIENTS THAT HAD BHR THA. THE COBALT MEAN WAS OF 15 PPB (MIN 0.83 PPB, MAX 119 PPB) AND CHROMIUM MEAN WAS OF 7.8 PPB (MIN 0.7 PPB, MAX 43PPB), IT IS UNKNOWN THE EXACT QUANTITY OF PATIENTS THAT HAD MORE LEVELS THAN ACCEPTABLE AND HOW THIS WAS ADDRESSED. FURTHER INFORMATION IS NOT KNOWN FOR THE MOMENT. |  | 2022 | 5 | 2022-05
3005975929-2022-00239 | 2022-05-07 | 2019-08-09 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "DOES MODULARITY OF METAL-ON-METAL HIP IMPLANTS INCREASE COBALT: CHROMIUM RATIO?", FROM 54 PATIENTS THAT HAD BHR THA, 25 PATIENTS HAD UNEXPLAINED PAIN THAT WAS ADDRESSED VIA REVISION SURGERY. THE PATIENTS OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00238 | 2022-05-07 | 2019-08-09 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "DOES MODULARITY OF METAL-ON-METAL HIP IMPLANTS INCREASE COBALT: CHROMIUM RATIO?", FROM 414 PATIENTS THAT HAD BHR, 14 PATIENTS PRESENTED FRACTURE AND WAS ADDRESSED VIA REVISION SURGERY. THE PATIENTS OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00234 | 2022-05-07 | 2019-08-09 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "DOES MODULARITY OF METAL-ON-METAL HIP IMPLANTS INCREASE COBALT: CHROMIUM RATIO?", BLOOD METAL IONS LEVELS WERE TAKE TO ALL 414 PATIENTS THAT HAD BHR. THE COBALT MEAN WAS OF 19 PPB (MIN 0 PPB, MAX 250 PPB) AND CHROMIUM MEAN WAS OF 16 PPB (MIN 0.3 PPB, MAX 343 PPB), IT IS UNKNOWN THE EXACT QUANTITY OF PATIENTS THAT HAD MORE LEVELS THAN ACCEPTABLE AND HOW THIS WAS ADDRESSED. FURTHER INFORMATION IS NOT KNOWN FOR THE MOMENT. |  | 2022 | 5 | 2022-05
3005975929-2022-00240 | 2022-05-07 | 2019-08-09 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "DOES MODULARITY OF METAL-ON-METAL HIP IMPLANTS INCREASE COBALT: CHROMIUM RATIO?", FROM 414 PATIENTS THAT HAD BHR, 11 PATIENTS HAD IMPINGEMENT, THIS WAS ADDRESSED VIA REVISION SURGERY. THE PATIENTS OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00230 | 2022-05-07 | 2019-08-09 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "DOES MODULARITY OF METAL-ON-METAL HIP IMPLANTS INCREASE COBALT: CHROMIUM RATIO?", FROM 414 PATIENTS THAT HAD BHR, 34 PATIENTS PRESENTED ASEPTIC LOOSENING THAT WAS ADDRESSED VIA REVISION SURGERY. THE PATIENTS OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00242 | 2022-05-07 | 2021-06-21 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "COMPARISON OF THE LONG-TERM CAUSE OF FAILURE AND SURVIVORSHIP OF FOUR HUNDRED AND TWENTY SEVEN METAL-ON-METAL HIP ARTHROPLASTIES: RESURFACING VERSUS LARGE HEAD TOTAL HIP ARTHROPLASTY", 7 PATIENTS SUFFERED FROM PAIN AND ELEVATED METAL ION LEVELS AFTER A PRIMARY BHR RESURFACING SURGERY. THE PATIENTS REQUIRED A REVISION SURGERY TO TREAT THE ADVERSE EVENT. THE OUTCOME OF THE PATIENTS IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00245 | 2022-05-07 | 2022-04-04 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "THE 20-YEAR RESULTS OF BIRMINGHAM HIP RESURFACING CONSECUTIVE COHORT SERIES, 1998-2001," 1 PATIENT SUFFERED FROM AN AVASCULAR NECROSIS (AVN) AT 2 MONTHS AND REQUIRED A REVISION SURGERY. THERE'S NO MORE INFORMATION AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00233 | 2022-05-07 | 2019-08-09 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "DOES MODULARITY OF METAL-ON-METAL HIP IMPLANTS INCREASE COBALT: CHROMIUM RATIO?", BLOOD METAL IONS LEVELS WERE TAKE TO ALL 414 PATIENTS THAT HAD BHR. THE COBALT MEAN WAS OF 19 PPB (MIN 0 PPB, MAX 250 PPB) AND CHROMIUM MEAN WAS OF 16 PPB (MIN 0.3 PPB, MAX 343 PPB), IT IS UNKNOWN THE EXACT QUANTITY OF PATIENTS THAT HAD MORE LEVELS THAN ACCEPTABLE AND HOW THIS WAS ADDRESSED. FURTHER INFORMATION IS NOT KNOWN FOR THE MOMENT. |  | 2022 | 5 | 2022-05
3005975929-2022-00228 | 2022-05-07 | 2019-08-09 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "DOES MODULARITY OF METAL-ON-METAL HIP IMPLANTS INCREASE COBALT: CHROMIUM RATIO?", FROM 414 PATIENTS THAT HAD BHR, 41 PATIENTS HAD HIGH METAL IONS IN BLOOD THAT WAS ADDRESSED VIA REVISION SURGERY. THE PATIENTS OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00235 | 2022-05-07 | 2019-08-09 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "DOES MODULARITY OF METAL-ON-METAL HIP IMPLANTS INCREASE COBALT: CHROMIUM RATIO?", BLOOD METAL IONS LEVELS WERE TAKE TO ALL 54 PATIENTS THAT HAD BHR THA. THE COBALT MEAN WAS OF 15 PPB (MIN 0.83 PPB, MAX 119 PPB) AND CHROMIUM MEAN WAS OF 7.8 PPB (MIN 0.7 PPB, MAX 43PPB), IT IS UNKNOWN THE EXACT QUANTITY OF PATIENTS THAT HAD MORE LEVELS THAN ACCEPTABLE AND HOW THIS WAS ADDRESSED. FURTHER INFORMATION IS NOT KNOWN FOR THE MOMENT. |  | 2022 | 5 | 2022-05
3005975929-2022-00224 | 2022-05-07 | 2019-08-09 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "DOES MODULARITY OF METAL-ON-METAL HIP IMPLANTS INCREASE COBALT: CHROMIUM RATIO?", FROM 54 PATIENTS THAT HAD BHR THA, 19 PATIENTS PRESENTED PSEUDOTUMOUR THAT WAS ADDRESSED VIA REVISION SURGERY. THE PATIENTS OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00214 | 2022-05-07 | 2022-02-16 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "LONG-TERM BLOOD METAL ION LEVELS AND CLINICAL OUTCOME AFTER BIRMINGHAM HIP ARTHROPLASTY ", 7 PATIENTS WITH BHR SURGERY HAD A REVISION SURGERY DUE TO PERIPROSTHETIC FRACTURE. FURTHER INFORMATION IS NOT KNOWN. |  | 2022 | 5 | 2022-05
3005975929-2022-00213 | 2022-05-07 | 2021-06-21 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "COMPARISON OF THE LONG-TERM CAUSE OF FAILURE AND SURVIVORSHIP OF FOUR HUNDRED AND TWENTY SEVEN METAL-ON-METAL HIP ARTHROPLASTIES: RESURFACING VERSUS LARGE HEAD TOTAL HIP ARTHROPLASTY," 7 PATIENTS SUFFERED FROM PAIN AND ELEVATED METAL ION LEVELS AFTER A PRIMARY BHR RESURFACING SURGERY. THE PATIENTS REQUIRED A REVISION SURGERY TO TREAT THE ADVERSE EVENT. THE OUTCOME OF THE PATIENTS IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00227 | 2022-05-07 | 2019-08-09 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "DOES MODULARITY OF METAL-ON-METAL HIP IMPLANTS INCREASE COBALT: CHROMIUM RATIO?", FROM 414 PATIENTS THAT HAD BHR, 7 PATIENTS HAD INSTABILITY THAT WAS ADDRESSED VIA REVISION SURGERY. THE PATIENTS OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00212 | 2022-05-07 | 2021-06-21 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "COMPARISON OF THE LONG-TERM CAUSE OF FAILURE AND SURVIVORSHIP OF FOUR HUNDRED AND TWENTY SEVEN METAL-ON-METAL HIP ARTHROPLASTIES: RESURFACING VERSUS LARGE HEAD TOTAL HIP ARTHROPLASTY", 1 PATIENT SUFFERED FROM INSTABILITY AFTER A PRIMARY BHR RESURFACING SURGERY. THE PATIENT REQUIRED A REVISION SURGERY TO TREAT THE ADVERSE EVENT. THE OUTCOME OF THE PATIENT IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00210 | 2022-05-07 | 2021-06-21 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "COMPARISON OF THE LONG-TERM CAUSE OF FAILURE AND SURVIVORSHIP OF FOUR HUNDRED AND TWENTY SEVEN METAL-ON-METAL HIP ARTHROPLASTIES: RESURFACING VERSUS LARGE HEAD TOTAL HIP ARTHROPLASTY", 2 PATIENTS SUFFERED FROM A FEMORAL NECK FRACTURE AFTER A PRIMARY BHR RESURFACING SURGERY. THE PATIENTS REQUIRED A REVISION SURGERY TO TREAT THE ADVERSE EVENT. THE OUTCOME OF THE PATIENTS IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00211 | 2022-05-07 | 2021-06-21 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "COMPARISON OF THE LONG-TERM CAUSE OF FAILURE AND SURVIVORSHIP OF FOUR HUNDRED AND TWENTY SEVEN METAL-ON-METAL HIP ARTHROPLASTIES: RESURFACING VERSUS LARGE HEAD TOTAL HIP ARTHROPLASTY", 6 PATIENTS SUFFERED FROM ADVERSE REACTION TO METAL DEBRIS (ARMD) AFTER A PRIMARY BHR RESURFACING SURGERY. THE PATIENTS REQUIRED A REVISION SURGERY TO TREAT THE ADVERSE EVENT. THE OUTCOME OF THE PATIENTS IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00215 | 2022-05-07 | 2022-02-16 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "LONG-TERM BLOOD METAL ION LEVELS AND CLINICAL OUTCOME AFTER BIRMINGHAM HIP ARTHROPLASTY", 2 PATIENTS WITH BHR SURGERY HAD A REVISION SURGERY DUE TO INFECTION. FURTHER INFORMATION IS UNKNOWN. |  | 2022 | 5 | 2022-05
3005975929-2022-00216 | 2022-05-07 | 2022-04-16 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "LONG-TERM BLOOD METAL ION LEVELS AND CLINICAL OUTCOME AFTER BIRMINGHAM HIP ARTHROPLASTY", 1 PATIENT WITH BHR THA SURGERY HAD A REVISION SURGERY DUE TO PAIN. FURTHER INFORMATION IS UNKNOWN. |  | 2022 | 5 | 2022-05
3005975929-2022-00219 | 2022-05-07 | 2022-02-16 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "LONG-TERM BLOOD METAL ION LEVELS AND CLINICAL OUTCOME AFTER BIRMINGHAM HIP ARTHROPLASTY", 2 PATIENTS WITH BHR SURGERY HAD A REVISION SURGERY DUE TO MALALIGNMENT. FURTHER INFORMATION IS UNKNOWN. |  | 2022 | 5 | 2022-05
3005975929-2022-00217 | 2022-05-07 | 2022-02-16 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "LONG-TERM BLOOD METAL ION LEVELS AND CLINICAL OUTCOME AFTER BIRMINGHAM HIP ARTHROPLASTY ", AFTER 1 PATIENT HAD BHR SURGERY, REVISION SURGERY WAS PERFORMED DUE TO PAIN. FURTHER INFORMATION IS UNKNOWN. |  | 2022 | 5 | 2022-05
3005975929-2022-00220 | 2022-05-07 | 2021-06-21 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "COMPARISON OF THE LONG-TERM CAUSE OF FAILURE AND SURVIVORSHIP OF FOUR HUNDRED AND TWENTY SEVEN METAL-ON-METAL HIP ARTHROPLASTIES: RESURFACING VERSUS LARGE HEAD TOTAL HIP ARTHROPLASTY", 6 PATIENTS SUFFERED FROM ADVERSE REACTION TO METAL DEBRIS (ARMD) AFTER A PRIMARY BHR RESURFACING SURGERY. THE PATIENTS REQUIRED A REVISION SURGERY TO TREAT THE ADVERSE EVENT. THE OUTCOME OF THE PATIENTS IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00221 | 2022-05-07 | 2022-04-04 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "THE 20-YEAR RESULTS OF BIRMINGHAM HIP RESURFACING CONSECUTIVE COHORT SERIES, 1998-2001," 3 PATIENTS SUFFERED AN EARLY FRACTURE AND REQUIRED A REVISION SURGERY. THERE'S NO MORE AVAILABLE INFORMATION. |  | 2022 | 5 | 2022-05
3005975929-2022-00187 | 2022-05-07 | 2021-03-26 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS,HIP,HEMI-,TRUNNIONBEARING,FEMORAL,METAL/POLYACETAL |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "CAN HIP RESURFACING BE SAFELY REVISED WITH SHORT-STEM TOTAL HIP ARTHROPLASTY? A CASE SERIES OF SIX PATIENTS," 1 PATIENT OPERATED IN 2007 WITH BHR PRESENTED PAIN, THE CLINICAL EXAMINATION SHOWED A HIP ROTATION STIFFNESS AND A HIP EXTENSION DEFICIT OF 10. THE IMAGING REVEALED A LOOSE PROSTHESIS WITH AN ACETABULAR BONE DEFECT AND METALLOSIS. |  | 2022 | 5 | 2022-05
3005975929-2022-00189 | 2022-05-07 | 2021-03-26 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS,HIP,HEMI-,TRUNNIONBEARING,FEMORAL,METALPOLYACETAL |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "CAN HIP RESURFACING BE SAFELY REVISED WITH SHORT-STEM TOTAL HIP ARTHROPLASTY? A CASE SERIES OF SIX PATIENTS", 1 PATIENT OPERATED IN 2004 WITH BHR PRESENTED LIMITED HIP MOBILITY AND PAIN AT HIP ROTATION SINCE 2016. AFTER DIAGNOSIS WITH ACETABULAR LOOSENING AND ACETABULAR BONE DEFECT, A REVISION USING COMPETITOR'S DEVICE WAS PERFORMED. AT LAST STUDY FOLLOW UP, THE PATIENT'S OUTCOME WAS EXCELLENT. A REVISION USING COMPETITOR'S DEVICE WAS PERFORMED. AT LAST STUDY FOLLOW UP, THE PATIENT'S OUTCOME WAS EXCELLENT. |  | 2022 | 5 | 2022-05
3005975929-2022-00188 | 2022-05-07 | 2021-03-26 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS,HIP,HEMI-,TRUNNION-BEARING,FEMORAL,METALPOLYACETAL |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "CAN HIP RESURFACING BE SAFELY REVISED WITH SHORT-STEM TOTAL HIP ARTHROPLASTY? A CASE SERIES OF SIX PATIENTS", 1 PATIENT OPERATED IN 2007 WITH BHR PRESENTED PAIN, THE CLINICAL EXAMINATION SHOWED A HIP ROTATION STIFFNESS AND A HIP EXTENSION DEFICIT OF 10. THE IMAGING REVEALED A LOOSE PROSTHESIS WITH AN ACETABULAR BONE DEFECT AND METALLOSIS. |  | 2022 | 5 | 2022-05
3005975929-2022-00209 | 2022-05-07 | 2022-02-16 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS,HIP,HEMI-,TRUNNION-BEARING,FEMORAL,METALPOLYACETAL |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "LONG-TERM BLOOD METAL ION LEVELS AND CLINICAL OUTCOME AFTER BIRMINGHAM HIP ARTHROPLASTY ", 7 PATIENTS WITH BHR SURGERY HAD A REVISION SURGERY DUE LOOSENING OF THE CUP. FURTHER INFORMATION IS UNKNOWN. |  | 2022 | 5 | 2022-05
3005975929-2022-00225 | 2022-05-07 | 2019-08-09 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP MODULAR HEAD (BHMH) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "DOES MODULARITY OF METAL-ON-METAL HIP IMPLANTS INCREASE COBALT: CHROMIUM RATIO?", FROM 54 PATIENTS THAT HAD BHR THA, 4 PATIENTS HAD INFECTION THAT WAS ADDRESSED VIA REVISION SURGERY. THE PATIENTS OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00222 | 2022-05-07 | 2019-08-09 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "DOES MODULARITY OF METAL-ON-METAL HIP IMPLANTS INCREASE COBALT: CHROMIUM RATIO?", FROM 54 PATIENTS THAT HAD BHR THA, 1 PATIENT PRESENTED INSTABILITY THAT WAS ADDRESSED VIA REVISION SURGERY. THE PATIENTS OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00223 | 2022-05-07 | 2019-08-09 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "DOES MODULARITY OF METAL-ON-METAL HIP IMPLANTS INCREASE COBALT: CHROMIUM RATIO?" FROM 414 PATIENTS THAT HAD BHR, 142 PATIENTS PRESENTED UNEXPLAINED PAIN THAT WAS ADDRESSED VIA REVISION SURGERY. THE PATIENTS OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00218 | 2022-05-07 | 2022-04-16 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "LONG-TERM BLOOD METAL ION LEVELS AND CLINICAL OUTCOME AFTER BIRMINGHAM HIP ARTHROPLASTY", 1 PATIENT WITH BHR THA SURGERY HAD A REVISION SURGERY DUE TO INFECTION. FURTHER INFORMATION IS UNKNOWN. |  | 2022 | 5 | 2022-05
3005975929-2022-00226 | 2022-05-09 | 2019-08-09 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "DOES MODULARITY OF METAL-ON-METAL HIP IMPLANTS INCREASE COBALT: CHROMIUM RATIO?", FROM 54 PATIENTS THAT HAD BHR THA, 5 PATIENTS HAD ASEPTIC LOOSENING THAT WAS ADDRESSED VIA REVISION SURGERY. THE PATIENTS OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00270 | 2022-05-18 | 2022-01-25 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "MIDTERM RESULTS OF THE BIRMINGHAM HIP RESURFACING: A SINGLE-SURGEON SERIES", 1 PATIENT SUFFERED FROM ELEVATED METAL ION LEVELS, OSTEOLYSIS AND ALTTR AFTER A BHR HIP RESURFACING PROCEDURE. THE EVENT WAS TREATED VIA REVISION SURGERY 3 YEARS POSTOPERATIVELY. THE OUTCOME OF THE PATIENT IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00266 | 2022-05-18 | 2022-01-25 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "MIDTERM RESULTS OF THE BIRMINGHAM HIP RESURFACING: A SINGLE-SURGEON SERIES," 7 PATIENTS PRESENTED HIGH METAL ION LEVELS AND ALTR AFTER A BHR HIP RESURFACING PROCEDURE, UNKNOWN FEMORAL BHR HEAD INVOLVED. THE EVENT WAS TREATED VIA REVISION SURGERY. THE OUTCOME OF THE PATIENTS IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00267 | 2022-05-18 | 2016-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Female | IT WAS REPORTED THAT ON LITERATURE REVIEW "HIP RESURFACING: A SINGLE SURGEON U.S. SERIES WITH MINIMUM TEN-YEAR FOLLOW-UP¿ 7.4 YEARS AFTER A BHR SURGERY, ONE (1) PATIENT HAD PAIN, HIGH LEVELS OF METAL IONS (47.8 PPB OF CO AND 47.5 PPB OF CR), CUP MALPOSITION WITH EXCESSIVE ANTERVERSION AND EFFUSION. REVISION SURGERY WAS PERFORMED TO EXPLANT BOTH IMPLANTS. COMPETITOR DEVICES WERE IMPLANTED. THE PATIENT OUTCOME AND CURRENT HEALTH STATUS IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00259 | 2022-05-18 | 2021-03-02 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Male | IT WAS REPORTED THAT ON LITERATURE REVIEW "CASE OF MISTAKEN IDENTITY: COBALT CARDIOMYOPATHY VERSUS AMYLOIDOSIS ON CARDIAC MRI", ONE PATIENT UNDERWENT A REVISION SURGERY OF THE RIGHT HIP ON 2017 DUE TO PAIN, SWELLING, METALLOSIS, PSEUDOTUMOR FORMATION AND EXTREME ACETABULAR BONE LOSS. THE PATIENT WAS DIAGNOSED WITH COBALT TOXICITY, WHICH ACCOUNTED FOR BOTH PERIPHERAL NEUROPATHY AND THE PROGRESSIVE CARDIOMYOPATHY. 6 WEEKS AFTER REVISION, THE CARDIAC FUNCTION HAD SIGNIFICANTLY IMPROVED, AND 6 MONTHS AFTER REVISION, IT DEMONSTRATED FURTHER IMPROVEMENT. THE PRIMARY RIGHT BHR SURGERY WAS PERFORMED IN 2003. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00265 | 2022-05-18 | 2022-01-25 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "MIDTERM RESULTS OF THE BIRMINGHAM HIP RESURFACING: A SINGLE-SURGEON SERIES", 7 PATIENTS PRESENTED HIGH METAL ION LEVELS AND ALTR AFTER A BHR HIP RESURFACING PROCEDURE, UNKNOWN BHR CUP COMPONENT INVOLVED. THE EVENT WAS TREATED VIA REVISION SURGERY. THE OUTCOME OF THE PATIENTS IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00264 | 2022-05-18 | 2022-01-25 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "MIDTERM RESULTS OF THE BIRMINGHAM HIP RESURFACING: A SINGLE-SURGEON SERIES", 1 PATIENT SUFFERED FROM PAIN SECONDARY TO IMPINGEMENT WITHOUT SIGNS OF LOOSENING AFTER A BHR HIP RESURFACING PROCEDURE. THE EVENT WAS TREATED VIA REVISION SURGERY. THE OUTCOME OF THE PATIENT IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00263 | 2022-05-18 | 2016-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 57 YR | Male | IT WAS REPORTED THAT ON LITERATURE REVIEW "HIP RESURFACING: A SINGLE SURGEON U.S. SERIES WITH MINIMUM TEN-YEAR FOLLOW-UP¿ AROUND 9 YEARS AFTER A BHR THR SURGERY, ONE PATIENT HAD A FEMORAL NECK FRACTURE AFTER TRAUMATIC BIKE ACCIDENT. REVISION SURGERY WAS PERFORMED TO EXCHANGE THE HEAD FOR AN ANTHOLOGY STEM WITH A DUAL MOBILITY INSERT FROM S+N (AND RETAINED THE ORIGINAL BHR CUP). THE PATIENT OUTCOME AND CURRENT HEALTH STATUS IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00262 | 2022-05-18 | 2022-01-25 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "MIDTERM RESULTS OF THE BIRMINGHAM HIP RESURFACING: A SINGLE-SURGEON SERIES", 1 PATIENT SUFFERED FROM ELEVATED METAL ION LEVELS AFTER A BHR HIP RESURFACING PROCEDURE. THE EVENT WAS TREATED VIA REVISION SURGERY. THE OUTCOME OF THE PATIENT IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00261 | 2022-05-18 | 2022-01-25 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "MIDTERM RESULTS OF THE BIRMINGHAM HIP RESURFACING: A SINGLE-SURGEON SERIES", 1 PATIENT SUFFERED FROM ELEVATED METAL ION LEVELS AFTER A BHR HIP RESURFACING PROCEDURE. THE EVENT WAS TREATED VIA REVISION SURGERY. THE OUTCOME OF THE PATIENT IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00258 | 2022-05-18 | 2021-03-02 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Male | IT WAS REPORTED THAT ON LITERATURE REVIEW "CASE OF MISTAKEN IDENTITY: COBALT CARDIOMYOPATHY VERSUS AMYLOIDOSIS ON CARDIAC MRI", ONE PATIENT UNDERWENT A REVISION SURGERY OF THE RIGHT HIP ON 2017 DUE TO PAIN, SWELLING, METALLOSIS, PSEUDOTUMOR FORMATION AND EXTREME ACETABULAR BONE LOSS. THE PATIENT WAS DIAGNOSED WITH COBALT TOXICITY, WHICH ACCOUNTED FOR BOTH PERIPHERAL NEUROPATHY AND THE PROGRESSIVE CARDIOMYOPATHY. 6 WEEKS AFTER REVISION, THE CARDIAC FUNCTION HAD SIGNIFICANTLY IMPROVED, AND 6 MONTHS AFTER REVISION, IT DEMONSTRATED FURTHER IMPROVEMENT. THE PRIMARY RIGHT BHR SURGERY WAS PERFORMED IN 2003. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00253 | 2022-05-18 | 2021-03-23 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 63 YR | Female | *US LEGAL* BILATERAL PATIENT. IT WAS REPORTED THAT, AFTER A RIGHT BHR ON (B)(6) 2010 DUE TO SEVERE ARTHRITIS. PLAINTIFF EXPERIENCES ELEVATED LEVELS OF COBALT AND CHROMIUM IONS. ADVANCED IMAGING DEMONSTRATED A COLLECTION OF FLUID AROUND THE RIGHT HIP. DUE TO FAILED IMPLANTS WITH METALLOSIS, PSEUDOTUMOR AND MALPOSITION OF COMPONENTS, PLAINTIFF UNDERWENT RIGHT REVISION SURGERY ON (B)(6) 2021 WHERE SIGNIFICANT RETRO ACETABULAR OSTEOLYSIS AND SIGNS OF BONE HOLES WERE FOUND. THE STATUS OF THE PATIENT IS UNKNOWN. |  | 2022 | 5 | 2022-05
3005975929-2022-00252 | 2022-05-18 | 2021-03-23 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 63 YR | Female | US LEGAL - BILATERAL PATIENT. IT WAS REPORTED THAT, AFTER A RIGHT BHR ON (B)(6) 2010 DUE TO SEVERE ARTHRITIS. PLAINTIFF EXPERIENCES ELEVATED LEVELS OF COBALT AND CHROMIUM IONS. ADVANCED IMAGING DEMONSTRATED A COLLECTION OF FLUID AROUND THE RIGHT HIP. DUE TO FAILED IMPLANTS WITH METALLOSIS, PSEUDOTUMOR AND MALPOSITION OF COMPONENTS, PLAINTIFF UNDERWENT RIGHT REVISION SURGERY ON (B)(6) 2021 WHERE SIGNIFICANT RETRO ACETABULAR OSTEOLYSIS AND SIGNS OF BONE HOLES WERE FOUND. THE STATUS OF THE PATIENT IS UNKNOWN. |  | 2022 | 5 | 2022-05
3005975929-2022-00269 | 2022-05-18 | 2022-01-25 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "MIDTERM RESULTS OF THE BIRMINGHAM HIP RESURFACING: A SINGLE-SURGEON SERIES", 1 PATIENT SUFFERED FROM ELEVATED METAL ION LEVELS, OSTEOLYSIS AND ALTTR AFTER A BHR HIP RESURFACING PROCEDURE. THE EVENT WAS TREATED VIA REVISION SURGERY 3 YEARS POSTOPERATIVELY. THE OUTCOME OF THE PATIENT IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00299 | 2022-05-19 | 2016-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "HIP RESURFACING: A SINGLE SURGEON U.S. SERIES WITH MINIMUM TEN-YEAR FOLLOW-UP¿ 2.7 YEARS AFTER A BHR SURGERY, ONE (1) PATIENT HAD AN ACETABULAR FRACTURE AFTER A FALL. REVISION SURGERY WAS PERFORMED TO EXCHANGE THE CUP FOR AN R3 CUP. THE PATIENT OUTCOME AND CURRENT HEALTH STATUS IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00300 | 2022-05-19 | 2016-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "HIP RESURFACING: A SINGLE SURGEON U.S. SERIES WITH MINIMUM TEN-YEAR FOLLOW-UP¿ 4.9 YEARS AFTER A BHR SURGERY, ONE (1) PATIENT HAD PAIN, HIGH LEVELS OF METAL IONS (47.6 PPB OF CO AND 25.2 PPB OF CR), CUP MALPOSITION WITH AN INCLINATION OF 56° AND EFFUSION. REVISION SURGERY WAS PERFORMED TO EXPLANT THE CUP. THE PATIENT OUTCOME AND CURRENT HEALTH STATUS IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00297 | 2022-05-19 | 2016-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 57 YR | Female | IT WAS REPORTED THAT ON LITERATURE REVIEW "HIP RESURFACING: A SINGLE SURGEON U.S. SERIES WITH MINIMUM TEN-YEAR FOLLOW-UP¿ 7.1 YEARS AFTER A BHR SURGERY, ONE (1) PATIENT HAD PAIN, HIGH LEVEL OF METAL IONS (44 PPB OF CO AND 15 PPB OF CR) AND CUP MALPOSITION WITH EXCESSIVE ANTEVERSION. REVISION SURGERY WAS PERFORMED TO EXPLANT BOTH X AND X. COMPETITOR DEVICES WERE IMPLANTED. THE PATIENT OUTCOME AND CURRENT HEALTH STATUS IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00296 | 2022-05-19 | 2016-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 42 YR | Female | IT WAS REPORTED THAT ON LITERATURE REVIEW "HIP RESURFACING: A SINGLE SURGEON U.S. SERIES WITH MINIMUM TEN-YEAR FOLLOW-UP¿ 11.1 YEARS AFTER A BHR SURGERY, ONE (1) PATIENT HAD PAIN, CUP MALPOSITION WITH EXCESSIVE ANTEVERSION. REVISION SURGERY WAS PERFORMED TO EXPLANT THE CUP. THE PATIENT OUTCOME AND CURRENT HEALTH STATUS IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00293 | 2022-05-19 | 2016-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "HIP RESURFACING: A SINGLE SURGEON U.S. SERIES WITH MINIMUM TEN-YEAR FOLLOW-UP¿ 9.3 YEARS AFTER A BHR SURGERY, ONE (1) PATIENT HAD PAIN, HIGH LEVELS OF METAL IONS (57 PPB OF CO AND 18 PPB OF CR), AND CUP MALPOSITION WITH EXCESSIVE ANTEVERSION. REVISION SURGERY WAS PERFORMED TO EXPLANT THE CUP. THE PATIENT OUTCOME AND CURRENT HEALTH STATUS IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00290 | 2022-05-19 | 2016-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "HIP RESURFACING: A SINGLE SURGEON U.S. SERIES WITH MINIMUM TEN-YEAR FOLLOW-UP¿ TWO YEARS AFTER A BHR SURGERY, ONE (1) PATIENT HAD PAIN, CUP LOOSENING AND MALPOSITION WITH EXCESSIVE ANTEVERSION, AND 3.5 PPB OF CO. REVISION SURGERY WAS PERFORMED TO EXCHANGE THE CUP FOR ANOTHER DYSPLASIA CUP FROM S+N. THE PATIENT OUTCOME AND CURRENT HEALTH STATUS IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00280 | 2022-05-19 | 2022-01-25 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "MIDTERM RESULTS OF THE BIRMINGHAM HIP RESURFACING: A SINGLE-SURGEON SERIES", 1 PATIENT SUFFERED FROM AN UNSEATED CUP IN THE EARLY POSTOPERATIVE PERIOD, WHICH RESULTED IN PAIN AND AN INCREASE IN LEG LENGTH AFTER A BHR HIP RESURFACING PROCEDURE. THE EVENT WAS TREATED VIA REVISION SURGERY. THE OUTCOME OF THE PATIENT IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00281 | 2022-05-19 | 2022-01-25 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "MIDTERM RESULTS OF THE BIRMINGHAM HIP RESURFACING: A SINGLE-SURGEON SERIES", 2 PATIENTS SUFFERED FROM FEMORAL LOOSENING AFTER A BHR HIP RESURFACING PROCEDURE. THE EVENT WAS TREATED VIA REVISION SURGERY. THE OUTCOME OF THE PATIENTS IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00285 | 2022-05-19 | 2022-01-25 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "MIDTERM RESULTS OF THE BIRMINGHAM HIP RESURFACING: A SINGLE-SURGEON SERIES", 11 PATIENTS HAD ELEVATED METAL ION LEVELS (MORE THAN 7 MCG/L) AFTER A BHR HIP RESURFACING PROCEDURE. IT WAS NOT REPORTED HOW THE ADVERSE EVENT WAS TREATED. THE OUTCOME OF THE PATIENTS IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00283 | 2022-05-19 | 2022-01-25 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "MIDTERM RESULTS OF THE BIRMINGHAM HIP RESURFACING: A SINGLE-SURGEON SERIES", 20 PATIENTS WERE DIAGNOSED WITH ALTR ON AN MRI AFTER A BHR HIP RESURFACING PROCEDURE. IT WAS NOT REPORTED HOW THE ADVERSE EVENT WAS TREATED. THE OUTCOME OF THE PATIENTS IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00291 | 2022-05-19 | 2016-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 47 YR | Female | IT WAS REPORTED THAT ON LITERATURE REVIEW "HIP RESURFACING: A SINGLE SURGEON U.S. SERIES WITH MINIMUM TEN-YEAR FOLLOW-UP¿ 8.3 YEARS AFTER A BHR SURGERY, ONE (1) PATIENT HAD PAIN, HIGH LEVELS OF METAL IONS (56.7 PPB OF CO AND 29.8 PPB OF CR), CUP MALPOSITION WITH EXCESSIVE ANTEVERSION AND EFFUSION. REVISION SURGERY WAS PERFORMED TO EXPLANT BOTH IMPLANTS. COMPETITOR DEVICES WERE IMPLANTED. THE PATIENT OUTCOME AND CURRENT HEALTH STATUS IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00289 | 2022-05-19 | 2016-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "HIP RESURFACING: A SINGLE SURGEON U.S. SERIES WITH MINIMUM TEN-YEAR FOLLOW-UP¿ FOUR YEARS AFTER A BHR SURGERY, ONE PATIENT HAD PAIN AND CUP MALPOSITION WITH AN INCLINATION OF 61 AND 2.2 PPB OF CO. REVISION SURGERY WAS PERFORMED TO EXCHANGE THE CUP FOR ANOTHER DYSPLASIA CUP FROM S+N. THE PATIENT OUTCOME AND CURRENT HEALTH STATUS IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00286 | 2022-05-19 | 2022-01-25 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "MIDTERM RESULTS OF THE BIRMINGHAM HIP RESURFACING: A SINGLE-SURGEON SERIES", 11 PATIENTS HAD ELEVATED METAL ION LEVELS (MORE THAN 7 MCG/L) AFTER A BHR HIP RESURFACING PROCEDURE. IT WAS NOT REPORTED HOW THE ADVERSE EVENT WAS TREATED. THE OUTCOME OF THE PATIENTS IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00282 | 2022-05-19 | 2022-01-25 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "MIDTERM RESULTS OF THE BIRMINGHAM HIP RESURFACING: A SINGLE-SURGEON SERIES", 20 PATIENTS WERE DIAGNOSED WITH ALTR ON AN MRI AFTER A BHR HIP RESURFACING PROCEDURE. IT WAS NOT REPORTED HOW THE ADVERSE EVENT WAS TREATED. THE OUTCOME OF THE PATIENTS IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 5 | 2022-05
3005975929-2022-00320 | 2022-05-23 | 2021-11-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL |  | 60 YR | Male | IT WAS REPORTED THAT, AFTER A LEFT BHR ON (B)(6) 2009 DUE TO LEFT OSTEOARTHRITIS. PLAINTIFF EXPERIENCES PAIN AND ELEVATED METAL IONS. PLAINTIFF UNDERWENT REVISION SURGERY ON LEFT HIP ON (B)(6) 2021. IT IS UNKNOWN THE STATUS OF THE PATIENT. |  | 2022 | 5 | 2022-05
3005975929-2022-00380 | 2022-06-03 | 2020-10-19 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 61 YR | Male | IT WAS REPORTED THAT, AFTER AN UNSPECIFIED METAL ON METAL HIP CONSTRUCT HAD BEEN IMPLANTED ON THE PATIENT¿S HIP ON (B)(6) 2009, THE PATIENT SUSTAINED AN UNSPECIFIED INJURY AND REQUIRED A REVISION SURGERY ON (B)(6) 2020. THE IDENTITY OF THE EXPLANTED DEVICES AND THE PROSTHESIS IMPLANTED IN EXCHANGE IS UNKNOWN. THE PATIENT¿S OUTCOME IS UNKNOWN. |  | 2022 | 6 | 2022-06
3005975929-2022-00378 | 2022-06-03 | 2018-10-31 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 77 YR | Female | IT WAS REPORTED THAT, AFTER AN UNSPECIFIED METAL ON METAL HIP CONSTRUCT HAD BEEN IMPLANTED ON THE PATIENT¿S HIP ON (B)(6) 2013, THE PATIENT SUSTAINED AN UNSPECIFIED INJURY AND REQUIRED A REVISION SURGERY ON (B)(6) 2018. THE IDENTITY OF THE EXPLANTED DEVICES AND THE PROSTHESIS IMPLANTED IN EXCHANGE IS UNKNOWN. THE PATIENT¿S OUTCOME IS UNKNOWN. |  | 2022 | 6 | 2022-06
3005975929-2022-00377 | 2022-06-03 | 2020-09-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) SLEEVED MODULAR | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 65 YR | Female | IT WAS REPORTED THAT THE PLAINTIFF UNDERWENT A FIRST REVISION SURGERY OF THE RIGHT HIP ON (B)(6) 2020 DUE TO ELEVATED METAL ION LEVELS IN BLOOD. THE OPERATION WAS COMPLETED WITHOUT COMPLICATIONS. AFTER THIS REVISION, THE PATIENT SUFFERED FROM A SECOND REVISION SURGERY (UNSPECIFIED IMPLANTS FROM AN UNSPECIFIED MANUFACTURER) AND KEEPS SUFFERING FROM CHRONIC PAIN AND PERMANENT REDUCTION IN MOBILITY AND DAY-TO-DAY ACTIVITIES. |  | 2022 | 6 | 2022-06
3005975929-2022-00379 | 2022-06-03 | 2021-03-03 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 53 YR | Male | IT WAS REPORTED THAT, AFTER AN UNSPECIFIED METAL ON METAL HIP CONSTRUCT HAD BEEN IMPLANTED ON THE PATIENT¿S HIP ON (B)(6) 2011, THE PATIENT SUSTAINED AN UNSPECIFIED INJURY AND REQUIRED A REVISION SURGERY ON (B)(6) 2021. THE IDENTITY OF THE EXPLANTED DEVICES AND THE PROSTHESIS IMPLANTED IN EXCHANGE IS UNKNOWN. THE PATIENT¿S OUTCOME IS UNKNOWN. |  | 2022 | 6 | 2022-06
3005975929-2022-00391 | 2022-06-10 | 2019-10-08 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Female | BILATERAL PATIENT. IT WAS REPORTED THAT AFTER A LEFT BHR PERFORMED IN MARCH 2013, THE PATIENT EXPERIENCED UNSPECIFIED SYMPTOMS THAT MADE A REVISION SURGERY NECESSARY, WHICH WAS PERFORMED ON (B)(6) 2019. THE PATIENT ALSO WENT TO A RIGHT BHR PERFORMED ON (B)(6) 2009. THE IDENTITY OF THE EXPLANTED DEVICES AND THE ONES IMPLANTED IN EXCHANGE IS UNKNOWN AT THIS TIME. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 6 | 2022-06
3005975929-2022-00392 | 2022-06-10 | 2020-10-27 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Female | IT WAS REPORTED THAT AFTER A BHR PERFORMED ON (B)(6) 2020, THE PATIENT EXPERIENCED UNSPECIFIED SYMPTOMS THAT MADE A REVISION SURGERY NECESSARY, WHICH WAS PERFORMED ON (B)(6) 2020. THE LIMB SIDE AND THE IDENTITY OF THE EXPLANTED DEVICES AND THE ONES IMPLANTED IN EXCHANGE IS UNKNOWN AT THIS TIME. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 6 | 2022-06
3005975929-2022-00393 | 2022-06-10 | 2021-05-03 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Female | *US LEGAL* IT WAS REPORTED THAT AFTER A BHR PERFORMED ON (B)(6) 2009, THE PATIENT EXPERIENCED UNSPECIFIED SYMPTOMS THAT MADE A REVISION SURGERY NECESSARY, WHICH WAS PERFORMED ON (B)(6) 2021. THE LIMB SIDE AND THE IDENTITY OF THE EXPLANTED DEVICES AND THE ONES IMPLANTED IN EXCHANGE IS UNKNOWN AT THIS TIME. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 6 | 2022-06
3005975929-2022-00394 | 2022-06-10 | 2018-10-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT AFTER A BHR PERFORMED ON (B)(6) 2012, THE PATIENT EXPERIENCED UNSPECIFIED SYMPTOMS THAT MADE A REVISION SURGERY NECESSARY, WHICH WAS PERFORMED ON (B)(6) 2018. THE LIMB SIDE AND THE IDENTITY OF THE EXPLANTED DEVICES AND THE ONES IMPLANTED IN EXCHANGE IS UNKNOWN AT THIS TIME. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 6 | 2022-06
3005975929-2022-00389 | 2022-06-10 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 69 YR | Female | BILATERAL PATIENT. IT WAS REPORTED THAT, AFTER AN UNSPECIFIED BHR PROSTHESIS CONSTRUCT HAD BEEN IMPLANTED ON THE PATIENT¿S LEFT HIP IN OR AROUND 2007, THE PATIENT HAS EXPERIENCED ONGOING PAIN AND OTHER UNSPECIFIED SYMPTOMS CONSISTENT WITH METAL-ON-METAL FAILURE. THE PATIENT¿S COBALT LEVELS WERE AROUND THE 50¿S (PRESUMABLY PPB) WHEN A REVISION SURGERY OF THE BHR RIGHT HIP IMPLANTS WAS PERFORMED ON (B)(6) 2020. THE CONTRIBUTION OF THIS HIP TO THE REPORTED COBALT LEVELS COULD NOT BE DISCARDED AT THIS MOMENT. IT IS UNKNOWN IF A REVISION PROCEDURE HAS ALREADY BEEN ADVISED OR PERFORMED BY A HEALTH CARE PROFESSIONAL. THE PATIENT¿S HEALTH STATUS IS NOT KNOWN. |  | 2022 | 6 | 2022-06
3005975929-2022-00388 | 2022-06-10 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 69 YR | Female | BILATERAL PATIENT. IT WAS REPORTED THAT, AFTER AN UNSPECIFIED BHR PROSTHESIS CONSTRUCT HAD BEEN IMPLANTED ON THE PATIENT¿S LEFT HIP IN OR AROUND 2007, THE PATIENT HAS EXPERIENCED ONGOING PAIN AND OTHER UNSPECIFIED SYMPTOMS CONSISTENT WITH METAL-ON-METAL FAILURE. THE PATIENT¿S COBALT LEVELS WERE AROUND THE 50¿S (PRESUMABLY PPB) WHEN A REVISION SURGERY OF THE BHR RIGHT HIP IMPLANTS WAS PERFORMED ON (B)(6) 2020, THE CONTRIBUTION OF THIS HIP TO THE REPORTED COBALT LEVELS COULD NOT BE DISCARDED AT THIS MOMENT. IT IS UNKNOWN IF A REVISION PROCEDURE HAS ALREADY BEEN ADVISED OR PERFORMED BY A HEALTH CARE PROFESSIONAL. THE PATIENT¿S HEALTH STATUS IS NOT KNOWN. |  | 2022 | 6 | 2022-06
3005975929-2022-00390 | 2022-06-10 | 2019-10-18 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Male | IT WAS REPORTED THAT AFTER A BHR PERFORMED ON (B)(6) 2017, THE PATIENT EXPERIENCED UNSPECIFIED SYMPTOMS THAT MADE A REVISION SURGERY NECESSARY, WHICH WAS PERFORMED ON (B)(6) 2019. THE LIMB SIDE AND THE IDENTITY OF THE EXPLANTED DEVICES AND THE ONES IMPLANTED IN EXCHANGE IS UNKNOWN AT THIS TIME. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 6 | 2022-06
3005975929-2022-00398 | 2022-06-12 | 2021-10-28 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, HEMI-, TRUNNION-BEARING, FEMORAL, METAL/POLYACETAL |  |  | Male | IT WAS REPORTED THAT, AFTER A BHR RESURFACING CONSTRUCT HAD BEEN IMPLANTED ON THE PATIENT¿S LEFT HIP ON (B)(6) 2009, THE PATIENT EXPERIENCED HIP PAIN, SWELLING, METALLOSIS, DIFFICULTY WALKING, SOFT TISSUE DAMAGE AND METAL-ON-METAL REACTION WITHIN THE LEFT HIP JOINT. A REVISION SURGERY WAS CONDUCTED ON 28-OCT-2021 TO ADDRESS THIS ADVERSE EVENT. THE IDENTITY OF THE EXPLANTED DEVICES AND THE ONES IMPLANTED IN EXCHANGE ARE NOT KNOWN. PATIENT¿S CURRENT HEALTH STATUS IS NOT KNOWN. ALL AVAILABLE INFORMATION HAS BEEN DISCLOSED. IF ADDITIONAL INFORMATION SHOULD BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. |  | 2022 | 6 | 2022-06
3005975929-2022-00401 | 2022-06-19 | 2021-10-28 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Male | IT WAS REPORTED THAT, AFTER A BHR RESURFACING CONSTRUCT HAD BEEN IMPLANTED ON THE PATIENT¿S LEFT HIP ON (B)(6) 2009, THE PATIENT EXPERIENCED HIP PAIN, SWELLING, METALLOSIS, DIFFICULTY WALKING, SOFT TISSUE DAMAGE AND METAL-ON-METAL REACTION WITHIN THE LEFT HIP JOINT. A REVISION SURGERY WAS CONDUCTED ON (B)(6) 2021 TO ADDRESS THIS ADVERSE EVENT. THE IDENTITY OF THE EXPLANTED DEVICES AND THE ONES IMPLANTED IN EXCHANGE ARE NOT KNOWN. PATIENT¿S CURRENT HEALTH STATUS IS NOT KNOWN. |  | 2022 | 6 | 2022-06
3005975929-2022-00402 | 2022-06-20 | 2021-07-21 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT, ON LITERATURE REVIEW "RESULTS OF THE FIRST U.S FDA - APPROVED HIP RESURFACING DEVICE AT 10 - YEAR FOLLOW - UP." THE NUMBER OF PATIENTS WHO INITIALLY UNDERWENT PRIMARY IMPLANTATION OF A BHR RESURFACING CONSTRUCT AND THAT EXPERIENCED POST-OPERATIVE COBALT AND/OR CHROMIUM LEVELS IN BLOOD GREATER THAN SEVEN (7) PARTS PER BILLION REMAINED SIMILAR OVER TIME: SEVEN (7) OUT OF TWO HUNDRED FIVE (205) SHOWED EVIDENCE OF THESE LEVELS AT ONE (1) YEAR AFTER THE INITIAL SURGERY, AND NINE (9) PATIENTS OUT OF TWO HUNDRED FOUR (204) SHOWED EVIDENCE DURING THE TEN (10) POST-OPERATIVE YEAR FOLLOW UP. ALTHOUGH IT IS UNKNOWN HOW THIS EVENT WAS ADDRESSED ON EACH PATIENT, A REVISION SURGERY WAS NOT REQUIRED. FURTHER INFORMATION WAS NOT PROVIDED. |  | 2022 | 6 | 2022-06
3005975929-2022-00404 | 2022-06-20 | 2021-07-21 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT, ON LITERATURE REVIEW "RESULTS OF THE FIRST U.S FDA - APPROVED HIP RESUFACING DEVICE AT 10 - YEAR FOLLOW - UP. ", OSTEOLYSIS WAS IDENTIFIED IN SEVEN (7) OUT OF TWO HUNDRED EIGHTEEN (218) ACETABULAR COMPONENTS IMPLANTED ON PATIENTS WHO UNDERWENT PRIMARY SURGERY WITH A BHR RESURFACING CONSTRUCT. ALTHOUGH IT IS UNKNOWN HOW THIS EVENT WAS ADDRESSED ON EACH PATIENT, A REVISION SURGERY WAS NOT REQUIRED. FURTHER INFORMATION WAS NOT PROVIDED. |  | 2022 | 6 | 2022-06
3005975929-2022-00403 | 2022-06-20 | 2021-07-21 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT, ON LITERATURE REVIEW "RESULTS OF THE FIRST U.S FDA-APPROVED HIP RESUFACING DEVICE AT 10 YEAR FOLLOW-UP.", THE NUMBER OF PATIENTS WHO INITIALLY UNDERWENT PRIMARY IMPLANTATION OF A BHR RESURFACING CONSTRUCT AND THAT EXPERIENCED POST-OPERATIVE COBALT AND/OR CHROMIUM LEVELS IN BLOOD GREATER THAN SEVEN (7) PARTS PER BILLION REMAINED SIMILAR OVER TIME: SEVEN (7) OUT OF TWO HUNDRED FIVE (205) SHOWED EVIDENCE OF THESE LEVELS AT ONE (1) YEAR AFTER THE INITIAL SURGERY, AND NINE (9) PATIENTS OUT OF TWO HUNDRED FOUR (204) SHOWED EVIDENCE DURING THE TEN (10) POST-OPERATIVE YEAR FOLLOW UP. ALTHOUGH IT IS UNKNOWN HOW THIS EVENT WAS ADDRESSED ON EACH PATIENT, A REVISION SURGERY WAS NOT REQUIRED. FURTHER INFORMATION WAS NOT PROVIDED. |  | 2022 | 6 | 2022-06
3005975929-2022-00405 | 2022-06-20 | 2021-07-21 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT, ON LITERATURE REVIEW "RESULTS OF THE FIRST U.S FDA - APPROVED HIP RESURFACING DEVICE AT 10 - YEAR FOLLOW - UP." OSTEOLYSIS WAS IDENTIFIED IN FOUR (4) OUT OF TWO HUNDRED EIGHTEEN (218) FEMORAL COMPONENTS IMPLANTED ON PATIENTS WHO UNDERWENT PRIMARY SURGERY WITH A BHR RESURFACING CONSTRUCT. ALTHOUGH IT IS UNKNOWN HOW THIS EVENT WAS ADDRESSED ON EACH PATIENT, A REVISION SURGERY WAS NOT REQUIRED. FURTHER INFORMATION WAS NOT PROVIDED. |  | 2022 | 6 | 2022-06
1020279-2022-03120 | 2022-06-27 | 2022-06-07 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW, INC. | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT, AFTER A BHR SURGERY HAD BEEN PERFORMED AROUND 16 YEARS AGO, THE RESURFACING HEAD LOOSENED. A REVISION SURGERY WAS PERFORMED TO ADDRESS THIS EVENT. INTRAOPERATIVELY, IT WAS ALSO NOTICED SOME BONE LOSS IN THE NECK. BOTH COMPONENTS WERE EXPLANTED TO CONVERT TO A THA WITH AN R3-ANTHOLOGY SYSTEM. PATIENT IS RECOVERING. |  | 2022 | 6 | 2022-06
3005975929-2022-00418 | 2022-07-10 | 2007-07-27 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 43 YR | Male | IT WAS REPORTED THAT, AFTER A BHR RESURFACING PRIMARY SURGERY HAD BEEN PERFORMED ON THE PATIENT¿S RIGHT HIP ON (B)(6) 2007, THE PATIENT REQUIRED AN IMMEDIATE WASHOUT PROCEDURE DUE TO A POSTOPERATIVE HEMATOMA. ADDITIONAL DETAILS REGARDING THIS INTERVENTION ARE NOT KNOWN. |  | 2022 | 7 | 2022-07
3005975929-2022-00420 | 2022-07-21 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | BIRMINGHAM HIP RESURFACING (BHR) CUP | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | ACCORDING TO DATA OBTAINED FROM THE SWEDISH HIP ARTHROPLASTY REGISTER - ANNUAL REPORT 2019 - A TOTAL OF 1277 PRIMARY RESURFACING PROCEDURES WITH BHR SYSTEM WERE STUDIED IN A PERIOD BETWEEN 1999 AND 2019. FROM THIS COHORT, ONE HUNDRED TWENTY FOUR (124) PATIENTS REQUIRED A REVISION SURGERY OF THE DUE TO UNSPECIFIED REASONS. AS THE INFORMATION WAS EXTRACTED FROM THE SWEDISH REGISTRY, NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 7 | 2022-07
3005975929-2022-00426 | 2022-07-22 | 2022-03-29 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Male | IT WAS REPORTED THAT THE PLAINTIFF UNDERWENT A RIGHT HIP REVISION SURGERY ON (B)(6) 2022 DUE TO ELEVATED COBALT AND CHROME LEVELS IN BLOOD, AND LOW-GRADE SYMPTOMS. BOTH METAL COMPONENTS WERE EXPLANTED DURING THE REVISION AND REPLACED WITH THA SYSTEM FROM A COMPETITOR´S COMPANY. IT IS UNKNOWN WHEN THE PRIMARY RIGHT HIP BHR RESURFACING SYSTEM WAS IMPLANTED. |  | 2022 | 7 | 2022-07
3005975929-2022-00425 | 2022-07-22 | 2022-03-29 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Male | IT WAS REPORTED THAT THE PLAINTIFF UNDERWENT A RIGHT HIP REVISION SURGERY ON (B)(6) 2022 DUE TO ELEVATED COBALT AND CHROME LEVELS IN BLOOD, AND LOW-GRADE SYMPTOMS. BOTH METAL COMPONENTS WERE EXPLANTED DURING THE REVISION AND REPLACED WITH THA SYSTEM FROM A COMPETITOR´S COMPANY. IT IS UNKNOWN WHEN THE PRIMARY RIGHT HIP BHR RESURFACING SYSTEM WAS IMPLANTED. |  | 2022 | 7 | 2022-07
3005975929-2022-00424 | 2022-07-22 | 2022-02-22 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Female | IT WAS REPORTED THAT, AFTER A BHR RESURFACING CONSTRUCT HAD BEEN IMPLANTED ON THE PATIENT¿S LEFT HIP ON AN UNSPECIFIED DATE, THE PATIENT EXPERIENCED ELEVATED METAL ION LEVELS IN SERUM. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2022 TO TREAT THIS ADVERSE EVENT. DURING THIS PROCEDURE, BOTH THE BHR RESURFACING HEAD AND SHELL WERE EXPLANTED AND REPLACED WITH A THR SYSTEM. THE PATIENT¿S OUTCOME IS UNKNOWN. |  | 2022 | 7 | 2022-07
3005975929-2022-00423 | 2022-07-22 | 2022-02-22 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Female | IT WAS REPORTED THAT, AFTER A BHR RESURFACING CONSTRUCT HAD BEEN IMPLANTED ON THE PATIENT¿S LEFT HIP ON AN UNSPECIFIED DATE, THE PATIENT EXPERIENCED ELEVATED METAL ION LEVELS IN SERUM. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2022 TO TREAT THIS ADVERSE EVENT. DURING THIS PROCEDURE, BOTH THE BHR RESURFACING HEAD AND SHELL WERE EXPLANTED AND REPLACED WITH A THR SYSTEM. THE PATIENT¿S OUTCOME IS UNKNOWN. |  | 2022 | 7 | 2022-07
3005975929-2022-00430 | 2022-07-24 | 2022-01-27 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 69 YR | Male | IT WAS REPORTED THAT THE PLAINTIFF UNDERWENT A LEFT TOTAL HIP ARTHROPLASTY METAL-ON-METAL ARTICULATION APPROXIMATELY 10 YEARS AGO. THE PATIENT EXPERIENCED SYMPTOMATIC MECHANICAL SOUNDS, SQUEAKING AND ELEVATED METAL COBALT AND CHROME LEVELS. THE PATIENT HAD A REVISION SURGERY ON (B)(6) 2022 TO TREAT THESE ADVERSE EVENTS. DURING THIS PROCEDURE, THE BHR MODULAR HEAD WAS EXPLANTED AND REPLACED WITH AN OXINIUM FEMORAL HEAD AND A COMPETITORS LINER. THE CURRENT HEALTH STATUS OF THE PATIENT IS UNKNOWN. |  | 2022 | 7 | 2022-07
3005975929-2022-00432 | 2022-07-26 | 2022-01-27 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Male | IT WAS REPORTED THAT THE PLAINTIFF UNDERWENT A LEFT TOTAL HIP ARTHROPLASTY METAL-ON-METAL ARTICULATION APPROXIMATELY 10 YEARS AGO. THE PATIENT EXPERIENCED SYMPTOMATIC MECHANICAL SOUNDS, SQUEAKING AND ELEVATED METAL COBALT AND CHROME LEVELS. THE PATIENT HAD A REVISION SURGERY ON (B)(6) 2022 TO TREAT THESE ADVERSE EVENTS. DURING THIS PROCEDURE, THE BHR MODULAR HEAD WAS EXPLANTED AND REPLACED WITH AN OXINIUM FEMORAL HEAD AND A COMPETITORS LINER. THE CURRENT HEALTH STATUS OF THE PATIENT IS UNKNOWN. |  | 2022 | 7 | 2022-07
3005975929-2022-00439 | 2022-07-28 | 2021-11-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | BIRMINGHAM HIP RESURFACING (BHR) CUP | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 71 YR | Male | IT WAS REPORTED THAT THE PATIENT UNDERWENT A REVISION SURGERY ON (B)(6) 2022 DUE TO A LEFT HIP PERIPROSTHETIC JOINT INFECTION. THE REVISION SURGERY CONSISTED IN: THE EXPLANTATION OF THE BHR-THA SYSTEM AND PLACEMENT OF AN ARTICULATING ANTIBIOTIC SPACER WITH ACETABULAR AND FEMORAL COMPONENTS. IRRIGATION AND DEBRIDEMENT PROCEDURE OF THE LEFT HIP SKIN, SUBCUTANEOUS TISSUE, MUSCLE, AND BONE. LEFT HIP PROXIMAL FEMUR AND EXTENDED TROCHANTERIC OSTEOTOMY. DURING THE REVISION SURGERY, PURULENCE WAS SEEN, AS WELL AS TRUNNION CORROSION. THE PATIENT TOLERATED THE PROCEDURE WELL. THIS PATIENT HAD A PREVIOUS LEFT HIP IMPLANT INFECTION IN (B)(6) 2021, WHICH WAS TREATED WITH AN IRRIGATION AND DEBRIDEMENT PROCEDURE AND IV ANTIBIOTIC THERAPY. THE PRIMARY LEFT HIP BHR THA PROCEDURE WAS PERFORMED ON (B)(6) 2010. |  | 2022 | 7 | 2022-07
3005975929-2022-00440 | 2022-07-28 | 2021-11-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 71 YR | Male | IT WAS REPORTED THAT THE PLAINTIFF UNDERWENT A REVISION SURGERY ON (B)(6) 2022 DUE TO A LEFT HIP PERIPROSTHETIC JOINT INFECTION. THE REVISION SURGERY CONSISTED IN: 1) THE EXPLANTATION OF THE BHR-THA SYSTEM AND PLACEMENT OF AN ARTICULATING ANTIBIOTIC SPACER WITH ACETABULAR AND FEMORAL COMPONENTS. 2) IRRIGATION AND DEBRIDEMENT PROCEDURE OF THE LEFT HIP SKIN, SUBCUTANEOUS TISSUE, MUSCLE, AND BONE. 3) LEFT HIP PROXIMAL FEMUR AND EXTENDED TROCHANTERIC OSTEOTOMY. DURING THE REVISION SURGERY, PURULENCE WAS SEEN, AS WELL AS TRUNNION CORROSION. THE PATIENT TOLERATED THE PROCEDURE WELL. THIS PATIENT HAD A PREVIOUS LEFT HIP IMPLANT INFECTION IN (B)(6) 2021, WHICH WAS TREATED WITH AN IRRIGATION AND DEBRIDEMENT PROCEDURE AND IV ANTIBIOTIC THERAPY. THE PRIMARY LEFT HIP BHR THA PROCEDURE WAS PERFORMED ON (B)(6) 2010. |  | 2022 | 7 | 2022-07
3005975929-2022-00450 | 2022-08-05 | 1901-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Female | IT WAS REPORTED THAT, AFTER A PRIMARY BHR SURGERY WAS PERFORMED ON (B)(6) 2007, THE PATIENT EXPERIENCED AN UNSPECIFIED ADVERSE EVENT THAT REQUIRES REVISION SURGERY. IT IS UNKNOWN IF THE REVISION SURGERY HAS ALREADY BEEN SCHEDULED. PATIENT'S CURRENT HEALTH STATUS IS UNKNOWN. |  | 2022 | 8 | 2022-08
3005975929-2022-00455 | 2022-08-17 | 2022-07-11 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, FEMORAL COMPONENT, CEMENTED, METAL |  |  | Unknown | IT WAS REPORTED THAT AFTER A BHR RESURFACING CONSTRUCT HAD BEEN IMPLANTED ON (B)(6) 2010, THE PATIENT EXPERIENCED METALLOSIS WITH A BIG PSEUDO-TUMOR. THIS ADVERSE EVENT WAS TREATED BY A REVISION SURGERY ON (B)(6) 2022, IN WHICH THE BHR FEMORAL HEAD WAS EXPLANTED. FURTHER INFORMATION IS UNKNOWN AT THE TIME. |  | 2022 | 8 | 2022-08
3005975929-2022-00456 | 2022-08-17 | 2022-07-11 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, FEMORAL COMPONENT, CEMENTED, METAL |  | 76 YR | Unknown | IT WAS REPORTED THAT AFTER A BHR RESURFACING CONSTRUCT HAD BEEN IMPLANTED ON (B)(6) 2010, THE PATIENT EXPERIENCED METALLOSIS WITH A BIG PSEUDO-TUMOR. THIS ADVERSE EVENT WAS TREATED BY A REVISION SURGERY ON (B)(6) 2022, IN WHICH THE BHR FEMORAL HEAD WAS EXPLANTED. FURTHER INFORMATION IS UNKNOWN AT THE TIME. |  | 2022 | 8 | 2022-08
3005975929-2022-00461 | 2022-08-18 | 2013-09-17 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 57 YR | Female | IT WAS REPORTED THAT, AFTER A BHR RESURFACING CONSTRUCT HAD BEEN IMPLANTED ON THE PATIENT¿S LEFT HIP IN 2009, THE PATIENT DEVELOPED HIP PAIN. DUE TO CONCERNS OF METALLOSIS, A REVISION SURGERY WAS CONDUCTED ON (B)(6) 2013 TO TREAT THIS ADVERSE EVENT. DURING THIS PROCEDURE, THE BHR RESURFACING HEAD WAS EXPLANTED AND REPLACED WITH A COMPETITORS THR SYSTEM. INTRAOPERATIVELY, MILD AMOUNTS OF INFLAMMATORY TISSUE WERE NOTICED. THE PATIENT TOLERATED WELL THE PROCEDURE AND WAS TRANSFERRED TO THE PACU UNIT IN STABLE CONDITION. ALL AVAILABLE INFORMATION HAS BEEN DISCLOSED. IF ADDITIONAL INFORMATION SHOULD BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. |  | 2022 | 8 | 2022-08
3005975929-2022-00458 | 2022-08-18 | 2018-06-26 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 61 YR | Female | IT WAS REPORTED THAT, AFTER A BHR RESURFACING CONSTRUCT HAD BEEN IMPLANTED ON THE PATIENT¿S RIGHT HIP, THE PATIENT DEVELOPED METALLOSIS AND INFLAMMATORY CONDITIONS. A REVISION SURGERY WAS CONDUCTED ON (B)(6) 2018 TO TREAT THIS ADVERSE EVENT. DURING THIS PROCEDURE, THE BHR ACETABULAR CUP WAS EXPLANTED AND REPLACED WITH A COMPETITORS THR SYSTEM. INTRAOPERATIVELY, SOME METALLOSIS DEBRIS WAS DEBRIDED. THE PATIENT WAS TRANSFERRED TO THE PACU UNIT IN STABLE CONDITION. ALL AVAILABLE INFORMATION HAS BEEN DISCLOSED. IF ADDITIONAL INFORMATION SHOULD BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. |  | 2022 | 8 | 2022-08
3005975929-2022-00459 | 2022-08-18 | 2018-06-26 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 61 YR | Female | IT WAS REPORTED THAT, AFTER A BHR RESURFACING CONSTRUCT HAD BEEN IMPLANTED ON THE PATIENT¿S RIGHT HIP, THE PATIENT DEVELOPED METALLOSIS AND INFLAMMATORY CONDITIONS. A REVISION SURGERY WAS CONDUCTED ON (B)(6) 2018 TO TREAT THIS ADVERSE EVENT. DURING THIS PROCEDURE, THE BHR RESURFACING HEAD WAS EXPLANTED AND REPLACED WITH A COMPETITORS THR SYSTEM. INTRAOPERATIVELY, SOME METALLOSIS DEBRIS WAS DEBRIDED. THE PATIENT WAS TRANSFERRED TO THE PACU UNIT IN STABLE CONDITION. ALL AVAILABLE INFORMATION HAS BEEN DISCLOSED. IF ADDITIONAL INFORMATION SHOULD BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. |  | 2022 | 8 | 2022-08
3005975929-2022-00460 | 2022-08-18 | 2013-09-17 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 57 YR | Female | IT WAS REPORTED THAT, AFTER A BHR RESURFACING CONSTRUCT HAD BEEN IMPLANTED ON THE PATIENT¿S LEFT HIP IN 2009, THE PATIENT DEVELOPED HIP PAIN. DUE TO CONCERNS OF METALLOSIS, A REVISION SURGERY WAS CONDUCTED ON (B)(6) 2013 TO TREAT THIS ADVERSE EVENT. DURING THIS PROCEDURE, THE BHR ACETABULAR CUP WAS EXPLANTED AND REPLACED WITH A COMPETITORS THR SYSTEM. INTRAOPERATIVELY, MILD AMOUNTS OF INFLAMMATORY TISSUE WERE NOTICED. THE PATIENT TOLERATED WELL THE PROCEDURE AND WAS TRANSFERRED TO THE PACU UNIT IN STABLE CONDITION. ALL AVAILABLE INFORMATION HAS BEEN DISCLOSED. IF ADDITIONAL INFORMATION SHOULD BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. |  | 2022 | 8 | 2022-08
3005975929-2022-00469 | 2022-09-08 | 2022-08-17 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT, AFTER BHR HIP RESURFACING WAS PERFORMED ON AN UNKNOWN DATE, THE PATIENT EXPERIENCED SEROMA. THIS EVENT WAS TREATED WITH REVISION SURGERY ON (B)(6) 2022, DURING WHICH THE HIP WAS WASHED OUT AND THE HEAD WAS EXCHANGED. PATIENT'S CURRENT HEALTH STATUS IS UNKNOWN. |  | 2022 | 9 | 2022-09
3005975929-2022-00473 | 2022-10-13 | 2013-03-25 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Female | IT WAS REPORTED THAT, AFTER A LEFT BHR RESURFACING CONSTRUCT HAD BEEN IMPLANTED IN 2010, THE PATIENT REQUIRED A REVISION SURGERY ON (B)(6) 2013 DUE TO INCREASING PAIN AND FAILURE OF THE DEVICE. THE IDENTITY OF THE EXPLANTED DEVICES AND THE ONES IMPLANTED IN EXCHANGE REMAINS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2022 | 10 | 2022-10
3005975929-2022-00479 | 2022-10-26 | 2019-03-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 54 YR | Male | IT WAS REPORTED THAT, AFTER A LEFT BHR CONSTRUCT HAD BEEN IMPLANTED ON (B)(6) 2014, THE PLAINTIFF REPORTED SIGNIFICANT UNSPECIFIED RISING LEVELS OF COBALT AND CHROMIUM IN HIS BODY. THE CONTRIBUTION OF THE METAL-ON-METAL IMPLANTS FROM THE LEFT HIP TO THE RISING LEVELS OF COBALT AND CHROMIUM CANNOT BE DISCARDED. THE PLAINTIFF¿S OUTCOME IS UNKNOWN. |  | 2022 | 10 | 2022-10
3005975929-2022-00480 | 2022-10-26 | 2019-03-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 54 YR | Male | IT WAS REPORTED THAT, AFTER A LEFT BHR CONSTRUCT HAD BEEN IMPLANTED ON (B)(6) 2014, THE PLAINTIFF REPORTED SIGNIFICANT UNSPECIFIED RISING LEVELS OF COBALT AND CHROMIUM IN HIS BODY. THE CONTRIBUTION OF THE METAL-ON-METAL IMPLANTS FROM THE LEFT HIP TO THE RISING LEVELS OF COBALT AND CHROMIUM CANNOT BE DISCARDED. THE PLAINTIFF¿S OUTCOME IS UNKNOWN. |  | 2022 | 10 | 2022-10
3005975929-2022-00491 | 2022-12-14 | 2008-12-03 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 45 YR | Male | IT WAS REPORTED THAT, AFTER A LEFT HIP BHR RESURFACING CONSTRUCT HAD BEEN IMPLANTED ON AN UNSPECIFIED DATE, THE PATIENT SUSTAINED A DISPLACED PERIPROSTHETIC FEMORAL NECK FRACTURE. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2008 TO CONVERT THE RESURFACING CONSTRUCT TO A TOTAL HIP ARTHROPLASTY. DURING THIS PROCEDURE, THE RESURFACING FEMORAL HEAD WAS EXPLANTED AND REPLACED WITH A BHR MODULAR SYSTEM. INTRAOPERATIVELY, THE ACETABULAR COMPONENT APPEARED TO BE STABLE AND IN EXCELLENT POSITION AND WAS LEFT IN SITU. THE PROCEDURE WAS COMPLETED UNEVENTFULLY. |  | 2022 | 12 | 2022-12
3005975929-2023-00001 | 2023-01-17 | 2022-05-18 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Female | IT WAS REPORTED THAT, AFTER A BHR-THA CONSTRUCT HAD BEEN IMPLANTED ON (B)(6) 2008, THE PLAINTIFF EXPERIENCED PAIN, METALLOSIS, ELEVATED METAL ION LEVELS. THE DEVICES IMPLANTED IN THE PRIMARY SURGERY WERE: BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI, BIRMINGHAM HIP RESURFACING (BHR) CUP, BIRMINGHAM HIP MODULAR HEAD (BHMH) SLEEVE AND A SYNERGY HIP STEM. THEREFORE THE PATIENT UNDERWENT A REVISION SURGERY ON (B)(6) 2022, IN WHICH THE BHR MODULAR HEAD AND THE BHR MODULAR SLEEVE WERE EXPLANTED; WHILE THE BHR DUAL MOBILITY POLY INSERT AND THE OXINIUM FEMORAL HEAD WERE IMPLANTED. INTRAOPERATIVE FINDINGS INCLUDED: METALLOSIS, WEAR OF THE DEVICE, METAL-STAINED FLUID, METAL-STAINED TISSUE, METAL-STAINED DEBRIS, AND DAMAGED SOFT TISSUE. PATIENT CURRENT HEALTH STATUS IS UNKNOWN. ALL AVAILABLE INFORMATION HAS BEEN DISCLOSED. IF ADDITIONAL INFORMATION SHOULD BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. |  | 2023 | 1 | 2023-01
3005975929-2023-00002 | 2023-01-17 | 2022-07-14 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Male | IT WAS REPORTED THAT THE PLAINTIFF UNDERWENT A LEFT HIP REVISION SURGERY ON (B)(6) 2022 DUE TO PAIN AND ELEVATED METAL ION LEVELS. PLAINTIFF¿S REVISION SURGERY RESULTED IN THE FOLLOWING INTRA-OPERATIVE FINDING: METALLOSIS, ADVERSE LOCAL TISSUE REACTION, BLACK TISSUE, TROCHANTERIC FEMUR FRACTURE, MASSIVE DEFECTS IN THE ACETABULUM AND PELVIS CAUSED BY ADVERSE LOCAL TISSUE REACTION, AND COMPLEX RECONSTRUCTION INCLUDING CAGE RECONSTRUCTION, BONE GRAFTS, AND A TROCHANTERIC PLATE. THE CURRENT STATE OF HEALTH OF THE PATIENT IS UNKNOWN. THE PRIMARY LEFT HIP SURGERY WAS PERFORMED ON (B)(6) 2019. ALL AVAILABLE INFORMATION HAS BEEN DISCLOSED. IF ADDITIONAL INFORMATION SHOULD BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. |  | 2023 | 1 | 2023-01
MW5115380 | 2023-03-01 | 2023-02-21 | I, n, j, u, r, y | Voluntary report | SMITH & NEPHEW, INC. | SMITH & NEPHEW BIRMINGHAM HIP RESURFACING FEMORAL HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED |  | 67 YR | Male | PT PRESENTED TO (B)(6) MEDICAL CENTER ON (B)(6) 2023 TO HAVE AN IMPLANTED SMITH & NEPHEW BIRMINGHAM HIP PROSTHESIS REMOVED STATING HE BELIEVES THE METAL ON METAL PRODUCT HAS CONTRIBUTED TO HIS OTHER HEALTH COMORBIDITIES LIKE HIS PACEMAKER, MOOD SWINGS AND ARRYTHMIAS. UNKNOWN DATE OF IMPLANT AT ANOTHER HOSPITAL. UNABLE TO DETERMINE IF IT IS DEFECTIVE DUE TO EXCESSIVE AMOUNT OF BIOBURDEN AND TISSUE UNABLE TO BE CLEANED. SURGEON NOTES FAILED HARDWARE, RIGHT HIP. SURGERY PERFORMED TO REMOVE WAS REVISION RIGHT TOTAL MAKO HIP ARTHROPLASTY. |  | 2023 | 3 | 2023-03
3005975929-2023-00022 | 2023-03-30 | 2023-02-21 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) IMPL | PROSTHESIS, HIP, FEMORAL COMPONENT, CEMENTED, METAL |  |  | Male | IT WAS REPORTED THAT, AFTER A BHR SYSTEM HAD BEEN IMPLANTED ON AN UNSPECIFIED DATE, THE PATIENT STARTED SUSPECTING THAT THE MOOD SWINGS AND ARRYTHMIAS EXPERIENCED ARE ATTRIBUTED TO THE METAL-ON-METAL CONSTRUCT. THE ISSUE WAS ADDRESSED VIA REVISION SURGERY ON (B)(6) 2023. THE PATIENT'S CURRENT HEALTH CONDITION IS UNKNOWN. |  | 2023 | 3 | 2023-03
3005975929-2023-00027 | 2023-04-18 | 2012-06-15 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 41 YR | Male | IT WAS REPORTED THAT, AFTER UNDERGOING BHR HIP RESURFACING ON AN UNKNOWN DATE, THE PATIENT EXPERIENCED UNKNOWN COMPLICATIONS THAT LED TO A REVISION SURGERY ON (B)(6) 2012. DURING THIS PROCEDURE, CONVERSION TO BHR-THA SYSTEM WAS PERFORMED. A MODULAR FEMORAL HEAD AND SLEEVE WERE PLACED, ALONG WITH AN ANTHOLOGY HIP STEM. THE CUP WAS MAINTAINED. |  | 2023 | 4 | 2023-04
3005975929-2023-00029 | 2023-04-21 | 2022-03-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP IMPLANT | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/CERAMIC/POLYMER, CEMENTED OR NON-POROUS |  |  | Male | IT WAS REPORTED THAT, AFTER A REVISION SURGERY WAS PERFORMED ON (B)(6) 2008 IN WHICH THE ZIMMER IMPLANTS WERE REPLACED AND EXCHANGED FOR A BHR HIP, THE PATIENT EXPERIENCED FLAKING FROM THE HIP. AFTER 2 ADDITIONAL BLOOD TESTS OVER THE PAST YEAR, EACH TEST RESULTS YIELDED AN INCREASE IN THE COBALT LEVELS IN BLOOD. A SECONDARY REVISION SURGERY IS SCHEDULED FOR (B)(6) 2023. PATIENT HAVE BECOME SYMPTOMATIC WITH PAIN & OTHER DIFFICULTIES. |  | 2023 | 4 | 2023-04
3005975929-2023-00031 | 2023-04-25 | 2021-01-14 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "HOW MUCH DOES A MEDICAL AND HEALTHCARE PRODUCTS REGULATORY AGENCY MEDICAL DEVICE ALERT FOR METAL-ON-METAL HIP ARTHROPLASTY PATIENTS REALLY COST?", ONE HUNDRED AND EIGHTY EIGHT (188) PATIENTS WERE REVISED AFTER UNDERGOING A BHR RESURFACING PROCEDURE. PATIENTS OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2023 | 4 | 2023-04
3005975929-2023-00032 | 2023-04-25 | 2021-01-14 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "HOW MUCH DOES A MEDICAL AND HEALTHCARE PRODUCTS REGULATORY AGENCY MEDICAL DEVICE ALERT FOR METAL-ON-METAL HIP ARTHROPLASTY PATIENTS REALLY COST?", FORTY THREE (43) PATIENTS PRESENTED BLOOD LEVELS OF COBALT OR CHROMIUM ABOVE 7 UG/L AFTER UNDERGOING A BHR RESURFACING PROCEDURE. PATIENTS OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2023 | 4 | 2023-04
3005975929-2023-00044 | 2023-04-25 | 2021-01-14 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 77 YR | Female | IT WAS REPORTED THAT ON LITERATURE REVIEW "HOW MUCH DOES A MEDICAL AND HEALTHCARE PRODUCTS REGULATORY AGENCY MEDICAL DEVICE ALERT FOR METAL-ON-METAL HIP ARTHROPLASTY PATIENTS REALLY COST?", ONE (1) PATIENT WAS CONFIRMED WITH LOOSENING OF THE CUP, ADVERSE REACTION TO METAL DEBRIS (ARMD) AND COBALT/CHROMIUM LEVELS AT 26,47 UG/L AND 35,22 UG/L RESPECTIVELY AFTER UNDERGOING A BHR "RESURFACING" PROCEDURE. THE EVENT WAS TREATED BY REVISION SURGERY. NO FURTHER INFORMATION IS AVAILABLE. |  | 2023 | 4 | 2023-04
3005975929-2023-00070 | 2023-05-15 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 78 YR | Male | IT WAS REPORTED THAT, AFTER A BHR SYSTEM HAD BEEN IMPLANTED ON AN UNSPECIFIED DATE, THE PLAINTIFF EXPERIENCED AN UNSPECIFIED ADVERSE EVENT THAT WILL BE ADDRESSED VIA REVISION SURGERY IN WHICH THE DATE IS STILL PENDING. NO ADDITIONAL INFORMATION WAS PROVIDED. |  | 2023 | 5 | 2023-05
MW5120631 | 2023-08-04 | 2023-02-01 | I, n, j, u, r, y | Voluntary report | SMITH & NEPHEW ORTHOPAEDICS LTD. | UNKNOWN (BIRMINGHAM HIP RESURFACING) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 52 YR | Male | HIP RESURFACED WITH SMITH AND NEPHEW DEVICE IN (B)(6) 2021. BEGAN EXPERIENCING RECURRING PAIN IN (B)(6) 2023. PROVIDER DETERMINED THAT DEVICE IS SHIFTING AND GRINDING, AND BLOOD COBALT AND CHROMIUM ARE EXTREMELY HIGH. DEVICE TO BE REPLACED (B)(6) 2023. |  | 2023 | 8 | 2023-08
3005975929-2023-00109 | 2023-08-10 | 2020-01-17 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 67 YR | Female | IT WAS REPORTED THAT THE PLAINTIFF UNDERWENT A LEFT HIP REVISION SURGERY ON (B)(6) 2020 DUE TO ELEVATED METAL ION LEVELS IN BLOOD AND PSEUDOTUMOR FORMATION. DURING THE PROCEDURE, IT WAS NOTICED THAT THERE WAS EXTENSIVE INFECTION OF THE PERIPROSTHETIC TISSUES WITH LOSS OF BONE SUBSTANCE ON THE POSTERIOR FEMORAL WALL AND EXPOSURE OF ABOUT HALF OF THE PROSTHESIS. AT THIS TIME, IT WAS DECIDED TO PERFORM A 2-STEP REVISION SURGERY AND THE M-O-M SYSTEM WAS REPLACED WITH ANTIBIOTIC SPACERS. THE SECOND STEP OF THE 2-STEP REVISION SURGERY WAS PERFORMED ON (B)(6) 2020, AND DURING THIS PROCEDURE, THE ANTIBIOTIC SPACERS WERE EXPLANTED AND REPLACED WITH AN UNSPECIFIED THA SYSTEM. DURING THE PROCEDURE, THE PATIENT SUFFERED FROM A FRACTURE OF THE GREATER TROCHANTERIC BONE WHICH WAS TREATED WITH A PLATE AND FIXED WITH METAL CLIPS. THE CURRENT STATE OF HEALTH OF THE PATIENT IS UNKNOWN. THE PRIMARY LEFT HIP PROCEDURE WAS PERFORMED ON (B)(6) 2008. |  | 2023 | 8 | 2023-08
3005975929-2023-00118 | 2023-09-03 | 2010-01-15 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) CUP | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 49 YR | Female | IT WAS REPORTED THAT A REVISION OF RESURFACING PROSTHESIS WITH TOTAL RIGHT HIP PROSTHESIS WAS PERFORMED ON (B)(6) 2010 DUE TO ACETABULAR CUP LOOSENING. DURING THIS PROCEDURE, IT WAS NOTICED THAT THE PATIENT PRESENTED SYNOVITIS AND METALLOSIS. AFTER THE PROCEDURE THE PATIENT WAS TAKEN TO THE RECOVERY ROOM IN STABLE CONDITION. |  | 2023 | 9 | 2023-09
3005975929-2023-00130 | 2023-09-05 | 2008-12-24 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) CUP INSTR | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 46 YR | Male | IT WAS REPORTED THAT, WHILE THE PATIENT WAS UNDERGOING A BHR SURGERY OF THE RIGHT HIP ON (B)(6) 2008, WHEN THE ACETABULAR COMPONENT TRIAL WAS IMPACTED AND INSERTED, A SMALL CRACK ON THE PATIENT'S ACETABULUM WAS NOTED. WHEN THE TRIAL WAS REMOVED IT WAS NOTED THAT THE FRACTURE HAD PROPAGATED AT LEAST 3CM PROXIMAL AND POSTERIORLY. DUE TO THIS, SURGEON ELECTED NOT TO PROCEED WITH THE PLACEMENT OF THE BHR CUP, INSTEAD, IT WAS DECIDED TO PROCEED WITH A BHR-THA SURGERY USING A CEMENT-LESS CUP WITH ADJUVANT SCREW FIXATION (R3 3 HOLE ACET SHELL 56MM). IT IS UNKNOWN IF THERE WERE ANY DELAYS TO THE SURGERY DUE TO THIS ISSUE. THE PATIENT WAS TRANSPORTED TO THE PAR AWAKE AND IN STABLE CONDITION HAVING TOLERATED THE PROCEDURE WELL. PATIENT UNDERWENT A REVISION SURGERY ON (B)(6) 2014. |  | 2023 | 9 | 2023-09
3005975929-2023-00143 | 2023-09-20 | 2023-02-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT, AFTER A BHR SYSTEM HAD BEEN IMPLANTED ON (B)(6) 2021, THE PATIENT BEGAN EXPERIENCING RECURRING PAIN IN AN UNSPECIFIED DATE IN 2023. THE HEALTHCARE PROVIDER DETERMINED THAT (I) THE DEVICE IS SHIFTING AND GRINDING; (II) BLOOD COBALT AND CHROMIUM LEVELS WERE REPORTED TO BE 'EXTREMELY' HIGH, ALTHOUGH THE CONCENTRATION WAS NOT MENTIONED; AND (III) THE IMPLANT IS TO BE REPLACED ON AN UNSPECIFIED DATE IN 2023. IT IS UNKNOWN IF THE REVISION SURGERY HAS BEEN ALREADY PERFORMED OR NOT. FURTHER INFORMATION WAS NOT PROVIDED. |  | 2023 | 9 | 2023-09
3005975929-2023-00160 | 2023-11-27 | 2023-11-07 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Unknown | IT WAS REPORTED THAT, AFTER A THR SURGERY HAD BEEN PERFORMED ON (B)(6) 2009, IT WAS NOTICED THAT THE NECK OF THE FEMORAL STEM WAS WORN OUT, HENCE THE SYSTEM NEEDED TO BE REMOVED. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2023 TO ADDRESS THIS EVENT, EXPLANTING ALL THE SYSTEM. CURRENT HEALTH STATUS OF PATIENT IS UNKNOWN. |  | 2023 | 11 | 2023-11
3005975929-2023-00162 | 2023-11-30 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  |  | Unknown | IT WAS REPORTED THAT, AFTER A LEFT THR HAD BEEN PERFORMED ON THE YEAR 2012, FOR THE PAST 2 YEARS THE PATIENT STARTED EXPERIENCING NUMBNESS AND PAIN. ON (B)(6) 2023, PATIENT WAS DETECTED WITH HIGH CHROMIUM AND COBALT LEVELS IN HIS BLOOD. PATIENT HAD UNDERGONE REVISION SURGERY ON HIS RIGHT HIP, AND IS ESTIMATED TO HAVE A REVISION SURGERY ON HIS LEFT HIP IN APPROXIMATELY 3 MONTHS. THE DOCTOR RECOMMENDS A COMPLETE RECOVERY OF HIS RIGHT HIP BEFORE THE NEXT SURGERY IS PERFORMED. CURRENT HEALTH STATUS OF PATIENT IS GOOD, BUT STILL RECOVERING. |  | 2023 | 11 | 2023-11
3005975929-2024-00024 | 2024-02-20 | 2016-03-21 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP IMPLANT | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 61 YR | Female | IT WAS REPORTED THAT, AFTER A THA-BHR WAS CARRIED OUT ON AN UNKNOWN DATE, THE PLAINTIFF WAS DIAGNOSED WITH METALLOSIS AND ACETABULAR CUP LOOSENING. THEREFORE, A REVISION WAS PERFORMED ON (B)(6) 2016, DURING WHICH, TISSUE SIGNIFICANTLY IMPREGNATED WITH METALLOSIS WAS FOUND. THE ACETABULUM WAS EASILY REMOVED, BOTH ACETABULAR SHELL AND FEMORAL HEAD WERE EXCHANGED FOR COMPETITOR AND S+N DEVICES. NO OTHER COMPLICATIONS WERE REPORTED. |  | 2024 | 2 | 2024-02
3005975929-2024-00031 | 2024-03-12 | 2019-11-26 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 72 YR | Male | IT WAS REPORTED THAT, A BILATERAL PLAINTIFF UNDERWENT A PRIMARY RIGHT R3-THA ON (B)(6) 2010. ALMOST TEN YEARS LATER, THEY WERE DIAGNOSED WITH EARLY RENAL FAILURE, ALTR, HIGH COBALT AND CHROMIUM ION LEVELS, AND PSEUDOPOLYPS WERE FOUND BY CROSSIMAGING. THEREFORE, A REVISION SURGERY WAS PERFORMED ON (B)(6) 2019, DURING WHICH DISCOLORED TISSUE TO GREYISH-BROWN COLOR AND TRUNNIONOSIS WAS FOUND. DURING SURGERY THE PSEUDOPOLYPS WERE DEBRIDED AND THE CUP WAS NOTICED TO BE IN GOOD CONDITION, THEREFORE, ONLY THE SHELL HOLE COVER, METAL LINER AND METAL MODULAR HEAD WERE EXCHANGED; A LATERALIZED XLPE LINER 36MM AND AN OXINIUM FEMORAL HEAD 36MM 0 WAS IMPLANTED. NO OTHER RELATED COMPLICATIONS WERE REPORTED. |  | 2024 | 3 | 2024-03
3005975929-2024-00054 | 2024-05-20 | 2023-12-11 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 86 YR | Male | IT WAS REPORTED THAT, AFTER A RIGHT BIRMINGHAM SURGERY ON (B)(6) 2007, THE PLAINTIFF WAS DIAGNOSED WITH HIGH CHROMIUM COBALT LEVELS IN (B)(6) 2023. TO TREAT THIS ADVERSE EVENT AND SYMPTOMS OF LONG-STANDING PAIN A REVISION SURGERY WAS PERFORMED ON (B)(6) 2023, DURING WHICH THE SHELL AND STEM HAD NO SIGNS OF WEAR AND WERE WELL FIXED, BUT BLACK CORROSION WAS FOUND ON THE TRUNNION. THE BIRMINGHAM HIP WAS CONVERTED TO AN OXINIUM DUAL MOBILITY BEARING. PATIENT WAS TRANSFERRED TO HOSPITAL BED AND AROUSED FROM ANESTHESIA WITHOUT COMPLICATIONS. |  | 2024 | 5 | 2024-05
3005975929-2024-00067 | 2024-05-26 | 2023-12-13 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 60 YR | Female | IT WAS REPORTED THAT THE PLAINTIFF UNDERWENT AN UNKNOWN METAL-ON-METAL HIP SURGERY ON AN UNKNOWN DATE. ON (B)(6) 2023 THE PATIENT OBTAINED RESULTS SHOWING HIGH LEVELS OF COBALT AND CHROMIUM IN THE BLOOD. IT IS UNKNOWN IF THIS ADVERSE EVENT WAS TREATED OR IF THERE IS A SCHEDULED REVISION SURGERY TO SOLVE THE PROBLEM OF THE CLAIMANT. CURRENT HEALTH STATUS OF THE PATIENT IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE AT THE MOMENT. |  | 2024 | 5 | 2024-05
3005975929-2024-00071 | 2024-05-27 | 2022-05-04 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Male | IT WAS REPORTED THAT THE PATIENT SUFFERED FROM LEFT HIP PAIN, CREPITUS, LIMITATIONS WITH HIS STRENGTH AND RANGE OF MOTION AFTER HAVING A BHR SURGERY. PATIENT WAS SCHEDULED FOR A LEFT HIP REVISION, AND EXCISION OF HETEROTROPHIC OSTEOPHYTE. IT IS UNKNOWN IF THE RIGHT HIP WAS REVISED. |  | 2024 | 5 | 2024-05
3005975929-2024-00069 | 2024-05-27 | 2022-12-22 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 29 YR | Male | IT WAS REPORTED THAT, A PLAINTIFF RECEIVED A BHR SYSTEM ON AN UNKNOWN DATE, AND LATER EXPERIENCED CHRONIC PAIN AND WAS DIAGNOSED WITH HIP HERNIAS, MINIMALLY ELEVATED METAL ION LEVEL, MRI SHOWED FLUID WITHIN HIP, AND EMGS SHOWED SEVERE POLYNEUROPATHY BELIEVED TO BE SECONDARY TO ELEVATED METAL IONS. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2022 DURING WHICH THE ACETABULAR SHELL WAS FOUND TO BE WELL FIXED, THIS WAS CONVERTED TO A DUAL MOBILITY HIP USING A S+N INSERT AND DEPUY FEMORAL COMPONENTS. NO OTHER COMPLICATIONS WERE REPORTED. |  | 2024 | 5 | 2024-05
3005975929-2024-00085 | 2024-06-23 | 2023-12-05 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 64 YR | Male | IT WAS REPORTED THAT, A BILATERAL PLAINTIFF UNDERWENT A RIGHT BHR SURGERY ON AN UNKNOWN DATE AND WAS LATER DIAGNOSED WITH HIGH CHROMIUM COBALT LEVELS ON (B)(6) 2022 AND ALSO EXPERIENCED PAIN. THEREFORE, A REVISION SURGERY WAS PERFORMED ON (B)(6) 2023, TO EXCHANGE THE BHR COMPONENTS FOR AN ENOVIS EMPOWR TOTAL HIP SYSTEM. THE PATIENT WAS TRANSFERRED TO THE RECOVERY ROOM IN STABLE CONDITION. |  | 2024 | 6 | 2024-06
3005975929-2024-00088 | 2024-06-26 | 2024-04-11 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) HEMI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 69 YR | Male | IT WAS REPORTED THAT A BILATERAL PATIENT HAD A PRIMARY RIGHT HIP METAL-ON-METAL THA SURGERY IN 2008. THE PATIENT EXPERIENCED ELEVATED METAL ION LEVELS, AND AFTER UNDERGOING A LEFT REVISION SURGERY ON (B)(6) 2022 DUE TO AN ADVERSE LOCAL TISSUE REACTION, THEIR METAL IONS REMAINED ELEVATED. THEREFORE, THEY ELECTED TO UNDERGO A REVISION SURGERY ON (B)(6) 2024, DURING WHICH, ALL BHR ACETABULAR AND MODULAR FEMORAL HEAD COMPONENTS WERE REPLACED FOR A COMBINATION OF S+N AND ZIMMER IMPLANTS. AFTER SURGERY THEY WERE TRANSFERRED TO THE RECOVERY ROOM IN STABLE CONDITION. |  | 2024 | 6 | 2024-06
3005975929-2024-00098 | 2024-07-03 | 2024-01-02 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "IS THERE STILL AN INDICATION FOR THE BIRMINGHAM HIP RESURFACING IN FEMORAL HEAD OSTEONECROSIS?", TWELVE (12) PATIENTS IN TOTAL PRESENTED OSTEOLYSIS AROUND THE FEMORAL COMPONENT. FOUR (4) PATIENTS HAD OSTEOLYSIS IN ZONE 1, FIVE (5) PATIENTS HAD OSTEOLYSIS IN ZONES 1 AND 3 AND THREE (3) PATIENTS HAD OSTEOLYSIS IN ALL THREE ZONES ACCORDING TO THE CLASSIFICATION SYSTEM BY AMSTUTZ. PATIENTS OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2024 | 7 | 2024-07
3005975929-2024-00096 | 2024-07-03 | 2024-01-02 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "IS THERE STILL AN INDICATION FOR THE BIRMINGHAM HIP RESURFACING IN FEMORAL HEAD OSTEONECROSIS?", ONE PATIENT UNDERWENT A REVISION SURGERY DUE TO LOOSENING OF THE FEMORAL RESURFACING HEAD. PATIENT OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2024 | 7 | 2024-07
3005975929-2024-00097 | 2024-07-03 | 2024-01-02 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "IS THERE STILL AN INDICATION FOR THE BIRMINGHAM HIP RESURFACING IN FEMORAL HEAD OSTEONECROSIS?", TWENTY NINE (29) PATIENTS IN TOTAL PRESENTED OSTEOLYSIS AROUND THE ACETABULAR COMPONENT. FIFTEEN (15) PATIENTS IN ZONE 1 AND FOURTEEN (14) PATIENTS IN ZONE 2 ACCORDING TO THE CLASSIFICATION SYSTEM BY DELEE AND CHARNLEY. PATIENTS OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2024 | 7 | 2024-07
3005975929-2024-00106 | 2024-07-08 | 2024-04-23 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "CLINICAL, FUNCTIONAL AND RADIOGRAPHIC LONG-TERM FOLLOW-UP (7-12 YEARS) OF BIRMINGHAM HIP RESURFACING, INCLUDING METAL IONS EVALUATION: A SINGLE SURGEON SERIES", AFTER FOLLOWING A BHR SURGERY, TWO (2) PATIENTS SUSTAINED A DEEP INFECTION REQUIRING REVISION SURGERY. PATIENTS¿ OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2024 | 7 | 2024-07
3005975929-2024-00110 | 2024-07-08 | 2024-04-23 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "CLINICAL, FUNCTIONAL AND RADIOGRAPHIC LONG-TERM FOLLOW-UP (7-12 YEARS) OF BIRMINGHAM HIP RESURFACING, INCLUDING METAL IONS EVALUATION: A SINGLE SURGEON SERIES", AFTER FOLLOWING A BHR SURGERY, TWO (2) PATIENTS SUSTAINED ELEVATED METAL ION CONCENTRATIONS IN BLOOD REQUIRING REVISION SURGERY. PATIENTS¿ OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2024 | 7 | 2024-07
3005975929-2024-00109 | 2024-07-08 | 2024-04-23 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "CLINICAL, FUNCTIONAL AND RADIOGRAPHIC LONG-TERM FOLLOW-UP (7-12 YEARS) OF BIRMINGHAM HIP RESURFACING, INCLUDING METAL IONS EVALUATION: A SINGLE SURGEON SERIES", AFTER FOLLOWING A BHR SURGERY, TWO (2) PATIENTS SUSTAINED OSTEONECROSIS REQUIRING REVISION SURGERY. PATIENTS¿ OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2024 | 7 | 2024-07
3005975929-2024-00107 | 2024-07-08 | 2024-04-23 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "CLINICAL, FUNCTIONAL AND RADIOGRAPHIC LONG-TERM FOLLOW-UP (7-12 YEARS) OF BIRMINGHAM HIP RESURFACING, INCLUDING METAL IONS EVALUATION: A SINGLE SURGEON SERIES", AFTER FOLLOWING A BHR SURGERY, THREE (3) PATIENTS SUSTAINED ALVAL (ASEPTIC LYMPHOCYTE-DOMINANT VASCULITIS-ASSOCIATED LESION) REACTION REQUIRING REVISION SURGERY. PATIENTS¿ OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2024 | 7 | 2024-07
3005975929-2024-00108 | 2024-07-09 | 2024-04-23 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "CLINICAL, FUNCTIONAL AND RADIOGRAPHIC LONG-TERM FOLLOW-UP (7-12 YEARS) OF BIRMINGHAM HIP RESURFACING, INCLUDING METAL IONS EVALUATION: A SINGLE SURGEON SERIES", AFTER FOLLOWING A BHR SURGERY, TWO (2) PATIENTS SUSTAINED ASEPTIC LOOSENING REQUIRING REVISION SURGERY. PATIENTS¿ OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2024 | 7 | 2024-07
3005975929-2024-00115 | 2024-07-09 | 2017-03-21 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW ""CLINICAL AND RADIOGRAPHIC OUTCOMES OF THE BIRMINGHAM HIP RESURFACING ARTHROPLASTY AT A MINIMUM FOLLOW-UP OF 10 YEARS: RESULTS FROM AN INDEPENDENT CENTRE"", 1 (ONE) PATIENT SUFFERED FROM PERIACETABULAR BONE RAREFACTION AND ELEVATED METAL ION LEVELS IN BLOOD AFTER A PRIMARY METAL-ON-METAL HIP RESURFACING PROCEDURE. HIP SCAN FOUND FLUID IN THE JOINT. TEST RESULTS INDICATED COBALT CONTENT OF 2.28 NG/ML AND CHROMIUM CONTENT OF 4.89 NG/ML. A REVISION SURGERY WAS RECOMMENDED TO ADDRESS THE ISSUE, BUT THE PATIENT REFUSED IT. THE OUTCOME OF THE PATIENT IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2024 | 7 | 2024-07
3005975929-2024-00116 | 2024-07-09 | 2015-11-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Male | IT WAS REPORTED THAT ON LITERATURE REVIEW "RETURN TO SPORTING ACTIVITY AFTER BIRMINGHAM HIP RESURFACING ARTHROPLASTY MID TERM RESULTS", 1 (ONE) PATIENT SUFFERED FROM A FRACTURE OF HIS RIGHT FEMORAL NECK 7 YEARS AFTER HIS INDEX PROCEDURE. THE INCIDENT OCCURRED SECONDARY TO A FALL FROM HIS PUSHBIKE. THE FRACTURE INCLUDED THE REGION OF BONE IN WHICH THE FEMORAL COMPONENT WAS FIXED HENCE THIS WAS REVISED TO A STEMMED PROSTHESIS. THE ACETABULAR COMPONENT WAS RETAINED. THE PRIMARY PROCEDURE WAS A BIRMINGHAM HIP RESURFACING ARTHROPLASTY. NO FURTHER INFORMATION IS AVAILABLE. |  | 2024 | 7 | 2024-07
3005975929-2024-00119 | 2024-07-10 | 2024-04-23 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "CLINICAL, FUNCTIONAL AND RADIOGRAPHIC LONG-TERM FOLLOW-UP (7-12 YEARS) OF BIRMINGHAM HIP RESURFACING, INCLUDING METAL IONS EVALUATION: A SINGLE SURGEON SERIES", AFTER FOLLOWING A BHR SURGERY, TWO (2) PATIENTS SUSTAINED A SUBCAPITAL FRACTURE REQUIRING REVISION SURGERY. PATIENTS¿ OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2024 | 7 | 2024-07
3005975929-2024-00120 | 2024-07-10 | 2024-04-23 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "CLINICAL, FUNCTIONAL AND RADIOGRAPHIC LONG-TERM FOLLOW-UP (7-12 YEARS) OF BIRMINGHAM HIP RESURFACING, INCLUDING METAL IONS EVALUATION: A SINGLE SURGEON SERIES", AFTER FOLLOWING A BHR SURGERY, ONE (1) PATIENT SUSTAINED AN UNSPECIFIED ADVERSE EVENT WHILE USING A BHR SYSTEM. PATIENT REQUIRED REVISION SURGERY (PERIACETABULAR OSTEOTOMY (PAO)). PATIENT OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2024 | 7 | 2024-07
3005975929-2024-00121 | 2024-07-10 | 2024-04-23 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "CLINICAL, FUNCTIONAL AND RADIOGRAPHIC LONG-TERM FOLLOW-UP (7-12 YEARS) OF BIRMINGHAM HIP RESURFACING, INCLUDING METAL IONS EVALUATION: A SINGLE SURGEON SERIES", AFTER FOLLOWING A BHR SURGERY, ONE (1) PATIENT SUSTAINED AN ACETABULAR MALPOSITION REQUIRING REVISION SURGERY. PATIENT OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2024 | 7 | 2024-07
3005975929-2024-00122 | 2024-07-10 | 2024-04-23 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "CLINICAL, FUNCTIONAL AND RADIOGRAPHIC LONG-TERM FOLLOW-UP (7-12 YEARS) OF BIRMINGHAM HIP RESURFACING, INCLUDING METAL IONS EVALUATION: A SINGLE SURGEON SERIES", AFTER FOLLOWING A BHR SURGERY, TWO (2) PATIENTS SUSTAINED A DEEP INFECTION WHILE USING A BHR SYSTEM. PATIENT¿S OUTCOME AND HOW THEY WERE TREATED IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2024 | 7 | 2024-07
3005975929-2024-00123 | 2024-07-10 | 2024-04-23 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "CLINICAL, FUNCTIONAL AND RADIOGRAPHIC LONG-TERM FOLLOW-UP (7-12 YEARS) OF BIRMINGHAM HIP RESURFACING, INCLUDING METAL IONS EVALUATION: A SINGLE SURGEON SERIES", AFTER FOLLOWING A BHR SURGERY, ONE (1) PATIENT SUSTAINED A HIP DISLOCATION WHILE USING A BHR SYSTEM. PATIENT OUTCOME AND HOW IT WAS TREATED IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2024 | 7 | 2024-07
3005975929-2024-00130 | 2024-08-12 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "LONG TERM RESULTS OF THE BIRMINGHAM HIP RESURFACING IMPLANT IN THE UNITED STATES: AN UPDATED ANALYSIS OF A SINGLE INSTITUTION¿S EXPERIENCE", ONE PATIENT UNDERWENT A TWO-STAGE REVISION SURGERY DUE TO AN INFECTION APPROXIMATELY ONE YEAR AFTER RECEIVING A BHR SYSTEM. THE FEMORAL AND ACETABULAR COMPONENTS WERE EXTRACTED AN ARTICULATING SPACER WAS PLACED. PATIENT OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2024 | 8 | 2024-08
3005975929-2024-00131 | 2024-08-12 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "LONG TERM RESULTS OF THE BIRMINGHAM HIP RESURFACING IMPLANT IN THE UNITED STATES: AN UPDATED ANALYSIS OF A SINGLE INSTITUTION¿S EXPERIENCE", ONE PATIENT UNDERWENT A TWO-STAGE REVISION SURGERY DUE TO AN INFECTION APPROXIMATELY FIVE YEARS AFTER RECEIVING A BHR SYSTEM. THE FEMORAL AND ACETABULAR COMPONENTS WERE EXTRACTED AN ARTICULATING SPACER WAS PLACED. PATIENT OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2024 | 8 | 2024-08
3005975929-2024-00132 | 2024-08-12 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "LONG TERM RESULTS OF THE BIRMINGHAM HIP RESURFACING IMPLANT IN THE UNITED STATES: AN UPDATED ANALYSIS OF A SINGLE INSTITUTION¿S EXPERIENCE", ONE PATIENT UNDERWENT A REVISION SURGERY DUE TO UNEXPLAINED PAIN APPROXIMATELY SIX YEARS AFTER RECEIVING A BHR SYSTEM. THE FEMORAL AND ACETABULAR COMPONENTS WERE EXTRACTED AND THE PATIENT RECEIVED A THA SYSTEM INSTEAD. PATIENT OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2024 | 8 | 2024-08
3005975929-2024-00133 | 2024-08-12 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "LONG TERM RESULTS OF THE BIRMINGHAM HIP RESURFACING IMPLANT IN THE UNITED STATES: AN UPDATED ANALYSIS OF A SINGLE INSTITUTION¿S EXPERIENCE", ONE PATIENT UNDERWENT A REVISION SURGERY DUE TO FEMORAL COMPONENT LOOSENING APPROXIMATELY SIX YEARS AFTER RECEIVING A BHR SYSTEM. THE FEMORAL COMPONENT WAS EXTRACTED AND THE SYSTEM WAS CONVERTED TO A DUAL MOBILITY. PATIENT OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2024 | 8 | 2024-08
3005975929-2024-00134 | 2024-08-12 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "LONG TERM RESULTS OF THE BIRMINGHAM HIP RESURFACING IMPLANT IN THE UNITED STATES: AN UPDATED ANALYSIS OF A SINGLE INSTITUTION¿S EXPERIENCE", ONE PATIENT UNDERWENT A REVISION SURGERY DUE TO FEMORAL COMPONENT LOOSENING TWELVE YEARS AFTER RECEIVING A BHR SYSTEM. THE FEMORAL COMPONENT WAS EXTRACTED AND THE SYSTEM WAS CONVERTED TO A DUAL MOBILITY. PATIENT OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2024 | 8 | 2024-08
3005975929-2024-00137 | 2024-08-12 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "LONG TERM RESULTS OF THE BIRMINGHAM HIP RESURFACING IMPLANT IN THE UNITED STATES: AN UPDATED ANALYSIS OF A SINGLE INSTITUTION¿S EXPERIENCE", ONE PATIENT UNDERWENT A REVISION SURGERY DUE TO ELEVATED METAL IONS TWELVE YEARS AFTER RECEIVING A BHR SYSTEM. THE FEMORAL COMPONENT WAS EXTRACTED AND THE SYSTEM WAS CONVERTED TO A DUAL MOBILITY. PATIENT OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2024 | 8 | 2024-08
3005975929-2024-00128 | 2024-08-12 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "LONG TERM RESULTS OF THE BIRMINGHAM HIP RESURFACING IMPLANT IN THE UNITED STATES: AN UPDATED ANALYSIS OF A SINGLE INSTITUTION¿S EXPERIENCE", ONE PATIENT UNDERWENT A REVISION SURGERY DUE TO A PSEUDOTUMOR AND ELEVATED METAL IONS THREE YEARS AFTER RECEIVING A BHR SYSTEM. THE FEMORAL COMPONENT WAS REVISED AND THE SYSTEM WAS CONVERTED TO A DUAL MOBILITY SYSTEM. PATIENT OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2024 | 8 | 2024-08
3005975929-2024-00129 | 2024-08-12 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "LONG TERM RESULTS OF THE BIRMINGHAM HIP RESURFACING IMPLANT IN THE UNITED STATES: AN UPDATED ANALYSIS OF A SINGLE INSTITUTION¿S EXPERIENCE", ONE PATIENT UNDERWENT A TWO-STAGE REVISION SURGERY DUE TO AN INFECTION APROXIMATELY FOUR YEARS AFTER RECEIVING A BHR SYSTEM. THE FEMORAL AND ACETABULAR COMPONENTS WERE EXTRACTED AN ARTICULATING SPACER WAS PLACED. PATIENT OUTCOME IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2024 | 8 | 2024-08
3005975929-2024-00138 | 2024-08-13 | 2015-04-24 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP MODULAR HEAD (BHMH) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 73 YR | Male | IT WAS REPORTED THAT, ON (B)(6) 2005 A BILATERAL PATIENT RECEIVED A LEFT BHR-THA SYSTEM. TEN YEARS LATER, THEY WERE DIAGNOSED WITH ELEVATED CHROMIUM COBALT IONS LEVELS, AND AN MRI-MARS SHOWED SIGNIFICANT FLUID COLLECTION IN BOTH HIPS. THEY ALSO STARTED DEVELOPING PAIN OVER THE LATERAL ASPECT OF THEIR LEG AND WHEN SITTING. THEREFORE, A REVISION SURGERY WAS PERFORMED ON (B)(6) 2015, TO TREAT AN ADVERSE LOCAL SOFT TISSUE REACTION. DURING SURGERY, FLUID COLLECTION, DEBRIS ASSOCIATED SYNOVITIS AND MACROSCOPIC EVIDENCE OF FRETTING AT THE TAPER JUNCTION BETWEEN NECK OF THE FEMORAL COMPONENT AND THE INSIDE OF THE MOM HEAD IMPLANT WAS FOUND. THE ACETABULAR COMPONENT WAS WELL FIXED IN APPROPRIATE POSITION AND ALIGNMENT, AS WELL AS THE FEMORAL COMPONENT, THEREFORE ONLY THE FEMORAL HEAD COMPONENTS WERE EXCHANGED FOR A STRYKER POLY AND OXINIUM MODULAR HEAD. NO OTHER COMPLICATIONS WERE REPORTED. |  | 2024 | 8 | 2024-08
3005975929-2024-00147 | 2024-09-27 | 2022-10-25 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Male | IT WAS REPORTED THAT, A PATIENT RECEIVED A RIGHT BHR CUP, MODULAR HEAD AND AN ANTHOLOGY STEM ON (B)(6) 2007. LATER THEY UNDERWENT MEDICALLY INDICATED REVISION SURGERY ON (B)(6) 2022. THE PATIENT HAD EXPERIENCED SEVERE PAIN, LIMITED MOBILITY, METALLOSIS, AND ELEVATED COBALT CHROMIUM LEVELS. DURING THE REVISION SURGERY, SIGNIFICANT METALLOSIS, DEBRIS, AND A LARGE PSEUDOTUMOR WAS FOUND, INDICATIVE OF MOM REACTION. THE BHR COMPONENTS WERE EXCHANGED FOR A ZIMMER FEMORAL HEAD AND CUP, AND A ZIMMER XLPE DUAL MOBILITY LINER. NO OTHER COMPLICATIONS WERE REPORTED. |  | 2024 | 9 | 2024-09
3005975929-2024-00148 | 2024-10-17 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP IMPLANT | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Male | IT WAS REPORTED THAT, AFTER RECEIVING A MOM HIP IMPLANT ON AN UNKNOWN DATE, THE PATIENT WAS DIAGNOSED WITH METALLOSIS. THEREFORE, THE IMPLANT WAS REVISED ON AN UNSPECIFIED DATE, AND THE PATIENT WAS INFORMED TO BE STILL RECOVERING FROM SUCH PROCEDURE. NO OTHER COMPLICATIONS WERE REPORTED. |  | 2024 | 10 | 2024-10
3005975929-2024-00155 | 2024-11-19 | 2024-02-06 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) CUP | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Male | IT WAS REPORTED THAT AFTER UNDERGOING A BHR RESURFACING TO BHR-THA CONVERSION IN HIS LEFT HIP JOINT ON (B)(6) 2007 DUE TO LEFT FEMORAL NECK FRACTURE, THE PATIENT EXPERIENCED PAIN, AN ADVERSE TISSUE REACTION TO METAL DEBRIS AND A PSEUDOTUMOR IN THE HIP JOINT. THIS ADVERSE EVENT WAS ADDRESSED BY CONDUCTING A REVISION SURGERY ON (B)(6) 2024. NO FURTHER INFORMATION IS AVAILABLE. |  | 2024 | 11 | 2024-11
3005975929-2024-00156 | 2024-11-20 | 2007-04-20 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Male | IT WAS REPORTED THAT AFTER UNDERGOING A BHR RESURFACING IN HIS LEFT HIP JOINT ON (B)(6) 2007, THE PATIENT SUSTAINED A LEFT FEMORAL NECK FRACTURE THAT REQUIRED A CONVERSION TO THA ON (B)(6) 2007. THE RESULTING SYSTEM INCLUDED A TRADITIONAL FEMORAL STEM, BHR CUP, AND MODULAR COBALT-CHROMIUM FEMORAL HEAD. NO FURTHER INFORMATION IS AVAILABLE. |  | 2024 | 11 | 2024-11
20733840 | 2024-11-20 | 2024-11-05 | M, a, l, f, u, n, c, t, i, o, n | User Facility report | SMITH & NEPHEW ORTHOPAEDICS LTD | BIRMINGHAM HIP RESURFACING (BHR) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 48 YR | Male | A REPRESENTATIVE FROM SMITH AND NEPHEW GAVE AN IMPLANT TO BE USED ON THE PATIENT FOR THE HIP RESURFACING PER SURGEON REQUEST. UPON INSERTION OF THE IMPLANT, PART OF THE PIECE BROKE WHILE USING IT. THE IMPLANT WAS INSERTED BUT NOT RETAINED. THIS REQUIRED THAT A NEW IMPLANT HAD TO BE OPENED, AND IT WAS DETERMINED THAT IT WAS A MANUFACTURER ERROR AFTER DISCUSSING IT WITH THE REPRESENTATIVE, SURGEON, AND SURGICAL TECH. THE REPRESENTATIVE STATED THAT HE WOULD DISCUSS IT WITH HIS COMPANY AND TAKE RESPONSIBILITY FOR THE IMPLANT. MANUFACTURER RESPONSE FOR PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING, BIRMINGHAM HIP RESURFACING (BHR) (PER SITE REPORTER). I REACHED OUT TO THE REPRESENTATIVE WHO IS IN POSSESSION OF THE FAILED PRODUCT FOR AN INVESTIGATIVE REPORT. NO FOLLOW UP AS OF 11/12/2024. |  | 2024 | 11 | 2024-11
3005975929-2024-00160 | 2024-12-04 | 2024-05-10 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Male | IT WAS REPORTED THAT, AFTER A RIGHT PRIMARY BHR PROCEDURE WAS PERFORMED ON AN UNKNOWN DATE, THE PATIENT WAS ELECTED FOR A REVISION SURGERY TO CONVERT INTO A THR, WITHOUT SPECIFIED SYMPTOMATOLOGY. THE REVISION WAS CARRIED OUT ON (B)(6) 2024, DURING WHICH SEVERE OA CHANGES IN THE HIP WITH MARKED DEFORMITY WERE NOTED, AND THE FEMORAL HEAD COMPONENT WAS EXCHANGED FOR A BIOLOX DELTA CERAMIC HEAD AND A SECURE FIT STEM. AFTER PROCEDURE, PATIENT WAS TRANSFERRED TO PACU IN STABLE STATE. |  | 2024 | 12 | 2024-12
3005975929-2024-00166 | 2024-12-05 | 2010-02-17 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 49 YR | Female | IT WAS REPORTED THAT, AFTER A LEFT BHR WAS PERFORMED ON (B)(6) 2008, PATIENT DEVELOPED PROXIMAL THIGH PAIN, LIMPING AND DIFFICULTY WITH DAILY ACTIVITIES. LATER AN X-RAY SHOWED MIGRATION OF FEMORAL STEM; THEREFORE, THEY UNDERWENT A REVISION SURGERY ON (B)(6) 2010. THE ACETABULAR SHELL WAS FOUND TO BE WELL INGROWN AND STABLE, AND A FIBROUS INTEGRATION ALONG HEAD AND NECK JUNCTION WAS OBSERVED. THEIR BHR FEMORAL COMPONENT WAS EXCHANGED FOR A MODULAR BHR HEAD, AND SYNERGY STEM. AFTER SURGERY, PATIENT WAS TAKEN TO RECOVERY ROOM IN SATISFACTORY CONDITION. |  | 2024 | 12 | 2024-12
3005975929-2024-00169 | 2024-12-20 | 2023-11-06 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 54 YR | Male | IT WAS REPORTED THAT, AFTER UNDERGOING A THA-MOM IN HIS RIGHT HIP JOINT ON AN UNKNOWN DATE, PATIENT EXPERIENCED OF A LARGE HEAD STEMMED METAL ON METAL TOTAL HIP ARTHROPLASTY WITH GALVANIC CORROSION AND ELEVATED CO/CR. THIS CAUSED SIGNIFICANT PAIN IN HIS HIP. ALSO, WORK-UP REVEALED LOW CONCERN FOR INFECTION. THIS ADVERSE EVENT WAS ADDRESSED BY CONDUCTING A REVISION SURGERY ON (B)(6) 2023, DURING WHICH, SURGEON EXCHANGED THE ACETABULAR CUP, HEAD TO DUAL MOBILITY AND RETAINED THE STEM. DURING PROCEDURE, THE ACETABULAR CUP WAS NOTED TO BE IN SIGNIFICANT RETROVERSION OF GREATER THAN 40 DEGREES. THE SURROUNDING NECROTIC SOFT TISSUES WERE EXCISED. PRIOR TO REMOVAL OF THE CUP, IT WAS NOTED THAT THERE WAS BONE LOSS AT THE POSTERIOR WALL. THE STEM WAS NOTED TO BE WELL FIXED. THE HIP WAS REDUCED. LEG LENGTH WAS NOTED TO BE APPROPRIATE. SHUCK TEST WAS APPROPRIATE ALBEIT SOFT DUE TO THE SOFT TISSUE DESTRUCTION. THE POSTOPERATIVE PLAN FOR THE PATIENT WAS WEIGHT BEARING AS TOLERATED, HIP ARTHROPLASTY POSTERIOR PRECAUTIONS, EARLY AMBULATION AND RISK ADJUSTED VTED PROPHYLAXIS. |  | 2024 | 12 | 2024-12
3005975929-2024-00170 | 2024-12-20 | 2023-12-07 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 55 YR | Male | IT WAS REPORTED THAT, AFTER UNDERGOING A THA-MOM IN HIS LEFT HIP JOINT ON AN UNKNOWN DATE, PATIENT EXPERIENCED OF A LARGE HEAD STEMMED METAL ON METAL TOTAL HIP ARTHROPLASTY WITH GALVANIC CORROSION AND ELEVATED CO/CR. THIS CAUSED SIGNIFICANT PAIN IN HIS HIP. ALSO, WORK-UP REVEALED LOW CONCERN FOR INFECTION. THIS ADVERSE EVENT WAS ADDRESSED BY CONDUCTING A REVISION SURGERY ON (B)(6) 2023, DURING WHICH, SURGEON EXCHANGED THE ACETABULAR CUP, HEAD TO DUAL MOBILITY AND RETAINED THE STEM. DURING PROCEDURE, IT WAS FOUND THERE WAS SIGNIFICANT ALTR WITH COPIOUS NECROTIC TISSUE AND GREY/GREEN FLUID, RETROVERTED MALPOSITIONED LINER WITH CORROSION BOTH AT THE LINER/CUP INTERFACE AND THE STEM/HEAD INTERFACE. THE ACETABULAR CUP WAS NOTED TO BE IN SIGNIFICANT RETROVERSION OF 40 DEGREES. THE SURROUNDING NECROTIC SOFT TISSUES WERE EXCISED. THERE WAS SIGNIFICANT BONE LOSS ANTERIORLY. THE STEM WAS NOTED TO BE WELL FIXED. THE HIP WAS REDUCED. LEG LENGTH WAS NOTED TO BE APPROPRIATE. SHUCK TEST WAS APPROPRIATE ALBEIT SOFT DUE TO THE SOFT TISSUE DESTRUCTION. THE POSTOPERATIVE PLAN FOR THE PATIENT WAS WEIGHT BEARING AS TOLERATED, HIP ARTHROPLASTY POSTERIOR PRECAUTIONS, EARLY AMBULATION AND RISK ADJUSTED VTED PROPHYLAXIS. |  | 2024 | 12 | 2024-12
3005975929-2024-00172 | 2024-12-26 | 2024-02-29 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP MODULAR HEAD (BHMH) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 67 YR | Female | IT WAS REPORTED THAT, A BILATERAL PATIENT UNDERWENT A PRIMARY RIGHT R3-THA SURGERY ON (B)(6) 2009. THEY LATER WERE DIAGNOSED WITH ELEVATED CHROMIUM COBALT LEVELS SINCE 2017. ADDITIONALLY, IN (B)(6) 2024 LABS DETECTED RENAL IMPAIRMENT AND BRAIN IMAGING SUGGESTED A COGNITIVE IMPAIRMENT AND A POSSIBLE ENCEPHALOPATHY, RELATED TO METAL TOXICITY. MRIS REVEALED NO CONVINCING EVIDENCE OF OSTEOLYSIS AND WERE NEGATIVE FOR ANY PSEUDOTUMOR. IT IS UNKNOWN IF THE PATIENT HAS BEEN SCHEDULED FOR A RIGHT HIP REVISION OR IF IT HAS BEEN PERFORMED. NO OTHER COMPLICATIONS WERE REPORTED. |  | 2024 | 12 | 2024-12
3005975929-2024-00173 | 2024-12-26 | 2024-05-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP MODULAR HEAD (BHMH) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 67 YR | Female | IT WAS REPORTED THAT, A BILATERAL PATIENT UNDERWENT A PRIMARY LEFT R3-THA SURGERY ON (B)(6) 2009. THEY LATER WERE DIAGNOSED WITH ELEVATED CHROMIUM COBALT LEVELS SINCE 2017. ADDITIONALLY, ON (B)(6) 2024 LABS DETECTED RENAL IMPAIRMENT AND BRAIN IMAGING SUGGESTED A COGNITIVE IMPAIRMENT AND A POSSIBLE ENCEPHALOPATHY, RELATED TO METAL TOXICITY. MRIS REVEALED A LEFT-SIDED ILIOPSOAS BURSITIS. THEREFORE, THE PATIENT WAS DIAGNOSED WITH METALLOSIS AND WAS REVISED ON (B)(6) 2024. DURING SURGERY, SIGNIFICANT AMOUNT OF DARK BROWN FLUID WAS FOUND, THE ACETABULAR SHELL HAD NO SIGNS OF WEAR, LOOSENING, OR OSTEOLYSIS; THEREFORE, THE METAL LINER AND FEMORAL HEMI HEAD WERE SWAPPED FOR A POLYETHYLENE LINER AND CERAMIC HEAD. PATIENT WAS LATER TAKEN TO RECOVERY ROOM IN STABLE CONDITION. |  | 2024 | 12 | 2024-12
3005975929-2025-00002 | 2025-01-14 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP IMPLANT | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Unknown | IT WAS REPORTED THAT, AFTER A BHR WAS PERFORMED ON AN UNKNOWN DATE, THE PATIENT HAS HAD TO UNDERGO EXTREME PAIN AND HAS REPORTED TO HAVE METAL IN ALL VITAL ORGANS. IT IS UNKNOWN HOW THIS HAS BEEN TREATED. |  | 2025 | 1 | 2025-01
3005975929-2025-00006 | 2025-02-14 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) CUP | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Unknown | IT WAS REPORTED THAT, ON THE LITERATURE REVIEW "REVISION OF METAL-ON-METAL HIP REPLACEMENTS WITH DUAL-MOBILITY BEARINGS AND ACETABULAR COMPONENT RETENTION", NINETY FOUR (94) METAL ON METAL BIRMINGHAM HIP REPLACEMENTS UNDERWENT REVISION SURGERY BETWEEN 2012 AND 2020. FROM THESE, TWO (2) HIPS WERE REVISED DUE TO ASEPTIC LOOSENING OF THE BHR ACETABULAR CUP. NO FURTHER INFORMATION IS AVAILABLE. |  | 2025 | 2 | 2025-02
3005975929-2025-00005 | 2025-02-14 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Unknown | IT WAS REPORTED THAT, ON THE LITERATURE REVIEW "REVISION OF METAL-ON-METAL HIP REPLACEMENTS WITH DUAL-MOBILITY BEARINGS AND ACETABULAR COMPONENT RETENTION", NINETY FOUR (94) METAL ON METAL BIRMINGHAM HIP REPLACEMENTS UNDERWENT REVISION SURGERY BETWEEN 2012 AND 2020. FROM THESE, FIVE (5) HIPS WERE REVISED DUE TO ASEPTIC LOOSENING OF THE FEMORAL HEAD COMPONENTS. NO FURTHER INFORMATION IS AVAILABLE. |  | 2025 | 2 | 2025-02
3005975929-2025-00007 | 2025-02-14 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP RESURFACING (BHR) CUP | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Unknown | IT WAS REPORTED THAT, ON THE LITERATURE REVIEW "REVISION OF METAL-ON-METAL HIP REPLACEMENTS WITH DUAL-MOBILITY BEARINGS AND ACETABULAR COMPONENT RETENTION", NINETY-FOUR (94) METAL ON METAL BIRMINGHAM HIP REPLACEMENTS UNDERWENT REVISION SURGERY BETWEEN 2012 AND 2020. FROM THESE, EIGHTY-FIVE (85) HIPS WERE REVISED DUE TO ARMD RELATED SYMPTOMS INCLUDING PAIN, RISING SERUM METAL IONS AND INCREASING SIZE OF ASSOCIATED COLLECTIONS. NO FURTHER INFORMATION IS AVAILABLE. |  | 2025 | 2 | 2025-02
3005975929-2025-00009 | 2025-02-28 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Unknown | IT WAS REPORTED THAT, AFTER A THR SURGERY HAD BEEN PERFORMED IN 2006, THE PATIENT STARTED HAVING PROBLEMS WITHIN A COUPLE MONTHS BUT NOTHING EXTREME. PROBLEMS GOT WORSE OVER TIME AND PATIENT HAD A TEST FOR METALLOSIS. PATIENT HAD ELEVATED LEVELS AND LIVED BACK ON PAIN MEDS TO DEAL WITH THE PROBLEMS FROM THE HIP. PATIENT ALSO DISCOVERED THAT SHE HAD PARTS WHICH WERE RECALLED WHICH LEFT THE PATIENT TO DISABILITY AND SUFFERING. CURRENT HEALTH STATUS OF PATIENT IS UNKNOWN. |  | 2025 | 2 | 2025-02
3005975929-2025-00022 | 2025-03-19 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "BIRMINGHAM HIP RESURFACING AT 25 YEARS", THREE (3) PATIENTS AFTER A BHR HIP RESURFACING PROCEDURE, SUFFERED A FEMORAL PERIPROSTHETIC FRACTURE AT A MEAN 15 YEARS AND 11 MONTHS FROM POST-SURGERY. IT IS UNKNOWN HOW PATIENTS WERE TREATED. THE OUTCOME OF THE PATIENTS IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2025 | 3 | 2025-03
3005975929-2025-00012 | 2025-03-19 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW " BLOOD METAL ION LEVELS AFTER HIP RESURFACING: A COMPARISON OF 2 DIFFERENT IMPLANTS", ONE (1) PATIENT AFTER A BHR HIP RESURFACING PROCEDURE HAD CO AND CR LEVELS OF 16,8 PPB AND 16,3 PPB AT 1-2 YEARS POSTOPERATIVE AND 16,7 PPB AND 16,1 PPB AT THE 3-5 YEARS FOLLOW UP. THE EVENT WAS TREATED VIA REVISION SURGERY. THE OUTCOME OF THE PATIENT IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2025 | 3 | 2025-03
3005975929-2025-00019 | 2025-03-19 | 2017-01-01 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 61 YR | Male | IT WAS REPORTED THAT ON LITERATURE REVIEW "BIRMINGHAM HIP RESURFACING AT 25 YEARS", ONE (1) PATIENT AFTER 19 YEARS OF A BHR HIP RESURFACING PROCEDURE, DEVELOPED A PAINFUL ADVERSE REACTION TO METAL DEBRIS (ARMD) DESPITE NORMAL SERUM METAL ION LEVELS, WITH A TENSE FLUID COLLECTION NOTED PRE- AND INTRAOPERATIVELY, COMBINED WITH LYSIS OF THE POSTERIOR COLUMN. PATIENT REQUIRED REVISION SURGERY TO A THA RETAINING THE ACETABULAR SHELL AND USING A CEMENTED FEMORAL STEM WITH A DUAL MOBILITY BEARING. NO FURTHER INFORMATION IS AVAILABLE. |  | 2025 | 3 | 2025-03
3005975929-2025-00017 | 2025-03-19 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 43 YR | Male | IT WAS REPORTED THAT ON LITERATURE REVIEW "BIRMINGHAM HIP RESURFACING AT 25 YEARS", ONE (1) PATIENT AFTER 14 YEARS OF A LEFT BHR HIP RESURFACING PROCEDURE DEVELOPED LYSIS AROUND THE ACETABULAR COMPONENT, WITH COMPONENT MIGRATION AND AN INCREASED INCLINATION ANGLE. PATIENT CONTINUED TO WORK WITH SOME DISCOMFORT. HIS OHS WAS REDUCED (OHS = 0) AND SERUM METAL IONS WERE ELEVATED (CR 2,521 NMOL/L, CO 926 NMOL/L). PATIENT REQUIRED REVISION SURGERY 15 YEARS AFTER INDEX SURGERY WITH BOTH COMPONENTS (ACETABULAR CUP AND FEMORAL HEAD) BEING REVISED. AT LAST FOLLOW-UP, METAL ION LEVELS REDUCED (CR 279 NMOL/L, CO 164 NMOL/L) AND OHS IMPROVED TO 15. NO FURTHER INFORMATION IS AVAILABLE. |  | 2025 | 3 | 2025-03
3005975929-2025-00013 | 2025-03-19 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Female | IT WAS REPORTED THAT ON LITERATURE REVIEW " BLOOD METAL ION LEVELS AFTER HIP RESURFACING: A COMPARISON OF 2 DIFFERENT IMPLANTS", ONE (1) PATIENT AFTER A RIGHT BHR HIP RESURFACING PROCEDURE HAD CO LEVELS OF 3,7 PPB AND CR LEVELS OF 3,0 PPB AT <3 YEARS POSTOPERATIVE AND AT 3-5 YEARS POSTOPERATIVE FOLLOW-UPS. THE EVENT WAS TREATED VIA REVISION SURGERY. THE OUTCOME OF THE PATIENT IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2025 | 3 | 2025-03
3005975929-2025-00014 | 2025-03-19 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Female | IT WAS REPORTED THAT ON LITERATURE REVIEW " BLOOD METAL ION LEVELS AFTER HIP RESURFACING: A COMPARISON OF 2 DIFFERENT IMPLANTS", ONE (1) PATIENT AFTER A LEFT BHR HIP RESURFACING PROCEDURE HAD CO LEVELS OF 3,7 PPB AND CR LEVELS OF 3,0 PPB AT <3 YEARS POSTOPERATIVE AND AT 3-5 YEARS POSTOPERATIVE FOLLOW-UPS. THE EVENT WAS TREATED VIA REVISION SURGERY. THE OUTCOME OF THE PATIENT IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2025 | 3 | 2025-03
3005975929-2025-00015 | 2025-03-19 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 41 YR | Female | IT WAS REPORTED THAT ON LITERATURE REVIEW "BIRMINGHAM HIP RESURFACING AT 25 YEARS", ONE (1) PATIENT AFTER A RIGHT BHR HIP RESURFACING PROCEDURE HAD A REVISION SURGERY DUE TO UNSPECIFIED REASONS. THE OUTCOME OF THE PATIENT IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. |  | 2025 | 3 | 2025-03
3005975929-2025-00020 | 2025-03-19 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 71 YR | Unknown | IT WAS REPORTED THAT ON LITERATURE REVIEW "BIRMINGHAM HIP RESURFACING AT 25 YEARS", ONE (1) PATIENT AFTER A BHR HIP RESURFACING PROCEDURE, SUFFERED A PERIPROSTHETIC NECK FRACTURE 13 YEARS AND TEN MONTHS POST-SURGERY. PATIENT REQUIRED REVISION SURGERY. NO FURTHER INFORMATION IS AVAILABLE. |  | 2025 | 3 | 2025-03
3005975929-2025-00016 | 2025-03-19 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 41 YR | Female | IT WAS REPORTED THAT ON LITERATURE REVIEW "BIRMINGHAM HIP RESURFACING AT 25 YEARS", ONE (1) PATIENT AFTER A LEFT BHR HIP RESURFACING PROCEDURE THE PATIENT COMPLAINT OF PAIN DURING WALKING AND PROLONGED SITTING AND HAD ELEVATED SERUM METAL IONS (CR 438 NMOL/L, CO 552 NMOL/L). PATIENT REQUIRED REVISION SURGERY 18 YEARS AND SIX MONTHS AFTER IMPLANTATION TO AN UNCEMENTED THA WITH CERAMIC-ON-POLYETHYLENE BEARING. PATIENT´S METAL IONS ONE-YEAR POST-SURGERY REDUCED TO NORMAL LEVELS (CR 44 NMOL/L, CO 17 NMOL/L), AND REMAINS WELL. NO FURTHER INFORMATION IS AVAILABLE. |  | 2025 | 3 | 2025-03
3005975929-2025-00027 | 2025-04-11 | 2025-02-19 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKNOWN BIRMINGHAM HIP IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Unknown | IT WAS REPORTED THAT, AFTER A THR PERFORMED IN 2007, IN (B)(6) 2024, AFTER 5 YEARS OF TERRIBLE GROIN PAIN, THE PATIENT HAD A ULTRASOUND WHICH A POCKET OF BLOOD WAS FOUND WITH METAL SHAVINGS. COBALT TESTING FOLLOWED WITH VERY HIGH LEVELS OF COBALT. POISONING STARTING AT 111 AND LAST TEST WAS AT 190. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2025 TO TREAT THE PATIENT'S PAIN. THE CURRENT HEALTH STATUS OF THE PATIENT IS UNKNOWN. |  | 2025 | 4 | 2025-04
3005975929-2025-00029 | 2025-04-23 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 55 YR | Unknown | Unique Complaint ID Number,Event Date,Date Entered,Manufacturer Aware Date,Brand Name,Generic Name,Model Number,Lot Number,Catalog Number,Serial Number,UDI Number,PMA / 510K Number,Date Returned to Manufacturer,Type of Reportable Event,Event Description,Manufacturer Narrative,Medical History,Patient Age,Patient Gender,Patient Weight,Date Implanted,Date Explanted,Health Effect Clinical Code,Health Effect Impact Code,Device Problem Code,Device Component Code,Investigation Type Code,Investigation Findings Code,Investigation Conclusion Code,Remedial Action Type,Latest Line Item Version;CASE-2025-00262305-1,Unknown,4/23/2025,3/21/2025,Birmingham Hip Resurfacing (BHR) Heads,Birmingham Hip Resurfacing (BHR) Heads,Unknown,Unknown,Unknown,Unknown,Unknown,P040033,N/A,IN,"It was reported that, in the Endoprothesenregister Joint Registry (EPRD) from Germany, a total of four hundred and twenty-six (426) hips underwent primary hip resurfacing procedures between 01-Jul-2014 and 31-Jan-2024, using Birmingham Hip Resurfacing (BHR) Heads. From these, one (1) patient required a revision surgery duo to infection.","Reporting Quarter: 1 (January 1 - March 31, 2025);;Summary of adverse events: It was reported that, in the Endoprothesenregister Joint Registry (EPRD) from Germany, a total of four hundred and twenty-six (426) hips underwent primary hip resurfacing procedures between 01-Jul-2014 and 31-Jan-2024, using Birmingham Hip Resurfacing (BHR) Heads. From these, twelve (12) hips were later revised due to the following complications: two (2) hips due to infection, one (1) hip due to aseptic loosening, one (1) hip due to wear and eight (8) hips due to other-unknown reasons. No further information is available. ;;Timeframe of Registry data: Implantations conducted between 01-Jul-2014 and 31-Jan-2024 in Germany.;;Analysis Conducted: Based on the most recent safety and performance evaluation, the Birmingham Hip Resurfacing (BHR) system presents a favorable benefit/risk assessment when used under the conditions and for the purposes intended by the manufacturer and with respect of the contraindications and precautions for use and warnings described in the information material supplied with the devices. This system is at least as safe and effective as all its therapeutic alternatives in the targeted indication. The intended purpose and risk reduction measures included in the information materials supplied by the manufacturer are adequate and suitable for the intended uses and users.;;According to this registry report, a total of four hundred and twenty-six (426) primary hip resurfacing procedures with BHR Resurfacing Heads have been performed in Germany between 01-Jul-2014 and 31-Jan-2024. The mean cumulative revision rate for BHR Resurfacing Heads were in line with the class device (elective THA with cemented stems) across 8 years of follow-up, based on the overlapping confidence intervals. The following cumulative revision rates with 95% confidence intervals are presented in this report: ;;o At 1st postoperative year: 1.4% (0.3%-2.6%) vs 2.4% (2.3%-2.5%) of the class.;o At 2nd postoperative year: 2.0% (0.6%-3.4%) vs 2.7% (2.6%-2.8%) of the class.;o At 3rd postoperative year: 2.3% (0.8%-3.8%) vs 2.9% (2.8%-3.0%) of the class.;o At 4th postoperative year: 2.7% (1.0%-4.3%) vs 3.2% (3.1%-3.3%) of the class.;o At 5th postoperative year: 2.7% (1.0%-4.3%) vs 3.4% (3.3%-3.5%) of the class.;o At 6th postoperative year: 2.7% (1.0%-4.3%) vs 3.6% (3.5%-3.8%) of the class.;o At 7th postoperative year: 4.4% (1.4%-7.4%) vs 3.9% (3.7%-4.0%) of the class.;o At 8th postoperative year: 4.4% (1.4%-7.4%) vs 4.1% (3.9%-4.3%) of the class.;;;Based on the review of other clinical sources such as clinical activities, complaint data, published literature and other joint registries, no increased risks to health or an increased trend in any of the adverse events summarized above have been identified. The reported adverse events relate to known inherent procedural risks that are appropriately documented in our risk files. No individual investigations into the reported adverse events are deemed necessary. ;;Additional Action(s) taken: No additional actions are deemed necessary at this time. Smith+Nephew will continue to monitor trends in accordance with our post-market surveillance process and take necessary action as required if anticipated severity and/or occurrence rates are exceeded.",Unknown,55,Unknown,Unknown,Unknown,Unknown,E1906,F1905,A24,G07001,B20;B22,C19,D12;D15,N/A,0;CASE-2025-00262305-1,Unknown,4/23/2025,3/21/2025,Birmingham Hip Resurfacing (BHR) Heads,Birmingham Hip Resurfacing (BHR) Heads,Unknown,Unknown,Unknown,Unknown,Unknown,P040033,N/A,IN,"It was reported that, in the Endoprothesenregister Joint Registry (EPRD) from Germany, a total of four hundred and twenty-six (426) hips underwent primary hip resurfacing procedures between 01-Jul-2014 and 31-Jan-2024, using Birmingham Hip Resurfacing (BHR) Heads. From these, one (1) patient required a revision surgery duo to infection.","Reporting Quarter: 1 (January 1 - March 31, 2025);;Summary of adverse events: It was reported that, in the Endoprothesenregister Joint Registry (EPRD) from Germany, a total of four hundred and twenty-six (426) hips underwent primary hip resurfacing procedures between 01-Jul-2014 and 31-Jan-2024, using Birmingham Hip Resurfacing (BHR) Heads. From these, twelve (12) hips were later revised due to the following complications: two (2) hips due to infection, one (1) hip due to aseptic loosening, one (1) hip due to wear and eight (8) hips due to other-unknown reasons. No further information is available. ;;Timeframe of Registry data: Implantations conducted between 01-Jul-2014 and 31-Jan-2024 in Germany.;;Analysis Conducted: Based on the most recent safety and performance evaluation, the Birmingham Hip Resurfacing (BHR) system presents a favorable benefit/risk assessment when used under the conditions and for the purposes intended by the manufacturer and with respect of the contraindications and precautions for use and warnings described in the information material supplied with the devices. This system is at least as safe and effective as all its therapeutic alternatives in the targeted indication. The intended purpose and risk reduction measures included in the information materials supplied by the manufacturer are adequate and suitable for the intended uses and users.;;According to this registry report, a total of four hundred and twenty-six (426) primary hip resurfacing procedures with BHR Resurfacing Heads have been performed in Germany between 01-Jul-2014 and 31-Jan-2024. The mean cumulative revision rate for BHR Resurfacing Heads were in line with the class device (elective THA with cemented stems) across 8 years of follow-up, based on the overlapping confidence intervals. The following cumulative revision rates with 95% confidence intervals are presented in this report: ;;o At 1st postoperative year: 1.4% (0.3%-2.6%) vs 2.4% (2.3%-2.5%) of the class.;o At 2nd postoperative year: 2.0% (0.6%-3.4%) vs 2.7% (2.6%-2.8%) of the class.;o At 3rd postoperative year: 2.3% (0.8%-3.8%) vs 2.9% (2.8%-3.0%) of the class.;o At 4th postoperative year: 2.7% (1.0%-4.3%) vs 3.2% (3.1%-3.3%) of the class.;o At 5th postoperative year: 2.7% (1.0%-4.3%) vs 3.4% (3.3%-3.5%) of the class.;o At 6th postoperative year: 2.7% (1.0%-4.3%) vs 3.6% (3.5%-3.8%) of the class.;o At 7th postoperative year: 4.4% (1.4%-7.4%) vs 3.9% (3.7%-4.0%) of the class.;o At 8th postoperative year: 4.4% (1.4%-7.4%) vs 4.1% (3.9%-4.3%) of the class.;;;Based on the review of other clinical sources such as clinical activities, complaint data, published literature and other joint registries, no increased risks to health or an increased trend in any of the adverse events summarized above have been identified. The reported adverse events relate to known inherent procedural risks that are appropriately documented in our risk files. No individual investigations into the reported adverse events are deemed necessary. ;;Additional Action(s) taken: No additional actions are deemed necessary at this time. Smith+Nephew will continue to monitor trends in accordance with our post-market surveillance process and take necessary action as required if anticipated severity and/or occurrence rates are exceeded.",Unknown,55,Unknown,Unknown,Unknown,Unknown,E1906,F1905,A24,G07001,B20;B22,C19,D12;D15,N/A,0;CASE-2025-00262305-1,Unknown,4/23/2025,3/21/2025,Birmingham Hip Resurfacing (BHR) Heads,Birmingham Hip Resurfacing (BHR) Heads,Unknown,Unknown,Unknown,Unknown,Unknown,P040033,N/A,IN,"It was reported that, in the Endoprothesenregister Joint Registry (EPRD) from Germany, a total of four hundred and twenty-six (426) hips underwent primary hip resurfacing procedures between 01-Jul-2014 and 31-Jan-2024, using Birmingham Hip Resurfacing (BHR) Heads. From these, one (1) patient required a revision surgery duo to aseptic loosening.","Reporting Quarter: 1 (January 1 - March 31, 2025);;Summary of adverse events: It was reported that, in the Endoprothesenregister Joint Registry (EPRD) from Germany, a total of four hundred and twenty-six (426) hips underwent primary hip resurfacing procedures between 01-Jul-2014 and 31-Jan-2024, using Birmingham Hip Resurfacing (BHR) Heads. From these, twelve (12) hips were later revised due to the following complications: two (2) hips due to infection, one (1) hip due to aseptic loosening, one (1) hip due to wear and eight (8) hips due to other-unknown reasons. No further information is available. ;;Timeframe of Registry data: Implantations conducted between 01-Jul-2014 and 31-Jan-2024 in Germany.;;Analysis Conducted: Based on the most recent safety and performance evaluation, the Birmingham Hip Resurfacing (BHR) system presents a favorable benefit/risk assessment when used under the conditions and for the purposes intended by the manufacturer and with respect of the contraindications and precautions for use and warnings described in the information material supplied with the devices. This system is at least as safe and effective as all its therapeutic alternatives in the targeted indication. The intended purpose and risk reduction measures included in the information materials supplied by the manufacturer are adequate and suitable for the intended uses and users.;;According to this registry report, a total of four hundred and twenty-six (426) primary hip resurfacing procedures with BHR Resurfacing Heads have been performed in Germany between 01-Jul-2014 and 31-Jan-2024. The mean cumulative revision rate for BHR Resurfacing Heads were in line with the class device (elective THA with cemented stems) across 8 years of follow-up, based on the overlapping confidence intervals. The following cumulative revision rates with 95% confidence intervals are presented in this report: ;;o At 1st postoperative year: 1.4% (0.3%-2.6%) vs 2.4% (2.3%-2.5%) of the class.;o At 2nd postoperative year: 2.0% (0.6%-3.4%) vs 2.7% (2.6%-2.8%) of the class.;o At 3rd postoperative year: 2.3% (0.8%-3.8%) vs 2.9% (2.8%-3.0%) of the class.;o At 4th postoperative year: 2.7% (1.0%-4.3%) vs 3.2% (3.1%-3.3%) of the class.;o At 5th postoperative year: 2.7% (1.0%-4.3%) vs 3.4% (3.3%-3.5%) of the class.;o At 6th postoperative year: 2.7% (1.0%-4.3%) vs 3.6% (3.5%-3.8%) of the class.;o At 7th postoperative year: 4.4% (1.4%-7.4%) vs 3.9% (3.7%-4.0%) of the class.;o At 8th postoperative year: 4.4% (1.4%-7.4%) vs 4.1% (3.9%-4.3%) of the class.;;;Based on the review of other clinical sources such as clinical activities, complaint data, published literature and other joint registries, no increased risks to health or an increased trend in any of the adverse events summarized above have been identified. The reported adverse events relate to known inherent procedural risks that are appropriately documented in our risk files. No individual investigations into the reported adverse events are deemed necessary. ;;Additional Action(s) taken: No additional actions are deemed necessary at this time. Smith+Nephew will continue to monitor trends in accordance with our post-market surveillance process and take necessary action as required if anticipated severity and/or occurrence rates are exceeded.",Unknown,55,Unknown,Unknown,Unknown,Unknown,E161201,F1905,A0103,G07001,B20;B22,C19,D12;D15,N/A,0;CASE-2025-00262305-1,Unknown,4/23/2025,3/21/2025,Birmingham Hip Resurfacing (BHR) Heads,Birmingham Hip Resurfacing (BHR) Heads,Unknown,Unknown,Unknown,Unknown,Unknown,P040033,N/A,IN,"It was reported that, in the Endoprothesenregister Joint Registry (EPRD) from Germany, a total of four hundred and twenty-six (426) hips underwent primary hip resurfacing procedures between 01-Jul-2014 and 31-Jan-2024, using Birmingham Hip Resurfacing (BHR) Heads. From these, one (1) patient required a revision surgery duo to wear.","Reporting Quarter: 1 (January 1 - March 31, 2025);;Summary of adverse events: It was reported that, in the Endoprothesenregister Joint Registry (EPRD) from Germany, a total of four hundred and twenty-six (426) hips underwent primary hip resurfacing procedures between 01-Jul-2014 and 31-Jan-2024, using Birmingham Hip Resurfacing (BHR) Heads. From these, twelve (12) hips were later revised due to the following complications: two (2) hips due to infection, one (1) hip due to aseptic loosening, one (1) hip due to wear and eight (8) hips due to other-unknown reasons. No further information is available. ;;Timeframe of Registry data: Implantations conducted between 01-Jul-2014 and 31-Jan-2024 in Germany.;;Analysis Conducted: Based on the most recent safety and performance evaluation, the Birmingham Hip Resurfacing (BHR) system presents a favorable benefit/risk assessment when used under the conditions and for the purposes intended by the manufacturer and with respect of the contraindications and precautions for use and warnings described in the information material supplied with the devices. This system is at least as safe and effective as all its therapeutic alternatives in the targeted indication. The intended purpose and risk reduction measures included in the information materials supplied by the manufacturer are adequate and suitable for the intended uses and users.;;According to this registry report, a total of four hundred and twenty-six (426) primary hip resurfacing procedures with BHR Resurfacing Heads have been performed in Germany between 01-Jul-2014 and 31-Jan-2024. The mean cumulative revision rate for BHR Resurfacing Heads were in line with the class device (elective THA with cemented stems) across 8 years of follow-up, based on the overlapping confidence intervals. The following cumulative revision rates with 95% confidence intervals are presented in this report: ;;o At 1st postoperative year: 1.4% (0.3%-2.6%) vs 2.4% (2.3%-2.5%) of the class.;o At 2nd postoperative year: 2.0% (0.6%-3.4%) vs 2.7% (2.6%-2.8%) of the class.;o At 3rd postoperative year: 2.3% (0.8%-3.8%) vs 2.9% (2.8%-3.0%) of the class.;o At 4th postoperative year: 2.7% (1.0%-4.3%) vs 3.2% (3.1%-3.3%) of the class.;o At 5th postoperative year: 2.7% (1.0%-4.3%) vs 3.4% (3.3%-3.5%) of the class.;o At 6th postoperative year: 2.7% (1.0%-4.3%) vs 3.6% (3.5%-3.8%) of the class.;o At 7th postoperative year: 4.4% (1.4%-7.4%) vs 3.9% (3.7%-4.0%) of the class.;o At 8th postoperative year: 4.4% (1.4%-7.4%) vs 4.1% (3.9%-4.3%) of the class.;;;Based on the review of other clinical sources such as clinical activities, complaint data, published literature and other joint registries, no increased risks to health or an increased trend in any of the adverse events summarized above have been identified. The reported adverse events relate to known inherent procedural risks that are appropriately documented in our risk files. No individual investigations into the reported adverse events are deemed necessary. ;;Additional Action(s) taken: No additional actions are deemed necessary at this time. Smith+Nephew will continue to monitor trends in accordance with our post-market surveillance process and take necessary action as required if anticipated severity and/or occurrence rates are exceeded.",Unknown,55,Unknown,Unknown,Unknown,Unknown,E2401,F1905,A040503,G07001,B20;B22,C19,D12;D15,N/A,0;CASE-2025-00262305-1,Unknown,4/23/2025,3/21/2025,Birmingham Hip Resurfacing (BHR) Heads,Birmingham Hip Resurfacing (BHR) Heads,Unknown,Unknown,Unknown,Unknown,Unknown,P040033,N/A,IN,"It was reported that, in the Endoprothesenregister Joint Registry (EPRD) from Germany, a total of four hundred and twenty-six (426) hips underwent primary hip resurfacing procedures between 01-Jul-2014 and 31-Jan-2024, using Birmingham Hip Resurfacing (BHR) Heads. From these, one (1) patient required a revision surgery duo to other-unknown reasons.","Reporting Quarter: 1 (January 1 - March 31, 2025);;Summary of adverse events: It was reported that, in the Endoprothesenregister Joint Registry (EPRD) from Germany, a total of four hundred and twenty-six (426) hips underwent primary hip resurfacing procedures between 01-Jul-2014 and 31-Jan-2024, using Birmingham Hip Resurfacing (BHR) Heads. From these, twelve (12) hips were later revised due to the following complications: two (2) hips due to infection, one (1) hip due to aseptic loosening, one (1) hip due to wear and eight (8) hips due to other-unknown reasons. No further information is available. ;;Timeframe of Registry data: Implantations conducted between 01-Jul-2014 and 31-Jan-2024 in Germany.;;Analysis Conducted: Based on the most recent safety and performance evaluation, the Birmingham Hip Resurfacing (BHR) system presents a favorable benefit/risk assessment when used under the conditions and for the purposes intended by the manufacturer and with respect of the contraindications and precautions for use and warnings described in the information material supplied with the devices. This system is at least as safe and effective as all its therapeutic alternatives in the targeted indication. The intended purpose and risk reduction measures included in the information materials supplied by the manufacturer are adequate and suitable for the intended uses and users.;;According to this registry report, a total of four hundred and twenty-six (426) primary hip resurfacing procedures with BHR Resurfacing Heads have been performed in Germany between 01-Jul-2014 and 31-Jan-2024. The mean cumulative revision rate for BHR Resurfacing Heads were in line with the class device (elective THA with cemented stems) across 8 years of follow-up, based on the overlapping confidence intervals. The following cumulative revision rates with 95% confidence intervals are presented in this report: ;;o At 1st postoperative year: 1.4% (0.3%-2.6%) vs 2.4% (2.3%-2.5%) of the class.;o At 2nd postoperative year: 2.0% (0.6%-3.4%) vs 2.7% (2.6%-2.8%) of the class.;o At 3rd postoperative year: 2.3% (0.8%-3.8%) vs 2.9% (2.8%-3.0%) of the class.;o At 4th postoperative year: 2.7% (1.0%-4.3%) vs 3.2% (3.1%-3.3%) of the class.;o At 5th postoperative year: 2.7% (1.0%-4.3%) vs 3.4% (3.3%-3.5%) of the class.;o At 6th postoperative year: 2.7% (1.0%-4.3%) vs 3.6% (3.5%-3.8%) of the class.;o At 7th postoperative year: 4.4% (1.4%-7.4%) vs 3.9% (3.7%-4.0%) of the class.;o At 8th postoperative year: 4.4% (1.4%-7.4%) vs 4.1% (3.9%-4.3%) of the class.;;;Based on the review of other clinical sources such as clinical activities, complaint data, published literature and other joint registries, no increased risks to health or an increased trend in any of the adverse events summarized above have been identified. The reported adverse events relate to known inherent procedural risks that are appropriately documented in our risk files. No individual investigations into the reported adverse events are deemed necessary. ;;Additional Action(s) taken: No additional actions are deemed necessary at this time. Smith+Nephew will continue to monitor trends in accordance with our post-market surveillance process and take necessary action as required if anticipated severity and/or occurrence rates are exceeded.",Unknown,55,Unknown,Unknown,Unknown,Unknown,E2401,F1905,A24,G07001,B20;B22,C19,D12;D15,N/A,0;CASE-2025-00262305-1,Unknown,4/23/2025,3/21/2025,Birmingham Hip Resurfacing (BHR) Heads,Birmingham Hip Resurfacing (BHR) Heads,Unknown,Unknown,Unknown,Unknown,Unknown,P040033,N/A,IN,"It was reported that, in the Endoprothesenregister Joint Registry (EPRD) from Germany, a total of four hundred and twenty-six (426) hips underwent primary hip resurfacing procedures between 01-Jul-2014 and 31-Jan-2024, using Birmingham Hip Resurfacing (BHR) Heads. From these, one (1) patient required a revision surgery duo to other-unknown reasons.","Reporting Quarter: 1 (January 1 - March 31, 2025);;Summary of adverse events: It was reported that, in the Endoprothesenregister Joint Registry (EPRD) from Germany, a total of four hundred and twenty-six (426) hips underwent primary hip resurfacing procedures between 01-Jul-2014 and 31-Jan-2024, using Birmingham Hip Resurfacing (BHR) Heads. From these, twelve (12) hips were later revised due to the following complications: two (2) hips due to infection, one (1) hip due to aseptic loosening, one (1) hip due to wear and eight (8) hips due to other-unknown reasons. No further information is available. ;;Timeframe of Registry data: Implantations conducted between 01-Jul-2014 and 31-Jan-2024 in Germany.;;Analysis Conducted: Based on the most recent safety and performance evaluation, the Birmingham Hip Resurfacing (BHR) system presents a favorable benefit/risk assessment when used under the conditions and for the purposes intended by the manufacturer and with respect of the contraindications and precautions for use and warnings described in the information material supplied with the devices. This system is at least as safe and effective as all its therapeutic alternatives in the targeted indication. The intended purpose and risk reduction measures included in the information materials supplied by the manufacturer are adequate and suitable for the intended uses and users.;;According to this registry report, a total of four hundred and twenty-six (426) primary hip resurfacing procedures with BHR Resurfacing Heads have been performed in Germany between 01-Jul-2014 and 31-Jan-2024. The mean cumulative revision rate for BHR Resurfacing Heads were in line with the class device (elective THA with cemented stems) across 8 years of follow-up, based on the overlapping confidence intervals. The following cumulative revision rates with 95% confidence intervals are presented in this report: ;;o At 1st postoperative year: 1.4% (0.3%-2.6%) vs 2.4% (2.3%-2.5%) of the class.;o At 2nd postoperative year: 2.0% (0.6%-3.4%) vs 2.7% (2.6%-2.8%) of the class.;o At 3rd postoperative year: 2.3% (0.8%-3.8%) vs 2.9% (2.8%-3.0%) of the class.;o At 4th postoperative year: 2.7% (1.0%-4.3%) vs 3.2% (3.1%-3.3%) of the class.;o At 5th postoperative year: 2.7% (1.0%-4.3%) vs 3.4% (3.3%-3.5%) of the class.;o At 6th postoperative year: 2.7% (1.0%-4.3%) vs 3.6% (3.5%-3.8%) of the class.;o At 7th postoperative year: 4.4% (1.4%-7.4%) vs 3.9% (3.7%-4.0%) of the class.;o At 8th postoperative year: 4.4% (1.4%-7.4%) vs 4.1% (3.9%-4.3%) of the class.;;;Based on the review of other clinical sources such as clinical activities, complaint data, published literature and other joint registries, no increased risks to health or an increased trend in any of the adverse events summarized above have been identified. The reported adverse events relate to known inherent procedural risks that are appropriately documented in our risk files. No individual investigations into the reported adverse events are deemed necessary. ;;Additional Action(s) taken: No additional actions are deemed necessary at this time. Smith+Nephew will continue to monitor trends in accordance with our post-market surveillance process and take necessary action as required if anticipated severity and/or occurrence rates are exceeded.",Unknown,55,Unknown,Unknown,Unknown,Unknown,E2401,F1905,A24,G07001,B20;B22,C19,D12;D15,N/A,0;CASE-2025-00262305-1,Unknown,4/23/2025,3/21/2025,Birmingham Hip Resurfacing (BHR) Heads,Birmingham Hip Resurfacing (BHR) Heads,Unknown,Unknown,Unknown,Unknown,Unknown,P040033,N/A,IN,"It was reported that, in the Endoprothesenregister Joint Registry (EPRD) from Germany, a total of four hundred and twenty-six (426) hips underwent primary hip resurfacing procedures between 01-Jul-2014 and 31-Jan-2024, using Birmingham Hip Resurfacing (BHR) Heads. From these, one (1) patient required a revision surgery duo to other-unknown reasons.","Reporting Quarter: 1 (January 1 - March 31, 2025);;Summary of adverse events: It was reported that, in the Endoprothesenregister Joint Registry (EPRD) from Germany, a total of four hundred and twenty-six (426) hips underwent primary hip resurfacing procedures between 01-Jul-2014 and 31-Jan-2024, using Birmingham Hip Resurfacing (BHR) Heads. From these, twelve (12) hips were later revised due to the following complications: two (2) hips due to infection, one (1) hip due to aseptic loosening, one (1) hip due to wear and eight (8) hips due to other-unknown reasons. No further information is available. ;;Timeframe of Registry data: Implantations conducted between 01-Jul-2014 and 31-Jan-2024 in Germany.;;Analysis Conducted: Based on the most recent safety and performance evaluation, the Birmingham Hip Resurfacing (BHR) system presents a favorable benefit/risk assessment when used under the conditions and for the purposes intended by the manufacturer and with respect of the contraindications and precautions for use and warnings described in the information material supplied with the devices. This system is at least as safe and effective as all its therapeutic alternatives in the targeted indication. The intended purpose and risk reduction measures included in the information materials supplied by the manufacturer are adequate and suitable for the intended uses and users.;;According to this registry report, a total of four hundred and twenty-six (426) primary hip resurfacing procedures with BHR Resurfacing Heads have been performed in Germany between 01-Jul-2014 and 31-Jan-2024. The mean cumulative revision rate for BHR Resurfacing Heads were in line with the class device (elective THA with cemented stems) across 8 years of follow-up, based on the overlapping confidence intervals. The following cumulative revision rates with 95% confidence intervals are presented in this report: ;;o At 1st postoperative year: 1.4% (0.3%-2.6%) vs 2.4% (2.3%-2.5%) of the class.;o At 2nd postoperative year: 2.0% (0.6%-3.4%) vs 2.7% (2.6%-2.8%) of the class.;o At 3rd postoperative year: 2.3% (0.8%-3.8%) vs 2.9% (2.8%-3.0%) of the class.;o At 4th postoperative year: 2.7% (1.0%-4.3%) vs 3.2% (3.1%-3.3%) of the class.;o At 5th postoperative year: 2.7% (1.0%-4.3%) vs 3.4% (3.3%-3.5%) of the class.;o At 6th postoperative year: 2.7% (1.0%-4.3%) vs 3.6% (3.5%-3.8%) of the class.;o At 7th postoperative year: 4.4% (1.4%-7.4%) vs 3.9% (3.7%-4.0%) of the class.;o At 8th postoperative year: 4.4% (1.4%-7.4%) vs 4.1% (3.9%-4.3%) of the class.;;;Based on the review of other clinical sources such as clinical activities, complaint data, published literature and other joint registries, no increased risks to health or an increased trend in any of the adverse events summarized above have been identified. The reported adverse events relate to known inherent procedural risks that are appropriately documented in our risk files. No individual investigations into the reported adverse events are deemed necessary. ;;Additional Action(s) taken: No additional actions are deemed necessary at this time. Smith+Nephew will continue to monitor trends in accordance with our post-market surveillance process and take necessary action as required if anticipated severity and/or occurrence rates are exceeded.",Unknown,55,Unknown,Unknown,Unknown,Unknown,E2401,F1905,A24,G07001,B20;B22,C19,D12;D15,N/A,0;CASE-2025-00262305-1,Unknown,4/23/2025,3/21/2025,Birmingham Hip Resurfacing (BHR) Heads,Birmingham Hip Resurfacing (BHR) Heads,Unknown,Unknown,Unknown,Unknown,Unknown,P040033,N/A,IN,"It was reported that, in the Endoprothesenregister Joint Registry (EPRD) from Germany, a total of four hundred and twenty-six (426) hips underwent primary hip resurfacing procedures between 01-Jul-2014 and 31-Jan-2024, using Birmingham Hip Resurfacing (BHR) Heads. From these, one (1) patient required a revision surgery duo to other-unknown reasons.","Reporting Quarter: 1 (January 1 - March 31, 2025);;Summary of adverse events: It was reported that, in the Endoprothesenregister Joint Registry (EPRD) from Germany, a total of four hundred and twenty-six (426) hips underwent primary hip resurfacing procedures between 01-Jul-2014 and 31-Jan-2024, using Birmingham Hip Resurfacing (BHR) Heads. From these, twelve (12) hips were later revised due to the following complications: two (2) hips due to infection, one (1) hip due to aseptic loosening, one (1) hip due to wear and eight (8) hips due to other-unknown reasons. No further information is available. ;;Timeframe of Registry data: Implantations conducted between 01-Jul-2014 and 31-Jan-2024 in Germany.;;Analysis Conducted: Based on the most recent safety and performance evaluation, the Birmingham Hip Resurfacing (BHR) system presents a favorable benefit/risk assessment when used under the conditions and for the purposes intended by the manufacturer and with respect of the contraindications and precautions for use and warnings described in the information material supplied with the devices. This system is at least as safe and effective as all its therapeutic alternatives in the targeted indication. The intended purpose and risk reduction measures included in the information materials supplied by the manufacturer are adequate and suitable for the intended uses and users.;;According to this registry report, a total of four hundred and twenty-six (426) primary hip resurfacing procedures with BHR Resurfacing Heads have been performed in Germany between 01-Jul-2014 and 31-Jan-2024. The mean cumulative revision rate for BHR Resurfacing Heads were in line with the class device (elective THA with cemented stems) across 8 years of follow-up, based on the overlapping confidence intervals. The following cumulative revision rates with 95% confidence intervals are presented in this report: ;;o At 1st postoperative year: 1.4% (0.3%-2.6%) vs 2.4% (2.3%-2.5%) of the class.;o At 2nd postoperative year: 2.0% (0.6%-3.4%) vs 2.7% (2.6%-2.8%) of the class.;o At 3rd postoperative year: 2.3% (0.8%-3.8%) vs 2.9% (2.8%-3.0%) of the class.;o At 4th postoperative year: 2.7% (1.0%-4.3%) vs 3.2% (3.1%-3.3%) of the class.;o At 5th postoperative year: 2.7% (1.0%-4.3%) vs 3.4% (3.3%-3.5%) of the class.;o At 6th postoperative year: 2.7% (1.0%-4.3%) vs 3.6% (3.5%-3.8%) of the class.;o At 7th postoperative year: 4.4% (1.4%-7.4%) vs 3.9% (3.7%-4.0%) of the class.;o At 8th postoperative year: 4.4% (1.4%-7.4%) vs 4.1% (3.9%-4.3%) of the class.;;;Based on the review of other clinical sources such as clinical activities, complaint data, published literature and other joint registries, no increased risks to health or an increased trend in any of the adverse events summarized above have been identified. The reported adverse events relate to known inherent procedural risks that are appropriately documented in our risk files. No individual investigations into the reported adverse events are deemed necessary. ;;Additional Action(s) taken: No additional actions are deemed necessary at this time. Smith+Nephew will continue to monitor trends in accordance with our post-market surveillance process and take necessary action as required if anticipated severity and/or occurrence rates are exceeded.",Unknown,55,Unknown,Unknown,Unknown,Unknown,E2401,F1905,A24,G07001,B20;B22,C19,D12;D15,N/A,0;CASE-2025-00262305-1,Unknown,4/23/2025,3/21/2025,Birmingham Hip Resurfacing (BHR) Heads,Birmingham Hip Resurfacing (BHR) Heads,Unknown,Unknown,Unknown,Unknown,Unknown,P040033,N/A,IN,"It was reported that, in the Endoprothesenregister Joint Registry (EPRD) from Germany, a total of four hundred and twenty-six (426) hips underwent primary hip resurfacing procedures between 01-Jul-2014 and 31-Jan-2024, using Birmingham Hip Resurfacing (BHR) Heads. From these, one (1) patient required a revision surgery duo to other-unknown reasons.","Reporting Quarter: 1 (January 1 - March 31, 2025);;Summary of adverse events: It was reported that, in the Endoprothesenregister Joint Registry (EPRD) from Germany, a total of four hundred and twenty-six (42 |  | 2025 | 4 | 2025-04
3005975929-2025-00039 | 2025-05-08 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) HEAD | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 71 YR | Male | IT WAS REPORTED THAT, AFTER THE PLAINTIFF UNDERWENT A LEFT METAL-ON-METAL BHR ON (B)(6) 2012, THE PATIENT INDICATED THAT HE HAS EXPERIENCED SEVERAL DISLOCATIONS PER YEAR. IT IS UNKNOWN THE SPECIFIC QUANTITY OF DISLOCATIONS, AND HOW OR WHEN THESE DISLOCATIONS WERE TREATED. THE CURRENT HEALTH STATUS OF THE PATIENT IS UNKNOWN. |  | 2025 | 5 | 2025-05
3005975929-2025-00053 | 2025-05-19 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 62 YR | Male | IT WAS REPORTED THAT, THE PLAINTIFF UNDERWENT A RIGHT METAL-ON-METAL BHR PERFORMED ON (B)(6) 2012 AND A LEFT BHR PERFORMED ON AN UNSPECIFIED DATE, AND AFTER THIS THE PATIENT HAS BEEN EXPERIENCING PAIN IN HIS RIGHT HIP AND HIGH LEVELS OF CHROMIUM AND COBALT. HE HAS BEEN DIAGNOSED WITH COBALT POISONING DUE TO HIS METAL-ON-METAL HIP REPLACEMENT, WHICH HAS ALSO AFFECTED HIS OTHER HIP. A REVISION SURGERY FOR HIS LEFT HIP HAS BEEN DISCUSSED BUT HAS NOT BEEN PERFORMED YET, AS THE LEFT HIP IS NOT CAUSING PAIN. THE CURRENT HEALTH STATUS OF THE PATIENT IS UNKNOW. |  | 2025 | 5 | 2025-05
3005975929-2025-00054 | 2025-05-30 | 2012-08-14 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 57 YR | Male | IT WAS REPORTED THAT, A PLAINTIFF WAS IMPLANTED A RIGHT BHR SYSTEM ON (B)(6) 2012, AND AFTER THIS A POST-OP VISIT CONDUCTED ON (B)(6) 2012, IN WHICH IT WAS MENTIONED THAT THE PATIENT WAS EXPERIENCING MILD PAIN TO HIP AND A SLIGHT CLICKING SENSATION WHEN ARISING FROM SEATED POSITION, PATIENT HAD DISCONTINUED ALL PAIN MEDICATIONS AT THE MOMENT. ADDITIONALLY, A DECREASED RIGHT HIP ABDUCTION WAS OBSERVED, NEVERTHELESS, RADIOGRAPH REVEALED A BHR IN GOOD CONDITION AND ALIGNMENT. IN (B)(6) 2024 THE BLOOD TESTS INDICATED THAT THE PATIENT HAD HIGH COCR LEVELS. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2024, IN WHICH S+N IMPLANTS WERE EXPLANTED AND REPLACED FOR COMPETITOR IMPLANTS. THE CURRENT HEALTH STATUS OF THE PATIENT IS UNKNOWN. |  | 2025 | 5 | 2025-05
3005975929-2025-00055 | 2025-06-05 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 56 YR | Female | IT WAS REPORTED THAT, AFTER THE PLAINTIFF UNDERWENT A RIGHT METAL-ON-METAL BHR PERFORMED IN 2006, SHE HAD A REVISION SURGERY PERFORMED ON (B)(6) 2007 DUE TO UNKNOWN REASONS. FEMORAL HEAD AND CUP WERE RETAINED. IT IS UNKNOWN THE CURRENT HEALTH STATUS OF THE PATIENT. |  | 2025 | 6 | 2025-06
3005975929-2025-00058 | 2025-06-11 | 2023-12-05 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | BIRMINGHAM HIP RESURFACING (BHR) CUP (UNKNOWN TYPE) | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 72 YR | Female | IT WAS REPORTED THAT, AFTER THE PLAINTIFF UNDERWENT A RIGHT METAL-ON-METAL BHR PERFORMED IN 2006 AND A FIRST REVISION SURGERY ON (B)(6) 2007 DUE TO UNKNOWN REASONS, SHE HAD AN ASEPTIC LYMPHOCYTE-DOMINANT VASCULITIS, ASSOCIATED TO A LESION IN THE RIGHT HIP. SHE ALSO HAS HISTORY OF ELEVATED COBALT AND CHROMIUM LEVELS. A REVISION SURGERY WAS PERFORMED ON (B)(6) 2023, IN WHICH ABUNDANT AMOUNT OF SYNOVITIS WAS NOTED WITH METAL-STAINING DEBRIS. THE OLD HEAD WAS REMOVED AND A MILD AMOUNT OF CORROSION WAS NOTED. THE STEM WAS NOTED TO BE IN EXCELLENT POSITION AND IT WAS WELL FIXED, SO IT WAS RETAINED. THE TRUNNION WAS CLEANED AND THE HYPERTROPHIC SYNOVIUM AND SOFT TISSUE WAS DEBRIDED FROM ABOUT THE CUP. AN EXTENSIVE AMOUNT OF BURNISHING WAS NOTED ABOUT THE CUP. WITH THE AID OF THE ACETABULAR OSTEOTOMES THE CUP WAS REMOVED. SEVERE OSTEOLYSIS WAS NOTED ABOUT THE RETROACETABULAR BONE, THIS WAS DEBRIDED. COMPETITOR DEVICES AND S+N OXINIUM FEMORAL HEAD WERE IMPLANTED. IT IS UNKNOWN THE CURRENT HEALTH STATUS OF THE PATIENT. |  | 2025 | 6 | 2025-06
3005975929-2025-00069 | 2025-07-08 | 2019-09-24 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 40 YR | Male | IT WAS REPORTED THAT A PATIENT UNDERWENT A LEFT BIRMINGHAM HIP RESURFACING (BHR) PROCEDURE IN 2007. TWELVE YEARS POST-IMPLANTATION, THE PATIENT EXPERIENCED PAIN IN THE LEFT HIP RADIATING TO THE GROIN, WHICH LIMITED MOBILITY. IMAGING CONFIRMED A LOOSE PROSTHESIS. LABORATORY ANALYSIS SHOWED ELEVATED METAL ION LEVELS (CHROMIUM 29.4 MCG/L, COBALT 24.6 MCG/L). A REVISION SURGERY WAS SCHEDULED FOR (B)(6) 2019. NO PROCEDURAL DELAY WAS NOTED. NO INTRAOPERATIVE FINDINGS OR OUTCOMES WERE REPORTED. |  | 2025 | 7 | 2025-07
3005975929-2025-00080 | 2025-07-30 |  | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | NA | Male | IT WAS REPORTED THAT, AFTER A BHR SURGERY PERFORMED 1 YEAR AGO, THE PATIENT HAS NOW METAL LEVELS ABOVE ACCEPTABLE RANGES. THE PATIENT IS CURRENTLY WAITING FOR A REVISION SURGERY. |  | 2025 | 7 | 2025-07
3005975929-2025-00083 | 2025-08-26 | 2009-02-13 | I, n, j, u, r, y | Manufacturer report | SMITH & NEPHEW ORTHOPAEDICS LTD | UNKN BIRMINGHAM HIP IMPL | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/METAL, RESURFACING |  | 56 YR | Female | IT WAS REPORTED THAT AFTER UNDERGOING A BHR ON (B)(6) 2009, PATIENT EXPERIENCED HIGH LEVELS OF COCR IN BLOOD TESTS. THE ADVERSE EVENT WAS SUBSEQUENTLY ADDRESSED THROUGH A REVISION SURGERY, PERFORMED ON AN UNSPECIFIED DATE. PATIENT'S CURRENT HEALTH STATUS IS UNKNOWN. |  | 2025 | 8 | 2025-08

## Monthly Summary
Year-Month | Report Count | Top Event Types
2020-01 | 4 | I, n, j, u, r, y
2020-03 | 2 | I, n, j, u, r, y
2020-05 | 18 | I, n, j, u, r, y
2020-06 | 11 | I, n, j, u, r, y
2020-07 | 11 | I, n, j, u, r, y
2020-08 | 23 | I, n, j, u, r, y
2020-09 | 9 | I, n, j, u, r, y
2020-10 | 15 | I, n, j, u, r, y
2020-11 | 18 | I, n, j, u, r, y
2020-12 | 27 | I, n, j, u, r, y
2021-01 | 32 | I, n, j, u, r, y
2021-02 | 19 | I, n, j, u, r, y
2021-03 | 10 | I, n, j, u, r, y
2021-04 | 7 | I, n, j, u, r, y
2021-05 | 49 | I, n, j, u, r, y
2021-06 | 5 | I, n, j, u, r, y
2021-07 | 8 | I, n, j, u, r, y
2021-08 | 9 | I, n, j, u, r, y
2021-09 | 9 | I, n, j, u, r, y, D, e, a, t, h
2021-11 | 21 | I, n, j, u, r, y
2021-12 | 9 | I, n, j, u, r, y
2022-01 | 4 | D, e, a, t, h, I, n, j, u, r, y
2022-02 | 5 | I, n, j, u, r, y
2022-03 | 9 | I, n, j, u, r, y
2022-04 | 24 | I, n, j, u, r, y
2022-05 | 83 | I, n, j, u, r, y
2022-06 | 18 | I, n, j, u, r, y
2022-07 | 10 | I, n, j, u, r, y
2022-08 | 7 | I, n, j, u, r, y
2022-09 | 1 | I, n, j, u, r, y
2022-10 | 3 | I, n, j, u, r, y
2022-12 | 1 | I, n, j, u, r, y
2023-01 | 2 | I, n, j, u, r, y
2023-03 | 2 | I, n, j, u, r, y
2023-04 | 5 | I, n, j, u, r, y
2023-05 | 1 | I, n, j, u, r, y
2023-08 | 2 | I, n, j, u, r, y
2023-09 | 3 | I, n, j, u, r, y
2023-11 | 2 | I, n, j, u, r, y
2024-02 | 1 | I, n, j, u, r, y
2024-03 | 1 | I, n, j, u, r, y
2024-05 | 4 | I, n, j, u, r, y
2024-06 | 2 | I, n, j, u, r, y
2024-07 | 15 | I, n, j, u, r, y
2024-08 | 9 | I, n, j, u, r, y
2024-09 | 1 | I, n, j, u, r, y
2024-10 | 1 | I, n, j, u, r, y
2024-11 | 3 | I, n, j, u, r, y, M, a, l, f, u, n, c, t, i, o, n
2024-12 | 6 | I, n, j, u, r, y
2025-01 | 1 | I, n, j, u, r, y
2025-02 | 4 | I, n, j, u, r, y
2025-03 | 9 | I, n, j, u, r, y
2025-04 | 2 | I, n, j, u, r, y
2025-05 | 3 | I, n, j, u, r, y
2025-06 | 2 | I, n, j, u, r, y
2025-07 | 2 | I, n, j, u, r, y
2025-08 | 1 | I, n, j, u, r, y

## Event Types
Event Type | Count
I, n, j, u, r, y | 560
D, e, a, t, h | 4
M, a, l, f, u, n, c, t, i, o, n | 1

## Yearly Summary
Year | Report Count
2020 | 138
2021 | 178
2022 | 165
2023 | 17
2024 | 43
2025 | 24

## Top Problems
Device Problem | Count
565